5065139 | WHEN A NEW QUALITY CONTROL (QC) RESULT CONTAINS AN ERROR CODE ON THE (B)(4), THE USER CAN EXCLUDE THE RESULT AND PROCEED WITH RUNNING QC. HOWEVER, IF A SECOND RESULT WITH AN ERROR CODE OCCURS AND THE USER TRIES TO EXCLUDE THE RESULT, THE ERROR FROM THE FIRST QC RESULT REPLACES THE SECOND ERROR. THE USER MUST THEN MANUALLY EXCLUDE THE ERROR RESULT A SECOND TIME. IF THE AUTO VERIFICATION FEATURE IS BEING USED, ALL RESULTS WILL BE HELD. THERE ARE NO KNOWN REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO THIS ISSUE.><SIEMENS HEALTHCARE DIAGNOSTICS HAS IDENTIFIED FIVE FEATURE ISSUES WITH THE (B)(4). THE ISSUES PERTAIN TO RESULT UNIT CONVERSIONS, QUALITY CONTROL PROCESSING, VIRUS PROTECTION, SYSTEM PERFORMANCE DEGRADATION, AND ORDERS RECEIVED FROM LIS BEING REJECTED. AN URGENT MEDICAL DEVICE CORRECTION (UMDC) ENTITLED (B)(4)" WAS SENT TO CUSTOMERS IN THE UNITED STATES IN AUGUST 2015. AN URGENT FIELD SAFETY NOTICE (UFSN) ENTITLED "(B)(4)" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN AUGUST 2015. THE UMDC AND UFSN DESCRIBE THE ISSUES AND PROVIDE ACTIONS THE CUSTOMER CAN TAKE TO PREVENT AND MITIGATE THE ISSUES." | 1
5164476 | COVIDIEN REFERENCE NUMBER:(B)(4). VERIFIED THE CUSTOMER COMPLAINT THAT THERE IS NO AUDIO ALARM DURING THE CUSTOMER DESCRIBED SITUATION OF THE WIRES EXPERIENCE STRESS" IT WAS REPORTED THAT THE UNIT HAS ALARMS THAT WORK PROPERLY. DURING TIMES WHEN THE SENSOR IS CONNECTED AND A CONSISTENT READING CAN NOT BE ASCERTAINED | 1
4612249 | ><DURING A LEFT HEART CATHETERIZATION PROCEDURE THE MCKESSON HEMODYNAMIC SYSTEM FROZE UP AND WE WERE UNABLE TO RECORD PRESSURES DURING THE END OF THE PROCEDURE. THE LEFT VENTRICULAR (LV) PRESSURES COULD BE VISUALLY SEEN BY THE DOCTOR BUT THAT PRESSURE WITH THE PULL BACK COULD NOT BE RECORDED. NO HARM TO PATIENT. THE PRESSURES WERE WITHIN NORMAL LIMITS. IS CAME AND REPLACED A CABLE AND WE PROCEEDED WITH THE NEXT CASE. | 1
4881673 | A NURSE (RN) REPORTED TO ENGINEERING A BURNING ELECTRICAL SMELL PERMEATING THROUGHOUT THE WING. ENGINEERING AND SECURITY RESPONDED, SEARCHED AND LOCATED A WORKSTATION ON WHEELS (WOW) WITH A FLICKERING CHARGER DISPLAY AND THE STRONG SMELL OF SMOKE NEXT TO AN UNOCCUPIED PATIENT BED. THE UNIT WAS UNPLUGGED AND TRANSPORTED TO THE GROUND FLOOR UTILITY AREA FOR STORAGE UNTIL REPAIRED. SEVERAL MINUTES LATER THE UTILITY AREA SMOKE ALARM ACTIVATED AND THE TEAM RESPONDED TO FIND THE LITHIUM-ION BATTERY EMITTING SMOKE. THE WOW WAS TRANSPORTED TO THE BACK LOADING DOCK AREA AND HELD TO DETERMINE SOURCE OF SMOKE AND EVALUATION. THE WOW'S BATTERIES THEN BEGAN TO SPARK AND EMIT LARGE FLAMES SEVERAL TIMES. THE TEAM THEN SPRAYED THE WOW WITH A DRY SODA BASED EXTINGUISHER AND PLACED THE DEVICE INTO AN EMPTY METALLIC BIN WHERE IT CONTINUED TO FLAME UP AND MELT THE BATTERY AND WORKSTATION. CITY FIRE WAS DISPATCHED AND BIOMEDICAL WAS CONTACTED TO DETERMINE NEXT STEPS APPROXIMATELY AN HOUR AND A HALF LATER. CITY FIRE CONTAINED EQUIPMENT UNTIL THERE WAS NO FURTHER BURNING OF THE DEVICE AND HANDED IT OFF TO HOSPITAL BIOMEDICAL AND IT COLLABORATED TO CONTACT MANUFACTURER AND SUBMIT ADVERSE EVENT REPORT. NO PATIENT HARM OR CARE AREA EVACUATION OCCURRED DUE TO THIS EVENT. THE MANUFACTURER IS DETERMINING FURTHER PLAN OF ACTION ON RETRIEVING DEVICES FOR ANALYSIS AND REPORT. | 1
4833688 | INFORMATION REGARDING THE NUMBER OF CASES FOR WHICH THE TISSUE SAMPLES WERE DIFFICULT TO MICROTOME AND SOME SLIDES WERE NOT SUFFICIENTLY READABLE TO MAKE A DIAGNOSIS"; AND CONFIRMATION AS TO WHETHER RE-BIOPSY OF ANY PATIENT(S) IS REQUIRED AND/OR HAS BEEN PERFORMED HAS NOT BEEN RECEIVED BY LEICA BIOSYSTEMS AS AT (B)(6) 2015.><BOTTLE 12 (IPA) WAS REMOVED FROM THE INSTRUMENT FOR 4 SECONDS AT 12:20PM ON (B)(6) 2015 IN RESPONSE TO INFO DISPLAYED IN ASSOCIATION WITH AN ERROR CODE INDICATING THAT A PROTOCOL COULD NOT BE SCHEDULED | 0
4810638 | ON (B)(6) 2015, A LEICA FIELD SERVICE SPECIALIST (FSS) MEASURED THE ETHANOL CONCENTRATION IN BOTTLE 5 USING A HYDROMETER. THE MEASURED ETHANOL CONCENTRATION IN BOTTLE 5 WAS 84.5%; AND THE ETHANOL CONCENTRATION CALCULATED BY THE INSTRUMENT SOFTWARE WAS 99.7%. PRELIMINARY INVESTIGATION BY LEICA BIOSYSTEMS FOUND THAT THE REAGENT IN BOTTLES 5 AND 14 WAS REPLACED PRIOR TO THE PROCESSING PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. BOTTLE 5 WAS USED FOR THE FIRST TIME IN THE FINAL DEHYDRATION STEP OF BOTH THE PROCESSING PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. THE DISCREPANCY BETWEEN THE MEASURED AND CALCULATED ETHANOL CONCENTRATION IN BOTTLE 5 INDICATES THAT A USE ERROR OCCURRED WHEN REPLACING THIS REAGENT. AS A CONSEQUENCE, THIS COMPLAINT REQUIRED REPORTING TO THE FDA AS A PRODUCT PROBLEM MALFUNCTION BASED ON THE TWO (2) YEAR PRESUMPTION RULE FOR THE MOST RECENT COMPLAINT WITH THE SAME ROOT CAUSE. INVESTIGATION OF THIS COMPLAINT BY LEICA BIOSYSTEMS REMAINS IN PROGRESS.><LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL PROCESSING OF TISSUE SAMPLES FROM TWO (2) PROCESSING RUNS USING THE 2HR GI BIOPSY PROTOCOL, WHICH WERE STARTED ON (B)(6) 2015. A LEICA FIELD SUPPORT SPECIALIST REPORTED THAT: THEY DESCRIBED THE BLOCKS AS POORLY INFILTRATED PER TAC. BOTH POORLY PROCESSED RUNS WERE REPROCESSED ON THE LAB'S VIP PROCESSOR AND APPEARED TO BE IN GOOD CONDITION. TISSUE STATUS IS STILL PENDING PATHOLOGIST'S REVIEW. I WILL BE VISITING THE SITE ON (B)(6) TO INVESTIGATE THE SOURCE OF THE POOR PROCESSING. THE CUSTOMER NOTED THAT THEY DID CHANGE A FEW PARAFFIN STATIONS IN A VERY UNCONVENTIONAL MANNER. THE SITE INDICATED THAT THEY DID NOT HAVE A PARAFFIN DRAIN HOSE AVAILABLE AT THE TIME NECESSARY TO REMOTE DRAIN THE PARAFFIN BATH. THEY STATED TO ME THAT THEY "SCOOPED" PARAFFIN MANUALLY OUT OF THE APPROPRIATE BATH AND REFILLED THE BATH WITH NEW PARAFFIN." ON 13 MAY 2015 | 0
4698928 | A LEICA FIELD SERVICE ENGINEER (FSE) ATTENDED THE LABORATORY ON (B)(4) 2015 IN ORDER TO INVESTIGATE THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. THE FSE FOUND THAT THE LIQUID LEVEL SENSORS WERE VERY DIRTY; AND REQUIRED CONSIDERABLE EFFORT TO CLEAN. THE FSE REPORTED THAT THE LIQUID LEVEL SENSORS FUNCTIONED TO SPECIFICATION AFTER CLEANING AND THAT NO PATIENT SAMPLE (S) WAS ADVERSELY AFFECT. EVALUATION OF THE INSTRUMENT LOGS BY GLOBAL TECHNICAL SUPPORT SHOWED THAT MULTIPLE EVENT/ERROR CODES ASSOCIATED WITH THE FUNCTION OF THE LLSS HAD BEEN RECORDED IN INSTRUMENT LOGS FOR BOTH RETORT A AND B, BUT PREDOMINANTLY FOR RETORT A.><LEICA BIOSYSTEMS RECEIVED A COMPLAINT THAT A LIQUID LEVEL SENSOR(S) WAS NOT FUNCTIONING CORRECTLY. THE COMPLAINANT DID NOT INDICATE THAT THE QUALITY OF TISSUE PROCESSING FOR ANY PATIENT SAMPLE (S) HAD BEEN ADVERSELY AFFECTED BY THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. | 0
5160843 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL PROCESSING OF TISSUE IN APPROXIMATELY 30 BLOCKS FROM TWO (2) PROCESSING RUNS, WHICH HAD BEEN PROCESSED USING EITHER RETORT A OR B. THE COMPLAINANT ADVISED THAT THE AFFECTED PROCESSING RUNS COMPLETED SUCCESSFULLY; AND THAT NO ERROR CODES HAD BEEN DISPLAYED. THE COMPLAINANT DESCRIBED THE AFFECTED TISSUE SAMPLES AS UNDER-PROCESSED". ON (B)(6) 2015 | 0
5055434 | BOTTLE 3 (ETHANOL) WAS REMOVED FROM THE INSTRUMENT ON TWO (2) OCCASION FOR 13 MINUTES AT 10:14AM ON (B)(6) 2015 AND 30 SECONDS AT 10:27AM ON (B)(6) 2015; AND THE PROPERTIES OF THE CORRESPONDING REAGENT STATION WERE RESET AT 10:27AM ON (B)(6) 2015. RESETTING THE REAGENT STATION SET THE CONCENTRATION TO THE DEFAULT VALUE CONFIGURED IN THE REAGENT TYPES DEFINITIONS (100% IN THIS INSTANCE); AND RESET THE NUMBER OF CASSETTES PROCESSED AND THE NUMBER OF CYCLES AND DAYS TO ZERO. HOWEVER, THE VARIANCE BETWEEN THE ETHANOL CONCENTRATION OF 77% WAS MEASURED BY THE LABORATORY USING A HYDROMETER AND THAT CALCULATED BY THE INSTRUMENT SOFTWARE OF 99.9%, INDICATES THAT AN ERROR OCCURRED DURING COMPLETION OF THIS ACTION. BOTTLE 3 (ETHANOL) WAS USED IN THE FINAL DEHYDRATION STEP OF THE BREAST AND FATTY" PROTOCOL STARTED IN RETORT B AT 10:47AM ON (B)(6) 2015 | 0
4995935 | ON 04 SEPTEMBER 2015, THE LEICA HISTOLOGY TECHNICAL SPECIALIST ADVISED THAT USER TRAINING HAS BEEN OFFERED TO THE LABORATORY. THE LABORATORY IS IN THE PROCESS OF PURCHASING ANOTHER PELORIS II TISSUE PROCESSOR AND USER TRAINING WILL BE PROVIDED UPON INSTALLATION OF THE INSTRUMENT. AS AT 04 SEPTEMBER 2015, LEICA BIOSYSTEMS HAS NOT RECEIVED THE PATIENT IDENTIFIER; AGE/DATE OF BIRTH AND GENDER OF THE PATIENT REQUIRING RE-BIOPSY, DESPITE SEVERAL REQUEST BEING MADE TO THE LABORATORY.><THE STATION PROPERTIES FOR BOTTLE 10 (ETHANOL) WERE RESET AT 07:57AM ON (B)(6) 2015. RESETTING THE REAGENT STATION SET THE CONCENTRATION TO THE DEFAULT VALUE CONFIGURED IN THE REAGENT TYPES DEFINITIONS (100% IN THIS INSTANCE); AND RESET THE NUMBER OF CASSETTES PROCESSED AND THE NUMBER OF CYCLES AND DAYS TO ZERO. INFORMATION PROVIDED BY THE LABORATORY INDICATES THAT PRIOR TO COMMENCEMENT OF THE PROTOCOL FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS IDENTIFIED, A USER HAD REPLACED THE REAGENT WITH 70% ETHANOL AND SET THE CONCENTRATION TO 100%. BASED ON THIS INFORMATION, IT IS LIKELY THAT BOTTLE 10 WAS REPLACED WITH 70% ETHANOL ON (B)(6) 2015, AND RESET TO 100% AT 07:57AM ON (B)(6) 2015. THE MINIMUM FINAL REAGENT CONCENTRATION REQUIRED FOR ETHANOL IS 98%. THE CONSEQUENCES OF USING ETHANOL AT A CONCENTRATION LESS THAN THE MINIMUM REQUIRED FOR THE FINAL DEHYDRATION STEP IN A PROTOCOL IS RE-INTRODUCTION OF WATER INTO THE TISSUE WHICH CANNOT BE DISPLACED IN SUBSEQUENT PROCESSING STEPS; AND CONTAMINATION OF REAGENTS USED IN SUBSEQUENT PROCESSING STEPS ULTIMATELY RESULTING IN SUB-OPTIMAL PROCESSING. INVESTIGATION OF THIS COMPLAINT FOUND THAT THE INSTRUMENT OPERATED WITHIN SPECIFICATION. THE ROOT CAUSE OF THE SUB-OPTIMAL TISSUE PROCESSING REPORTED WAS A USE ERROR. SPECIFICALLY PER THE INFORMATION PROVIDED BY THE COMPLAINANT, A USER REPLACED THE REAGENT WITH 70% ETHANOL, BUT AFFIRMED IN ERROR THAT THE CONCENTRATION WAS TO BE SET TO 100% IN THE INSTRUMENT SOFTWARE. DETAILS OF THE EVENT WERE ALSO BROADCAST IN THE BROADCAST AT THE FOLLOWING LINK: A TRANSCRIPT OF THE BROADCAST IS AVAILABLE UPON REQUEST.><LEICA BIOSYSTEMS RECEIVED A COMPLAINT THAT SUB-OPTIMAL TISSUE PROCESSING HAD BEEN IDENTIFIED ON (B)(6) 2015 FROM A PROCESSING RUN, WHICH STARTED ON (B)(6) 2015. A LEICA REPRESENTATIVE REPORTED THAT THE FOLLOWING INFO REGARDING THE CIRCUMSTANCES INVOLVED IN THE COMPLAINT HAD BEEN PROVIDED BY GRAND RIVER HOSPITAL: WE WERE INFORMED THAT A STAFF MEMBER PULLED AN ETHANOL BOTTLE" (THEY ONLY USE PURE REAGENT IN ALL OF THEIR STATIONS) AND REPLACED THE REAGENT WITH 70% ALCOHOL AT THE DISCRETION THAT THERE WAS NOT A DEGRADED ALCOHOL ON BOARD AT THE TIME WITH A LOW PURITY. THE STATION DETAILS FOR THIS BOTTLE WERE THEN INCORRECTLY UPDATED TO 100% PURE. (IN FACT THE BOTTLE CONTAINED 70% ALCOHOL). MOST LIKELY | 0
4380025 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUBOPTIMAL PROCESSING OF APPROXIMATELY (B)(4) TISSUE BLOCKS COMPRISING BIOPSY SAMPLES FROM A PROTOCOL RUN. THE COMPLAINANT DESCRIBED THE TISSUE SAMPLES AS BURNT". ON (B)(6) 2014 | 0
5149133 | ON 17 SEPTEMBER 2015, LEICA BIOSYSTEMS WAS ADVISED THAT TISSUE BADLY PROCESSED". ON-SITE ASSESSMENT OF THE OPERATION AND FUNCTION OF THE INSTRUMENT WAS CONDUCTED BY A LEICA FIELD SERVICE ENGINEER (FSE) ON 17 SEPTEMBER 2015. THE FSE MEASURED THE ETHANOL CONCENTRATION IN THE BOTTLES DESIGNATED FOR THIS REAGENT USING A HYDROMETER; AND PERFORMED A TEST FOR WATER ON REAGENT BOTTLES DESIGNATED AS XYLENE. THE FSE FOUND THAT THE VARIANCE BETWEEN THE MEASURED ETHANOL CONCENTRATION AND THAT CALCULATED BY INSTRUMENT SOFTWARE WAS EXCEEDED THE ACCEPTABLE LIMIT FOR BOTTLE 9. THE ETHANOL CONCENTRATION IN BOTTLE 9 WAS 50% AND THE CALCULATED CONCENTRATION BY THE INSTRUMENT SOFTWARE WAS 100%. THE FSE PERFORMED SYSTEM VERIFICATION TESTS | 0
4751410 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING INTERMITTENT PROBLEMS WITH THE INSTRUMENT WITH RUNNING PROCESSING PROTOCOLS OVERNIGHT. ON 03/26/2015, LEICA BIOSYSTEMS RECEIVED INFO THAT NO TISSUE SAMPLE(S) WAS ADVERSELY AFFECTED BY THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. | 0
4420883 | ><AFTER WEARING THE EARLOOP MASK AS ASSISTANT, STARTED TO EXPERIENCE A COUGH, FELT ITCHY AND SUFFERED FROM AN ASTHMA ATTACK. THE ASSISTANT WAS USING A SPRAY TO HELP TREAT HER ASTHMA. THE TYPE OF SPRAY USED WAS NOT DISCLOSED NOR IF ANY MEDICAL ATTENTION WAS REQUIRED. | 0
4843605 | PORTABLE UNIT WAS USED TO OBTAIN PATIENT IMAGE. THE IMAGE COULD NOT BE TRANSFERRED TO IMAGE STORAGE SOFTWARE DUE TO A GLITCH" IN THE SOFTWARE. THIS CAUSED THE TECHNICIAN TO TAKE A SECOND IMAGE. A SECOND SIMILAR EVENT OCCURRED SUBSEQUENTLY ON A SUBSEQUENT DATE AND WAS CORRECTED BY THE SOFTWARE VENDOR WITH A VERSION UPGRADE. APPEARED PROBLEM RELATED TO ANNOTATIONS MADE ON IMAGE BY THE TECHNICIAN CAUSED FILE CORRUPTION." | 1
4699887 | IN THE NEW PYXIS ES SYSTEM, A DOSE OF ROCEPHIN WAS ORDERED BY A PROVIDER AND DID NOT SHOW UP ON THE PATIENT'S PROFILE. THE MEDICATION WAS DELAYED AND THE PATIENT DECIDED TO LEAVE BEFORE IT WAS ADMINISTERED. IT WAS LATER DETERMINED THAT THE MEDICATION WAS ENTERED BY A PHYSICIAN AND AUTOVERIFIED, HOWEVER THE NURSE WAS NOT AWARE THAT THE MEDICATION WITH TWO COMPONENTS WOULD NOT SHOW ON THE PROFILE FOR REMOVAL. (B)(6). | 1
4996470 | A FIELD SAFETY CORRECTIVE ACTION HAS BEEN INITIATED BY THE MANUFACTURER. LEICA BIOSYSTEMS IS ISSUING THE FIELD SAFETY NOTICE(FSN) TO INFORM THE CUSTOMER ABOUT A FIELD SAFETY CORRECTIVE ACTION (FSCA) THAT WE ARE INITIATING, TO ADDRESS AN ISSUE WE HAVE EXPERIENCED WITH OUR TISSUE PROCESSOR, (B)(4).><THE AFFECTED DEVICE HAS AN INCORRECT INTERNAL WIRING TO THE REMOTE ALARM PLUG, WHICH CAN CAUSE THE REMOTE ALARM TO NOT FUNCTION CORRECTLY IN THE CASE OF AN INSTRUMENT FAILURE. IN CASE A CUSTOMER HAVE AN INSTRUMENT FAILURE DURING PROCESSING AND THE CUSTOMER CURRENTLY RELY ON THE REMOTE ALARM, THE MISSING ALARM SIGNAL COULD LEAD TO SUBOPTIMAL QUALITY PROCESSING. IN SOME CASES, A RE-BIOPSY OF THE PATIENT COULD BE NECESSARY IN ORDER TO OBTAIN ANOTHER VIABLE TISSUE SAMPLE. | 0
4699888 | AFTER THE NEW PYXIS ES SYSTEM WAS INSTALLED, A DILAUDID ORDER WAS ENTERED IN THE ED AND THE ORDER WAS SENT TO THE PHARMACY FOR VERIFICATION. THE VERIFICATION WAS DELAYED AND THE MEDICATION WAS NOT GIVEN PROMPTLY TO THE PATIENT. THE ED WAS USING A PROFILED SYSTEM AND THE MEDICATION WAS NOT OVERRIDEABLE. (B)(4). | 1
4755149 | FENESTRATED BIPOLAR WAS BEING USED DURING DAVINCI NISSEN WHEN IT STOPPED WORKING PROPERLY. THE JAWS WERE NOT CLOSING TOGETHER IN ALIGNMENT. INSTRUMENT REMOVED FROM FIELD. NO HARM TO PATIENT. | 0
5004603 | FOLLOW-UP: AFTER FURTHER INVESTIGATION, DAKO HAS DISCOVERED THAT THE PROBLEM COMPLAINT DESCRIBED IN THIS MDR, 2022180-2015-0002, HAD NOTHING TO DO WITH THE WT1 ANTIBODY. THE ISSUE WAS DISCOVERED TO BE RELATED TO THE ENVISION FLEX+ PRODUCT PROBLEM THAT WAS OUTLINED IN MDR 9610099-2015-2015. THIS INVESTIGATION IS CLOSED.><ON (B)(6) 2015, DAKO RECEIVED A COMPLAINT REGARDING THE ROUTINE USE OF THE FLEX MONOCLONAL MOUSE ANTI-HUMAN WILMS' TUMOR 1 (WT1) PROTEIN CLONE 6F-H2 ON THE AUTOSTAINER LINK 48 FOR DIAGNOSTIC USE. INITIALLY THE CUSTOMER REPORTED FALSE POSITIVE STAINING, CLAIMING THAT THIS STAINING CONTRIBUTED TO A DIAGNOSTIC ERROR LEADING TO DISTRESS, UNNECESSARY SURGERY AND OTHER INCORRECT TREATMENT. ON AUGUST 12, 2015, DAKO RECEIVED FOLLOW UP INFORMATION REGARDING THIS INCIDENT. THE CUSTOMER REPORTED THAT THEY HAVE REVALIDATED THE SAME BATCH OF WT1 USED IN THE ORIGINAL TESTING AND HAVE CONFIRMED EXPECTED STAINING RESULTS. THE CUSTOMER REPORTED THAT A RETEST WAS ALSO CARRIED OUT AT ANOTHER LABORATORY, ON SECTIONS FROM THE SAME BIOPSY, AND RESULTS WERE ALSO POSITIVE. THE CUSTOMER REPORTED THAT A SEPARATE TEST CAN ON SURGICAL TEST RESECTION TISSUE DONE AT ANOTHER HOSPITAL WAS NEGATIVE. THE CUSTOMER REPORTED THAT AS A RESULT OF THESE CONCLUSIONS, THE CUSTOMER HAS INFORMED DAKO THAT THEY ARE NO LONGER PURSUING A COMPLAINT ABOUT DAKO REGARDING FALSE POSITIVE STAINING WITH THE WT1 ANTIBODY. DAKO IS IN THE PROCESS OF OBTAINING DETAILS OF THIS COMPLAINT FOR FURTHER INVESTIGATION.><PLEASE SEE ADDITIONAL MANUFACTURE NARRATIVE. | 0
5297723 | THE PROCEDURE WAS AN ELECTIVE TOTAL LAPAROSCOPIC HYSTERECTOMY, BILATERAL SALPINGO-OOPHORECTOMY WITH DAVINCI ASSISTANCE, AND POSSIBLE LAPAROTOMY. FOLLOWING SPECIMEN REMOVAL TRANSVAGINALLY, THE DAVINCI ROBOTIC SYSTEM SUFFERED A SERIES OF FAULTS. TROUBLESHOOTING OF FAULTS WAS OVERRIDDEN SEVERAL TIMES RESULTING IN A SYSTEM OVERRIDE. INTUITIVE SURGICAL WAS CALLED BY THE CIRCULATING NURSE. INSTRUCTIONS WERE PROVIDED TO SHUT THE SYSTEM DOWN AND REBOOT WHICH WAS PERFORMED TWICE AND THE FAULT WAS UNRECOVERABLE. AT THIS POINT, THE PHYSICIAN COULD NOT PROCEED ANY FURTHER. THE DEVICE WAS UNDOCKED AND PHYSICIAN DECIDED TO CLOSE THE VAGINA TRANSVAGINALLY. | 1
5167784 | ON (B)(6) 2015, AS AN INPATIENT ON (B)(6), I REPORTED AN APPROXIMATE 40 POINT DISCREPANCY BETWEEN A MECHANICAL AND A DIGITAL BLOOD PRESSURE MACHINE TO THE HOSPITAL'S PATIENT REPRESENTATIVE. I NOTICED THIS DISCREPANCY ON (B)(6) 2015 AT 4 AM WHEN AN OVERNIGHT NURSE CHOSE FOR THE PURPOSES OF ADMINISTERING MEDICATION, TO RELY ON A LOWER READING OF MY BLOOD PRESSURE (160+/96) FROM A MECHANICAL MACHINE AFTER TAKING TWO CONSECUTIVE B.P. READINGS (R. ARM AND L. ARM) WITH A DIGITAL B.P. MACHINE THAT INDICATED MY SYSTOLIC WAS ABOVE 200 AND MY DIASTOLIC WAS OVER 100. THIS DIGITAL B.P. MACHINE WAS USED AGAIN AT 8:00 AM TO READ THE BLOOD PRESSURE OF MANY OTHER INPATIENTS ON UNIT 48. MY HIGH BLOOD PRESSURE MEDICATION WAS LATER DOUBLED DUE TO HIGH READINGS ON THE DIGITAL B.P. MACHINE. I WAS NEVER INFORMED OF THE OUTCOME OF MY REPORT AND IN A CONVERSATION WITH THE PATIENT CARE MANAGER HE STATED THAT THE MACHINE WAS ROUTINELY CALIBRATED AND IN PERFECT WORKING ORDER. | 0
4699897 | A NEW PYXIS MACHINE CALLED PYXIS ES WAS INSTALLED IN OUR FACILITY. THE MACHINE SHOULD INDICATE THE PROPER AMOUNT OF MEDICATION TO REMOVE SO THAT THE NURSE HAS THE RIGHT DOSAGE. AN ORDER FOR ATIVAN AND ROBITUSSIN AC WERE SENT TO THE SCREEN BUT THE MACHINE INDICATED THE WRONG AMOUNT FOR REMOVAL. THE NURSE CAUGHT THE ERROR IN TIME AND THE CORRECT AMOUNT WAS REMOVED FOR THE PATIENT. THE PROBLEM WAS SENT TO THE IT TEAM FOR TROUBLESHOOTING BY THE PHARMACIST WORKING THAT DAY. A PROBLEM IN THE TRANSMISSION OF THE ORDER WAS IDENTIFIED. (B)(6). | 1
4735655 | POSSIBLE FAILURE OF DAVINCI INTUITIVE STAPLES TO FIRE AND CLOSE BOWEL RESULTING IN LEAKAGE OF BOWEL CONTENTS INTO PT'S ABDOMEN. THE SURGEON IRRIGATED THE ABDOMEN AT TIME OF THE PROCEDURE WITH 2L OF NACL AND THEN PLACED THE DRAIN IN. THE SURGEON DISCLOSED THE STAPLE LINE ISSUE WITH FAMILY IMMEDIATELY POST OPERATIVELY, AS NOTED IN HIS OPERATIVE REPORT. THE MANUFACTURER OF THE STAPLES WAS NOTIFIED. IT IS UNKNOWN WHETHER THE ISSUE IS WITH THE STAPLER OR DAVINCI DISPOSABLE BLUE STAPLER RELOAD (MODEL #: 41645B LOT# DS1135118 EXP 12/2015) OR DAVINCI DISPOSABLE GREEN STAPLER RELOAD (MODEL #: 41445G LOT # DS1134430 EXP 11/2015). | 0
5067746 | THERE WAS INTERVAL INCREASE OF POTASSIUM LEVEL IN A CASE ON BOTH POTASSIUM SUPPLEMENTS AND A POTASSIUM SPARING DIURETIC FOR END STAGE HEART FAILURE. A POTASSIUM LEVEL WAS 5.7 YET BOTH THE POTASSIUM AND THE DIURETIC WERE ADMINISTERED BECAUSE THE HIGH POTASSIUM LEVEL WAS NOT SEEN HAVING BEEN ELECTRONICALLY SENT TO THE MDDS WITHOUT NOTICE OR WARNING. ADDITIONAL THERAPY WAS NEEDED TO CORRECT THE RISK. | 1
5297731 | THE TELEMETRY SERVER REGULARLY AND UNEXPECTEDLY DROPS A PATIENT FROM THE LIST AND REBOOTS, RESULTING IN LOSS OF DATA AND GAPS IN MONITORING OF PATIENTS IN THE CARDIAC REHAB TELEMETRY SERVICES DEPARTMENT. THE VENDOR WAS CONTACTED ABOUT THIS ISSUE AND IT WAS IDENTIFIED THAT THE ISSUE HAS BEEN EXPERIENCED AT OTHER SITES AS WELL. SCOTTCARE IS CONDUCTING TESTING TO FIND OUT THE UNDERLYING ISSUES BUT THUS FAR NO FINAL RESOLUTION HAS BEEN ACHIEVED AND THE PROBLEMS CONTINUE. | 1
5243445 | ON (B)(6) PATIENT HAD RUNS OF VTACH ONE 25 BEAT ONE 22 BEAT RUN, NEITHER VENT RECORDED/ALARMED BY GE MONITOR. FOUND BY NURSE GOING THROUGH PREVIOUS DAYS EVENTS. NO HARM. | 1
4722052 | DURING THE ROBOTIC TOTAL HYSTERECTOMY, THE JAWS ON THE COBRA FAILED TO WORK. | 0
4397155 | PATIENT WAS RECEIVING TREATMENT AND DURING THE PROCEDURE ONE OF THE PLASTIC TUBES DEVELOPED A SMALL CRACK. THE MD WAS NOTIFIED AND RINSE BACK DONE. DURING RINSEBACK ANOTHER MACHINE WAS PRIMED WITH NEW TUBING. ONCE THE RINSE BACK WAS COMPLETED THE PATIENT WAS DISCONNECTED FROM ONE MACHINE AND IMMEDIATELY CONNECTED TO ANOTHER AND THE TREATMENT RESUMED UNTIL COMPLETION. THIS WAS NOT A MALFUNCTION OF THE MACHINE BUT A CRACK IN THE DISPOSABLE TUBING. | 0
4631443 | PT UNDER GENERAL ANESTHESIA FOR A ROBOTICALLY ASSISTED HYSTERECTOMY WHEN THE MACHINE MALFUNCTIONED. THE LAPAROSCOPIC SCISSORS WERE NOT WORKING, AND THE ENTIRE ARM #2 WAS RIGID AND NOT MOVING. THE TROCAR WAS NOT LOCKING AND WAS DISCONNECTING FROM THE ROBOT ARM WITH ATTEMPTED ARM MANIPULATION FROM THE CONSOLE. STAFF IN THE ROOM REPORTED THAT THIS IS A RECURRING ISSUE WITH ARM #2. CASE WAS DELAYED >1 HOUR IN ATTEMPT TO TROUBLE SHOOT AND CORRECT THE MALFUNCTION.======================MANUFACTURER RESPONSE FOR SURGICAL ROBOT, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================NOTHING AT THIS TIME. | 1
5322313 | ><AN INCIDENT OCCURRED ON THE GANTRY MANUFACTURING LINE. THE WORKER NOTICED THAT THE MAIN BEARING WAS MAKING NOISE WHEN IT WAS ROTATED. UPON SOME MANIPULATION OF THE BEARING, THE BEARING'S THREADED JOINT SEPARATED AND THE ROTATING SIDE FELL TO THE FLOOR. NO ONE WAS INJURED. THE BEARING SYSTEM IS USED ON THE LIGHTSPEED FAMILY OF CTS. INVESTIGATION IDENTIFIED THAT THE ROOT CAUSE WAS THE IMPROPER MOUNTING OF THE THREAD CUTTING TOOL AT THE BEARING MANUFACTURER RESULTING IN THE PITCH DIAMETER BEING OUTSIDE OF SPECIFICATIONS. A REVIEW OF BEARING MANUFACTURER RECORDS INDICATED THAT TWO OTHER BEARINGS HAVE BEEN DELIVERED TO GE WITH OUT OF SPECIFICATION THREADS. ONE WAS STILL IN MANUFACTURING AND THUS REMOVED AND THE OTHER WAS IN A LIGHTSPEED PLUS SYSTEM IN THE FIELD. GE CONTACTED THE CUSTOMER WITH THE SUSPECT BEARING AND HAS REPLACED THEIR GANTRY. THE FIELD UNIT DID NOT HAVE ANY REPORTED PROBLEMS AND DID NOT HAVE A FAILURE. | 0
4580187 | DEVICE WAS RETURNED TO THE COMPANY FOR FURTHER ANALYSIS.><CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR WHICH MAY AFFECT PT MONITORING. NO PT INJURY WAS REPORTED. | 1
4567432 | CUSTOMER REPORTED AN ISSUE WITH THE PM9000 EXPRESS MONITOR WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CO2. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4580201 | CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR WHICH MAY AFFECT PT MONITORING. NO PT INJURY WAS REPORTED.><DEVICE WAS RETURNED TO THE COMPANY FOR FURTHER ANALYSIS. | 1
4495610 | CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR THAT MAY HAVE AFFECTED MONITORING. NO PT INJURY WAS REPORTED.><DEVICE WAS RETURNED TO THE COMPANY FOR FURTHER INVESTIGATION. | 1
4645814 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S SPO2 BOARD. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4567464 | CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR V MONITOR'S DISPLAY, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S LCD DISPLAY. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4437829 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S COMMUNICATION BOARD. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. | 1
4567434 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S SPO2 MODULE AND RESEATING THE CABLES. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT V MONITOR WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
4567425 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S FRONT BEZEL ASSEMBLY. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 7 MONITOR'S DISPLAY, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 1
4437823 | CUSTOMER REPORTED A COMMUNICATION LOSS BETWEEN THE PANORAMA CENTRAL STATION AND AMBULATORY TELEPACK, WHICH MAY HAVE AFFECTED TELEMETRY MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE SYSTEM AND DROPPIT WAS CONFIRMED. ISSUE OCCURS DURING AUTOSYNCH WITH THE FACILITY'S SERVERS. SYSTEM WAS RECONFIGURED AND COMMUNICATIONS WERE REESTABLISHED. | 1
4479718 | CUSTOMER REPORTED A PATIENT HAD A VENTRICULAR TACHYCARDIA EVENT AND V-SERIES MONITOR DID NOT ALARM A VENTRICULAR TACHYCARDIA, INSTEAD THE MONITOR ONLY ALARMED A HIGH HEAR RATE. NO PATIENT INJURY WAS REPORTED.><DOCUMENTATION FROM THE REPORTED EVENT IS CURRENTLY UNDER REVIEW. | 1
4645809 | DEVICE WAS EVALUATED. CORRECTIONS INCLUDED REPLACING CO2 MODULE. UNIT CALIBRATED, PERFORMANCE AND SAFETY TESTED TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 5 MONITOR WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
4619102 | DEVICE WAS EVALUATED. CORRECTIONS INCLUDED REPLACING ISOLATED HOST COMM BOARD. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
4567525 | CUSTOMER WAS PROVIDED WITH A LOANER, WHILE THE UNIT'S MPM MODULE IS BEING REPAIRED.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR'S MPM MODULE, WHICH MAY HAVE AFFECTED ECG MONITORING. NO PT INJURY WAS REPORTED. | 1
4437830 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6/7 MONITOR, WHICH MAY HAVE AFFECTED O2 MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT AND VERIFIED THE PROBLEM. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S O2 SENSOR. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS. | 1
4580188 | DEVICE WAS RETURNED TO THE COMPANY FOR FURTHER ANALYSIS.><CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR WHICH MAY AFFECT PT MONITORING. NO PT INJURY WAS REPORTED. | 1
4479764 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED ECG, RESPIRATION, AND TEMPERATURE MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S PATIENT BOARD, CONNECTOR CABLE, U7 ON THE CPU BOARD AND GROUND CABLE. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4459451 | DURING ROBOTIC PROCEDURE THE DAVINCI COBRA GRASPER FAILED TO OPERATE. UPON EXAMINATION, FOUND THAT A CABLE WAS BROKEN AT THE TIP OF THE INSTUMENT. SIX LIVES WERE REMAINING ON INSTRUMENT. | 0
4644597 | THE CUSTOMER HAS CONFIGURED HANGING PROTOCOL (HP) CONFIGURATION OF FIT TO WINDOW" AS THE ZOOM FACTOR. THE CUSTOMER STATED THAT THERE WAS AN INTERMITTENT ISSUE IN THAT SOME SERIES WERE ZOOMED AND SOME WERE NOT WHEN AXIAL IMAGES ARE LINKED | 1
4398113 | INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. (B)(4).><A CUSTOMER REPORTED THE SYSTEM PRODUCED INCORRECT BIOMETRY RESULTS. THERE WAS NO PT HARM. | 1
4928792 | AN ORDER WAS ENTERED VIA CPOE FOR A BLOOD TEST TO MEASURE ELECTROLYTES THE NEXT MORNING. TWO DAYS LATER, THERE WAS AN ELEVATED POTASSIUM LEVEL APPEARING IN THE RESULTS SECTION; BUT THIS WAS NOT SEEN IN A TIMELY MANNER AND CAME AS A SHOCK TO THE ATTENDING PHYSICIAN WHO HAD MADE ARRANGEMENTS TO DISCHARGE THE PATIENT FROM THE HOSPITAL. THERE IS A TWOFOLD PROBLEM HERE. THE FIRST IS THAT THE ORDER WAS NOT CARRIED OUT BY THE ANCILLARY DEPARTMENT IN A TIMELY MANNER, AND WHEN THE TEST WAS FINALLY DONE, THE RESULTS CAME BACK SILENTLY TO THE MDDS EHR, NO ONE KNEW THEY WERE THERE AND GROSSLY ABNORMAL, WITH POTENTIALLY LIFE THREATENING IMPLICATIONS. THE DESIGN OF THE EHR AND ITS ORDERING DEVICE ARE FLAWED. THERE IS NO A MECHANISM FOR RECONCILING THE EXECUTION OF ORDERS THAT PROVIDES A LIST OF THOSE THAT WERE NOT DONE; AND THERE IS NOT ANY INDICATION THAT NEW RESULTS HAVE COME IN TO THE MDDS, BOTH OF THESE CAN CAUSE DEATH AND SHOULD BE REMEDIED IMMEDIATELY. | 1
4438777 | ><INTRA-PROCEDURAL FAILURE OF HEMODYNAMIC ,ECG AND VITAL SIGNS MONITORING CAPABILITY. WHILE PERFORMING A INTRA-VASCULAR PROCEDURE THE MCKESSON HEMODYNAMICS MONITORING SYSTEM SHUT DOWN AND FROZE UPON RETURN TO FUNCTIONING. THIS IS AN ONGOING ISSUE WITH MULTIPLE DIFFERENT INCIDENTS OF FAILURE. MANUFACTURE HAS BEEN WORKING TO RESOLVE WITHOUT SUCCESSFUL RESOLUTION. | 1
4928763 | THIS IS A COMMON AND WIDESPREAD PROBLEM AFFECTING MANY CASES. WHEN THE PATIENT'S EHR IS OPEN AND BEING SEARCHED FOR RADIOLOGY RESULTS, THERE IS A CAMERA ICON ON WHICH TO CLICK WHICH OPENS THE RADIOGRAPH IMAGES TO BE SEEN. THE PATIENT'S CHEST RADIOGRAPH WAS THE SUBJECT OF THIS QUERY. THE ICON WAS CLICKED AND A NORMAL APPEARING CHEST RADIOGRAPH APPEARED AND OTHER RESULTS WERE THEN ACCESSED. THE PROBLEM IS THAT THE RADIOGRAPH THAT APPEARED AND WAS NORMAL, WAS NOT OF THIS PATIENT, BUT OF SOME OTHER PATIENT THAT WAS NOT EVEN ON MY PATIENT LIST. THE PATIENT IN QUESTION HAD A RADIOGRAPH THAT INDICATED PULMONARY EDEMA AND WAS SOUGHT AFTER THE PHYSICAL LUNG EXAM WAS FOUND TO BE INCONSISTENT WITH A NORMAL CHEST RADIOGRAPH. THE DEFECTIVE PRESENTATION OF ANOTHER PATIENT'S RADIOGRAPH IS COMMON WITH THIS INTERFACE OF DEVICES BUT IF THE IMAGE WAS ALLOWED TO STAY ON THE SCREEN, THE CORRECT IMAGE OF THE RIGHT PATIENT WOULD LOAD. THIS DEFECT CAUSES MISDIAGNOSIS AND WRONG TREATMENTS. | 1
4884268 | PT WAS ADMITTED WITH MULTIPLE ALLERGIES DOCUMENTED INCLUDING DAIRY AND MILK. THEY WERE ON A HOME MEDICATION OF A DAIRY-FREE PROBIOTIC POWDER. PER OUR POLICY WE DO NOT CONTINUE HOME POWDERS/LIQUIDS THAT WE CANNOT IDENTIFY (COMPARED TO A TABLET WITH MARKINGS FOR IDENTIFICATION). THEY INITIATED OUR FORMULARY PROBIOTIC (FLORANEX). AFTER THE FIRST DOSE, THE PT EXPERIENCED SYMPTOMS THAT COULD BE RELATED TO HER ALLERGY BUT IT WAS NOT DISCOVERED UNTIL THE SECOND DOSE THAT THE FLORANEX CHEWABLE CONTAINS MILK. THERE WAS NO ALERT IN OUR CPOE SYSTEM ((B)(6)) TO FLAG THE PROVIDER OR PHARMACIST OF THIS ALLERGY WHEN THE FLORANEX WAS ORDERED/VERIFIED. IN ADDITION, WHEN TRYING TO INVESTIGATE IF FLORANEX COULD HAVE BEEN THE CAUSE OF THE ALLERGIC REACTION, THE PROVIDER/ PHARMACIST WENT TO (B)(6) WHICH DOES NOT HAVE THIS LISTED AS A WARNING, PRECAUTION OR CONTRAINDICATION (AS IT DOES WITH OTHER FOOD/DRUG ALLERGIES SUCH AS PROPOFOL WITH EGG/SOY). IT IS ONLY LISTED UNDER DOSAGE FORMS WHICH IS NOT THE FIRST PLACE THE PROVIDER OR PHARMACIST WHO TYPICALLY SEARCH FOR THIS INFO. WE ARE NOTIFYING BOTH (B)(6) TO SUGGESTION CHANGES TO THEIR DATA BASE. UNTIL THIS IS DONE, WE ARE ADDING A CUSTOM ALERT IN (B)(6) TO NOTIFY THE PHARMACIST. WE WANTED TO SHARE THIS INFO AS IT IS VULNERABILITY IN A CPOE SYSTEM THAT MANY HOSPITALS USE. IT PRESENTS A SAFETY CONCERN AS MORE PTS COME INTO HOSPITALS ON HOME DIETARY SUPPLEMENTS" THAT ARE NOT FDA REGULATED AND MAY NOT HAVE SAFETY CHECKS SUCH AS ALLERGY AND INTERACTION CHECKING BUILT INTO CPOE DATABASES. I THOUGHT THIS MIGHT BE A GOOD ITEM FOR THE (B)(6) NEWSLETTER. PLEASE LET ME KNOW IF YOU NEED ANY ADD'L INFO. (B)(6)." | 1
4790395 | DAVINCI SI MONOPOLAR SCISSORS BROKE WHILE INSIDE OF PATIENT, STAFF NOTICED AND SCISSORS WHERE REMOVED AND REPLACED WITH NEW ONES. | 0
5090830 | .><ICU RN IN PATIENT'S ROOM WITNESSED THE MONITOR ABOVE THE PATIENT GO BLANK. SMOKE STARTED COMING OUT OF THE COMPUTER. THE NURSING STAFF IMMEDIATELY EVACUATED THE PATIENT AND CONTACTED BIOMED STAFF TO RESPOND. BIOMED UNPLUGGED THE DEVICE AND RE-PLUGGED THE DEVICE, WHICH CAUSED THE SMOKE TO REOCCUR. THE DEVICE WAS SEQUESTERED. THE PATIENT WAS NOT HARMED. | 1
5156379 | PHYSIOLOGICAL MONITOR DASH 3000 FAILED TO ALARM AND PATIENT EXPIRED. INVESTIGATION CONTINUING. | 1
5152216 | PHYSIOLOGICAL MONITOR DASH 3000 FAILED TO ALARM AND PT EXPIRED. INVESTIGATION CONTINUING. | 1
4557254 | IN THE MONTH OF DECEMBER, CENTRAL TELEMETRY MONITORING SYSTEM IF EXPERIEMCING SOME ITERFERENCE WITH THE TELEMETRY SIGNAL RESULTING IN SIGNAL DROP OUT. LOSS OF SIGNAL WAS OBSERVED ON VARIOUS UNITS TROUBLE SHOOTING TO IMPROVE SIGNAL AT BEDSIDE NOT SUCCESSFUL, TRANSCIEVERS WERE FUNCTIONAL. THUS IS WAS DETERMINED TO BE A NETWORK ACCESS POINT (AP) ISSUE. SEVERAL ACCESS POINTS WERE NOT SYNCING CORRECTLY, AFTER REBOOTING THE ACCESS POINT SIGNAL LOSS ISSUE APPEARED TO BE SOLVED BUT CONTINUED MONITORING SHOWED THAT SIGNAL LOSS WAS RECURRENT.======================MANUFACTURER RESPONSE FOR TELEMETRY MONITORING NETWORK ACCESS POINT, PATIENT NET ACCESS POINT (PER SITE REPORTER).======================GE RECIEVED NETWORK RECORDS AND IDENTIFIED AP RELATED ERRORS, BAD GHOST" AP | 1
4730617 | PRODUCT QUALITY COMPLAINT FOLLOW UP RECEIVED: ON (B)(6) 2014, A PREGNANT FEMALE CONSUMER REPORTED A POSSIBLE DRUG EXPOSURE DURING PREGNANCY. THE PREGNANT FEMALE CONSUMER OF UNKNOWN AGE STARTED PRANACTIN-CITRIC, 1 DOSE, ONCE, ON (B)(6) 2014 FOR A CHRONIC SORE THROAT." SHE HAD TAKEN 1 DOSE OF PRANACTIN-CITRIC ON (B)(6) 2014 AND STATED "I JUST FOUND OUT THAT I AM PREGNANT." IT WAS UNKNOWN IF THE CONSUMER WAS PREGNANT WHEN SHE RECEIVED BREATH TEK. IT WAS UNKNOWN IF THERE WAS ANY RELEVANT LABORATORY DATA. IT WAS REPORTED IN THIS CASES THAT AS OF (B)(6) 2014 | 0
4623890 | ><UNIT BEGAN SKIPPING SHOCKS THEN DISPLAYED ERROR 212 - HIGH VOLTAGE ERROR. FIELD SERVICE ENGINEER DISPATCHED. (B)(4). PATIENT UNDER ANESTHESIA WHEN MACHINE WENT DOWN. CASE POSTPONED AND COMPLETED LATER THAT SAME DAY. | 0
5008104 | ><FIELD SERVICE ENGINEER FOUND THE COUPLING PUMP WEAK. REPLACED COUPLING PUMP AND WAS ABLE TO FILL THERAPY HEAD AND TEST FIRE SHOCKS. | 0
4795200 | A SPECIFIC ROOT CAUSE COULD NOT BE DETERMINED BASED ON THE PROVIDED INFORMATION. THE REASON FOR THE SAMPLES NOT BEING CENTRIFUGED COULD NOT BE FULLY CLARIFIED AS THE NECESSARY DATA WAS REQUESTED, BUT NOT AVAILABLE FOR INVESTIGATION.><THE CUSTOMER INITIALLY REPORTED THAT THEY WERE GETTING ERRORS ON THEIR MPA PREANALYTICS SYSTEM. THE CUSTOMER RESET THE SYSTEM AND PERFORMED A SHUT DOWN. THE CUSTOMER ALSO REPORTED THAT THERE WERE SIX PATIENT SAMPLES THAT WERE SENT THROUGH THE SYSTEM AND WERE NEVER CENTRIFUGED. THESE SIX SAMPLES THEN PASSED ON TO A C501 ANALYZER FOR PROCESSING. AS A RESULT OF THE ISSUE, THE CUSTOMER SAID THAT ION SELECTIVE ELECTRODE (ISE) POTASSIUM AND CREATININE RESULTS FOR FIVE PATIENT SAMPLES WERE QUESTIONED. THE CUSTOMER STATED THAT THEY ISSUED CORRECTED REPORTS FOR ALL FIVE SAMPLES. OF THESE FIVE SAMPLES, IT WAS DETERMINED THAT THREE SAMPLES HAD ERRONEOUS ISE POTASSIUM RESULTS THAT WERE REPORTED OUTSIDE OF THE LABORATORY. THE FIRST SAMPLE INITIALLY WAS PROCESSED ON THE MPA SYSTEM AND RESULTED AS 5.04 MEQ/L. A VALUE OF 5.0 MEQ/L VALUE WAS REPORTED OUTSIDE OF THE LABORATORY. THE PRIMARY TUBE OF THE SAMPLE WAS REPEATED AND RESULTED AS 4.27 MEQ/L. THE REPEAT RESULT WAS BELIEVED TO BE CORRECT AND WAS REPORTED OUTSIDE OF THE LABORATORY AS 4.3 MEQ/L. THE SECOND SAMPLE, FROM A (B)(6) MALE BORN ON (B)(6) 1963, WAS INITIALLY PROCESSED ON THE MPA SYSTEM AND RESULTED AS 4.71 MEQ/L. A 4.7 MEQ/L VALUE WAS REPORTED OUTSIDE OF THE LABORATORY. THE PRIMARY TUBE OF THE SAMPLE WAS REPEATED AND RESULTED AS 3.93 MEQ/L. THE REPEAT RESULT WAS BELIEVED TO BE CORRECT AND REPORTED OUTSIDE OF THE LABORATORY AS 3.9 MEQ/L. THE THIRD SAMPLE, FROM A (B)(6) MALE BORN ON (B)(6) 1922, WAS INITIALLY PROCESSED ON THE MPA SYSTEM AND RESULTED AS 5.08 MEQ/L. A 5.1 MEQ/L VALUE WAS REPORTED OUTSIDE OF THE LABORATORY. THE PRIMARY TUBE OF THE SAMPLE WAS REPEATED AND RESULTED AS 4.33 MEQ/L. THE REPEAT RESULT WAS BELIEVED TO BE CORRECT AND WAS REPORTED OUTSIDE OF THE LABORATORY AS 4.3 MEQ/L. THE PATIENTS WERE NOT ADVERSELY AFFECTED. THE FIELD SERVICE REPRESENTATIVE DETERMINED THAT A STOPPER ASSEMBLY WAS MISADJUSTED ON THE CENTRIFUGE UNIT OF THE PREANALYTICS SYSTEM. THE MISADJUSTMENT CAUSED BOTH CENTRIFUGE UNITS TO STOP. WHEN THE CENTRIFUGE UNIT WAS RESET, THE PARTIALLY CENTRIFUGED SAMPLES WERE THEN SENT ON TO THE C501 ANALYZER. HE ADJUSTED THE STOPPER ASSEMBLY. THE CUSTOMER THEN RAN A NUMBER OF RACKS THROUGH THE SYSTEM AND ENCOUNTERED NO ISSUES. THERE HAVE BEEN NO FURTHER ISSUES REPORTED WITH THE MPA SYSTEM.><IT WAS UNKNOWN IF THE INITIAL REPORTER SENT A REPORT TO THE FDA. RESULTS-DEVICE SUBASSEMBLY= STOPPER ASSEMBLY. | 0
4868571 | THE CUSTOMER HAD AN ISSUE WITH A DAMAGED BAR CODE LABEL ON A PARENT TUBE THAT CAUSED THE ALIQUOT TO BE MISLABELED AND TESTED AS ANOTHER PATIENT. A SCRATCH ON THE BARCODE CAUSED A NUMBER 4 TO BE READ AS A NUMBER 7. THE CUSTOMER WAS ABLE TO REPRODUCE THE INCORRECT READING OF THE BARCODE. THE FIELD SERVICE REPRESENTATIVE FOUND THE GRIPPER FINGERS HAD DAMAGED THE LABEL AND HE REPLACED THE GRIPPER FINGERS. OPERATIONAL AND MECHANICAL CHECKS PASSED AND ADJUSTMENTS WERE OK. THE CUSTOMER STATED CHEMISTRY AND TSH" WERE TESTED ON THE SAMPLE. THE CUSTOMER REFUSED TO PROVIDE SPECIFIC DATA. THE INCORRECT RESULTS WERE REPORTED OUTSIDE THE LABORATORY AND A CORRECTED REPORT WAS SENT. THE CUSTOMER DID NOT BELIEVE THE PATIENT WAS ADVERSELY AFFECTED.><IT WAS UNKNOWN IF THE INITIAL REPORTER SENT REPORT TO THE FDA.><A SPECIFIC ROOT CAUSE COULD NOT BE IDENTIFIED. ADDITIONAL INFORMATION FOR FURTHER INVESTIGATION WAS REQUESTED BUT WAS NOT PROVIDED." | 0
4544960 | IT IS NOT KNOWN IF THE INITIAL REPORTER HAS OR INTENDS TO REPORT THE EVENT TO FDA.><REPORTER STATED THAT A PATIENT SAMPLE WAS TESTED ON THE URISYS 1100 WHICH RETURNED A NEGATIVE RESULT FOR LEUKOCYTES. WITHIN 4 HOURS, THE SAME SAMPLE WAS REPORTEDLY TESTED ON A LABORATORY INSTRUMENT WHICH REPORTED A VALUE OF ++ FOR LEUKOCYTES (APPROXIMATELY 300 WBC/HPF). NO ADVERSE EVENT WAS REPORTED. THE MANUFACTURER REQUESTED THE RETURN OF THE SUSPECT PRODUCT FOR EVALUATION. | 0
5140112 | .><THE CAMERA ON THE ROBOT WOULD NOT ALIGN CORRECTLY. PATIENT WAS IN THE ROOM, UNDER ANESTHESIA WHILE THE NURSES AND TECH TROUBLESHOOTED. THE SCOPE, CAMERA CORD, AND LIGHT CORD WERE ALL REPLACED. THE ROBOT WAS REBOOTED 3 TIMES. NONE OF THE FOLLOWING RESOLVED THE ISSUE. THE INTUITIVE HOTLINE NUMBER WAS CALLED FOR FURTHER TROUBLESHOOTING HOWEVER IT WAS DETERMINED THAT AN ENGINEER WOULD NEED TO COME OUT AND LOOK AT THE ROBOTIC TOWER. NOTHING COULD BE DONE BY THE OPERATING ROOM STAFF TO FIX THE PROBLEM. OUR PROCESS IS TO HAVE 3 OPERATING ROOM STAFF IN THE ROOM FOR ROBOTIC CASES DURING TURNOVER, HOWEVER, WE ONLY HAD TWO OPERATING ROOM STAFF MEMBERS TO START THIS CASE. THIS CAUSED A DELAY IN GETTING EQUIPMENT SETUP. PROCEDURE CANCELLED AND RESCHEDULED. | 0
4643649 | ><THE LITHOTRIPSY MACHINE WAS CHECKED BY THE TECHNICIAN BEFORE THE PATIENT ENTERED. PATIENT WAS PUT UNDER GENERAL ANESTHESIA. THE EQUIPMENT FAILED AT THE START OF THE PROCEDURE. PATIENT AWAKENED AND TRANSFERRED TO POST-ANESTHESIA CARE UNIT (PACU). NO EXTRACORPOREAL SHOCK WAVE LITHOTRIPSY (ESWL) PERFORMED.======================MANUFACTURER RESPONSE FOR LITHOTRIPSY MACHINE - LITHOTRON GOLD, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================COMPANY TECH PRESENT. | 0
4914828 | THE DOCTOR WAS INTERRUPTED MID PROCEDURE DUE TO MALFUNCTION OF PK FORCEPS. HE STATED THE WIRE BROKE THAT CONTROLS THE ACTION OF THE JAWS. FIVE OUT OF TEN LIVES REMAINED ON FORCEPS. LOT# N10150220 014. | 0
5009867 | DURING ROBOTIC PARTIAL NEPHRECTOMY CASE, ROBOTIC LARGE NEEDLE HOLDER STOPPED WORKING DURING CASE. NO PT INJURY. INSTRUMENT REPLACED WITH ANOTHER ONE.><. | 0
5309806 | .><WHILE DOING THE RIGHT PELVIC LYMPHADENECTOMY, AN INJURY TO THE RIGHT EXTERNAL ILIAC VEIN WAS NOTED TO BE APPROXIMATELY 8 MM IN LENGTH. IMMEDIATELY, PNEUMOPERITONEUM PRESSURE WAS INCREASED AND THE VASCULAR INJURY WAS CLOSED WITH SUTURE, OBTAINING HEMOSTASIS. IT WAS THEN DISCOVERED THAT THE GROUNDING SHEET OF THE ROBOTIC MONOPOLAR INSTRUMENT HAD A HOLE IN IT WHICH CAUSED THE INJURY. THE INSTRUMENT WAS REMOVED FROM THE ABDOMEN AND DISCARDED. THE TIP COVER WITH THE HOLE IS USED WITH THE DAVINCI XI SURGICAL SYSTEM ROBOTIC ARM MONOPOLAR CURVED SCISSORS (8MM) VER.-12 REF 470179. I HAVE PICTURES OF THE COVER WITH THE HOLE IN IT THAT I CAN FORWARD TO YOU. | 0
4716951 | DURING DAVINCI TUBAL ANASTOMOSIS, THE PK WAS REQUESTED TO BE USED. PK WAS PLACED INTO ARM BUT NEVER SHOWED UP ON THE INVENTORY MANAGEMENT LOG AS BEING USED AND IT NEVER LET THE DOCTOR USE THE INSTRUMENT. A NEW PK WAS OPENED. NO HARM TO THE PATIENT. | 0
4767945 | VESSEL SEALER STARTED TO ALARM AND STOPPED WORKING. | 0
4680358 | WHILE SUTURING WITH A MEGA SUTURE CUT NEEDLE DRIVER DURING A ROBOTIC NISSEN FUNDIPLICATION, ONE SIDE OF THE INSTRUMENT FELL OFF. THE INSTRUMENT TIP WAS RETREIVED BY THE SURGEON. NO HARM WAS EVIDENT, PHYSICAL OR OTHERWISE. | 0
5294627 | THE INITIAL MDR 2432235-2015-00580 WAS FILED ON DECEMBER 14, 2015. THE FIRST SUPPLEMENTAL MDR 2432235-2015-00580_S1 WAS FILED ON DECEMBER 30, 2015. ADDITIONAL INFORMATION (01-06-2016): AFTER FURTHER INVESTIGATION, SIEMENS DETERMINED THAT A USER ERROR CONTRIBUTED TO THE ISSUE. THE CUSTOMER WAS INCORRECTLY CONFIGURING THEIR SIDS. IN ADDITION TO ADJUSTING THE CONFIGURATION OF THE MISPL TO USE A SHORTER PURGE INTERVAL, THE CUSTOMER HAS BEEN INFORMED OF THE PROPER USE AND CONFIGURATION OF SIDS.><A SIEMENS CUSTOMER SERVICE ENGINEER (CSE) PERFORMED TROUBLESHOOTING REMOTELY. THE CSE DISCOVERED THAT THE ISSUE WAS A CONFIGURATION OF THE PURGE MISPL RULE. THE PURGE INTERVAL ALLOWED FOR A SAMPLE ID FROM (B)(6) 2015 TO BE RE-USED BEFORE THE OLD DATA WAS PURGED. THIS CAUSED THE WRONG PATIENT INFORMATION TO BE SENT TO THE DIMENSION VISTA INSTRUMENT. THE CSE DISCOVERED THAT THE SAMPLE IN QUESTION WAS NOT PURGED AS IT REMAINED IN A PENDING TO REVIEW" STATUS. THE ORDER ON THE DAY OF THE EVENT (B)(6) 2015 FOR THIS SAMPLE WAS REJECTED | 1
4517920 | THE CUSTOMER REPORTED THE CENTRALINK AUTO VERIFICATION RULE IS NOT FUNCTIONING CORRECTLY. THE CENTRALINK RULE FOR SAMPLE RERUNS IS SET TO NOT AUTO VALIDATE SAMPLE RERUNS THOUGH ARE STILL BEING AUTO VALIDATED ON THE CUSTOMERS ADVIA 2400 SYSTEM. THERE WERE NO REPORTS OF ADVERSE HEALTH CONSEQUENCES OR KNOWN PATIENT INTERVENTION DUE TO THE AUTO VERIFICATION RULE NOT FUNCTIONING CORRECTLY.><A SIEMENS CUSTOMER SERVICE ENGINEER (CSE) WAS DISPATCHED TO THE CUSTOMER SITE. THE CSE ANALYZED THE INSTRUMENT AND INSTRUMENT DATA AND DETERMINED THAT RESULTS WERE HELD FOR VALIDATION IN THE LABORATORY INFORMATION SYSTEM. THE CSE DETERMINED THAT NO INCORRECT RESULTS WERE REPORTED TO THE PHYSICIAN(S). THE CSE INCREASED THE SEVERITY SETTING THAT CENTRALINK ASSIGNS TO THE CUSTOMERS ADVIA 2400 INSTRUMENT FLAG R" (REPEATED RESULT) TO 2 IN ORDER TO HOLD THOSE RESULTS THAT ARE AUTO REPEATED IN THIS INSTRUMENT. THE CAUSE OF THE AUTO VERIFICATION RULE NOT FUNCTIONING CORRECTLY IS UNKNOWN. THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. FURTHER EVALUATION OF THE DEVICE IS NOT REQUIRED." | 1
4539737 | JAWS OF THE FORCEPS DO NOT MEET TOGETHER; THE INSTRUMENT HAS 8 USES LEFT.MANUFACTURER RESPONSE FOR ROBOTIC BIPOLAR FORCEP, (BRAND NOT PROVIDED) (PER SITE REPORTER)======================TOOK DESCRIPTION OF WHAT HAPPENED AND ISSUED RGA#. | 0
4797323 | THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT><THE CUSTOMER REPORTED A PROBLEM WITH A FALSE NEGATIVE RESULT FOR SSCII TEST (P1+P2 BOTH SHOULD HAVE ++ AND THE RESULT WAS - FOR BOTH) WHEN TESTING ON TANGO OPTIMO. A FIELD SERVICE ENGINEER INSPECTED THOROUGHLY THE AFFECTED TANGO INSTRUMENT AND FOUND THE CAMERA VALUES TO BE OUT OF SPECIFICATION. POSITIVE TEST RESULTS WOULD BE INTERPRETED AS BEING NEGATIVE BY THE CAMERA. THE CAMERA WAS READJUSTED TO BE WITHIN SPECIFICATION. SINCE THEN THE INSTRUMENT HAS BEEN RUNNING WITHOUT ANY OTHER COMPLAINT FROM THIS CUSTOMER. | 0
4896945 | COVIDIEN REFERENCE: (B)(4). ONE PULSE OXIMETER WAS RECEIVED FOR EVALUATION. A VISUAL INSPECTION FOUND NO ANOMALIES. PERFORMANCE TESTS WERE CONDUCTED, AND OXYGEN SATURATION(SPO2) WAS TESTED WITH DIFFERENT SETTINGS. NO ERRORS OR ALARMS OCCURRED DURING THE TESTS. POWER ON AND OFF TESTS WERE PERFORMED SEVERAL TIMES IN AN ATTEMPT TO REPRODUCE THE REPORTED FAILURE. NO ERRORS OR ALARMS WERE OBSERVED. THE CUSTOMER REPORTED MALFUNCTION WAS NOT VERIFIED AS THE DEVICE WAS FOUND TO BE FUNCTIONING AS INTENDED.><(B)(4)><IT WAS REPORTED THAT A PULSE OXIMETER READING REACHED THE LOWER LIMIT AND THEN READ ZERO AS THE PATIENT&#39;S OXYGEN SATURATION DECREASED. THE CUSTOMER STATED THAT IT WAS DIFFICULT TO ASSESS THE PATIENT&#39;S CONDITION TO DETERMINE WHETHER TO TAKE ANY FURTHER ACTION. THE NEONATAL TEAM WAS UNABLE TO PROVIDE PATIENT INFORMATION. THE MONITOR WAS BEING USED WITH MAX NI SENSORS, WHICH WERE DISCARDED BY THE HOSPITAL STAFF. THERE IS NO REPORT OF SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT. | 1
4847064 | DAVINCI INSTRUMENT, BIPOLAR FORCEP REF# (B)(4) FAILED TO DESICCATE TISSUE. THE GROUNDING PAD AND BIPOLAR CORD CONNECTIONS WERE CHECKED WITHOUT RESOLVE. THE POLAR CORD WAS REPLACED STILL WITHOUT RESOLVE. ANOTHER MARYLAND BIPOLAR INSTRUMENT REPLACED THE ORIGINAL AND THE ISSUE WAS RESOLVED. NO PATIENT INJURY OCCURRED AND A PROCEDURAL DELAY OF APPROX 2 MIN RESULTED. THE SURGEON WAS PRESENT AT THE TIME OF INSTRUMENT MALFUNCTION. | 0
4538405 | STAPLER WORKED TO START WITH HOWEVER, AFTER 15 USES IT WOULD NOT ENGAGE WITH THE ROBOTIC ARM.MANUFACTURER RESPONSE FOR ROBOTIC STAPLER, (BRAND NOT PROVIDED) (PER SITE REPORTER)======================UNKNOWN; THIS WAS REPORTED BY THE OR STAFF. | 0
5001138 | NO ITEMS OF EXTERNAL VISUAL DAMAGE NOTED. NEGATIVE BATTERY CONNECTION IN BATTERY COMPARTMENT WAS DEFORMED AND PRESSED IN AS RECEIVED. MONITOR DID NOT FUNCTION AS RECEIVED, AS NEGATIVE BATTERY TERMINAL WAS DEFORMED, AND CONTACT WOULD NOT BE MADE WITH THE NEGATIVE POST OF THE BATTERY. THE BATTERY CONNECTION COULD BE FORCED BY INSERTING AN APPROXIMATELY 1/4 INCH THICK NUT BETWEEN THE BOTTOM OF THE BATTERY AND THE BATTERY COMPARTMENT. USER DAMAGE IN INSERTING A FOREIGN OBJECT IN THE BATTERY COMPARTMENT, POSSIBLY IN AN ATTEMPT TO ESTABLISH BATTERY CONNECTION, DEFORMING BATTERY TERMINALS. REPLACEMENTS WERE PROVIDED TO FACILITY.><ON 05/05/2015, (B)(4) RECEIVED REPORT FROM A DISTRIBUTOR THAT A UMP PULL-STRING MONITOR FAILED TO ALARM WHEN A RESIDENT GOT OUT OF BED AND FELL. FURTHER INFORMATION GATHERED INDICATED TWO SUCH INCIDENTS HAD OCCURRED. IN THE FIRST INCIDENT, OCCURRING (B)(6) 2015, THE RESIDENT WAS MONITORED IN BED DURING NIGHT SHIFT AND FELL, SUFFERING A BROKEN HIP REQUIRING SURGERY TO REPAIR. IN THE SECOND INCIDENT, OCCURRING (B)(6) 2015, THE RESIDENT WAS MONITORED IN A CHAIR DURING NIGHT SHIFT AND FELL, SUFFERING A BROKEN TIBIA REQUIRING SURGERY TO REPAIR. FACILITY REPORTED THAT BOTH SURGERIES WERE SUCCESSFUL. | 1
5123040 | THE MANUFACTURER'S INVESTIGATION CONCLUDED THAT THE INCIDENT HAPPENED WHEN BOTH THE ELEKTA OY SERVICE ENGINEER AND PRODUCT DEVELOPMENT ENGINEERS WERE ON SITE TO PERFORM AN AFTER-INSTALLATION SERVICE TO THE PILOT INTERNAL HELIUM RECYCLER (IHR). WHEN SERVICING THE COMPRESSOR ASSEMBLY IN THE CABINET, THE ENGINEER FOLLOWED THE SERVICE INSTRUCTIONS FOR CHANGING THE GAS COLLECTION PUMP. THE COMPRESSOR ASSEMBLY RESEMBLES MECHANICALLY THE GAS COLLECTION PUMP ASSEMBLY, BUT DIFFERS IN RESPECT OF POWERING ARRANGEMENT: THE MAIN POWER CABLE OF THE CABINET IS CONNECTED TO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY FROM WHICH IT CONTINUES TO UPS. THE UPS POWERS THE CABINET. OUTPUTS OF THE UPS WERE DETACHED ACCORDING TO THE INSTRUCTION, BUT THE MAIN POWER CABLE WAS STILL IN CONTACT WITH MAINS POWER AND THUS ALSO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY WAS HAVING THE MAINS VOLTAGE. THE ELECTRIC SHOCK SOURCE WAS THE MAIN POWER CABLE. THE SERVICE ENGINEER DETACHED THE MAIN POWER CABLE FROM THE DIN RAIL, AND THE BARE END OF THE CABLE TOUCHED HIS ARM. THE MAIN POWER CABLE COULD HAVE BEEN DEPOWERED BY SWITCHING THE POWER OFF FROM THE MAIN SWITCH. POWER SUPPLY OF THE DEVICE IS FED THROUGH A MEDICAL SAFETY ISOLATION TRANSFORMER. IN THIS CASE THE MEDICAL SAFETY ISOLATION TRANSFORMER PREVENTED WORSE OUTCOME, AS INTENDED. ONLY TRAINED ELEKTA SERVICE ENGINEERS ARE ALLOWED TO WORK ON THE IHR. IN THE BEGINNING OF THE SERVICE MANUAL THERE IS A GENERAL INSTRUCTION TO ALWAYS DEPOWER THE SYSTEM FROM THE MAIN SWITCH, BUT THIS INSTRUCTION IS NOT REPEATED SEPARATELY FOR EACH TASK. AS THERE WERE BOTH SERVICE AND PRODUCT DEVELOPMENT PARTICIPATING TO THE SERVICE, THERE MIGHT HAVE BEEN SOME CONFUSION REGARDING WHO IS ULTIMATELY RESPONSIBLE FOR ENSURING THAT THE MAINS POWER IS POWERED OFF. THE ULTIMATE ROOT CAUSE FOR THE INCIDENT WAS HUMAN ERROR. ELEKTA OY HAS OPENED (B)(4) TO HANDLE THE CORRECTIVE AND PREVENTIVE ACTIONS. AS THE ACTIONS INCLUDED AN UPDATE IN THE FIELD, ELEKTA OY HAS REPORTED THE FIELD SAFETY CORRECTIVE ACTION UNDER REFERENCE NUMBER FCA-OY-0001 ON 30TH OCTOBER 2015. THE CORRECTIVE AND PREVENTIVE ACTIONS ARE LISTED AS FOLLOWS:- - REFRESHER TRAINING ON ELECTRICAL SAFETY HAS BEEN GIVEN TO ELEKTA OY SERVICE ENGINEERS. ALSO, ALL ELEKTA OY EMPLOYEES WHO WORK ON ELECTRICAL INSTALLATIONS HAVE ATTENDED ADDITIONAL EXTERNAL ELECTRICAL SAFETY TRAINING. - ELEKTA OY WILL CREATE A PROCEDURE FOR ELECTRICAL WORK. ALL APPLICABLE EMPLOYEES WILL BE TRAINED. ESTIMATED COMPLETION BY END OF JANUARY 2016. - THE DESIGN OF ALL NEW IHRS HAS BEEN CHANGED, SO THAT THE MAIN POWER CABLE IS RE-ROUTED NOT GO THROUGH THE COVER OF THE COMPRESSOR ASSEMBLY. IN ADDITION, A WARNING STICKER HAS BEEN ADDED TO THE DESIGN TO REMIND THE SERVICE ENGINEER TO ALWAYS DEPOWER THE MAINS POWER BEFORE SERVICING. THIS DESIGN CHANGE WILL BE IMPLEMENTED TO THE DEVICES ON THE FIELD THROUGH THE FCA-OY-0001. NOTE THAT THE WARNING STICKER HAS ALREADY BEEN ATTACHED TO THE DEVICES IN THE FIELD. - THE IHR SERVICE MANUAL WILL BE UPDATED SO THAT THE INSTRUCTION TO DEPOWER THE INSTALLATION IS STATED IN THE BEGINNING OF EACH TASK. ESTIMATED COMPLETION BY END OF JANUARY 2016. - ELEKTA OY MANUAL VALIDATION INSTRUCTION HAS BEEN UPDATED TO INCLUDE A REVIEW OF AN ELECTRICAL SAFETY EXPERT FOR SERVICE AND INSTALLATION MANUALS. ESTIMATED RELEASE BY END OF JANUARY 2016. THE DESIGN CHANGE WILL BE IMPLEMENTED BY THE END OF APRIL 2016. THE MANUFACTURER IS NOT AWARE OF SIMILAR INCIDENTS WITH THIS TYPE OF MEDICAL DEVICE WITH A SIMILAR ROOT CAUSE. THE DEVICE HAS NOT BEEN DISTRIBUTED IN THE US.><AN ELEKTA SERVICE ENGINEER WAS PERFORMING A SERVICE OPERATION OF THE HELIUM RECYCLER CABINET OF THE INTERNAL HELIUM RECYCLER (IHR). DURING THE SERVICE, HE GOT AN ELECTRIC SHOCK IN HIS ARM DUE TO WHICH THE ARM WAS TEMPORARILY NUMB BUT THERE WAS NO PERMANENT INJURY. THE ENGINEER VISITED THE DOCTOR AND HAD A HEALTH CHECK THE FOLLOWING DAY. EVERYTHING WAS OK. WHEN SERVICING THE COMPRESSOR ASSEMBLY IN THE CABINET, THE ENGINEER FOLLOWED THE SERVICE INSTRUCTIONS FOR CHANGING THE GAS COLLECTION PUMP. THE POWER CABLE OF THE CABINET IS CONNECTED TO DIN RAIL OF THE COMPRESSOR ASSEMBLY FROM WHICH IT CONTINUES TO THE UPS (UNINTERRUPTIBLE POWER SUPPLY). THE UPS POWERS THE CABINET. OUTPUTS OF THE UPS WERE DETACHED ACCORDING TO THE INSTRUCTION, BUT THE MAIN POWER CABLE WAS STILL IN CONTACT WITH MAINS POWER AND THUS ALSO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY WAS HAVING THE MAINS VOLTAGE. THE ELECTRIC SHOCK SOURCE WAS THE MAIN POWER CABLE, WHICH WAS CONNECTED TO THE DIN RAIL. THE MAIN POWER CABLE COULD HAVE BEEN DEPOWERED BY SWITCHING THE POWER OFF FROM THE MAIN SWITCH.><AN ELEKTA SERVICE ENGINEER WAS PERFORMING A SERVICE OPERATION OF THE HELIUM RECYCLER CABINET OF THE INTERNAL HELIUM RECYCLER (IHR). DURING THE SERVICE, HE GOT AN ELECTRIC SHOCK IN HIS ARM DUE TO WHICH THE ARM WAS TEMPORARILY NUMB BUT THERE WAS NO PERMANENT INJURY. THE ENGINEER VISITED THE DOCTOR AND HAD A HEALTH CHECK THE FOLLOWING DAY. EVERYTHING WAS OK. WHEN SERVICING THE COMPRESSOR ASSEMBLY IN THE CABINET, THE ENGINEER FOLLOWED THE SERVICE INSTRUCTIONS FOR CHANGING THE GAS COLLECTION PUMP. THE POWER CABLE OF THE CABINET IS CONNECTED TO DIN RAIL OF THE COMPRESSOR ASSEMBLY FROM WHICH IT CONTINUES TO THE UPS (UNINTERRUPTIBLE POWER SUPPLY). THE UPS POWERS THE CABINET. OUTPUTS OF THE UPS WERE DETACHED ACCORDING TO THE INSTRUCTION, BUT THE MAIN POWER CABLE WAS STILL IN CONTACT WITH MAINS POWER AND THUS ALSO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY WAS HAVING THE MAINS VOLTAGE. THE ELECTRIC SHOCK SOURCE WAS THE MAIN POWER CABLE, WHICH WAS CONNECTED TO THE DIN RAIL. THE MAIN POWER CABLE COULD HAVE BEEN DEPOWERED BY SWITCHING THE POWER OFF FROM THE MAIN SWITCH.><THE MAIN DEVICE (ELEKTA NEUROMAG TRIUX) IS A MAGNETOENCEPHALOGRAPH, THE MAGNETIC COUNTER PART OF EEG. THE DEVICE MEASURES MAGNETIC FIELD, WHICH IS VERY WEAK, AND CAN ONLY BE MEASURED WITH SUPER CONDUCTING SENSORS. THUS THE SENSORS ARE KEPT IN LIQUID HELIUM. THE PART INVOLVED IN THE INCIDENT WAS THE INTERNAL HELIUM RECYCLER (IHR), WHICH IS AN OPTIONAL NON-MEDICAL DEVICE BUT WILL BECOME A PART OF MEDICAL DEVICE WHEN INSTALLED. THE PURPOSE OF THE IHR IS TO REDUCE THE HELIUM CONSUMPTION. THE IHR IS AN OPTION FOR THE MAIN SYSTEM. THERE ARE ONLY TWO INSTALLATIONS WORLD WIDE, BOTH IN (B)(4). THE MANUFACTURER'S PRELIMINARY INVESTIGATION FOUND THAT THE INCIDENT HAPPENED WHEN BOTH THE ELEKTA OY SERVICE ENGINEER AND PRODUCT DEVELOPMENT ENGINEERS WERE ON SITE TO PERFORM AFTER-INSTALLATION SERVICE TO THE PILOT DEVICE. IN THE FUTURE, ONLY ELEKTA SERVICE ENGINEERS ARE ALLOWED TO SERVICE IHR, THUS HOSPITAL ENGINEERS WILL NEVER SERVICE THE IHR. AT PRESENT, THERE ARE ONLY FOUR SERVICE ENGINEERS TRAINED TO WORK ON IHR, ALL ELEKTA OY EMPLOYEES. WHEN SERVICING THE COMPRESSOR ASSEMBLY IN THE CABINET, THE ENGINEER FOLLOWED THE SERVICE INSTRUCTIONS FOR CHANGING THE GAS COLLECTION PUMP. THE POWER CABLE OF THE CABINET IS CONNECTED TO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY FROM WHICH IT CONTINUES TO UPS. THE UPS POWERS THE CABINET. OUTPUTS OF THE UPS WERE DETACHED ACCORDING TO THE INSTRUCTION, BUT THE MAIN POWER CABLE WAS STILL IN CONTACT WITH MAINS POWER AND THUS ALSO THE DIN RAIL OF THE COMPRESSOR ASSEMBLY WAS HAVING THE MAINS VOLTAGE. THE ELECTRIC SHOCK SOURCE WAS THE MAIN POWER CABLE, WHICH WAS CONNECTED TO THE DIN RAIL. THE MAIN POWER CABLE COULD HAVE BEEN DEPOWERED BY SWITCHING THE POWER OFF FROM THE MAIN SWITCH. IN THE BEGINNING OF THE SERVICE MANUAL THERE IS A GENERAL INSTRUCTION TO ALWAYS DEPOWER THE SYSTEM FROM THE MAIN SWITCH, BUT THIS INSTRUCTION IS NOT REPEATED SEPARATELY FOR EACH TASK. THUS THE ULTIMATE ROOT CAUSE FOR THE INCIDENT WAS HUMAN ERROR. AS THERE WERE BOTH SERVICE AND PRODUCT DEVELOPMENT PARTICIPATING, THERE MIGHT HAVE BEEN SOME CONFUSION REGARDING WHO IS ULTIMATELY RESPONSIBLE FOR ENSURING THAT THE MAINS POWER IS POWERED OFF. POWER SUPPLY OF THE DEVICE IS FED THROUGH A MEDICAL SAFETY ISOLATION TRANSFORMER. IN THIS CASE THE MEDICAL SAFETY ISOLATION TRANSFORMER PREVENTED WORSE OUTCOME, AS INTENDED. REFRESHER TRAINING ON ELECTRICAL SAFETY HAS BEEN GIVEN TO ELEKTA OY SERVICE ENGINEERS. ALL ELEKTA OY EMPLOYEES WHO WORK ON ELECTRICAL INSTALLATIONS WILL ATTEND ADDITIONAL ELECTRICAL SAFETY TRAINING. AN INVESTIGATION HAS BEEN INITIATED TO ASSESS THE NEED FOR FURTHER CORRECTIVE ACTIONS, DETAILS OF WHICH WILL BE INCLUDED IN THE FINAL REPORT. | 0
4861735 | ><THE CUSTOMER REPORTED THAT WHEN USING PARAGON LABORATORY MANAGEMENT RELEASE 12.1 AFTER UPGRADING TO (B)(6) SQL SERVER 2012, THEY HAD A PROBLEM GENERATING AN ACCURATE FINAL CUMULATIVE REPORT. THE DATA DISPLAY IN THE FINAL CUMULATIVE REPORT DID NOT MATCH THE DATA DISPLAY EXPECTED FROM THE DATABASE INFO. THE TEST RESULTS WERE CORRECT AND THE DATA/TIME STAMP VALUES WERE CORRECT, BUT THE REPORT DISPLAY SHOWED A MISMATCH IN THE TEST RESULT VALUE REPORTED AGAINST THE DATE/TIME STAMP. | 1
4888351 | ><DESCRIPTION: INACCURATE OR INAPPROPRIATE ALLERGIES ARE BEING DOCUMENTED IN PT CHARTS RELATED TO SEASONAL ALLERGIES". THE TERM "SEASONALE" (BRAND NAME OF AN ORAL CONTRACEPTIVE) IS BEING ACCIDENTLY SELECTED. IN OUR HIS SEASONAL ALLERGIES IS A DIAGNOSIS AND THE ALLERGY PROFILE RECORDS ONLY ACTUAL SUBSTANCES | 1
4541588 | DIGLO2 INTERNATIONAL CO LTD REGISTRATION NO (B)(4) DOES NOT HAVE BLOOD PRESSURE PRODUCT WITH PREMARKET NOTIFICATION 510 (K). BUT THEY SHOW THE PRODUCT WITH FDA" LOGO IN THE FOLLOWING WEBSITE: (B)(4) ILLEGALLY SELL THE BLOOD PRESSURE PRODUCT WITHOUT FDA VERIFICATION AND PUT "FDA" LOGO ON THE WEBSITE." | 0
5273197 | PHILIPS FIELD SERVICE ENGINEER (FSE) WAS ONSITE AND GAVE INSTRUCTIONS FOR REVIEWING MX ALARM MESSAGES. THERE WERE NO PRESSURE ALARM MESSAGES POSTED FOR THIS TIMEFRAME; THE EARLIEST MESSAGE AVAILABLE WAS AT 4:07 AM - ABP REDUCE SIZE". THE PIIC ALARM AND EVENT REVIEW DID NOT HAVE PRESSURE ALARMS POSTED. THE TREND REVIEW SHOWED NO PRESSURE VALUES FOR THE TIME FRAME PRIOR TO WHEN THE PATIENT CODED. THERE IS NO EXPLICITLY DEFINED INDICATION AS TO HOW THE REPORTED PROBLEM WAS RESOLVED | 1
4738989 | THE CUSTOMER REPORTED GETTING A SPEAKER MALFUNCTION ON AN MP2 THAT HAD THE SPEAKER REPLACED BY A FIELD SERVICE ENGINEER. NO PATIENT INVOLVEMENT WAS REPORTED.><THE CUSTOMER REPORTED THAT THE INTELLIVUE MP2 PATIENT MONITOR DISPLAYS THE ERROR MESSAGE 'SPEAKER MALFUNCTION.'" NO DEATH OR PATIENT/USER HARM WAS REPORTED.><CONTACT OFFICE CORRECTION: (B)(4).><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL | 1
4739419 | THE CUSTOMER REPORTED THAT HE RECEIVED A SPEAKER FAILURE INOP ON THE SCREEN OF THE MX700 WITH NO SOUND. NO PATIENT HARM WAS REPORTED.><.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. SUBMISSION OF A REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, DISTRIBUTOR, MANUFACTURER, OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT.><IT WAS REPORTED THAT THE CUSTOMER WAS GETTING A SPEAKER MALFUNCTION; IT WOULD NOT ALARM. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED; THERE WAS NO PATIENT OR USER HARMED. | 1
5305396 | A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><.><THE CUSTOMER REPORTED THAT A PATIENT HAD AN IDIOVENTRICULAR RHYTHM BUT AN ALARM WAS NOT PROVIDED ON (B)(6) 2015 AT 22:14:24. NO PATIENT HARM WAS REPORTED. | 1
4632530 | (B)(4).><THE CUSTOMER STATED THAT THERE WAS NO SOUND" COMING FROM THE SPEAKER. NO PATIENT INVOLVEMENT WAS REPORTED.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION." | 1
5197016 | .><THE CUSTOMER REPORTED A SPEAKER MALFUNCTION WITH LOSS OF AUDIO." THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM.><(B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE." | 1
5003235 | .><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER STATED THE HAD A SPEAKER MALFUNCTIONING ERROR" ON THE DEVICE. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM.><THE CUSTOMER STATED THAT HE HAD A "SPEAKER MALFUNCTIONING ERROR" ON THE DEVICE. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM." | 1
4617055 | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><(B)(4).><THE CUSTOMER REPORTED THAT THE INTELLIVUE MP2 PATIENT MONITOR DISPLAYS THE ERROR MESSAGE 'SPEAKER MALFUNCTION'". NO DEATH OR PATIENT/USER HARM WAS REPORTED.><THE CUSTOMER REPORTED A SPEAKER MALFUNCTION WITH LOSS OF SOUND. NO PATIENT INVOLVEMENT WAS REPORTED." | 1
5108046 | .><THE CUSTOMER REPORTED THAT M2740A ((B)(4)) MOUNTING HARDWARE ((B)(4)) BROKEN WHILE IN SITU. ARM WAS SNAGGED BY THE ELECTRIC BIRTHING BED AS IT WAS BEING RAISED. THE PIVOT ARM ACTUALLY SNAPPED DUE TO THE FORCE EXERTED AND THE AVALON FM30 FETAL MONITOR FELL TO THE FLOOR. NO ONE WAS INJURED AND NO INDICATION OF ANY KIND OF A PATIENT INVOLVEMENT."><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. (B)(4)." | 1
4599757 | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><(B)(4)><THE CUSTOMER REPORTED THAT THE MX800 IS GIVING AN ERROR MESSAGE STATING SPEAKER MALFUNCTION. THERE WAS NO PATIENT INVOLVEMENT. | 1
5104816 | A (B)(6) WAS BEING RESUSCITATED AFTER WHICH PHILIPS RECEIVED FEEDBACK INDICATING LACK OF ALARMS. THE (B)(6) PATIENT NEEDED TO BE RESUSCITATED. THE CUSTOMER MENTIONED LATER THAT THE PATIENT HAD SURVIVED.><THE CUSTOMER DECLINED TO SUPPORT THE INVESTIGATION FROM PHILIPS. THEREFORE IT IS NOT POSSIBLE TO MAKE A CLEAR CONCLUSION AS TO THE CAUSE OF THIS REPORT. NOTE THAT THE PRESSURE DIFFERENCES WERE NOTED DURING CPR. IT IS NOT UNREASONABLE THAT THE DIFFERENT TYPES OF PRESSURE MEASURMENTS WOULD BE DIFFERENT DURING CPR. THERE WAS NO TROUBLE FOUND WITH THE DEVICE AND NO REPAIR WAS WARRANTED. NO FURTHER INVESTGATION OR ACTION ARE WARRANTED.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. (B)(4).><A (B)(6) CHILD WAS BEING RESUSCITATED AFTER WHICH PHILIPS RECEIVED FEEDBACK INDICATING LACK OF ALARMS. THE USERS NOTED THAT THE PRESSURE READINGS FROM THE X2 WERE 45 MMHG DIFFERENT FROM THE PICCO MODULE. THE (B)(6) PATIENT NEEDED TO BE RESUSCITATED. THE CUSTOMER MENTIONED LATER THAT THE PATIENT HAD SURVIVED. | 1
5103743 | .><THE CUSTOMER REPORTED THAT THERE WAS A LOUDSPEAKER ERROR ON THE DEVICE AND NO SOUND COMING FROM THE LOUDSPEAKER. THERE WAS NO PATIENT INVOLVEMENT.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. (B)(4). | 1
5030883 | .><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER REPORTED THERE WAS A FAULT WITH EITHER THE (MP70) MONITOR OR MODULE X2 MONITOR SATURATION WENT DOWN TO 90. THERE WAS NO ALARM SOUND | 1
5154292 | (B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER REPORTED THERE WAS NO SOUND ON ALARM. THE DEVICE WAS IN CLINICAL USE AT THE TIME OF THE EVENT. THERE WAS NO HARM TO THE PATIENT OR USER DURING THIS ALLEGED MALFUNCTION.><THE CUSTOMER REPORTED THERE WAS NO SOUND ON ALARM. NO DEATH, SERIOUS INJURY OR ADVERSE EVENT OCCURRED OR WAS REPORTED AS A RESULT OF THIS ISSUE.><. | 1
4687124 | .><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><THE CUSTOMER STATED THAT HE REPLACED THE SPEAKER BUT STILL GETTING SPEAKER MALFUNCTION. NO PATIENT HARM WAS REPORTED. | 1
4761745 | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><THE CUSTOMER REPORTED THAT THE SCREEN NEEDS REPLACING AND THE UNIT DOES NOT ALARM PROPERLY. THERE WAS NO ALLEGATION OF A PATIENT OR USER HARM.><(B)(4).><THE CUSTOMER REPORTED THAT THE SCREEN NEEDS REPLACING AND THE UNIT DOES NOT ALARM PROPERLY. NO PATIENT HARM WAS REPORTED. | 1
4944843 | .><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><THE CUSTOMER STATED THAT THERE WAS A SPEAKER MALFUNCTION INOP ON THE DEVICE. THERE WAS NO SOUND COMING FROM THE SPEAKER. THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM. | 1
4539175 | THE CUSTOMER REPORTED THAT THE DEVICE ISSUED A :SPEAKER MALFUNCT." INOP. NO PATIENT HARM WAS REPORTED.><THE CUSTOMER REPORTED THAT THE DEVICE ISSUED A ¿SPEAKER MALFUNCT.¿ INOP. NO PATIENT HARM WAS REPORTED.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><(B)(4)." | 1
4990153 | THE DEVICE WAS NOT RETURNED TO THE FACTORY FOR EVALUATION AS THE FIELD SERVICE ENGINEER WAS ALREADY ONSITE AND TESTED THE DEVICE AND ESTABLISHED THE DEVICE FUNCTIONED ACCORDING TO MANUFACTURER SPECIFICATION PROVIDING ALARMS BOTH AUDIBLY AND VISUALLY. THE ALARM LOGS WERE REVIEWED AND SHOW THAT THE DEVICE DID SOUND AN AUDIBLE DESAT ALARM DURING THE DATE/TIMES IN QUESTION FOR BED (B)(4). THE DEVICE ALARMED SEVERAL TIMES FOR DESAT AND VENT FIB/TACH AND VTACH. THERE WAS NO MALFUNCTION OF THE DEVICE. THE PROVIDED LOGS SHOW THAT THE DEVICE ALARMED AS SPECIFIED FOR DESAT ALARMS AND RED ALARMS. THE DEVICE WAS TESTED AND WAS FOUND TO BE PERFORMING TO SPECIFICATIONS. DEVICE LABELING DESCRIBES ARRHYTHMIA ALARM CHAINING WHERE ONLY THE HIGHEST PRIORITY ALARM CONDITION IN EACH CHAIN IS ANNOUNCED. LOWER PRIORITY ALARMS IN THE SAME CHAIN WILL NOT BE ANNOUNCED WHILE AN ALARM IS ACTIVE OR DURING THE CONFIGURED TIMEOUT PERIOD. IF ALARM CONDITIONS OF EQUAL SEVERITY FROM DIFFERENT CHAINS ARE DETECTED, THE ALARM CONDITION THAT OCCURRED MOST RECENTLY IS ANNOUNCED. VTACH IS LOWER IN THE CHAIN THAN VENT FIB/TACH AND THEREFORE EITHER THE ORIGINAL VTACH ENDED ON ITS OWN OR SOMEONE SILENCED IN THE ROOM TO ALLOW A LOWER PRIORITY ALARM TO SOUND. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED.><THE CUSTOMER REPORTED THAT CUSTOMER HAS AN URGENT REQUEST TO HAVE THE LOGS REVIEWED FOR DESAT ALARMS THAT OCCURRED ON (B)(6) 2015 BETWEEN THE HOURS OF 0900-1000 ON THE SCCT61 CLASSIC PIIC FOR BED (B)(4). STAFF STATES THAT DURING THIS TIME THE PATIENT HAD A DESAT AND THE MONITOR DID NOT AUDIBLY ALARM." THE PATIENT WAS CODED WITH MEDICATION AND WAS DEFIBRILLATED TWO TIMES.><THE CUSTOMER REPORTED THAT "CUSTOMER HAS AN URGENT REQUEST TO HAVE THE LOGS REVIEWED FOR DESAT ALARMS THAT OCCURRED ON (B)(6) 2015 BETWEEN THE HOURS OF 0900-1000 ON THE SCCT61 CLASSIC PIIC FOR BED 6S14. STAFF STATES THAT DURING THIS TIME THE PATIENT HAD A DESAT AND THE MONITOR DID NOT AUDIBLY ALARM." THE PATIENT WAS CODED WITH MEDICATION AND WAS DEFIBRILLATED TWO TIMES.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. (B)(4)." | 1
4605983 | THE CUSTOMER EVALUATED THE DEVICE.><THE CUSTOMER STATED THAT THE SPO2 DROPPED OUT. NO PATIENT INVOLVEMENT WAS REPORTED.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | 1
5038294 | THE CUSTOMER REPORTED A SPEAKER MALFUNCTION. THERE WAS NO ALLEGATION OF A PATIENT OR USER HARM.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. (B)(6).><. | 1
4944841 | THE CUSTOMER STATED THAT THERE WAS NO SOUND" COMING FROM THE SPEAKER. THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM.><.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION." | 1
4944842 | (B)(4).><THE CUSTOMER REPORTED THAT THE NURSE'S ARE SAYING THIS UNIT IS NOT ALARMING CORRECTLY. NO PATIENT HARM WAS REPORTED.><A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><THE CUSTOMER REPORTED THAT THE NURSE'S ARE SAYING THIS UNIT IS NOT ALARMING CORRECTLY. NO PATIENT HARM WAS REPORTED. | 1
5088110 | THE CUSTOMER REPORTED THAT THE DEVICE ISSUED A SPEAKER MALFUNCT." INOP. THERE WAS NO ALLEGATION OF AN ADVERSE EVENT OR ANY PATIENT OR USER HARM.><.><THE CUSTOMER REPORTED THAT THE DEVICE ISSUED A "SPEAKER MALFUNCT." INOP. AT THE TIME OF THE ALLEGED MALFUNCTION | 1
5060232 | THE CUSTOMER REPORTED THAT THE DEVICE SHOWS A SPEAKER MALFUNCTION INOP AT THE CENTRAL AND BEDSIDE. NO PATIENT OR CUSTOMER HARM WAS REPORTED.><. | 1
5096835 | THE CUSTOMER REPORTED THAT THE BLUE CLIP (HOLD MMS) IS BROKEN OFF; THE MMS CAME OUT AND HIT AN EMPLOYEE". THE CUSTOMER REPORTED THE MMS LOCK "FRACTURED" AND DROPPED THE MMS ON A STAFF MEMBER | 0
5136655 | ><DOCTOR WAS PERFORMING SURGERY ON PATIENT AND THE ANESTHESIA MACHINE COMPLETELY QUIT WORKING. THE CRNA QUICKLY STARTED TO MANUALLY BAG THE PATIENT AND GIVE HIM SUPPLEMENTAL IV PROPOFOL TO KEEP THE PATIENT SEDATED FOR THE SURGERY. THE SCREEN THAT GIVES ALL THE READING ON HOW MUCH ANESTHESIA IS BEING ADMINISTERED WENT COMPLETELY BLANK WITH AND ERROR NUMBER. THE CRNA TURNED THE MACHINE POWER BUTTON OFF AND THEN TURNED IT BACK ON AND THE MACHINE BEGAN TO WORK AGAIN. CRNA WAS ABLE TO KEEP THE PATIENT SAFE BY BAGGING THE PATIENT AND GIVING SUPPLEMENTAL SEDATING MEDICATION TO KEEP THE PATIENT SEDATED UNTIL THE MACHINE DID POWER BACK UP AND CRNA WAS ABLE TO FINISH THE ANESTHESIA FOR THIS PATIENT. GE WAS IMMEDIATELY NOTIFIED AND THEY SENT A REPAIR TECH OUT THE NEXT DAY AND REPLACED THE POWER CONTROLLER BOARD AND POWER SUPPLY IN THE MACHINE. MANUFACTURER RESPONSE FOR ANESTHESIA MACHINE, AISYS CS2 (PER SITE REPORTER): THEY IMMEDIATELY DISPATCHED A FSE TO REPAIR THE PROBLEM. | 1
4903206 | ADDITIONAL INFORMATION RECEIVED FROM THE USER INDICATES THERE WAS A DELAY IN THE DISPLAY OF THE EKG ON THE RECORDING SYSTEM. THE DISPLAY INDICATED THE PATIENT WAS A PERSISTENT ARRHYTHMIA AND A SHOCK WAS DELIVERED; HOWEVER, REVIEW INDICATED THE PATIENT WAS IN SINUS RHYTHM PRIOR TO THE SHOCK. REBOOTING THE SYSTEM RESOLVED THE DISPLAY ISSUE.><(B)(4). THE DEVICE WAS NOT RETURNED FOR EVALUATION. EXAMINATION OF THE WORKMATE LOGS RETURNED CONFIRMED A DEFIBRILLATION SHOCK WAS PERFORMED AND A SECOND ONE APPROXIMATELY 30 SECONDS LATER. THE WORKMATE LOGS ALSO CONFIRM THAT THE SYSTEM WAS RESTARTED ON THE REPORTED DATE. REVIEW OF THE STUDY INDICATED THE PATIENT WAS IN NORMAL SINUS RHYTHM AT STUDY START. COMMENTS IN THE STUDY INDICATED THE INTRODUCTION OF ATRIAL FIBRILLATION FOLLOWED BY WHAT APPEAR AS ATTEMPTS TO PACE OUT OF THE INDUCED RHYTHM PRIOR TO THE CARDIOVERSION. BASED ON THE STUDY PROVIDED, IT APPEARS THAT ATRIAL FIBRILLATION WAS STILL PRESENT PRIOR TO THE FIRST CARDIOVERSION RETURNING TO NORMAL SINUS RHYTHM AFTER. FURTHER, THERE ARE INTERMITTENT PACS WITH OTHERWISE NORMAL SINUS RHYTHM PRIOR TO THE SECOND CARDIOVERSION WITH NORMAL SINUS RHYTHM AFTER. WHILE WE ARE UNABLE TO CONFIRM, IT IS POSSIBLE THAT THE REPLAY ACTION WAS INITIATED FROM THE REVIEW SCREEN. SELECTING THE REPLAY BUTTON WILL REPLAY THE CURRENTLY DISPLAYED REVIEW IMAGE ON THE REAL-TIME MONITOR AT THE CURRENT DISPLAYED SPEED. SIGNALS WILL CONTINUE TO DISPLAY UNTIL THE REPLAY BUTTON IS PRESSED AGAIN OR THE RECORDING ENDS. BASED ON THE INFORMATION PROVIDED TO ST. JUDE MEDICAL, THE CAUSE OF THE REPORTED DISPLAY DELAY REMAINS UNKNOWN.><(B)(4).><DURING AN ELECTROPHYSIOLOGY PROCEDURE, AN UNNECESSARY SHOCK WAS DELIVERED WHILE THE PATIENT WAS IN SINUS RHYTHM. THE PATIENT REMAINED IN SINUS RHYTHM AND THERE WERE NO ADVERSE CONSEQUENCES. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT YET RECEIVED. | 0
4820745 | DEVICE WORKED INTERMITTANTLY, A LOOSE WIRE ON THE END OF THE INSTRUMENT FENESTRATED BIPOLAR FORCEP COULD BE SEEN ON THE MONITOR.INSTRUMENT WAS REMOVED AND REPLACED WITH ANOTHER. NO HARM TO THE PATIENT. | 0
5207093 | .><A PHILIPS PATIENT MONITORING SURVEILLANCE STATION RESTARTED RESULTING IN AN OUTAGE OF SURVEILLANCE MONITORING. THIS OCCURRED DUE TO AN ATTEMPT TO MANUALLY RESIZE ONE OF THE SECTORS DISPLAYING A SINGLE PATIENT'S VITAL SIGNS AND PHYSIOLOGIC WAVEFORMS. THIS IS AN OFTEN-USED FEATURE THAT HAD NOT PREVIOUSLY RESULTED IN AN APPLICATION RESTART. THE OUTAGE LASTED FOR 5 MINUTES, A DURATION FOR WHICH NO CRITICAL ALARMS WERE AUDIBLE. THE ROOT CAUSE OF THIS ISSUE WAS FOUND TO BE A SOFTWARE GLITCH IN THE NEW SOFTWARE REVISION B.01. | 1
4784363 | DA VINCI GRASPING RETRACTOR CABLE SNAPPED DURING USE. INSTRUMENT WAS VISUALLY EXAMINED-NO INSTRUMENTATION WAS MISSING. REPLACED WITH NEW INSTRUMENT. | 0
5207094 | DURING AND CONTINUING AFTER INSTALLATION OF PHILIPS INTELLIVUE MX800 PATIENT MONITORS, IT WAS NOTICED THAT A NUMBER OF PATIENT MONITORS WERE CEASING TO SEND PHYSIOLOGIC DATA TO THE SURVEILLANCE STATION EVERY 2 MINUTES. THE DEVICES WOULD DROP OFF THE SECTOR, WAIT FOR A PERIOD OF TIME DURING WHICH CRITICAL ALARMS WOULD NOT BE AUDIBLE AT THE SURVEILLANCE STATION, AND THEN RETURN. THIS WAS FOUND TO BE CAUSED BY A SETTING IN THE PHILIPS INTELLIVUE INFORMATION CENTER IX (PIIC IX) PRIMARY SERVER IDENTIFYING AN ADMISSIONS DATABASE TARGET (ADT) IN THE INTELLIBRIDGE ENTERPRISE (IBE) SYSTEM WHEN THE ADT FEED TO THE IBE WAS NOT ACTIVE. THE ISSUE RESOLVED ONCE THE TARGET WAS TURNED OFF IN THE PRIMARY SERVER.><. | 1
5207092 | A PHILIPS PATIENT MONITORING SURVEILLANCE STATION RESTARTED RESULTING IN AN OUTAGE OF SURVEILLANCE MONITORING. THIS OCCURRED DUE TO AN ATTEMPT TO RETRIEVE A DISCHARGED PATIENT THAT HAD BEEN ADMITTED MULTIPLE TIMES, MANUALLY, WITHOUT AN MRN (MEDICAL RECORD NUMBER). THE OUTAGE LASTED FOR 5 MINUTES, A DURATION FOR WHICH NO CRITICAL ALARMS WERE AUDIBLE. THE ROOT CAUSE OF THIS ISSUE WAS FOUND TO RESULT FROM A PATIENT DATABASE MIGRATION FROM PIIC IX A.02 TO PIIC IX B.01, DURING WHICH TWO FIELDS IN THE PATIENT RECORD THAT ARE NEW TO PIIC IX B.01 WERE LEFT BLANK. RETRIEVING ONE OF THESE PATIENTS USING THE PATIENT (B)(6) SEARCH APPLICATION RESULTED IN A SOFTWARE EXCEPTION ON THE SURVEILLANCE STATION, CAUSING AN APPLICATION RESTART.><. | 1
5294474 | THIS COMPLAINT WAS OPENED AS A REQUIREMENT FOR THE RECALL OF THE LUMIFY SOFTWARE. PD27418 WAS COMPLETED AND CLOSED PREVIOUS TO THE COMPLAINT BEING OPENED, AND A RECALL IS UNDER WAY. FCO79500372 AND RECALL 2-0596-2016.><SOFTWARE DESIGN DEFECT. THE COLOR FLOW IS BACKWARDS IN ALL PRESETS WITH PRODUCT LUMIFY - 795216 1.0 RELEASE SW. THE KEY AND SCALE THAT ARE DISPLAYED ARE CORRECT, THE ACTUAL COLOR IS WRONG. NO CUSTOMER COMPLAINT NO INJURY REPORTED.><A RECENT SOFTWARE UPDATE TO THE TRACKWISE E-MDR HAS RESULTED IN THE ADDRESSES FOR THE MANUFACTURER'S CONTACT PERSON TO BE SOURCED FROM A NEW DATA LOCATION. THE ADDRESS ISSUE HAS BEEN CORRECTED. (B)(4).><FOUND INTERNALLY BY PHILIPS THE LUMIFY - 795216 COLOR FLOW IS BACKWARDS 1.0 RELEASE SW. | 1
5116376 | .><I-STAT IN DOCKING STATION, PICU STAFF SMELLED SMOKE, FOUND I-STAT TOO HOT TO TOUCH AND COVERED WITH BLACK SOOT. MANUFACTURER RESPONSE FOR I-STAT 1 WIRELESS, I-STAT 1 WIRLESS ANALYZER (PER SITE REPORTER): REPLACEMENT TO BE SENT OVERNIGHT. REPLACING 2 RECHARGING BATTERIES DOCKING STATION, AND I-STAT. | 0
5038568 | ><HAD IVC FILTER IMPLANT PLACED DUE TO SPONTANEOUS PRODUCTION OF DVT IN LEGS AND PULMONARY EMBOLUS IN LUNGS IN (B)(6) 2007. SINCE PLACEMENT OF DEVICE, PATIENT HAS HAD NO OBVIOUS INCIDENTS REQUIRING TREATMENT. HOWEVER, MR (B)(6) HAD RECEIVED NOTIFICATION VIA MAIL THAT HIS DEVICE IS IN THE RECALL GROUP, GREENFIELD FILTER CATALOG #52501, LOT# 9375737, REFERENCE# M001500110. ALSO, THE LAWYERS OFFICE ADVISED MR (B)(6) TO REPORT THE IMPLANT TO THE FDA FOR DOCUMENTATION. MR (B)(6) ALSO REQUIRES WARFARIN THERAPY AS WELL. | 0
4567462 | A PATIENT UNDERGOING ESOPHAGOGASTRODUODENOSCOPY (EGD) WITH THE NINEPOINT MEDICAL NVISIONVLE IMAGING SYSTEM AND WITH THE COVIDIEN BARRX RADIOFREQUENCY ABLATION (RFA) SYSTEM. IMAGING WAS COMPLETED WITH A 20 MM NVISIONVLE OPTICAL PROBE. THE BALLOON ON THE OPTICAL PROBE WAS DEFLATED AND REMOVED FROM ENDOSCOPE WORKING CHANNEL. THE BARRX RFA SYSTEM WAS THEN INTRODUCED AND REMOVED FROM THE ENDOSCOPE WORKING CHANNEL AFTER THE RFA WAS COMPLETED. WHILE THE ENDOSCOPE WAS STILL IN PT, THE SAME 20 MM OPTICAL PROBE WAS REINTRODUCED THROUGH THE WORKING CHANNEL FOR POST RFA IMAGING. THE OPTICAL PROBE'S BALLOON WAS RE-INFLATED AND THE PT WAS THEN RE-IMAGED WITH THE NVISION IMAGING SYSTEM. THE TEAM THEN ATTEMPTED TO DEFLATE THE BALLOON FOR REMOVAL FROM THE ENDOSCOPE WORKING CHANNEL, HOWEVER THE BALLOON WOULD NOT DEFLATE. MULTIPLE ATTEMPTS TO DEFLATE THE BALLOON WERE MADE WITH NO RESULTS. THE CONNECTIONS ON NVISION INFLATION ACCESSORY WERE CHECKED AND FOUND TO BE SECURE. THE TEAM ALSO ATTEMPTED TO DEFLATE THE BALLOON WITH A 35 ML LUER-LOCK SYRINGE. AT THIS POINT, THE PHYSICIAN FELT THAT THE BALLOON HAD DEFLATED ENOUGH, AND WOULD BE ABLE TO PULL IT THROUGH THE ENDOSCOPE WORKING CHANNEL. HOWEVER, WHEN THE OPTICAL PROBE WAS REMOVED COMPLETELY FROM ENDOSCOPE CHANNEL, THE BALLOON WAS NOT ATTACHED TO OPTICAL PROBE. THE ENDOSCOPE WAS REINTRODUCED INTO PT AND THE BALLOON WAS OBSERVED IN ESOPHAGUS. A SNARE WAS INTRODUCED THROUGH ENDOSCOPE WORKING CHANNEL AND USED TO REMOVE THE BALLOON. THE BALLOON AND ENDOSCOPE WERE REMOVED FROM PATIENT AT SAME TIME. NO ADVERSE EFFECTS ON PT WERE NOTED BY THE PHYSICIAN.><THE INVESTIGATION OF THE RETURNED OPTICAL PROBE SHOWED THAT THE BALLOON COMPLETELY DETACHED FROM THE OPTICAL PROBE'S SHEATH AND THAT IT WAS STILL PARTIALLY INFLATED. FOUR MAJOR KINKS WERE NOTED ALONG THE OPTICAL PROBE SHEATH. THE INNER IMAGING LUMEN OF THE OPTICAL PROBE WAS STRETCHED DISTALLY, WHILE THE PROXIMAL PORTION OF THE OPTICAL PROBE WAS FOUND TO BE INTACT. THE DETACHED BALLOON WAS PRESSURIZED AND IT WAS UNABLE TO BE INFLATED NOR DEFLATED WITH A TOUHY BORST ADAPTER AND SYRINGE. THE MOST PROBABLE CAUSE OF THE BALLOON SEPARATION FROM THE OPTICAL PROBE WAS THE ATTEMPT TO REMOVE IT THROUGH THE ENDOSCOPE WORKING CHANNEL WHILE IT WAS STILL PARTIALLY INFLATED. THE SHAFT OF THE OPTICAL PROBE WOULD MOVE FREELY IN THE ENDOSCOPE WORKING CHANNEL, BUT THE BALLOON WOULD BECOME STUCK AT THE WORKING CHANNEL OPENING DUE TO ITS PARTIALLY-INFLATED SIZE. THUS, CAUSING THE OPTICAL PROBE'S SHEATH TO BECOME STRETCHED AND THE BALLOON TO BECOME DETACHED FROM THE ASSEMBLY. | 0
5098917 | ><REPORTER DECLINED REQUEST TO RETURN DEVICE FOR EVALUATION, ALTHOUGH PROVIDED ONE PHOTOGRAPH. EVALUATION OF THE PHOTOGRAPH IS CONSISTENT WITH REPORTER'S STATEMENT THAT THE PATIENT BIT DOWN ON THE DEVICE. PRODUCTION HISTORY RECORDS FOR THE REPORTED LOT NUMBER WERE REVIEWED AND ALL IN-PROCESS CHECKS INDICATED PASSING RESULTS. THE PACKAGING LABEL CONTAINS PRECAUTIONARY STATEMENTS ADVISING THE END USER TO USE A BITE BLOCK WHEN PERFORMING ORAL CARE ON PATIENTS WITH ALTERED LEVELS OF CONSCIOUSNESS, OR THOSE WHO CANNOT COMPREHEND COMMANDS, AND ENSURING THE FOAM HEAD IS INTACT. BASED ON THE INVESTIGATION, AND EVALUATION OF THE PHOTOGRAPH, THERE IS INSUFFICIENT EVIDENCE TO CONCLUDE THAT THE REPORTED ISSUE WAS ATTRIBUTABLE TO A PRODUCT DEFECT OR MANUFACTURING OPERATIONS. FACILITY PROVIDED ONE PHOTOGRAPH. | 0
5056383 | ><DURING RIGHT HEART CATHETERIZATION IN CATH LAB ROOM ONE, WE WERE NOT ABLE TO MEASURE CARDIAC OUTPUT, NON-INVASIVE BLOOD PRESSURE (NIBP), OR PULSE OXIMETRY. | 1
5112576 | .><THERE ARE APPROXIMATELY 170 RUBBERMAID M40 MOBILE CARTS THAT ARE USED IN ALL AREAS OF OUR INSTITUTION. DURING THE LAST 8 DAYS, WE HAVE EXPERIENCE 3 FIRES ASSOCIATED WITH THESE DEVICES. ALL FIRES WERE EXPERIENCED IN THE CRITICAL CARE UNITS. UPON OPENING THE UNITS AFTER THE FIRE, IT APPEARS THAT SOME OF THE COMPONENTS ON THE CIRCUIT BOARD ARE GETTING REALLY HOT AND CATCHING ON FIRE. AT THIS TIME, WE ARE NOT SURE IF THE ACTUAL FIRE IS A RESULT OF A FAULTY ELECTRICAL COMPONENT SUCH AS A CAPACITOR/TRANSISTOR OR IF IT IS THE RESULT OF HEAT BUILD UP DUE TO EXCESSIVE LINT IN THE POWER SUPPLY AREA. ALL FAILED UNITS WERE MANUFACTURED IN 2011. | 1
4878661 | IN (B)(6) 2014, A PATIENT CAME IN TO HAVE A HOOKWIRE PLACED IN HER BREAST. MICRO CALCIFICATIONS WERE SEEN ON THE CRANIOCAUDAL (CC) REGULAR MAGNIFICATION VIEW BUT ON NO OTHER VIEWS. THE CUSTOMER USED THE SAME PARAMETERS AS USUAL AND APPROPRIATE COMPRESSION. THERE WAS NO VISIBLE LUMP, DISCOLORATION, OR OTHER SYMPTOMS. THREE MONTHS LATER, THE PATIENT WENT TO ANOTHER FACILITY WHERE A BIOPSY OF THE SAME AREA SHOWED MALIGNANT MICRO CALCIFICATIONS. SUBSEQUENTLY, THE PATIENT HAD A 2.5CM CLUSTER OF MALIGNANT MICRO-CALCIFICATIONS REMOVED THAT WAS ORIGINALLY DIAGNOSED AS BENIGN INSTEAD OF CANCEROUS. ACCORDING TO THE CUSTOMER, THEY DID MISDIAGNOSE THE PATIENT ORIGINALLY DUE TO IMAGING INCAPABILITY AND POOR VISUALIZATION OF MICRO-CALCIFICATIONS IN MAGNIFICATION MAMMOGRAPHY VIEWS.><THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE, THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. ADDITIONAL MANUFACTURER NARRATIVE: A FOLLOW-UP MDR WILL BE SUBMITTED SHOULD FURTHER INFORMATION BECOME AVAILABLE.><THE GE HEALTHCARE FIELD ENGINEER CONFIRMED THE SYSTEM WAS FUNCTIONING PROPERLY AND PASSED ALL IMAGE QUALITY TESTS. THE PRODUCT WAS DETERMINED TO BE WITHIN SPECIFICATIONS AND PERFORMING AS DESIGNED FOR ITS INTENDED USE. GE HEALTHCARE ENGINEERING ANALYZED THE PATIENT IMAGES AND CONFIRMED THAT THE MICRO CALCIFICATIONS ARE VISIBLE ON CRANIAL-CAUDAL (CC) PROJECTION, BUT NOT VISIBLE ON STANDARD MEDIO-LATERAL-OBLIQUE (MLO) AND THE 3 MAGNIFICATION PROJECTIONS. GOOD CLINICAL PRACTICES WOULD HAVE REQUIRED TO RE-IMAGE THE PATIENT IN MAGNIFICATION MODE AT THE ORIGINAL PROJECTION CC WHERE THE CALCIFICATION WAS INITIALLY LOCATED. RADIOLOGIST DID ACTUALLY PERFORM MAGNIFICATION VIEWS, BUT IN MLO, WHEREAS THE CALCIFICATION WAS ONLY VISIBLE IN CC. THE MAGNIFICATION VIEWS SHOWED A LOT OF PATIENT MOVEMENTS DURING ACQUISITIONS WHICH CAUSED TO THE DEGRADED IMAGE QUALITY. THESE IMAGES SHOULD HAVE BEEN RETAKEN BY THE DOCTOR PER STANDARD CLINICAL PRACTICES. RADIOLOGIST STATED SHE USED THE SUBOPTIMAL IMAGES AND MISREAD THEM RESULTING IN THE DIAGNOSTIC DELAY, HENCE USE" ERROR. THE MISDIAGNOSIS WAS DUE TO USER ERROR. THE RADIOLOGIST PERFORMED AN INCOMPLETE DIAGNOSTIC IN CC VIEW (MICRO CALCIFICATION SEEN BUT NOT DIAGNOSED). IN ADDITION | 0
4881521 | ><PLASTIC PARTS ON THE HEAD WALL HAVE FAILED, AS HAS A DOOR LATCH AND THE FOOT WALL PLASTIC FASTENER. | 0
5128338 | MEDTRONIC RECEIVED INFORMATION VIA LITERATURE REVIEW OF A STUDY PERFORMED TO COMPARE OUTCOMES AFTER TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) AND SURGICAL AORTIC VALVE REPLACEMENT (SAVR) IN PATIENTS WITH SEVERE SYMPTOMATIC AORTIC STENOSIS. THE STUDY POPULATION INCLUDED 405 PATIENTS (MEAN AGE 79.8 YEARS), 276 OF WHICH UNDERWENT TAVR FROM NOVEMBER 2007 THRU FEBRUARY 2011 AND 129 OF WHICH UN DERWENT SAVR FROM AUGUST 2003 THRU JUNE 2008. A PROPENSITY SCORE MATCHING BASED ON CLINICAL CHARACTERISTICS AND SURGICAL RISK SCORES WAS THEN UTILIZED TO COMPARE 222 PATIENTS (111 FROM EACH THE TAVR AND SAVR GROUPS). OF THE 111 TAVR PATIENTS, 41 RECEIVED A MEDTRONIC TRANSCATHETER BIOPROSTHETIC VALVE (SERIAL NUMBERS NOT REPORTED). OF THE 111 SAVR PATIENTS, 26 RECEIVED A MEDTRONIC SURGICAL BIOPROSTHETIC VALVE (SERIAL NUMBERS NOT REPORTED). AMONG ALL 222 MATCHED PATIENTS, 16 TOTAL CARDIOVASCULAR DEATHS (7 TAVR; 9 SAVR) OCCURRED WITHIN ONE YEAR FOLLOW-UP. NO DEATHS WERE DIRECTLY ATTRIBUTED TO MEDTRONIC PRODUCTS. AMONG ALL 222 MATCHED PATIENTS, PERI-PROCEDURAL AND IN-HOSPITAL ADVERSE EVENTS INCLUDED: 150 LIFE-THREATENING/MAJOR BLEEDING COMPLICATIONS (68 TAVR; 82 SAVR), 38 ACUTE KIDNEY INJURY (9 TAVR; 29 SAVR), 16 PERMANENT PACEMAKER (13 TAVR; 3 SAVR), 16 MAJOR VASCULAR COMPLICATIONS (TAVR ONLY), 4 REPEATED INTERVENTION FOR UNSPECIFIED VALVE-RELATED MALFUNCTION (3 TAVR; 1 SAVR), AND 2 SURGICAL INTERVENTIONS (SAVR ONLY). ADVERSE EVENTS UP TO ONE YEAR POST-IMPLANT INCLUDED: 15 STROKE/TRANSIENT ISCHEMIC ATTACK (5 TAVR; 10 SAVR), AND 2 MYOCARDIAL INFARCTION (SAVR ONLY). THE AUTHORS NOTED A SIGNIFICANTLY HIGHER INCIDENCE OF PERMANENT PACEMAKER REQUIRED FOR MEDTRONIC TRANSCATHETER BIOPROSTHETIC VALVES (10/41) AS COMPARED WITH NON-MEDTRONIC TRANSCATHETER BIOPROSTHETIC VALVES (3/70). APART FROM THIS STATEMENT, NO OTHER ADVERSE EVENTS WERE DIRECTLY ATTRIBUTED TO MEDTRONIC PRODUCTS. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION HAS BEEN REQUESTED.><REQUESTS FOR ADDITIONAL INFORMATION PROVIDED NO FURTHER DETAILS.><NO DEVICE WAS RETURNED AND NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED. WITHOUT UNIQUE DEVICE IDENTIFIER INFORMATION A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE PERFORMED AND ROOT CAUSE COULD NOT BE DETERMINED.><(B)(4). TITLE: TRANSCATHETER VS SURGICAL AORTIC VALVE REPLACEMENT IN INTERMEDIATE-SURGICAL-RISK PATIENTS WITH AORTIC STENOSIS: A PROPENSITY SCORE¿MATCHED CASE-CONTROL STUDY CITATION: AM HEART J. 2012 DEC;164(6):910-7 (DOI: 10.1016/J.AHJ.2012.09.005) AUTHORS: AZEEM LATIB, MB CHB, FRANCESCO MAISANO, MD, LETIZIA BERTOLDI, MD, ANDREA GIACOMINI, MD, JOANNE SHANNON, MD, MICAELA CIONI, MD, ALFONSO IELASI, MD, FILIPPO FIGINI, MD, KENSUKE TAGAKI, MD, ANNALISA FRANCO, MD, REMO DANIEL COVELLO, MD, ANTONIO GRIMALDI, MD, PIETRO SPAGNOLO, MD, GILL LOUISE BUCHANNAN, MD, MAURO CARLINO, MD, ALAIDE CHIEFFO, MD, MATTEO MONTORFANO, MD, OTTAVIO ALFIERI, MD, AND ANTONIO COLOMBO, MD DATE ACCEPTED BY PUBLISHER USED FOR EVENT DATE. | 0
5313881 | THE POSSIBLE AFFECT OF THE REPORTED CONDITION IS THAT THE PROBE COULD DRIP EITHER WASH SOLUTION A OR B INTO A REAGENT VIAL AND DILUTE THE REAGENT T RESULTING IN REDUCED REACTION IN A TEST. WASH SOLUTION A IS USED TO WASH AND RINSE THE INSIDE OF THE PROBE. WASH SOLUTION B IS USED TO WASH THE OUTSIDE OF THE PROBE. WASH SOLUTION B CONTAINS A DETERGENT WHICH COULD ALSO HAVE LYSING EFFECT DESTROYING THE RED CELLS IN A REAGENT PRODUCT. THE AFFECTED REAGENT WAS A PANEL CELL #1, BEING USED WITH MTS ANTI-IGG CARD FOR ANTIBODY SCREENING. USE OF A DILUTED RED CELL REAGENT COULD RESULT IN A WEAKENED REACTION WHICH COULD LEAD THE PROVUE TO REPORT A FALSE NEGATIVE REACTION GRADING. A FALSE NEGATIVE ANTIBODY SCREEN AMY LEAD TO THE INADVERTENT TRANSFUSION OF ANTIBODY-CONTAINING PLASMA THAT MAY LEAD TO A HEMOLYTIC TRANSFUSION REACTION IN A SUSCEPTIBLE PACTION. THE LIKELIHOOD OF SERIOUS INJURY IS NOT REMOTE. BASED ON THE AVAILABLE INFORMATION , MTS IS REPORTING THIS EVENT BASED ON THE POTENTIAL FOR INJURY SHOULD THE EVENT RECUR WITHOUT DETECTION BY THE USER. (B)(4)><CUSTOMER STATED THAT THEY NOTICED AN INCREASE IN THE VOLUME OF TH REAGENT CONTAINED IN THE PANEL CELL #1 VIAL AFTER BEING USED ON THE PROVUE, POSSIBLE CONTAMINATION FROM SYSTEM WAS FLUIDS. | 0
4743783 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS AND A REVIEW OF LABELING. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE, THE OPERATOR'S MANUAL STATES WHEN A RESULT IS FLAGGED WITH A PATIENT OR PANIC LIMITS ALERT, IT IS RECOMMENDED THAT THE LABORATORY'S REVIEW CRITERIA IS FOLLOWED, WHICH MAY INCLUDE REVIEW OF A STAINED SMEAR TO VERIFY THE RESULT AND TO CHECK FOR THE PRESENCE OF ANY ADDITIONAL ABNORMALITY. IT WAS DETERMINED THAT THE QC RESULTS WERE ALSO OUT OF RANGE AFTER THE ERRONEOUS RESULTS WERE NOTED BY THE CUSTOMER. QUALITY CONTROL SPECIMENS MUST BE RUN AND RESULTS CONFIRMED TO BE WITHIN ACCEPTABLE LIMITS BEFORE REPORTING PATIENT RESULTS. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED.><THE CUSTOMER OBSERVED FALSELY ELEVATED PLATELET RESULTS ON THE CELL-DYN EMERALD ANALYZER. THE FOLLOWING DATA WAS PROVIDED (K/UL): PATIENT 1: INITIAL PLATELET 983, REPEAT 265 THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
4587017 | (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER WAS EXPERIENCING DISCREPANT HEMOGLOBIN RESULTS GENERATED BY THE CELLDYN SAPPHIRE ANALYZER. A REVIEW OF THE COMPLAINT INFORMATION AND SUBMITTED DATA FOUND THAT SEVERAL OF THE PARAMETERS WERE MARKED AS SUSPECT OR INVALID. IN ADDITION, DATA FAULTS WERE NOTED FOR MOST OF THE SAMPLES. THESE FAULTS ALERT THE OPERATOR TO DATA CONDITIONS THAT REQUIRE FURTHER EXAMINATION. FIELD SERVICE WAS PERFORMED. UPON REPLACEMENT OF THE PISTON PUMP ASSEMBLY (WORN FROM NORMAL USE), THE ISSUE WAS RESOLVED. PREVENTIVE MAINTENANCE WAS ALSO PERFORMED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE CELLDYN SAPPHIRE OPERATION MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A DEFICIENCY.><THE CUSTOMER WAS EXPERIENCING DISCREPANT HEMOGLOBIN RESULTS GENERATED BY THE CELLDYN SAPPHIRE ANALYZER. PATIENT (B)(6) GENERATED HEMOGLOBIN RESULTS OF 8.26, 8.07, 6.78 AND 8.00 G/DL. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. | 0
5264064 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, REVIEW OF INSTRUMENT LOGS, A REVIEW OF LABELING AND INSTRUMENT SERVICE. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. FIELD SERVICE CONFIRMED THAT THE SYSTEM WAS IN IDLE/READY STATUS AT THE TIME OF TROUBLESHOOTING, NOT IN STANDBY; THEREFORE, THE INSTRUMENT WAS IN A PRESSURIZED STATE. THE CUSTOMER WAS ONLY WEARING A LAB COAT FOR PERSONAL PROTECTIVE EQUIPMENT. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION NO PRODUCT DEFICIENCY WAS IDENTIFIED.><THE CUSTOMER REPORTED A SPLASH TO HER LIP DURING TROUBLESHOOTING OF THE CD1800 CELL-DYN ANALYZER. THE CUSTOMER DISCONNECTED TUBING AND LIQUID SPLASHED ON HER LIP, NO MEDICAL TREATMENT WAS SOUGHT. THE AREA WAS CLEANED WITH SOAP AND WATER.><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). AN EVALUATION IS IN PROCESS. | 0
4751098 | THE CUSTOMER STATED THAT THE CELL-DYN EMERALD GENERATED FALSELY ELEVATED PLATELET RESULTS FROM FOUR PATIENT SAMPLES. ONE PATIENT SAMPLE GENERATED A RESULT OF 170 K/UL, WHICH WAS REPEATED AT 85 K/UL WHEN THE PATIENT SAMPLE WAS TESTED AT THE HOSPITAL LAB. HOWEVER, CHEMOTHERAPY WAS ADMINISTERED TO THE PATIENT BASED ON THE FALSELY ELEVATED RESULT. NO CONSEQUENCES OR IMPACT TO THE PATIENT OVER-ALL HEALTH WAS REPORTED.><(B)(4). PRODUCT EVALUATION IS IN PROCESS AND THE RESULTS WILL BE SUBMITTED IN A FOLLOW-UP REPORT.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER STATED THAT THE CELL-DYN EMERALD GENERATED FALSELY ELEVATED PLATELET RESULTS ON FOUR PATIENT SAMPLES. PATIENT #4 WAS DISCHARGED FROM THE HOSPITAL ON THE MORNING OF (B)(6) 2015 AND HAD BLOOD DRAWN AS PART OF THE DISCHARGE ORDER. THE PLATELET COUNT WAS 85 K/UL (NORMAL RANGE 150 - 450 K/UL). AFTER DISCHARGE FROM THE HOSPITAL, THE PATIENT WALKED ACROSS THE STREET TO COAST HEMATOLOGY ONCOLOGY FOR CHEMOTHERAPY TREATMENT. A SECOND BLOOD SAMPLE WAS COLLECTED AT COAST AS PART OF ROUTINE BLOOD WORK PRIOR TO ADMINISTERING CHEMOTHERAPY. THE PLATELET COUNT GENERATED BY THE CELL-DYN EMERALD WAS 170 K/UL (NORMAL RANGE 150 - 390 K/UL). CHEMOTHERAPY WAS ADMINISTERED BASED ON THE RESULTS GENERATED. NO CONSEQUENCES OR IMPACT TO THE PATIENT OVER-ALL HEALTH WAS REPORTED.><FIELD SERVICE WAS DISPATCHED AND REPLACED THE COUNTING CHAMBERS W/O APERTURE DUE TO INTERMITTENT ISSUES OBSERVED AT THE TIME OF SERVICE. IN ADDITION, THE RBC COUNTING HEAD WAS REPLACED DUE TO NORMAL WEAR. NO FURTHER PLATELET ISSUES WERE OBSERVED. THE SERVICE ACTIVITY PERFORMED IS CONSISTENT WITH ADDRESSING A HIGH PLATELET BACKGROUND OR RESULTS ISSUE. A PRODUCT TREND REVIEW AND ALL CUSTOMER COMPLAINTS RECEIVED FOR THIS ISSUE WERE REVIEWED. OUR REVIEW OF THIS DATA DID NOT IDENTIFY ANY ADVERSE TRENDS OR ABNORMAL COMPLAINT ACTIVITY. PER THE COMPLAINT TEXT INFORMATION, IT DOES NOT APPEAR THAT THE CUSTOMER WAS PERFORMING MONTHLY MAINTENANCE AS PRESCRIBED IN THE CELL-DYN EMERALD OPERATOR'S MANUAL. THE RESULTS INVOLVED IN THIS COMPLAINT ISSUE ARE FROM SAMPLES WHERE THE PATIENTS ARE ON SOME TYPE OF THERAPY, SUCH AS CHEMO, PLATELET, ETC. IF THE CELL-DYN EMERALD SYSTEM IS NOT PROPERLY MAINTAINED, THERAPEUTIC DRUGS CAN POTENTIALLY BUILD UP ON AND AROUND THE RBC/PLT PATH, COUNTING HEAD, ETC. THIS BUILD UP CAN POTENTIALLY AFFECT COUNTING DUE TO CONTAMINATION AND/OR SIZING DUE TO GROSS BUILD UP. FOR THIS REASON, THE CELL-DYN EMERALD HAS A MAINTENANCE SCHEDULE TO HELP ELIMINATE THIS BUILD UP. THE CUSTOMER WAS INSTRUCTED TO PERFORM SCHEDULED MAINTENANCE AS OUTLINED IN THE CELL-DYN EMERALD OPERATOR'S MANUAL TO ENSURE OPTIMUM PERFORMANCE. | 0
4796430 | THERE WERE NO RETURNS FROM THE CUSTOMER SITE FOR THIS EVALUATION. AN ABBOTT FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE AND FOUND THAT THE CELL-DYN EMERALD ANALYZER WAS POSITIONED ON A COUNTER TOP BETWEEN TWO CENTRIFUGES. THE FSE HAD TO STABILIZE THE ANALYSER BECAUSE IT WOBBLED" ON THE COUNTER TOP. THE FSE ALSO REPLACED THE INSTRUMENT'S COUNTING CHAMBER. IT WAS ALSO NOTED THAT THE CUSTOMER TESTED PATIENT SAMPLES ALMOST IMMEDIATELY AFTER THEY WERE DRAWN. NO OTHER INSTRUMENT ISSUES WERE NOTED. THE FSE PERFORMED MAINTENANCE PROCEDURES AND RETESTED THE SUSPECT SAMPLE | 0
4662370 | (B)(4). THE CELL-DYN 3700 OPERATOR'S MANUAL HAZARDS SECTION PROVIDES PRECAUTIONARY WARNINGS AND INFORMATION NECESSARY FOR SAFE USE OF THE CELL-DYN 3700 SYSTEM. TICKET TRACKING AND TRENDING WAS REVIEWED WITH NO ADVERSE TRENDS NOTED. BASED ON THE COMPLAINT INFORMATION, THE OPERATOR DID NOT WEAR PROTECTIVE EYEWEAR THROUGHOUT THE ENTIRE MAINTENANCE ACTIVITY; THEREFORE, THIS EVENT IS DETERMINED TO BE A USE ERROR. THE CELL-DYN 3700 ANALYZER IS PERFORMING ACCEPTABLY.><AN OPERATOR WAS PERFORMING MAINTENANCE AND REPLACING THE RBC METERING TUBE ASSEMBLY ON THE CELL-DYN 3700 WHEN CELL-DYN DILUENT SPLASHED INTO HIS EYE FROM TUBING THAT CAME LOOSE FROM A SOLENOID VALVE WHICH ASPIRATES CELL-DYN DILUENT. WHEN PERFORMING MAINTENANCE, THE OPERATOR WAS WEARING PROTECTIVE EYEWEAR AND GLOVES WHILE THE INSTRUMENT WAS OFF; HOWEVER, WHEN TURNING THE INSTRUMENT BACK ON TO VERIFY THE MENISCUS OF THE RBC METERING TUBE, THE OPERATOR HAD REMOVED HIS PROTECTIVE EYEWEAR. THE TUBING CONNECTED TO SOLENOID VALVE 4-3 CAME LOOSE, AND THE OPERATOR WAS SPLASHED IN THE EYE WITH DILUENT REAGENT. THE OPERATOR'S EYE WAS WASHED, CORTICOID EYE DROPS WERE GIVEN, AND A PATCH COVERED THE EYE. THE OPERATOR'S EYE RECOVERED. NO SPECIFIC OPERATOR INFORMATION WAS PROVIDED. | 0
5310404 | (B)(4). TEST RESULTS, INACCURATE. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS AND WHEN COMPLETE, A FINAL REPORT WILL BE SENT.><AN ABBOTT CUSTOMER TECHNICAL ADVOCATE WAS CONTACTED WITH REGARD TO THE CELL-DYN 4000 ANALYZER ASPIRATING AND IDENTIFYING A PT SAMPLE WHEN RUN IN AUTOLOADER MODE. THE ACCOUNT STATED THAT A SPECIMEN IN RACK NUMBER 13 POSITION 3 WAS CORRECTLY ASPIRATED, THEN A BARCODED TUBE IN RACK POSITION 4 WAS READ, BUT THE RACK MOVED BACK 2 SPACES, THEN FORWARD 1 SPACE, SO THE TUBE IN RACK POSITION 3 WAS ASPIRATED AGAIN, AND THE RESULTS ASSIGNED TO THE SPECIMEN ID FOR TUBE POSITION 4. THE ACCOUNT ALSO NOTED THAT A MIXHEAD FAILED TO DESCEND FAULT WAS GENERATED. WHEN THE ACCOUNT CHECKED THE MIXHEAD THEY FOUND THAT THE TUBE IN POSITION 1 OF THE MIXHEAD WAS ATTEMPTING TO BE DROPPED IN RACK POSITION 3 WHICH ALREADY HAD A TUBE IN IT. NO INCORRECT RESULTS WERE REPORTED FROM THE LAB. NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
5054111 | A DILUENT REAGENT SPRAY DOES NOT POSE HARM TO THE USER BECAUSE THE REAGENT IS A NACL SOLUTION (SALINE) AND WITH NO HARMFUL INGREDIENT. INVESTIGATION INCLUDED REVIEW OF PRODUCT HISTORICAL DATA AND PRODUCT LABELING. REVIEW OF PRODUCT HISTORICAL DATA DID NOT FIND A SIMILAR INCIDENT. THE CELL-DYN SAPPHIRE OPERATOR'S MANUAL PROVIDES ADEQUATE GUIDANCE, PROCEDURES, AND CAUTIONARY MESSAGES WITH WARNING SYMBOLS FOR SAFELY OPERATING, MAINTAINING, AND TROUBLESHOOTING THE CELL-DYN SAPPHIRE INSTRUMENT. THE MANUAL STATES THAT AN OPERATOR SHOULD WEAR PROPER PROTECTIVE EQUIPMENT, LIKE LABORATORY COAT, GLOVES AND PROTECTIVE EYEWEAR. PER COMPLAINT TEXT INFORMATION, THE OPERATOR WAS ONLY WEARING LABORATORY COAT. BASED ON THE INVESTIGATION, THE INCIDENT IS COVERED IN THE CELL-DYN SAPPHIRE PRODUCT LABELING AND WAS DEEMED AN ISOLATED INCIDENT INVOLVING ONE SPECIFIC CELL-DYN SAPPHIRE ANALYZER WITH A USER NOT WEARING PROPER PROTECTIVE EQUIPMENT, IN THIS CASE, GLOVES AND SAFETY EYEGLASSES. A PRODUCT DEFICIENCY WAS NOT IDENTIFIED.><WHILE TROUBLESHOOTING WBC DILUTION CUP ERRORS ON THE CELL-DYN SAPPHIRE ANALYZER, THE CUSTOMER WAS SPRAYED IN THE FACE, INTACT SKIN AND ON THEIR LABCOAT WITH DILUENT LEAKING FROM THE ANALYZER. THE CUSTOMER RINSED HER FACE USING THE EYE/FACE WASH STATION IN THE LABORATORY AND WAS MONITORED BY HER COWORKERS FOR ANY POTENTIAL REACTION. THE INGREDIENTS OF THE DILUENT WERE CHECKED AND NO HAZARDOUS CONTENTS WERE NOTED. THE CUSTOMER FILLED OUT A HOSPITAL SUPPLIED INCIDENT REPORT. NO MEDICAL TREATMENT WAS SOUGHT AS IT WAS NOT DEEMED NECESSARY BY THE THREE HOSPITAL EMPLOYEES PRESENT. NO SERIOUS INJURY WAS REPORTED.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). | 0
5207727 | A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SERVICE HISTORY REVIEW, ABBOTT FIELD SERVICE EVALUATION, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. A SERVICE HISTORY REVIEW FOR THIS INSTRUMENT REVEALED NO SUBSEQUENT COMPLAINTS OF DISCREPANT RESULTS BEING REPORTED. AN ABBOTT FIELD SERVICE ENGINEER (FSE) COMPLETED TROUBLESHOOTING OF THE INSTRUMENT. THIS INVOLVED PART REPLACEMENTS PER NORMAL TROUBLESHOOTING PROCEDURES, WHICH RESOLVED THE ISSUE. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY OF CELL-DYN EMERALD ANALYZER, LIST NUMBER 09H39 WAS IDENTIFIED.><THE CUSTOMER INDICATED THAT FALSELY ELEVATED PLATELET RESULTS WERE GENERATED WHILE USING THE CELL-DYN EMERALD ANALYZER. THE CUSTOMER PROVIDED PATIENT RESULT DATA ON (B)(6) 2015, AS FOLLOWS (REFERENCE RANGE (B)(4) FOR THE INITIAL RESULT, RETEST REFERENCE RANGE IS (B)(4)) SID RASP: INITIAL 42, RETEST 52, RETESTED AT ANOTHER LAB: 16. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
4625078 | THE CUSTOMER OBSERVED A FALSELY DEPRESSED HEMOGLOBIN RESULT FOR ONE PATIENT ON THE CELL-DYN CD1800 ANALYZER. THE INITIAL RESULT WAS 2.5 G/DL. THE PATIENT WAS SENT TO THE HOSPITAL AND REDRAWN, THE RESULTS WERE NORMAL (NO SPECIFIC DATA PROVIDED). TWO ADDITIONAL FALSELY DEPRESSED RESULTS WERE OBTAINED: 7.2 AND 7.4 G/DL. THERE WAS NO FURTHER IMPACT TO PATIENT MANAGEMENT REPORTED.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND A SITE VISIT. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. DURING A VISIT FROM FIELD SERVICE, THE 100 UL SYRINGE WAS FOUND TO BE WORN OUT FROM NORMAL USE AND WAS REPLACED. A LARGE CLOT WAS FOUND IN THE SAMPLE TUBE OF THE PATIENT SAMPLE WHICH GENERATED THE ABNORMALLY LOW HGB RESULT. THE CLOT WAS REMOVED, AND THEN THE SAMPLE WAS RUN TWICE IN THE CELL-DYN 1800 ANALYZER. THE ABNORMALLY LOW RESULTS WERE REPORTED TO THE PHYSICIAN. THE RESULTS FROM THE REDRAWN SAMPLES WERE NORMAL. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
4945814 | CLIENT WAS HAVING APNEA SPELLS. THE OXIMETER ALERTED CAREGIVER/NURSES THAT CLIENT WAS EXPERIENCING DIFFICULTIES. UPON ARRIVAL IN CLIENT'S ROOM, IT WAS NOTED THAT CLIENT WAS NOT BREATHING. CLIENT WAS/IS ON BIPAP. AGNES MONITOR WAS NOTED TO NOT BE ALARMING CLIENT'S CHEST WAS NOT RISING. DME COMPANY WAS NOTIFIED, MONITOR EXCHANGED FOR NEW MONITOR. DME COMPANY WAS TO DOWNLOAD INFO AND SEND TO PHYSICIAN. | 1
4465600 | ><UPON INVESTIGATION, IT WAS DETERMINED THAT WHEN THE ANESTHESIA CARE RECORD (ACR) APPLICATION IS INSTALLED ON A WORKSTATION RUNNING (B)(4), THE ACR APPLICATION WILL FAIL TO RETAIN CASE DATA IF THE USER HOVERS HIS/HER MOUSE OVER THE NOTES (COMMENTS) PANE, THE MESSAGES PANE, OR THE FLUID INTAKE/OUTPUT SCREEN AND A TOOL TIP IS DISPLAYED. AS A RESULT, DURING THIS PERIOD OF INACTIVITY, CASE DATA FROM THE PHYSIOLOGICAL MONITORS ARE NOT RECORDED AND CASE DATA UPDATES RECEIVED FROM OTHER SYSTEMS (VIA INTERFACES AND INTEGRATION) ARE NOT RETAINED ON THE WORKSTATION. WHEN THE USER¿S MOUSE IS RE-POSITIONED, THE APPLICATION WILL RETURN TO ITS ACTIVE STATE AND BEGIN TO RETAIN CASE DATA. THE REPORTING HOSP SITE WAS NOTIFIED OF THE RESULTS OF THE INVESTIGATION AND THE POTENTIAL OF THE ACR SCREEN TO BECOME UNRESPONSIVE AND DATA TO NOT RECORD. TYPICALLY THE ACR APPLICATION IS USED AS TOUCH-SCREEN APPLICATION WHERE THE MOUSE IS RARELY UTILIZED. THIS EVENT WILL NOT OCCUR ON A WORKSTATION WHEN THE USER IS USING A TOUCH-SCREEN MONITOR WITHOUT A MOUSE. THERE WAS NO INJURY ON THE ACTUAL PT AND/OR NO DELAY IN PT CARE. | 1
4399462 | ><ECG MONITOR ALARM NOT SOUNDING, AUDIBLE ALARM VOLUME SET AT 70%. | 1
4461100 | SAME CASE AS MDR 2134265-2015-00590. IT WAS REPORTED THAT UNINTENDED RF ENERGY WAS DELIVERED TO THE PATIENT. DURING AN ABLATION PROCEDURE, 10HZ NOISE WAS SEEN ON THE ABLATION CHANNEL WHEN NO RF ENERGY WAS APPLIED. SOMEHOW THERE WAS PACING CAPTURED IN THE ATRIUM FOR THE ABLATION CATHETER (10HZ), INDICATING THE ABLATION CATHETER WAS SENDING ENERGY INTO THE PATIENT. NO PATIENT COMPLICATIONS WERE REPORTED AND THE PATIENT IS IN GOOD CONDITION. NO ADDITIONAL TREATMENT WAS REQUIRED.><DEVICE EVALUATED BY MFR. THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED. (B)(4).><DEVICE EVALUATED BY MFR: THE RESULTS OF THE TESTING INDICATED THAT THE RETURNED CLEARSIGN AMPLIFIER AS WELL AS THE ASSOCIATED COMPLAINT CARTO CABLE PASSED INCOMING INSPECTION TESTS, INCLUDING NOISE AND LEAKAGE TESTING. THE POSSIBILITY OF INTERACTION EFFECTS BETWEEN THE TWO COMPONENTS AS A CAUSE WAS ELIMINATED BY TESTING THE RETURNED COMPLAINT AMPLIFIER AND CABLE TOGETHER AND FOUND NO ISSUES. BOTH DEVICES PERFORMED AS EXPECTED. THE MANUFACTURING BATCH RECORD REVIEW CONFIRMED THAT THE DEVICE MET ALL MATERIAL, ASSEMBLY AND PERFORMANCE SPECIFICATIONS. THE MOST PROBABLE CAUSE OF THE REPORTED DIFFICULTIES MAY BE DUE INTERACTION WITH ANOTHER DEVICE. (B)(4).><SAME CASE AS MDR 2134265-2015-00590. IT WAS REPORTED THAT UNINTENDED RF ENERGY WAS DELIVERED TO THE PATIENT. DURING AN ABLATION PROCEDURE, 10HZ NOISE WAS SEEN ON THE ABLATION CHANNEL WHEN NO RF ENERGY WAS APPLIED. SOMEHOW THERE WAS PACING CAPTURED IN THE ATRIUM FOR THE ABLATION CATHETER (10HZ), INDICATING THE ABLATION CATHETER WAS SENDING ENERGY INTO THE PATIENT. NO PATIENT COMPLICATIONS WERE REPORTED AND THE PATIENT IS IN GOOD CONDITION. NO ADDITIONAL TREATMENT WAS REQUIRED. | 0
4895124 | I WORK AT A COMMUNITY HOSPITAL IN XXXXXXX XXXX AND WE USE CERNER AS OUR COMPUTER SYSTEM. WHEN A PATIENT IS ADMITTED, THE PHARMACIST OR NURSE DOCUMENTS A MEDICATION HISTORY FOR THE PROVIDER TO REVIEW AND CONVERT TO AN INPATIENT ORDER. WE HAVE CHOSEN TO HAVE THIS LIST SHOW UP ON EVERY ENCOUNTER FOR THE MD/RPH/RN REVIEW WITH PATIENTS AND MODIFY IF NEEDED. WE ADMITTED AN (B)(6) YEAR OLD LADY WHO HAD SOTALOL 80MG PO BID ORIGINALLY ENTERED IN OUR HOME MED LIST ON (B)(6) 2014. SHE WAS READMITTED TO OUR HOSPITAL ON (B)(6) 2015 (2 MONTHS LATER) AND THIS MEDICATION WAS CONTINUED PER PATIENT VERIFICATION. THE PHYSICIAN CONVERTED THE ORDER TO AN INPATIENT ORDER. OUR NIGHT PHARMACIST WAS REVIEWING THE ORDER AND QUESTIONED THE DOSE. WE HAVE THE ABILITY TO REVIEW AN EXTERNAL PHARMACY HISTORY. MEDICATIONS FROM OUTPATIENT PHARMACIES POPULATE THIS FIELD WHICH ALLOWS US TO REVIEW MEDICATIONS, DIRECTIONS FOR USE AND REFILL HISTORY. THE PHARMACIST NOTICED THE DIRECTIONS AS: TK SS T PO BID. PHARMACISTS ARE VERY AWARE OF THE LATIN TERMINOLOGY AND WE KNOW THAT SS MEANS HALF. HOWEVER, NURSING STAFF AND NEWER PHYSICIANS ARE NOT AWARE OF SOME OF THESE LATIN TERMS AND THE DOSE WAS CONTINUED AT 80MG BID. (B)(6) DOES NOT RECOMMENDED WE USE SS AND THIS IS THE PERFECT REASON WHY. THE PHARMACIST DOUBLE CHECKED AND THE PATIENT WAS ONLY TAKING HALF OF AN 80MG TABLET TWICE DAILY AND THE ORDER WAS CHANGED. I HAVE CONTACTED CERNER ASKING THEM IF THEY COULD LOOK AT THEIR PROGRAM AND HAVE IT POPULATE THE ACTUAL PATIENT INSTRUCTIONS ON THE BOTTLE THAT IS GENERATED FROM WHAT THE PHARMACY STAFF ENTER IN THEIR SYSTEM TO GET THE DESIRED PATIENT LABEL. I HAVE NOT HEARD BACK FROM THEM YET. PATIENT'S AGE: OLDER ADULT (65-84 YEARS). (B)(6). | 1
5207084 | .><PHILIPS MX40 TELEMETRY DEVICE DROPPED OFF NETWORK: IT WAS REPORTED THAT AN INTELLIVUE MX40 WEARABLE PATIENT MONITOR CEASED TO TRANSMIT DATA TO THE SURVEILLANCE STATION AFTER BEING BROUGHT OUT OF STANDBY MODE. THE LACK OF DATA TRANSMISSION PERSISTED FOR HALF AN HOUR, WITH FOUR 1-SECOND LONG WAVEFORM SEGMENTS TRANSMITTED INTERMITTENTLY. THE BATTERIES WERE TAKEN OUT OF THE MX40, AND THEN RETURNED TO THE DEVICE, IN ORDER TO CYCLE POWER TO THE DEVICE. UPON RESTARTING, THE MX40 SUCCESSFULLY CONNECTED TO THE SURVEILLANCE STATION. THE MX40 WAS THEN REPLACED WITH ANOTHER DEVICE AS A PRECAUTIONARY MEASURE. ROOT CAUSE IS UNKNOWN. | 1
5096488 | PATIENT COMPLAINED OF RU OF ABDOMINAL BURNING (HOT). WE TELEMETRY MONITOR REMOVED BY PATIENT'S FAMILY. BOX HOT TO TOUCH - FAMILY CALLED NURSE INTO ROOM. THE BOX ON OVER BED TABLE. HOT TO TOUCH NOTED BY NURSE. BLISTER NOTED ON PATIENT'S ABDOMEN. | 0
5238597 | ><DURING MATERNAL AND FETAL MONITORING IN THE BIRTHCARE ROOM, WHEN THE RN PLUGGED IN THE MATERNAL PULSE OXIMETRY SPO2 CABLE, THE FETAL MONITOR SPEAKER EMITTED A MIDDLE FREQUENCY 60 CYCLE HUM" THE NOISE WAS PRONOUNCED ENOUGH TO INTERFERE WITH THE CLINICIAN'S ABILITY TO CLEARLY HEAR THE FETAL HEART SOUNDS FROM THE ULTRASOUND FETAL TRANSDUCER. WHEN THE CLINICIAN DISCONNECTED THE SPO2 CABLE FROM THE MONITOR | 1
5243483 | PT HAD FALSE ASYSTOLE AND PAUSE ALARMS THROUGHOUT THE WEEKEND. FRIDAY (B)(6) WAS THE WORST WITH 233, THEN SATURDAY (B)(6) 38 AND SUNDAY (B)(6) 47. GE WAS NOTIFIED AND ATTEMPTED TO TROUBLESHOOT THIS PACEMAKER PT'S FALSE ALARMS INCLUDING MULTIPLE PATCH CHANGES AND ALTERNATE LEAD PLACEMENT WHICH ONLY TEMPORARILY HELPED. GE NOTES THIS TO BE A PROBLEM WITH THE PT'S INTRINSIC QRS COMPLEX BEING TOO SQUATTY" BUT WE DID NOT HAVE THESE FALSE ALARM ISSUES WITH OUR PREVIOUS TELEMETRY MONITORING SYSTEM." | 1
4613521 | A HEMOCUE HB 201 SYSTEM HAS PROVIDED A RESULT THAT IS OUTSIDE SPEC WHEN COMPARED TO RESULTS PROVIDED BY LABORATORY METHODS. THIS COULD POTENTIALLY LEAD TO A MISSED CLINICAL ACTION IF IT WAS TO RECUR. HEMOCUE: 116 G/L; LAB METHOD 1: 70 G/L; LAB METHOD 2: 72 G/L; NO PT IMPACT REPORTED. THE USER DID NOT TAKE ANY ACTION BASED ON THE RESULT FROM THE HEMOCUE SYSTEM.><THE HEMOCUE HB 201 MICROCURVETTES RETURNED BY THE CUSTOMER FROM THE SAME LOT THAT WERE USED AT THE EVENT HAVE BEEN INVESTIGATED AT HEMOCUE. NO MALFUNCTION WAS DETECTED. NOTE THAT SIMILAR EVENTS AT THE SAME HOSP WITH THE SAME MICROCURVETTES HAVE BEEN REPORTED PREVIOUSLY: MDR NUMBER 300304483-2015-00002 AND MDR NUMBER 300304483-2015-00003. | 0
5207089 | .><IT WAS REPORTED THAT PHILIPS PATIENT MONITORING WAS LOST FOR AN ENTIRE CARDIAC TELEMETRY UNIT. PHILIPS MX40 DEVICES REMAINED OPERATIONAL AT THE BEDSIDE BUT WERE NOT VISIBLE AT THE SURVEILLANCE STATION, PREVENTING CRITICAL ALARMS FROM BEING AUDIBLE TO CAREGIVERS OUTSIDE OF PATIENT ROOMS. THE (B)(6) SWITCH TO WHICH THE WIRELESS ACCESS POINT (AP) UNITS WERE CONNECTED TO WAS FOUND TO HAVE BEEN MANUALLY RESET VIA A PUSH-BUTTON MECHANISM TO THE FACTORY DEFAULT CONFIGURATION, THUS DISCONNECTING THE WIRELESS BRIDGE BETWEEN THE TELEMETRY DEVICES AND THE SURVEILLANCE STATION. RESTORATION OF THE PHILIPS-INSTALLED CONFIGURATION DID NOT BRING THE APS BACK ONLINE. THE SYSTEM WAS REVERTED BACK TO THE PREVIOUSLY-INSTALLED GE PLATFORM. | 1
4494815 | ISSUE WITH EPIC SYSTEM IN WHICH DOSE IS CLEARED UPON ORDER VERIFICATION. THIS SPECIFICALLY OCCURS WHEN PHARMACISTS CHANGE THE PRODUCT THAT THE DOSE CLEARS. THIS IS A PROBLEM WITH THE LATEST VERSION OF EPIC AND HAS AFFECTED OTHER HOSPITALS AS WELL. EPIC KNOWS ABOUT THE ISSUE BUT THE FIX IS TAKING A LONG TIME. WE FIRST NOTIFIED THEM IN XXXXXXX AND THEY ALREADY KNEW ABOUT IT. EPIC SUGGESTED A TEMPORARY FIX BUT THAT CREATED A LOT MORE PROBLEMS SO WE COULDN'T IMPLEMENT IT. THERE HAVE BEEN SEVERAL CASES AT OUR HOSPITAL WHERE PHARMACISTS DIDN'T REALIZE THE DOSE DISAPPEARED AND VERIFIED THE ORDER. THE NURSE THEN RECEIVES AND ORDER WITHOUT A DOSE. ACCESS NUMBER:(B)(4). | 1
4637890 | RECENTLY, THE FACILITY RECEIVED A FIELD SAFETY NOTICE TWO DAYS AFTER THE DATE OF THE NOTICE REGARDING TWO ISSUES WITH THE MICROTEK MEDICAL SURGICAL DRAPES. THESE TWO ISSUES INCLUDE A CLOUDY/WAXY APPEARANCE TO THE DRAPE AND POTENTIAL TEARS TO THE SYSTEM DRAPES. THESE DRAPES ARE USED WITH THE DA VINCI STANDARD S, AND SI SURGICAL SYSTEMS. THE CONCERN IS THAT MANUFACTURER IS REQUIRING THE FACILITY TO PHYSICALLY OPEN EACH STERILE PACKAGE TO INSPECT THE DRAPES PRIOR TO USE. THE MANUFACTURER WILL NOT ACCEPT RETURN OF THE DRAPES UNLESS THE STERILE PACKAGE IS OPENED AND THE DRAPE IS INSPECTED. THIS SITUATION LEAVES PATIENTS POTENTIALLY AT RISK AS ALL OF CURRENT STOCK COULD BE IMPACTED AND THE ABILITY TO QUICKLY GET REPLACEMENT STOCK COULD BE IMPEDED. FDA HAS BEEN MADE AWARE OF THE IMPACTED LOT NUMBERS VIA CORRESPONDENCE PROVIDED BY THE MANUFACTURER. | 0
5025168 | ><THERE WAS NO PATIENT INVOLVEMENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. | 1
4784252 | DURING THE SURGICAL PROCEDURE, THE LARGE ROBOTIC NEEDLE DRIVER INSTRUMENT HAD THE GRASPING FORCEPS CABLE DETACH, CAUSING LOSS OF CONTROL IN OPENING AND CLOSING THE FORCEPS. THE INSTRUMENT WAS EXCHANGED WITH A NEW UNIT AND THE SURGICAL PROCEDURE WAS COMPLETED WITHOUT ADDITIONAL FAILURE. THE INSTRUMENT WAS RETURNED TO BIOMED FOR EVALUATION AND SHIPPING TO THE MANUFACTURER. THE MANUFACTURER PROVIDED AN EXCHANGE AND THE FAILED UNITS WERE RETURNED. | 0
5253442 | IT WAS REPORTED BY THE PHYSICIAN THAT A PILLCAM SB3 CAPSULE STUDY RAN SHORT. PHYSICIAN REMOVED THE CAPSULE ENDOSCOPICALLY INSTEAD OF ALLOWING THE CAPSULE TO PASS. THE PATIENT SWALLOWED A SECOND CAPSULE. THERE WAS NO REPORTED HARM TO THE PATIENT OR USER.><(B)(4). | 0
5018380 | PATIENT WITH MULTIPLE MYELOMA ON CHEMOTHERAPY. AT ROUTINE CLINIC VISIT NOTED TO BE WEAK AND SWEATY, SO TRANSFERRED TO ER FOR EVAL AND TREATMENT. IN ER FOUND TO BE HYPOTENSIVE, TACHYCARDIC, CONFUSED, FEBRILE, PRESUMABLY DUE TO SEPSIS / SEPTIC SHOCK. DURING RESUSCITATION, UNABLE TO ENTER ORDERS FOR FLUIDS, MEDICATIONS, LAB WORK, X-RAYS IN THE TACHYCARDIC BECAUSE CHART IS OPEN BY (B)(6) IN (B)(6) CLINIC WORKSTATION (B)(6)". CALLED CLINIC BUT CALL WENT TO VOICE MAIL. CALLED IT HELP DESK WHO SAID CALL CLINIC | 1
4799529 | QUESTION IF UPGRADE FOR RECALL UNITS FIXED PROBLEM IDENTIFIED IN RECALL. (B)(6) 2015 - DEVICE 1, DIGITAL MOTION CONTROL UPGRADE". (B)(6) 2015 - WHEN UNPLUGGED ALL LIGHTS ON LEFT BATTERY BAR GO OUT. ALSO BEEPING AND WILL NOT EXPOSE. THE PORTABLE WAS CHARGED FOR 7 HOURS UNINTERRUPTED PRIOR TO THIS. (B)(6) 2015 - PORTABLE FIXED. (B)(6) 2015 - CALLED SERVICE - BEEPING AND WILL NOT EXPOSE. SOURCE ONE/MERRY X-RAY SERVICE ENGINEER ARRIVED THAT AFTERNOON. HE SAID WE NEED NEW BATTERIES. (B)(6) 2015 - CHANGED BATTERIES | 1
4655621 | THE MARYLAND BIPOLAR FORCEPS BROKE; A WIRE SNAPPED AND THE DEVICE COULD NO LONGER BE UTILIZED SAFELY. | 0
5013703 | .><BLADE WAS NOT SEALING OR CAUTERIZING WELL, AND NO ALARM OR ERROR MESSAGE SOUNDED ON THE SCREEN. IT WAS REPLACED WITH A NEW DEVICE AND THERE WAS NO HARM TO THE PATIENT. | 0
4501387 | HARMONIC INSERT JAW LINING DISLODGED DURING DAVINCI LAP CHOLECYSTECTOMY. INSTRUMENT WAS REMOVED AND REPLACED WITH A NEW ONE. NO INJURY TO PATIENT. | 0
4620264 | DURING INTERNAL REVIEW, MCKESSON DETERMINED THAT THIS ISSUE IS SIMILAR TO 8022257-2013-00004, AND IS REPORTABLE UNDER PER 21 CFR 803. MCKESSON HAS IDENTIFIED AN ISSUE WHERE APPLICATION LOGIC LEADS TO DISCREPANCIES BETWEEN THE IMAGE INDEX AND THE CONTENT OF THE IMAGE FOLDER WHICH CAN CAUSE IMAGE LOSS. THE IMAGE FOLDER IS A FILE SYSTEM DIRECTORY WHERE IMAGES ARE STORED; IT ALSO CONTAINS THE IMAGE INDEX. THE IMAGE INDEX INDICATES TO THE APPLICATION A LIST OF IMAGE FILES TO EXPECT FOR USE WITHIN THE APPLICATION. IF THE IMAGE INDEX AND THE CONTENTS OF THE IMAGE FOLDER ARE NOT SYNCHRONIZED, THE FOLLOWING MAY OCCUR: IMAGE MAY BE LOST OR UNAVAILABLE, LEADING TO CLINICAL DECISIONS BEING MADE ON INCOMPLETE STUDIES AND/OR THE STUDY CANNOT BE ARCHIVED. IN THE EVENT OF A HARDWARE FAILURE, UNARCHIVED STUDIES MAY BE LOST LEADING TO PRIOR STUDIES NOT BEING AVAILABLE FOR COMPARISON. THE FOLLOWING FACTORS MAY CONTRIBUTE TO THE OCCURRENCE OF THE DISCREPANCY: NETWORK STABILITY ISSUES (FOR EXAMPLE: POOR PERFORMANCE, INSTABILITY AND COMMUNICATION ERRORS (DROPPED CONNECTIONS, PACKET LOSS)); DISK WRITE FAILURES (CANNOT WRITE TO DISK). THE PRODUCT MAY FAIL TO HANDLE THE OUTCOME OF SUCH FACTORS RESULTING IN THE DISCREPANCIES DESCRIBED ABOVE.><WHILE PERFORMING PROACTIVE MAINTENANCE OF UNARCHIVED STUDIES, MCKESSON IDENTIFIED 3 UNREPORTED STUDIES FROM THE SHORT TERM STORAGE IN THE HMI PACS THAT FAILED TO ARCHIVE TO LONG TERM STORAGE. NO PATIENT HARM WAS REPORTED. | 1
4553023 | THE DEVICE WOULD NOT RESPOND WHEN NEEDED. ANOTHER DEVICE WAS OBTAINED TO COMPLETE THE PROCEDURE. | 0
4485164 | .><THIS MEDIAL DEVICE REPORT IS BEING SUBMITTED DUE TO AN ADDITIONAL SIMILAR ISSUE THAT WAS REPORTED IN FOLLOW-UP MDR REPORTS FOR FDA # 9616389-2014-00029, 00031, 00032, 00033 AND 9616389-2015-00002. THIS REPORT IS FOR A 6TH EVENT FOR THE ELECTROSTATIC DISCHARGE ISSUE RESULTING IN UNINTENDED MOVEMENT WHEN USING A DX-D100 MOBILE UNIT. ON (B)(6) 2015, A SITE TECHNICIAN WAS DRIVING THE DX-D100 UNIT INTO AN ELEVATOR WHEN THE UNIT STOPPED. AS THE TECHNICIAN ATTEMPTED TO BACK OUT OF THE ELEVATOR, THE UNIT STARTED MOVING FORWARD QUICKLY AND SPINNING RIGHT IN A CIRCULAR MOTION AND ULTIMATELY CRASHING. THE TECHNICIAN WAS ABLE TO USE THE EMERGENCY SWITCH TO STOP THE SYSTEM. NO ONE WAS HARMED DUE TO THIS EVENT. THE TECHNICIAN MOVED THE DX-D100 UNIT BACK TO THE BIOMED DEPARTMENT AND TESTED THE SYSTEM MOVEMENT. THE BIO-MED DEPARTMENT WAS UNABLE TO RECREATE THE ISSUE AND THE DX-D100 SYSTEM WORKED FINE. THIS IS THE ONLY EVENT OF UNINTENDED MOVEMENT THE CUSTOMER HAS EXPERIENCED WITH THIS SPECIFIC DX-D100 UNIT. AGFA SERVICE CONFIRMED THE DX-D100 UNIT HAD THE REV H DIGITAL MOTION CONTROL BOARD AND THE CORRECT FIRMWARE INSTALLED. THE VOLTAGES WERE AT THE CORRECT SPECIFICATION AT 2.5 VDC. THERE WAS NO DAMAGE TO ANY CABLES AND THE CABLE CONNECTIONS APPEARED TO BE PLACED CORRECTLY. THE ENCODER MOTOR HAD THE CONNECTORS SEATED PROPERLY. AGFA SERVICE PLANS TO REPLACE THE DMC BOARD AND THE LEFT AND RIGHT GAUGES. THE CUSTOMER IS WILLING TO WAIT UNTIL THE PARTS ARRIVE BEFORE PUTTING THIS DX-D100 UNIT BACK INTO SERVICE. THE CUSTOMER HAS TWO OTHER DX-D100 SYSTEMS TO USE WHILE THIS UNIT IS OUT OF SERVICE. SEDECAL, AGFA'S SUPPLIER, HAS CONFIRMED ELECTROSTATIC DISCHARGE MAY OCCUR WHEN DRIVING OVER METAL THRESHOLDS OR WHILE ENTERING/LEAVING AN ELEVATOR. SEDECAL WILL BE IMPLEMENTING A CORRECTIVE ACTION RELATED TO A NEW ARRESTOR AND WILL COMMUNICATE THE CORRECTION TO AGFA. CORRECTIVE ACTIONS FOR THIS ISSUE WILL BE DOCUMENTED VIA PROBLEM RECORD, (B)(4). | 0
5270607 | AN EVENT OF UNINTENDED MOVEMENT OF ONE DX-D100 UNIT OCCURRED AT THE SITE ON (B)(6) 2015. THE CUSTOMER CONTACTED AGFA AND EXPLAINED WHEN USING THE DX-D100 UNIT IN THE INTENSIVE CARE UNIT OF THE HOSPITAL AND WHILE AN OPERATOR HAD THEIR HANDS GRIPPED ONTO THE UNIT, THE UNIT AUTOMATICALLY TOOK OFF AT FULL SPEED DOWN THE HALLWAY. A STUDENT TECHNICIAN THOUGHT QUICKLY AND HIT THE EMERGENCY BUTTON. NO ONE WAS INJURED DURING THIS EVENT. THE UNIT WAS REMOVED FROM SERVICE. AT THE ONSET OF THE EVENT, THE CUSTOMER'S BIOMED SERVICE COULD NOT REPRODUCE THE PROBLEM. AGFA AND OUR SUPPLIER, (B)(4), BEGAN FURTHER INVESTIGATION IMMEDIATELY. THE UNIT WAS THOROUGHLY CHECKED AND ALL REQUIRED UPGRADES, INCLUDING PROPER MOUNTING OF THE RESISTIVE ARRESTOR, HAD BEEN PERFORMED ON THIS UNIT AND THE UNIT HAD BEEN WORKING AS INTENDED, UNTIL THIS EVENT HAPPENED. THE MOVEMENT PROBLEM COULD NOT BE REPRODUCED. DURING THE INVESTIGATION, IT WAS DISCOVERED THE CUSTOMER HAD INSTALLED A METALLIC CHAIN ON THE BOTTOM OF THE MACHINE BETWEEN THE FRAME AND THE FLOOR WHICH BYPASSED THE CORRECTION AGFA HAD IMPLEMENTED TO AVOID UNCONTROLLED ESD. THE CUSTOMER THOUGHT BY ADDING THE CHAIN THEY COULD ALSO HELP AVOID ESD DISCHARGES. WITH THE ADDED METALLIC CHAIN, THE RESISTANCE BETWEEN THE FRAME AND GROUND WOULD BE 0 OHMS AND WITH THE RESISTIVE ARRESTOR PROPERLY INSTALLED WOULD BE 1M OHM. DUE TO THIS ADDED CHAIN, THERE WAS NOT ANY RESISTANCE BETWEEN THE FRAME OF THE EQUIPMENT AND ANY METALLIC PART PLACED ON THE FLOOR THAT COULD BE CONNECTED TO GROUND. IF THE EQUIPMENT IS CHARGED WITH HIGH VOLTAGE (ESD), DUE TO THE FRICTION BETWEEN THE WHEELS AND A NON-CONDUCTIVE FLOOR, THE DMC BOARD WOULD NOT BE PROTECTED AND COULD BE AFFECTED BY THE SPARK PRODUCED IF THIS VOLTAGE IS DISCHARGED TO THE GROUND THROUGH THE CHAIN. UNINTENDED MOVEMENT COULD OCCUR. EVEN THOUGH THE RESISTIVE ARRESTOR HAD BEEN PROPERLY INSTALLED BY AGFA, IT HAS BEEN DETERMINED THE EVENT WAS CAUSED BY THE CUSTOMER'S INSTALLATION OF THE METAL. THE PROBLEM WAS NOT RELATED TO THE DESIGN OF THE DX-D100 AND THIS EVENT IS AN ISOLATED CASE. THE METAL CHAIN WAS REMOVED FROM THE UNIT BY THE CUSTOMER'S BIOMED DEPARTMENT AND THE UNIT CHECKED BY AGFA. THE ISSUE COULD NOT BE REPRODUCED. THE UNIT WAS PLACED BACK INTO NORMAL USE AND RECOMMENDATION MADE TO NOT INSTALL ANY CHAINS TO THE AGFA EQUIPMENT. THERE HAS BEEN NO REPORTED HARM TO PATIENT OR USER DURING THIS EVENT.><. | 0
5123344 | PICIS CLINICAL SOLUTIONS RECEIVED AN ENHANCEMENT REQUEST ON (B)(6) 2015: A SENTENIAL EVENT OCCURRED RELATED TO PATIENT WEIGHT IN LBS BEING ENTERED IN THE KG BOX ON TRIAGE. I AM REQUESTING THAT WEIGHT ENTRY BE AN OPTION ON THE SITES PAGE ALLOWING THE ADMINISTRATOR TO SELECT WEIGHT ENTRY IN KG ONLY" THE PATIENT WEIGHT WAS DOCUMENTED INCORRECTLY IN THE ED PULSECHECK APPLICATION BY A CLINICIAN TRIAGING THE PATIENT. THE TRIAGE NURSE OBTAINED A VALUE FOR THE PATIENT'S WEIGHT MEASURED IN LBS AND INCORRECTLY DOCUMENTED IT INTO THE ED PULSECHECK FIELD LABELED KG. THIS RESULTED IN A DOCUMENTED WEIGHT MORE THAN TWICE THE PATIENT'S ACTUAL WEIGHT. PATIENT WEIGHT WAS DOCUMENTED AS 53.8 KG WHEN ACTUAL WEIGHT WAS (B)(6) KG. IT WAS FURTHER IDENTIFIED THAT A WEIGHT BASED MEDICATION (PHENOBARBITOL) DOSE WAS ADMINISTERED TO THE PATIENT BASED ON THE INCORRECT WEIGHT VALUE ENTERED. THE DOSE WAS CALCULATED BY THE ORDERING PROVIDER BASED UPON WEIGHT RECORDED WITHIN ED PULSECHECK. THE WEIGHT BASED MEDICATION ADMINISTERED TO PATIENT MAY HAVE CAUSED THE NEED FOR THE PATIENT TO BE TRANSFERRED TO ANOTHER FACILITY FOR FURTHER TREATMENT. THERE WAS NO CONFIRMATION FROM THE RECEIVING HOSPITAL PROVIDER THAT THE MEDICATION ADMINISTERED WAS THE CAUSE OR CONTRIBUTING FACTOR TO THE PATIENT'S CONDITION. THE ED PULSECHECK PRODUCT CURRENTLY HAS TWO TEMPLATE FIELDS FOR THE DOCUMENTATION OF A PATIENT'S WEIGHT | 1
4375994 | ACCORDING TO THE COMPLAINT, A (B)(6) YEAR OLD CHILD RECEIVED A SMALL ABRASION/SKIN REACTION OR BURN ON THE ARM AFTER USING THE TCM4 FOR A SLEEP STUDY. AFTER A WEEK IT WAS STILL NOT HEALED. MORE INFO, INCLUDING A PICTURE OF THE INJURY, HAS BEEN REQUESTED IN ORDER TO FURTHER SPECIFY THE TYPE OF INJURY.><THE MONITOR HAS BEEN RECEIVED AND TESTED BY RADIOMETER AND ALL TESTS WERE PASSED. THE FOLLOWING WARNING IS INCLUDED IN THE OPERATORS MANUAL P 6-2: WARNING-RISK OF BURNS: DO NOT ALLOW THE TCPCO2/TCPO2 SENSOR TEMPERATURE TO EXCEED 43 DEGREES CELSIUS FOR NEONATES AND 44 DEGREES CELSIUS FOR ADULTS IF SENSORS ARE ATTACHED TO SKIN FOR MORE THAN FOUR HOURS AS THIS MAY OTHERWISE CAUSE BURNS. THE TEMPERATURE OF THE SENSOR WAS 43 DEGREES CELSIUS. INFO REGARDING HOW LONG THE SENSOR WAS ATTACHED TO THE SKIN HAS BEEN REQUESTED TO DETERMINE IF AND IF THIS EVENT COULD BE CAUSED BY A USER ERROR. IT WAS MENTIONED IN THE COMPLAINT THAT THE PATIENT ROLLED ONTO THE SIDE WHERE THE SENSOR WAS ATTACHED. THIS COULD ENHANCE THE HEATING EFFECT OF THE SENSOR. NO INFO OF TREATMENT APPLIED HAS BEEN RECEIVED. THIS INJURY IS EVALUATED AS REPORTABLE | 0
4663595 | THESE INSTRUMENTS STOPPED FUNCTIONING WHILE IN USE; IN BOTH CASES THE GENERATOR DID NOT ACTIVATE. THERE WAS NO EFFECT ON TISSUE, AS WELL AS THERE WAS NO SOUND INDICATING THE GENERATOR WAS ACTIVATING. (NOTE: TWO DIFFERENT LOT NUMBERS). THIS INCIDENT DID NOT AFFECT THE TISSUE, AND THE SURGERY WAS ABLE TO BE COMPLETED WITH A NEW INSTRUMENT. THE PATIENT TOLERATED THE PROCEDURE WELL, AND WAS TAKEN TO THE RECOVERY ROOM IN STABLE CONDITION.WE HAVE A CD VIDEO OF THIS SURGERY; THE INSTRUMENT ISSUES ARE SEEN AT 8 MIN 50 SECONDS, AND AT 32 MINUTES. A COPY OF THIS CD WAS FEDEX-ED TO THE VENDOR, ALONG WITH THE INSTRUMENTS.WE HAVE A CD OF THIS SURGERY; THE INSTRUMENT ISSUES ARE SEEN AT 8 MIN 50 SECONDS, AND AT 32 MINUTES. A COPY OF THIS CD WAS FEDEXED TO THE VENDOR, ALONG WITH THE INSTRUMENTS. | 0
4376022 | ACCORDING TO THE COMPLAINT, A (B)(4) YEAR OLD CHILD RECEIVED A SMALL ABRASION/SKIN REACTION OR BURN ON THE ARM AFTER USING THE TCM4 FOR A SLEEP STUDY. AFTER A WEEK IT WAS STILL NOT HEALED. MORE INFO, INCLUDING A PICTURE OF THE INJURY, HAS BEEN REQUESTED IN ORDER TO FURTHER SPECIFY THE TYPE OF INJURY. | 0
5204443 | .><DURING A ROBOTIC ASSISTED BILATERAL INGUINAL HERNIA REPAIR THE PROGRASP ARM STOPPED WORKING. ONCE CLOSED, THE JAWS WOULD NOT REOPEN. NO HARM CAME TO THE PATIENT. THE DEFECTIVE ARM WAS REMOVED FROM THE FIELD AND REPLACED WITH A FUNCTIONING ARM. TRACKING NUMBER WAS ISSUED BY INTUITIVE SURGICAL. FAILURE ANALYSIS IS SENT TO OUR FACILITY BY INTUITIVE SURGICAL QUARTERLY. | 0
4806194 | AN EVALUATION OF THIS FAILURE MODE FROM THE DESIGN FAILURE MODES AND EFFECTS ANALYSIS OF THIS PRODUCT, THAT THE USE OF THE PRODUCT IN A MANNER LIKELY TO CAUSE ADVERSE HEALTH CONSEQUENCE IS IMPROBABLE. HEREWITH THE INVESTIGATION PLAN AND REPORT AS ATTACHMENTS. THE PRESCREEN RESULT WAS WITHIN ACCEPTABLE CRITERIA, THE ROOT CAUSE OF THIS CASE IS NOT RELATED TO THE FAILURE OF DEVICE PERFORMANCE, HOWEVER, WHETHER THERE IS OTHER FACTOR (SUCH AS, HUMAN FACTOR) NEED TO BE VERIFIED, WILL BASE ON CONTINUED MONITORING OF THE MARKET INFO.><(B)(4). | 0
4987806 | (B)(4) DURING A PULMONARY VEIN ISOLATION (PVI) PROCEDURE, THE PHYSICIAN ISOLATED ALL FOUR PULMONARY VEINS AND SINCE THE PATIENT WAS IN ATRIAL FLUTTER, A ROOF LINE WAS PERFORMED. THE PROCEDURE WAS FINISHED WITHOUT PATIENT CONSEQUENCES. THREE WEEKS POST-PROCEDURE THIS PATIENT PRESENTED TO THE HOSPITAL WITH EPILEPTIC SEIZURES AND FEVER; BLOOD CULTURES WERE POSITIVE. CEREBRAL CT SCAN SHOWED NO ABNORMALITIES. THORACO-ABDOMINAL CT-SCAN REVEALED COAGULATED BLOOD IN STOMACH. THE DECISION WAS TAKEN TO PERFORM A GASTROSCOPY TO IDENTIFY SOURCE OF BLEEDING. DURING GASTROSCOPY, BLOOD PRESSURE DROPPED AND PATIENT WENT INTO CARDIAC ARREST, DUE TO A HEMORRHAGIC STROKE. SUBSEQUENTLY THE PATIENT EXPIRED. THE PHYSICIAN¿S OPINION REGARDING THE CAUSALITY OF THIS EVENT WAS PROCEDURE RELATED. THIS EVENT WAS ORIGINALLY REPORTED UNDER SMART TOUCH CATHETER, REPORT# 9673241-2014-00158 ON JULY 8TH 2015, ADDITIONAL INFORMATION WAS RECEIVED. THIS CASE WAS PRESENTED DURING HEART RHYTHM SOCIETY SYMPOSIUM AND THE PHYSICIAN SHOWED HIS USE OF CARTO 3 SYSTEM AND VISITAG MODULE AND FORCE TIME INTEGRAL. THE PHYSICIAN¿S FORCE TIME INTEGRAL TARGET VALUES WERE 800G. HIS FORCE TIME INTEGRAL VALUES ARE BEYOND ST. JUDE MEDICAL RECOMMENDATIONS. ST. JUDE MEDICAL RECOMMENDS FORCE TIME INTEGRAL OF 400G. THE PHYSICIAN¿S OPINION REGARDING THE CAUSALITY OF THIS EVENT WAS PROCEDURE RELATED. PER CARTO 3 SYSTEM IFU: THE CARTO VISITAGT MODULE PROVIDES ACCESS TO DATA COLLECTED DURING THE APPLICATION OF RF ENERGY. THE DATA DOES NOT INDICATE THE EFFECTIVENESS OF RF ENERGY APPLICATION. LESION FORMATION DEPENDS ON MANY FACTORS THAT INCLUDE BUT ARE NOT LIMITED TO CATHETER STABILITY, POWER, TIME, AND APPLIED FORCE. THE CARTO VISITAG TM MODULE PROVIDES VISUAL DISPLAY FOR TRACKING THESE PARAMETERS. THE VISITAGT MODULE LOCATIONS CAN BE COLORED ACCORDING TO FTI | 0
4745092 | DURING PLACEMENT OF SINGLE SITE PORT FOR A SINGLE SITE ROBOTIC CHOLECYSTECTOMY, THE PORT (GEL) TORE, RENDING IT DEFECTIVE AND UNABLE TO USE. ALL OTHER PARTS INTACT. | 0
4905031 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. WHITE BLOOD CELL (WBC) COUNT IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND EVALUATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE REPORTED RESULTS COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
5021121 | INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKO REDUCTION FAILURE MAY BE DONOR RELATED. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOODCELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE REPORTED RESULTS COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4638508 | INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR)WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(6). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. | 0
4346296 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4426985 | INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIM A ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4839460 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND EVALUATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE DISPOSABLE SET COULD NOT BE EVALUATED AS IT WAS NOT RETURNED. RDF ANALYSIS DETAILED A NUMBER OF ALERTS AND ALARMS AS WELL AS EVIDENCE OF CLUMPING/AGGREGATES AT THE START OF THE PROCEDURE. THESE ALARMS, ALERTS AND CLUMPING COULD CONTRIBUTE TO THE HIGHER THAN EXPECTED RESIDUAL WBCS AS THEY COULD HAVE AFFECTED THE STABILITY OF THE INTERFACE. THE SIGNALS IN THE RUN DATA FILE ARE CONSISTENT WITH SOME POSSIBLE CLUMPING/AGGREGATION DURING THE BEGINNING OF PLATELET COLLECTION THAT COULD HAVE AFFECTED THE SEPARATION OF CELLS IN THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE CLUMPING/AGGREGATION AND RESULTING LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. WHILE THE TRIMA ACCEL SYSTEM IS TYPICALLY ROBUST AGAINST NUMEROUS ACCESS ALERTS AND ADJUSTMENTS, IT IS ALSO POSSIBLE THAT THE HIGH NUMBER OF ALERTS AND ADJUSTMENTS OCCURRING AT THE BEGINNING OF PLATELET COLLECTION DISRUPTED THE STEADY-STATE OF THE SYSTEM AND CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS PROCEDURE.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(6). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. | 0
4555582 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALIYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. | 0
4657091 | THE CUSTOMER REPORTED THAT THE DRAW SPEED WAS SLOW AFTER A PROCEDURE HAD STARTED. THE OPERATOR NOTICED A BLOOD CLOT IN THE NEEDLE AND THEY ENDED THE PROCEDURE. THE OPERATOR CONFIRMED THAT THE DONOR WAS 'OK'. PATIENT (DONOR) INFORMATION IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><THE CUSTOMER DECLINED TO PROVIDE PATIENT (DONOR) INFORMATION. PATIENT (DONOR) GENDER AND WEIGHT WERE OBTAINED FROM THE RUN DATA FILE.><THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDRS TO IDENTIFY RECORDS IN WHICH AN EVENT OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED.><INVESTIGATION: A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THIS OCCURRENCE CANNOT BE DETERMINED AT THIS TIME AS THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN. REVIEW OF THE RUN DATA FILE CONFIRMED THAT THE ¿CONNECT DONOR PROMPT¿ SCREEN WAS NEVER REACHED AND ¿START DRAW¿ BUTTON WAS NEVER PRESSED. BASED ON THE COMPLAINT DESCRIPTION AND ANALYSIS OF THE RDF, IT IS LIKELY THAT THE DONOR WAS CONNECTED EARLY WITH THE CLAMPS ON THE NEEDLE LINE CLOSED SO THE SYSTEM COULD NOT DETECT IT. THE NEEDLE CLAMP WAS THEN OPENED AND THERE WAS NO BLOOD FLOW DUE TO THE ELAPSED TIME OF UP TO 29 MINUTES WHEN THE PAUSE BUTTON WAS PRESSED.><INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
4973929 | INVESTIGATION: A SERVICE CALL WAS CONDUCTED FOR THE MACHINE INVOLVED WITH THIS INCIDENT. AN AUTOTEST TO CHECK THE CONFIGURATION AND OPERATIONAL FUNCTIONS WAS DONE. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT, THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THE WBC COUNT IS UNKNOWN AT THIS TIME.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. THE RUN WAS EVENT FREE, WITH NO ADJUSTMENTS OR CHANGES IN PUMP SPEED MADE DURING THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKO REDUCTION FAILURE MAY BE DONOR RELATED. ANALYSIS OF THE RBC DETECTOR SIGNAL SHOWED THAT PLATELETS TOOK LONGER THAN EXPECTED TO INITIALLY EXIT THE LRS CHAMBER. THIS MEANS THAT THERE WAS A DELAY BETWEEN WHEN THE TRIMA ACCEL DEVICE PREDICTED THAT PLATELETS WERE COLLECTING AND WHEN THEY ACTUALLY STARTED COLLECTING. THE CONCENTRATION OF PLATELETS GOING BY THE RBC DETECTOR WAS CONTINUALLY LOWER THAN WHAT THE TRIMA ACCEL DEVICE PREDICTED THROUGHOUT THE RUN. THIS LOW YIELD COLLECTION MAY HAVE BEEN DUE TO INACCURATE ENTRY OF THE DONOR INFORMATION, IMPROPER LOADING OF THE DISPOSABLE IN THE CENTRIFUGE OR AN INACCURATE YSF VALUE. ALTHOUGH NOT CONCLUSIVE, A DONOR RELATED PHENOMENON MAY HAVE CONTRIBUTED TO THIS LOW YIELD FAILURE. | 0
5061185 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(4). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND EVALUATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4791742 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. TERUMO BCT IS AWAITING RETURN OF THE DISPOSABLE SET FOR EVALUATION.><.><ADDITIONAL EVALUATION METHOD CODE: STERILIZATION PROCESS REVIEW INVESTIGATION: A TRIMA SET WAS RETURNED FOR INVESTIGATION. VISUAL INSPECTION CONFIRMED THE SET WAS ASSEMBLED CORRECTLY WITH NO KINKS, OCCLUSIONS OR MISSING CLAMPS. FLOW OF SOLUTION WAS OBSERVED THROUGH THE LRS CHAMBER. THERE WERE NO SIGNS OF DAMAGE TO THE CHANNEL AND THERE WERE NO WITNESS MARKS ON THE LOOP ASSEMBLY WHICH SUGGESTED THE SET WAS LOADED CORRECTLY. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE INVESTIGATION OF THE RETURNED SET DID NOT LEAD TO A CONCLUSIVE ROOT CAUSE. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. REVIEW OF THE RUN DATA FILE DID NOT SHOW A DEFINITIVE ROOT CAUSE FOR THE HIGHER THAN EXPECTED WBC COUNT OF THE PLATELET PRODUCT. SIGNALS SHOWED THAT PLATELETS EXITED THE CHAMBER 6 MINUTES AFTER THE VALVE WAS OPENED AND THE CONCENTRATION WAS CONSISTENT DURING THE COLLECTION. IT IS POSSIBLE THAT AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING THE COLLECTION CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. BASED ON THE AVAILABLE INFORMATION, IT IS ALSO POSSIBLE THAT THIS LEUKO REDUCTION FAILURE MAY BE DONOR RELATED. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
4871009 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THE WBC COUNT IS UNAVAILABLE AT THIS TIME.><INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. REPEATED ATTEMPTS TO GET THE RUN DATA FILE FROM THE CUSTOMER WERE UNSUCCESSFUL. ROOT CAUSE: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND EVALUATION. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE MACHINE WAS SUCCESSFULLY VERIFIED TO BE WORKING AT CHECKOUT.><.><INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE MACHINE WAS CHECKED OUT AT THE CUSTOMER SITE BY A TERUMO BCT SERVICE TECHNICIAN. ALL CALIBRATIONS WERE SUCCESSFULLY VERIFIED BY THE TECHNICIAN. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
4725252 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOODCELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4973727 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THE WBC COUNT IS UNKNOWN AT THIS TIME.><INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. HOWEVER, IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT WBCS MAY HAVE BEEN ESCAPING THE LRS CHAMBER ALONG WITH THE PLATELETS DURING THE COLLECTION. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE MAY BE DONOR RELATED.><INVESTIGATION: A SERVICE CALL WAS CONDUCTED FOR THE MACHINE INVOLVED WITH THIS INCIDENT. AN AUTOTEST TO CHECK THE CONFIGURATION AND OPERATIONAL FUNCTIONS WAS DONE. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
4871750 | THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDRS TO IDENTIFY RECORDS IN WHICH AN EVENT OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED.><INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER DECLINED TO GIVE THE PATIENT INFORMATION. THE OPERATOR CONFIRMED THAT THE DONOR WAS OK.><INVESTIGATION: THE DISPOSABLE SET WAS NOT AVAILABLE FOR RETURN FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE ROOT CAUSE FOR THIS PARTICULAR ISSUE IS UNKNOWN AT THIS TIME. POSSIBLE CAUSES INCLUDE BUT ARE NOT LIMITED TO:- DONOR ACCESS PROBLEM- KINKED OR CLAMPED DRAW LINE- DEFECTIVE OR MIS-CALIBRATED ACCESS PRESSURE SENSOR- DECOUPLED ACCESS PRESSURE SENSOR DIAPHRAGM ON TUBING SET><THE CUSTOMER REPORTED THAT THEY RECEIVED A 'DRAW PRESSURE LOW' ALARM APPROXIMATELY 5 MINUTES INTO A COLLECTION PROCEDURE. WHILE THE OPERATOR MADE ADJUSTMENTS TO THE NEEDLE POSITION, THEY NOTICED A CLOT IN THE NEEDLE. THE MACHINE ALARMED AND 'END RUN' WAS THE ONLY OPTION. PATIENT (DONOR) INFORMATION IS UNAVAILABLE AT THIS TIME. PATIENT'S (DONOR) OUTCOME IS UNAVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. | 0
4683869 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(4). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE REPORTED RESULTS COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION,OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4403919 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4471458 | INVESTIGATION: THE DISPOSABLE SETS WERE UNAVAILABLE FOR RETURN AND EVALUATION. ALSO, THERE WERE MULTIPLE UNSUCCESSFUL ATTEMPTS TO OBTAIN THE RUN DATA FILES (RDF) FOR EVALUATION OF THESE INCIDENTS. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. THE PLASMA LINE MAY HAVE BEEN OCCLUDED DURING A PORTION OF THE PROCEDURE. THE PLASMA LINE MAY BECOME OCCLUDED IF THE CENTRIFUGE COLLAR IS NOT LOADED INTO THE LATCH IN THE CORRECT ORIENTATION. UNDER THESE CONDITIONS, THE RBC LINE MAY LIE ON TOP OF THE PLASMA LINE AND, UNDER CERTAIN FLOW CONDITIONS, MAY CAUSE THE PLASMA LINE TO PINCH OFF. THIS IN TURN CAUSES HIGHER FLOW THROUGH THE LRS CHAMBER, WHICH MAY LEAD TO ELEVATED LEVELS OF WBCS IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. WBC COUNT IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><THE CUSTOMER DID NOT PROVIDE THE WBC COUNT.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
5140380 | INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE REPORTED RESULTS COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOODCELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. (B)(6). THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. | 0
4956587 | INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE GREATER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(4). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
4748735 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(4). THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE REPORTED RESULTS COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING,CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4804271 | AFTER HEMATOPOIETIC PROGENTIOR CELL (HPC) PRODUCT WAS COLLECTED AND DONOR WAS DISCONNECTED FROM AMICUS, I STRIPPED THE PRODUCT TUBING WITH MANUAL STRIPPERS TO PUT ALL CELLS INTO BAG. UPON DOING THIS, THE INJECTION PORT'S RUBBER TYPE TOP BECAME DISCONNECTED AND SOME DROPS OF HPC PRODUCT WERE NOTED TO HAVE LEAKED OUT BEFORE ROLLER CLAMP BELOW SITE WAS CLAMPED. I DO NOT BELIEVE THE ROLLER CLAMP WAS CLOSED, BUT POTENTIALLY THE TUBING MAY HAVE BEEN KINKED OR OCCLUDED, WHICH CAUSED THE OCCURRENCE. | 0
5085913 | AN MDR FOLLOW-UP REPORT WILL BE SUBMITTED AFTER AN ELEVATED COMPLAINT INVESTIGATION CONCLUDES.><THE ABBOTT VP 2000 PROCESSOR IS A DEVICE DESIGNED TO AUTOMATE AND STANDARDIZE SLIDE SPECIMEN PROCESSING AND ROUTINE SLIDE STAINING FOR THE LABORATORY. THE CUSTOMER REPORTED A SLOW SMALL LEAK FROM A DRAIN FITTING OF THE WATER BASIN. THE FSE REPLACED THE RINSE TANK ASSEMBLY WHICH RESOLVED THE LEAKAGE ISSUE. NO INJURY WAS REPORTED. THIS MEDICAL DEVICE REPORT (MDR) IS SUBMITTED ON THE BASIS FOR POTENTIAL HARM SHOULD THE MALFUNCTION RECUR. WATER LEAKING ON THE FLOOR HAS THE POTENTIAL TO CREATE A SLIPPERY FLOOR SURFACE WITH AN ASSOCIATED HARM OF SLIPPING OR FALLING. SLIPPING AND/OR FALLING HAS THE POTENTIAL TO CAUSE SERIOUS INJURY OR DEATH .><INVESTIGATION INTO THIS COMPLAINT FOR MDR 3005248192-2015-00019 FOLLOW-UP REPORT 1 INCLUDED AN EXISTING DATA REVIEW, QUALITY DATA REVIEW, A PRODUCT/SYSTEM/INSTRUMENT EVALUATION, AND A COMPLAINT HISTORY REVIEW. EXISTING DATA REVIEW: THE ABBOTT VP 2000 PROCESSOR SERVICE MANUAL (30-608308/R1 (B)(4) 2008) CONTAINS INSTRUCTIONS FOR REPLACEMENT OF THE RINSE TANK ASSEMBLY. QUALITY DATA REVIEW (DEVICE HISTORY REVIEW): CAPA / NON-CONFORMANCE REVIEW: A NONCONFORMANCE AND CAPA INVESTIGATION SEARCH WAS PERFORMED FOR THE RINSE TANK ASSEMBLY. THERE WERE 0 RELATED NONCONFORMANCES OR CAPA INVESTIGATION IDENTIFIED PER THE SEARCH RESULTS. PRODUCT/SYSTEM/INSTRUMENT EVALUATION: SERVICE HISTORY REVIEW: IT WAS IDENTIFIED THAT THIS OCCURRENCE IS THE FIRST REPLACEMENT OF THE RINSE TANK ASSEMBLY FOR THE SUBJECT INSTRUMENT. COMPLAINT HISTORY REVIEW: A COMPLAINT SEARCH IDENTIFIED EIGHT ADDITIONAL COMPLAINT RELATED TO LEAKING OF THE RINSE TANK ASSEMBLY AT THE DRAIN FITTING. THREE OF THE COMPLAINTS WERE INVESTIGATED AND CONFIRMED AND THREE WERE ADDRESSED PER RISK ASSESSMENT. TWO COMPLAINTS WERE CLOSED TO TROUBLESHOOTING. PRODUCT DEFICIENCY DECISION BASED ON THE RESULTS OF THE INVESTIGATION ELEMENTS A PRODUCT DEFICIENCY FOR RINSE TANK ASSEMBLY ON THE VP2000 INSTRUMENT WAS NOT IDENTIFIED. THEREFORE, THIS COMPLAINT WILL BE DISPOSITIONED AS UNCONFIRMED. | 0
5051098 | THE DEVICE WAS NOT RETURNED FOR AN EVALUATION. HOWEVER, THE CUSTOMER WAS ABLE TO PROVIDE COPIES OF THE MIXCHECK REPORT, WHICH PROVIDED DETAILS ABOUT THE COMPOUNDING PROCESS FOR THE ORDERS. THE MIXCHECK REPORT SHOWED THE BAGS WERE WITHIN RANGE AND THE FINAL INGREDIENTS IN THIS EVENT WERE WITHIN ACCEPTABLE LIMIT. THERE WAS NO INDICATION THE COMPOUNDER CONTRIBUTED TO THIS EVENT. THIS EVENT IS LOGGED UNDER COMPLAINT FILE NUMBER (B)(4).><THE CUSTOMER HAD CALLED TO REPORT BAGS PRODUCED ON A EXACTAMIX COMPOUNDER COMPLETED DELIVERY BEFORE THE SCHEDULED COMPLETION DURING PATIENT INFUSION. THE REPORTER WAS UNSURE OF THE CAUSE OF THE EVENT SINCE THE BAGS PASSED INSPECTION PRIOR TO RELEASE AND THE FACILITY ADDS OVERFILL TO THE BAGS IN ORDER TO COMPENSATE FOR PRIMING. THERE WAS NO REPORTED PATIENT INJURY OR MEDICAL INTERVENTION. THIS IS REPORT 1 OF 3 FOR THIS EVENT/PATIENTS. NO ADDITIONAL INFORMATION IS AVAILABLE. | 0
4642782 | AS THE REPORTED EVENT WAS A SOFTWARE ISSUE AND RESOLVED BY OVER THE PHONE BY BAXTER TECHNICAL SERVICES, NO PRODUCT WAS RETURNED FOR EVALUATION. A CAPA HAS BEEN OPENED TO ADDRESS THE REPORTED ISSUE. SHOULD ADDITIONAL, RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PRODUCT NOT RETURNED FOR EVALUATION.><IT WAS REPORTED THAT WHEN A TPN BAG LABEL WAS SCANNED, THE WRONG ORDER WAS DISPLAYED ON THE BAG LABEL. THIS ISSUE WAS RESOLVED BY BAXTER TECHNICAL SERVICES BY HAVING THE CUSTOMER MANUALLY INCREMENT THE ORDER NUMBER BEYOND THE LAST ORDER SAVED IN THE ABACUS DATABASE. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION IS AVAILABLE. | 1
4642765 | IT WAS REPORTED THAT AFTER RESTORING THE ABACUS SOFTWARE DATABASE, INCORRECT COMPOUNDING ORDERS WERE PULLED WHEN SCANNING THE ABACUS CREATED LABELS. THIS ISSUE WAS RESOLVED BY BAXTER TECHNICAL SERVICES BY HAVING THE CUSTOMER MANUALLY INCREMENT THE ORDER NUMBER BEYOND THE LAST ORDER SAVED IN THE ABACUS DATABASE. THERE WAS NO PATIENT INVOLVEMENT. NO ADDITIONAL INFORMATION IS AVAILABLE.><AS THE REPORTED EVENT WAS A SOFTWARE ISSUE AND RESOLVED BY OVER THE PHONE BY BAXTER TECHNICAL SERVICES, NO PRODUCT WAS RETURNED FOR EVALUATION. THERE WERE NO NON-CONFORMANCES ASSOCIATED WITH THE PRODUCTION OF THIS LOT NUMBER. A CAPA HAS BEEN OPENED TO ADDRESS THE REPORTED ISSUE. SHOULD ADDITIONAL, RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. PRODUCT NOT RETURNED FOR EVALUATION. | 1
4594276 | THE EVALUATION OF THE DATA PROVIDED FROM THE TWO EM1200 COMPOUNDERS AT CLINICAL SPECIALTIES SHOWED NO EVIDENCE THAT THE EM1200 COMPOUNDER CONTRIBUTED TO THE BAGS PREMATURELY RUNNING DRY DURING INFUSION. THE CURLIN PUMP INSTRUCTIONS PROVIDED TO PATIENT B: INFUSE 2200ML IV ONCE A DAY OVER 12 HOURS WITH A 1 HOUR TAPER UP AND 1 HOUR TAPER DOWN VIA CURLIN PUMP. THE CURLIN INFUSION PUMPS USED TO INFUSE THE BAGS IN THIS EVENT WILL BE REPORTED AND REVIEWED BY THE MANUFACTURER. BAXTER HEALTHCARE IS NOT THE MANUFACTURER OF THE CURLIN INFUSION PUMPS. EVALUATION CODES METHOD: DATA VALIDATION ORDER ENTRY LABELS, COMPOUNDER MIXCHECK REPORTS AND THE BLACK-BOX DATA WERE REVIEWED AND EVALUATED. RESULTS:USER DEVICE INTERFACE IT IS UNKNOWN HOW OR WHAT CAUSED THE NUTRITION BAGS TO RUN DRY. CONCLUSION: NO FAILURE DETECTED, DEVICE OPERATED WITHIN SPECIFICATION. THE INFORMATION FROM THE DATA PROVIDED FOUND THE EM1200 COMPOUNDER DID NOT CONTRIBUTE TO THIS EVENT. DEVICE NOT AVAILABLE FOR RETURN><ON (B)(6) 2015, THE SENIOR PHARMACY TECHNICIAN AT CLINICAL SPECIALTIES HAD CONTACTED BAXTER TECHNICAL SERVICES IN (B)(6), TO REPORT A TOTAL OF 6 BAGS THAT HAD COMPLETED DELIVERY BEFORE THE SCHEDULED COMPLETION. THE BAGS WERE PUMPED AND INFUSED TO 6 PATIENTS. ALL PATIENTS ARE HOME CARE PATIENTS. THE PHARMACY HAD CREATED A BATCH OF TOTAL PARENTERAL BAGS FOR EACH PATIENT TO BRING HOME FOR THE WEEK. IT WAS REPORTED BY THE NURSES THAT THE INFUSION PUMPS USED TO INFUSE THE BAGS STARTED ALARMING. WHEN THE NURSES CHECKED WHY THE INFUSION PUMPS WERE ALARMING, THE NURSES NOTICED THE BAGS HAD RUN DRY. IN REGARDS TO PATIENT STATUS, THE SENIOR PHARMACY TECHNICIAN INDICATED THE 6 PATIENTS ARE STABLE AND DOING FINE. THIS REPORT IS FOR THE (B)(6) MALE, PATIENT B; 2 OF 6 REPORTS FROM THE SAME EVENT AS IT INVOLVED 6 DIFFERENT PATIENTS. | 1
4594391 | ON (B)(6) 2015, THE SENIOR PHARMACY TECHNICIAN AT (B)(6) HAD CONTACTED BAXTER TECHNICAL SERVICES IN (B)(4) TO REPORT A TOTAL OF (B)(4) BAGS THAT HAD COMPLETED DELIVERY BEFORE THE SCHEDULED COMPLETION. THE BAGS WERE PUMPED AND INFUSED TO (B)(6) PATIENTS. ALL PATIENTS ARE HOME CARE PATIENTS. THE PHARMACY HAD CREATED A BATCH OF TOTAL PARENTERAL BAGS FOR EACH PATIENT TO BRING HOME FOR THE WEEK. IT WAS REPORTED BY THE NURSES THAT THE INFUSION PUMPS USED TO INFUSE THE BAGS STARTED ALARMING. WHEN THE NURSES CHECKED WHY THE INFUSION PUMPS WERE ALARMING, THE NURSES NOTICED THE BAGS HAD RUN DRY. IN REGARDS TO PATIENT STATUS, THE SENIOR PHARMACY TECHNICIAN INDICATED THAT THE (B)(6) PATIENTS ARE STABLE AND DOING FINE. THIS REPORT IS FOR THE (B)(6) FEMALE, PATIENT F;(B)(4) REPORTS FROM THE SAME EVENT AS IT INVOLVED (B)(6) DIFFERENT PATIENTS.><THE EVALUATION OF THE DATA PROVIDED FROM THE TWO EM1200 COMPOUNDERS AT CLINICAL SPECIALTIES SHOWED NO EVIDENCE THE EM1200 COMPOUNDER CONTRIBUTED TO THE BAGS PREMATURELY RUNNING DRY DURING INFUSION. CURLIN INFUSION PUMP INSTRUCTIONS PROVIDED TO PATIENT (B)(4): INFUSE AT 2500ML IV FOUR DAYS PER WEEK (SUN, TUES, THURS, SAT) OVER 10 HOURS WITH A 1 HOUR TAPER UP AND 1 HOUR TAPER DOWN VIA CURLIN PUMP. THE CURLIN INFUSION PUMPS USED TO INFUSE THE BAGS IN THIS EVENT WILL BE REPORTED AND REVIEWED BY THE MANUFACTURER. BAXTER HEALTHCARE IS NOT THE MANUFACTURER OF THE CURLIN INFUSION PUMP. EVALUATION METHOD: DATA VALIDATION. ORDER ENTRY LABELS, COMPOUNDER MIXCHECK REPORTS AND THE BLACK-BOX DATA WERE REVIEWED AND EVALUATED. RESULTS: USER DEVICE INTERFACE. IT IS UNKNOWN HOW OR WHAT CAUSED THE NUTRITION BAGS TO RUN DRY. CONCLUSION: NO FAILURE DETECTED, DEVICE OPERATED WITHIN SPECIFICATION. THE INFORMATION FROM THE DATA PROVIDED FOUND THAT THE EM1200 COMPOUNDER DID NOT CONTRIBUTE TO THIS EVENT. DEVICE NOT AVAILABLE FOR RETURN. | 1
4655618 | THE MARYLAND BIPOLAR FORCEPS 8MM LONGER WIRE SNAPPED, MAKING THE DEVICE UNUSABLE. | 0
5207990 | ><IT WAS REPORTED THAT THE CATHETERS ARE STUCK TO THE PACKING WITH THE ADHESIVE AND THEY CAN'T USE THEM, CATHETER WALL IS ADHERING TO THE INNER PACKAGING. | 0
4976720 | ONSITE TESTING OF THE INVOLVED DEVICES PERFORMED BY A SPACELABS FIELD SERVICE ENGINEER (FSE) CONFIRMED THE EQUIPMENT PERFORMED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. REVIEW OF THE CUSTOMER PROVIDED PATIENT RETROSPECTIVE DATABASE WAS CONDUCTED BY A SPACELABS LEAD SOFTWARE ENGINEER. THE EVENT TIME WAS LOCATED IN THE DATABASE. TWO ECG LEADS WERE RECORDED FOR THIS EVENT. ALL NINE BEATS OF THE EVENT WERE DETECTED; HOWEVER, NOT ALL BEATS WERE CLASSIFIED AS VENTRICULAR. FOUR BEATS IN THE SERIES FROM ONE LEAD WERE CLASSIFIED AS NORMAL SINCE THEIR SHAPE AND AMPLITUDE MATCHED THE NORMAL BEATS FOR THIS LEAD. THERE WAS NO ALARM FOR THIS EVENT SINCE ALARM CRITERIA WAS NOT MET. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY RECEIVER MODULE MODEL 90478 WITH ARIATELE TRANSMITTER MODEL 96281 AND CENTRAL MONITOR MODEL 91387-38 FAILED TO GENERATE AN ALARM FOR NINE BEATS OF VENTRICULAR TACHYCARDIA (VTACH) ON (B)(6) 2015 AT 3:22 A.M. NO ONE WAS INJURED AS A RESULT OF THIS EVENT. | 1
4741367 | SPACELABS RECEIVED A REPORT THAT ON (B)(6) 2015 AT 3:17 A.M. A TELEMETRY PATIENT MONITORED WITH TRANSMITTER MODEL 90343-5, RECEIVER MODULE MODEL 90478 AND CENTRAL MONITOR MODEL 91387-38 EXPERIENCED VENTRICULAR TACHYCARDIA (VTACH). THERE WAS NO ALARM GENERATED AT THE CENTRAL MONITOR AND NO RECORDING IN THE ALARMS TAB FOR ICS-G2 FOR THIS EVENT. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><USING THE PATIENT RETROSPECTIVE DATABASE PROVIDED BY THE CUSTOMER AND REVIEWED BY A SPACELABS LEAD SOFTWARE ENGINEER, A 25 BEAT EPISODE OF VENTRICULAR TACHYCARDIA (VTACH) WAS LOCATED FOR THE REPORTED EVENT. THERE WAS AN ASSOCIATED ALARM FOR THIS EVENT. THERE WAS A 13 BEAT VTACH EPISODE FOLLOWING THE 25 BEAT EPISODE WITH A 5.5 SECOND SEPARATION OF NORMAL BEATS. WHEN THERE IS LESS THAN 10 SECONDS SEPARATING TWO EPISODES OF VTACH, THE DATABASE REPORTS IT AS ONE LONG EPISODE INSTEAD OF TWO. THIS IS BY DESIGN. THERE WAS NO MALFUNCTION. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED.><ONSITE TESTING BY A SPACELABS FIELD SERVICE ENGINEER (FSE) WAS DECLINED BY THE CUSTOMER SINCE THE EQUIPMENT WAS STILL IN USE. SPACELABS HAS LAUNCHED AN INVESTIGATION INTO THIS ISSUE AND WILL FILE A SUPPLEMENTAL REPORT WHEN THE INVESTIGATION IS COMPLETE. PLACEHOLDER. | 1
5051451 | SPACELABS HAS LAUNCHED AN INVESTIGATION IN THIS EVENT AND WILL FILE A SUPPLEMENTAL REPORT WHEN THE INVESTIGATION IS COMPLETE.><THE PRODUCT WAS RECEIVED AT SPACELABS¿ EQUIPMENT SERVICE CENTER FOR REPAIR. THE REPORTED PROBLEM WAS VERIFIED. THE BATTERY ASSEMBLY PN 146-0039-00 WAS NOT FUNCTIONING; THEREFORE, THE PART WAS REPLACED. THE REPAIRED UNIT PASSED ALL FUNCTIONAL TESTS AND WAS RETURNED TO THE CUSTOMER. LACK OF POWER FOR THE MODULES WAS THE WARNING SIGN THAT THE CUSTOMER NOTICED AND THAT PROMPTED THE ACTION OF REMOVING THE PRODUCT FROM USE. THIS REPORT IS COMPLETE AND THIS PARTICULAR ISSUE IS CONSIDERED CLOSED. SPACELABS MONITORS ITS COMPLAINTS FOR SIMILAR PRODUCT ISSUES TO DETERMINE NEED FOR ADDITIONAL INVESTIGATION.><SPACELABS RECEIVED A REQUEST FOR SERVICE REPAIR FOR TELEMETRY MODULE HOUSING MODEL 90479 THAT LOST POWER DURING USE ON (B)(6) 2015 AND WOULD NOT RESTART. NO ONE WAS INJURED AS A RESULT OF THIS EVENT. | 1
4578612 | ONSITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THE EQUIPMENT PERFORMED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. A SPACELABS LEAD SOFTWARE ENGINEER REVIEWED THE PATIENT RETROSPECTIVE DATABASE PROVIDED BY THE CUSTOMER. THE DATABASE REVIEW CONFIRMED THERE WAS A 12 BEAT RUN OF VTACH AT THE REPORTED TIME. THE MONITOR DID NOT PRODUCE AN ALARM FOR THIS EVENT SINCE THE CUSTOMER HAD DISABLED THE AUTOMATIC LEAD SWITCH (ALS) FEATURE. THE RECEIVER MODULE PROCESSES ALL AVAILABLE LEADS FROM THE TRANSMITTER FOR ARRHYTHMIA AND ALARM MONITOR DISPLAY AS DETERMINED BY THE ECG ALGORITHM. THE CUSTOMER CHOSE TO VIEW ONE ECG LEAD AT THE MONITOR AND COULD NOT SEE THAT ANOTHER LEAD HAD INTERMITTENTLY LOST PATIENT CONTACT DURING THE REPORTED EVENT. A TEXT MESSAGE WAS DISPLAYED AT THE MONITOR TO ¿CHECK C¿ AND THE HEART RATE DISPLAYED WAS ¿???¿. THIS MESSAGE IS TO ALERT THE CUSTOMER TO CHECK THE PATIENT LEAD. HAD THE ALS FEATURE BEEN ENABLED, THE ECG ALGORITHM WOULD HAVE BEEN DIRECTED TO IGNORE THE LOST LEAD AND ALARM FOR VTACH. THERE WAS NO EQUIPMENT MALFUNCTION. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE IS CLOSED. PLACEHOLDER.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITORED WITH TRANSMITTER MODEL 90343-05, RECEIVER MODULE MODEL 90478, AND CENTRAL MONITOR MODEL 91387-38 DID NOT GENERATE AN ALARM FOR A WITNESSED RUN OF VENTRICULAR TACHYCARDIA (VTACH) AT 11:31 P.M. ON (B)(6) 2015. NO ONE WAS INJURED AS A RESULT OF THIS EVENT. | 1
4406202 | FURTHER INVESTIGATION BY THE SPACELABS FIELD SERVICE ENGINEER (FSE) CONFIRMED THE FACILITY BIOMED HAD INCORRECTLY PROGRAMMED THE TWO PATIENT TRANSMITTERS TO THE SAME CHANNEL NUMBER (FREQUENCY). CONSEQUENTLY, THE RECEIVER MODULE COULD NOT CORRECTLY IDENTIFY DATA FROM TWO DIFFERENT PATIENTS FOR VALID DATA PROCESSING AND CLINICAL PRESENTATION OR ALARMS NOTIFICATION. THIS WAS OPERATOR ERROR AND NO FAULT WAS FOUND WITH THE INVOLVED DEVICES. THE FSE WORKED WITH THE FACILITY BIOMED TO RESOLVE THE ISSUE AND CORRECTLY PROGRAM THE TRANSMITTERS TO INDEPENDENT CHANNEL NUMBERS (FREQUENCIES). THIS REPORT IS CONSIDERED COMPLETE AND THE ISSUE CLOSED. PLACEHOLDER.><SPACELABS RECEIVED A REPORT THAT TELE BED 13 AND TELE BED 25 HAVE TWO (2) DIFFERENT PATIENTS ADMITTED, HOWEVER BOTH OF THOSE BEDS RECEIVERS ARE PROGRAMMED TO THE SAME CHANNEL NUMBER. ONE PATIENT PASSED AWAY.><THE TELEMETRY TRANSMITTER SENDS PATIENT ECG SIGNALS TO THE TELEMETRY RECEIVER MODULE WHICH CONTAINS THE ECG ALGORITHM FOR PROCESSING THE SIGNALS. THE DATA IS ANALYZED FOR DISPLAY AND DETERMINES ALARM STATUS, SENDING ALARM MESSAGING TO THE TELEMETRY CENTRAL MONITOR. ONSITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THE EQUIPMENT WORKED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. SPACELABS IS EVALUATING THIS REPORTED EVENT AND WILL FILE A FOLLOW UP REPORT WHEN OUR EVALUATION IS CONCLUDED. | 1
5228153 | SPACELABS RECEIVED A REPORT THAT ON (B)(6) 2015 A TELEMETRY RECEIVER MODULE HOUSING MODEL 90479 FAILED WHILE IN USE. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><SPACELABS HAS LAUNCHED AN INVESTIGATION INTO THIS EVENT AND WILL FILE A SUPPLEMENTAL REPORT WHEN THE INVESTIGATION IS COMPLETE. PLACEHOLDER.><THE PRODUCT WAS RECEIVED AT SPACELABS¿ EQUIPMENT SERVICE CENTER FOR REPAIR. THE REPORTED PROBLEM WAS VERIFIED. POWER SUPPLY (B)(4) WAS NOT FUNCTIONING; THEREFORE THE BOARD WAS REPLACED. THE REPAIRED UNIT PASSED ALL FUNCTIONAL TESTS AND WAS RETURNED TO THE CUSTOMER. LACK OF POWER WAS THE WARNING SIGN THAT THE CUSTOMER NOTICED AND THAT PROMPTED THE ACTION OF REMOVING THE PRODUCT FROM USE. THIS REPORT IS COMPLETE AND THIS PARTICULAR ISSUE IS CONSIDERED CLOSED. SPACELABS MONITORS ITS COMPLAINTS FOR SIMILAR PRODUCT ISSUES TO DETERMINE NEED FOR ADDITIONAL INVESTIGATION. | 1
4974122 | SPACELABS RECEIVED A SERVICE REPAIR REQUEST FOR INTERMITTENT DISPLAY AND REBOOT FOR BEDSIDE MONITOR MODEL 91369 WHILE IN USE ON (B)(6) 2015. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><THE PRODUCT WAS RECEIVED AT SPACELABS¿ EQUIPMENT SERVICE CENTER FOR REPAIR. THE REPORTED PROBLEM WAS VERIFIED. THE INTEGRATED CHIP NOVRAM DS1646 COMPONENT WAS NOT FUNCTIONING; THEREFORE, THE COMPONENT WAS REPLACED. THE REPAIRED UNIT PASSED ALL FUNCTIONAL TESTS AND WAS RETURNED TO THE CUSTOMER. LACK OF A DISPLAY WAS THE WARNING SIGN THAT THE CUSTOMER NOTICED AND THAT PROMPTED THE ACTION OF REMOVING THE PRODUCT FROM USE. THIS REPORT IS COMPLETE AND THIS PARTICULAR ISSUE IS CONSIDERED CLOSED. SPACELABS MONITORS ITS COMPLAINTS FOR SIMILAR PRODUCT ISSUES TO DETERMINE NEED FOR ADDITIONAL INVESTIGATION.><SPACELABS HAS LAUNCHED AN INVESTIGATION INTO THIS EVENT AND WILL FILE A COMPLETE REPORT WHEN THE INVESTIGATION IS FINISHED. | 1
4317424 | QUALITY SYSTEM REGULATIONS AND OUR QUALITY MANAGEMENT SYSTEM WE INITIATED AN EVALUATION OF THE INCIDENT AS DESCRIBED IN THE MEDWATCH REPORT. WE IDENTIFIED A CUSTOMER COMPLAINT RECORD THAT DOCUMENTED THIS INCIDENT ON (B)(6) 2014 - NOTE THAT ACCORDING TO CLIENT'S MEDWATCH REPORT THE INCIDENT OCCURRED ON (B)(6) 2014. THE SUPPORTING DOCUMENTATION OF THE CASE CONTAINS CLIENT'S STATEMENT ABOUT KEY SETTINGS THAT WERE PURPOSEFULLY MODIFIED BY SOMEONE FROM THEIR STAFF. TECHNICAL SERVICE SPECIALIST ADVISED CLIENT TO INTRODUCE USER RESTRICTIONS BY CHANGING THE XLSECURITY SETTINGS IN THE CONFIGURATION OF THE NURSE STATION, IT WAS EXPLAINED TO THE CLIENT THAT DEVICE SOFTWARE PROVIDES ADEQUATE AUTHORIZATION, ACCESS CONTROL, AND ACCOUNTABILITY FEATURES THAT ARE PART OF THE BEST SECURITY PRACTICES. TO MODIFY THE ALARM SETTINGS THE USER REQUIRE TO FOLLOW A SEQUENCE OF STEPS THAT CAN BE PERFORMED IF THE USER DELIBERATELY INTENDS TO PERFORM SO. BASED ON THE INFORMATION WE HAVE REVIEWED WE CONCLUDE THAT THE DEVICE DID PERFORM AS INTENDED.><ON (B)(6) 2014 WE BECAME AWARE OF VOLUNTARY REPORT (REPORT #22007710000-2014-8038) SUBMITTED THROUGH FDA MEDWATCH PROGRAM INVOLVING ONE NATUS ELECTROENCEPHALOGRAPHY SYSTEM FOR LONG TERM MONITORING USED WITHIN AN EPILEPSY MONITORING UNIT AT (B)(6) HOSPITAL. REPORTEDLY THE EEG TECHNICIAN NOTIFIED THAT THE AUDIBLE ALARMS WERE NOT GOING OFF WHEN SHE PRESSED THE PATIENT EVENT BUTTON ATTACHED TO THE ACQUISITION SYSTEM, FURTHER ANALYSIS OF THE INCIDENT PERFORMED BY THE CLIENT CORROBORATED THAT THE ALARM SETTINGS WERE MODIFIED AT THE NURSE STATION MONITORING SYSTEM. SYSTEM PERFORMED AS INTENDED HOWEVER CLIENT EXPRESSED CONCERN ABOUT THE CAUSE OF THE PROBLEM LYING ON THE DEVICE INSTEAD OF THE STAFF INADVERTENTLY HAVING CHANGED THE SETTINGS. | 1
4920939 | A VESSEL SEALER WAS BEING USED BY SURGEON DURING THE PROCEDURE. THE INSTRUMENT DID NOT WORK PROPERLY. THE INTENDED PURPOSE OF SEALING THE TISSUE DID NOT WORK. THERE WAS NO HARM TO THE PATIENT. | 0
4522890 | ><THE NURSE (CNRA) WAS PREPPING FOR A CASE IN OPERATING ROOM. THE GE AISYS MACHINE HAD FINISHED ITS ROUTINE MACHINE CHECK AND PASSED. AS THE PATIENT WAS BEING BROUGHT INTO THE ROOM, THE INTERNAL ERROR MESSAGE POPPED UP ON THE SCREEN. THE MACHINE WAS SWAPPED OUT WITH ANOTHER. THE MANUFACTURER REPRESENTATIVE CAME IN TO EVALUATE THE MACHINE. THE REPRESENTATIVE RAN THE TROUBLESHOOTING PROGRAM SOFTWARE AND IT COMES UP AS A RARE ERROR THAT SAYS REBOOT AND THIS SHOULD BE FINE. THE SOFTWARE SAYS IF THIS ERROR OCCURS AGAIN TO REPLACE TWO COMPUTER BOARDS IN THE MACHINE BUT PROBABILITY OF SEEING THIS ERROR AGAIN IS EXTREMELY MINIMAL. MACHINE HAS BEEN PLACED BACK IN SERVICE. | 1
5109169 | FIELD SERVICE TECHNICIAN INDICATED THE UPS ON THE DEVICE POPPED AND SMOKED WHEN IT WAS CONNECTED TO POWER DURING THE INITIAL INVESTIGATION OF THE SERVICE CALL. THE DEVICE WAS DISCONNECTED FROM POWER AND REPLACED. THE DEVICE PLACED BACK IN SERVICE AFTER SUCCESSFUL TESTING. FLUID WAS PRESENT INSIDE THE UPS.><CUSTOMER REPORTED A NON-FUNCTIONING PAS DEVICE CAUSED BY A FLUID INGRESSION OF IV FLUIDS. THERE WAS NO PATIENT PRESENT AND NO HARM OCCURRED. | 0
4502020 | OLYMPUS WAS INFORMED THAT TWO PATIENTS SUSTAINED COLON PERFORATION DURING A DIAGNOSTIC COLONOSCOPY PROCEDURE. IT WAS REPORTED THAT BOTH PATIENTS WERE EXAMINED WITH THE SAME COLONOSCOPE. BOTH PATIENTS WERE TRANSFERRED TO OPERATING ROOM FOR SURGICAL INTERVENTION. BOTH PATIENTS ARE REPORTEDLY DOING FINE. THIS IS ONE OF TWO REPORTS. .><THE DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. THE EVALUATION REVEALED THAT THERE WAS EVIDENCE OF THIRD-PARTY REPAIRS ON THE DEVICE, AS THE ENTIRE INSERTION TUBE, AND LIGHT GUIDE TUBE WERE NON-OLYMPUS PARTS. THERE WAS NO SHARP EDGES OR SURFACES THAT WOULD LIKELY CAUSE OR CONTRIBUTE TO THE PATIENTS' OUTCOME. BASED ON THE INVESTIGATION, THE MOST LIKELY CAUSE OF THE REPORTED EVENT IS ATTRIBUTED TO USER HANDLING. ADDITIONALLY, THE INSTRUCTIONS FOR USE STATES: PROHIBITION OF IMPROPER REPAIR AND MODIFICATION. THIS INSTRUMENT DOES NOT CONTAIN ANY USER-SERVICEABLE PARTS. DO NOT DISASSEMBLE | 0
5129877 | THE ANALYZER ALERTED THE OPERATOR TO POSSIBLE SAMPLE ABNORMALITY. THE USER REPORTED ERRONEOUSLY ELEVATED WHITE BLOOD CELL (WBC) VALUE AS A PRELIMINARY REPORT AWAITING VERIFICATION. THE DOCTOR ADMINISTERED ANTIBIOTIC THERAPY. THE USER PERFORMED MANUAL DIFFERENTIAL AND OBSERVED THE PRESENCE OF NUCLEATED RED BLOOD CELLS (MRBC). THE USER MANUALLY CALCULATED THE CORRECTED WBC VALUE AND ISSUED A CORRECTED REPORT AND INFORMED THE DOCTOR. ANTIBIOTIC THERAPY WAS CONTINUED AFTER CORRECTED REPORTS WERE ISSUED. IT CANNOT BE CONFIRMED IF ANTIBIOTIC THERAPY WAS BASED ON THE ERRONEOUSLY ELEVATED PRELIMINARY WBC ALONE OR IF OTHER CLINICAL SIGNS AND SYMPTOMS WARRANTED THE TREATMENT. THERE WAS NO HARM TO THE PT. THE ANALYZER JUDGED THE SAMPLE AS POSITIVE" WITH INTERPRETIVE PROGRAM (IP) MESSAGING ALERTING THE OPERATOR TO POSSIBLE SAMPLE ABNORMALITY | 0
5337787 | FENESTRATED BIPOLAR ROBOTIC INSTRUMENT MALFUNCTIONED WHILE WORKING IN PATIENT'S ABDOMEN. A WIRE BROKE ON THE FENESTRATED BIPOLAR. WIRE STAYED ATTACHED TO INSTRUMENT.><. | 0
4625652 | ><OUR HEALTH SYSTEM RECENTLY UPGRADED THE ANESTHESIA MACHINE INVENTORY AT TWO FACILITIES. AT BOTH SITES GE AISYS CS2 MACHINES WERE INSTALLED (ABOUT 3 MONTHS AGO AT ONE SITE AND LAST MONTH AT THIS SITE). FOUR DIFFERENT AND RECENTLY INSTALLED MACHINES EXPERIENCED COMPLETE VENTILATOR FUNCTION FAILURE DURING OPERATIONS IN THEIR RESPECTIVE OPERATING ROOMS. IN EACH INSTANCE, THE SCREEN DISPLAY FOR VENTILATOR FUNCTIONS WENT BLANK SUDDENLY AND WITHOUT WARNING, AND AN ERROR MESSAGE WAS DISPLAYED. THIS MEANS THAT MECHANICAL VENTILATION CEASED AS WELL AS DELIVERY OF INHALATIONAL ANESTHETIC GASES, AND THE DISPLAY OF ALL VENTILATOR FUNCTION ALONG WITH IT. IN EACH INSTANCE, THE ANESTHESIA CARE TEAM MEMBERS QUICKLY SWITCHED OVER FROM MACHINE SUPPLIED VENTILATOR FUNCTION TO MANUAL VENTILATOR FUNCTION WITH WALL SUPPLY OXYGEN AND CONVERSION FROM INHALATIONAL ANESTHESIA TECHNIQUE TO INTRAVENOUS ANESTHESIA. DURING THESE EPISODES, NO PATIENT HARM WAS DETECTED OR OBSERVED AND THE SEPARATE PHYSIOLOGIC MONITOR ATTACHED TO THE AISYS CS2, THE GE CARESCAPE V850 MONITOR, CONTINUED TO FUNCTION APPROPRIATELY, DEMONSTRATING THAT EACH PATIENT'S CONDITION WAS STABLE AND UNCHANGED. IN EACH CASE, THE MACHINES WERE POWERED DOWN FOR LESS THAN 15 SECONDS AND THEN TURNED BACK ON, WHICH REBOOTED THE ELECTRONICS AND RETURNED ALL FUNCTIONS VENTILATOR FUNCTIONS | 1
5049708 | THE CUSTOMER STATED THAT THE DEVICE WOULD TURN ON, TAKE VITALS, STAY ON FOR A FEW MINUTES AND THEN BEEP A FEW TIMES AND SHUT OFF.><WELCH ALLYN IS REPORTING THIS IN AN ABUNDANCE OF CAUTION FOR A POTENTIAL UNEXPECTED DEVICE SHUTDOWN. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE DEVICE EVALUATION IS COMPLETE.><WELCH ALLYN PRODUCT SERVICE CONFIRMED THE DEVICE WAS RECEIVED WITH A LOW CAPACITY BATTERY, DUE TO BATTERY AGE, WHICH CAUSED THE ISSUES AS SPECIFIED BY THE CUSTOMER. A KNOWN GOOD BATTERY WAS FOUND TO RESOLVE THE ISSUE. NO FAILURE WAS FOUND WITHIN THE DEVICE ITSELF. THE OPERATOR'S MANUAL PROVIDES AN ANNUAL BATTERY TEST TO PREVENT THIS TYPE OF FAILURE. NO FURTHER INVESTIGATION WILL BE CONDUCTED. | 1
4466577 | CRITICALLY ILL PT IN RESUSCITATION. ORDERS COULD NOT BE ENTERED IN (B)(6) SYSTEM BECAUSE CHART WAS OPEN ELSEWHERE IN THE FACILITY. IT SERVICES STATED THEY COULD DO NOTHING TO CORRECT THE PROBLEM. DEADLOCK WAS EVENTUALLY RESOLVED, BUT NOT BY SPECIFIC ACTION (OTHER USER EVENTUALLY RELINQUISHED THE CHART). | 1
5093864 | ><GIRAFFE INCUBATOR MATTRESS; WHEN POSITIONED AT OPTIMAL ELEVATION FOR VENT ASSOCIATED PNEUMONIA PREVENTION (20-30%) (EVIDENCE BASED) THE DOOR TO ACCESS THE PATIENT IS BLOCKED AND THEREFORE NOT USABLE. USE OF THIS BED PREVENTS ABILITY TO PERFORM EVIDENCE BASED BEST PRACTICE FOR HEAD OF BED ON INTUBATED NEONATES. | 0
5301537 | ON (B)(6) 2015, OLYMPUS WAS INFORMED THE FOLLOWING INFORMATION. DURING AN UNSPECIFIED PROCEDURE, THE TRANSFER OF AN ENDOSCOPIC IMAGE FROM UWIT-TX TO UWIT-RX CONNECTED TO THE SUB MONITOR IS INTERRUPTED. THE FACILITY REPLACED THE UWIT-RX AND THE MONITOR WITH SPARE DEVICES AND THE PHYSICIAN COMPLETED THE PROCEDURE. THERE WAS NO REPORT OF THE PATIENT'S INJURY REGARDING THIS EVENT.><THE SUBJECTED DEVICE WAS RETURNED TO OLYMPUS MEDICAL SYSTEMS CORP.(OMSC) FOR EVALUATION. OMSC EVALUATED THE DEVICE BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED. IN A STATE UWIT-TX AND UWIT-RX ARE FACE-TO-FACE HORIZONTALLY, OMSC SWITCHED THE CHANNELS OF UWIT-TX AND UWIT-RX 500 TIMES. BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED. IN A STATE THEY ARE NOT FACE-TO-FACE HORIZONTALLY, OMSC SWITCHED THE CHANNELS OF UWIT-TX AND UWIT-RX 500 TIMES. BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED. OMSC CHECKED THE DEVICE HISTORY RECORD OF THE SUBJECT DEVICE, AND THERE WAS NO IRREGULARITY FOUND. THE UWIT-TX AND UWIT-RX INSTRUCTION MANUALS STATE THE CORRESPONDING METHOD WHEN THERE IS AN ABNORMALITY IN THE ENDOSCOPIC IMAGE. THERE WERE NO FURTHER DETAILS PROVIDED. IF SIGNIFICANT ADDITIONAL INFORMATION IS RECEIVED, THIS REPORT WILL BE SUPPLEMENTED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. THIS REPORT IS ONE OF TWO REPORTS. CROSS REFERENCE MFR. REPORT NUMBER 8010047-2015-01260.><OMSC EVALUATED THE DEVICE CONTINUOUSLY BUT THE REPORTED PHENOMENON WAS NOT REPRODUCED.OMSC SURMISED THAT THE CAUSE OF THE PHENOMENON IS THE FOLLOWINGS IN THEORY.FACTOR1 - THE INSTALLATION STATUS OF UWIT-TX AND UWIT-RX IS INAPPROPRIATE. UWIT-TX AND UWIT-RX CANNOT COMMUNICATE. FACTOR2 - UWIT-TX AND/OR UWIT-RX GO WRONG. UWIT-TX AND UWIT-RX WERE NOT SYNCHRONIZING.FACTOR3 - UNDER THE INFLUENCE OF THE EXTERNAL ENVIRONMENT, UWIT-TX AND/OR UWIT-RX RAN THE SYSTEM RESET. UWIT-TX AND UWIT-RX WERE NOT SYNCHRONIZING.OMSC CHECKED THIS DEVICE, DID NOT FIND ANY ABNORMALITY.OMSC CHECKED THE INSTALLATION STATUS OF UWIT-TX AND UWIT-RX IN THE FACILITY. THE DISTANCE BETWEEN UWIT-TX AND UWIT-RX WAS FROM 1 - 5 M, THIS IS WITHIN THE RANGE OF IFU. THE DIRECTION OF UWIT-TX AND UWIT-RX WAS NOT SLIGHTLY HORIZONTAL. BASED ON THESE AND THE INVESTIGATION RESULT OF SUBJECT DEVICES, OMSC SURMISED THAT THIS PHENOMENON MIGHT BE CAUSED BY FACTOR3". "FACTOR3" MIGHT BE CAUSED BY ANY OF THE FOLLOWING MECHANISMS.- THE DCDC CONVERTER OF UWIT-TX AND/OR UWIT-RX DID NOT OUTPUT THE ELECTRIC VOLTAGE ENOUGH | 1
4946407 | AUTISTIC INDIVIDUAL WITH HISTORY OF CONSTIPATION BECAME UNRESPONSIVE IN GROUP HOME. HE RECEIVED EMERGENCY MEDICAL CARE. DEATH WAS PRONOUNCED IN THE EMERGENCY ROOM. AN AUTOPSY FOUND PERFORATION OF THE CECUM ASSOCIATED WITH A FOREIGN BODY: A MEDICAL DEVICE, A WIRELESS ESOPHAGEAL PH MONITORING DEVICE (BRAVO CAPSULE) WAS FOUND IN THE CECUM AT AUTOPSY, ADJACENT TO THE PERFORATION SITE. THE PERITONEAL CAVITY CONTAINED BLOOD AND FECAL MATTER. DATES OF USE: PRESUMED DAYS. DIAGNOSIS OR REASON FOR USE: PLACED IN ESOPHAGUS FOR PH MONITORING- LATER FOUND IN CECUM. | 0
4489968 | THE CASE WAS PROCEEDING WELL, AND THERE WAS NO FLUID DEFICIT (MINUS THE 100ML ON FLOOR). THE INTRAUTERINE TISSUE REMOVAL PORTION OF THE CASE WAS STARTED AND IT WAS NOTED THAT THE AMOUNT OF POSITIVE DEFICIT (MEANING MORE FLUID THAN LESS) WAS NOTED TO CLIMB STEADILY. THE PATIENT WAS STABLE AND THERE WERE NO CONCERN FROM DOCTOR AS THE FLUID WAS CLEAR WITH ONLY A HINT OF BLOOD NOTED IN FLUID. THE CIRCULATOR CONTINUED TO MONITOR AND KEEP SURGEON INFORMED. THE INTAKE AND OUTPUT NOTED VIA OBSERVATION AND CALCULATION SHOWED THERE WAS NO ADDITIONAL FLUID. AT THAT POINT, A SCALE WAS OBTAINED AND ALL FLUIDS WERE WEIGHED AND IT WAS CALCULATED THAT THERE WAS A 300ML DEFICIT AND NOT AN ADDITION OF FLUID. A WORK ORDER WAS PUT IN TO HAVE THE FLUID MANAGEMENT SYSTEM CHECKED AND CALIBRATED. | 0
4901506 | DURING A ROBOTIC LOBECTOMY THE ROBOT SYSTEM BEGAN TO FREEZE WITH ERROR MESSAGES DISABLING THE EQUIPMENT. AFTER RESTARTING THE SYSTEM, IT FROZE AGAIN. THE SURGEON CONVERTED THE ROBOTIC LOBECTOMY TO AN OPEN SURGICAL PROCEDURE. THE ROBOTIC HOTLINE WAS CALLED AND THE ROBOTIC ENGINEER FLEW IN THAT EVENING TO REPLACE PARTS; THE UCC COMMUNICATION BOARD AND MICROCHIPS. ROBOTIC NOW FUNCTIONING AGAIN.SUBSEQUENTLY, SIX DAYS LATER, DURING A ROBOTIC ASSISTED LAPAROSCOPIC PARTIAL NEPHRECTOMY, THE DA VINCI XI ROBOT HAD AN IRRECOVERABLE FAULT". THE SURGICAL TEAM WAS UNABLE TO RESET THE MACHINE. THE ROBOTIC CASE HAD TO BE CONVERTED TO AN OPEN PROCEDURE." | 1
4660629 | WHEN DOING A ROBOTIC INCISIONAL/UMBILICAL HERNIA REPAIR WITH MESH, THE MEGA SUTURECUT DAVINCI INSTRUMENT HAD A WIRE NEAR THE TIP OF INSTRUMENT THAT BENT OUTWARD AND 1 SMALL WIRE CAME UNWRAPPED WITH END POINTING OUT. INSTRUMENT REMOVED AND REPLACED. NO VISIBLE HARM TO THE PATIENT. THE INSTRUMENT HAD 4 USES LEFT AND WAS NOT REPROCESSED. | 0
5041755 | FENESTRATED BIPOLAR BEING USED WHEN STAFF NOTICED SMALL WIRE STICKING OUT. A 45MM STAPLER FAILED TO FIRE AND STAFF STATED THAT THERE MAY BE A SHARP/BLADE STICKING OUT OF INSTRUMENT.><. | 0
4477748 | THE PROTECTIVE WHITE PORTION OF THE SHEARS CAME OFF DURING THE PROCEDURE. | 0
4902600 | THE BUG WAS DISCOVERED DURING TESTING INTERNALLY IN THE MANUFACTURER'S FACILITY. THE BUG WAS EVALUATED TO HAVE PATIENT RISK. THE BUG CAUSES BOLD ACTIVATIONS MAPS TO BE VISUALIZED AS OVERLAYS WITHOUT TAKING THE COREGISTRATION INTO ACCOUNT. ANY OUTPUT CREATED FROM THESE ACTIVATION MAPS AS OVERLAYS WILL NOT BE ADJUSTED ACCORDING TO THE COREGISTRATION. THE BUG ONLY AFFECTS THE BOLD FMRI ANALYSIS AND NO OTHER MODULES OR ANALYSES. FURTHERMORE, THE PROBLEM WILL NOT OCCUR IF THE BOLD/DTI WIZARD IS USED, OR IF COREGISTRATION IS DONE BEFORE THE BOLD MODULE ANALYSIS IS STARTED.><DEVELOPMENT TEAM DURING INTERNAL TESTING DISCOVERED A BUG CONCERNING HANDLING OF COREGISTRATION FOR BOLD SERIES. THIS BUG WAS VERIFIED IN CURRENT VERSION 2.3.14. ALL USERS WILL BE NOTIFIED TO STOP USING THE BOLD MODULE IN NORDICICE AND CONTACT NORDICNEUROLAB FOR FURTHER INSTRUCTIONS FOR BLOCKING THE BOLD MODULE FOR FURTHER USE. AFFECTED USERS WILL BE OFFERED A TRANSITION TO NORDICBRAINEX ((B)(4)) WHICH IS OPTIMIZED FOR THE BOLD WORKFLOW. RISK EVALUATION: THE PROBLEM CAN POTENTIALLY CAUSE A MISALIGNMENT OF THE BOLD ACTIVATIONS IN RELATIONS TO THE STRUCTURAL UNDERLAY. GIVEN THAT THE HEAD MOTION CORRECTED BY THE COREGISTRATION IS USUALLY SMALL, THE EFFECT OF THIS BUG IS EQUALLY SMALL. THE MOST COMMON SIZE OF MISALIGNMENTS ARE ON THE SAME ORDER OF MAGNITUDE AS THE INHERENT UNCERTAINTIES OF THE ANALYSIS ITSELF, WHICH MAKES IT UNLIKELY THAT IS SHOULD INTRODUCE ANY ADDED RISK PATIENT HARM IN THE LAST STEP. FOR LARGER MISALIGNMENTS, IT IS UNLIKELY THAT THE ERROR IS NOT DETECTED BY THE USER. | 1
5339192 | UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR.><CUSTOMER REPORTED THAT DURING A CASE VITALS CAPTURE (FOR DOCUMENTING VITAL SIGNS DATA IN THE CHRONOLOGICAL RECORD FOR THE PROCEDURE) STOPPED WORKING DUE TO NIBP FAILURES. THE USER EXPECTS THAT WHEN AN AUTOMATIC INTERVAL IS SET FOR VITALS CAPTURE IT SHOULD BE FOLLOWED BY THE SYSTEM. HOWEVER, IN THIS CASE, THE STAFF DIDN'T REALIZE THE VITALS CAPTURE HAD STOPPED. THE PATIENT DIED AFTER LEAVING THE CATH LAB AND THE CUSTOMER DOES NOT FEEL MERGE HEMO CONTRIBUTED TO THE PATIENT'S DEATH. THE PATIENTS' VITAL SIGNS (MINUS THE NIPB) PLUS THE LIVE ECG AND INVASIVE PRESSURES WERE ALWAYS VISIBLE ON THE HEMO SCREEN. ALSO, THE SYSTEM IS DESIGNED TO STOP VITALS CAPTURE DURING USE IF NIBP MEASUREMENTS STOP. | 0
5339170 | UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR.><INR VALUES PRESENTING INCORRECTLY IN HEARTSUITE HEMO - DECIMAL IS BEING DROPPED RESULTING IN ERRONEOUS VALUE. IT USED AN ALGORITHM THAT LOOKED FOR A .0" AND REPLACED IT WITH AN EMPTY STRING. FOR EXAMPLE: 1.0 WOULD SHOW AS 1. IF THE INR ON PRE-PROCEDURE IS 1 | 1
5338447 | UPON RETROSPECTIVE REVIEW, THIS ISSUE WAS DETERMINED TO BE REPORTABLE AS AN MDR.><MERGE CARDIO 10.1 REPORTS DO NOT DISPLAY PATIENT IDENTIFIERS IN THE HEADER OF EACH PAGE, THEY ARE ONLY DISPLAYED ON THE FIRST PAGE. IF REPORTS ARE PRINTED TO HARD COPY, PAGES COULD POTENTIALLY BECOME MIXED UP BETWEEN DIFFERENT PATIENTS. PRINTING REPORTS TO HARD COPY SO THEY CAN BE FAXED IS NOT A COMMON METHOD OF DELIVERY BUT IS USED BY THE CUSTOMER THAT REPORTED THIS ISSUE. USER ERROR IS THE MAIN CONTRIBUTING FACTOR. THE FREQUENT USE OF MASS FAXES AND PRINTOUTS IS MAKING IT MORE LIKELY THAT PAGES GET MIXED UP. THIS WOULD NOT BE AN ISSUE IF ALL REPORTS ARE ELECTRONICALLY DISTRIBUTED AS PDFS. | 1
5250004 | ><REVIEW OF THE RELEASE DOCUMENTATION AND PAPERWORK FOR THE TWO ORDERS REVEALED THAT GE BATCH NUMBER (B)(4) (BMS LOT NUMBER: 14100304-1446) THINSTRAND SUTURE MATERIAL WAS USED FOR BOTH ORDERS. A NONCONFORMANCE REPORT WAS INITIATED TO DOCUMENT DETERIORATION OF THIS LOT OF MATERIAL, WHICH HAS AFFECTED ITS INTENDED USE. THERE HAD BEEN DIFFICULTY EXPERIENCED IN THE BUILDING OF THE STRANDS LEADING TO AN INCREASE IN FAILURES DURING THE HEATING PROCESS OF BUILDING TSRX ORDERS. THE TWO ORDERS REPORTED IN THE COMPLAINT WERE BOTH PROCESSED ON (B)(6) 2015 AND SHIPPED ON 15-0CT-2015. THE MATERIAL, MANUFACTURED IN OCT-2014, DID NOT CONTAIN AN EXPIRATION DATE ON ITS LABELING PROVIDED BY THE SUPPLIER. PHYSICAL TESTING WAS PERFORMED ON THE MATERIAL CONFIRMING ITS DETERIORATING CONDITION. THIS HAS LED THEM TO BELIEVE THAT THE CONDITION OF THE MATERIAL LIKELY CONTRIBUTED TO THE JAMS EXPERIENCED BY THE CUSTOMER FOR THE TWO ORDERS REPORTED IN THE ORIGINAL COMPLAINT. AN INCIDENT REPORT WAS PROVIDED DETAILING IMMEDIATE ACTIONS TAKEN WITH REGARDS TO THIS MATTER, INCLUDING ASSIGNING A 6 MONTH EXPIRATION PERIOD TO THE MATERIAL FROM THE TIME IT IS RECEIVED AT (B)(4). A NEW LOT HAS BEEN RECEIVED THAT HAS BEEN IN PROCESS FOR SEVERAL ORDERS TO THIS DATE WITHOUT EXPERIENCING ANY ISSUES OF DETERIORATION OR PROCESS FAILURES. | 0
4855723 | DURING SURGERY, THE CABLE BROKE ON THE ROBOT ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER. THE INSTRUMENT WAS REMOVED AND REPLACED. THE CABLE CAME OFF ITS TRACK ON 2ND ROBOT ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER. THE INSTRUMENT WAS REMOVED AND REPLACED AGAIN. | 0
5220636 | GE ALARM FALSE ALARMING VTACH IN A PATIENT WITH A PACEMAKER AND WIDE QRS COMPLEX. | 1
5245492 | INITIAL ANALYSIS SUGGESTED TWO CAUSES A: BELT SLIPS/JUMPING B: LOCK SOLENOID PIN DISLODGE/DISENGAGED. STUDIES SHOWED IF THE SOLENOID LOCKING PIN WAS NOT ENGAGED OR BECAME DISLODGED FROM THE SLOT, IT'S POSSIBLE THE BRAKES ARE NOT SUFFICIENT TO PREVENT THE MIDDLE ARM OR CRADLE FROM MOVING OR ARREST AN ARM IN MOTION, AND THE BELT THAT LINKS THE MIDDLE ARM AND CRADLE TO THE DRIVE AND BRAKING MECHANISMS CAN JUMP TEETH, ALLOWING MOVEMENT. A VISIT TO THE CUSTOMER SITE WAS CONDUCTED ON 02-FEB-2016. UPON INSPECTION OF THE SOLENOID LIMIT SWITCH PCB, IT WAS OBSERVED THAT THE SWITCH WAS INCORRECTLY SET UP AND IT WAS POSSIBLE FOR THE ROUND TIP OF THE SOLENOID PIN TO NOT FULLY CLEAR THE SLOT. IN THIS SCENARIO IT IS POSSIBLE FOR THE PIN TO DISENGAGE FROM THE SLOT. IT WAS NOT POSSIBLE TO REPLICATE THE FAULT, HOWEVER FOR THE ISSUE TO OCCUR A FAULT MUST BE PRESENT IN THAT THE SOLENOID PIN DISLODGES FROM THE CRADLE LOCKING SLOT, AND THE BRAKE/BELT MECHANISM DOES NOT HAVE SUFFICIENT HOLDING POWER TO HOLD THE CRADLE AND MIDDLE ARM, ALLOWING THE CRADLE TO EXTEND. TWO INDEPENDENT FAULTS ARE REQUIRED FOR THIS FAILURE TO OCCUR, AND IN NORMAL OPERATION THE SOLENOID PIN AND LATCHING ARM ARE SUFFICIENT TO HOLD THE ARM IN THE DESIRED POSITION. RISK ASSESSMENT: TWO INDEPENDENT FAULTS ARE REQUIRED FOR THIS FAILURE TO OCCUR, AND IN NORMAL OPERATION THE SOLENOID PIN AND LATCHING ARM ARE SUFFICIENT TO HOLD THE ARM IN THE DESIRED POSITION. IN A SINGLE FAULT (E.G. IF THE SOLENOID PIN FAILS) THE RISK INCREASES, AND THEREFORE MUST BE CONSIDERED AS A SINGLE FAULT. SEVERITY: IF THE PANEL WERE TO IMPACT WITH A PATIENT THIS WOULD CAUSE A MAJOR INJURY. LIKELIHOOD OF HARM: THE PROBABILITY THE HAZARD RESULTING IN HARM IS DEPENDENT ON THE GANTRY ANGLES AND REQUIRES THE PATIENT TO BE ON THE TABLE, AND IS THEREFORE CONSIDERED UNLIKELY. THE IVIEWGT IFU STATE NOT TO OPEN OR CLOSE THE PANEL AT 180DEGREES +/- 45DEGREES, ESPECIALLY WHEN A PATIENT IS ON THE TABLE. THE XVI IFU DOES NOT SPECIFY THE SAME RECOMMENDATION, THEREFORE DOES NOT RULE OUT THE POSSIBILITY A USER MAY CHOOSE TO OPEN/CLOSE THE GANTRY POSITIONS. THESE FACTORS COMBINE TO GIVE A LIKELIHOOD OF REMOTE. THEREFORE TH RISK VALUE BASED ON MAJOR/REMOTE IS CALCULATED AS TOLERABLE.><A VISIT TO THE CUSTOMER SITE WAS CONDUCTED ON THE 02-FEB-2016. UPON INSPECTION OF THE SOLENOID LIMIT SWITCH PCB, IT WAS OBSERVED THAT THE SWITCH WAS INCORRECTLY SET UP AND IT WAS POSSIBLE FOR THE ROUND TIP OF THE SOLENOID PIN TO NOT FULLY CLEAR THE SLOT. IN THIS SCENARIO IT IS POSSIBLE FOR THE PIN TO DISENGAGE FROM THE SLOT. HAD THIS FAILURE OCCURRED WITH A PATIENT ON THE PRECISE TABLE AND THE GEOMETRIC SET UP BEEN IN SUCH A WAY THAT THE PANEL CAME INTO CONTACT WITH THE PATIENT THIS WOULD RESULT IN A MAJOR INJURY. THIS COULD HAPPEN AT GANTRY ANDLES BETWEEN 180±30 FOR IVIEW GT AND 90±30 FOR XVI. IT WAS NOT POSSIBLE TO REPLICATE THE FAULT, HOWEVER FOR THE ISSUE TO OCCUR A FAULT MUST BE PRESENT IN THAT THE SOLENOID PIN DISLODGES FROM THE CRADLE LOCKING SLOT, AND THE BRAKE/ BELT MECHANISM DOES NOT HAVE SUFFICIENT HOLDING POWER TO HOLD THE CRADLE AND MIDDLE ARM, ALLOWING THE CRADLE TO EXTEND. TWO INDEPENDENT FAULTS ARE REQUIRED FOR THIS FAILURE TO OCCUR, AND IN NORMAL OPERATION THE SOLENOID PIN AND LATCHING ARM ARE SUFFICIENT TO HOLD THE ARM IN THE DESIRED POSITION. IT SHOULD BE NOTED THAT EVEN WHEN THE SWITCH WAS SET TO THE WORST CASE POSITION THE SOLENOID PIN FAILURE COULD NOT BE REPRODUCED. ALONG WITH THE 2 INDEPENDENT FAILURES IN ORDER FOR THIS TO CAUSE HARM TO A PATIENT THEN THE FAILURE MUST OCCUR AT SPECIFIC ANGLES (180±30 FOR IVIEW GT AND 90±30 FOR XVI). INFORMATION GATHERED FROM CLINICAL SPECIALISTS SUGGEST THAT FOLDING AND RETRACTING THE PANEL WHILE THE PATIENT IS ON THE TABLE IS INFREQUENT. THE IVIEWGT INSTRUCTIONS FOR USE ALSO STATE NOT TO OPEN OR CLOSE THE PANEL AT 180±45[?] COMBINING ALL THESE FACTORS ELEKTA RATE THE LIKELIHOOD OF HARM TO BE IMPROBABLE. COMBINING ALL THESE FACTORS THE LIKELIHOOD OF HARM IS RATED IMPROBABLE. A MAJOR AND IMPROBABLE CAUSE CALCULATES A TOLERABLE RISK VALUE. THE MANUFACTURER IS CURRENTLY INVESTIGATING WHETHER THERE COULD BE IMPROVEMENTS MADE TO THE CORRECTIVE MAINTENANCE MANUAL OR THE TEST REQUIREMENTS SPECIFICATION WHICH WOULD CLARIFY THE IMPORTANCE OF SETTING THE MICRO-SWITCH ON THE SOLENOID CORRECTLY.><IT WAS REPORTED THAT DURING A HEXAPOD ISOCENTRE CALIBRATION, THE IVIEWGT IMAGING ARM WAS RETRACTED AND WHILST PART-WAY THROUGH FOLDING INTO THE CLOSED POSITION, THE CRADLE AND MID ARM SECTIONS UNCONTROLLABLY EXTENDED. THE GANTRY WAS AT 180 DEGREE AND THE ARM WAS APPROXIMATELY 45 DEGREE THROUGH THE FOLDING ACTION. THERE WERE NO INJURIES REPORTED.><THE MANUFACTURER'S INVESTIGATION IS ONGOING. FURTHER DETAILS WILL BE PROVIDED IN THE FINAL REPORT. | 0
4559654 | THIS INSTRUMENT HAS BEEN INVESTIGATED. ON (B)(4) 2015 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE. TO ENSURE PROPER CONFIGURATION OF THE HANDHELD SCANNER, THE FE PERFORMED MOD 35 TO RECONFIGURE THE HANDHELD SCANNER. THE FE COMPLETED THE PROCEDURE AND VERIFICATION OF SAMPLE BARCODES AGAIN WITHOUT ISSUES. THE FE INSPECTED THE SAMPLE CAMERA IN THE ANALYZER AND FOUND NO ISSUES. THE FE REVERIFIED THE CAMERAS ALIGNMENTS AND SETTING AS WELL AS THE OPTICS WITH NO ISSUES. THE FE SCANNED OCD AND CUSTOMER BARCODES WITHOUT ISSUES. THE FE MOVED THE SAMPLE CAMERA FROM M1 TO M2 MOUNTING LOCATION AND PERFORMED ALL ADJUSTMENTS AS WELL AS ENSURED THE DISTANCE SETTING WAS SET TO 130. THE FE RE-PERFORMED ALL ADJUSTMENTS AND TROUBLESHOOTING STEPS WITH NO RECURRENCE OF THE ISSUES REPORTED. THE FE LEFT THE CAMERA AT MOUNTING M2 POST REPAIRS. THE FE RAN ALL SAMPLE CAMERA DIAGNOSTICS AND IN PROVUE ST WITH NO ISSUES AND ALL LABELS READ AS INDICATED. THE CUSTOMER RAN QC TO VERIFY THE ANALYZER PERFORMANCE POST REPAIRS WITH NO FURTHER ISSUES AND ALL RESULTS PASSING AND REPORTING WITHIN LAB SPECS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION.><THE CUSTOMER CALLED TO REPORT ONE INCIDENT OF A SAMPLE TUBE BARCODE MISREAD BY THE PROVUE CAMERA. THE SAMPLE TUBE BARCODE WAS MISREAD AS 564390 INSTEAD OF 564330. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4). | 1
5252194 | THE FE PERFORMED THE READER CAMERA VALUE ADJUSTMENT AND PLACED CAM0 VALUE BACK INTO SPECIFICATION OF 110 ((B)(4)). REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION.><THE FE DISCOVERED THAT THE PROVUE CAMERA SETTING WAS OUT OF SPECIFICATION. THE LOG SHOWED THE SETTING TO BE AT 100 WHICH IS OUT OF SPECIFICATION. THE CUSTOMER IS UNABLE TO CONFIRM NO ERRONEOUS RESULTS HAVE BEEN REPORTED DUE TO THE LENGTH OF TIME THE OUT OF SPECIFICATION CONDITION HAS EXISTED. THE FE INFORMED THE CUSTOMER OF THE OUT OF SPECIFICATION FINDING. (B)(4). | 0
4534488 | CUSTOMER REPORTING FALSE NEGATIVE REACTIONS ON PROVUE WITH A CAP SAMPLE KNOWN TO CONTAIN ANTI-FYA TO THE (B)(6) RESOLVE PANEL A LOT# VRA216.><ON (B)(6) 2015 AND (B)(6) 2015 AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE. FE CONTACTED SECOND LEVEL SUPPORT FOR ASSISTANCE. AFTER TROUBLESHOOTING, SECOND LEVEL SUPPORT REQUESTED THE FE TO VERIFY THE TUBE DIAMETER SETTING. THE LOG SHOWS THAT THE PROVUE IDENTIFIED THE SAMPLE TUBE DIAMETER AS 15 MM AND ANOTHER TIME AS 12 MM. IF THE TUBE DIAMETER IS NOT CORRECTLY IDENTIFIED THIS COULD HAVE AN IMPACT ON THE ASPIRATION OF THE SAMPLE. FROM THE LOGS, IT WAS OBSERVED THAT TUBE WAS INITIALLY IDENTIFIED AS 12MM AND USER MANUALLY CHANGED TO 15MM. DUE TO THE DIAMETER OF THE TUBE BEING CHANGED, NOT ENOUGH SAMPLE IS ASPIRATED, AS A RESULT, SAMPLE MAY NOT BE DISPENSED TO THE FIRST FEW WELLS RESULTING IN THE UNEXPECTED VOLUME ERROR ~". IN THESE CASES | 0
5126683 | THIS TESTING WAS FOR CORRELATION STUDIES AND UNUSUAL TESTING WAS PURPOSELY PERFORMED PER PROTOCOL (PATIENT IS TYPE B AND DONORS ARE TYPES A AND AB). CUSTOMERS WOULD NOT INTENTIONALLY CHOOSE INCOMPATIBLE BLOOD GROUPING UNITS AND PERFORM AN IMMEDIATE SPIN CROSSMATCH. VARIATIONS IN CELL CONCENTRATION CAN AFFECT THE SENSITIVITY OF TEST RESULTS. WHEN THE CELL CONCENTRATION OF RED BLOOD CELLS IS MODIFIED FOR USE IN CERTAIN APPLICATIONS, SUCH AS THE GEL TEST OR COLUMN AGGLUTINATION METHODS, THE ACCURACY OF THE CELL CONCENTRATION BECOMES MORE CRITICAL. WHEN CELLS ARE TOO CONCENTRATED, WEAKER RESULTS MAY BE OBTAINED DUE TO THE INCREASE IN THE ANTIGEN/ANTIBODY RATIO. WHEN THE CELL CONCENTRATION IS TOO LOW, THE RED CELLS MAY BE DIFFICULT TO SEE, POTENTIALLY CAUSING A WEAK REACTION TO BE MISSED. A DEFINITIVE ROOT CAUSE WAS NOT FOUND ALTHOUGH, THE COMPATIBLE CROSSMATCH MAY BE DUE TO IMPROPER CELL CONCENTRATION OF THE DONOR CELLS OR THAT THE DONORS TESTED MAY BE WEAK SUBGROUPS. NO ERRONEOUS RESULTS WERE REPORTED DUE TO THIS INCIDENT.><CUSTOMER REPORTING TWO FALSE NEGATIVE RESULTS DURING IMMEDIATE SPIN CROSSMATCH USING THE MTS BUFFERED GEL CARDS. ACCORDING TO CUSTOMER, FACILITY WAS PERFORMING CORRELATION STUDIES BETWEEN TUBE METHOD AND BUFFERED GEL CARDS TESTED ON PROVUE. (B)(4). | 0
5239497 | ON 10-26-15 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE TO SERVICE THE INSTRUMENT AND CONFIRMED THE IMAGE ISSUES BY REVIEWING THE LOG FILES. FE PERFORMED ADJUSTMENTS TO THE GRIPPER AND VERIFIED CAMERA BRIGHT VALUE IS WITHIN SPECIFICATION. TO VERIFY PROPER OPERATION ON THE PROVUE, FE RAN CARD READ DIAGNOSTICS AND CUSTOMER RAN QC. CUSTOMER ACCEPTED SERVICE AND RETURNED INSTRUMENT INTO SERVICE. PATIENT RESULTS WERE MANUALLY ADJUSTED BY THE OPERATOR TO REFLECT THE WEAK POSITIVE REACTIONS BEFORE REPORTING. NO INCORRECT OR ERRONEOUS RESULTS WERE REPORTED AS A RESULT OF THIS INCIDENT. THERE WAS NO RISK PRESENT AT THE TIME OF THE INCIDENT. THERE WAS NO HARM TO ANY PATIENT. THE INVESTIGATION DETERMINED THAT THE MOST LIKELY ROOT CAUSE OF THIS EVENT IS A COMBINATION BETWEEN INSTRUMENT RELATED ISSUE AND PATIENTS' ANTIBODIES BEING WEAK AND/OR AT THE DETECTION LIMIT OF THE TECHNIQUE AND REAGENTS USED.><ACCOUNT REPORTS RECEIVING ?" DURING ANTIBODY SCREENING ON FOUR PATIENT SAMPLES. AS PER CUSTOMER'S POLICY | 0
5204072 | THE FE DISCOVERED THAT THE PROVUE CAMERA SETTING WAS OUT OF SPECIFICATION. THE LOG SHOWED THE SETTING TO BE AT 130 WHICH IS OUT OF SPECIFICATION. THE CUSTOMER IS UNABLE TO CONFIRM NO ERRONEOUS RESULTS HAVE BEEN REPORTED DUE TO LENGTH OF TIME OUT OF SPECIFICATION CONDITION HAS EXISTED. THE FE INFORMED THE CUSTOMER OF THE OUT OF SPEC FINDING. (B)(4).><THE FE PERFORMED ALL THE REQUIRED CAMERA ADJUSTMENTS. AFTER THE ADJUSTMENTS THE FE RECEIVED A CAM0 READING OF 104. THE FE RAN THE DIAGNOSTICS WITH NO ERRORS. THE CUSTOMER SUCCESSFULLY RAN AND ACCEPTED QC. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. | 0
4486424 | THIS INSTRUMENT HAS BEEN INVESTIGATED. ON (B)(6) 2015 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE. ON (B)(6) 2015 THE FE WAS ON SITE AND FOUND THE SAMPLE CAMERA IMAGE SHAKING. THE FE ORDERED A NEW SAMPLE CAMERA. ON (B)(6) 2015 THE FE WENT BACK ON SITE WITH A NEW SAMPLE CAMERA. THE FE REPLACED THE SAMPLE CAMERA AND VERIFIED IT WITH ST AND WADDIANA. THE FE TESTED THE SAMPLE CAMERA USING THE CUSTOMER SAMPLE BAR CODES, IT WORKS OK. THE CUSTOMER RAN QC AND ACCEPTED RESULTS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION.><THE CUSTOMER STATES THAT THE SAMPLE CAMERA WAS MISREADING BARCODE LABELS ON THE SAMPLE TUBES. A REVIEW OF THE ORTHO PROVUE LOG FILES DETERMINED THAT 7 SAMPLE TUBES WERE MISREAD. NO ERRONEOUS RESULTS WERE REPORTED. THIS IS 6 OF 7. SAMPLE TUBE 185547-1 MISREAD AS 185546-1. (B)(4). | 1
4841434 | LOG FILES WERE PROVIDED TO (B)(4) FOR REVIEW. DURING REVIEW, IT WAS IDENTIFIED THAT THE READER CAMERA BRIGHTNESS WAS OUT OF SPECIFICATION (CAMERA BRIGHTNESS SETTING AT 131 AND SHOULD BE BETWEEN 101 - 128). CAMERA OUT OF SPEC FROM (B)(6). (B)(4).><ON (B)(6) 2015 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND PROVIDED SECOND LEVEL ENGINEERING WITH BEFORE AND AFTER IMAGES OF THE REFERENCE CARD. ALSO RECORDED BRILLO READING OF 111. THE FE DID OPTICS ALIGNMENT AND ADJUSTMENTS, INCLUDING FINE ADJUSTMENT. THE FE GENERATED A NEW REFERENCE IMAGE. THE FE RAN CONTROLS TO VERIFY QC AND CORRECT INSTRUMENT OPERATION. CONTROLS AND QC VERIFIED AND ACCEPTED BY THE CUSTOMER. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. | 0
4751464 | THE CUSTOMER CONTACTED OCD TO REPORT THE DISCOVERY OF SWITCHED WASH TUBING LINES BETWEEN THE WASH SOL A AND WASTE CONTAINER. TESTING WAS PERFORMED WHILE THE WASH TUBING WAS CONNECTED IMPROPERLY. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4).><THE CUSTOMER SWITCHED THE LINES BACK TO THE PROPER CONNECTION AND REPEATED THE MAINTENANCE PROCEDURES. THE CUSTOMER DISCARDED THE OLD WASH A SOLUTION AND MADE A NEW BATCH. ALL TEST WERE REPEATED AND THE RESULTS WERE ACCEPTABLE. NO ERRONEOUS RESULTS WERE REPORTED AND THERE WAS NO BIOHAZARD EXPOSURE. THE CUSTOMER STATED ALL TESTING INCLUDING QC WAS ACCEPTABLE. | 0
4444924 | DURING A LUNG SIMULATION ON THE CYBERKNIFE, IT WAS NOTED THAT THE SIMULATION SOFTWARE YIELDED AN UNEXPECTED RESULT FOR THE TUMOR TRACKING METHOD. THE CYBERKNIFE SOFTWARE INDICATED THAT THE TUMOR TRACKING METHOD SHOULD BE A 1-VIEW A" INSTEAD OF THE MOST LOGICAL CHOICE OF "1-VIEW B". THE RESULT CONTRADICTED THE VISUAL ANALYSIS BY THE THERAPISTS | 1
4960304 | A PHYSICIST ANNUAL REVIEW OF A TOSHIBA FLUOREX IMAGING SYSTEM SHOWED A VERY HIGH RADIATION READING OF THE MAXIMUM OUTPUT OF THE FLUORO FUNCTION ON ALL MA STATIONS. TOSHIBA CAME IN TO WORK WITH OUR 3RD PARTY AND INTERNAL BIOMED ON (B)(6) 2015. (REPORT WAS NOT RECEIVED BY OUR FACILITY FROM PHYSICIST UNTIL END OF JUNE). TOSHIBA DETERMINED THAT MA CALIBRATIONS WERE OFF BY A SIGNIFICANT AMOUNT AND RE-CALIBRATED ALL. AN MAC (MA CONTROL BOARD) WAS REPLACED IN THE SYSTEM EARLIER IN (B)(6) 2015 AND CALIBRATION WAS COMPLETED AT THAT TIME PER TOSHIBA PUBLISHED PROCEDURES. NO CAUSE FOR THE MISCALIBRATION COULD BE DETERMINED. | 1
4369936 | THE POTENTIAL EXISTS FOR USER EXPOSURE TO BLOOD-BORNE PATHOGENS. TREATMENT MAY HAVE INCLUDED ANTI-VIRAL POST-EXPOSURE PROPHYLAXIS. IT IS NOT KNOWN IF ANY TESTING FOR (B)(6) WAS ORDERED OR PERFORMED. THE DECISION TO REPORT THIS ISSUE WAS MADE BECAUSE IT IS UNKNOWN IF ANTI-VIRAL PROPHYLAXIS WAS CONSIDERED NECESSARY BY THE USER'S EMPLOYER. THE COMPLAINANT TEXT STATED THAT A PRODUCT MALFUNCTION OCCURRED. THE ANALYZER CONTRIBUTED TO THE EVENT REPORTED. WASTE FLUID LIKELY IS CONTAMINATED WITH BIO-HAZARDOUS MATERIAL THAT MAY CONTAIN BLOOD-BORNE PATHOGENS. NO EVIDENCE, SUCH AS AN ERROR LOG OR PHOTOGRAPHS WERE SUBMITTED TO ILLUSTRATE THAT AN ACTUAL DEVICE FAILURE OCCURRED OR IF THE FLUID BEGAN TO LEAK BECAUSE A PIECE OF TUBING BECAME WORN AND SLIPPED OFF OF A CONNECTION. THIS INCIDENT IS AN ACCIDENTAL EXPOSURE. ANYONE PERFORMING MAINTENANCE ON EQUIPMENT IS REQUIRED TO FOLLOW ALL DIRECTIVE IN THE SYSMEX OPERATOR'S MANUALS AS WELL AS ANY OSHA REQUIREMENTS FOR PPE; HOWEVER, BECAUSE THE USER INCURRED AN EYE INJURY THAT HAD POTENTIAL FOR EXPOSURE TO BLOOD-BORNE PATHOGENS AND POTENTIALLY WAS GIVEN ANTI-VIRAL PROPHYLAXIS, THE INJURY IS CONSIDERED SERIOUS, AND CARE WOULD EXCEED SIMPLE FIRST-AID MEASURES.><(B)(4) WAS CONNECTED ON (B)(6) 2014 REGARDING AN INCIDENT OF POSSIBLE EXPOSURE AFTER A MALFUNCTION OCCURRED AT OR NEAR THE SOLENOID VALVE OF A CA-650 AUTOMATED BLOOD COAGULATION ANALYZER (SERIAL NUMBER (B)(4)) IN (B)(6). OVERFLOW DEVELOPED FROM THE WASHING SPITZ DURING THE PIPETTE WASHING SEQUENCE AND LIQUID WASTE BEGAN TO ACCUMULATE IN THE TRASH BOX WHERE USED REACTION TUBES NORMALLY FALL. THE USER WAS SPLASHED IN THE EYE WITH FLUID CONTAINING CA CLEAN I, DETERGENT (A SOLUTION CONTAINING SODIUM HYPOCHLORITE), AND LIQUID WASTE WHEN THE DISPOSAL BOX WAS REMOVED TO DISCARD THE USED REACTION TUBES. THE ISSUE OCCURRED ON (B)(6) AND WAS REPORTED TO SYSMEX AMERICA, INC. (SAI) ON (B)(6). IT WAS NOT KNOWN IF THE TECHNOLOGIST FLUSHED HIS/HER EYES WITH WATER FOR AT LEAST 15 MINUTES AS DIRECTED IN THE INSTRUCTIONS FOR USE (IFU), BUT DID RECEIVE MEDICAL ATTENTION. THE USER'S EYES BECAME RED AND SWOLLEN, AND THE USER WAS GIVEN AN OPTHALMIC DROP TO USE IN THE EYES, THE REDNESS AND SWELLING SUBSIDED THE FOLLOWING DAY. | 0
5103775 | THE CENTRAL MONITOR HAS 3 DISPLAYS THAT ARE USED TO SHOW PATIENT VITALS. DURING USE THE MOUSE SHOULD MOVE FROM EACH DISPLAY TO THE OTHER SEAMLESSLY. DURING THIS INCIDENT THE MOUSE BECAME STUCK ON THE FIRST DISPLAY AND THE TECH WAS UNABLE TO CONTROL OR USE THE OTHER 2 DISPLAYS. AT OTHER TIMES THE MOUSE CURSOR WOULD FREEZE IN PLACE AND NOT MOVE AT ALL. OTHER INSTANCES HAVE CAUSED THE KEYBOARD TO BECOME UNRESPONSIVE. MANUFACTURER RESPONSE FOR PHILIPS CENTRAL MONITOR, (BRAND NOT PROVIDED) (PER SITE REPORTER): PHILIPS HAS PERFORMED A SOFTWARE UPDATE AS WELL AS RELOCATING THE HARDWARE FOR THE SYSTEM TO MINIMIZE EMI PROBLEMS.><. | 1
4884222 | I HAVE REPORTED THIS BEFORE. INAPPROPRIATE CHOICES FOR MEDICATIONS ORDERING. MCKESSON PARAGON CPOE SOFTWARE USING FDB DATABASE. I HAVE REPORTED TO MCKESSON AND THEY SAY IT'S AN ISSUE WITH FDB. WILL UPLOAD THE SCREENSHOT. DIGOXIN ORAL ORDERING. SUGGESTS ONLY WEIGHT BASED DOSING. EVEN FURTHER DOWN THE LIST IT DOES EVEN NOT OFFER DIGOXIN 0.125 MG OR DIGOXIN 0.25 MG. (B)(6). | 1
5292436 | ><THE GE DOCTOR'S MONITOR THAT IS USED DURING THE PERFORMANCE OF INTERVENTIONAL RADIOLOGY (IR) GUIDED BIOPSIES HAS A DEFECT ON THE SCREEN THAT COULD POTENTIALLY RESULT IN AN ERROR OF WRONG SITE PROCEDURE. THE GE MONITOR CONTAINS FONTS THAT ARE SANS SERIF (WITHOUT FEET)- THIS CAUSES THE MONITORING SCREEN LETTERS AND NUMBERS TO BE CONFUSED. THE EXAMPLE THAT ALMOST CREATED AN ERROR A COUPLE OF WEEKS AGO, WAS THAT THE LETTER I" (STANDS FOR INFERIOR)AND THE NUMBER "1" HAVE THE SAME APPEARANCE ON THE SCREEN WHEN IT IS IN USE. APPROXIMATELY TWO WEEKS AGO | 1
4516353 | PT INITIALLY REPORTED TO THE HEART TRANSPLANT TREATMENT CENTER WITH A DRIVELINE FAULT" ALARM CONDITION SOME TIME IN (B)(6) AND WAS TOLD TO SWITCH TO HIS BACKUP CONTROLLER BY HOSPITAL PERSONNEL. UPON ARRIVAL TO THE CLINIC | 1
4962996 | REVIEW OF DATABASE LOGS WAS USED TO EVALUATE THE SOFTWARE AT THE HEALTH SYSTEM. THE COMPLAINT WAS INVESTIGATED THE DAY IT WAS REPORTED. ONCE THE DATE OF THE EVENT WAS CORRECTED TO BE (B)(6) 2011, THE AVAILABLE SYSTEM LOGS IDENTIFIED THAT THE SOFTWARE BEHAVED ACCORDING TO DESIGN PER THE PREDEFINED HR RULES. THE RULES FOR TRIGGERING AN HR ALERT ARE: 2 CONSECUTIVE 1-MIN HR VALUES > 150 OR 10-MIN HR > 150 OR 2 CONSECUTIVE 1-MIN HR VALUES < 40 OR 10-MIN HR < 40 OR 10-MIN HR VALUE > PT UPPER LIMIT (130) OR 10-MIN HR VALUE < PT LOWER LIMIT (50). THE 5 MINUTE AND 10 MINUTE LOGS INDICATE THAT DEVICE FUNCTIONED AS INTENDED. THE 1 MINUTE LOGS HAD BEEN PURGED OUT OF THE SYSTEM AND USER FACILITY DID NOT PROVIDE THE BACKUP LOGS FOR PHILIPS VISICU TO RESEARCH. THIS REPORT IS BEING FILED AS THE ECARE MANAGER DEVICE WAS USED DURING AN EVENT THAT WAS FOLLOWED BY A PT'S DEATH. ADDITIONAL INFO BECAME KNOWN ON 07/28/2015 DURING A RETROSPECTIVE REVIEW OF COMPLAINTS. PHILIPS VISICU ATTEMPTED TO CONFIRM THE PT'S DEATH WITH USER FACILITY, HOWEVER, THE USER FACILITY WAS UNABLE TO PROVIDE CONFIRMATION. THERE IS NO DATA THAT INDICATES THIS EVENT WAS RELATED TO A DEVICE DEFECT OR MALFUNCTION. ALL AVAILABLE INFO HAS BEEN PLACED ON FILE IN QUALITY ASSURANCE FOR APPROPRIATE TRACKING AND TRENDING.><PHILLIPS VISICU IDENTIFIED A REPORTABLE EVENT DURING A RETROSPECTIVE REVIEW OF COMPLAINTS. THE USER FACILITY REPORTED ON 12/13/2011 THAT THE SMART ALERTS IN ECARE MANAGER DID NOT FIRE ON ONE PT AND THEY SHOULD HAVE. THE USER FACILITY STATED THAT THE HR ALERT SHOULD HAVE FIRED AT 7:06 ON (B)(6) 2011. FURTHER INFO PROVIDED STATED THAT THE PT WAS CODING WHEN THE SMART ALERTS SHOULD HAVE GONE OFF. LATER THE USER FACILITY CONFIRMED THAT THE CORRECT DATE OF THE EVENT WAS (B)(6) 2011. ON 07/28/2015 DURING RETROSPECTIVE REVIEW, IT WAS IDENTIFIED THROUGH PHILIPS VISICU'S DATA LOGS THAT THE PT HAD EXPIRED SOMETIME AFTER THE INCIDENT. WE WERE UNABLE TO CONFIRM THE EXACT TIME. | 1
5149822 | THIS ISSUE OCCURRED DURING INTERNAL ENGINEERING TESTS. THIS ISSUE IS CURRENTLY BEING INVESTIGATED.><AN OCD EMPLOYEE REPORTED THAT THE ORTHOVISION MTS SOFTWARE FAILED TO MANAGE AN ALREADY USED DILUTION TRAY. AFTER DATABASE CLEANING WAS DONE THE DILUTION TRAY WAS REUSED. NO ERRONEOUS PATIENT RESULTS WERE REPORTED DUE TO THIS ISSUE. THIS ISSUE OCCURRED DURING INTERNAL ENGINEERING TESTS. (B)(4). | 1
4516363 | ><ACCUFLO (B)(4) LOT#100925 USED TO ADMINISTER A DOSE OF 5-FLUOROURACIL OVER 23 HOURS. THE DEVICE IS AN ELASTOMERIC INFUSION DEVICE WITH 4MG/HR RATE AND 120ML VOLUME. THE DEVICE WAS PLACED ON THE PT AT 1300 (B)(6) 2015 TO ADMINISTER THE DOSE OVER THE 23 HOURS. THE PT CALLED THE INFUSION CENTER ON (B)(6) 2015 STATING THAT THE DEVICE WAS EMPTY AT 0134 ON (B)(6) 2015. THE INFUSION SHOULD HAVE BEEN RUNNING UNTIL APPROX 1200 ON (B)(6) 2015. THE ENTIRE 23 HOUR INFUSION WAS INFUSED OVER APPROX 12 HOURS. | 1
5311386 | PATIENT INSTRUCTED ON PUMP USE TO MAINTAIN MILK SUPPLY. PATIENT DENIED PAIN. NURSE LEAVES THE ROOM TO ALLOW FOR PRIVACY. PATIENT NOTICES BLOOD IN MILK AND CALLS LACTATION CONSULTATION TO THE ROOM. LARGE BLOOD FILLED BLISTERS WERE NOTED ON THE PATIENT'S NIPPLE. PATIENT STATED SHE THOUGHT THE PAIN WAS NORMAL WITH PUMPING. DIAGNOSIS OR REASON FOR USE: BREASTFEEDING. EVENT ABATED AFTER USE STOPPED OR DOSE REDUCED: YES. | 0
5177087 | A DA VINCI HARMONIC CURVED SHEARS GRAY TORQUE WRENCH PIECE BROKE OFF AND THE SHEARS WOULD NOT OPEN. THE TORQUE WRENCH PIECE SHOULD NOT HAVE COME OFF AND THE SHEARS SHOULD HAVE OPENED. THE SHEARS WERE REPLACED AND THE CASE WAS COMPLETED.><. | 0
4516364 | PT REPORTED TO THE TRANSPLANT TREATMENT CLINIC COMPLAINING OF AN ALARM CONDITION HE RECEIVED FROM HIS POCKET CONTROLLER THAT TOLD HIM TO CONTACT THE HOSP". UPON BEING CONNECTED TO THE POWER MODULE AND BROWSING THE PT'S CONTROLLER HISTORY | 1
5336562 | THE CUSTOMER USED A SCENARIO WHICH WAS NOT INDICATED IN THE DEXIS USER MANUAL. A SOFTWARE EVALUATION WAS PERFORMED ON DEXIS 10.1.2. EVALUATION REVEALED THAT THE REPORTED ISSUE IS ONLY OBSERVED WHEN MULTIPLE INSTANCES OF DEXIS (PATIENT) ADMINISTRATION ARE ACTIVE, WHICH IS NOT CONSIDERED TO BE A NORMAL SCENARIO FOR THE USE OF DEXIS. THE ISSUE CAN BE RESOLVED BY RELOADING THE PATIENT LIST IN THE ACTIVE DEXIS ADMINISTRATION WINDOW. THE USER IS ABLE TO DETERMINE THE CORRECT PATIENT IMAGE AND NO INJURIES HAVE BEEN REPORTED. THIS CONCLUDES OUR INVESTIGATION.><THE CUSTOMER REPORTED THAT INCORRECT PATIENT IS DISPLAYED WHEN CLICKING DIRECTLY ON A PATIENT NAME IN THE DEXIS (PATIENT) ADMINISTRATION SCREEN. NO INJURIES HAVE BEEN REPORTED. | 1
4619894 | VESSEL SEALER #1 WAS IN USE DISSECTING THE VAGINAL MARGIN/CUFF WHEN ERROR APPEARED STATING BLADE EXPOSED | 0
4377201 | (2) DA VINCI GRASPING RETRACTORS WIRES BROKE DURING USE. NO INJURY TO PT. | 0
4826769 | ><THE SEER EVENT MONITORS HAVE A CONSISTENT PROBLEM WITH EITHER STAYING IN CALIBRATION MODE OR NOT TRANSMITTING. WE ARE HAVING REPEATED PROBLEMS WITH THESE DEVICES. THERE ARE NO COMMON VARIABLES REGARDING WHEN OR IF THE MONITORS WILL BECOME STUCK IN CALIBRATION MODE. ONCE IN CALIBRATION MODE AND STUCK, THE MONITOR BEEPS/ALARMS CONTINUOUSLY THAT IT IS IN CALIBRATION MODE. STAFF ARE UNABLE TO UN-STICK" THE MONITOR: BIOMED IS CONTACTED AND REMOVES/REPLACES THE MONITOR. THE MONITORS ARE THEN SENT BACK TO THE MANUFACTURER TO HAVE THEIR PASSWORD RESET. WE ARE SEEING THIS PROBLEM IN MULTIPLE MONITORS CONSISTENTLY FOR THE LAST THREE MONTHS. SOME DEVICES ARE LESS THAN ONE YEAR OLD AND STILL UNDER WARRANTY. OTHERS ARE OLDER. TO DATE | 1
5261154 | THE COMPLAINT OCCURRED DURING AN INSTALLATION PERIOD FOR THREE NEWLY PURCHASED UNITS. THE INITIAL UNIT WAS REMOVED FROM INSTALLATION AND THE OTHER TWO WERE WITHHELD FROM INSTALLATION UNTIL ISSUE WAS RESOLVED. WE WORKED WITH THE CUSTOMER TO DETERMINE THE CAUSE. CUSTOMER PROVIDED NETWORK DATA RECORDING TO ASSIST IN PROBLEM REPRODUCTION TO HELP IN DETERMINING THE CAUSE OF THE PROBLEM. CUSTOMER SET DATE OF (B)(6) 2015 TO GO TO THEIR FACILITY TO INVESTIGATE. DURING THIS VISIT WE VERIFIED FREEZE MALFUNCTION AND DETERMINED THE STATE OF THE UNIT DURING THE FREEZE PERIOD. WE RETURNED TO FACTORY WITH CUSTOMER UNITS. ALL CUSTOMERS WITH THIS DEVICE REVISION WERE CONTACTED TO SEE IF ANYONE ELSE WAS EXPERIENCING THIS BEHAVIOR. NO OTHER CUSTOMER EXPERIENCED THIS ISSUE TO DATE. THROUGH IN-DEPTH TESTING AND COMMUNICATION WITH MANUFACTURER OF OUR NETWORK CONTROLLER DEVICE, AS WELL AS WEB SEARCHES, WE DETERMINED THE ROOT CAUSE OF THE MALFUNCTION, A DEFECT IN THE DEVICE'S NETWORK CONTROLLER CHIP. IT TOOK 6 WEEKS TO DEVELOP AND VALIDATE A SOLUTION THAT COULD BE USED FOR A CORRECTION FOR NEWLY MANUFACTURED DEVICES. THE SOLUTION WAS VALIDATED AND THE CUSTOMER'S 3 UNITS WERE RETURNED ON 9/9/2015. VIDCO HAS NEVER HAD A DEVICE INVOLVED IN AND RESPONSIBLE FOR AN ADVERSE EVENT. WE WERE AND STILL ARE UNCLEAR WHETHER REPORTING THIS MALFUNCTION WAS NECESSARY, BUT DECIDED TO ACT ON THE SIDE OF CAUTION EVEN THOUGH WE WERE SEVERAL MONTHS PAST THE INITIAL DISCOVERY OF THE MALFUNCTION.><CUSTOMER AT (B)(6) HOSPITAL CALLED TO REPORT A DEVICE THEY WERE TRYING TO INSTALL WAS PERIODICALLY FREEZING UP AND COULD ONLY BE MADE OPERATIONAL AGAIN THROUGH CYCLING POWER TO THE DEVICE. | 1
5002087 | INVESTIGATION IS PENDING.><INVESTIGATION/CONCLUSION: THERE WERE NO FALSE POSITIVE MAMP, BZO, MTD, APAP, OPI OR THC RESULTS OBSERVED FROM ANY RETAIN DEVICES TESTED WITH TOX-DS 0.5X CALIBRATORS OR ON THE RETURNED TRIAGE METERPRO METER WITH TOX DS LEVEL 1 CONTROLS. THE MANUFACTURING RECORDS FOR THE LOT WERE REVIEWED AND THE LOT MET OUS RELEASE SPECIFICATIONS. THERE WAS NO SAMPLE RETURN, THEREFORE, SAMPLE SPECIFIC INTERFERENCE WAS UNABLE TO BE RULED OUT AS A POTENTIAL CAUSE FOR THE FALSE POSITIVE RESULTS. THERE WAS NO PRODUCT DEFICIENCY ESTABLISHED AND NO CORRECTIVE ACTION REQUIRED.><THE CUSTOMER ALLEGED POTENTIAL (B)(6) RESULTS ON THE TRIAGE TOX DRUG SCREEN PANEL ON A PATIENT IN (B)(6). THE TRIAGE TOX DRUG SCREEN WAS PERFORMED USING A FRESH URINE SAMPLE AND THE FOLLOWING RESULTS WERE POSITIVE: ACETAMINOPHEN (APAP), BENZODIAZEPINES (BZO), TETRAHYDROCANNABINOL (THC), METHAMPHETAMINES (MAMP), METHADONE (MTD), OPIATES (OPI). THE SERVICE WHERE THE PATIENT WAS HOSPITALIZED DISAGREED WITH THE RESULTS. THE SAME SAMPLE WHICH WAS STORED FOR A FEW HOURS WAS RETESTED ON THE SAME STRIP LOT AND ALL THE RESULTS WERE NEGATIVE. THERE WAS NO REPORTED ADVERSE PATIENT SEQUELA. THERE WAS NO ADDITIONAL INFORMATION PROVIDED. (NOTE: THE TRIAGE TOX DRUG SCREEN PANEL DEVICE IS NOT AVAILABLE IN THE UNITED STATES; HOWEVER, THIS MDR FILING IS DUE TO THE A SAME OR SIMILAR DEVICE BEING AVAILABLE IN THE UNITED STATES.) | 0
4637623 | THIS INCIDENT WAS CONSERVATIVELY REPORTED ON MARCH 27, 2015 AS A SERIOUS INJURY REPORT TYPE SINCE GATHERING INFORMATION FROM THE SITE TOOK TIME (LEGAL CASE), HOWEVER, THE ADDITIONAL INFORMATION RECEIVED THE SAME DAY CONFIRMED THAT THERE WAS NO INJURY, NO DELAY IN CARE OR MISDIAGNOSIS - THIS INCIDENT WAS DUE TO AN OPERATOR ERROR (SEE BELOW FOR MORE DETAILS): GE HEALTHCARE (GEHC) RECEIVED ADDITIONAL INFORMATION ON MARCH 27, 2015, AFTER INITIAL MDR WAS SUBMITTED, WHICH INDICATES THAT ON (B)(6) 2014 A PATIENT WITH SUSPECTED MALIGNANCY IN BOTH BREASTS UNDERWENT A MAMMOGRAPHY EXAM AT (B)(6), ON A (GEHC) SENO ESSENTIAL MAMMOGRAPHY UNIT. WORKLIST SERVER USED TO IMPORT PATIENT ON MAMMOGRAPHY SYSTEM WAS DOWN, HOWEVER, GE HEALTHCARE ACQUISITION SYSTEM WAS OPERATING PER SPECIFICATIONS, THEREFORE, THE RADIOLOGIST CREATED THE PATIENT'S NAME MANUALLY FOR HER IMAGE REPORT AND THIS HAS RESULTED IN AN INCORRECT PATIENT'S NAME. THE MISLABELED RADIOLOGY REPORT WAS SENT AND ARCHIVED INTO THE PACS SYSTEM. IN (B)(6) 2015 , WHEN THE PATIENT RETURNED FROM THE MAMMOGRAPHY EXAM AT (B)(6) HOSPITAL TO SEE HER SURGEON (THE PATIENT IS PRESUMED TO HAVE TAKEN A COPY OF HER RESULTS TO HER PHYSICIAN/SURGEON), THE SURGEON REVIEWED THE PATIENT'S REPORT AFTER WHICH HE RECOMMENDED AND SCHEDULED THAT THE PATIENT HAVE A REPEAT EXAM AT ANOTHER HOSPITAL ((B)(6)). IT APPEARS THAT THE PATIENT IS NOW BEING TREATED AT (B)(6) (THE POST EXAM TREATMENT IS UNKNOWN). THIS WOULD PRESUME THERE WOULD BE NO TREATMENT DELAY FOR THE PATIENT. IN (B)(6) 2015, THE RADIOLOGIST OF (B)(6) HOSPITAL DISCOVERED THE MISTAKE WHEN HE DID A ROUTINE FOLLOW UP TO SEE WHY THE PATIENT HAD NOT RETURNED TO THE HOSPITAL AND NOTIFIED GE HEALTHCARE. THIS TRIGGERED THE RADIOLOGIST TO REVIEW THE MAMMOGRAPHY IMAGES AND THEN THE ISSUE OF THE PATIENT NAME VS IMAGES CAME TO LIGHT WHICH COULD EXPLAIN THE DELAY BETWEEN EXAM DATE AND DISCOVERY OF THE MIX UP. IT IS CONFIRMED THAT THE DATE OF BIRTHDAY , PATIENT ID, EXAM DATE CORRESPONDS TO THE PATIENT THEREFORE, THE REPORT CORRESPONDS TO THE PATIENT , ONLY THE NAME IS WRONG. BASED ON THE ABOVE WORKFLOW, GE HEALTHCARE (B)(4) PERFORMED ON MARCH 27, 2015 CONCLUDED THAT THERE IS NO INJURY DUE TO THE WRONG PATIENT NAME SINCE ALL PLANNED EXAMS AND TREATMENT OCCURRED WITHOUT ANY DELAY OR ALLEGED MISDIAGNOSIS. MOREOVER, IF SUCH USE ERROR WERE TO REOCCUR IN ASSOCIATION WITH CORRECT PATIENT IDENTIFIERS, PATIENT DATE OF BIRTH, OTHER IMAGE EXAMS, KNOWN PATIENT HISTORY, PHYSICAL EXAM, AND PATIENT FEEDBACK TO CORROBORATE POSSIBLE PATIENT SYMPTOMS, A SERIOUS INJURY WOULD NOT BE LIKELY. GE HEALTHCARE INVESTIGATION HAS BEEN COMPLETED. INVESTIGATION ANALYSIS CONFIRMED THAT THE WRONG NAME IN THE REPORT IS A SITE SPECIFIC ISSUE DUE TO AN OPERATOR ERROR WHEN PATIENT WAS MANUALLY CREATED. THIS ERROR IS OBVIOUS TO THE USER SINCE THE NAME OF THE PATIENT IS DISPLAYED DURING ACQUISITION ON THE IMAGE. GOOD CLINICAL PRACTICES UNDERLINE THAT CLINICIANS, RADIOLOGISTS, AND MAMMOGRAPHIC TECHNICIANS ASK THE NAME OF THE PATIENT DURING THE EXAM AND REVIEW OF IMAGES TO CONFIRM THAT THE IMAGES CORRESPOND TO THE CORRECT PATIENT. ALL APPLICATION TRAININGS WERE PROVIDED TO TECHNICIANS AND PHYSICIST. A LETTER WAS SENT TO THE RADIOLOGIST OF (B)(6) ON MARCH 27, 2015 WHICH EVIDENCES THAT THE GE HEALTHCARE EQUIPMENT WAS FUNCTIONING ACCORDING TO SPECIFICATION ON THE DAY OF THE EXAMINATIONS. SEVERAL SAFETY RECOMMENDATION AND INSTRUCTIONS REGARDING THE CREATION OF A PATIENT WITH WORKLIST OR WITH A MANUAL ENTRY ARE CONTAINED WITHIN THE OPERATING MANUAL AND INDICATES BEFORE STARTING THE EXAMINATION | 1
5273138 | .><ENDO WRIST ONE VESSEL SEALER #410322; LOT M10150728 FAULTED DURING A PROCEDURE. BLADE JAMMED WHILE ATTEMPTING TO SEAL THIN TISSUE (TUBE). NON RECOVERABLE FAULT ALARMED, ALARM RESET, INSTRUMENT UNPLUGGED; RE-PLUGGED INTO UNIT, 2ND FAULT ALARM, INSTRUMENT REPLACED. | 0
5313658 | THE SUV VALUES THAT ARE BEING CALCULATED IN THE PET/CT FUSION TOOL ARE INCORRECT. THE ANOMALY PRESENTED ITSELF WHEN THE SPHERE ANNOTATION WAS DRAWN OVER THE IMAGE, THE VALUES FOR MINIMUM, AVERAGE, AND MAXIMUM WERE ALL EQUAL OR IN SOME CASES VERY CLOSE. THIS IS AN INDICATION THAT THE VALUES ARE INCORRECT. THE VOI (VOLUME OF INTEREST) ONLY REFLECTS THE SMALL REGION WHERE THE USER FIRST BEGAN DRAGGING THE RIGHT-MOUSE BUTTON TO DRAW THE SPHERE. THIS COULD LEAD TO MISDIAGNOSIS IN PET CASES WHERE THE SUV (MAXIMUM) WOULD DISPLAY A LOWER VALUE THAN IT SHOULD.><CUSTOMERS HAVE BEEN NOTIFIED OF THE MALFUNCTION AND HAVE BEEN GIVEN CLEAR INSTRUCTIONS ON HOW TO PREVENT THE MALFUNCTION FROM OCCURRING AS WELL AS INSTRUCTIONS FOR REQUESTING A FREE UPGRADE TO A CORRECTED VERSION OF THE SOFTWARE. | 1
5136210 | WE WERE RECENTLY MADE AWARE OF A VIRTUOSO SOFTWARE PRODUCT ISSUE THAT COULD POTENTIALLY IMPACT PATIENT SAFETY FOR CUSTOMERS USING VERSION 5.4 OR 5.5. THE SOFTWARE SCANS SLIDES AT 4X,10X, 20X, OR 40X MAGNIFICATIONS. FOR IMAGES SCANNED AT 4X, 10X, AND 40X MAGNIFICATION, THE LINEAR AND AREA MEASUREMENT TOOLS DISPLAY A VALUE THAT IS 1/5, ½ AND 2X THE ACTUAL MEASUREMENT, RESPECTIVELY. THE MEASUREMENT WILL BE DISPLAYED INCORRECTLY REGARDLESS OF WHAT MAGNIFICATION YOU ARE VIEWING THE IMAGE AT. FOR IMAGES THAT ARE SCANNED AT 20X MAGNIFICATION, THE MEASUREMENT TOOLS WORK PROPERLY IN ALL VERSIONS OF VIRTUOSO INCLUDING VERSIONS 5.4 AND 5.5. FOR CUSTOMERS IN THE EU/EEA AND CANADA WHO USE VIRTUOSO V5.4 AND V5.5 FOR PRIMARY DIAGNOSIS, THE MEASUREMENT TOOLS FOR SLIDES THAT WERE SCANNED AT 4X, 10X, OR 40X, COULD POTENTIALLY CAUSE INCORRECT ESTIMATION OF CLINICAL STAGE LEADING TO IMPROPER TREATMENT. MOST PATHOLOGY LABS SCAN AT 20X MAGNIFICATION, WHICH IS NOT IMPACTED BY THIS ERROR. USE OF THE MEASUREMENT TOOL IN THE VIRTUOSO SOFTWARE IS CONSIDERED 'PRIMARY DIGITAL DIAGNOSIS', WHICH IS NOT ON-LABEL IN THE UNITED STATES. THEREFORE, US-BASED LABORATORIES SHOULD PERFORM VALIDATION OF THE MEASUREMENT TOOL PRIOR TO USE OF THE SOFTWARE FOR CLINICAL CASES.><CUSTOMER FROM (B)(6) REPORTED WHEN USING SOFTWARE TO SELECT THE SAME AREA IN 2 BIF PICTURES SCANNED AT 20X AND 40X RESPECTIVELY, THE SIZE OF THE SCANNED TISSUE AS REPORTED BY THE SOFTWARE ARE DIFFERENT, DEPENDING ON SCAN MAGNIFICATION: AT 40X THE SIZE VALUE IS DOUBLE THE REAL SIZE. CLINICAL EVALUATION REGARDING CHANGES IN THERAPEUTIC MANAGEMENT FOR UP TO 18 PATIENTS HAS NOT BEEN PROVIDED; HOWEVER EARLY INFORMATION SPECIFIES CHANGES IN T OR N STAGING, LEADING TO VARIATIONS IN ASSESSING PROGNOSIS. MICROSCOPIC TUMOR MEASUREMENTS ARE MOST DIRECTLY TIED TO PATIENT STAGING AND THERAPEUTIC INTERVENTION IN THE DIAGNOSIS OF MELANOMA. | 1
5018020 | ><ULTRASOUND TECHNICIAN WAS PERFORMING A TRANSVAGINAL ULTRASOUND ON PATIENT. THE PATIENT COMPLAINED THAT SHE COULD FEEL AN ELECTRICAL CURRENT UPON FEELING THE TRANSDUCER TOUCHING HER SKIN. THE VOLUSON WAS REMOVED FROM SERVICE. A SERVICE CALL WAS PLACED TO GENERAL ELECTRIC. GENERAL ELECTRIC'S MODALITY SERVICE ENGINEER EXAMINED THE DEVICE. THE VOLUSON SYSTEM (MACHINE AND TRANSDUCER)PASSED ALL TESTS AND APPEARED TO BE IN PROPER WORKING CONDITION. NO UNTOWARD PATIENT EFFECTS. PER REPORTER'S RESPONSE FROM MANUFACTURER. CUSTOMER'S STATED PROBLEM COULD NOT BE REPRODUCED. DID AN ELECTRICAL SAFETY TEST ON THE MACHINE AND IT WAS WITHIN TOLERANCE, WITH GROUND RESISTANCE OF 0.05 OHMS AND CHASSIS LEAKAGE OF 2 MICROAMPS. DID LEAKAGE CURRENT TEST ON THE RIC5-9H PROBE AND THE PROBE WAS WITHIN TOLERANCE WITH THE VALUE OF 2 MICROAMPS. DID A VISUAL INSPECTION ON THE PROBE AND IT HAD NO MARKS OR CUTS ON THE MEMBRANE, STRAIN RELIEF, OR CABLE. CHECKED THE PINS ON THE CONNECTOR OF THE PROBE AND THEY WERE INTACT (NOT BENT). DID A FUNCTIONAL CHECKOUT WITH THE PROBE CONNECTED TO THE MACHINE AND IT FUNCTIONED PROPERLY. RESEATED THE CONNECTIONS TO THE MAIN MONITOR DISPLAY AND RAN THE MONITOR ADJUSTMENT PROCEDURES PER THE PSM DIRECTION KTI105990, REV 3, SECTION 6-3 AND THE MONITOR DISPLAY WAS WITHIN SPECIFICATION. CONSOLE PASSED ALL TESTS AND IS NOW READY FOR USE. | 0
4570218 | ><INTUBATED PATIENT IN ICU, ALERT, AWAKE AND ORIENTATED. ON (B)(6) 2015 THE SAGE COVERED YANKER (PRODUCT # 6914) WAS USED FOR THE PATIENT; THE PATIENT WAS SELF SUCTIONING ORAL SECRETIONS. ON (B)(6) 2015 THE PATIENT COMPLAINED OF GENERALIZED THROAT PAIN. ON (B)(6) /2015 THE PATIENT HAD A WITNESSED CARDIAC ARREST, DURING RESUSCITATION, ANESTHESIOLOGY ATTEMPTED RE-INTUBATION BUT WERE UNABLE TO ADVANCE THE ETT TUBE. BAG VALVE MASK OXYGENATION WAS BEING PROVIDED, AND O2 SATURATION IMPROVED FROM 59 PERCENT TO 70 PERCENT. DURING THE ETT EXCHANGE ANESTHESIOLOGY FOUND A CIRCULAR CANISTER HAD BROKEN OFF THE SAFE COVERED YANKER AND WAS LODGED AT THE BACK OF THE PATIENT CORD. THE DEVICE WAS REMOVED. PATIENT WAS RE-INTUBATED AND THE PATIENT'S OXYGENATION LEVEL RETURNED TO 96 PERCENT. DIAGNOSIS OR REASON FOR USE: TO FACILITATE REMOVAL OF ORAL SECRETIONS. | 0
4404575 | (B)(4). GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED.><INVESTIGATION REVEALED THE TRAM MODULE WAS APPROXIMATELY 8 YEARS OLD. THE TRAM WAS TESTED IN DEPOT REPAIR AND FOUND TO READ 60/80 MMHG ON BP2 WHEN CONNECTED TO A SIMULATOR SET TO 120/80 MMHG. BASED ON THE REPAIR INFORMATION PROVIDED, THE TRAM 451N DAS MODULE (P/N 417121-004) WAS REPLACED TO CORRECT THE REPORTED PROBLEM. FURTHER INVESTIGATION INTO THE REPORTED FAILURE COULD NOT BE COMPLETED AS THE DAS MODULE WAS SCRAPPED. THERE ARE NO CIC OR OTHER TRAM 451N LOGS FOR INVESTIGATION. ROOT CAUSE WAS IDENTIFIED TO BE A FAILED TRAM DAS MODULE BASED ON THE SERVICE INFORMATION PROVIDED BY DEPOT REPAIR. FURTHER INVESTIGATION INTO THE ROOT CAUSE COULD NOT BE COMPLETED SINCE THE FAILED DAS MODULE WAS SCRAPPED. CORRRECTION: DATE OF EVENT: WAS: UNKNOWN - IS: (B)(6) 2014.><THE CUSTOMER REPORTED THAT THE TRAM MODULE PROVIDED AN INACCURATE LOW INVASIVE BLOOD PRESSURE. THERE WAS NO PATIENT CONSEQUENCE AS THE ISSUE WAS DISCOVERED PRIOR TO PATIENT USE | 1
4612917 | THE CUSTOMER REPORTED A FAILURE OF THE DEFIB SYNC PORT ON THE PATIENT DATA MODULE. THE ISSUE WAS DISCOVERED DURING PREVENTATIVE MAINTENANCE. NO PATIENT INVOLVEMENT.><THERE WAS NO PATIENT INVOLVEMENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><INVESTIGATION REVEALED THE LOSS OF ECG DEFIB SYNC OUTPUT FUNCTION ON THE PDM DEFIB SYNC CONNECTOR WAS CAUSED BY FAILURE OF THE PDM DEFIB SYNC OUTPUT HARDWARE. THE DEFIB SYNC HARDWARE WAS REPLACED.. | 1
5332095 | BASED ON THE REPAIR TECHNICIAN STATEMENTS, INVESTIGATION DETERMINED THAT THE FAILURE WAS A PDM DAS CPU HARDWARE FAILURE. THE FAILED DAS HARDWARE WAS REPLACED.><THERE WAS NO PT INVOLVEMENT. DATE OF INCIDENT UNKNOWN AT THIS TIME. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN.><THE CUSTOMER REPORTED A LOSS OF FUNCTION OF THE DEFIB SYNC CONNECTOR OF THE DEVICE. THERE WAS NO PATIENT INVOLVEMENT. | 1
4950842 | THE COMPLAINT WAS RECEIVED FROM ONE CLINICAL CUSTOMER SITE IN (B)(6). THE CUSTOMER REPORTED A BACT/ALERT® FA PLUS BOTTLE (P/N 410851) THAT, UPON BEING ENTERED INTO THE OBSERVA SYSTEM, WAS LINKED TO ANOTHER PATIENT IN THE SYSTEM. THE EXACT LOT OF THE BOTTLES INVOLVED IN THIS EVENT WAS NOT AVAILABLE FROM THE CUSTOMER. THE POTENTIAL LOTS IN USE AT THE TIME OF THE EVENT WERE MANUFACTURED IN FEBRUARY, MARCH, AND APRIL OF 2015, AND THE CUSTOMER COMPLAINT WAS RECEIVED IN (B)(6) 2015. THE BOTTLE WAS ENTERED INTO THE OBSERVA SYSTEM ON (B)(6) 2015. HOWEVER, IT IS UNKNOWN IF THE BOTTLE WAS SCANNED OR MANUALLY ENTERED (TYPED) INTO THE OBSERVA SYSTEM FOR THE LOADING EVENT ON (B)(6) 2015. A PROBABLE SCENARIO OF THE DUPLICATE LABEL IDENTIFICATION NUMBER IS THAT, DURING THE ORIGINAL LOADING EVENT ON (B)(6) 2015, THE BOTTLE ID WAS MANUALLY TYPED INTO THE INSTRUMENT AND A TYPOGRAPHICAL ERROR WAS CREATED AT THAT TIME. THE EVENT WOULD NOT HAVE BEEN DISCOVERED, HOWEVER, UNTIL THE SECOND EVENT A WEEK LATER ON (B)(6) 2015 WHEN THE BOTTLE IMPRINTED WITH THAT ID ON THE BARCODE WAS SCANNED INTO THE OBSERVA SYSTEM. QUALITY SYSTEMS CONFIRMED THAT THE BOTTLE LABEL SUPPLIER HAS CONTROLS IN PLACE TO ASSURE THAT IN THE EVENT A DUPLICATE LABEL WAS GENERATED, THAT IT WOULD BE IDENTIFIED, PROPERLY QUARANTINED, AND NEVER SENT TO BIOMERIEUX. THE SUPPLIER SCANS 100% OF THE LABELS GENERATED AND RETAINS AN ELECTRONIC FILE OF EACH LABEL (IDENTIFICATION). THEIR LABEL GENERATION PROCESS, SOFTWARE APPLICATIONS AND QUALITY SYSTEM HAVE BEEN AUDITED BY BIOMERIEUX IN ORDER TO CONFIRM THAT THE VALIDATED STATE INDICATED IN THEIR INVESTIGATION RESPONSE MEETS BIOMERIEUX STANDARDS. A SEARCH OF THE SUPPLIER'S GENERATION FILES ASSOCIATED WITH THE LOTS IN QUESTION AS WELL AS THE INSPECTION FILES WERE SEARCHED TO DETERMINE IF THE BOTTLE ID IN QUESTION WAS PRODUCED IN DUPLICATE. THAT SEARCH DEEMED THAT THE BOTTLE ID WAS NOT PRINTED IN DUPLICATE. THIS COMPLAINT CANNOT BE CONFIRMED BASED ON THE INFORMATION AVAILABLE.><ON (B)(6) 2015, A CUSTOMER CONTACTED BIOMERIEUX TO REPORT THE SAME BOTTLE ID ON 2 BOTTLES OF BACT/ALERT® FA PLUS CULTURE BOTTLE. THE FIRST BOTTLE WAS LOADED AND COMPLETED TESTING WITHOUT INCIDENT AND WAS DISCARDED. THE 2ND BOTTLE WAS ENTERED INTO THE OBSERVA DATA MANAGEMENT COMPUTER, THE TECH REALIZED THAT THE BOTTLE ID MATCHED A PREVIOUS BOTTLE AND THEREFORE THIS NEW BOTTLE ATTEMPTED TO ATTACH TO THE PREVIOUS BOTTLE'S PATIENT INFORMATION. TO ENSURE THE BOTTLE ATTACHED TO THE PROPER PATIENT'S RECORD, THE 2ND BOTTLE WAS LOADED INTO 3D USING A GENERIC BOTTLE ID AND NOT THE DUPLICATE BOTTLE ID. WHEN REQUESTED TO PROVIDE SPECIFIC DETAILS REGARDING THE INCIDENT, THE CUSTOMER INDICATED NO DEATH NOR INJURY WAS ASSOCIATED WITH THIS EVENT. AN INVESTIGATION INTO THIS INCIDENT HAS BEEN INITIATED WITHIN BIOMERIEUX. | 1
4824086 | (B)(6) (A PROVINCIAL BLOOD BANK IN (B)(6)) INFORMED (B)(4) OF A FALSE NEGATIVE RESULT IN ASSOCIATION WITH THE BACT/ALERT BPA CULTURE BOTTLE (REF. (B)(4), LOT 1040464, EXP. 21JUL15). A PLATELET POOL SAMPLE WAS INOCULATED TO THE BACT/ALERT BPA CULTURE BOTTLE ON (B)(4) 2015. ON (B)(4) 2015 (THREE DAYS INTO THE 7-DAY BACT/ALERT BPA CULTURE BOTTLE TEST), THE PLATELET POOL ASSOCIATED WITH THE BACT/ALERT BPA CULTURE BOTTLE WAS TRANSFUSED TO A PATIENT. THE PATIENT EXPERIENCED A TRANSFUSIONAL REACTION AS A CONSEQUENCE OF THE TRANSFUSION; THE TRANSFUSION WAS STOPPED. THE PATIENT RECEIVED TREATMENT FOR THE REACTION AND EXPERIENCED NO FURTHER INCIDENT. THOUGH THE HOSPITAL STAFF STATES A BACILLUS SPECIES WAS IDENTIFIED IN THE TRANSFUSION BAG AND IN A BLOOD SPECIMEN COLLECTED FROM THE PATIENT, THE BACT/ALERT BPA CULTURE BOTTLE WAS NEGATIVE AT 7 DAYS. AN INVESTIGATION WILL BE INITIATED BY BIOMÉRIEUX. HOWEVER,A PLATELET POOL SAMPLE IS UNAVAILABLE FOR TESTING.><DEVICE NOT RETURNED TO MANUFACTURER.><DUE TO PROHIBITIVE SHELF-LIFE, THE CUSTOMER DID NOT SUBMIT PLATELET POOL ISOLATES FOR INVESTIGATIONAL TESTING. A BACT/ALERT DATA BACK-UP FILE WAS OBTAINED FROM THE CUSTOMER. THE BOTTLE HISTORIES AND READINGS FOR THE BACT/ALERT BPA BOTTLE WERE EXTRACTED FROM THE DATA BACKUP. THE HISTORIES WERE NORMAL AND SHOWED THE BOTTLE FLAGGED NEGATIVE AFTER 7 DAYS. THE READINGS WERE INTERPOLATED WITH AN ALGORITHM TOOL, AND WERE DETERMINED NEGATIVE. STATISTICAL EVALUATION OF THE BOTTLE READINGS INDICATES THAT BOTH THE FAST AND SLOW SLOPE MEASUREMENTS SHOW AN INITIAL SPIKE BUT THEN DIMINISHED OVER TIME, AND DID NOT ACCUMULATE ENOUGH MOMENTUM FOR A SIGNAL. THE INVESTIGATION EXAMINED THE BACT/ALERT BPA LOT 1040464 MEDIA FORMULATION, FILLING, AND PACKAGING RECORDS, INCLUDING THE QUALITY CONTROL ACCEPTABLE QUALITY LIMIT INSPECTION DOCUMENTATION AND ALL THE RESULTS WERE WITHIN SPECIFICATION. RETAINED SAMPLES FROM BACT/ALERT BPA LOT 1040464 WERE CHALLENGED IN AN INTERNAL STUDY TO CONFIRM RECOVERY OF B. CEREUS ATCC 11778. THE BACT/ALERT BPA LOT 1040464 PERFORMED AS EXPECTED ACCORDING TO SPECIFICATION. THE BACT/ALERT BPA INSTRUCTIONS FOR USE (IFU) WERE REVIEWED. THE PLATELET SPECIMEN COLLECTION AND PREPARATION SECTION STATES THE LEUKOCYTE REDUCED PLATELET SPECIMEN MUST BE COLLECTED USING STERILE PROCEDURES TO REDUCE THE RISK OF CONTAMINATING THE SAMPLING SITE AND SAMPLING SITE COUPLER. THE PLATELET SPECIMEN SHOULD BE TAKEN AT LEAST 24 HOURS AFTER COLLECTION TO ALLOW FOR NATURAL PROLIFERATION IN THE PLATELET PRODUCT." IN THE PROCEDURAL NOTES AND PRECAUTIONS SECTION OF THE IFU | 0
4577220 | .><THE INCIDENT OCCURRED AT THE (B)(6) BD BIOSCIENCES FACSCANTO CLINICAL SOFTWARE COMPLAINT (B)(4). A (B)(6) PT RECEIVED A T-CELL DEPLETED HSCT AND ENGRAFTED WITH A VERY LOW T CELL COUNT. DURING FOLLOW UP A SAMPLE (UNCLEAR WHETHER MARROW ASPIRATE OR PB) WAS GATED AUTOMATICALLY BY THE CANTO CLINICAL SOFTWARE SUCH THAT A VERY HIGH LEVEL OF T CELL POSITIVITY WAS DEMONSTRATED. THIS TURNED OUT TO BE AN INCORRECT RESULT AS THE PT CONTINUED TO HAVE LESS THAN 1% T CELLS. UPON REVIEW, A VERY SMALL CD 3+ SUBSET WAS SEEN; HOWEVER, THE SOFTWARE ALGORITHM DID NOT RECOGNIZE THIS SMALL POPULATION AS CD3+. INSTEAD, IT MISTAKENLY DREW THE GATE IN THE MIDDLE OF A CD3- POPULATION RESULTING IN A FALSELY ELEVATED CD3+ POPULATION. ACCORDING TO DR (B)(6), THE ATTENDING MD FOR THE FLOW LAB, THERE WAS NO ADVERSE CLINICAL IMPACT TO THE PT AS A CONSEQUENCE OF THIS ERRONEOUS RESULT. THE INCORRECT RESULT WAS MANUALLY RE-GATED IN DIVA TO REVEAL THE CONTINUED PRESENCE OF T-CELL LYMPHOPENIA. SPECIFICALLY, ON (B)(6) 2014 AND AGAIN ON (B)(6) 2014, THE BD FACSCANTO CLINICAL SOFTWARE CATEGORIZED MOST CD3 LYMPHOCYTE POPULATION AS POSITIVE ON THE PT SAMPLES. THE RESULT DID NOT MATCH PT MEDICAL HISTORY. AFTER MANUAL ADJUSTMENT, THE CD3% CHANGED FROM 92.91% TO 0.66%. THE NEGATIVE CD3% WAS CONFIRMED BY USING BD FACS DIVA SOFTWARE. | 1
4884460 | (B)(4). THE FOLLOWING CORRECTIVE AND PREVENTIVE ACTION HAS BEEN TAKEN: ABBOTT DIAGNOSTICS DIVISION CONTACTED THE FOUR IMPACTED CUSTOMERS TO EXPLAIN THE SITUATION AND INSTRUCTED THEM TO IMMEDIATELY DISCONTINUE USE OF THESE ARCHITECT C8000 INSTRUMENTS. ALL IMPACTED CUSTOMERS ACKNOWLEDGED RECEIPT AND UNDERSTANDING OF THIS ISSUE ON MAY 29, 2015. ALL FOUR IMPACTED CUSTOMERS RECEIVED REPLACEMENT ARCHITECT C8000 ANALYZERS.><THE ARCHITECT C8000 ANALYZER, SERIAL NUMBER (B)(4), AT THE CUSTOMER SITE WAS REPLACED DUE TO TUBING IDENTIFIED AS HAVING AN INCORRECT DIAMETER. NO INJURY OR IMPACT TO PATIENT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. A PRODUCT DEFICIENCY WAS IDENTIFIED AND FIELD ACTION (B)(4) WAS TAKEN TO ADDRESS THE ISSUE. | 0
5054210 | AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><NO ARCHITECT INSTRUMENT PROBLEMS WERE NOTED. AFTER REPLACING THE PRE-TRIGGER REAGENT WITH ANOTHER NEW BOTTLE OF THE SAME LOT AND FLUSHING THE SYSTEM, THE ISSUE WAS RESOLVED. ALL ASSAY CONTROLS WERE WITHIN RANGE AND ALL RETESTED PATIENT SAMPLES RETURNED EXPECTED RESULTS. HISTORICAL COMPLAINT SEARCHES DETERMINED THAT THERE IS NO UNUSUAL ACTIVITY FOR THE LIKELY CAUSE LOT. THIS IS THE ONLY COMPLAINT AGAINST THE LOT. COMPLAINT TRENDING REPORT REVIEW DETERMINED THAT THERE IS NO NON STATISTICAL OR ADVERSE TREND FOR THE ISSUE FOR THE PRODUCT. NO TESTING OF THE LIKELY CAUSE LOT WAS REQUIRED AS A NEW BOTTLE OF THE SAME LOT PERFORMED AS EXPECTED AND SHOWED NO ISSUES AT THE CUSTOMER SITE. ALSO NO OTHER COMPLAINTS HAVE BEEN RECEIVED FOR THE PRE-TRIGGER LOT. THIS INDICATES THE ISSUE OBSERVED MAY BE ATTRIBUTED TO CONTAMINATION / MISHANDLING. THE ARCHITECT SYSTEMS OPERATIONS MANUAL PROVIDES INSTRUCTIONS ON REPLACEMENT OF THE PRE-TRIGGER AND ADVISES THAT RESULTS AND INVENTORY STATUS CAN BE ADVERSELY AFFECTED IF NOT LOADED CORRECTLY, INCLUDING MISALIGNMENT OF THE LEVEL SENSOR. A REVIEW OF THE MANUFACTURING DOCUMENTATION DID NOT IDENTIFY ANY ISSUES ASSOCIATED WITH THE CUSTOMER OBSERVATION. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS COMPLAINT INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED. THERE IS NOT ENOUGH INFORMATION TO REASONABLY SUGGEST A MALFUNCTION AS REPLACEMENT WITH A NEW BOTTLE OF THE SAME LOT RESOLVED THE ISSUE, BUT NO EVIDENCE OF WHY THE FIRST BOTTLE FAILED. A REVIEW OF LABELING CONCLUDED THAT THE ISSUE IS SUFFICIENTLY ADDRESSED.><THE CUSTOMER OBSERVED DISCREPANT RESULTS AFTER REPLACING THE PRE-TRIGGER SOLUTION ON THE ARCHITECT I2000SR ANALYZER. INITIAL NEGATIVE TOTAL B-HCG RESULTS OF <1.20 MIU/ML WERE GENERATED ON (B)(6) 2015 FOR TWO PATIENTS, SAMPLE ID (B)(4), THAT RETESTED POSITIVE AT 140.171 AND 81.18 MIU/ML RESPECTIVELY ON (B)(6) 2015. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 1
5271464 | FIELD SERVICE REPLACED THE INCORRECT TUBING ON THE ARCHITECT I2000SR ANALYZER PER (B)(4). THERE WAS NO IMPACT TO PATIENT RESULTS.><(B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED ON ALL IMPACTED I2000 SR INSTRUMENTS. THE MANDATORY TSB INSTRUCTS FIELD SERVICE TO REPLACE THE INCORRECT WASTE SYSTEM, VACUUM SYSTEM, AND SOME BUFFER SYSTEM TUBINGS. | 0
4789427 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND INSTRUMENT SERVICE. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE: A PRODUCT LABELING REVIEW FOUND THE ARCHITECT SYSTEM OPERATIONS MANUAL ADDRESSES INFORMATION FOR REGARDING ELECTRICAL HAZARDS AND ADDRESSES PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR ERROR CODE 3000 INCLUDING A POOR CONNECTION OR FAILING ANTENNAE BOARD. A REVIEW FOUND THE SAFETY HAZARDS MEMO CONTAINS ADEQUATE INFORMATION NOTING ABBOTT DIAGNOSTIC EQUIPMENT AND ACCESSORIES ARE CERTIFIED TO THE APPROPRIATE SAFETY STANDARDS, AND ADEQUATE PROTECTION IS PROVIDED FOR THE OPERATOR AGAINST SPREAD OF FIRE FROM THE EQUIPMENT. PART R1/R2 LLS ANTENNA BOARD (PART NUMBER (B)(4)) MALFUNCTIONED AND WAS REPLACED AND THE INSTRUMENT IS FUNCTIONING AS INTENDED. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION NO PRODUCT DEFICIENCY WAS IDENTIFIED.><THE CUSTOMER OBSERVED EVIDENCE OF FIRE UNDER THE ANTENNA ON THE ARCHITECT I2000SR ANALYZER. THERE WERE NO INJURIES. THE CUSTOMER ALSO EXPERIENCED ERROR CODE 3000 UNABLE TO PROCESS TEST, LIQUID NOT FOUND FOR (R2 PIPETTOR) AT (REAGENT BOTTLE POSITIONS).><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). | 1
5141813 | AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><THE CUSTOMER STATED THAT AN ARCHITECT MAGNESIUM RESULT OF >3.9 MMOL/L WAS GENERATED FOR A PATIENT SAMPLE ((B)(4)) THAT UPON RETESTING GENERATED A NORMAL RESULT (EXACT VALUE NOT PROVIDED). THE CUSTOMER'S NORMAL RANGE IS 0.66 TO 1.07 MMOL/L. FIELD SERVICE IDENTIFIED A LEAKING PROBE ON THE CUVETTE WASHER. CONTROL VALUES AND PRECISION TESTING WAS WITHIN SPECIFICATIONS AFTER SERVICE REPLACED THE PROBE. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><THE ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO TROUBLESHOOT THE ISSUE. UPON INSPECTION OF THE ANALYZER, THE FSR OBSERVED THAT CUVETTE WASH NOZZLE #5 HAD BEEN LEAKING. THE FSR REPLACED THE NOZZLE, AND THEN INITIATED A CUVETTE WASH. THERE WAS NO EVIDENCE OF LEAKING OR DRIPPING AFTER THE NOZZLE WAS REPLACED. THE FSR THEN RECALIBRATED MAGNESIUM AND PERFORMED A PRECISION RUN TO VERIFY INSTRUMENT PERFORMANCE. NO ADDITIONAL ERRATIC MAGNESIUM RESULTS WERE REPORTED ON ARCHITECT (B)(4). A SERVICE TICKET REVIEW FOR (B)(4) REVEALED NO ADDITIONAL ERRATIC OR DISCREPANT RESULTS, NOR DID IT IDENTIFY ANY SERVICE OR COMPLAINT ISSUES THAT MAY HAVE CONTRIBUTED TO THIS ISSUE. QUALITY METRICS WERE REVIEWED. NO SYSTEMIC ISSUES OR ADVERSE TRENDS FOR THE ISSUE ASSOCIATED WITH THIS TICKET. (ERRATIC / DISCREPANT RESULTS) WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE ARCHITECT SERVICE AND SUPPORT MANUAL PROVIDE ADEQUATE INFORMATION, TROUBLESHOOTING, AND MAINTENANCE CONCERNING ERRATIC/ DISCREPANT RESULTS. BASED ON THE AVAILABLE INFORMATION, A DEFICIENCY OF THE SYSTEM WAS NOT IDENTIFIED. THERE IS NO EVIDENCE TO REASONABLY SUGGEST A MALFUNCTION OCCURRED FOR THE CUVETTE WASH NOZZLES. THE ISSUE WAS RESOLVED BY THE FIELD SERVICE REPRESENTATIVE THROUGH STANDARD TROUBLESHOOTING PROCEDURES. | 0
5036558 | AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><A RETURN WAS NOT REQUIRED NOR REQUESTED FOR THIS EVALUATION AS THE ABBOTT FIELD SERVICE ENGINEER (FSE) PROVIDED SUFFICIENT INFORMATION REGARDING THE DAMAGED EXTERNAL WASTE PUMP. A REVIEW OF THE INFORMATION FROM THE CUSTOMER SITE DOCUMENTS THAT THE EXTERNAL WASTE PUMP (EWP) WAS PLUGGED INTO AN ELECTRICAL OUTLET AND PERSONAL PROTECTIVE EQUIPMENT (PPE) WAS NOT BEING USED. THE CUSTOMER WAS ATTEMPTING TO TROUBLESHOOT THE EWP DUE TO DIFFICULTY EMPTYING AND THE AUDIBLE ALARM SOUNDING. SUBSEQUENTLY THE CUSTOMER OBSERVED THE FLOAT SENSOR APPARATUS WAS BROKEN FROM WHERE IT CONNECTS TO THE PUMP HOUSING EXPOSING THE INTERNAL WIRES. AN ABBOTT FIELD SERVICE ENGINEER (FSE) EXAMINATION FOUND NO BARE WIRES. THE CUSTOMER WAS UNABLE TO STATE THE CONDITION OF THE FLOAT SENSOR APPARATUS PRIOR TO OR AFTER THEY CLEANED THE FLOATS. NEITHER THE EWP INSTRUCTION MANUAL NOR THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES OPERATOR/CUSTOMER INSTRUCTIONS FOR OPENING AND/OR CLEANING THE FLOAT SENSORS. THE EWP INSTRUCTION MANUAL PROVIDES THE OPERATOR/CUSTOMER WITH TECHNICAL SPECIFICATIONS, ELECTRICAL SAFETY PARAMETERS, AND PUMP INSTALLATION INSTRUCTIONS. THE TECHNICAL ASSISTANCE SECTION STATES CALL YOUR AREA ABBOTT CUSTOMER SUPPORT CENTER IF YOU CANNOT CORRECT AN OBSERVED PROBLEM, IF THE PROBLEM OCCURS REPEATEDLY, OR IF YOU OBSERVE ABNORMAL OPERATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES THE OPERATOR INFORMATION FOR MAINTENANCE, COMPONENT REPLACEMENT (INCLUDING THE EXTERNAL WASTE PUMP), AND TROUBLESHOOTING THE REPORTED ISSUE. THE EXTERNAL WASTE PUMP INSTRUCTION MANUAL (PART NUMBER 90297-101) PROVIDES TECHNICAL SPECIFICATIONS, ELECTRICAL SAFETY PARAMETERS, AND PUMP INSTALLATION INSTRUCTIONS. THE TECHNICAL ASSISTANCE SECTION STATES CALL YOUR AREA ABBOTT CUSTOMER SUPPORT CENTER IF YOU CANNOT CORRECT AN OBSERVED PROBLEM, IF THE PROBLEM OCCURS REPEATEDLY, OR IF YOU OBSERVE ABNORMAL OPERATION. THE ARCHITECT C8000 SERVICE MANUAL PROVIDES SERVICE AND MAINTENANCE PROCEDURES TO ADDRESS THE CURRENT CUSTOMER ISSUE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. BASED ON THE RESULTS OF THIS EVALUATION AND INFORMATION FROM THE CUSTOMER SITE, A SPECIFIC CAUSE OF THE SHOCK AND/OR THE DAMAGE TO THE FLOAT SENSOR APPARATUS COULD NOT BE DETERMINED. THE ISSUE WAS RELATED TO A USE ERROR AS THE OPERATOR/CUSTOMER DID NOT USE PPE OR UNPLUG THE POWER CORD PRIOR TO SERVICING THE EWP. THERE IS INSUFFICIENT INFORMATION TO INDICATE A MALFUNCTION OR A DEFICIENCY OF THE EXTERNAL WASTE PUMP OR THE ARCHITECT C8000 ANALYZER.><THE CUSTOMER REPORTS THAT AN EMPLOYEE AT THEIR SITE RECEIVED AN ELECTRICAL SHOCK WHILE TROUBLESHOOTING AN ISSUE WITH AN EXTERNAL WASTE PUMP CONNECTED TO AN ARCHITECT C8000 ANALYZER. THE EMPLOYEE WAS CLEANING THE FLOATS ON THE PUMP AND CAME INTO CONTACT WITH BARE ELECTRICAL WIRES AT THE FLOAT ASSEMBLY. THERE WERE NO INJURIES SUSTAINED BY THE OPERATOR RESULTING FROM THIS ISSUE. THE PUMP ASSEMBLY FOR THIS ANALYZER HAS BEEN REPLACED. THERE HAVE BEEN NO FURTHER ISSUES REPORTED WITH SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS ACCEPTABLE. | 0
5114556 | FIELD SERVICE REPLACED THE INCORRECT TUBING ON THE ARCHITECT I2000SR ANALYZER PER THE MANDATORY TECHNICAL SERVICE BULLETIN (TSB 116-132). THERE WAS NO IMPACT TO PATIENT RESULTS.><(B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED ON ALL IMPACTED I2000 SR INSTRUMENTS. THE MANDATORY TSB INSTRUCTS FIELD SERVICE TO REPLACE THE INCORRECT WASTE SYSTEM, VACUUM SYSTEM, AND SOME BUFFER SYSTEM TUBINGS. | 0
4348319 | CUSTOMERS EXPERIENCED REAGENT CARRYOVER DUE TO PROTEIN BUILDUP IN THE R2 PROBE. IN SOME CASES, CUSTOMERS EXPERIENCED DISCREPANT ELEVATED MICROALBUMIN RESULTS DUE TO UNEXPECTED CREATININE R2 CARRYOVER. NO INCIDENTS OF IMPACT TO PATIENT MANAGEMENT WERE REPORTED.><THERE WERE A TOTAL OF 203 COMPLAINTS RECEIVED FOR THIS PRODUCT PROBLEM WHITH THE FIRST COMPLAINT RECEIVED ON (B)(6) 2005 AND THE LAST COMPLAINT RECEIVED ON (B)(6) 2012. THERE WERE NO DEATHS OR SERIOUS INJURIES (ADVERSE EVENTS) REPORTED.><(B)(4) - NO CONSEQUENCES OR IMPACT TO PATIENT. HIGH TEST RESULTS: A PRODUCT DEFICIENCY WAS IDENTIFIED AND PRODUCT INFORMATION LETTERS WERE RELEASED GUIDING THE CUSTOMERS THROUGH EXISTING PROCEDURES TO DETECT AND TROUBLESHOOT POTENTIAL CARRYOVER. THE INSTALLATION OF ARCHITECT SYSTEM SOFTWARE VERSION 8.0 VIA TECHNICAL SERVICE BULLETINS, (ADDING MAINTENANCE AND DIAGNOSTIC PROCEDURES) PROVIDED CUSTOMERS WITH ADDITIONAL MAINTENANCE PROCEDURES FOR PROBE CLEANING TO MINIMIZE POTENTIAL CARRYOVER. THERE WERE A TOTAL OF (B)(4) COMPLAINTS RECEIVED FOR THIS PRODUCT PROBLEM WITH THE FIRST COMPLAINT RECEIVED ON (B)(4) 2008 AND THE LAST COMPLAINT RECEIVED ON (B)(4) 2010. THERE WERE NO DEATHS OR SERIOUS INJURIES (ADVERSE EVENTS) REPORTED. | 0
5048112 | (B)(4). A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED AND PERFORMED PROBE PRECISION TESTS, REVIEWED INSTRUMENT LOGS AND OBSERVED AN ISSUE WITH THE BHCG CARRYOVER, PERFORMED DECONTAMINATION PROCEDURES WITH A PASSING CARRYOVER TEST. NO REOCCURRENCE OF BHCG CARRYOVER HAS BEEN OBSERVED SINCE THE FSR SITE VISIT. THE FSR DETERMINED THE BHCG RESULT ISSUE WAS CAUSE BY CARRYOVER FROM THE STAT PROBE WHICH WAS CONTAMINATED/DIRTY. THE ISSUE WAS RESOLVED BY THE FSR PERFORMING AS NEEDED MAINTENANCE PROCEDURE, FLUSH FLUIDS. THE ARCHITECT ISR07687 RESULT LOG WAS REVIEWED AND CONFIRMED THE INITIAL STAT BHCG (ASSAY ID 30) RESULTS AS SPECIFIED IN THE COMPLAINT TEXT FOR SAMPLE IDENTIFICATION (SID) (B)(4). THE RESULT LOG SHOWS THESE SAMPLES WERE RETESTED FOR ROUTINE BHCG (ASSAY ID 651) INSTEAD OF STAT BHCG, AND ALL THE RESULTS WERE <7000.00 MIU/ML (DILUTION 1:15). ADDITIONALLY, THE RESULT COMMENTS CONFIRMED THAT SERUM SAMPLES WERE USED FOR THE STAT BHCG TESTING. A REVIEW OF THE 12 MONTH SEARCH FOR SIMILAR COMPLAINTS DID NOT IDENTIFY AN ADVERSE TREND OF THE PROBE. A REVIEW OF THE ARCHITECT PRODUCT MONITORING IDENTIFIED NO TRIGGER ASSOCIATED WITH THE ARCHITECT I2000SR ERRATIC RESULT RATE WITH NO TRENDS WERE IDENTIFIED FOR THE I2000SR. THE BHCG PACKAGE INSERTS ADDRESS SAMPLE HANDLING/ PERFORMANCE CHARACTERISTICS, INCLUDING THE RECOMMENDATION OF PLASMA SAMPLES FOR STAT TESTING. THE PACKAGE INSERT ALSO ADDRESSES HOW A DISPLAYED RESULT OF < 7000.00 MIU/ML (1:15 AUTOMATED DILUTION PROTOCOL) INDICATES THE NEED TO RETEST THE SAMPLE AT A LOWER DILUTION OR UNDILUTED. THE RESULT AND INTERPRETATION SHOULD NOT BE REPORTED. THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES INFORMATION REGARDING TROUBLESHOOTING OF THE DESCRIBED ISSUE. PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR ERRATIC ASSAY RESULTS INCLUDE BUT ARE NOT LIMITED TO: THE PROBE IS DIRTY OR PARTIALLY CLOGGED WITH THE CORRESPONDING CORRECTIVE ACTION OF PERFORMING AS NEEDED MAINTENANCE PROCEDURE, FLUSH FLUIDS. THE ARCHITECT I2000SR IS PERFORMING ACCEPTABLY.><THE ACCOUNT GENERATED FALSE POSITIVE BHCG RESULTS ON 7 PATIENTS WHO REPEATED NEGATIVE WHEN PROCESSING ON THE ARCHITECT I2000SR ANALYZER. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. DATA PROVIDED: (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 64.2 MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 53.3 MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 25 MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 52.1MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 86.1MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 61.0 MIU/ML, REPEAT NEGATIVE OF 4.5 MIU/ML. (B)(6), INITIAL POSITIVE ARCHITECT BHCG OF 65.4 MIU/ML, REPEAT NEGATIVE OF <1.2 MIU/ML. | 0
4474841 | THE ABBOTT FIELD SERVICE DISCOVERED THE INCOMING WATER WAS TOO COLD, ABOUT 11 DEGREES C. THE CUSTOMER WAS ADVISED TO HAVE A MIXING VALVE INSTALLED TO BLEND COLD AND HOT WATER TO ACCEPTABLE TEMPERATURE. THE ISSUE WAS RESOLVED AFTER INSTALLING THE VALVE. A REVIEW OF THE INSTRUMENT SERVICE TICKET HISTORY DID NOT IDENTIFY ANY OTHER FACTORS THAT MAY HAVE CONTRIBUTED TO THE ERRATIC RESULT. A HISTORICAL DATA REVIEW DID NOT IDENTIFY ANY SPECIFIC ISSUES THAT REQUIRED FURTHER ACTION, AND NO ADVERSE TRENDS FOR THE ARCHITECT C4000 WERE IDENTIFIED. THE CURRENT LABELING WAS REVIEWED AND PROVIDES ADEQUATE ASSISTANCE FOR TROUBLESHOOTING ERRATIC RESULTS. THE ISSUE WAS RESOLVED BY INSTALLING A MIXING VALVE TO BRING THE INCOMING WATER SUPPLY TEMPERATURE WITHIN THE RANGE SPECIFIED IN THE ARCHITECT C4000 PRE-INSTALLATION CHECKLIST. BASED ON THE EVALUATION, THE ARCHITECT C4000 IS PERFORMING ACCEPTABLY.><THE ACCOUNT GENERATED ELEVATED CO2 OF >30 MEQ/L ON A PATIENT SAMPLES THAT REPEATED 22 MEQ/L WHEN PROCESSED ON THE ARCHITECT C4000 ANALYZER. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><NO CONSEQUENCES OR IMPACT TO PATIENT; (B)(4) - HIGH TEST RESULTS. AN EVALUATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS. | 0
5088879 | AFTER PERFORMING DECONTAMINATION OF THE ARCHITECT I2000SR, AN EMPLOYEE WAS PERFORMING CALIBRATIONS WHEN SHE NOTICED LIQUID COMING OUT OF DOUBLED WASTE TUBING. THE OPERATOR ADJUSTED THE TUBING AND IMMEDIATELY FELT ILL, WAS DIZZY, VOMITED AND STARTED TO HAVE TROUBLE BREATHING. THE OPERATOR WAS TAKEN TO THE EMERGENCY ROOM DUE TO INHALATION OF THE BLEACH VAPORS FROM THE DECONTAMINATION MIXED WITH THE WORK BUFFER SOLUTION. A SECOND EMPLOYEE FINISHED THE CALIBRATIONS ON THE ARCHITECT I2000SR AND ALSO VOMITED MANY TIMES WITH BLOOD IN IT. THE SECOND OPERATOR WAS ALSO TAKEN TO THE EMERGENCY ROOM. BOTH EMPLOYEES RECEIVED RESPIRATORY THERAPY AND HYDRATION THERAPY WITH INTRAVENOUS FLUIDS. BOTH OPERATORS HAVE RETURNED TO WORK WITH NO ADVERSE EFFECTS FROM RECEIVING THE RESPIRATORY AND HYDRATION THERAPY.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS.><A MDR FOR THE SECOND OPERATOR WAS INITIALLY SUBMITTED UNDER THIS MANUFACTURER REPORT NUMBER 1628664-2015-00238. UPON REVIEW, IT WAS DISCOVERED THAT A SEPARATE MDR REPORT SHOULD HAVE BEEN CREATED FOR THE SECOND OPERATOR EVENT. REFER TO MANUFACTURER REPORT NUMBER 1628664-2015-00286 FOR ALL FURTHER INFORMATION REGARDING THE SECOND OPERATOR. THE CUSTOMER PERFORMED A DECONTAMINATION PROCEDURE USING 2% INSTEAD OF 0.5% SODIUM HYPOCHLORIDE (BLEACH) ON THE ARCHITECT I2000SR. THIS PROCEDURE IS SIMILAR TO THE INTERNAL DECONTAMINATION PROCEDURE WHICH IS AN AS-NEEDED MAINTENANCE TO BE PERFORMED BY FIELD SERVICE USING 0.5% SODIUM HYPOCHLORITE. AFTER THE PROCEDURE WAS COMPLETED, THE CUSTOMER OBSERVED LIQUID WASTE COMING OUT OF THE INSTRUMENT AND NOTICED THE OUTLET WASTE TUBING HAD BECOME DISCONNECTED FROM THE EXTERNAL WASTE PUMP. DURING THE DECONTAMINATION PROCEDURE, THE OUTLET WASTE TUBING HAD BEEN REMOVED BY THE CUSTOMER FROM A NON-ABBOTT WASTE CONTAINER IN ORDER TO EMPTY THE CONTAINER. WHEN THE TUBING WAS PLACED BACK INTO THE CONTAINER, THE CUSTOMER DID NOT NOTICE THE END OF THE TUBING CONNECTED TO THE EXTERNAL WASTE PUMP HAD BECOME BENT CUTTING OFF FLOW. WHILE RUNNING THE INSTRUMENT, THE PRESSURE FROM THE FLOW OF LIQUID WASTE FROM THE EXTERNAL WASTE PUMP CAUSED THE TUBING TO BECOME DISCONNECTED, AND LIQUID WASTE LEAKED ONTO THE FLOOR. THE CUSTOMER ADJUSTED THE TUBING AND RECONNECTED TO THE WASTE PUMP. THE LABORATORY SPACE FOR THE ARCHITECT I2000SR WAS DESCRIBED AS BEING 4 METERS X 4 METERS AND NOT WELL VENTILATED. THE ARCHITECT SYSTEM OPERATIONS MANUAL ADDRESSES LIQUID WASTE REQUIREMENTS AND THE HANDLING/DISPOSING OF WASTE, INCLUDING THE USE OF THE EXTERNAL WASTE PUMP WITH THE ARCHITECT I2000SR WHEN A FLOOR DRAIN IS UNAVAILABLE AND A SINK DRAIN IS USED; CHEMICAL HAZARD PRECAUTIONS WHEN USING SODIUM HYPOCHLORITE (BLEACH) AND OTHER DISINFECTANTS, INCLUDING THE USE OF PROPER FACILITY/EQUIPMENT/OTHER PROTECTIVE MEASURES AND THE USE OF PERSONAL PROTECTIVE EQUIPMENT (PPE) WHEN PERFORMING DECONTAMINATION PROCEDURES; INSTRUCTIONS ON PERFORMING DAILY MAINTENANCE; INSTRUCTIONS WHEN OPERATING THE EXTERNAL WASTE PUMP; AND THE PRECAUTION THAT MAINTENANCE NOT SPECIFIC IN DOCUMENTATION PROVIDED BY ABBOTT LABORATORIES SHOULD NOT BE ATTEMPTED ALONG WITH INSTRUCTIONS TO THE OPERATOR ON CONTACTING AREA CUSTOMER SUPPORT. A REVIEW OF THE ARCHITECT PRODUCT MONITORING FOUND NO SIMILAR ISSUES AS DESCRIBED IN THIS COMPLAINT. BASED ON THE RESULTS OF THIS EVALUATION, THERE IS NO INDICATION OF A MALFUNCTION OF THE EXTERNAL WASTE PUMP OR THE ARCHITECT I2000SR, AND NO DEFICIENCY OF THE PART OR THE ARCHITECT I2000SR WAS FOUND. | 0
5243888 | A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE ACCOUNT. THE FSE DETERMINED THE WZ MANIFOLD (PART NUMBER 7-96176-05) TO BE THE LIKELY CAUSE DUE TO EXCESSIVE BROWN BUILD UP FOUND ON THE NOZZLES OF WZ MANIFOLD 1. THE MANIFOLD AND THE MANIFOLD VALVES WERE REPLACED TO RESOLVE THE ISSUE. A REVIEW OF THE ARCHITECT I2000SR (SN (B)(4)) SERVICE HISTORY WAS PERFORMED. WEEKLY SITE VISITS WERE CONDUCTED THROUGH DECEMBER 12, 2015 TO INSPECT BOTH WZ MANIFOLDS, AND NO ISSUES WERE OBSERVED. NO CONTRIBUTING FACTOR ON OR AROUND THE DATE OF THE EVENT WAS IDENTIFIED, AND NO SUBSEQUENT OCCURRENCES OF DISCREPANT RESULTS WERE FOUND AFTER THE REPLACEMENT OF WZ MANIFOLD 1 AND THE 4 WZ MANIFOLD VALVES BY THE FSE. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE OR NON-STATISTICAL TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEMS OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A MALFUNCTION / DEFICIENCY.><THE CUSTOMER STATED THAT A FALSE POSITIVE ARCHITECT B-HCG RESULT OF 132 IU/L WAS GENERATED. THE SAMPLE WAS REPEATED AND A RESULT OF <1.2 IU/L WAS GENERATED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
5098415 | (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER STATED THAT THE ARCHITECT ANALYZER GENERATED CONTROLS OUT OF RANGE HIGH (OORH) FOR ALL ASSAYS. THROUGH TROUBLESHOOTING THE CUSTOMER REPLACED THE WASH BUFFER AND ALL CONTROLS GENERATED WITHIN ACCEPTABLE RANGE. THE LAB REPEATED PATIENT SAMPLES RUN PRIOR TO THE QC OOR WITH SUSPECT WASH BUFFER AND FOUND 2 SAMPLES REQUIRED CORRECTED REPORTS. ONE FALSELY ELEVATED PHENOBARBITAL SAMPLE - RESULTS PT#2 = 43.1 REPEAT = 8.0 AND ONE HBSAG RESULT WHICH WAS INITIALLY NONREACTIVE BUT UPON REPEAT IS REACTIVE AND WAS SENT FOR CONFIRMATION. THERE WAS NO REPORTED IMPACT TO PATIENT MANAGEMENT. THERE WAS NO ADDITIONAL PATIENT INFORMATION PROVIDED.><EVALUATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. THE CUSTOMER REPORTED CONTROLS OUT OF RANGE FOR ALL ASSAYS BEING TESTED AND THEY ALSO OBSERVED (B)(6) PATIENT RESULTS FOR THE FOLLOWING ASSAYS; PHENOBARBITAL AND (B)(6). THE CUSTOMER PREPARED AND REPLACED THE CONCENTRATED WASH BUFFER WHICH RETURNED THE CONTROL RESULTS TO NORMAL RANGE. THE CUSTOMER RETESTED PATIENT SAMPLES RAN PRIOR TO THE CONTROLS OUT OF RANGE AND FOUND THAT 2 SAMPLES REQUIRED CORRECTED REPORTS (PHENOBARBITAL AND (B)(6)). THERE WERE NO RETURNS AVAILABLE FOR THIS INVESTIGATION. A REVIEW OF COMPLAINTS DID NOT IDENTIFY ELEVATED COMPLAINT ACTIVITY FOR THE ISSUE. THERE IS NOT ENOUGH INFORMATION TO REASONABLY SUGGEST A MALFUNCTION. PER THE COMPLAINT TEXT THERE MAY HAVE BEEN SOME ISSUE WITH THE PREPARATION OR HANDLING OF THE ARCHITECT CONCENTRATED WASH BUFFER. A REVIEW OF LABELING WAS ALSO PERFORMED AND FOUND ADEQUATE INFORMATION ON DETECTING AND RESOLVING THE CUSTOMER'S ISSUE. BASED ON ALL AVAILABLE INFORMATION AND THIS COMPLAINT INVESTIGATION, THE ASSAY PERFORMED AS INTENDED AND NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
4935377 | (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER OBSERVED A FALSELY DECREASED CREATININE RESULT WHILE USING THE ARCHITECT C16000 ANALYZER. THE CUSTOMER PROVIDED THE FOLLOWING DATA (CUSTOMER PROVIDED RANGE 0.57 - 1.11) MG/DL: (B)(6) 2015: INITIAL 1.63, RETEST 5.38 THE RESULT OF 1.63 WAS QUESTIONED BY THE PHYSICIAN. THIS PATIENT HAD PREVIOUSLY GENERATED A RESULT OF 5.19 ON (B)(6) 2015. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT DATA, TROUBLESHOOTING BY ABBOTT FIELD SERVICE, A REVIEW OF THE SYSTEM SERVICE HISTORY, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. RETURN MATERIAL WAS NOT AVAILABLE. AN ABBOTT FIELD SERVICE ENGINEER (FSE) COMPLETED TROUBLESHOOTING OF THE INSTRUMENT. THIS INVOLVED CLEANING OF A PART BY THE FSR THROUGH STANDARD TROUBLESHOOTING PROCEDURES AND DAILY MAINTENANCE, WHICH RESOLVED THE ISSUE. REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT IDENTIFY ANY OTHER ISSUES THAT MAY HAVE CONTRIBUTED TO THE CURRENT EVENT. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND FOR THE PART WHICH WAS CLEANED ON THE ARCHITECT C16000 ANALYZER. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY OF THE ARCHITECT C16000 ANALYZER, LIST NUMBER 03L77, WAS IDENTIFIED. | 0
5074967 | THE CUSTOMER OBSERVED A FALSELY ELEVATED CA19-9 RESULTS ON THE ARCHITECT I2000 ANALYZER. THE FOLLOWING DATA WAS PROVIDED (UNIT OF MEASURE U/ML): INITIAL 89, REPEAT 19. THE CUSTOMER HAS PREVIOUS RESULTS AROUND 15-20. AN ADDITIONAL RESULT WAS PROVIDED AT A LATER DATE FOR (B)(4) INITIAL 60, REPEATS 84, 12, 11.3. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND INSTRUMENT SERVICE TESTING. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. A SINGLE DEFINITIVE CAUSE OF THE ISSUE WAS NOT IDENTIFIED. FIELD SERVICE REPLACED MULTIPLE PARTS THROUGH NORMAL TROUBLESHOOTING ACTIVITIES: PROBE LIST NUMBER 08C94-42, PROBE, WZ LIST NUMBER 08C94-35, AND TUBING/SENSOR, TEMPERATURE, WZ LIST NUMBER 08C94-87. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION, THE ANALYZER PERFORMED AS INTENDED AND NO PRODUCT DEFICIENCY WAS IDENTIFIED.><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). AN EVALUATION IS IN PROCESS. | 0
4558147 | INITIAL REPORTER ADDRESS: (B)(6). DEVICE CODE: (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. AN EVALUATION IS IN-PROCESS.><THE CUSTOMER GENERATED ELEVATED MAGNESIUM RESULTS WHILE USING THE ARCHITECT C16000 ANALYZER. THE CUSTOMER PROVIDED THE FOLLOWING RESULTS: INITIAL RESULT: 3.5 MMOL/L, RETEST MAGNESIUM RESULT: 1.0 MMOL/L THE RESULTS WERE NOT REPORTED FROM THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT DATA, TROUBLESHOOTING BY ABBOTT FIELD SERVICE, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. ABBOTT FIELD SERVICE (FS) INVESTIGATED THE ISSUE ON SITE, INDICATED THAT INTERMITTENT BUBBLING WAS OCCURRING AT THE CUVETTE WASH NOZZLES AND REPLACED TUBING, THIS RESOLVED THE ISSUE. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY OF THE ARCHITECT C16000 ANALYZER, LIST NUMBER 03L77, WAS IDENTIFIED. | 0
4863179 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT DATA, TROUBLESHOOTING BY ABBOTT FIELD SERVICE, A REVIEW OF THE SYSTEM SERVICE HISTORY, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. RETURN MATERIAL WAS NOT AVAILABLE. AN ABBOTT FIELD SERVICE ENGINEER (FSE) COMPLETED TROUBLESHOOTING OF THE INSTRUMENT. THIS INVOLVED A PART REPLACEMENT PER NORMAL TROUBLESHOOTING PROCEDURES, WHICH RESOLVED THE ISSUE. REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT IDENTIFY ANY OTHER ISSUES THAT MAY HAVE CONTRIBUTED TO THE CURRENT EVENT. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY OF THE ARCHITECT C16000 ANALYZER, LIST NUMBER 03L77, WAS IDENTIFIED.><(B)(6). (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER OBSERVED A FALSELY ELEVATED POTASSIUM RESULT WHILE USING THE ARCHITECT C16000 PROCESS MODULE. THE RETEST RESULT ALIGNED WITH THE PATIENT CLINICAL PICTURE. THE CUSTOMER PROVIDED THE FOLLOWING (MMOL/L): POTASSIUM: INITIAL 6.9, RETEST: 3.5 MMOL/L. NO IMPACT TO PATIENT MANAGEMENT WAS INDICATED. | 0
5330646 | THE CUSTOMER STATED THAT THE ARCHITECT ANALYZER GENERATED SAMPLE PRESENTATION AND SAMPLE QUEUE ERRORS ON A PATIENT SAMPLE ((B)(4)) ON THE AMS (OMNILAB MIDDLEWARE) BUT NOT ON THE ARCHITECT ANALYZER. THE ARCHITECT TOOK SAMPLE FROM ONE TUBE WHILE A DIFFERENT TUBE WAS EXPECTED TO BE SAMPLED AND THE RESULTS DID NOT GO TO EXCEPTIONS. THE RESULTS GENERATED WERE FOR FOLATE - INITIAL 6.9 / REPEAT 3.3. THERE WAS NO REPORTED IMPACT TO PATIENT MANAGEMENT. THERE WAS NO ADDITIONAL PATIENT INFORMATION PROVIDED.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><AN EVALUATION WAS PERFORMED BY REVIEWING THE COMPLAINT TEXT, OTHER CUSTOMER COMPLAINTS FOR SIMILAR ISSUES, REVIEW OF LABELING, AND REVIEW OF HISTORICAL DATA. ERROR CODE 8276 SAMPLE PRESENTATION ERROR WAS REPORTED TO HAVE OCCURRED ON ARCHITECT SYSTEM (B)(4). REVIEW OF LOGS BY OMNILAB (MANUFACTURER OF THE AMS MIDDLEWARE) NOTED THAT ONE OF THE ARCHITECT INSTRUMENTS SENT ERROR 8276 FOR SAMPLE (B)(4), AND AMS RULES PROPERLY MANAGED THE ERROR. SAMPLE PRESENTATION (SPE) AND SAMPLE QUEUE (SQE) ERROR WAS GENERATED ON (B)(4) (FOLATE TEST) ON THE APS I2000SR INTERFACE MODULE A WEEK LATER. THE ERROR WAS SEEN IN THE AMS MIDDLEWARE ONLY, NO ERROR WAS FOUND IN THE MESSAGE HISTORY OF ARCHITECT (B)(4) OR IN THE APS HISTORY. WHEN A SAMPLE MOVES INTO POSITION FOR TESTING, THE SAMPLE CARRIER ID IS READ, TESTS ARE ASPIRATED, AND THE CARRIER ID IS READ AGAIN BEFORE THE SAMPLE IS RELEASED. IF THE INITIAL AND FINAL READS DO NOT MATCH, THE FIRST SAMPLE GENERATES THE PRESENTATION ERROR AND THE SECOND SAMPLE GENERATES THE QUEUE ERROR. THESE ERRORS ARE GENERATED BY THE APS SYSTEM. SAMPLE PRESENTATION ERRORS DUE TO ERROR CODE 8275 / 8276 ARE GENERATED BY THE ARCHITECT. THESE ERRORS OCCUR AS A RESULT OF ISSUES WITH THE COMMUNICATION BETWEEN THE APS AND ARCHITECT SYSTEMS. FIELD SERVICE REPLACED A MOTOR, PART NUMBER 8-203835-01, WHICH INTERMITTENTLY FORCED PUCKS PAST THE SAMPLE GATE TO ADDRESS THE ISSUE. NO ISSUES WITH THE APS COULD BE IDENTIFIED. THE RESULT LOG FOR ARCHITECT I2000SR SN (B)(4) CONFIRMED THAT (B)(4) WAS RUN ON (B)(6) 2015 AT 13:57 AND 18:10, NO ERROR MESSAGES WERE ASSOCIATED WITH THE SAMPLE AND NO ERRORS WERE FOUND IN THE HISTORY LOG. A REVIEW OF TICKET HISTORY FOR THE ANALYZER DID NOT IDENTIFY ANY ISSUES THAT MAY HAVE CONTRIBUTED TO THE CURRENT COMPLAINT. THERE WAS NO TREND IDENTIFIED FOR MOTOR, PART NUMBER 8-203835-01, OR FOR THE APS SYSTEM RELATED TO THIS ISSUE. A REVIEW OF THE ARCHITECT AND APS OPERATIONS MANUALS DESCRIBES PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR SPE AND SQE ERRORS. A REVIEW OF COMPLAINTS FOR THE OCCURRENCE OF SPE/SQE ERRORS OR ERROR CODE 8275/8276, FOUND THE ERRORS WERE ADDRESSED BY REPLACING OR ADJUSTING VARIOUS PARTS AND RERUNNING SAMPLES AS NECESSARY. BASED ON THE AVAILABLE INFORMATION, A DEFICIENCY WAS NOT IDENTIFIED. ERROR MESSAGES WERE PROPERLY GENERATED TO ALERT THE OPERATOR. | 0
4665666 | (B)(4). PRODUCT EVALUATION IS IN PROCESS AND THE RESULTS WILL BE SUBMITTED IN A FOLLOW-UP REPORT.><PRODUCT EVALUATION WAS PERFORMED IN ORDER TO INVESTIGATE THIS ISSUE. THE ISSUE WAS DETERMINED TO BE DUE TO THE INSTRUMENT PROBE. NO ADVERSE TREND OF THE INSTRUMENT PROBE WAS FOUND AND NO TREND ASSOCIATED WITH THE ARCHITECT I1000SR ERRATIC RESULT RATE WAS IDENTIFIED. A REVIEW OF THE CUSTOMER'S INSTRUMENT (I1SR55087) SERVICE HISTORY WAS PERFORMED, AND THE REVIEW OF TICKETS IDENTIFIED NO ADDITIONAL OCCURRENCES OF DISCREPANT RESULTS AND IDENTIFIED NO CONTRIBUTING FACTOR ON OR AROUND THE DATE OF THE EVENT. BASED ON THE RESULTS OF THE INVESTIGATION, THERE IS NO INDICATION OF A DEFICIENCY OF THE INSTRUMENT PROBE, THERE IS INSUFFICIENT INFORMATION TO REASONABLY SUGGEST A MALFUNCTION OF THE PROBE. THE CUSTOMER WAS REFERRED TO THE ARCHITECT SYSTEMS OPERATIONS MANUAL, WHICH ADDRESSES OPERATIONAL PRECAUTIONS AND LIMITATIONS AND ADDRESSES TROUBLESHOOTING OF ELEVATED AND ERRATIC RESULTS, INCLUDING THE INADEQUATE DISPENSE OF WASH BUFFER BY THE PIPETTOR AND A DAMAGED OR INCORRECTLY POSITIONED PROBE.><THE CUSTOMER OBSERVED FALSE POSITIVE PATIENT RESULTS FOR THE ARCHITECT STAT TROPONIN-I ASSAY. THE RESULTS WERE REPORTED OUT AND QUESTIONED BY PHYSICIANS. THE PATIENTS THEN WERE ADMITTED AND TRANSFERRED TO ANOTHER FACILITY WHERE THE PATIENTS WERE RETESTED AND NEGATIVE RESULTS WERE GENERATED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. THE CUSTOMER INDICATED THAT LOW CONTROLS WERE TRENDING HIGH (STILL WITHIN RANGE), WHEN THE FALSE POSITIVE PATIENT RESULTS WERE GENERATED. THE CUSTOMER'S REFERENCE RANGE IS 0.00-0.02 NG/ML. THE CUSTOMER PROVIDED THE FOLLOWING PATIENT DATA: SID (B)(6) GENERATED A POSITIVE RESULT OF 0.021 NG/ML, WHICH WAS REPEATED NEGATIVE AT 0.00 SID (B)(6) GENERATED A POSITIVE RESULT OF 0.022 NG/ML, WHICH WAS REPEATED NEGATIVE AT 0.013 NG/ML THE CUSTOMER PROVIDED THE FOLLOWING DATA FROM ONE PATIENT SAMPLE: SID (B)(6) GENERATED A POSITIVE RESULT OF 0.032 NG/ML, WHICH WAS REPEATED POSITIVE AT 0.023 NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
4925064 | (B)(4). THE EVALUATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS.><THE ACCOUNT USES A NORMAL MAGNESIUM RANGES OF 0.70 TO 0.86 MMOL/L ((B)(6) YEARS OLD) AND 0.70 TO 0.91 MMOL/L ((B)(6) YEARS OLD).><THE ACCOUNT GENERATED A FALSELY ELEVATED MAGNESIUM OF 3.5057MMOL/L ON ID 728603 USING THE ARCHITECT C4000. THE SAMPLE REPEATED IN NORMAL RANGE OF 0.967, 0.9599, 0.9481, 0.9533 MMOL/L USING THE SAME ANALYZER AND A ANOTHER ANALYZER. THE ACCOUNT USES A NORMAL MAGNESIUM RANGE OF 0.60 TO 1.1 MMOL/L. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED FOR ID (B)(6). NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><(B)(4). A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS.><A REVIEW OF THE ARCHITECT SERIAL (B)(4) HISTORY LOG IN ABBOTTLINK DID NOT IDENTIFY MESSAGES ASSOCIATED WITH THE INCONSISTENT RESULTS. THE ARCHITECT SERIAL (B)(4) SERVICE HISTORY SINCE NOVEMBER 2014 INCLUDES SEVERAL INSTANCES OF INCONSISTENT HIGH MAGNESIUM RESULTS WHICH WERE ADDRESSED BY ABBOTT FIELD SERVICE ON MULTIPLE DATES BY PERFORMING MAINTENANCE AND TROUBLESHOOTING WHICH INCLUDED REPLACING AND/OR CLEANING AND/OR ADJUSTING VARIOUS PARTS. NO ADVERSE TRENDS WERE IDENTIFIED IN A REVIEW OF TICKET TRENDING DATA FOR MAGNESIUM ASSAY. A REVIEW OF ARCHITECT C4000 TRACKING AND TRENDING DATA REVEALED NO SYSTEMIC ISSUES OR ADVERSE TRENDS ASSOCIATED WITH MAGNESIUM RESULTS. THE ARCHITECT C4000 ERRATIC RESULT RATE IS WITHIN ACCEPTABLE LIMITS WITH NO TRENDS IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES INFORMATION WITH REGARD TO SPECIMEN HANDLING, MAINTENANCE, AND TROUBLESHOOTING THE REPORTED ISSUE. PROBABLE CAUSES INCLUDE, BUT ARE NOT LIMITED TO SCHEDULED MAINTENANCE IS DUE, CUVETTE WASHER IS MALFUNCTIONING, REAGENT OR SAMPLE CARRYOVER, SAMPLE INTEGRITY, AND HARDWARE FAILURE FOR WHICH THE CUSTOMER IS ADVISED TO CONTACT AREA CUSTOMER SUPPORT. A SPECIFIC CAUSE FOR THE INTERMITTENT MAGNESIUM FLIERS HAS NOT BEEN IDENTIFIED AT THIS TIME. THERE IS INSUFFICIENT INFORMATION TO REASONABLY SUGGEST THAT A DEVICE MALFUNCTION HAS CAUSED THE ISSUE AND NO PRODUCT DEFICIENCY IS IDENTIFIED. | 0
4884943 | THE ARCHITECT C8000 ANALYZER, SERIAL NUMBER (B)(4), AT THE CUSTOMER SITE WAS REPLACED DUE TO TUBING POTENTIALLY HAVING AN INCORRECT DIAMETER. NO INJURY OR IMPACT TO PATIENT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. A PRODUCT DEFICIENCY WAS IDENTIFIED AND FIELD ACTION (B)(6)2015 WAS TAKEN TO ADDRESS THE ISSUE.><(B)(4). THE FOLLOWING CORRECTIVE AND PREVENTIVE ACTION HAS BEEN TAKEN: ABBOTT DIAGNOSTICS DIVISION CONTACTED THE FOUR IMPACTED CUSTOMERS TO EXPLAIN THE SITUATION AND INSTRUCTED THEM TO IMMEDIATELY DISCONTINUE USE OF THESE ARCHITECT C8000 INSTRUMENTS. ALL IMPACTED CUSTOMERS ACKNOWLEDGED RECEIPT AND UNDERSTANDING OF THIS ISSUE ON (B)(6) 2015. ALL FOUR IMPACTED CUSTOMERS RECEIVED REPLACEMENT ARCHITECT C8000 ANALYZERS. | 0
4600829 | PRODUCT EVALUATION WAS PERFORMED IN ORDER TO INVESTIGATE THIS ISSUE. FIELD SERVICE REVIEWED THE INSTRUMENT LOGS AND FOUND THE TEMPERATURE WAS IRREGULAR FOR THE WASH ZONE 1 AND 2. AS A RESULT, THE FOLLOWING PARTS WERE REPLACED AND CODED AS BEING WORN OUT FROM NORMAL USE: THE WASH ZONE 1 AND 2 VALVES FOR A TOTAL OF 8 VALVES. TWO VALVE, VACUUM TWO VALVE, DRAIN, VACUUM VESSEL ONE VACUUM PUMP REPAIR KIT THE PARTS REFERENCED ABOVE WERE CONSIDERED TO HAVE RESOLVED THE ISSUE. SERVICE HISTORY FOR THE ARCHITECT (B)(4) REVEALED NO SUBSEQUENT COMPLAINTS OF DISCREPANT RESULTS. 12 MONTH SEARCH FOR SIMILAR COMPLAINTS DID NOT IDENTIFY AN ADVERSE TREND FOR THE FOLLOWING PARTS REPLACED. NO TRENDS WERE IDENTIFIED FOR THE REFERENCED PARTS. THE ARCHITECT SYSTEM OPERATIONS MANUAL ADDRESSES TROUBLESHOOTING FOR THE DESCRIBED ISSUE. THE B-HCG PACKAGE INSERT ADDRESSES SAMPLE HANDLING AND PERFORMANCE CHARACTERISTICS. A SINGLE DEFINITIVE CAUSE OF THE ISSUE WAS NOT IDENTIFIED. FIELD SERVICE REPLACED MULTIPLE PARTS THROUGH NORMAL TROUBLESHOOTING TO RESOLVE THE ISSUE. ALL PARTS WERE DETERMINED TO BE WORN OUT FROM NORMAL USE. THERE IS INSUFFICIENT INFORMATION TO REASONABLY SUGGEST A MALFUNCTION OF THE REPLACED PARTS. NO PRODUCT DEFICIENCY WAS IDENTIFIED.><THE CUSTOMER STATED THAT ONE PATIENT SAMPLE GENERATED A (B)(6), WHICH WAS REPEATED (B)(6) THE NEXT DAY. A NEW SAMPLE FROM THE SAME PATIENT WAS ALSO TESTED AND (B)(6) RESULTS WERE GENERATED. BOTH SAMPLES WERE THEN RETESTED AND (B)(6) RESULTS WERE GENERATED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><(B)(4). PRODUCT EVALUATION IS IN PROCESS AND THE RESULTS WILL BE SUBMITTED IN A FOLLOW-UP REPORT. | 0
5166710 | FIELD SERVICE REPLACED THE INCORRECT TUBING ON THE ARCHITECT I2000SR ANALYZER PER THE MANDATORY TECHNICAL SERVICE BULLETIN ((B)(4)). THERE WAS NO IMPACT TO PATIENT RESULTS.><(B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED ON ALL IMPACTED I2000 SR INSTRUMENTS. THE MANDATORY TSB INSTRUCTS FIELD SERVICE TO REPLACE THE INCORRECT WASTE SYSTEM, VACUUM SYSTEM, AND SOME BUFFER SYSTEM TUBINGS. | 0
5233612 | THE CUSTOMER NOTICED AN ELECTRICAL ODOR WHILE USING THE ACCELERATOR APS CENTRIFUGE MODULE. ABBOTT FIELD SERVICE INSPECTED THE MODULE AND NOTED A MELTED COMPONENT (R100 RESISTOR) ON THE HETTICH MAIN BOARD. FIELD SERVICE REPLACED THE BOARD TO RESOLVE THE ISSUE. NO SMOKE OR FIRE WAS REPORTED. NO INJURY OCCURRED.><FIELD SERVICE FOUND A BURNED R100 COMPONENT ON THE MAIN ELECTRONICS BOARD, PART NUMBER (B)(4). THE BOARD WAS REPLACED TO RESOLVE THE ISSUE. THE ACCELERATOR APS OPERATIONS MANUAL AND THE CENTRIFUGE OPERATIONS MANUAL WERE REVIEWED. TROUBLESHOOTING AND SAFETY INFORMATION IS PROVIDED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND AN INCREASE IN FAILURES FOR THE MAIN ELECTRONICS BOARD, PART (B)(4) (WHICH REPLACED PART (B)(4)) TRIGGERED AN ALERT WHICH SUGGESTS AN ADVERSE TREND. BASED ON THIS INFORMATION, A PRODUCT DEFICIENCY WAS IDENTIFIED AND FURTHER INVESTIGATION IS BEING PERFORMED. A FOLLOW UP REPORT WILL BE SUBMITTED AT THE COMPLETION OF THE INVESTIGATION.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><A PRODUCT DEFICIENCY WAS IDENTIFIED DUE TO AN INCREASE IN COMPLAINTS FOR FAILURES OF THE R100 COMPONENT OF THE MAIN ELECTRONICS BOARD (PARTS 8-35003665-01 AND 8-206623-01). THE SUPPLIER, (B)(4), DETERMINED THAT THE ISSUE IS FROM THE R100 COMPONENT OF THE HETTICH CENTRIFUGE BOARD WHICH WAS TOO SMALL TO WITHSTAND A CURRENT SPIKE. HETTICH RELEASED AN ENHANCED MAIN BOARD REVISION 06 WITH A DIFFERENT R100 RESISTOR. IT IS CONSIDERED ACCEPTABLE TO CONTINUE USING THE AUTOMATION SYSTEM BECAUSE THE ISSUE DOES NOT IMPACT OPERATOR NOR PATIENT SAFETY. THE ACCELERATOR APS OPERATIONS MANUAL AND THE CENTRIFUGE OPERATIONS MANUAL WERE REVIEWED AND SHOW THAT TROUBLESHOOTING AND SAFETY INFORMATION IS PROVIDED. THERE HAVE BEEN NO REPORTED INJURIES FOR THIS ISSUE. IF THE ELECTRICAL COMPONENTS OF THE MAIN ELECTRONICS BOARD FAIL, IT WILL SHUT DOWN THE CENTRIFUGE MODULE. IF THE CENTRIFUGE SHUTS DOWN, THE OPERATOR WILL RECEIVE AN ERROR MESSAGE. THE FREQUENCY OF THIS HAZARD WAS ASSESSED AND IT WAS DETERMINED THAT THOUGH THERE WAS A TREND IN COMPLAINTS, THE OVERALL FREQUENCY OF LOW HAD NOT CHANGED FROM THE PREDICTED FREQUENCY AS OUTLINED IN THE RISK MANAGEMENT FILE. | 1
5267358 | FIELD SERVICE REPLACED THE INCORRECT TUBING ON THE ARCHITECT I2000SR ANALYZER PER THE MANDATORY TECHNICAL SERVICE BULLETIN (TSB 116-132A). THERE WAS NO IMPACT TO PATIENT RESULTS.><(B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED ON ALL IMPACTED I2000 SR INSTRUMENTS. THE MANDATORY TSB INSTRUCTS FIELD SERVICE TO REPLACE THE INCORRECT WASTE SYSTEM, VACUUM SYSTEM, AND SOME BUFFER SYSTEM TUBINGS. | 0
4466670 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND INSTRUMENT SERVICE. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. FIELD SERVICE REPLACED THE NOZZLE NUMBER 4 ((B)(4)) OF THE CUVETTE WASHER WHICH ULTIMATELY RESOLVED THE ISSUE. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION, THE INSTRUMENT PERFORMED AS INTENDED AND NO PRODUCT DEFICIENCY WAS IDENTIFIED.><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><THE LIST NUMBER OF THE CONCOMITANT MEDICAL PRODUCT (CLINICAL CHEMISTRY MAGNESIUM ASSAY LOT 67612UN14) CHANGED FROM LN 07D70-31 TO LN 03P68-21. AN EVALUATION IS STILL IN PROCESS.><THE CUSTOMER OBSERVED FALSELY ELEVATED MAGNESIUM RESULTS ON THE ARCHITECT C4000 ANALYZER. THE FOLLOWING DATA WAS PROVIDED: INITIAL 2.8181, REPEAT 0.9262 MMOL/L. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
5335987 | THIS DEVICE WAS NOT RETURNED FOR EVALUATION. THIS ISSUE OF ANTICOAGULANT (AC) DEPLETION HAS BEEN INVESTIGATED UNDER A CORRECTIVE ACTION. THE RESULTS OF THAT INVESTIGATION DETERMINED THE LIKELY CAUSE WAS THE HARSH CLEANING SOLUTION USED TO CLEAN THE PUMP ROLLERS AT THE CUSTOMER SITE CAN CAUSE DAMAGE TO THE PUMP ROLLERS WHICH MAY LEAD TO A DEVICE MALFUNCTION. NEW ROTORS WERE SENT TO THE CUSTOMER FOR THE ON-SITE TECHNICIAN TO REPLACE. HAEMONETICS FILED A MEDICAL DEVICE SAFETY ALERT AND IS AWAITING THE FDA CORRECTION/REMOVAL REPORTING NUMBER. NOT RETURNED.><HAEMONETICS RECEIVED A COMPLAINT ON (B)(6) 2015 FOR A REPORT OF ANTICOAGULANT DEPLETION DURING A PLASMAPHERESIS. THERE WAS NO REPORT OF ANY DONOR INJURY OR REACTION. | 0
5335158 | THIS PUMP MOTOR WAS RETURNED FOR EVALUATION. THE PART WAS TESTED AND RAN ON A PCS2 DEVICE WITH NO ISSUES NOTED. THE ISSUE OF ANTICOAGULANT (AC) DEPLETION HAS BEEN INVESTIGATED UNDER A CORRECTIVE ACTION. THE RESULTS OF THAT INVESTIGATION DETERMINED THE LIKELY CAUSE WAS THE HARSH CLEANING SOLUTION USED TO CLEAN THE PUMP ROLLERS AT THE CUSTOMER SITE CAN CAUSE DAMAGE TO THE PUMP ROLLERS WHICH MAY LEAD TO A DEVICE MALFUNCTION. THE MOST LIKELY ROOT CAUSE OF THE REPORTED ISSUE IS DUE TO THIS ISSUE WITH THE PUMP ROLLERS. HAEMONETICS FILED A MEDICAL DEVICE SAFETY ALERT AND IS AWAITING THE FDA CORRECTION/REMOVAL REPORTING NUMBER.><HAEMONETICS RECEIVED A COMPLAINT ON (B)(6) 2015 FOR A REPORT OF ANTICOAGULANT DEPLETION DURING A PLASMAPHERESIS DUE TO THE MOTOR NOT TURNING FREELY. THERE WAS NO REPORT OF ANY DONOR INJURY OR REACTION. | 0
5334340 | THIS DEVICE WAS NOT RETURNED FOR EVALUATION. THIS ISSUE OF ANTICOAGULANT (AC) DEPLETION HAS BEEN INVESTIGATED UNDER A CORRECTIVE ACTION. THE RESULTS OF THAT INVESTIGATION DETERMINED THE LIKELY CAUSE WAS THE HARSH CLEANING SOLUTION USED TO CLEAN THE PUMP ROLLERS AT THE CUSTOMER SITE CAN CAUSE DAMAGE TO THE PUMP ROLLERS WHICH MAY LEAD TO A DEVICE MALFUNCTION. NEW ROTORS WERE SENT TO THE CUSTOMER FOR THE ON-SITE TECHNICIAN TO REPLACE. (B)(4).><HAEMONETICS RECEIVED A COMPLAINT ON (B)(6) 2015 FOR A REPORT OF ANTICOAGULANT DEPLETION DURING A PLASMAPHERESIS. THERE WAS NO REPORT OF ANY DONOR INJURY OR REACTION. | 0
4619889 | ><CRYOCARE MACHINE FOR PROSTATE CRYOABLATION MALFUNCTIONED WHILE IN THE MIDDLE OF THE PROCEDURE. THE MACHINE WAS UNABLE TO BE RESTARTED. THE CASE WAS ABORTED. | 0
4826878 | ><THE TREADMILL WILL RANDOMLY STOP DURING A BRUCE CYCLE. NO PATIENTS HAVE BEEN HARMED, BUT THERE IS THE POTENTIAL FOR PATIENT HARM WHEN THE TREADMILL STOPS SUDDENLY AND WITHOUT PRIOR WARNING WHILE THE PATIENT IS USING IT. THIS FACILITY HAS FOUR OF THESE DEVICES; TWO OF THE DEVICES HAVE HAD THIS PROBLEM AND THE OTHER TWO HAVE THE POTENTIAL TO HAVE THIS PROBLEM. THERE IS NO WARNING THAT SIGNALS THE STAFF OR THE PATIENT THAT THE PROBLEM WILL OCCUR OR IS ABOUT TO OCCUR. WE HAVE BEEN TRYING TO CONTACT THE MANUFACTURER X2 MONTHS REGARDING THIS ISSUE. WHEN CALLING, WE ARE DIRECTED TO AN AUTOMATED MESSAGE ON A SPECIAL PHONE LINE THAT THE MANUFACTURER HAS SET UP FOR THIS SPECIFIC PROBLEM. THERE HAVE BEEN NO MANUFACTURER RETURN CALLS, NO CORRESPONDENCE, AND NO SERVICE TECHNICIANS ON SITE FOR THIS ISSUE. THIS FACILITY BEGAN HAVING PROBLEMS WITH THIS DEVICE PRIOR TO THE MANUFACTURER ISSUING AN ACTION LETTER. ALTHOUGH CALLS AND REQUESTS FOR SERVICE SUPPORT WERE SENT TO THE MANUFACTURER PRIOR TO THE ACTION NOTICE, THIS FACILITY STILL HAS NOT RECEIVED RETURN COMMUNICATION FROM THE MANUFACTURER ON THIS ISSUE.======================MANUFACTURER RESPONSE FOR STRESS MACHINE, T2100 (PER SITE REPORTER).======================GE HEALTHCARE HAS ISSUED A PRODUCT ALERT FOR OUR STRESS MACHINES. WE CANNOT AT THIS MOMENT GET A HOLD OF GE TO COME IN AND RECERTIFY THE TREADMILLS. | 0
5311058 | THE CANNULA DID NOT DEPLOY ON THE OMNI POD. | 0
5157355 | IT WAS REPORTED THAT A PHYSICIAN'S TABLET PROGRAMMER WAS EXHIBITING AN ERROR MESSAGE OF UNABLE TO OPEN PORT" WHEN ATTEMPTING TO INTERROGATE A DEMO DEVICE. THE PROGRAMMER USB SERIAL ADAPTER CABLE WAS REPLACED AND INTERROGATION WAS THEN SUCCESSFUL. IT WAS SUBSEQUENTLY REPORTED THAT THE SAME ERROR MESSAGE OCCURRED WHEN USING THE TABLET PROGRAMMER WITH THE REPLACEMENT USB SERIAL ADAPTER CABLE. THE PHYSICIAN DID NOT HAVE ANOTHER SPARE USB CABLE TO USE SO A NEW TABLET PROGRAMMER WAS PROVIDED TO THE PHYSICIAN. THE ORIGINAL TABLET PROGRAMMER AND BOTH USB SERIAL ADAPTER CABLES HAVE BEEN RETURNED TO THE MANUFACTURER AND ARE CURRENTLY UNDERGOING PRODUCT ANALYSIS. THE 1ST SUSPECTED USB CABLE FAILURE IS REPORTED IN MFR. REPORT# 1644487-2015-06127. THE 2ND SUSPECTED USB SERIAL ADAPTER CABLE FAILURE IS REPORTED IN THIS REPORT.><PRODUCT ANALYSIS FOR THE 2ND RETURNED USB CABLE INDICATES THAT THE SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. PRODUCT ANALYSIS OF THE TABLET PROGRAMMER REVEALED THAT THE TABLET WAS UNABLE TO PERFORM AN INTERROGATION. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A DEFECTIVE MAIN BOARD (USB PORT NOT FUNCTIONING). ONCE THE MAIN BOARD WAS REPLACED WITH A KNOWN GOOD MAIN BOARD | 1
5124559 | .><THE SUSPECT SOFTWARE INSTALL FLASH DRIVE WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS. THERE WERE NO OBSERVED ANOMALIES WITH THE USB DRIVE OR UPGRADE SOFTWARE APPLICATION FILE. THE CAUSE FOR THE INCORRECT OPERATION WAS DUE TO USER-ERROR AS THE BATCH FILE NAME WAS NOT TYPED IN CORRECTLY DURING THE UPGRADE PROCESS. USE OF THE PROPER BATCH FILE NAME RESULTED IN SUCCESSFUL UPGRADE FUNCTIONALITY.><IT WAS REPORTED THAT THE PHYSICIAN'S TABLET WAS BEING UPGRADED BUT COULD NOT PROCEED PAST THE SOFTWARE UPGRADE STEP. THE SCREEN WAS FROZEN" AT THE SOFTWARE DEPLOYMENT PART OF THE PROCEDURE. THE TABLET WAS RESTARTED BUT ONCE THE RESTART COMPLETED | 1
4913648 | ANALYSIS WAS COMPLETED. IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE SCREEN ALIGNMENT UTILITY. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><.><IT WAS REPORTED THAT THE PHYSICIAN¿S PROGRAMMING HANDHELD WAS NOT WORKING. THE MANUFACTURER¿S REPRESENTATIVE ATTEMPTED TO USE THE HANDHELD, BUT IT WOULD NOT RESPOND TO HIS TOUCH OF HIS FINGER OR THE STYLUS. A HARD RESET WAS PERFORMED, BUT THE HANDHELD STILL WOULD NOT RESPOND. THE HANDHELD AND SOFTWARE HAVE BEEN RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE. | 1
5290207 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD SCREEN WAS OPERABLE; HOWEVER, IT WAS TOO DIM TO READ. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON 03/22/2016. NO ANOMALIES ASSOCIATED WITH THE DISPLAY WERE IDENTIFIED DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON 03/22/2016. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 1
4449834 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><THE PROGRAMMING HANDHELD AND RELATED SOFTWARE WERE RECEIVED BY THE MANUFACTURER FOR ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD, AND THE REPORTED ALLEGATION WAS NOT VERIFIED. THE HANDHELD WAS ABLE TO POWER ON AND OFF WITH NO OBSERVED ANOMALIES. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS.><IT WAS REPORTED THAT THE PHYSICIAN¿S PROGRAMMING HANDHELD WOULD NOT TURN ON. IT WAS REPORTED THAT THE HANDHELD WAS PLUGGED IN AND CHARGED, BUT THE HANDHELD WOULD STILL NOT TURN ON. GOOD FAITH ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. A REPLACEMENT PRODUCT WAS PROVIDED AS REPLACEMENT. THE HANDHELD HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE.><. | 1
5022700 | .><A TABLET SERIAL CABLE WAS RETURNED DUE TO DAMAGE CAUSED BY WATER FLOODING. AN ANALYSIS WAS PERFORMED ON THE SERIAL CABLE AND IT WAS DETERMINED THAT THE TABLET WAS UNABLE TO ESTABLISH COMMUNICATION WITH A KNOWN GOOD WAND AND GENERATOR. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A BROKEN WIRE CONNECTION IN THE SERIAL CABLE DB9 HOOD ASSEMBLY. NO FURTHER ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS. | 1
4354004 | IT WAS REPORTED THAT THE PROGRAMMING SYSTEM WAS NOT FUNCTIONING SO THE VNS PATIENT'S DEVICE SETTINGS COULD NOT BE INCREASED. FOLLOW-UP REVEALED THAT THE HANDHELD DEVICE WOULD NOT CHARGE DESPITE BEING CONNECTED TO THE POWER ADAPTOR FOR OVER TWO HOURS. THE HANDHELD WAS RESET BUT WOULD FREEZE AT STARTUP. THE SUSPECT HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE.><. | 1
4900070 | ANALYSIS OF THE RETURNED TABLET WAS COMPLETED ON 07/15/2015 AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE TABLET PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><.><IT WAS REPORTED THAT THE PHYSICIAN'S TABLET HAD A LOOSE USB PORT. THE PHYSICIAN REQUESTED A NEW PROGRAMMING TABLET. THE FAULTY TABLET WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 1
5247155 | IT WAS REPORTED ON (B)(6) 2015 FROM THAT A PHYSICIAN'S TABLET WAS MALFUNCTIONING. IT WAS STATED THAT THE SCREEN APPEARED FUZZY AS IF THE TABLET HAS BEEN DROPPED. THE TABLET WAS WORKING FINE UNTIL IT WAS SENT FOR UPGRADE FOR 11.0 SOFTWARE UPGRADE AND AS SOON AS IT WAS TURNED ON THE SCREEN APPEARED FUZZY. IT WAS ATTEMPTED TO TURN THE DEVICE ON AND OFF 4 TIMES BUT THIS DID NOT HELP TO RESOLVE THE ISSUE. IT WAS ALSO NOTED THAT THE TABLET APPEARS DAMAGED WHERE THE PEN HOLDER IS WHICH IS BELIEVED TO POSSIBLE BE WHY THE SCREEN IS FUZZY. THE FUZZY SCREEN WAS DESCRIBED AS NOT SHOWING ANY WRITING ON IT AS IT WOULD WHEN NORMALLY TURNED ON. IT IS WHITE STATIC/FUZZY. THE TABLET IS EXPECTED TO BE RETURNED FOR ANALYSIS BUT HAS NOT BEEN RECEIVED TO DATE.><.><THE TABLET WAS RECEIVED FOR ANALYSIS ON 11/30/2015. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><PRODUCT ANALYSIS FOR THE TABLET WAS COMPLETED AND APPROVED ON 12/16/2015. AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND NO ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE TABLET PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
4899403 | (B)(4).><THE USB TO DB9 SERIAL ADAPTER CABLE WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS. THE ANALYSIS WAS COMPLETED AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE CIRCUIT BOARD, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE MANUFACTURER'S SALES REPRESENTATIVE WAS RECEIVING AN UNABLE TO OPEN PORT" MESSAGE UPON ATTEMPTING TO INTERROGATE HER DEMO GENERATOR. TROUBLESHOOTING WAS UNABLE TO RESOLVE THE ISSUE. THE SUSPECT DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR ANALYSIS TO DATE.><DEVICE FAILURE OCCURRED | 1
4366900 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2013. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 1
4027546 | THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ELECTRONIC SUBMISSION OF THIS REPORT WAS NOT POSSIBLE DUE TO A PROBLEM WITH THE MFR'S SOFTWARE USED FOR EMDR SUBMISSIONS. METHOD: REVIEW OF PROGRAMMING HISTORY PERFORMED.><IT WAS REPORTED THAT DURING THE PT'S OFFICE VISIT THAT THE PT'S MAGNET ACTIVATIONS WERE NOT DISPLAYING CORRECTLY. IT WAS REPORTED THAT THE PT ONLY HAD ONE MAGNET SWIPE RECENTLY, WITH ONLY A TOTAL OF 8 REGISTERING. HOWEVER, FOR THE ONE ACTIVATION ON (B)(6) 2014, THE TABLET MAGNET HISTORY WAS DISPLAYING 15 TIMES. THE TOTAL SWIPES REGISTERED DISPLAYED AS 86. A COPY OF THE PHYSICIAN'S PROGRAMMING TABLET DATA WAS OBTAINED. MAGNET HISTORY COINCIDES WITH ONLY ONE SWIPE REGISTERING ON (B)(6) 2014 WITH NO INDICATION OF MULTIPLE MAGNET ACTIVATIONS. IT WAS OBSERVED THAT TOTAL OPERATING TIME ROLLOVER OF THE DEVICE OCCURRED BETWEEN THE LAST APPOINTMENT ON (B)(6) 2012 AND (B)(6) 2014. AS A RESULT, THE MAGNET HISTORY ON THE TABLET WAS SHOWING 15 ACTIVATIONS ON (B)(6) 2014, BECAUSE NO OTHER ACTIVATIONS WERE AVAILABLE IN THE HISTORY DUE TO ROLLOVER. THE 15 MOST RECENT MAGNET ACTIVATIONS STORED WITHIN THE GENERATORS MEMORY SPAN THE TIME PERIOD BEFORE/AFTER THE ROLLOVER EVENT ARE DISPLAYED AND SINCE THERE WAS ONLY ONE SWIPE, 15 ACTIVATIONS WERE DISPLAYED ON THE TABLET ON (B)(6) 2014. IT IS KNOWN THAT THE DEVICE REED SWITCH CAN BE ACTIVATED BY SEVERAL CONSUMER DEVICES/OTHER DEVICES THAT HAVE MAGNETIC FIELD (I.E. MULTIPLE ACCIDENTAL MAGNET SWIPES, ETC.). HOWEVER, ATTEMPTS FOR FURTHER INFO FROM THE PHYSICIAN WERE UNSUCCESSFUL.><THIS REPORT IS TO CORRECT THE INITIAL MFR. REPORT THAT WAS INADVERTENTLY SENT WITH A DUPLICATE REPORT #. THE REPORT WAS SUBMITTED VIA EMDR ON 08/21/2014 REGARDING A HANDHELD SCREEN ISSUE AND WAS INADVERTENLY SENT WITH REPORT # 1644487-2014-02046. PLEASE CORRECT THAT INITIAL REPORT TO REFLECT THE NEW NUMBER ATTACHED TO THIS FOLLOW-UP REPORT.><THIS REPORT IS TO CORRECT THE MFR. REPORT # OF THE INITIAL MFR. REPORT. | 1
5193806 | IT WAS REPORTED THAT THEY PHYSICIAN'S TABLET SERIAL CABLE WAS FAULTY AS IT LED TO COMMUNICATION ERRORS. AN UNABLE TO ESTABLISH COMMUNICATION" MESSAGE WAS RECEIVED. UPON REPLACING THE SERIAL CABLE WITH ANOTHER | 1
4920682 | DURING AN INTERNAL REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS IDENTIFIED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2013. THE SETTINGS WERE NOT CORRECTED ON THE SAME DATE. NO PATIENT ADVERSE EVENTS WERE REPORTED.><REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY. | 1
4835226 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN¿S (B)(4) X50 HANDHELD IS NOT TURNING ON EVEN WITH A FULL CHARGE. IT WAS STATED THAT THE LOCK BUTTON IS OFF. THE HANDHELD AND SOFTWARE WAS RECEIVED FOR ANALYSIS ON 05/20/2015. PRODUCT ANALYSIS FOR THE RETURNED HANDHELD WAS COMPLETED AND APPROVED ON 06/09/2015. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD SHOWED NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. PRODUCT ANALYSIS FOR THE RETURNED SOFTWARE WAS COMPLETED AND APPROVED ON 06/09/2015. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
4628497 | IT WAS REPORTED THAT THE PHYSICIAN'S TABLET WAS GIVING A FAILURE TO OPEN PORT" MESSAGE. IT WAS MENTIONED THAT THE USB SERIAL CONNECTOR ON THE PROGRAMMER WAS "REALLY LOOSE" AND THEY FELT THAT WAS LIKELY THE ISSUE. A NEW USB SERIAL CABLE WAS REQUESTED AND SENT BUT THIS DID NOT FIX THE ISSUE | 1
5098522 | IT WAS REPORTED THAT A TABLET PROGRAMMING SYSTEM RECEIVED AN UNABLE TO OPEN PORT" ERROR. IT WAS BELIEVED THAT THE SERIAL CABLE FAILED. THE TABLET SERIAL CABLE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><PRODUCT ANALYSIS WAS COMPLETED FOR THE TABLET SERIAL CABLE. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH TWO DISCONNECTED WIRE CONNECTIONS IN THE RETURNED SERIAL CABLE. ONCE THE WIRES WERE SOLDERED ONTO THE SERIAL CABLE PCB | 1
4671982 | AN ANALYSIS WAS PERFORMED ON THE RETURNED USB TO DB9 CABLE AND A MECHANICAL PROBLEM WITH THE SERIAL CABLE WAS VERIFIED. THE CAUSE FOR THE REPORTED ISSUE IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE PROGRAMMING TABLET WAS REPORTING A FAILURE TO OPEN PORT" ERROR. NO PATIENT WAS AFFECTED OR INVOLVED. TROUBLESHOOTING WAS PERFORMED. TROUBLESHOOTING INCLUDED PERFORMING A HARD RESET AND RE-SEATING CONNECTIONS. IT WAS NOTED THAT THE CONNECTION BETWEEN THE USB CABLE AND TABLET WAS LOOSE. A NEW USB SERIAL ADAPTER CABLE WAS SENT TO SITE. THE TABLET USB SERIAL ADAPTER CABLE WAS RECEIVED BY THE MANUFACTURER. HOWEVER | 1
4641640 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><PRODUCT ANALYSIS WAS COMPLETED FOR THE HANDHELD DEVICE ON 04/20/2015. THE CAUSE FOR THE FAILURE TO POWER ON IS ASSOCIATED WITH A DEFECTIVE MAIN BOARD. ONCE THE BOARD WAS REPLACED, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. PRODUCT ANALYSIS WAS COMPLETED FOR THE SOFTWARE FLASHCARD ON 04/20/2015. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WOULD NOT POWER ON AFTER BEING PLUGGED INTO AN ELECTRICAL OUTLET FOR APPROXIMATELY 10 MINUTES. IT WAS REPORTED THAT THE POWER LIGHT ILLUMINATED ORANGE INDICATING THAT THERE WAS POWER TO THE HANDHELD. A HARD RESET WAS PERFORMED WHICH DID NOT RESOLVE THE ISSUE. THE HANDHELD WAS LEFT PLUGGED TO THE POWER OUTLET FOR A WHILE LONGER AND THEN TROUBLESHOOTING WAS AGAIN ATTEMPTED. THE HANDHELD WOULD STILL NOT POWER ON SO THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING COMPUTER. THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 1
5135916 | .><THE TABLET WAS RECEIVED FOR ANALYSIS ON 10/12/2015. PRODUCT ANALYSIS FOR THE TABLET WAS COMPLETED AND APPROVED ON 11/04/2015: AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE TABLET WAS UNABLE TO RECOGNIZE THE SD CARD. AS A RESULT, THE DATABASE UTILITIES BUTTON WAS DISABLED IN THE VNS SOFTWARE. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A HARDWARE FAILURE ON THE MAIN BOARD. THE ANOMALY WILL PREVENT THE ABILITY TO PERFORM A SOFTWARE UPGRADE. ALTHOUGH THE DATABASE UTILITIES BUTTON WAS DISABLED, THE VNS SOFTWARE STILL HAD LIMITED FUNCTIONALITY AND WAS ABLE TO INTERROGATE, PROGRAM, AND PERFORM DIAGNOSTIC TESTS DURING THE ANALYSIS.><IT WAS REPORTED ON (B)(6) 2015 FROM AN UPGRADE SPECIALIST THAT HE WAS ABOUT TO START THE UPGRADE PROCESS FOR THIS TABLET, BUT THE DATABASE UTILITIES BUTTON WAS GRAYED OUT. THIS ISSUE WAS DISCOVERED ON (B)(6) 2015. HE INDICATED THAT HE TRIED REBOOTING ABOUT 5-6 TIMES AND EVEN USED 3 DIFFERENT SD CARDS BUT THE BUTTON WOULD NOT WORK. THE TABLET THAT WAS HAVING ISSUES IS EXPECTED TO BE RETURNED FOR ANALYSIS BUT HAS NOT BEEN RECEIVED TO DATE.><AN ADDENDUM TO PRODUCT ANALYSIS OF THE TABLET WAS COMPLETED AND APPROVED ON (B)(6) 2016: VISUAL INSPECTION OF THE SD MEMORY CARD SLOT LEADS IDENTIFIED NO ANOMALIES. | 1
4733491 | DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE HANDHELD PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD DOES NOT RESPOND TO CHARGING OR HARD RESETTING. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING COMPUTER. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><AN ANALYSIS ON THE RETURNED HANDHELD CONFIRMED THAT THERE WERE NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WAS RETURNED WITHOUT A BATTERY COVER. AS A RESULT, THE HANDHELD WOULD NOT POWER ON. ONCE A KNOWN BATTERY COVER WAS INSTALLED, THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH THE FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING ANALYSIS OF THE RETURNED FLASHCARD.><.><THE HANDHELD COMPUTER AND THE SOFTWARE FLASHCARD WERE RECEIVED BY THE MANUFACTURER. ANALYSIS IS UNDERWAY BUT IT HAS NOT BEEN COMPLETED TO DATE. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE HANDHELD PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. | 1
4953732 | .><(B)(4). SUSPECT DEVICE UDI, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY DID NOT INCLUDE THE UDI.><PRODUCT ANALYSIS WAS COMPLETED FOR THE TABLET SERIAL CABLE. ANALYSIS REVEALED A DISCONNECTED WIRE CONNECTION IN THE SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE TABLET DEVICE SHOWED AN UNABLE TO OPEN PORT" ERROR MESSAGE. TROUBLESHOOTING WAS PERFORMED BUT THE ISSUES DID NOT RESOLVE UNTIL A REPLACEMENT USB SERIAL CABLE WAS PROVIDED. THE SUSPECT USB SERIAL CABLE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY." | 1
5334861 | PRODUCT ANALYSIS FOR THE TABLET WAS COMPLETED AND APPROVED ON (B)(6) 2016. AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE USB PORT WAS DAMAGED. AS A RESULT, THE TABLET WAS UNABLE TO ESTABLISH COMMUNICATION. NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED ON (B)(6) 2015 THAT THE USB PORT ON THE PHYSICIAN'S TABLET HAS BEEN DAMAGED AND THE PROGRAMMING WAND CAN NO LONGER PLUG INTO IT. NO PATIENT CARE WAS IMPACTED BECAUSE THEY WERE ABLE TO USE A BACKUP SYSTEM. THE TABLET WAS RECEIVED FOR ANALYSIS ON 12/10/2015. PRODUCT ANALYSIS IS UNDERWAY BUT HAS NOT BEEN COMPLETED TO DATE.><. | 1
5239418 | IT WAS REPORTED BY A COMPANY REPRESENTATIVE THAT HIS TABLET USB SERIAL CABLE HAD FAILED. IT WAS REPORTED THAT ERROR MESSAGES WERE RECEIVED WHILE ATTEMPTING TO PROGRAM PATIENT DATA PRE-OPERATIVELY. TROUBLESHOOTING WAS PERFORMED BY TURNING THE TABLET OFF, AND THEN ON AGAIN, REMOVING THE CABLE AND ENSURING THE CABLE WAS TIGHTENED. SWITCHING USB SERIAL CABLES RESOLVED THE ISSUE. THE COMPANY REPRESENTATIVE ACCIDENTALLY DISCARDED THE SUSPECTED SERIAL CABLE, AND IT IS UNABLE TO BE RETURNED FOR ANALYSIS.><. | 1
4895864 | A COMPUTER SOFTWARE ERROR HAS BEEN VERIFIED, ON A PATIENT'S MODEL 102 GENERATOR, FROM A REVIEW ON (B)(6) 2015 OF THE PROGRAMMING HISTORY DATABASE SPREADSHEET AND PERIODIC REVIEW ALGORITHM. ON (B)(6) 2012 THE DEVICE WAS INTERROGATED, PROGRAMMED, A SYSTEM DIAGNOSTIC WAS PERFORMED (RESULTS WERE OK") | 1
4924673 | DURING AN INTERNAL REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS IDENTIFIED THAT AN UNINTENDED CHANGE IN DEVICE SETTINGS OCCURRED DURING AN OFFICE VISIT BETWEEN (B)(6) 2012 AND (B)(6) 2014. THE DEVICE WAS INTERROGATED ON (B)(6) 2014 AND IT WAS FOUND TO BE DISABLED. THE SETTINGS WERE NOT CORRECTED. REVIEW OF THE AVAILABLE DIAGNOSTICS HISTORY FOUND NO INTERRUPTED SYSTEM TESTS THAT COULD LEAD TO THE UNINTENDED SETTINGS. NO PATIENT ADVERSE EVENTS WERE REPORTED.><REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY. | 1
4660263 | THE PHYSICIAN PROVIDED INFORMATION ON (B)(6) 2015 THAT HE DID INTENTIONALLY LEAVE THE GENERATOR SETTINGS PROGRAMMED TO ZERO.><REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><PROGRAMMING AND DIAGNOSTIC HISTORY FOR THE VNS WAS RECEIVED WHICH INDICATED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2014. THE PATIENT¿S DEVICE SETTINGS WERE NOT CORRECTED AT THE OFFICE VISIT. | 1
4786640 | (B)(4). SUSPECT DEVICE UDI, CORRECTED DATA: THE FOLLOW-UP REPORT #1 INADVERTENTLY REPORTED THE INCORRECT UDI. SERIAL #, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY REPORTED THE INCORRECT PRODUCT INFORMATION. DEVICE MANUFACTURE DATE, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY REPORTED THE INCORRECT PRODUCT INFORMATION.><IT WAS REPORTED THAT THE PHYSICIAN¿S PROGRAMMING TABLET WAS INTERMITTENTLY HAVING COMMUNICATION DIFFICULTIES WITH THE FAILURE TO OPEN PORT" ERROR MESSAGE. IT WAS REPORTED THE PHYSICIAN WOULD RETRY INTERROGATION ONCE TO A FEW TIMES | 1
5318465 | .><ANALYSIS OF THE ADAPTER CABLE WAS COMPLETED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED ON 12/01/2015 THAT THE PHYSICIAN WAS RECEIVING A ¿FAILURE TO OPEN PORT¿ ERROR MESSAGE. THIS ISSUE WAS ATTRIBUTED TO FAULTY ADAPTER CABLES AND THUS THEY WERE REPLACED WHICH RESOLVED THE PROBLEM. THE SUSPECT USB TO DB9 SERIAL ADAPTER CABLE HAS BEEN RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE. | 1
4841843 | SUSPECT DEVICE UDI: (B)(4).><THE SUSPECT USB CABLE WAS RETURNED FOR PRODUCT ANALYSIS. THE ANALYSIS PERFORMED ON THE RETURNED USB TO DB9 CABLE REVEALED A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE TABLET DEVICE SHOWED AN ¿UNABLE TO OPEN PORT¿ ERROR MESSAGE WHEN ATTEMPTING TO PERFORM INTERROGATIONS. THE SUSPECT USB SERIAL CABLE HAS NOT BEEN RETURNED TO DATE.><DATE OF EVENT, CORRECTED DATA: THE DATE OF THE EVENT WAS (B)(6) 2015 AND NOT 5/21/2015 AS ORIGINALLY REPORTED. DATE OF THIS REPORT, CORRECTED DATA: THE DATE OF THE INITIAL REPORT WAS (B)(4) 2015 AND NOT 5/21/2015 AS ORIGINALLY REPORTED. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY. | 1
4396183 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><.><IT WAS REPORTED THAT THE PHYSICIAN NOTICED THAT THE TABLET USB ADAPTER CABLE CONNECTION TO ONE OF HIS TABLETS WAS VERY LOOSE. A REPLACEMENT ADAPTER CABLE WAS PROVIDED TO THE PHYSICIAN TO SEE IF THIS WILL RESOLVE THE ISSUE. THE SUSPECT ADAPTER CABLE HAS NOT BEEN RECEIVED TO DATE.><DATE OF EVENT, CORRECTED DATA: ADDITIONAL INFORMATION WAS RECEIVED THAT THE DATE OF EVENT WAS (B)(6) 2014. THIS WAS INADVERTENTLY NOT REPORTED PREVIOUSLY.><THE TABLET SERIAL DATA CABLE WAS RECEIVED BY THE MANUFACTURER. AN ANALYSIS WAS PERFORMED ON THE RETURNED USB TO DB9 CABLE, AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. NO LOOSE FITTING CONNECTIONS WERE IDENTIFIED DURING THE ANALYSIS. | 1
5176641 | .><THE POWER CORD WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS ON (B)(6) 2015. ANALYSIS WAS PERFORMED ON THE RETURNED AC ADAPTER AND NO ANOMALIES ASSOCIATED WITH THE AC ADAPTER PERFORMANCE WERE NOTED DURING TESTING USING A KNOWN GOOD TABLET. THE AC ADAPTER PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED BY THE PHYSICIAN THAT THE POWER CORD FOR THE PHYSICIAN'S TABLET DOES NOT WORK. THE POWER CORD WAS TESTED BY THE COMPANY REPRESENTATIVE WITH A KNOWN WORKING GOOD TABLET AND THE CABLE ONCE AGAIN DID NOT WORK. A KNOWN WORKING POWER CABLE WAS CONNECTED TO THE PHYSICIAN'S TABLET AND THE TABLET WAS ABLE TO CHARGE. THE PHYSICIAN HAS BEEN PROVIDED WITH A NEW POWER CABLE. THE COMPLAINT POWER CORD IS EXPECTED TO BE RETURNED TO THE MANUFACTURER FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE. | 1
5291254 | THE HANDHELD PROGRAMMER, SOFTWARE FLASHCARD, AND PROGRAMMING WAND WERE RECEIVED ON 12/28/2015. ANALYSIS OF THE PROGRAMMING WAND WAS COMPLETED ON 01/07/2016. NO VISUAL OR MECHANICAL ANOMALIES WERE IDENTIFIED WITH THE WAND. CONTINUITY TESTING OF THE SERIAL DATA CABLE AND BATTERY CABLE PASSED. THE WAND PERFORMED ACCORDING TO ALL FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE HANDHELD PROGRAMMER AND SOFTWARE FLASHCARD HAS NOT BEEN COMPLETED TO DATE.><.><IT WAS REPORTED THAT A PHYSICIAN'S HANDHELD PROGRAMMER WAS NO LONGER HOLDING A CHARGE. TROUBLESHOOTING WAS PERFORMED, INCLUDING CONFIRMING THAT IT WAS PLUGGED IN PROPERLY, SOFT AND HARD RESETS, AND RESEATING THE CONNECTIONS, BUT THE HANDHELD WOULD STILL DIE AFTER A FEW MOMENTS OF USE AFTER BEING UNPLUGGED. NO PATIENTS WERE AFFECTED BY THE HANDHELD NOT HOLDING A CHARGE. THE HANDHELD IS EXPECTED FOR RETURN, BUT IT HAS NOT BEEN RECEIVED TO DATE.><ANALYSIS WAS COMPLETED ON THE HANDHELD PROGRAMMER AND THE FLASHCARD ON 01/26/2016. NO ANOMALIES WERE IDENTIFIED WITH THE HANDHELD BATTERY, BUT IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO CHARGE THE MAIN BATTERY. THE CAUSE FOR THE ANOMALY WAS ASSOCIATED WITH BROKEN SOLDER CONNECTIONS ON THE HANDHELD MAIN BOARD. NO OTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
4650841 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><.><IT WAS REPORTED THAT THE USB CONNECTOR OF THE TABLET DEVICE APPEARED TO BE DEFECTIVE. ANOTHER USB CONNECTOR WAS USED WITH THE PROGRAMMING SYSTEM AND THE ISSUES RESOLVED. A REPLACEMENT USB CONNECTOR WAS SENT TO THE HEALTHCARE PROFESSIONAL. THE SUSPECTED USB CABLE WAS RETURNED TO THE MANUFACTURER ON (B)(6) 2015. ANALYSIS IS UNDERWAY BUT IT HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS OF THE RETURNED USB CABLE WAS COMPLETED ON 04/16/2015 AND IT CONFIRMED A DISCONNECTED WIRE CONNECTION IN THE SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. | 1
4961313 | PRODUCT ANALYSIS OF THE RETURNED SUSPECT USB SERIAL CABLE IDENTIFIED AN INTERNAL DISCONNECTED WIRE CONNECTION. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PRINTED CIRCUIT BOARD, NO FURTHER ANOMALIES WERE IDENTIFIED.><UDI: (B)(4).><IT WAS REPORTED THAT A USER'S TABLET PROGRAMMER REPEATEDLY DISPLAYED AN ERROR MESSAGE OF, ERROR ESTABLISHING COMMUNICATION." SWITCHING WANDS DID NOT SOLVE THE PROBLEM. THE USER REPLACED THE TABLET'S USB CABLE WITH AN ALTERNATE USB CABLE AND NORMAL OPERATION WAS OBSERVED. THE SUSPECT USB CABLE HAS BEEN RETURNED TO THE MANUFACTURER AND IS CURRENTLY UNDERGOING ANALYSIS." | 1
4763180 | REVIEW OF DEVICE PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2011 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2011. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2011. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2011 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2012. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 1
5064623 | TYPE OF DEVICE, NAME, CORRECTED DATA: PROGRAMMING SOFTWARE. THE SUSPECT DEVICE NAME WAS PROVIDED AS PROGRAMMING COMPUTER IN THE INITIAL MDR. HOWEVER PRODUCT ANALYSIS DETERMINED THAT THE SUSPECT DEVICE WAS PROGRAMMING SOFTWARE, WHICH WAS INADVERTENTLY LEFT OUT IN THE SUPPLEMENTAL MDR #1.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO HARDWARE ANOMALIES ASSOCIATED WITH THE TABLET WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS CONFIRMED THAT THE TABLET WOULD NOT EMIT A COMPLETION NOISE FOLLOWING AN INTERROGATION. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH THE DEVICE DRIVER NOT BEING PROPERLY INSTALLED DURING THE OS INSTALLATION. THIS ANOMALY IS A KNOWN ISSUE, AND IS ADDRESSED IN THE UPGRADE PROCEDURE. REINSTALLING THE OS IS THE RECOMMENDED FIX FOR THIS CONDITION. NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE.><IT WAS REPORTED THAT THE PHYSICIAN'S PROGRAMMING TABLET WAS UPGRADED BY A FIELD TECHNICIAN WITHOUT ISSUE. HOWEVER, WHEN A DEMO DEVICE WAS INTERROGATED SUCCESSFULLY, THERE WAS NO SOUND CUE EMITTED FROM THE TABLET TO INDICATE INTERROGATION SUCCESSFUL. IT IS UNKNOWN IF THE INTERROGATION SUCCESSFUL CUE SOUND WAS EVER HEARD PRIOR TO THE UPGRADE PROCESS. THE TABLET WAS RECEIVED FOR ANALYSIS BUT IT HAS NOT BEEN COMPLETED YET.><. | 1
5253787 | IT WAS REPORTED THAT THE HANDHELD DEVICE FROZE ON THE INTERROGATION SCREEN. FOLLOW-UP REVEALED THAT THE ISSUES HAD OCCURRED IN THE PAST AND THAT THEY WOULD RESOLVE AFTER PERFORMING A HARD RESET. THE PROGRAMMING SYSTEM WAS RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><.><ANALYSIS OF THE HANDHELD AND SOFTWARE FLASHCARD CONCLUDED THAT THE DEVICES PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
4804071 | AN ANALYSIS WAS PERFORMED ON THE RETURNED USB AND TWO DISCONNECTED WIRE CONNECTIONS WERE FOUND. ONCE THE WIRES WERE SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><=DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE MEDICAL PROFESSIONAL COULD NOT INTERROGATE PULSE GENERATORS WITH THE PROGRAMMING SYSTEM. THE MEDICAL PROFESSIONAL USED A DIFFERENT PROGRAMMING COMPUTER WITH THE WAND, AND INTERROGATIONS COULD BE COMPLETED. FURTHER TROUBLESHOOTING WAS RUN USING THE MEDICAL PROFESSIONAL¿S WAND WITH A KNOWN GOOD TABLET AND INTERROGATIONS COULD BE COMPLETED. ANOTHER KNOWN GOOD WAND WAS USED WITH THE MEDICAL PROFESSIONAL¿S TABLET AND INTERROGATION COULD NOT BE COMPLETED, WHICH INDICATED THAT THE SOURCE OF THE PROBLEM WAS THE IN THE MEDICAL PROFESSIONAL¿S TABLET OR THE ACCOMPANYING USB CABLE. FURTHER INFORMATION WAS LATER RECEIVED INDICATING THAT FURTHER TROUBLESHOOTING HAD DETERMINED THAT THE MEDICAL PROFESSIONAL¿S USB CABLE WAS THE SOURCE OF THE COMMUNICATION ISSUES. THE SUSPECT USB CABLE WAS RETURNED TO THE MANUFACTURER. ANALYSIS OF THE USB CABLE IS UNDERWAY, BUT IT HAS NOT BEEN COMPLETED TO DATE.><. | 1
5015297 | (B)(4).><IT WAS REPORTED THAT AN ERROR MESSAGE UNABLE TO OPEN PORT" WAS BEING RECEIVED ON A PHYSICIAN'S TABLET FOR MULTIPLE PATIENTS. THE PHYSICIAN CALLED THE COMPANY REPRESENTATIVE | 1
4816608 | (B)(4): DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><DURING THE INVESTIGATION IN MANUFACTURER REPORT # (B)(4) INVOLVING A SOFTWARE FLASHCARD, ANALYSIS OF THE RETURNED HANDHELD DEVICE IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ESTABLISH COMMUNICATION WITH A KNOWN GOOD WAND AND GENERATOR. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH 3 BROKEN WIRE CONNECTIONS IN THE SERIAL CABLE. ONCE THE WIRES WERE SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. | 1
4664532 | DURING REVIEW OF THE PATIENT¿S PROGRAMMING HISTORY AVAILABLE IN THE MANUFACTURER¿S DATABASE, IT WAS IDENTIFIED THAT A PROGRAMMING ANOMALY OCCURRED ON (B)(6) 2014. A SYSTEM DIAGNOSTIC TEST WAS PERFORMED DURING PATIENT¿S VNS IMPLANT SURGERY AND THE PATIENT LEFT THE HOSPITAL AT THE UNINTENDED SETTINGS. THE PATIENT¿S GENERATOR WAS INTERROGATED FINALLY BUT THE SETTINGS WERE NOT CORRECTED THE SAME VISIT. MANUFACTURER LABELING INDICATES TO PERFORM A FINAL INTERROGATION TO CONFIRM AND CORRECT THE SETTINGS PRIOR TO PATIENTS LEAVING THE APPOINTMENT.><ANALYSIS OF PROGRAMMING HISTORY. | 1
4913626 | .><THE NURSE REPORTED THAT SHE WAS UNABLE TO ADVANCE TO THE INTERROGATION SCREEN FROM THE MAIN MENU. SHE WAS UNABLE TO ADVANCE PAST THE FIRST SCREEN ON HER HANDHELD. DURING THE REPORT, THE SCREEN LOCK WAS SWITCHED TO THE LOCK POSITION AND THEN BACK TO THE UNLOCK POSITION. A HARD RESET WAS ATTEMPTED, AND THE EDGES OF THE SCREEN WERE CLEANED. ANOTHER HARD RESET WAS PERFORMED. IT WAS CONFIRMED THAT THE STYLUS WAS BEING USED, BUT WAS STILL UNABLE TO ADVANCE PAST THE SCREEN. THE HANDHELD AND RELATED SOFTWARE WERE RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS WAS COMPLETED ON THE HANDHELD AND SOFTWARE, AND THE REPORTED ALLEGATION WAS CONFIRMED. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 1
5254940 | A HANDHELD DEVICE WHICH WAS RETURNED DUE TO UPGRADE TO THE LATEST PROGRAMMING SYSTEM WAS ANALYZED. DURING PRODUCT ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A SWOLLEN MAIN BATTERY. THE SWOLLEN BATTERY BENT THE BATTERY COVER CAUSING THE BATTERY LATCH SWITCH TO REGISTER THAT THE LATCH WAS UNLOCKED, THUS PREVENTING THE HANDHELD FROM POWERING ON. THE DEVICE FUNCTIONED NORMALLY WHEN USING A MODIFIED COVER WHICH LATCHED PROPERLY. NO FURTHER ANOMALIES WERE IDENTIFIED.><. | 1
4786633 | THE TABLET WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><.><IT WAS REPORTED THAT THE PROGRAMMING TABLET WAS DISPLAYING A LINE DOWN THE MIDDLE AFTER HE TURNS IT ON. THEN AFTER THE COMPANY SOFTWARE WAS DISPLAYED, THE TABLET WOULD SHUT DOWN INADVERTENTLY. HE CONFIRMED THAT HE COULD TURN THE DEVICE BACK ON, AND THE SAME ISSUE WOULD OCCUR. THIS WAS REPEATEDLY HAPPENING BEFORE HE WAS ABLE TO USE THE VNS SOFTWARE. A REPLACEMENT TABLET WAS PROVIDED. THE SUSPECT TABLET HAS NOT BEEN RECEIVED TO DATE.><.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATIONS WERE NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE TABLET PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS . A VISUAL INSPECTION OF THE TABLET DID VERIFY THAT THE CASE WAS DAMAGED. A CAUSE FOR THE DAMAGE COULD NOT BE IDENTIFIED BASED ON THE AVAILABLE INFORMATION. THE DAMAGE TO THE CASE HAD NO FUNCTIONAL IMPACT IN THE TABLET PERFORMANCE. | 1
4911500 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING AN INTERNAL REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS IDENTIFIED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2014. THE SETTINGS WERE NOT CORRECTED ON THE SAME DATE. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 1
4407912 | MANUFACTURER DEVICE HISTORY RECORDS WERE REVIEWED.REVIEW OF THE HANDHELD DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY SPECIFICATIONS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN¿S HANDHELD DID NOT HOLD A CHARGE. NO TROUBLESHOOTING WAS REPORTED. THE PHYSICIAN HAS SEVERAL PROGRAMMING HANDHELDS, SO IT IS NOT BELIEVED THAT ANY PATIENTS WERE AFFECTED. THE HANDHELD AND RELATED SOFTWARE WERE RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD, AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ADDITIONALLY, NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. | 1
4530117 | REVIEW OF PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2013 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2012. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2012 DID NOT IDENTIFY ANY DIAGNOSTICS. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 1
4929792 | THE USB SERIAL CABLE WAS RECEIVED, AND ANALYSIS WAS COMPLETED AND APPROVED. NO ANOMALIES WERE NOTED, AND THE USB SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT A PHYSICIAN'S NEWLY DELIVERED USB CABLE FOR HIS TABLET WAS FOUND NON-FUNCTIONAL (MFR. REPORT # 1644487-2015-05191). TROUBLESHOOTING WAS PERFORMED AND IDENTIFIED THE USB CABLE AS THE ISSUE. A COMPANY REPRESENTATIVE LOANED HIS USB CABLE TO THE PHYSICIAN, WHICH ALSO BEGAN NOT FUNCTIONING PROPERLY (MFR. REPORT # 1644487-2015-05186). TROUBLESHOOTING WAS AGAIN PERFORMED AND IDENTIFIED THE USB CABLE AS THE ISSUE. THE PHYSICIAN WAS PROVIDED A NEW USB CABLE THAT WAS ALSO FOUND TO BE NON-FUNCTIONAL. THE COMPANY REPRESENTATIVE ATTEMPTED TO USE THE PHYSICIAN'S NEW USB CABLE WITH HIS PROGRAMMING SYSTEM, BUT THE USB CABLE WAS STILL NON-FUNCTIONAL. THE PHYSICIAN AND COMPANY REPRESENTATIVE WERE PROVIDED NEW USB CABLES THAT FUNCTIONED PROPERLY. THE PHYSICIAN'S SECOND FAULTY USB CABLE HAS BEEN RECEIVED, BUT PRODUCT ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. (B)(4). | 1
4764437 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><THE TABLET WAS RETURNED TO THE MANUFACTURER. AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND IT WAS NOTED THAT THE DEVICE DID NOT HAVE THE POWER BUTTON. NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE.><(B)(4). DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE TABLET PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE POWER BUTTON OF A MEDICAL PROFESSIONAL¿S TABLET HAD DETACHED FROM THE DEVICE AND IT WAS LOST. THE TABLET COULD NOT BE SWITCHED ON AFTERWARDS. DEVICE MANUFACTURING RECORDS CONFIRMED THAT THE TABLET PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. THE TABLET HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE. | 1
5260456 | DURING AN INTERNAL REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT A PROGRAMMING ANOMALY OCCURRED ON A VNS PATIENT¿S DEVICE THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS FOUND IN OFFICE VISIT ON (B)(6) 2015. THE NORMAL MODE AND MAGNET MODE WERE CHANGED TO 0 MA. NO PATIENT ADVERSE EVENTS WERE REPORTED.><. | 1
4434364 | NO FAILURE OCCURRED IN THE FIELD WITH THE HANDHELD, ONLY DURING PRODUCT ANALYSIS WAS IT OBSERVED THAT THE BATTERY COVER WAS MISSING.><DESCRIBE EVENT OR PROBLEM; CORRECTED DATA: NO FAILURE OCCURRED IN THE FIELD WITH THE HANDHELD AND THERE IS NO FAILURE OF THE DEVICE.><ON (B)(6) 2015 PRODUCT ANALYSIS WAS COMPLETED ON A RETURNED HANDHELD. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WAS RETURNED WITHOUT A BATTERY COVER. AS A RESULT, THE HANDHELD WOULD NOT POWER ON. ONCE A KNOWN BATTERY COVER WAS INSTALLED, THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 1
4913436 | IT WAS REPORTED ON (B)(6) 2015 THAT THE PHYSICIAN¿S HANDHELD WAS NOT WORKING. IT WAS INITIALLY REPORTED THAT THE PROGRAMMING SYSTEMS WANTED THE PHYSICIAN¿S OFFICE TO ¿REPROGRAM IT,¿ BUT IT WOULD NOT LET THE OFFICE STAFF PERFORM IT. THE VNS FIELD REPRESENTATIVE FOLLOWED-UP, AND REPORTED THAT THE HANDHELD WOULD NOT POWER ON. THE VNS FIELD REPRESENTATIVE HAD THE HANDHELD PLUGGED INTO THE POWER OUTLET, BUT IT WOULD STILL NOT POWER ON. THE SUSPECT HANDHELD HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR ANALYSIS TO DATE.><THE INITIAL REPORT INADVERTENTLY REPORTED THIS DATA INCORRECTLY.><THE HANDHELD DEVICE AND RELATED SOFTWARE WERE RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><.><ANALYSIS WAS COMPLETED, AND THE FAILURE TO PROGRAM THE HANDHELD ON WAS VERIFIED. THE CAUSE FOR THE REPORTED ANOMALY IS ASSOCIATED WITH A SWOLLEN MAIN BATTERY. THE SWOLLEN BATTERY BENT THE BATTERY COVER CAUSING THE BATTERY LATCH SWITCH TO REGISTER THAT THE LATCH WAS UNLOCKED, THUS PREVENTING THE HANDHELD FROM POWERING ON. THIS CONDITION CAN ALSO CAUSE THE HANDHELD TO POWER OFF INTERMITTENTLY. THE CAUSE IS NOT KNOWN BUT MAY BE RELATED TO WEAR AND/OR USAGE OF BATTERY. NO ANOMALIES WERE IDENTIFIED WITH THE SOFTWARE. | 1
5238129 | IT WAS REPORTED THAT A HANDHELD DEVICE BACKLIGHT WAS NOT WORKING. THE DEVICE WOULD POWER ON BUT THE BACKLIGHT WAS REPORTEDLY NOT FUNCTIONING PROPERLY. THE DEVICE WAS RETURNED TO THE MANUFACTURER BUT WAS INADVERTENTLY DISPOSED OF BEFORE PRODUCT ANALYSIS COULD BE PERFORMED.><. | 1
4669865 | AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION WAS VERIFIED. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE USB PORT WAS DAMAGED. AS A RESULT, THE TABLET WAS UNABLE TO ESTABLISH COMMUNICATION. NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED PROGRAMMING TABLET WAS BROKEN. THE USB PORT ON THE SIDE OF THE TABLET WAS REPORTED TO BE DAMAGED. A REPLACEMENT TABLET WAS PROVIDED. THE PHYSICIAN¿S ASSISTANT DOES NOT KNOW WHAT CAUSED THE DAMAGE. HOWEVER, THE WAND CANNOT CONNECT TO THE TABLET AS A RESULT, BUT THE TABLET STILL TURNS ON. THE TABLET WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 1
4874210 | IT WAS REPORTED THAT THE PROGRAMMING TABLET DID NOT WORK DUE TO RECEIVING A PORT ERROR MESSAGE. NO KNOWN TROUBLESHOOTING WAS PERFORMED. A REPLACEMENT USB SERIAL ADAPTER CABLE WAS PROVIDED. THE SUSPECT TABLET USB SERIAL CABLE HAS NOT BEEN RECEIVED BY THE MANUFACTURER. IT IS THOUGHT THAT THE CABLE MAY HAVE BEEN DISCARDED BY THE PHYSICIAN'S OFFICE.><. | 1
5226333 | .><THE PHYSICIAN REPORTED THAT SHE BELIEVES THAT THE FAULTED SYSTEM DIAGNOSTICS TEST WAS LIKELY DUE TO THE FACT THAT SHE WAS JUST LEARNING TO USE THE SYSTEM. THE PATIENT'S SETTINGS WERE RE-TESTED ON (B)(6) 2015 AND CONFIRMED NORMAL AS WELL AS ON (B)(6) 2015.><ON (B)(6) 2015 THE PATIENT'S PROGRAMMING HISTORY WAS REVIEWED. IT WAS FOUND THAT ON THE VISIT DATED (B)(6) 2015 THE PATIENT'S INITIAL INTERROGATION SHOWED SETTINGS INDICATIVE OF A FAULTED SYSTEM DIAGNOSTICS TEST. AT THE PATIENT'S PREVIOUS VISIT IN THE PROGRAMMING HISTORY ON (B)(6) 2013 A FINAL INTERROGATION WAS PERFORMED WHICH SHOWED THE PATIENT NOT PROGRAMMED TO THESE FAULTED SETTINGS. IT IS UNKNOWN WHEN THE FAULTED DIAGNOSTIC TEST OCCURRED. GOOD FAITH ATTEMPTS FOR ADDITIONAL PROGRAMMING HISTORY FROM THE PHYSICIAN WERE UNSUCCESSFUL. | 1
4925696 | .><IT WAS REPORTED THAT WHEN THE PATIENT'S DEVICE WAS INTERROGATED ON (B)(6) 2015 AND THAT THE SETTINGS IDENTIFIED WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PATIENT'S PREVIOUS VISIT ON (B)(6) 2015. THE COMPANY REPRESENTATIVE REPORTED THAT WHEN THE DEVICE WAS ATTEMPTED TO BE INTERROGATED ON (B)(6) 2015 THE TABLET BATTERY DIED (MFR. REPORT # 1644487-2015-05257) AND THE INTERROGATION WAS NOT COMPLETED. THE SETTINGS WERE CORRECTED ON (B)(6) 2015. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE TABLET DEVICE MENTIONED IN THE PREVIOUSLY SUBMITTED MDR WAS NOT INVOLVED IN THE PROGRAMMING SESSION ON (B)(4) 2015. | 1
4847311 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WOULD NOT POWER ON AFTER BEING PLUGGED IN AND CHARGING. THE CHARGING INDICATOR LIGHT WAS ILLUMINATED WHEN THE HANDHELD WAS PLUGGED IN TO AN ELECTRICAL OUTLET. A COMPANY REPRESENTATIVE PERFORMED TROUBLESHOOTING BY ENSURING THE SCREEN LOCK WAS NOT ENGAGED AND THE BATTERY COVER LATCH WAS CONFIRMED TO BE IN THE LOCKED POSITION. THE BATTERY WAS ATTEMPTED TO BE REMOVED BY UNLOCKING THE LATCH; HOWEVER, THE COVER WAS UNABLE TO BE REMOVED. THE LATCH LOCK WAS AGAIN PLACED IN THE LOCK POSITION AND THE HANDHELD WAS ATTEMPTED TO BE POWERED ON; HOWEVER, THE HANDHELD WOULD STILL NOT POWER ON. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET AND THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS OF THE HANDHELD WAS COMPLETED ON 07/09/2015. NO ANOMALIES ASSOCIATED WITH THE BATTERY LATCH WERE IDENTIFIED DURING THE ANALYSIS. THE RETURNED HANDHELD WAS ABLE TO POWER ON AND OFF. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE HANDHELD DISPLAY WOULD NOT SHOW AN IMAGE WHEN THE DEVICE WAS POWERED ON. THE BACKLIGHT WOULD TURN ON, BUT THE DISPLAY IMAGE WAS SOLID WHITE. THE CAUSE FOR THE IDENTIFIED ANOMALY IS ASSOCIATED WITH A DEFECTIVE DISPLAY. NO FURTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON 07/09/2015. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 1
5060553 | ANALYSIS OF THE SOFTWARE OF THE HANDHELD WAS COMPLETED AND IT CONFIRMED PROPER FUNCTIONALITY OF BOTH DEVICES. NO ANOMALIES WERE OBSERVED. | 1
5050242 | IT WAS REPORTED BY THE PHYSICIAN THE HANDHELD PROGRAMMER HAD ISSUES WITH THE SCREEN FREEZING OVER THE PAST YEAR, INTERMITTENTLY, WHILE CHECKING PATIENTS. IT WAS NOTED THE HANDHELD HAD TO HAVE A HARD RESTART MULTIPLE TIMES. THE PHYSICIAN HAS REQUESTED A REPLACEMENT AND THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET. THE HANDHELD PROGRAMMER IS EXPECTED TO BE RETURNED TO THE MANUFACTURER, BUT HAS NOT BEEN RECEIVED TO DATE.><THE HANDHELD DEVICE AND RELATED SOFTWARE WERE RECEIVED ON 09/23/2015. PRODUCT ANALYSIS IS EXPECTED BUT HAS NOT BEEN COMPLETED TO DATE.><THIS INFORMATION WAS ORIGINALLY INCORRECTLY REPORTED ON THE INITIAL MFR. REPORT.><.><ANALYSIS FOR THE RETURNED HANDHELD DEVICE REVEALED NO ANOMALIES WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD DEVICE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS FOR THE RETURNED SOFTWARE WAS COMPLETED AND CONFIRMED THE EVENT OF THE SCREEN FREEZING DURING INTERROGATION. THE SCREEN FREEZING IS A KNOWN ISSUE THAT HAS BEEN PREVIOUSLY INVESTIGATED BY THE MANUFACTURER AND THE ROOT CAUSE WAS DETERMINED TO BE DUE TO A FAILURE IN THE API FUNCTION THAT RESULTS IN CONTROL OF THE DEVICE NOT RETURNING TO THE CYBERONICS SOFTWARE. THE MOST PROBABLE CAUSE IS AN ANOMALY IN THE INTERFACE BETWEEN THE MICROSOFT PROVIDED SOFTWARE AND THE DELL HARDWARE WHICH RESULTS IN THE OS'S INABILITY TO LOCATE SPECIFIC MEMORY DIRECTORIES WITHIN THE FILE SYSTEM. NO ADDITIONAL ISSUES WERE OBSERVED WITH THE RETURNED FLASHCARD. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
5297142 | .><THE SUSPECT DEVICE WAS RECEIVED BY THE MANUFACTURER. ANALYSIS OF THE DEVICE IS UNDERWAY, BUT IT HAS NOT BEEN COMPLETED TO DATE.><IT WAS REPORTED THAT A MOTION TABLET WITH V11.0 SOFTWARE COULD NOT DO AN SR DOWNLOAD. THE DEVICE DISPLAYED THE FOLLOWING ERROR MESSAGE WHEN ATTEMPTING TO PERFORM AN EXTENDED GENERATOR MEMORY DOWNLOAD: ¿FILE SAVE FAILURE. PLEASE ENSURE A SECURE DIGITAL CARD WITH SUFFICIENT FREE SPACE IS INSERTED IN THE DEVICE BEFORE ATTEMPTING TO SAVE A FILE. NOTE: ONLY CYBERONICS APPROVED SD CARDS SHOULD BE USED.¿ THE TABLET IN QUESTION IS A DEMO PRODUCT, AND TROUBLESHOOTING WAS ATTEMPTED WHICH INCLUDED RETRYING, REBOOTING, VERIFYING THE SD CARD WAS NOT LOCKED, AND VERIFYING THAT THERE WAS SUFFICIENT SPACE AVAILABLE ON THE CARD; HOWEVER, THE ISSUE PERSISTED. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE TABLET PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. NO ADDITIONAL RELEVANT INFORMATION WAS RECEIVED TO DATE.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET IT WAS IDENTIFIED THAT THE TABLET WAS UNABLE TO RECOGNIZE THE SD CARD. AS A RESULT, THE DATABASE UTILITIES BUTTON WAS DISABLED IN THE VNS SOFTWARE. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A HARDWARE FAILURE ON THE MAIN BOARD. THE ANOMALY WILL ALSO PREVENT THE TABLET (OR VNS SOFTWARE) FROM ACCESSING AND WRITING TO THE SD CARD. ALTHOUGH THE DATABASE UTILITIES BUTTON WAS DISABLED, THE VNS SOFTWARE STILL HAD LIMITED FUNCTIONALITY AND WAS ABLE TO INTERROGATE, PROGRAM, AND PERFORM DIAGNOSTIC TESTS DURING THE ANALYSIS. | 1
4608040 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF THE PATIENT¿S PROGRAMMING HISTORY AVAILABLE IN THE MANUFACTURER¿S DATABASE, IT WAS IDENTIFIED THAT A PROGRAMMING ANOMALY OCCURRED. A SYSTEM DIAGNOSTIC TEST WAS PERFORMED, AND A FINAL INTERROGATION WAS NOT PERFORMED AND THE PATIENT LEFT THE VISIT. MANUFACTURER LABELING INDICATES TO PERFORM A FINAL INTERROGATION TO CONFIRM INTENDED SETTINGS PRIOR TO PATIENTS LEAVING THE CLINIC. UPON INITIAL INTERROGATION ON THE SUBSEQUENT VISIT, THE DEVICE WAS AT UNINTENDED SETTINGS. | 1
4753050 | THE SUSPECT TABLET WAS RECEIVED ON 08/17/2016. THE TABLET WAS REPORTED TO HAVE BEEN DROPPED IN (B)(6) 2016, RESULTING IN THE CRACKED SCREEN. THE TABLET WAS RETURNED FOR THE CRACKED SCREEN AND NOT FOR THE PREVIOUSLY REPORTED FAILURE TO TURN ON ISSUE. AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION OF UNABLE TO POWER ON OR OFF WAS NOT VERIFIED. THE TABLET WAS ABLE TO POWER ON AND OFF WITH NO OBSERVED ANOMALIES. THE REPORTED ALLEGATION OF CRACKED SCREEN WAS VERIFIED. A VISUAL INSPECTION OF THE TABLET VERIFIED THAT THE DISPLAY WAS CRACKED. ALTHOUGH THE SCREEN WAS CRACKED, THE TABLET WAS STILL ABLE TO COMPLETE TESTING WITH NO OBSERVED ANOMALIES. NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET WERE IDENTIFIED DURING THE ANALYSIS.><IT WAS REPORTED THAT A TABLET WAS UNABLE TO BE TURNED ON DESPITE HAVING PLUGGED INTO THE WALL. THE BATTERY ICON LIGHT WAS LIT WHILE PLUGGED INTO THE OUTLET. THE A/C CHARGER WAS CONFIRMED TO BE WORKING AS IT DISPLAYED A GREEN LIGHT WHILE BEING CONNECTED TO POWER SOURCE. THE POWER BUTTON WAS CONFIRMED TO BE CORRECTLY ORIENTED. A REPLACEMENT TABLET WAS SENT BUT THE SUSPECT DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE.><(B)(4).><ADDITIONAL INFORMATION WAS RECEIVED THAT THE TABLET WAS WORKING PROPERLY. THE TABLET IS NOT EXPECTED TO BE RETURNED AS THE REPORTED ISSUE WAS CONFIRMED INVALID. | 1
5167988 | .><IT WAS REPORTED THAN A MEDICAL PROFESSIONAL COULD NOT INTERROGATE GENERATORS WHEN USING HIS PROGRAMMING SYSTEM. TROUBLESHOOTING WAS PERFORMED AND THE CONNECTION OF THE USB CABLE TO THE TABLET WAS SUSPECTED TO BE DEFECTIVE. THE RETURN OF THE SUSPECT USB CABLE IS EXPECTED BUT IT HAS NOT BEEN RECEIVED TO DATE. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE TABLET PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION.><THE SUSPECTED USB CABLE WAS RETURNED TO THE MANUFACTURER ON 10/23/2015. ANALYSIS IS UNDERWAY BUT IT HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS OF THE RETURNED USB CABLE WAS COMPLETED AND CONFIRMED BROKEN LEAD AT DB9 INTERFACE OF DB9 TO USB SERIAL ADAPTER COMMUNICATION CABLE. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><LOT#; CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR INADVERTENTLY PROVIDED THE WRONG SUSPECT DEVICE FOR THE EVENT. | 1
4772575 | THE USB TO DB9 CABLE WAS RETURNED TO THE MANUFACTURER ON 05/21/2015. ANALYSIS WAS COMPLETED ON 06/17/2015 ON THE RETURNED USB TO DB9 CABLE AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE SERIAL CABLE PERFORMANCE WERE NOTED DURING TESTING. THE SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><.><IT WAS REPORTED THAT THE TABLET DEVICE DISPLAYED AN ¿UNABLE TO OPEN PORT¿ ERROR MESSAGE INDICATING A POTENTIAL ISSUE WITH THE USB SERIAL CABLE. A REPLACEMENT CABLE WAS PROVIDED WHICH WORKED INITIALLY; HOWEVER, THE ERROR MESSAGE AGAIN APPEARED. THE TABLET DEVICE WAS POWERED OFF AND ON AND THE REPLACEMENT USB SERIAL CABLE WAS DISCONNECTED FOR 15 SECONDS AND REINSERTED INTO THE USB PORT; HOWEVER, THE ISSUE DID NOT RESOLVE. THE CABLE WAS TESTED AGAIN WITH A DIFFERENT TABLET AND PROGRAMMING THAT WERE KNOWN TO FUNCTIONING PROPERLY, AND THE TABLET DEVICE STILL DISPLAYED THE SAME ERROR MESSAGE. THIS MANUFACTURER REPORT INVOLVES THE REPLACEMENT SERIAL CABLE. THE ORIGINAL SERIAL CABLE WILL BE CAPTURED IN MANUFACTURER REPORT #1644487-2015-04713. THE SERIAL CABLE INVOLVED IN THIS MANUFACTURER REPORT HAS NOT BEEN RETURNED TO DATE. | 1
4578054 | IT WAS REPORTED THAT THE PHYSICIAN&#39;S HANDHELD HAD FREQUENT BLACK SCREENS. FREQUENT TROUBLESHOOTING WAS PERFORMED AND IT WAS REPORTED THAT THE HANDHELD IS NO LONGER RELIABLE. A NEW PROGRAMMING TABLET WAS PROVIDED TO THE PHYSICIAN. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><THE HANDHELD AND THE SOFTWARE FLASHCARD WERE RECEIVED BY THE MANUFACTURER ON 03/10/2015. ANALYSIS OF THE HANDHELD FOUND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WHEN USED WITH THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD SOFTWARE FOUND NO ANOMALIES. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE MANUFACTURING RECORDS WERE REVIEWED.REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION. | 1
5054159 | .><ADDITIONAL INFORMATION WAS RECEIVED THAT THE FACILITY DISCARDED THE CABLE. THE DEVICE WAS REPORTED TO HAVE SHOWN UNABLE TO OPEN PORT" ERROR MESSAGE. TROUBLE SHOOTING WAS HANDLED BY USING AN ADDITIONAL CABLE FROM BACK UP DEVICE. PRODUCT EVALUATION COULD NOT BE PERFORMED AS THE SUSPECT DEVICE WAS DISCARDED.><THE SUSPECTED USB CABLE WAS RECEIVED ON 10/2/2015. NO ANOMALIES ASSOCIATED WITH THE SERIAL CABLE PERFORMANCE WERE NOTED DURING PRODUCT ANALYSIS. THE SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><(B)(4).><IT WAS REPORTED THAT A PHYSICIAN HAS A MALFUNCTIONING SERIAL ADAPTER CABLE ASSOCIATED WITH A TABLET. NO PATIENTS WERE AFFECTED AS THEY HAD A BACK UP TABLET TO FIX THE SITUATION. ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION WERE UNSUCCESSFUL TO DATE." | 1
4933281 | IT WAS REPORTED THAT THE MANUFACTURER'S SALES REPRESENTATIVE'S VNS PROGRAMMING SYSTEM WAS RECEIVING AN UNABLE TO OPEN PORT"; ERROR MESSAGE WHILE TRYING TO INTERROGATE A DEMO GENERATOR. THE ISSUE WAS LOCALIZED TO THE USB TO DB9 SERIAL ADAPTER CABLE BY USING A KNOWN GOOD PROGRAMMING TABLET AND WAND AND REPRODUCING THE ERROR. A NEW USB TO DB9 SERIAL ADAPTER CABLE WAS SENT TO THE SALES REPRESENTATIVE WHICH RESOLVED THE ISSUE. THE SUSPECT SERIAL ADAPTER CABLE HAS BEEN RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER | 1
5328978 | AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION WAS VERIFIED. NO ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE TABLET PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><(B)(4).><IT WAS REPORTED THAT A MOTION TABLET FOR A MEDICAL PROFESSIONAL WAS NOT TURNING ON. IT WAS REPORTED THAT THE WAND BATTERY TEST WAS FINE. TROUBLESHOOTING WAS PERFORMED WITHOUT SOLVING THE ISSUE. BY SWITCHING THE TABLET ON, IT WAS LOCKED ON A BLANK SCREEN WITH A SMALL ARROW-CURSOR IN THE CENTRE, AND THE TABLET WAS UNRESPONSIVE. IT WAS REPORTED THAT THE TABLET HAD BEEN PREVIOUSLY LEFT ON CHARGE. RETURN TO THE MANUFACTURER OF THE SUSPECTED TABLET IS EXPECTED BUT IT HAS NOT BEEN RECEIVED TO DATE. REVIEW OF MANUFACTURING RECORDS CONFIRMED ALL TESTS PASSED FOR THE DEVICE PRIOR TO DISTRIBUTION. NO NONCONFORMANCES WERE OBSERVED.><DESCRIBE EVENT OR PROBLEM; CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR INCORRECTLY INDICATED THAT THE REPORTED ALLEGATION ON THE SUSPECT DEVICE WAS VERIFIED.><THE SUSPECT TABLET WAS RECEIVED BY THE MANUFACTURER. ANALYSIS OF THE DEVICE IS UNDERWAY, BUT IT HAS NOT BEEN COMPLETED TO DATE.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION WAS NOT VERIFIED. | 1
4956624 | THE HANDHELD DEVICE AND SOFTWARE FLASHCARD HAVE BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY. THE REASON FOR RETURN WAS STATED AS IT WOULD GET AN ERROR MESSAGE AND FREEZE".><AN ANALYSIS OF THE HANDHELD IDENTIFIED THAT THE MAIN BATTERY WAS UNABLE TO HOLD A CHARGE AND POWER THE HANDHELD. THE CAUSE FOR THE BATTERY FAILURE IS ASSOCIATED WITH A SWOLLEN BATTERY WHICH WAS CAUSED BY WEAR AND AGE OF THE BATTERY. NO FURTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS; THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ERROR MESSAGES OR SCREEN FREEZES OCCURRED DURING THE PRODUCT ANALYSIS.><IT WAS REPORTED THAT THE HANDHELD DEVICE WAS HAVING DIFFICULTY HOLDING A CHARGE. THE HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE. NO FURTHER INFORMATION RELEVANT TO THE EVENT HAS BEEN RECEIVED TO DATE.><.><A COMPANY REPRESENTATIVE REPORTED THAT THE USER WAS RECEIVING ERROR MESSAGES. THE ISSUE WAS NOT RESOLVED AND A REPLACEMENT SYSTEM WAS REQUESTED." | 1
5001183 | DURING THE INVESTIGATION IN MANUFACTURER REPORT #1644487-2015-04778, A REPLACEMENT USB SERIAL CABLE FOR THE TABLET DEVICE WAS PROVIDED TO THE PHYSICIAN; HOWEVER, THE ISSUES DID NOT RESOLVE. ANOTHER SERIAL CABLE THAT WAS KNOWN TO BE FUNCTIONING NORMALLY WAS USED WITH THE SUSPECT TABLET DEVICE BUT THE ISSUES CONTINUED. A REPLACEMENT TABLET DEVICE WAS PROVIDED AND THE SUSPECT TABLET DEVICE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET AND THE REPORTED ALLEGATION WAS VERIFIED. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE USB PORT WAS DAMAGED WHICH WAS DETERMINED TO MOST LIKELY BE USER RELATED BUT ULTIMATELY THE CAUSE COULD NOT BE ESTABLISHED. AS A RESULT, THE TABLET WAS UNABLE TO ESTABLISH COMMUNICATION. NO FURTHER ANOMALIES WERE IDENTIFIED.><(B)(4). | 1
4537977 | .><IT WAS REPORTED THAT THE PHYSICIAN¿S OFFICE HAD A PROGRAMMING HANDHELD THAT WOULD NOT CHARGE WHEN PLUGGED INTO THE POWER CABLE. IT WAS THEREFORE BELIEVED THAT THE ISSUE MAY BE WITH THE POWER CABLE AND A NEW POWER CABLE WAS SENT TO THE PHYSICIAN, BUT THIS DID NOT RESOLVE THE ISSUE WITH THE PROGRAMMING HANDHELD. IT WAS REPORTED THAT THE PROGRAMMING HANDHELD WOULD BE RETURNED FOR PRODUCT ANALYSIS BUT IT HAS NOT BEEN RECEIVED TO DATE. IT WAS REPORTED THAT THE HANDHELD MAY HAVE GOTTEN ¿ZAPPED¿.><THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2015. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE MAIN BATTERY WAS INSTALLED BACKWARDS. AS A RESULT, THE HANDHELD WAS UNABLE TO CHARGE THE MAIN BATTERY. ADDITIONALLY, IT WAS IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON USING THE AC ADAPTER. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH DAMAGED SYNC CONNECTOR LEADS TO THE PCB. BECAUSE OF THE DAMAGE, THE HANDHELD WAS UNABLE TO RECEIVE POWER FROM THE AC ADAPTER. THE CAUSE FOR THE DAMAGE TO THE CONNECTOR LEADS AND SOLDER CONNECTIONS IS MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE AS IT APPEARS THE CONNECTOR DETENTS ARE NOT BEING COMPRESSED WHEN REMOVING OR MANIPULATING THE SERIAL DATA CABLE, THUS PLACING AN EXTENSIVE AMOUNT OF FORCE ON THE PCB AND ATTACHED CABLE RECEPTACLE. NO OTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2015. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
5099796 | IT WAS FOUND DURING INTERNAL REVIEW OF PROGRAMMING HISTORY DATA THAT THE PATIENT'S GENERATOR PRESENTED A DUTY CYCLE VALUE OF 1% ON (B)(6) 2014. THE DEVICE SETTINGS WERE CORRECTED ON (B)(6) 2014. NO PATIENT ADVERSE EVENTS WERE REPORTED.><CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR INADVERTENTLY PROVIDED THE WRONG EVALUATION CODES FOR THE EVENT.><. | 1
4524519 | THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2015. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON USING THE AC ADAPTER. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH DAMAGED LEADS ASSOCIATED WITH THE HANDHELD SYNC CONNECTOR. BECAUSE OF THE DAMAGE, THE HANDHELD WAS UNABLE TO RECEIVE POWER FROM THE AC ADAPTER. THE CAUSE FOR THE DAMAGE TO THE CONNECTOR LEADS AND SOLDER CONNECTIONS IS MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE AS IT APPEARS THE CONNECTOR DETENTS ARE NOT BEING COMPRESSED WHEN REMOVING OR MANIPULATING THE SERIAL DATA CABLE, THUS PLACING AN EXTENSIVE AMOUNT OF FORCE ON THE PCB AND ATTACHED CABLE RECEPTACLE. NO OTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2015. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(6) 2015 IT WAS REPORTED THAT THE HANDHELD MAY HAVE GOTTEN ¿ZAPPED¿.><.><IT WAS REPORTED THAT THE PHYSICIAN WAS HAVING DIFFICULTY TURNING ON THE PROGRAMMING HANDHELD. THE SCREEN LOCK BUTTON, THE BATTERY LATCH, AND THE WALL OUTLET WERE ALL CHECKED, WHICH DID NOT RESOLVE THE ISSUE. A SOFT RESET WAS PERFORMED BUT THE HANDHELD ONLY FLASHED THE RED LIGHT ON THE POWER BUTTON AND NOTHING APPEARED ON THE SCREEN. A HARD RESET ONLY RESULTED IN A ¿¿BATTERY CRITICALLY LOW¿ MESSAGE APPEARING ON THE SCREEN. THE OFFICE HAD ANOTHER PROGRAMMING HANDHELD AND WHEN THEY TRIED TO PLUG THAT ONE INTO THE POWER CABLE OF THE FIRST PROGRAMMING HANDHELD, THERE WAS AN INABILITY TO CHARGE THAT ONE AS WELL. IT WAS THEREFORE BELIEVED THAT THE ISSUE MAY BE WITH THE POWER CABLE AND A NEW POWER CABLE WAS SENT TO THE PHYSICIAN, BUT THIS DID NOT RESOLVE THE ISSUE WITH THE PROGRAMMING HANDHELD NOT TURNING ON. IT WAS REPORTED THAT THE PROGRAMMING HANDHELD WOULD BE RETURNED FOR PRODUCT ANALYSIS BUT IT HAS NOT BEEN RECEIVED TO DATE. | 1
4723831 | IT WAS REPORTED THAT THE TABLET DEVICE WAS ¿NOT WORKING.¿ THE SUSPECT TABLET DEVICE HAS NOT BEEN RETURNED TO DATE. ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><AN ANALYSIS WAS COMPLETED ON THE RETURNED PROGRAMMING TABLET. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE TABLET COULD NOT CHARGE THE MAIN BATTERY, AND AS A RESULT WAS ALSO NOT ABLE TO BE POWERED USING THE MAIN BATTERY. THE CAUSE FOR THE IDENTIFIED ANOMALY IS ASSOCIATED WITH A LOOSE BATTERY CABLE TO THE MOTHERBOARD. ONCE THE CABLE WAS RESEATED, NO FURTHER ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS.><.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><ADDITIONAL INFORMATION WAS RECEIVED THAT THE PHYSICIAN REPORTED THAT THE SCREEN WOULD SOMETIMES SWITCH IN THE MIDDLE OF INTERROGATION AND THAT THE SCREEN FREEZES. THE PHYSICIAN TROUBLESHOOTING USING DIFFERENT WANDS AND SERIAL CABLES ISOLATED THE ISSUE TO THE TABLET. THE TABLET WAS NOT PLUGGED INTO THE ELECTRICAL OUTLET DURING USE. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET WHICH WAS REPORTED TO BE FUNCTIONING AS INTENDED. THE SUSPECT TABLET HAS BEEN RETURNED FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE. | 1
4486900 | AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET USB TO DB9 SERIAL CABLE, AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE PROGRAMMING TABLET WAS REPORTING A FAILURE TO OPEN PORT" ERROR. NO PATIENT WAS AFFECTED AS AN ALTERNATIVE PROGRAMMING SYSTEM HAD BEEN USED. TROUBLESHOOTING WAS PERFORMED. TROUBLESHOOTING INCLUDED PERFORMING A HARD RESET | 1
5271395 | IT WAS REPORTED THAT THE SERIAL CABLE WAS DISCARDED; THEREFORE, NO ANALYSIS CAN BE PERFORMED.><THE COMPANY REPRESENTATIVE INDICATED THAT HE NEEDED A NEW USB CABLE FOR HIS TABLET AS HE WAS RECEIVING AN ERROR MESSAGE. A NEW USB CABLE WAS PROVIDED. THE FAULTY USB CABLE HAS NOT BEEN RECEIVED FOR ANALYSIS TO DATE.><(B)(4). | 1
5127405 | (B)(4).><IT WAS REPORTED THAT THE VNS GENERATOR WAS UNABLE TO BE INTERROGATED PRIOR TO A REPLACEMENT SURGERY. THE ISSUE RESOLVED AFTER REPLACING THE USB TO DB9 SERIAL ADAPTER CABLE ON THE PROGRAMMING TABLET. THE USB TO DB9 SERIAL ADAPTER CABLE HAS SINCE BEEN DISCARDED AND IS THUS UNABLE TO BE RETURNED FOR PRODUCT ANALYSIS. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN OBTAINED TO DATE. | 1
5000133 | AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND IT WAS IDENTIFIED THAT THE MAIN BATTERY WAS INSTALLED INCORRECTLY. ADDITIONALLY, IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD PROGRAMMER SCREEN IS FREEZING FROM TIME TO TIME BUT IS RESOLVED WITH A HARD RESET. THE PHYSICIAN IS HOWEVER DISSATISFIED WITH THE DEVICE AND THEREFORE A REPLACEMENT WAS PROVIDED. ADDITIONAL INFORMATION WAS RECEIVED THAT VARIOUS DISPLAY SCREENS HAVE BEEN FREEZING FOR APPROXIMATELY 9 MONTHS. HARD RESETS WERE PERFORMED TO TROUBLESHOOT THE ISSUE DURING THOSE TIMES. THE DEVICE WAS RECEIVED BUT ANALYSIS HAS NOT BEEN COMPLETED.><. | 1
5196155 | .><IT WAS REPORTED THAT A PHYSICIAN'S TABLET WAS NOT PROGRESSING PAST A BLACK SCREEN, WHICH PRESENTED AN ERROR MESSAGE WHICH PREVENTED FURTHER ACTION DURING AN ATTEMPTED SYSTEMS DIAGNOSTICS CHECK. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO-DATE. THE TABLET HAS BEEN RECEIVED AND IS PENDING ANALYSIS, WHICH HAS NOT BEEN COMPLETED TO-DATE.><ANALYSIS WAS PERFORMED ON THE RETURNED TABLET. NO ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE TABLET PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 1
4425692 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><ANALYSIS OF THE RETURNED SERIAL CABLE WAS COMPLETED ON 01/29/2015 AND IT FOUND THAT THE CAUSE FOR THE SERIAL ADAPTER FAILURE IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE.><.><IT WAS REPORTED THAT THE TABLET DEVICE DISPLAYED A FAILURE TO OPEN PORT ERROR" MESSAGE. TROUBLESHOOTING CONFIRMED PROPER FUNCTIONALITY OF THE WAND WITH A DEMO GENERATOR. THE TABLET USB ADAPTOR HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY." | 1
5250758 | IT WAS REPORTED THAT THE PHYSICIAN'S VNS PROGRAMMING TABLET HAD MECHANICAL DAMAGE. NO ONE AT THE OFFICE KNEW HOW THE TABLET WOUND UP TO BE THE STATE IT WAS IN. NO PATIENTS WERE NEGATIVELY IMPACTED BY THE TABLET'S CONDITION AS IT WAS STATED TO STILL BE FUNCTIONAL. THE TABLET WAS RECEIVED FOR ANALYSIS. ANALYSIS CONFIRMED THAT THE TABLET POWER SWITCH BUTTON WAS ORIENTED UPSIDE-DOWN APPARENTLY DUE TO USER ERROR, RESULTING IN ACCIDENTAL REMOVAL AND INCORRECT REINSERTION. THERE WAS ALSO SOME OBSERVED DEBRIS THAT PREVENTED THE TABLET LID FROM CLOSING. ONCE THE POWER BUTTON WAS INSTALLED CORRECTLY AND THE DEBRIS WAS REMOVED, NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET WERE IDENTIFIED DURING THE ANALYSIS. DEVICE MANUFACTURING RECORDS WERE REVIEWED AND FOUND ALL SPECIFICATIONS MET PRIOR TO DISTRIBUTION. SINCE THE MANUFACTURING PROCESSES INVOLVE POWERING THE TABLET ON, AND NO SUCH ISSUES WERE NOTED ON THE DEVICE RECORD HISTORY, IT IS BELIEVED THAT THIS ISSUE (POWER BUTTON REPLACED UPSIDE DOWN) OCCURRED OUT IN THE FIELD.><. | 1
5271585 | REVIEW OF DEVICE PROGRAMMING HISTORY IDENTIFIED THAT THE GENERATOR HAS BEEN DISABLED DUE TO A BURST WATCHDOG TIMEOUT ASSOCIATED WITH THE VERSION 7.1 SOFTWARE. INVESTIGATION DETERMINED THAT A FLAW WITHIN THE PROGRAMMING SOFTWARE RESULTS IN A LONGER THAN EXPECTED RAMP UP/RAMP DOWN PERIODS ASSOCIATED WITH MAGNET STIMULATION AT CERTAIN PROGRAMMED PARAMETERS. THE PROGRAMMING SYSTEM SOFTWARE WAS UPGRADED AS A RESULT.><. | 1
5050153 | .><A HANDHELD WAS RETURNED DUE TO WATER DAMAGE FROM FLOODING AT A HOSPITAL. ANALYSIS FOUND THAT IT DID NOT COMMUNICATE DUE TO A DAMAGED SERIAL CABLE, WHICH WAS NOT RELATED TO THE WATER DAMAGE. IT WAS IDENTIFIED THAT THERE WAS A BROKEN WIRE CONNECTION TO THE PCB. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. | 1
4932643 | THE ISSUE WAS INVESTIGATED BY SIEMENS EXPERTS AND THE ROOT CAUSE FOR THE REPORTED ISSUE WAS IDENTIFIED AS FOLLOW. WHEN A NEW ORDER ARRIVES AND THE HIS-RIS (HOSPITAL INFORMATION SYSTEM) INTERFACE JOB INSERTS A ROW INTO THE VISIT_ORDER" TABLE. BEFORE ADDITIONAL ROWS CAN BE INSERTED BY HIS-RIS INTO THE "VISIT_ACTIVITY" AND "VISIT_ACTIVITY_INFO" TABLES | 1
5004154 | KEYMED LTD HAVE RECEIVED THE SUBJECT DEVICE. THE SUBJECT DEVICE HAS BEEN EVALUATED. THE EVALUATION DETERMINED THAT ONE SOCKET OUTLET ON THE CENTRAL STRIP WAS SEVERELY ERODED AROUND THE LEFT HAND POWER CONTACT DUE TO EXTREME HEAT. INTERNALLY THERE WAS SOME DISTORTION OF THE MOULDING WHERE THE CONTACT IS SECURED, BUT NOT TO THE SAME EXTENT AS EXTERNALLY. THE EQUIPMENT LEAD PLUG HAS NOT BEEN FULLY ENGAGED WITH THE SOCKET, SO THAT ITS LEFT HAND PIN WAS ONLY TOUCHING THE SOCKET CONTACT AT ITS TIP RATHER THAN BEING GRIPPED ON BOTH SIDES AS IT WOULD NORMALLY BE WHEN THE PLUG IS FULLY ENGAGED. THE SINGLE POINT OF CONTACT WILL PRODUCE CONSIDERABLE HEAT WHEN CURRENT IS FLOWING. THE IFU ALREADY CONTAINS CAUTIONS TO ENSURE THAT CONNECTORS ARE FULLY ENGAGED WHEN THE SYSTEM IS CONFIGURED AND TO CHECK THEM DURING ROUTINE MAINTENANCE.><KEYMED LTD HAVE REQUESTED THE RETURN OF THE SUBJECT DEVICE FROM (B)(6), SO THAT A FULL INVESTIGATION CAN TAKE PLACE TO ESTABLISH THE ROOT CAUSE OF THIS MALFUNCTION. AWAITING RETURN OF SUBJECT DEVICE.><THE OLYMPUS WM-NP2 MOBILE WORKSTATION IS INTENDED FOR USE IN MEDICAL FACILITIES UNDER THE DIRECTION OF A TRAINED PHYSICIAN, AND HAS BEEN DESIGNED TO BE USED WITH A RANGE OF OLYMPUS EQUIPMENT TO FACILITATE GI ENDOSCOPY, ENDOSCOPIC ULTRASOUND, RESPIRATORY AND SURGICAL ENDOSCOPIC PROCEDURES. THE SUBJECT DEVICE, A SEPARATION TRANSFORMER (B)(4), IS LOCATED ON THE UNDERSIDE OF THE BASE SHELF. THE SEPARATION TRANSFORMER CONNECTS TO THE MAINS POWER SUPPLY AND HAS A BANK OF OUTPUT SOCKETS WHICH ALLOWS FOR THE ELECTRICAL CONNECTION OF OLYMPUS MEDICAL DEVICES WHICH ARE SITUATED ON THE WORKSTATION. KEYMED (MEDICAL &INDUSTRIAL EQUIPMENT) LTD, HAS BEEN MADE AWARE OF AN EVENT IN (B)(6) WHERE DURING THE PREPARATION FOR USE, A SNAPPING NOISE OCCURRED AS THE FACILITY STAFF POWERED ON A (B)(4) LIGHT SOURCE WHICH WAS INSTALLED ON THE WORKSATATION. FACILITY STAFF CHECKED THE (B)(4) SEPARATION TRANSFORMER AND FOUND THAT THE OUTPUT SOCKET OF THE TRANSFORMER WHICH CONNECTED TO THE (B)(4) LIGHT SOURCE HAD MELTED. THE HEALTH FACILITY USED ANOTHER WORKSTATION TO PERFORM THE PROCEDURE. THERE IS NO REPORT OF INJURY TO PATIENT OR USER AND THEREFORE THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. | 1
5029102 | WHILE PERFORMING AN X-RAY PROCEDURE ON A PT, THE TECH USED THE AUTO BUTTON TO CENTER THE EQUIPMENT. WHEN THE BUTTON WAS RELEASED, THE XRAY TUBE CONTINUED TO DESCEND AND THEN STRUCK THE PT IN THE CHEST AREA. THE OTHER TECH TRIED TO CATCH THE TUBE BUT WAS A SECOND TOO LATE. THERE WERE NO INJURIES NOTED ON THE PT'S CHEST AREA. THE VENDOR HAD SINCE ASSESSED AND COMPLETED THE REPAIR OF OVERHEAD XRAY TUBE SUSPENSION. | 0
5274066 | MINDRAY SERVICE TEAM HAS INSPECTED THE DEVICE , THE REPORTED ISSUE COULD NOT DUPLICATED, THE ECG FUNCTION OF THE DEVICE MEET THE SPECIFICATION. MINDRAY (B)(4) TEAM HAS ANALYZED THE RETURNED LOG FILES. WE FOUND THE SUSPECTED SYSTEOLE WAVEFORM WAS CONTINUED FOR 6 SECONDS, NOT AS LONG AS THE CUSTOMER REPORTED FOR 25 SECINFS. THERE ARE THREE QRS WAVEFORM EXISTING IN THIS 6 SECONDS SEGMENT OF WAVEFORM, AS A RESULT, IT IS NOT A ASYSTOLE WAVEFORM FROM THE ECG ALGORITHM ANALYSIS. AND MINDRAY T1 MONITOR HAS TRIGGGERED ALARMS OF HR TOO LOW" TO REFLECT THE PATIENT SITUATION DURING THE EVENT. (B)(4).><MINDRAY SERVICE TEAM HAS PERFORMED ON-SITE INVESTIGATION WITH THE REPORTED PROBLEM. THE COMPLAINED UNIT WAS SENT BACK TO THE WORKSHOP OF MINDRAY (B)(4) OFFICE FOR INSPECTION | 1
5242407 | CUSTOMER REPORTED SMOKE COMING FROM THE DPM 7 MONITOR. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND FOUND BURNED INVERTER BOARD IN THE UNIT'S FRONT PANEL. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S FRONT PANEL ASSEMBLY. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. (B)(4). | 0
4435909 | MACHINE READ IDENTIFYING DEVICE" | 0
5101924 | ON 08/25/2015, VITAL CUSTOMER SUPPORT RECEIVED A CALL FROM (B)(6) REGARDING DISPLAY OF INCORRECTLY INVERTED BRAIN PERFUSION COLOR MAPS (CBF AND MTT) ON OUR VITREA CT BRAIN PERFUSION 4D SOFTWARE APPLICATION. THE REPORTER INDICATED THAT A DELAY IN PATIENT TREATMENT HAD OCCURRED BECAUSE OF THE ISSUE. AGAIN ON 08/26/2015, THE SAME REPORTER INFORMED OUR CUSTOMER SUPPORT TEAM REGARDING ANOTHER PATIENT ISSUE WITH THE SOFTWARE IN WHICH THE CT BRAIN PERFUSION RESULT FROM OUR SOFTWARE WAS NOT IN AGREEMENT WITH THE PERFORMED MR RESULTS. THE MRI SHOWED AN OBVIOUS STROKE (INFARCT), WHILE THE CT PERFUSION RESULTS ON OUR SOFTWARE DID NOT. THE PATIENT WAS TREATED BASED ON MRI RESULTS.><ON (B)(6) 2015, VITAL CUSTOMER SUPPORT RECEIVED A CALL FROM (B)(6) REGARDING DISPLAY OF INCORRECTLY INVERTED BRAIN PERFUSION COLOR MAPS (CBF AND MTT) ON OUR VITREA CT BRAIN PERFUSION 4D SOFTWARE APPLICATION. THE REPORTER INDICATED THAT A DELAY IN PATIENT TREATMENT HAD OCCURRED BECAUSE OF THE ISSUE. AGAIN ON (B)(6) 2015, THE SAME REPORTER INFORMED OUR CUSTOMER SUPPORT TEAM REGARDING ANOTHER PATIENT ISSUE WITH THE SOFTWARE IN WHICH THE CT BRAIN PERFUSION RESULT FROM OUR SOFTWARE WAS NOT IN AGREEMENT WITH THE PERFORMED MR RESULTS. THE MRI SHOWED AN OBVIOUS STROKE (INFARCT), WHILE THE CT PERFUSION RESULTS ON OUR SOFTWARE DID NOT. THE PATIENT WAS TREATED BASED ON MRI RESULTS.><THE CUSTOMER HAS PROVIDED US TWO DATASETS TO INVESTIGATE. THE RESULTS OF THE INVESTIGATION IDENTIFIED TWO SEPARATE ISSUES. ISSUE #1: UNEXPECTED MTT/CBF MAPS. THIS CUSTOMER STATED THE MTT/CBF MAPS APPEAR TO BE INVERTED. THE INVESTIGATION SHOWS THE DATASET HAS A TRUNCATED SCAN, THE TIME-DENSITY GRAPH SHOWS THAT THERE IS NOT ENOUGH WASHOUT FOR OUR SOFTWARE TO PRODUCE ACCURATE RESULTS. FROM OUR NEURO EDUCATION AND REFERENCE GUIDE (VPMC-13505 A, PAGE 122), WE STATE CAUTION: FACTORS THAT CAN AFFECT THE ACCURACY OF THE RESULTS ARE: ·NOT FULLY CAPTURING THE CONTRAST BOLUS (INCLUDING AT LEAST 2 SECONDS OF PRE-ENHANCEMENT IMAGES AND POST-ENHANCEMENT)." ALSO | 1
4636131 | THE INCIDENT WAS CAUSED BY THE EMPLOYEE RAISING THEMSELVES UP FROM UNDER THE KEYBOARD OF THEIR AUTOMATE 2500 SAMPLE PROCESSING SYSTEM AND HITTING THEIR HEAD AGAINST THE BASE OF THE KEYBOARD TRAY. THERE WERE NO INSTRUMENT ISSUES IN CONNECTION WITH THIS EVENT.><A CUSTOMER IN (B)(6) REPORTED TO BECKMAN COULTER (BEC) THAT ONE OF THEIR EMPLOYEES HAD INJURED THEIR HEAD AS A RESULT OF RAISING THEMSELVES UP FROM UNDER THE KEYBOARD OF THEIR AUTOMATE 2500 SAMPLE PROCESSING SYSTEM. THE EMPLOYEE HAD HIT THEIR HEAD AGAINST THE BASE OF THE KEYBOARD TRAY WHICH RESULTED IN A CONCUSSION AND TIME OFF WORK. THE EMPLOYEE WAS ADMITTED TO THE EMERGENCY ROOM (ER) WHERE CONCUSSION WAS DIAGNOSED. THE EMPLOYEE WAS SENT DIRECTLY HOME AFTER THE ER VISIT. IT WAS STATED BY THE CUSTOMER THAT THE EMPLOYEE SELF-CERTIFIED AND LATER WAS SIGNED OFF WORK BY A FAMILY DOCTOR (GP). AS OF THE (B)(6) 2015 THE EMPLOYEE WAS CONFIRMED TO STILL BE OFF WORK AS A RESULT OF THE INCIDENT. | 0
4469573 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE WAS ASSOCIATED WITH A DEVICE THAT WAS TURNING ITSELF ON AND OFF AND ALARMING IMPROPER SHUTDOWN." IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT.><(B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVAL. THE INVESTIGATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED." | 0
4407826 | (B)(4). THE BATTERY WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED AND REPRODUCED THE REPORTED WATER ON THE CIRCUIT BOARD. FURTHER EVALUATION FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PCB AS A RESULT OF FLUID INTRUSION WHICH CAUSED THE COMPONENTS TO FAIL. THE FLUID INTRUSION RENDERED THE UNIT UN-REPAIRABLE.><IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY HAD WATER ON THE CIRCUIT BOARD. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INVOLVEMENT. | 0
4536032 | IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY MODULE WOULD NOT CHARGE, RESULTING IN THE PUMP TURNING OFF. IT WAS ALSO REPORTED THAT THERE WAS NO PT INVOLVEMENT.><(B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVAL AND THEREFORE AN EVAL COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVAL WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
4464820 | IT WAS REPORTED THAT A WIRELESS BATTERY HAD HEAVY FLUID INTRUSION. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY.><MANUFACTURER REF NO.: (B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVALUATION AND THEREFORE AN EVALUATION COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVALUATION WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
4558648 | MFR REF NO: (B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVAL. THE EVAL HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT. | 0
4707279 | MANUFACTURER REF NO.: (B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE DEVICE WAS FOUND OUT OF SPECIFICATION IN RELATION TO THE REPORTED FLUID INTRUSION. THE EVALUATION FOUND THE DEVICE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE. IT WAS PAIRED WITH A KNOWN GOOD SPECTRUM PUMP, WHICH FOUND THAT THE DEVICE'S WIRELESS CONNECTION CAPABILITIES ARE INOPERABLE DUE TO FLUID INTRUSION. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION. THE DEVICE WAS RETIRED FROM SERVICE.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INTRUSION AND PRINTED CIRCUIT BOARD DAMAGE. IT WAS ALSO REPORTED THERE WAS NO PATIENT INVOLVEMENT. | 0
4489771 | IT WAS REPORTED THAT A SPECTRUM PUMP HAD AN UNK BATTERY ERROR. IT WAS ALSO REPORTED THERE WAS NO PT INJURY.><(B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVALUATION. THEREFORE, AN EVALUATION COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVAL WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
4693695 | IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INGRESS DURING BIOMED TESTING. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY.><(B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVALUATION. THE EVALUATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED. | 0
4558650 | IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT.><(B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE DEVICE WAS FOUND OUT OF SPEC IN RELATION TO THE REPORTED FLUID INTRUSION WHICH WAS CONFIRMED BUT NOT REPRODUCED. THE DEVICE WAS FOUND TO HAVE CORROSION ON THE RADIO PRINTED CIRCUIT BOARD (PCB) AS A RESULT OF FLUID INTRUSION. THE DAMAGE WAS REMOVED FROM SERVICE AS IT WAS DETERMINED TO BE UNREPAIRABLE DUE TO THE MAC ADDRESS BEING ASSIGNED AS THE PRODUCT SERIAL NUMBER. | 0
4608016 | (B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVALUATION AND THEREFORE AN EVALUATION COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVALUATION WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT. | 0
4441079 | (B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE DEVICE WAS FOUND OUT OF SPEC IN RELATION TO THE REPORTED FLUID INTRUSION. THE EVAL FOUND MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE. THE MODULE WAS PAIRED WITH A KNOWN GOOD SPECTRUM PUMP, WHICH FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE DUE TO A FLUID INTRUSION. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION, WHICH CAUSED THE COMPONENTS TO FAIL. THE MODULE WAS RETIRED FROM SERVICE.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PT INJURY. | 0
4716311 | IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PATIENT INJURY.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED AND REPRODUCED THE REPORTED SYMPTOM FLUID INTRUSION" | 0
4582256 | (B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVAL AND THEREFORE AN EVAL COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVAL WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS MODULE DISPLAYED DO NOT LOW ONTO POWER UP IT COMPLIT." IT WAS ALSO REPORTED THERE WAS NO PT INJURY." | 0
4449230 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE G SUFFERED FROM FLUID INTRUSION. ANY PT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNK.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVAL WAS PERFORMED. THE EVAL CONFIRMED AND REPRODUCED THE REPORTED SYMPTOM FLUID INTRUSION. THE EVAL FOUND THAT THE WIRELESS MODULE'S WIRELESS CONNECTION CAPABILITIES WERE INOPERABLE DUE TO THE REPORTED SYMPTOM. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION, RENDERING THE UNIT IRREPARABLE. THE DEVICE WAS RETIRED FROM SVC. | 0
4605164 | IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY SMOKED DUE TO AN ELECTRICAL SHORT. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INVOLVEMENT.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED THE ELECTRICAL SHORT DUE TO FLUID INTRUSION. CORROSION WAS FOUND ON THE RADIO PRINTED CIRCUIT BOARD (PCB) AND WIRELESS MODULE FLEX, CAUSING THE COMPONENTS TO FAIL AND RENDERING THE DEVICE IRREPARABLE. THE DEVICE WAS REMOVED FROM SERVICE. | 0
4763041 | IT WAS REPORTED THAT A SPECTRUM PUMP HAD BURN MARKS INSIDE BATTERY BOARDS. ANY PATIENT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNKNOWN.><(B)(4). THE DEVICE WAS NOT RETURNED TO BAXTER FOR EVAL. THEREFORE, AN EVAL COULD NOT BE COMPLETED. IF THE DEVICE IS RETURNED, AN EVAL WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
4469470 | (B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVAL. THE INVESTIGATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED.><IT WAS REPORTED THAT A WIRELESS BATTERY MODULE FAILED A RUN TIME TEST. ANY PT INVOLVEMENT IS UNK. | 0
4693783 | (B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVAL. THE INVESTIGATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED.><DURING BAXTER'S EVAL OF A SPECTRUM PUMP, THE UNIT WOULD NOT TURN ON WITH THE WIRELESS MODULE THAT IT WAS RECEIVED WITH. THERE WAS NO PT INVOLVEMENT. | 0
4753272 | (B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVALUATION. THE EVALUATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED.><IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PATIENT INVOLVEMENT. | 0
4702989 | IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY HAD WATER DAMAGE. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE REPORTED WATER DAMAGE (BATTERY NOT WORKING) WAS CONFIRMED DURING EVALUATION. THE EVALUATION FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE DUE TO A CHECK BATTERY" CONDITION. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PCB AS A RESULT OF FLUID INTRUSION | 0
4626683 | IT WAS REPORTED A WIRELESS MODULE FOR A SPECTRUM PUMP HAD FLUID INTRUSION. ANY PATIENT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNK.><(B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE REPORTED FLUID INTRUSION WAS CONFIRMED DURING THE DEVICE INVESTIGATION. THE EVALUATION FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD (PCB) DUE TO FLUID INTRUSION. THE DEVICE WAS REMOVED FROM SERVICE. | 0
4648130 | IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY HAD POSSIBLE FLUID INTRUSION. IT WAS ALSO REPORTED THAT THERE WAS NO PT INVOLVEMENT.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVAL WAS PERFORMED. THE EVAL CONFIRMED THE REPORTED FLUID INTRUSION. THE EVAL FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD, AS A RESULT OF FLUID INTRUSION, CAUSING THE COMPONENTS TO FAIL. THE FAILED COMPONENTS RENDERED THE DEVICE IRREPARABLE AND THE DEVICE WAS REMOVED FROM SVC. | 0
4589216 | IT WAS REPORTED THAT A WIRELESS MODULE FOR A SPECTRUM PUMP HAD FLUID INTRUSION. IT WAS FURTHER REPORTED THIS CONDITION OCCURRED DURING TESTING. IT WAS ALSO REPORTED THERE WAS NO PATIENT INVOLVEMENT.><(B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE REPORTED FLUID INTRUSION WAS CONFIRMED THROUGH THE DEVICE INVESTIGATION. THE EVALUATION FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD (PCB) AS A RESULT OF FLUID INTRUSION. THE FLUID INTRUSION CAUSED THE COMPONENTS TO FAIL AND DEEMED THE DEVICE IRREPARABLE. THE DEVICE WAS REMOVED FROM SERVICE. | 0
4530133 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PATIENT INJURY.><(B)(4). THE DEVICE HAS BEEN RECEIVED BY BAXTER FOR EVALUATION. THE EVALUATION HAS NOT BEEN COMPLETED AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE PROVIDED WHEN THE INVESTIGATION IS COMPLETED. | 0
4468946 | (B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED THE REPORTED SYMPTOM FLUID INTRUSION. EVALUATION FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION. THE DEVICE COULD NOT BE REPAIRED AND WAS RETIRED FROM SERVICE.><IT WAS REPORTED THAT A WIRELESS BATTERY MODULE G HAS FLUID INTRUSION. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INVOLVEMENT. | 0
4707117 | IT WAS REPORTED THAT A SPECTRUM PUMP WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT.><MANUFACTURER REF NO.: (B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE DEVICE WAS FOUND OUT OF SPECIFICATION IN RELATION TO THE REPORTED FLUID INTRUSION. THE DEVICE WAS PAIRED WITH A KNOWN GOOD SPECTRUM PUMP, WHICH FOUND THAT THE DEVICE'S WIRELESS CONNECTION CAPABILITIES ARE INOPERABLE. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION. THE DEVICE WAS RETRIED FROM SERVICE. | 0
4462656 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE G EXPERIENCED ARCING. IT WAS REPORTED THERE WAS NO PT INVOLVEMENT.><MANUFACTURER REF NO.: (B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE EVAL WAS UNABLE TO REPRODUCE THE REPORTED SYMPTOM ARCING, BUT FOUND EVIDENCE THAT ARCING OCCURRED. THE DEVICE WAS RETIRED FROM SERVICE. | 0
4863196 | IT WAS REPORTED VIA A PRESENTATION AT A REGIONAL GASTROENTEROLOGIST MEETING THAT THE PATIENT WAS UNDERGOING AN EGD PROCEDURE, WITH SEDATION ADMINISTERED WITH THE SEDASYS SYSTEM. THE EVENT REQUIRED BAG-MASK VENTILATION DUE TO LARYNGOSPASM ASSOCIATED WITH UNDERSEDATION. ADDITIONAL INFORMATION FROM THE SITE INDICATED THAT THEY USED THE TERM LARYNGOSPASM TO DESCRIBE A PATIENT THAT WAS AGITATED/COUGHING WHILE THE ENDOSCOPIST WAS INSERTING THE ENDOSCOPE WHICH ELICITED THE ¿GAG REFLEX.¿ THIS LED TO TRANSIENT OXYGEN DESATURATION THAT WAS RESOLVED WITH BAG-MASK VENTILATION. THE EVENT WAS CAUSED BY THE STIMULATION OF THE VOCAL CORDS BY THE ENDOSCOPE DURING THE EGD. THIS CAN OCCUR WITH OTHER MODES OF SEDATION. THE EVENT DID NOT REQUIRE CALLING FOR A CODE TEAM OR INVOLVING AN ANESTHESIOLOGIST TO MANAGE THE AIRWAY. POSITIVE PRESSURE WAS NOT REQUIRED TO ¿BREAK¿ THE LARYNGOSPASM AND NO FURTHER AIRWAY MANEUVERS WERE REQUIRED. THERE WAS NO PATIENT INJURY RESULTING FROM THE EVENT AND THE PATIENT WAS ROUTINELY DISCHARGED FROM THE FACILITY THE DAY OF THE PROCEDURE.><(B)(4). CONCLUSION: UNIT NOT RETURNED FOR EVALUATION. ADDITIONAL INFORMATION: THE CLINICIANS AT THE SITE INDICATED THE EVENT WAS NOT CAUSED BY THE SEDASYS SYSTEM RATHER THE PROCEDURE ITSELF. THE SITE REVIEWED THE EVENT WITH THEIR ANESTHESIA DEPARTMENT, WHICH RECOMMENDED THEY START USING A TOPICAL ANESTHETIC TO NUMB THE VOCAL CORDS PRIOR TO INSERTING THE ENDOSCOPE. SINCE IMPLEMENTING THAT PRACTICE IN (B)(6), THE SITE HAS OBSERVED NO LARYNGOSPASMS. AFTER REVIEWING THE AVAILABLE INFORMATION WE HAVE CONCURRED WITH THE SITE, THAT THE LARYNGOSPASM WAS CAUSED BY THE PROCEDURE; SPECIFICALLY STIMULATION OF THE VOCAL CORDS BY THE ENDOSCOPE. WE HAVE ELECTED TO REPORT THE EVENT DUE TO THE BAG-MASK VENTILATION TO BE CONSISTENT WITH OUR POST-APPROVAL COMMITMENT TO FDA. SPECIFICALLY, THE INCIDENCE OF BAG-MASK VENTILATION IS A SECONDARY ENDPOINT OF ONE OF TWO REQUIRED POST-APPROVAL STUDIES. THE EVENT REPORTED DID NOT OCCUR IN THAT CLINICAL STUDY, AS THAT STUDY IS NOT SLATED TO START UNTIL LATE 2016. | 0
5035340 | PROCEDURE: INSERTION OF INTRACRANIAL PRESSURE MONITOR DEVICE. WHEN THE DEVICE WAS PREPARED FOR THE CASE THE DEVICE WAS SWITCH ON, AND NOTICED THAT THE DEVICE HAD A MESSAGE SYSTEM FAILURE ON THE SCREEN. SAME THING ERROR MESSAGE HAPPENED TO THE SECOND DEVICE. THEY CHANGED THE PROCEDURE TO AN INSERTION OF EXTERNAL VENTRICULAR DRAIN. TWENTY MINUTES DELAY IN SURGERY. NO ADVERSE EVENT TO PATIENT.><IT WAS INITALLY REPORTED THAT THE DEVICE WOULD NOT BE RETURNED FOR EVALUATION. HOWEVER, THE DEVICE WAS LATER MADE AVAILABLE. IT WAS FOUND THAT THERE WAS ELECTRICAL DAMAGE TO THE ANALOG BOARD, THE LCD WAS NOT FUNCTIONAL, THE CLAMP WAS CORRODED AND THE BATTERY WAS WEAK AND WOULD NOT HOLD A CHARGE. THE SUPPLIER REPLACED THE U24, U8, U30 ON THE ANALOG BOARD, LCD, CLAMP AND BATTERY. THE INSTURMENT WAS TESTED AND PASED ALL SPECIFICATIONS. THE INSTURMENT WAS TESTED AND PASED ALL SPECIFICATIONS THE SUPPLIER WAS NOT ABLE TO DIRECLY DETERMINE THE CAUSE OF FAILURE. HOWEVER, THE DEVICE WAS MANUFACTURED IN 2006 AND IS THEREFORE OUTSIDE OF WARRANTY. IT IS LIKELY THAT THE AGE OF THE DEVICE CONTRIBUTED TO THE REPORTED ISSUE. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED.><(B)(4). IT HAS BEEN COMMUNICATED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. WITHOUT THE DEVICE IT IS NOT POSSIBLE FOR CODMAN TO CONDUCT A PROPER INVESTIGATION. SINCE A LOT NUMBER HAS BEEN PROVIDED A REVIEW OF THE MANUFACTURING RECORDS WILL BE REVIEWED. WE ANTICIPATE THAT THE EVALUATION WILL REVEAL THAT THE DEVICE CONFORMED TO SPECIFICATIONS PRIOR TO RELEASE. IF ANYTHING OTHERWISE IS FOUND THEN A FOLLOW UP REPORT WILL BE FILED. IF AT SOME POINT THE DEVICE DOES BECOME AVAILABLE, THIS COMPLAINT WILL BE RE-OPENED, EVALUATED AND A FOLLOW UP REPORT WILL BE FILED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR COMPLAINTS. AT THE PRESENT TIME THIS COMPLAINT IS CONSIDERED CLOSED. REFER TO MDR 1226348-2015-10483 FOR INFORMATION REGARDING THE SECOND DEVICE REPORTED IN THIS EVENT. | 1
5068978 | A CUSTOMER COMPLAINT WAS REPORTED THROUGH THE NATUS NEUROLOGY (B)(6) DISTRIBUTOR. AN (B)(6) CUSTOMER WAS USING A NATUS ENDEAVOR CR DESKTOP SYSTEM ON A PATIENT TO MONITOR EEG/EMG FUNCTIONS. THE CUSTOMER WAS PERFORMING PATIENT STIMULATION IN A NON-RECURRENT MODE AND REPORTED THAT A DOUBLE STIMULATION OCCURRED ON THE PATIENT USING A COMPUTER MOUSE AND BY CLICKING ON THE ASSOCIATED SYSTEM COMPUTER ICON. AFTER PURPOSELY PERFORMING THE FIRST STIMULATION, A SECOND STIMULATION PULSE OCCURRED SUPPOSEDLY WITHOUT FURTHER INPUT FROM THE DOCTOR. FOLLOWING THE SECOND STIMULATION, THE PATIENT EXPERIENCED A SEIZURE WHICH REQUIRED DRUG INTERVENTION TO CONTROL.><(B)(4). DEVICE WAS NOT RETURNED TO MANUFACTURER FOR EVALUATION (B)(4). THE SYSTEM MANUFACTURING RECORDS WERE REVIEWED. ALL SYSTEM COMPONENT INSPECTION TESTS WERE IN SPECIFICATION AND NO ISSUES ASSOCIATED WITH THE PRODUCT MANUFACTURER WERE IDENTIFIED THAT MAY HAVE CONTRIBUTED TO THE COMPLAINANTS CLAIM. EVALUATION PERFORMED BY THE MANUFACTURER USING DUPLICATE EQUIPMENT COULD NOT REPRODUCE THE REPORTED MULTIPLE STIMULATION FAULT EXCEPT THROUGH REPEATED ACTIVATION OF THE MOUSE/ICON INTERFACE, AN EXPECTED FUNCTION. THE SYSTEM SOFTWARE CODE WAS ALSO REVIEWED CONFIRMING THAT A DOUBLE STIMULATION CANNOT OCCUR ON ITS OWN. EVEN IF THIS PURPOSELY OCCURS THROUGH REPEATED ACTIVATION, STIMULATION OF THIS KIND USING SUCH EQUIPMENT IS NOT CONSIDERED HAZARDOUS, IS NOT CONSIDERED ADEQUATE TO INITIATE A PATIENT SEIZURE, NOR IS SUCH A RESPONSE REPORTED IN LITERATURE AS A POTENTIAL RISK IN ELECTRO-DIAGNOSTIC MEDICINE. | 1
4974698 | ARTERIAL BLOOD GAS WAS ORDERED VIA CPOE SYSTEM BECAUSE OF OBTUNDATION IN PT WITH A HISTORY OF RESPIRATORY FAILURE. MORE THAN 24 HOURS LATER, THE TEST WAS LISTED AS IN PROGRESS" | 1
4423697 | MY WIFE, (B)(6), HAD BEEN TAKING WARFARIN FOR AFIB SINCE 2010. SHE BEGAN USING THE ALERE INRATIO PT/INR MONITOR SYSTEM AND TEST STRIPS IN 2012 DOING SELF TEST AT HOME. ON (B)(6) 2014, (B)(6) HAD COLON SURGERY. SHE HAD STAGE 3 COLON CANCER. WHILE IN THE HOSPITAL, SHE WAS GIVEN HEPARIN. WHEN SHE LEFT THE HOSPITAL, THE DOCTOR TOLD HER TO START TAKING WARFARIN IMMEDIATELY. SHE DID. HER TEST RESULTS WERE 2.5-3.5. (B)(6) WAS RECOVERED FROM HER OPERATION. SHE DID HAVE INFECTIONS. ON (B)(6) 2014, HER TEST RESULT WAS 2.8. ON (B)(6) 2014, I HAD AN AMBULANCE TAKE (B)(6) TO EMERGENCY FOR SEVERE HEADACHES. HER INR READING WAS OVER 5. SHE HAD BLEEDING ON THE BRAIN AND A STROKE. (B)(6) DID NOT COME BACK HOME. SHE DIED ON (B)(6) 2014 IN (B)(6). | 0
4834775 | WHEN ENTERING A WEIGHT FOR DOSING MEDICATION, THE DEVICE DOES NOT SPECIFY THE UNITS IE KG FOR WEIGHT ENTRY. INNUMERABLE HEALTH CARE PROFESSIONALS ENTERED THE WEIGHT IN POUNDS RESULTING IN POTENTIAL OVERDOSES. | 1
4604419 | PDM MALFUNCTION - INTERMITTENT LEAD FAILURE LED TO INTERMITTENT DISPLAY OF WAVEFORM THROUGHOUT CASE. ALL WAVEFORMS (EKG, ARTERIAL LINE BP, PULSE) DISAPPEARED EVEN THOUGH EKG LEADS WERE THE ONLY ONES TO GO OUT. THIS OCCURRED INTRAOPERATIVE DURING OPEN HEART SURGERY. | 1
4992907 | AT APPROXIMATELY 11:57 AM ALARMS STARTING SOUNDING. THE PATIENT MONITOR ON THE ANESTHESIA MACHINE IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PATIENT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATION ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL PATIENT MONITOR. AFTER ABOUT 15 MINUTES IT WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PATIENT MONITOR WAS REMOVED FROM THE NETWORK AND THE PATIENT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
5045814 | CUSTOMER NOTIFIED HOLOGIC OF A SLIDE WITH ABNORMAL CELLS THAT WERE NOT LOCATED IN THE 22 FIELDS OF VIEW (FOV). ALTHOUGH THE CUSTOMER REPORTED THE EVENT TO HOLOGIC ON (B)(6) 2015, THE SLIDE WAS FROM (B)(6) 2014. THE CUSTOMER DISCOVERED THE SLIDE DURING QC. THIS IS A POTENTIAL DELAY IN DIAGNOSIS. HOLOGIC'S CYTOLOGY APPLICATION SPECIALIST (CAS) REVIEWED THE SLIDE WITH A MIX OF OTHER SLIDES. THE CAS AGREED THAT THERE WERE NO ABNORMAL CELLS IN THE 22 FOV TO PROMPT AN AUTOSCAN (FULL REVIEW) OF THE SLIDE. CUSTOMER DOES NOT WANT TO SUBMIT THE SLIDE TO HOLOGIC FOR FURTHER REVIEW. | 0
5313240 | ><DURING A DOBUTAMINE STRESS TEST THE CARDIAC ULTRASOUND EQUIPMENT FAILED. THE SCREEN SHUT DOWN AND WOULD NOT RESPOND RESULTING IN CANCELLATION OF THE PROCEDURE. | 1
4604302 | PDM MALFUNCTION OF EKG CONNECTION FOR ABOUT 3 SECONDS DURING INTRAOPERATIVE OPEN HEART SURGERY AND AFTER SOFTWARE UPGRADE. | 1
4712895 | PATIENT PRESENTED TO EMERGENCY DEPARTMENT ON (B)(6) 2014 WITH SEVERE SHORTNESS OF BREATH. PATIENT WAS ADMITTED WITH CHF EXACERBATION AND HYPERTENSION ON (B)(6) 2014. PLACED ON CPAP AND GIVEN DIURETICS. PATIENT WAS IMPROVING AND EXPECTED TO BE DISCHARGED ON (B)(6)2014. DURING THE NIGHT THE CARDIAC RHYTHM CHANGED (BEGAN TO SLOW) LEADING TO A CARDIOPULMONARY ARREST. RESUSCITATION WAS NOT SUCCESSFUL AND THE PATIENT TIME OF DEATH WAS (B)(6) 2014 AT 00:33. LATE IN THE EVENT, WHEN THE PATIENT WAS IN COMPLETE HEART BLOCK WITH A VERY SLOW ESCAPE RHYTHM AND VISIBLE NON-CONDUCTED ATRIAL ELECTRICAL SPIKES ON THE ELECTROGRAMS, THE MONITOR'S PRINTOUT REPORTS INDICATE THAT THE MONITOR INCORRECTLY REPORTED OUT HEART RATES THAT THE MONITOR INCORRECTLY REPORTED OUT HEART RATES THAT WERE NUMERICALLY RAPID WHEN THE PATIENT'S ACTUAL HEART RATE WAS VERY SLOW. | 1
5004379 | MONITOR IN ICU ROOM (B)(6) LOST ALL INVASIVE PRESSURES WHILE ON HEART PATIENT CONNECTED TO MONITOR SP CABG. NURSE HAD TO RE-SELECT PRESSURE ON THE MONITOR. | 1
5179913 | INITIAL SAMPLE IN (B)(6) 2015 DEMONSTRATED LOW PLATELET (PLT) VALUES. SAMPLE OBTAINED IN (B)(6) 2015 GENERATED DECREASING PLT VALUES. PHYSICIAN ORDERED A BONE MARROW PROCEDURE DUE TO DECREASING PLATELET VALUES. BONE MARROW EVAL DETERMINED: THROMBOCYTOPENIA AND PLT CLUMPING. THE OPERATOR REVIEWED THE ORIGINAL SLIDE FROM JULY AND NO PLT CLUMPS WERE OBSERVED, HOWEVER, PLT ESTIMATE APPEARED ADEQUATE DESPITE LOW VALUE. BOTH ANALYSES GENERATED THROMBOCYTOPENIA" FLAGS ALERTING THE OPERATOR OF A SAMPLE ABNORMALITY. VERIFICATION OF ACCURATE RESULTS IS REQUIRED PRIOR TO REPORTING TO THE DOCTOR. PLT CLUMPING IS A KNOWN SAMPLE INTERFERENCE IN EDTA ANTICOAGULATION THAT FALSELY DECREASES PLT COUNTS (PSEUDOTHROMBOCYTOPENIA). AFTER THE BONE MARROW PROCEDURE | 0
5004382 | HAD RECEIVED OHS PATIENT FROM OR MINDRAY MONITOR SHUT DOWN AND WAS NOT ABLE TO BE RESTARTED. SCREEN WENT BLACK AND WHEN I TRIED TO RESTART, IT KEPT SAYING STARTING" AND THEN GO BLACK AGAIN. CHANGED BOXES WITHOUT AVAIL. SHUT DOWN WHOLE COMPUTER AND RESTARTED TO REBOOT SYSTEM AND THEN IT WORKED." | 1
4628452 | ><PATIENT INFORMATION UNAVAILABLE. THIS RECORD REPRESENTS PATIENT 2 OF 2. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | 1
4706450 | IT WAS REPORTED THAT THE COROMETRICS 172 FETAL MONITOR WAS NOT CONFIGURED PROPERLY DURING THE MANUFACTURING PROCESS FOR THE FETAL HEART RATE (FHR) LOWER ALARM LIMIT. THE FHR WAS CONFIGURED TO 50 BEATS PER MINUTE (BPM) INSTEAD OF 120 BPM ON THE DEVICE THAT WAS CONFIGURED AT THE GE BORDER OPERATIONS FACTORY. THERE WAS NO PATIENT INVOLVEMENT SINCE THE ISSUE WAS DISCOVERED BY GE HEALTHCARE PERSONNEL DURING INSTALLATION OF THE MONITOR PRIOR TO PATIENT USE.><GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE PROVIDED AFTER THE INVESTIGATION HAS BEEN COMPLETED. NO REPORT OF PATIENT INVOLVEMENT. | 1
5069836 | .><A SMALL CAPSULE STUDY DATA RECORDER BATTERY STOPPED AFTER 5 HOURS. THE BATTERY SHOULD LAST AT LEAST 7-8 HOUR. IT CAUSED AN INCOMPLETE STUDY FOR THE PATIENT. | 1
5238428 | THE MANUFACTURER DID NOT RECEIVE DEVICES, X-RAYS, OR OTHER SOURCE DOCUMENTS FOR REVIEW AS IT UNKNOWN IF A REVISION ALREADY TOOK PLACE AS NO LOT NUMBERS WERE PROVIDED FOR THE DEVICES, THE DEVICE HISTORY RECORDS COULD NOT BE REVIEWED. A CAUSE FOR THIS SPECIFIC EVENT CANNOT BE ASCERTAINED FROM THE INFORMATION PROVIDED. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE AND AN INVESTIGATION RESULT BE AVAILABLE, THAT CHANGED THIS ASSESSMENT, AN AMENDED MEDICAL DEVICE REPORT WILL BE SUBMITTED. ZIMMER'S REFERENCE NUMBER OF THIS FILE IS (B)(4).><DHR REVIEW: THE DHR CHECK COULD NOT BE PERFORMED AS THE LOT NUMBER WAS NOT AVAILABLE. TREND ANALYSIS: TREND ANALYSIS COULD NOT BE PERFORMED AS NO PRODUCT INFORMATION WAS AVAILABLE. COMPATIBILITY CHECK: THE COMPATIBILITY CHECK COULD NOT BE PERFORMED AS NO PRODUCT INFORMATION WAS PROVIDED. REVIEW OF INCOMING INFORMATION: IT WAS REPORTED THAT A PATIENT HAS METAL REACTIONS AND SWELLING AFTER IMPLANTATION OF AN ANKLE FIX 4.0 PLATE. IT WAS STATED THAT THE PATIENT HAD METAL REACTIONS AND SWELLING. FIXATION PROBLEMS EVEN WITH LOCKING SCREW WAS REPORTED. ALSO THAT THE SCREWS COME OUT OF THE PLATE AND THE HOLE CONSTRUCT IS FAILING". NO FURTHER DOCUMENTS WERE PROVIDED. REVIEW OF INTERNAL DOCUMENTS: THE SURGICAL TECHNIQUE ANKLE FIX SYSTEM 4.0 CONTAINS INFORMATION AND ILLUSTRATIONS ABOUT ALL THE STEPS NEEDED FOR THE IMPLANTATION OF A ANKLE FIX PLATE. NEITHER X-RAYS | 0
4506478 | ><THE EQUIPMENT WAS JUST INSTALLED; IT HAD NOT BEEN TURNED OVER TO THE HOSPITAL, AND WAS IN THE PROCESS OF BEING TESTED, WHEN A FAULTY ELECTRONIC PANEL CAUGHT FIRE. A REPAIR TECH WAS CALLED AND CAME TO REPLACE THE FAULTY ELECTRONIC PANEL. GE IS INVESTIGATING AND WILL SUBMIT A REPORT. | 0
4626281 | THE CUSTOMER REPORTED SPONTANEOUSLY CHANGING ALARM DEFAULTS. PER THE CUSTOMER, THERE WAS NO SPECIFIC PATIENT AFFECTED AND PATIENT INFORMATION IS NOT AVAILABLE.><ENGINEERING ATTEMPTED TO REPRODUCE THE REPORTED ISSUE BY SIMULATING VARIOUS WORKFLOWS INCLUDING: PERFORMING NEW CASES, POWER CYCLES AND COMBINATIONS OF INTERMIXING POWER CYCLES, SPO2 PROBE ALARM CONDITIONS AND PLACING THE MONITOR INTO STANDBY. THE REPORTED EVENT WAS UNABLE TO BE REPRODUCED. THE ROOT CAUSE OF THE SPO2 NO CRISIS AUDIBLE ALARM ISSUE COULD NOT BE DETERMINED. ENGINEERING WAS UNABLE TO REPRODUCE THE ISSUE WITH THE SPO2 ALARM LEVEL CHANGING FROM CRISIS TO MESSAGE LEVEL DESPITE NUMEROUS ATTEMPTS.><PER THE CUSTOMER, THERE WAS NO SPECIFIC PATIENT AFFECTED AND PATIENT INFORMATION IS NOT AVAILABLE. GE HEALTHCARE'S INVESTIGATION IN ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
4826533 | ><THE DEPARTMENT IS DOWN TO THREE SETS OF LEADS AND GE HAS NO PLANS OF STOCKING THE EKG LEADS, NOR DOES THE SECONDARY SUPPLIER COMPANY PLAN ON HAVING ANY OF THESE TYPES OF LEADS FOR AT LEAST A MONTH. GE GAVE NO WARNING TO THE RELENQUISHMENT OF SALES TOWARDS THE LEADS AND JUST STATED, NO ETA ON REPLENISHMENT." THIS PROBLEM CONTINUES. THE MANUFACTURER DID NOT PROVIDE NOTICE TO THE HOSPITAL AND WE ARE UNABLE TO GET ADDITIONAL LEADS.======================MANUFACTURER RESPONSE FOR CARDIAC MONITOR LEADS | 0
4992889 | AT APPROXIMATELY 11:57 AM ALARMS STARTING SOUNDING. THE PATIENT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PATIENT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATION ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL PATIENT MONITOR. AFTER ABOUT 15 MINUTES IT WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PATIENT MONITOR WAS REMOVED FROM THE NETWORK AND THE PATIENT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
4790554 | THE DA VINCI ROBOTIC MONOPOLAR INSTRUMENT WAS INSERTED INTO THE ROBOTIC ARM AND IT DID NOT HAVE THE GREY WINGS"/LATCH TO GIVE THE ASSISTANT THE ABILITY TO REMOVE THE INSTRUMENT WHEN FINISHED USING IT. IT WAS NOTICED WHEN RETRIEVING THE INSTRUMENT THAT THIS INSTRUMENT WAS UNABLE TO BE REMOVED. THE EMERGENCY RELEASE KIT WHICH OPENS THE JAW OF THE INSTRUMENT MANUALLY WAS USED TO UNDOCKTHE ARM AND RE-DRAPE THE ARM SWITCHING THE INSTRUMENT OUT AND PROCEEDING WITH THE CASE | 0
4656695 | ><WHEN OPENING A SEALED CARTON OF THE SAGE # 6000 MOUTH CARE KITS, WE NOTICED IT WAS INFESTED WITH DEAD INSECTS. IT APPEARED NONE WERE IN THE PACKAGES ITSELF, JUST IN THE SHIPPING CONTAINER. WE FEEL THE LOT NUMBER COULD BE EITHER 47020 OR 47301. | 0
4992900 | AT APPROXIMATELY 11:57 AM ALARMS STARTING SOUNDING. THE PATIENT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PATIENT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATION ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL PATIENT MONITOR. AFTER ABOUT 15 MINUTES IT WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PATIENT MONITOR WAS REMOVED FROM THE NETWORK AND THE PATIENT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
5225960 | ON (B)(6) 2015, THE CUSTOMER REPORTED THAT THEY HAD AN ISSUE ON (B)(6) 2015, WHERE THEY WERE UNABLE TO LOG-IN TO THE PHILIPS XPER WORKSTATION DURING CUSTOMER NETWORK OUTAGE. BECAUSE THEY COULD NOT LOG IN THEY WERE UNABLE TO MONITOR THE PATIENT WITH OUR EQUIPMENT. BASIC MONITORING WAS AVAILABLE USING ALTERNATE EQUIPMENT. CUSTOMER CHOOSE TO DISCONTINUE THE CASE AND THE PATIENT WAS TRANSPORTED TO ANOTHER FACILITY. IT WAS LATER REPORTED THE PATIENT PASSED A FEW DAYS LATER. THE INVESTIGATION IS CURRENTLY ON-GOING.><THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. | 1
4501337 | THE NEEDLE DRIVER STOPPED WORKING. IT HAD 8 USED LEFT... THIS IS A REPOSABLE INSTRUMENT.MANUFACTURER RESPONSE FOR ROBOTIC NEEDLE DRIVER, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================TOOK DESCRIPTION OF PROBLEM AND ISSUED RGA#. | 0
4493985 | FRONT DOOR PANEL LATCH WAS BROKEN AND THE PART WAS REPLACED. WHILE TURNING VENT ON TO RUN A DEVICE CHECK, A HARD DRIVE FAILURE OCCURRED. THIS SEEMS TO BE A COMMON ISSUE WITH THESE VENTS. ORDERED A REPLACEMENT HARD DRIVE, AS TECH SUPPORT RECOMMENDED, THEN RELOADED SOFTWARE. THIS DID NOT FIX PROBLEM. THE NEXT RECOMMENDATION FROM TECH SUPPORT WAS TO LOAD A NEWER SOFTWARE VERSION FROM 2.31 TO 2.41. NO SUCCESS WITH THIS EITHER. NOW VENT IS AWAITING A FIELD SERVICE TECH TO PERFORM THE LATEST DEVICE CORRECTION CONSISTING OF UPGRADING HARD DRIVE AND BATTERIES.======================MANUFACTURER RESPONSE FOR VENTILATOR, C500/VN500 BABYLOG (PER SITE REPORTER).======================FIELD SERVICE TECH IS COMING IN TO FIX VENT. | 0
5249665 | THE MACHINE WAS INSTALLED BY THE VENDOR AND THE WIRING FROM THE MACHINE TO THE VISUAL CUE (X-RAY IN USE) WAS INSTALLED BY AN UNKNOWN CONTRACTOR. WHEN THE MACHINE WAS BEING ADJUSTED FOR THE PATIENT, IT SEEMS TO HAVE PULLED THE ENERGIZED WIRES AWAY FROM THE INTERNAL CIRCUIT BOARD CAUSING IT TO SHORT. THE ELECTRICAL SHORT CAUSED SMOKE AND SPARKS WHICH WERE WITNESSED BY THE CLINICAL TEAM. MANUFACTURER RESPONSE FOR CARESTREAM DENTAL PANOREX X-RAY MACHINE, (BRAND NOT PROVIDED) (PER SITE REPORTER): UNKNOWN.><. | 0
4959641 | OUR TELEMETRY SYSTEM IS COMPRISED OF SIX CICS (CLINICAL INFORMATION CENTER). THEY ARE CONNECTED TO MONITORS WHICH GIVE OUR NURSING STAFF THE VITAL SIGNS NEEDED FOR THAT PATIENT. ONCE A WEEK A MEMBER OF TELEMETRY OR BIOMED NEED TO SHAKE THE MOUSE THAT IS CONNECTED TO THE CIC OR THE CIC WILL FREEZE UP OR IT WILL NOT GIVE CERTAIN VITAL SIGNS FOR SOME PATIENTS: THE VITAL SIGNS DO NOT SHOW UP ON THE CIC. WE HAVE CONTACTED THE MANUFACTURER BUT THEY ARE UNABLE TO DETERMINE THE CAUSE OR A SOLUTION TO THE PROBLEM OTHER THAN SHAKING THE MICE. OTHER CIC'S AT THIS FACILITY WITHOUT THE TELEMETRY APPLICATION DO NOT HAVE THIS PROBLEM.====================== MANUFACTURER RESPONSE FOR CIC, CLINICAL INFORMATION CENTER (PER SITE REPORTER)======================THEY TOLD US THAT THEY DO NOT HAVE A SOLUTION AT THIS POINT. | 1
5038355 | .><WHILE USING THE MARYLAND FORCEP FOR THE DAVINCI ROBOT SI, NOTICED ARCING FROM THE FORCEP. NO HARM TO PATIENT OR STAFF. | 0
5337740 | .><DURING THE SURGERY THE INSTRUMENT ALERTED WITH THE MESSAGE BLADE EXPOSED | 0
5096419 | PT INTUBATED BY ANESTHESIA ON CATH LAB TABLE FOR CARDIAC ABLATION PROCEDURE. BIOSENSE WEBSTER REP IN ROOM WHILE RECEIVING ERRATIC ERROR" MESSAGE FROM COMMUNICATION OF LOCATION (MAGNET) PAD. SHE TRIED REBOOTING SEVERAL TIMES AND PLACED A CALL TO BIOSENSE WEBSTER | 1
5031041 | FENESTRATED BIPOLAR FORCEPS BROKE WHILE INSIDE THE PATIENT. A SMALL PIECE OF PLASTIC CAME OFF WHERE THE INSTRUMENT ARTICULATES. THE SMALL PIECE WAS RECOVERED FROM THE PATIENT. NO HARM CAME TO THE PATIENT. UNFORTUNATELY NEITHER THE DEVICE OR THE PACKAGING WAS SAVED FROM THE EVENT. MANUFACTURER RESPONSE FOR FENESTRATED BIPOLAR FORCEPS, DA VINCI SI 5 MM FENESTRATED BIPOLAR FORCEPS (PER SITE REPORTER): THE MANUFACTURER HAS BEEN RESPONSIVE.><. | 0
5291986 | THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) CRASHED AND THEY HAD TO REPLACE THE HARD DRIVE. THEY DID NOT HAVE ANY SCREEN SHOTS OR BACK UP SETTINGS FOR THE CNS AND THEY NEED HELP CONFIGURING THE MONITOR.><THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) CRASHED AND THEY HAD TO REPLACE THE HARD DRIVE. THEY DID NOT HAVE ANY SCREEN SHOTS OR BACK UP SETTINGS FOR THE CNS AND THEY NEED HELP CONFIGURING THE MONITOR. NIHON KOHDEN SENT THE CUSTOMER THE CNS DEFAULT SHEETS VIA EMAIL. NIHON KOHDEN ASSISTED THE CUSTOMER IN CONFIGURING THE CNS. REQUESTED CUSTOMER FILL OUT THE DEFAULT SHEET AND RETURN TO NIHON KOHDEN. THE CNS IS NOW FULLY FUNCTION. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED. | 1
4384663 | THE HARD DRIVE WAS REIMAGED AND THE SYSTEM WAS DOWNGRADED TO A MORE STABLE SOFTWARE VERSION. THIS HISTORICAL COMPLAINT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF COMPLAINT FILES IN RESPONSE TO A RECENT FDA INSPECTION. THERE IS NO CHANGE TO THE ACTUAL PERFORMANCE OF ANY PRODUCT AND THIS REPORT ONLY REPRESENTS AN ENHANCEMENT TO THE REPORTING CRITERIA GOING FORWARD.><(B)(4). | 1
4922402 | AFTER TROUBLESHOOTING WITH THE CUSTOMER, IT WAS DETERMINED THAT THE DISPLAY IS DEFECTIVE. THE CENTRAL MONITOR SYSTEM IS FUNCTIONING. THE CUSTOMER PUT A REPLACEMENT MONITOR IN PLACE AND WILL BE ORDERING ANOTHER MONITOR.><(B)(4). | 1
4918125 | OUR INSTALLATION SPECIALIST WENT OUT TO THE SITE AND REDUCED THE AMOUNT OF ATTENUATION TO THE TELEMETRY ANTENNAS FOR THE IMAGING DEPARTMENT PART OF THE ANTENNA COVERAGE AREA AND CORRECTING THE ANTENNA SYSTEM SETUP. CUSTOMER ISSUE WAS RESOLVED.><IMP# (B)(4). | 1
5126701 | THE CUSTOMER REPORTED THAT THE TOUCH SCREEN ON THE CNS (CENTRAL MONITORING SYSTEM) WAS NOT WORKING.><.><NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED.><ADDITIONAL MANUFACTURER NARRATIVE: CUSTOMER REQUESTED PART NUMBER FOR NEW 24 TOUCH SCREEN. CUSTOMER WAS PROVIDED WITH PART NUMBER AND WAS TRANSFERRED TO CUSTOMER SERVICE." | 1
5209732 | THE BIOMEDICAL ENGINEER REPORTED THAT THE HOSPITAL DID A GENERATOR TEST AND THE CNS (CENTRAL MONITORING SYSTEM) AND MONITOR POWERED OFF. THE BIOMEDICAL ENGINEER STATES HE CHANGED THE BATTERIES IN THE UPS A MONTH AGO. THE CNS WILL NOW NO LONGER MOVE PAST THE SPLASH SCREEN UPON START UP. IT STATES TO PLEASE WAIT A MINUTE. THE BIOMEDICAL ENGINEER WAITED 4 HOURS AND THE CNS NEVER CAME BACK UP. THE BIOMEDICAL ENGINEER HAS PUT A SPARE CNS IN PLACE.><THE BIOMEDICAL ENGINEER REPORTED THAT THE HOSPITAL DID A GENERATOR TEST AND THE CNS (CENTRAL MONITORING SYSTEM) AND MONITOR POWERED OFF. THE BIOMEDICAL ENGINEER STATES HE CHANGED THE BATTERIES IN THE UPS A MONTH AGO. THE CNS WILL NOW NO LONGER MOVE PAST THE SPLASH SCREEN UPON START UP. IT STATES TO PLEASE WAIT A MINUTE. THE BIOMEDICAL ENGINEER WAITED 4 HOURS AND THE CNS NEVER CAME BACK UP. THE BIOMEDICAL ENGINEER HAS PUT A SPARE CNS IN PLACE. THE DEVICE WAS RETURNED TO NIHON KOHDEN, EVALUATED, AND THE REPORTED ISSUE WAS CONFIRMED. BOTH OF THE HARD DRIVES WERE REPLACED AND THE DEVICE WAS RETURNED TO THE CUSTOMER. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | 1
5055369 | THE CUSTOMER REPORTED THEY WERE EXPERIENCING A SIGNAL LOSS.><INVESTIGATION IS CURRENTLY UNDERWAY AT NIHON (B)(4). A SUPPLEMENTAL REPORT WILL BE FILED. DEVICE NOT RETURNED. | 1
4792826 | ADDITIONAL MANUFACTURER NARRATIVE: ATTEMPTED TO DETERMINE WHETHER THE CUSTOMER REPLACED THE DISPLAY MONITOR AND IF THAT RESOLVED THEIR ISSUE. CUSTOMER DID NOT RETURN OUR PHONE CALL REQUESTS FOR ADDITIONAL INFORMATION. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE.><PLEASE REFER IMPORTER REPORT#: (B)(4).><THE CUSTOMER WAS GOING TO REPLACE THE MONITOR FOR THE CENTRAL MONITORING SYSTEM THEMSELVES. WE CONTACTED THE CUSTOMER BY EMAIL TWICE AND LEFT A MESSAGE SO THAT WE CAN CONFIRM THAT THEY REPLACED THE DISPLAY MONITOR FOR THE SYSTEM AND THAT RESOLVED THEIR ISSUE. WE HAVE NOT RECEIVED ANY COMMUNICATION FROM THEM. | 1
5158596 | NIHON (B)(4) CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON (B)(4) WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE TELE DEVICES ON THE CNS (CENTRAL MONITORING SYSTEM) ARE SHOWING SIGNAL LOSS FOR PERIODS OF 1-20 SECONDS INTERMITTENTLY ON DIFFERENT SECTORS. | 1
5312713 | NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><ADDITIONAL INFORMATION: THE CUSTOMER WAS SENT A HARD DRIVE TO FIX THE ISSUE. CUSTOMER WAS EMAILED THE PASSWORD THAT WAS NEEDED FOR ELECTRONIC REGISTRATION.><THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) BOOTS INTO A BLUE ERROR SCREEN THAT READS CANNOT FIND BOOT DEVICE". PATIENTS ARE ON HARD WIRED BEDSIDE." | 1
4421015 | (B)(4).><PER THE CUSTOMER (B)(6), THE ISSUE IS WITH THE SPLITTER BOX AS THE ISSUE FOLLOWS THE SPLITTER BOX. AWAITING REQUESTED DEVICE FOR FAILURE INVESTIGATION. | 1
5225233 | THE BIOMEDICAL ENGINEER REPORTED THE CNS (CENTRAL MONITORING SYSTEM) IS EXPERIENCING A COMMUNICATION LOSS.><NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. | 1
4949335 | PLEASE REFER TO (B)(4).><WE CONTACTED THE CUSTOMER AND EXPLAINED TO THEM THAT THE NETWORK IS SET UP ON THE HOSPITAL STRUCTURE, AND THAT WE CAN NOT SEE HOW THE SWITCHES AND RUNS ARE SET UP. CUSTOMER WILL CHECK WITH THEIR IT DEPARTMENT TO SEE IF THERE IS A CENTRALIZED SWITCH THAT IS LOSING POWER OR FAULTY. OUR SUPPORT TEAM IS WORKING WITH THE CUSTOMER AND INVESTIGATION IS STILL UNDER PROCESS. | 1
4902145 | CONFIRMED CUSTOMER COMPLAINT. THIS UNIT CAME IN WITH DEGRADED HARD DRIVES. THE CNS HARD DRIVES COULD NOT BE RECOVERED THEY FROZE AT 16% COMPLETION. BOTH HARD DRIVES WERE REPLACED AND THE REPAIRED UNIT WAS RETURNED TO THE CUSTOMER.><IMP # (B)(4). | 1
4792848 | PLEASE REFER IMPORTER REPORT#: (B)(4).><ADDITIONAL MANUFACTURER NARRATIVE: THE BME CONTACTED NIHON KOHDEN AND STATED THAT HE PURCHASED A (B)(4) MOTHERBOARD, WHICH HE CLAIMS IS ABOUT THE SAME AS THE ONE RECOMMENDED BY THE MANUFACTURER. HE STATED THIS RESOLVED THE ISSUE AND THAT THERE WAS NO USE FOR THE LOANER. NIHON KOHDEN VERBALLY EXPLAINED TO THE CUSTOMER THAT THIS DEVICE WAS NOW CONSIDERED ADULTERATED AS IT DOES NOT HAVE THE MOTHERBOARD SPECIFIED BY THE MANUFACTURER.><CUSTOMER BELIEVES THAT THE PROBLEM IS WITH THE MOTHERBOARD AFTER TESTING WITH ANOTHER POWER SUPPLY. WE FOLLOWED UP WITH THE CUSTOMER AND LEFT A MESSAGE FOR ADDITIONAL INFO ON THE DEVICE SUCH AS IF THEY HAD REPLACED THE MOTHERBOARD OR IF OTHER REPAIRS WERE DONE AND IF THE ISSUE WITH THE DEVICE WAS RESOLVED. THE CUSTOMER HAS NOT RESPONDED BACK. | 1
4390874 | THE CUSTOMER RE-INITIALIZED THE BEDSIDE MONITORS, PROGRAMMED SETTINGS AND CONNECTED BACK ON THE AG-920RA, ISSUE HAS NOT RETURNED AFTER FULL DAYS TESTING. ISSUE RESOLVED. THIS HISTORICAL COMPLAINT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF COMPLAINT FILES IN RESPONSE TO A RECENT FDA INSPECTION. THERE IS NO CHANGE TO THE ACTUAL PERFORMANCE OF ANY PRODUCT AND THIS REPORT ONLY REPRESENTS AN ENHANCEMENT TO THE REPORTING CRITERIA GOING FORWARD.><(B)(4). | 1
5177214 | EVENT PROBLEM AND EVALUATION CODES MANUFACTURER NARRATIVE: THE CUSTOMER REPORTED THAT THE EMERGENCY DEPARTMENT BED 9 IS SHOWING AN ALARM FROM EMERGENCY DEPARTMENT BED 4 ON THE CNS (CENTRAL MONITORING SYSTEM). THE BIOMEDICAL ENGINEER ALSO STATES THAT THE CLINICAL STAFF SAID AN ALARM DIDN'T RING ON THE PATIENT. CUSTOMER WAS INSTRUCTED TO CHECK INTERBED SETTINGS, STOP MONITORING BED 4, AND POWER CYCLE BED 4 TO SEE IF THE ALARM MESSAGE WOULD GO AWAY. MESSAGE WENT AWAY ON BOTH BSMS. CUSTOMER WAS INSTRUCTED TO CHECK ALARM HISTORY IN THE TREND TAB BECAUSE IF THE ALARM IS SHOWING THEN IT ALARMED. EXPLAINED TO CUSTOMER THAT ALARMS CAN BE SILENCED AT THE BEDSIDE OR VIA INTERBED AND IF SOMEONE USED THE LARGE ALARM SILENCE ON THE CNS IT COULD HAVE SILENCED MULTIPLE ALARMS AT THE SAME TIME. NO FURTHER ASSISTANCE WAS NEEDED. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE EMERGENCY DEPARTMENT BED 9 IS SHOWING AN ALARM FROM EMERGENCY DEPARTMENT BED 4 ON THE CNS (CENTRAL MONITORING SYSTEM). THE BIOMEDICAL ENGINEER ALSO STATES THAT THE CLINICAL STAFF SAID AN ALARM DIDN'T RING ON THE PATIENT.><THE CUSTOMER REPORTED THAT THE EMERGENCY DEPARTMENT BED 9 IS SHOWING AN ALARM FROM EMERGENCY DEPARTMENT BED 4 ON THE CNS (CENTRAL MONITORING SYSTEM). THE BIOMEDICAL ENGINEER ALSO STATES THAT THE CLINICAL STAFF SAID AN ALARM DIDN'T RING ON THE PATIENT. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED. | 1
4792847 | THE HOSPITAL REBOOTED THE SYSTEM AND MANY UNITS WERE BACK ONLINE. HOWEVER, ALL OF THE FACILITIES TELEMETRY REMAINED IN A STATE OF COMMUNICATION LOSS. THEY FOUND A SWITCH THAT WAS INDICATING NO COMMUNICATION LINK. A POWER CYCLE DID NOT HELP. AT THAT TIME, THE BIOMED HAD AN EXTRA SWITCH AND PLACED IT INTO SERVICE. ALL ORG RECEIVERS THAT WERE ON THIS SWITCH IMMEDIATELY CAME OUT OF COMMUNICATION LOSS. WE FOLLOWED UP WITH THE CUSTOMER AND THE BIOMED STATES THE SYSTEM IS OPERATING NORMALLY.><PLEASE REFER IMPORTER REPORT#: (B)(4). | 1
4384614 | (B)(4).><WE HAD THE CUSTOMER RESTART POWER TO THE CNS, AND THE CUSTOMER FOUND THAT WHILE MOVING CABLES FOUND THE ISSUE WAS RESOLVED. ADVISED USER TO SECURE ALL CABLES. UNIT WAS NOT RETURNED AS ISSUE WAS RESOLVED. THIS HISTORICAL COMPLAINT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF COMPLAINT FILES IN RESPONSE TO A RECENT FDA INSPECTION. THERE IS NO CHANGE TO THE ACTUAL PERFORMANCE OF ANY PRODUCT AND THIS REPORT ONLY REPRESENTS AN ENHANCEMENT TO THE REPORTING CRITERIA GOING FORWARD. | 1
5252565 | (B)(4). CUSTOMER WAS SENT A NEW HARD DRIVE WHICH RESOLVED THE ISSUE.><THE CUSTOMER REPORTED THAT THERE WAS A POWER OUTAGE AND WHEN THE GENERATOR CAME ON THE CNS (CENTRAL MONITORING SYSTEM) REBOOTED AND IS STUCK AT THE WINDOWS 2000 SCREEN WITH AN ERROR MESSAGES STATING WINDOWS 2000 CANNOT LOAD DUE TO A MISSING OR CORRUPT FILE". THE CNS WAS POWER CYCLED SEVERAL TIMES. CUSTOMER STATES HE BELIEVES THE UPS WAS NON-OPERATIONAL AND THAT THE HARD DRIVE MOST LIKELY HAS NEVER BEEN REPLACED OR THE CNS WAS NEVER POWER CYCLED. CUSTOMER STATES HE WOULD RATHER PURCHASE A NEW HARD DRIVE INSTEAD OF SENDING THE DEVICE IN FOR REPAIR. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THERE WAS A POWER OUTAGE AND WHEN THE GENERATOR CAME ON THE CNS (CENTRAL MONITORING SYSTEM) REBOOTED AND IS STUCK AT THE WINDOWS 2000 SCREEN WITH AN ERROR MESSAGES STATING "WINDOWS 2000 CANNOT LOAD DUE TO A MISSING OR CORRUPT FILE". THE CNS WAS POWER CYCLED SEVERAL TIMES. CUSTOMER STATES HE BELIEVES THE UPS WAS NON-OPERATIONAL AND THAT THE HARD DRIVE MOST LIKELY HAS NEVER BEEN REPLACED OR THE CNS WAS NEVER POWER CYCLED. CUSTOMER STATES HE WOULD RATHER PURCHASE A NEW HARD DRIVE INSTEAD OF SENDING THE DEVICE IN FOR REPAIR." | 1
4850286 | (B)(4). NIHON KOHDEN RECEIVED A CALL BY (B)(6), BIOMED. (B)(6) (BIOMED) THROUGH THE PROCEDURE FOR DELETING AND REINSTALLING THE PRINTER DRIVER. PRINTED A TEST PAGE. UPON COMPLETION OF THE PROCEDURE, REQUESTED BIOMED HAVE CLINICAL STAFF PRINT A FULL DISCLOSURE PAGE. NO ISSUES NOTED AT THIS TIME. BIOMED EXPRESSED CONCERN THAT THIS ISSUE MAY NOT BE RESOLVED. I REQUESTED THAT HE WAIT AFTER A PERIOD OF TIME AND THEN PRINT ANOTHER FULL DISCLOSURE PAGE FROM A DIFFERENT PATIENT. IF ISSUE STILL EXISTS TO CALL TECH SUPPORT BACK. NO FURTHER CONTACT FROM CUSTOMER.><(B)(4).><WE CALLED TWICE AND LEFT A MESSAGE WITH THE CUSTOMER THAT THIS COULD BE AN ISSUE WITH THE PRINTER DRIVER AND TO DELETING AND REINSTALLING THE DRIVERS COULD ELIMINATE THIS PROBLEM. WE FOLLOWED UP WITH THE CUSTOMER AND WAS TOLD THAT THEY HAD SOME COMMUNICATION LOSS. THIS WAS CONFIRMED BY THE STRIPS THAT WAS PROVIDED BY THE CUSTOMER. WE ASKED THEM IF THEY HAVE TRIED DELETING AND REINSTALLING THE DRIVERS AND IF THAT RESOLVED THEIR ISSUE. THEY STATED THAT THEY HAVE NOT TRIED THAT YET AND WILL CALL WHEN THEY ARE READY TO DO THAT. | 1
5102727 | THE CUSTOMER REPORTED THAT ONE OF THE HARD DRIVES ON THE CNS (CENTRAL MONITORING SYSTEM) HAD A FAILURE. THE BIOMED REPLACED THE HARD DRIVE AND THE NEW HARD DRIVE FAILED. THE BIOMED WAS ABLE TO GET THE CNS WORKING BUT IT WAS REPORTED TO HAVE BEEN FREEZING. THE CUSTOMER WAS ADVISED THAT THEY SHOULD NOT USE THE CNS UNTIL THE HARD DRIVE REPLACEMENTS WERE RECEIVED FROM NIHON (B)(4) AND WERE INSTALLED. TWO HARD DRIVES WERE CONFIGURED AND PROVIDED TO THE CUSTOMER. NIHON (B)(4) RECEIVED AND EVALUATED THE RETURNED HARD DRIVES AND THE PROBLEM WAS DUPLICATED. NIHON (B)(4) WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) HAD A HARD DRIVE FAILURE.><. | 1
4888352 | (B)(4).><THE CENTRAL MONITORING SYSTEM WAS REBOOTED AND THE ISSUE HAS BEEN RESOLVED. | 1
4668470 | WE CALLED THE CUSTOMER BIOMED. HE INDICATED THAT THEY PURCHASED A NEW DISPLAY AND PUT IT INTO SERVICE AND THAT RESOLVED THE ISSUE. ALSO, THE FAILED DISPLAY WAS DISCARDED ON SITE BY THE CUSTOMER.><(B)(4) . | 1
4834650 | AN EXCHANGE RNS SERVER WAS SENT TO THE CUSTOMER WHICH RESOLVED THE CUSTOMER ISSUE. REPORT OF SECOND OF 2 RNS UNITS HAVING COMMUNICATION LOSS ISSUES.><THE CUSTOMER REPORTED THAT BOTH REMOTE NETWORK STATIONS (RNS OR REMOTE MONITORING SYSTEM) ARE EXPERIENCING COMMUNICATION LOSS. | 1
5328175 | MANUFACTURER NARRATIVE: THE CUSTOMER REPORTS THAT THE HOSPITAL LOST POWER AND THE CNS (CENTRAL MONITORING SYSTEM) REBOOTED. DURING THE BOOT PROCESS, THE CNS FREEZES ON THE SPLASH SCREEN. CUSTOMER ALLOWED UNIT TO WORK FOR APPROXIMATELY ONE HOUR AND IT NEVER LEFT THE SPLASH SCREEN. THE CUSTOMER TRIED A HARD BOOT (REMOVED POWER, THEN RESTARTED AFTER ABOUT A MINUTE) AND THE DEVICE HAD THE SAME RESULT. THE DEVICE DOES NOT COMPLETE BOOTING INTO THE MONITORING SOFTWARE. THE CUSTOMER REPORTS THAT THE UPS WAS NOT FUNCTIONING PROPERLY WHEN THE HOSPITAL LOST POWER. THE UNIT WAS EVALUATED AND THE REPORTED PROBLEM WAS DUPLICATED. THE DEVICE HAS BEEN REPAIRED AND RETURNED TO THE CUSTOMER. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE.><NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTS THAT THE HOSPITAL LOST POWER AND THE CNS (CENTRAL MONITORING SYSTEM) REBOOTED. DURING THE BOOT PROCESS, THE CNS FREEZES ON THE SPLASH SCREEN. CUSTOMER ALLOWED UNIT TO WORK FOR APPROXIMATELY ONE HOUR AND IT NEVER LEFT THE SPLASH SCREEN. THE CUSTOMER TRIED A HARD BOOT (REMOVED POWER, THEN RESTARTED AFTER ABOUT A MINUTE) AND THE DEVICE HAD THE SAME RESULT. THE DEVICE DOES NOT COMPLETE BOOTING INTO THE MONITORING SOFTWARE. THE CUSTOMER REPORTS THAT THE UPS WAS NOT FUNCTIONING PROPERLY WHEN THE HOSPITAL LOST POWER. | 1
4384645 | THE SYSTEM WAS DOWNGRADED TO A MORE STABLE SOFTWARE VERSION AS A PART OF THE RECALL INITIATED (B)(6) 2014. THIS HISTORICAL COMPLAINT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF COMPLAINT FILES IN RESPONSE TO A RECENT FDA INSPECTION. THERE IS NO CHANGE TO THE ACTUAL PERFORMANCE OF ANY PRODUCT AND THIS REPORT ONLY REPRESENTS AN ENHANCEMENT TO THE REPORTING CRITERIA GOING FORWARD.><(B)(4). | 1
5190446 | THE CUSTOMER REPORTED THAT THE LCD IS COMPLETELY BLANK (BLACK) ON THE BSM (BEDSIDE MONITOR) BUT THE HARD KEY BUTTONS ARE STILL RESPONSIVE. THE BIOMEDICAL ENGINEER IS REQUESTING A REPLACEMENT INVERTER BOARD.><THE CUSTOMER REPORTED THAT THE LCD IS COMPLETELY BLANK (BLACK) ON THE BSM (BEDSIDE MONITOR) BUT THE HARD KEY BUTTONS ARE STILL RESPONSIVE. THE BIOMEDICAL ENGINEER IS REQUESTING A REPLACEMENT INVERTER BOARD. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED.><MANUFACTURER NARRATIVE: THE CUSTOMER REPORTED THAT THE LCD IS COMPLETELY BLANK (BLACK) ON THE BSM (BEDSIDE MONITOR) BUT THE HARD KEY BUTTONS ARE STILL RESPONSIVE. THE BIOMEDICAL ENGINEER REQUESTED A REPLACEMENT INVERTER BOARD. CUSTOMER WAS PROVIDED WITH PART NUMBER FOR INVERTER BOARD AND TRANSFERRED TO CUSTOMER SERVICE TO ORDER AN INVERTER BOARD. FOLLOWED UP WITH CUSTOMER. ON (B)(6) 2015 CUSTOMER CALLED BACK AND ADVISED THE INVERTER BOARD FIXED THE ISSUE. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | 1
5019831 | THE CUSTOMER REPORTED THAT THE CENTRAL MONITORING SYSTEM (CNS) WAS DISPLAYING RANDOM SOFTWARE ERRORS WHICH IS CAUSING THE CNS TO REBOOT.><THE CUSTOMER ORDERED AND RECEIVED A REPLACEMENT HARD DRIVE BUT THEY HAVE NOT INSTALLED IT YET. THE CUSTOMER STATED THAT THEY HAVE NEVER REPLACE THE HARD DRIVE ON THIS UNIT. REPLACING THE HARD DRIVE EVERY TWO YEARS IS A PART OF THE PREVENTATIVE MAINTENANCE NOTED IN THE SERVICE MANUAL. HOSPITAL HAS NOT RETURNED HARD DRIVE. | 1
4384748 | (B)(4).><THE CUSTOMER PURCHASED A NEW TOUCH SCREEN DISPLAY. NO FURTHER INFO WAS PROVIDED. THIS HISTORICAL COMPLAINT IS BEING FILED AS PART OF A RETROSPECTIVE REVIEW OF COMPLAINT FILES IN RESPONSE TO A RECENT FDA INSPECTION. THERE IS NO CHANGE TO THE ACTUAL PERFORMANCE OF ANY PRODUCT AND THIS REPORT ONLY REPRESENTS AN ENHANCEMENT TO THE REPORTING CRITERIA GOING FORWARD. | 1
4991228 | WE SPOKE TO OUR ONCALL TECH SUPPORT TO CONFIRM THE DETAILS OF THE COMPLAINT, AND THE FOLLOWING ARE THOSE NOTES. THE NURSE REPORTED THAT THE CENTRAL MONITORING SYSTEM HAD A WINDOWS ERROR AND THAT THE UNIT WAS REBOOTED. THEY STATED THAT EVERYTHING WAS WORKING FINE. WE TOLD THEM TO NOTIFY THEIR BIOMED OF THE ERROR AND TO HAVE THE BIOMED CALL US IN THE MORNING. THEREFORE, THE UNIT DID NOT REBOOT BY ITSELF. ALSO, THE CUSTOMER STATED THAT THERE WAS NO DATA LOST EXCEPT FOR DURING THE TIME THE NURSE HAD REBOOTED THE UNIT. THERE IS NO FURTHER INFO AT THIS TIME.><(B)(4). | 1
4991241 | THE CUSTOMER BIOMED CHECKED THE CENTRAL MONITORING SYSTEM TO SEE IF IT ALARMED THERE TO CONFIRM IF THIS WAS AN ISSUE WITH THE BEDSIDE VITAL SIGNS MONITOR ONLY. THEY STATED THAT THE DEVICE IS SHOWING VENTRICULAR TACHYCARDIA BUT NOT ALARMING. THE MONITOR IS ALSO TURNING OFF. WHEN THE MONITOR WAS TURNED BACK ON THE MONITOR KEPT ALARMING EVEN AFTER IT WAS SILENCED. THE UNIT WAS RETURNED FOR EVAL AND THE ISSUE WAS NOT DUPLICATED. THE SIMILAR ISSUE WAS ONLY DUPLICATED WHEN THE BEDSIDE WAS WARMING UP. THE MONITOR WILL NOT ALLOW ANY ALARMS TO ACTIVATE WHEN FIRST TURNED ON FOR 2 MINUTES. THIS PROBLEM WOULD NORMALLY BE FROM TH AY INPUT MODULE. NOTIFIED THE CUSTOMER THAT WE COULD NOT DUPLICATE THE ISSUE, AND THEY STATED IT WAS FINE TO RETURN THE UNIT. WE CALLED THE CUSTOMER TO REQUEST ADD'L INFO FOR THE INVESTIGATION. AS OF (B)(6) 2015 THE CUSTOMER HAS NOT CALLED.><(B)(4). | 1
5241825 | THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) HAS COMMUNICATION LOSS IN ALL SECTORS. THEY PERFORMED A SYSTEM RESTART AND ALL THE SECTORS CAME BACK UP. ADVISED THE CUSTOMER THAT THEY SHOULD BE PERFORMING RESTARTS EVERY 90 DAYS AND REPLACING THE HARD DRIVE PER MAINTENANCE SCHEDULE. ALSO ADVISED THE HARD DRIVE MAY NEED TO BE REPLACED. CUSTOMER WAS PROVIDED WITH NIHON KOHDEN'S NOTIFICATION NUMBER IN CASE HE HAS ANY FURTHER ISSUES. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED.><THE CUSTOMER REPORTED THAT THE CNS (CENTRAL MONITORING SYSTEM) HAS COMMUNICATION LOSS IN ALL SECTORS. THEY PERFORMED A SYSTEM RESTART AND ALL THE SECTORS CAME BACK UP. ADVISED THE CUSTOMER THAT THEY SHOULD BE PERFORMING RESTARTS EVERY 90 DAYS AND REPLACING THE HARD DRIVE PER MAINTENANCE SCHEDULE. ALSO ADVISED THE HARD DRIVE MAY NEED TO BE REPLACED. PROVIDED NIHON KOHDEN CASE NUMBER. | 1
5270597 | MANUFACTURER NARRATIVE: THE CUSTOMER REPORTED THAT THE BSM (BEDSIDE MONITOR) WOULD NOT ALLOW HER TO DISCHARGE A PATIENT THAT EXPIRED ON THE DEVICE. THIS WAS A DO NOT RESUSCITATE" PLANNED EVENT. PATIENT WAS TAKEN OFF OF LIFE SUPPORT. THE DEVICE ALARMED ASYSTOLE LIKE IT WAS SUPPOSED TO. WHENEVER THE NURSE TRIES TO DISCHARGE THE PATIENT AT THE CNS (CENTRAL MONITORING SYSTEM) OR THE BSM (BEDSIDE MONITOR) THE DEVICE DISPLAYS THE ARRHYTHMIA RECALL AND WILL NOT DISCHARGE. THE CUSTOMER WAS TOLD TO UNPLUG THE ECG CABLES AND THEN SILENCE ALARM AND TRY DISCHARGE. AFTER CLEARING THE ALARM THE PATIENT WAS ABLE TO BE DISCHARGED. THE DEVICE WAS NEVER SENT IN FOR AN EVALUATION AS THE PROBLEM WAS RESOLVED. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE BSM (BEDSIDE MONITOR) WOULD NOT ALLOW HER TO DISCHARGE A PATIENT THAT EXPIRED ON THE DEVICE. THIS WAS A "DO NOT RESUSITATE" PLANNED EVENT. PATIENT WAS TAKEN OFF OF LIFE SUPPORT. THE DEVICE ALARMED ASYSTOLE LIKE IT WAS SUPPOSED TO. WHENEVER THE NURSE TRIES TO DISCHARGE THE PATIENT AT THE CNS (CENTRAL MONITORING SYSTEM) OR THE BSM (BEDSIDE MONITOR) THE DEVICE DISPLAYS THE ARRHYTHMIA RECALL AND WILL NOT DISCHARGE.><THE CUSTOMER REPORTED THAT THE BSM (BEDSIDE MONITOR) WOULD NOT ALLOW HER TO DISCHARGE A PATIENT THAT EXPIRED ON THE DEVICE. THIS WAS A "DO NOT RESUSITATE" PLANNED EVENT. PATIENT WAS TAKEN OFF OF LIFE SUPPORT. THE DEVICE ALARMED ASYSTOLE LIKE IT WAS SUPPOSED TO. WHENEVER THE NURSE TRIES TO DISCHARGE THE PATIENT AT THE CNS (CENTRAL MONITORING SYSTEM) OR THE BSM (BEDSIDE MONITOR) THE DEVICE DISPLAYS THE ARRHYTHMIA RECALL AND WILL NOT DISCHARGE. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. DEVICE NOT RETURNED." | 1
5139493 | THE CUSTOMER REPORTED THAT TWO BSM'S (BEDSIDE MONITORS) WENT INTO MOMENTARY COMMUNICATION LOSS WITH THE CNS (CENTRAL MONITORING SYSTEM).><NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. | 1
4840767 | THE CUSTOMER REPORTED THAT BOTH REMOTE NETWORK STATIONS (RNS OR REMOTE MONITORING SYSTEM) ARE EXPERIENCING COMMUNICATION LOSS ON ALL BEDS FOR AN UNKNOWN AMOUNT OF TIME AND THEN CAME OUT OF COMMUNICATION LOSS.><OUR TECHNICAL SUPPORT WALKED THE NURSE THROUGH RESTARTING THE REMOTE NETWORK STATIONS, AND THE ISSUE WAS RESOLVED. | 1
5040473 | THIS UNIT WILL BE EXCHANGED AND BE REMOVED FROM SERVICE THIS WEEK. NOT RETURNED TO MANUFACTURER.><THE CUSTOMER REPORTED THAT THE REMOTE NETWORK STATION (RNS OR REMOTE MONITORING SYSTEM) HAS LOW VOLUME ALARM. | 1
4790550 | DURING THE PROCEDURE THE ROBOTIC MARYLAND DISSECTOR FELL APART, EMPLOYEE WENT TO PULL OUT THE CORD THAT ATTACHES TO THE INSTRUMENT AND THE PRONGS CAME OUT WITH THE CORD, RENDERING THE INSTRUMENT UNUSABLE. IT WAS EXCHANGED AND PATIENT WAS NOT HARMED. | 0
4514884 | MONOPOLAR CURVED SCISSORS HAD A TORN PROTECTIVE SHEATH. WHEN PLACING THE MONOPOLAR TIP COVER ON THE SHEATH WAS TORN. | 0
4838156 | HEARING AID - EAR CANAL PORTION BROKE OFF INSIDE CANAL. REMOVAL WAS PERFORMED, OUTPATIENT, BY PHYSICIAN. | 0
4684404 | THERE WAS AN UNEXPECTED CRASH OF THE EHR WITH ALL TERMINAL SCREENS BLANK AND NO RECORDS AVAILABLE ON ANY PATIENT IN THE HOSP. WIDESPREAD DELAYS IN CARE OCCURRED. THIS PATIENT SUFFERED A COMPLICATION OF A SIMPLE PROCEDURE, THERE WERE DELAYS AND COMMUNICATION FAILURES DIRECTLY ATTRIBUTED TO THE CRASH OF THE EHR , CONCOMITANT MAYHEM, AND HARRIED NURSES. THESE DELAYS SET IN MOTION A CASCADE OF ADVERSE EVENTS INCLUDING BUT NOT LIMITED TO RESPIRATORY FAILURE AND DELIRIUM, ULTIMATELY RESULTING IN DEATH. | 1
4417590 | ><NOTICED AFTER 3RD ATTEMPT OF INTUBATION THAT THE EXHALATION SIDE OF A AIRWAY CIRCUIT CONNECTED TO COMMON GAS PORT. MOVED THE EXHALATION SIDE OF AIRWAY AND CONNECTED TO EXHALATION PORT. THE CIRCUITS FOR THIS MACHINE ARE INTERCHANGEABLE. PT SAFETY HAZARD. CIRCUITS AND PORTS SHOULD NOT BE INTERCHANGEABLE. RESULTED IN BAROTRAUMA AND LEFT TENSION PNEUMOTHORAX. | 0
4834530 | ON (B)(6) 2015, BLOOD PRESSURE MONITOR MALFUNCTIONED BY OVER INFLATING CUFF TO 243 BEFORE REMOVING FROM ARM. THIS HIGH PRESSURE WAS VERY PAINFUL. | 0
4618267 | THE CUSTOMER NOTICED A BURNING SMELL WAS COMING FROM THE AUTOMATE SAMPLE PROCESSING SYSTEM AND FOUND THE POWER SWITCH WAS IN THE OFF" POSITION. THERE WAS NO EXPOSURE | 0
4992662 | ><WHILE AT (B)(6) CANCER CENTER FOR TREATMENT, A MEDI-TAK PATCH WAS UTILIZED TO WARM PATIENT'S VEIN TO MAKE IT MORE ACCESSIBLE FOR AN IV. A 3RD DEGREE BURN WAS NOTICED AT SITE. CURRENTLY NURSES ARE CHANGING BANDAGES DAILY AND A REFERRAL TO A PLASTIC SURGEON HAS BEEN MADE. | 0
5242289 | .><THE CUSTOMER HAS ALLEGED A POTENTIAL SAFETY ISSUE REGARDING A DRUG DOSAGE (AMIODARONE) IN A STANDARD ORDER DID NOT DISPLAY. THE PATIENT HAD SEVERE CARDIOGENIC SHOCK AND WAS IN MULTI ORGAN FAILURE AND DIED A FEW DAYS LATER.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | 1
4738914 | PHILIPS HEALTHCARE RECEIVED A REQUEST FROM THE HOSPITAL FOR INFORMATION AS TO WHETHER OR NOT IT WAS POSSIBLE TO DETERMINE IF THE DEVICE HAD BEEN SHUT OFF. THE COMPLAINT STATED THAT, WHEN THE DEVICE ALARMED AT 5:00AM ON (B)(6) 2015, THE PATIENT WAS CHECKED AND DETERMINED TO BE FINE. HOWEVER, WHEN THE PARENT CHECKED ON THE PATIENT AT 7:20AM, THE PATIENT WAS REPORTEDLY IN CARDIAC ARREST AND THE DEVICE WAS NOT ALARMING. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL AND PASSED AWAY LATER THAT NIGHT.><THE COMPLAINT STATED THAT, WHEN THE DEVICE ALARMED AT 5:00AM ON (B)(4) 2015, THE PATIENT WAS CHECKED AND DETERMINED TO BE FINE. HOWEVER, WHEN THE PARENT CHECKED ON THE PATIENT AT 7:20AM, THE PATIENT WAS REPORTEDLY IN CARDIAC ARREST AND THE DEVICE WAS NOT ALARMING. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL AND PASSED AWAY LATER THAT NIGHT. PER PHONE CONVERSATION WITH THE HOSPITAL¿S SENIOR PROFESSIONAL CARE MANAGER, THE HOSPITAL DOES NOT FEEL THAT THE DEVICE MALFUNCTIONED, BUT THEY WANTED TO KNOW IF THERE WAS A WAY TO TELL IF THE DEVICE WAS TURNED OFF OR IF IT SHUT DOWN ITSELF. THIS QUESTION WAS BASED ON AN INITIAL REVIEW OF A DOWNLOAD FROM THE DEVICE THAT INDICTED THAT THE DEVICE WAS NOT TURNED ON FROM 12:00AM TO 7:20AM ON (B)(6) 2015. THE HOSPITAL AND THE DEVICE MANUFACTURER REVIEWED THE DOWNLOAD AND IT SHOWED A LOOSE LEAD ALARM AT 12:00AM. FOLLOWING THE LOOSE LEAD ALARM, THE DOWNLOAD SHOWED THAT THE DEVICE WAS MANUALLY SHUT OFF BY THE USER AND WAS NOT TURNED BACK ON. BASED ON THE AVAILABLE DOWNLOAD INFORMATION, THE PARENT¿S CLAIM THAT THE DEVICE ALARMED AT 5:00AM CANNOT BE CONFIRMED. FURTHER REVIEW OF THE DEVICE DOWNLOAD SHOWED THAT UP UNTIL 12:00AM ON (B)(4) 2015, EVERY TIME THE DEVICE ALARMED, THE DEVICE WAS MANUALLY TURNED OFF AND TURNED BACK ON. THE ¿IF AN ALARM SOUNDS¿ SECTION CONTAINS THE FOLLOWING: ¿IF AN ALARM SOUNDS WHILE YOU ARE MONITORING YOUR BABY, CHECK YOUR BABY FIRST. THEN FOLLOW THE INSTRUCTIONS BELOW TO RESPOND TO LIGHTS AND ALARMS. ALWAYS CHECK YOUR BABY¿S SKIN COLOR. IS IT NORMAL? ALWAYS CHECK TO SEE IF YOUR BABY IS BREATHING. IF YOUR BABY IS NOT BREATHING, FOLLOW THE RESPOND LIKE THIS INSTRUCTION.¿ THE PARENTS¿ GUIDE THEN PROVIDES A TABLE THAT EXPLAINS THE VARIOUS ALARM LIGHTS, WHAT KIND OF ALARM IT IS (INTERMITTENT OR CONTINUOUS), A DESCRIPTION OF THE PATIENT¿S CONDITION AND INSTRUCTIONS FOR RESPONDING TO THE ALARM. UNDER THE LOOSE LEADS ALARM, THE CUSTOMER IS PROVIDED WITH TROUBLESHOOTING LIKE CHECKING CONNECTIONS AND THEN RESETTING THE DEVICE. THERE IS ALSO A NOTE THAT STATES THAT ¿PATIENT ALARMS CANNOT BE SILENCED WITH THE RESET BUTTON. THE ALARM WILL STOP ONLY WHEN THE PATIENT SIGNALS ARE WITHIN THE ALARM LIMITS.¿ PHILIPS HEALTHCARE HAS RECEIVED THE DEVICE BACK FOR EVALUATION AND NO OBSERVATIONS SUBSTANTIATING A MALFUNCTION OR OTHER OPERATION OUTSIDE OF DESIGN SPECIFICATIONS WERE MADE DURING THE EVALUATION. THE DEVICE WAS FOUND TO OPERATE AND ALARM WITHIN SPECIFICATIONS TO NOTIFY A CAREGIVER OF A POTENTIAL EVENT. A 12 MONTH COMPLAINT REVIEW HAS BEEN CONDUCTED AND IT IS CONCLUDED THAT THIS IS THE ONLY COMPLAINT FOR THIS ISSUE. PHILIPS HEALTHCARE HAS DETERMINED THAT THE COMPLAINT ISSUE ALLEGED BY THE CUSTOMER WAS NOT OBSERVED DURING THE MANUFACTURER¿S EVALUATION OF THE PRODUCT. THE INVESTIGATION ALSO FOUND THAT THE DEVICE WAS SHUT OFF BY THE USER IN RESPONSE TO AN ALARM CONDITION AND NOT TURNED BACK ON AFTERWARDS. IT IS CONCLUDED THAT USE OF THE DEVICE DOES NOT PRESENT AN INCREASED RISK TO THE END USER OR PATIENT AND THAT NO CORRECTIVE ACTION OR ADDITIONAL INVESTIGATION ACTIVITY IS NECESSARY AT THIS TIME. | 1
4812697 | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><.><THE CUSTOMER REPORTED THAT THE SPEAKER WAS MALFUNCTIONING. NO PATIENT HARM WAS REPORTED. | 1
5318050 | THE CUSTOMER REPORTED THE ROLLSTAND BROKE. THERE WAS NO REPORT OF PATIENT HARM.><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER STATED THE ROLLSTAND SNAPPED IN HALF. NO PATIENT OR USER HARM WAS REPORTED.><. | 0
4794409 | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><.><THE CUSTOMER REPORTED THAT THE MONITOR WAS DISPLAYING A SPEAKER MALFUNCTION ERROR MESSAGE. NO PATIENT HARM WAS REPORTED. | 1
5177331 | .><A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER HAD AN ISSUE WITH THE MX40 WHERE THE DEVICE DID NOT MONITOR ECG ON A PATIENT FOR THREE HOURS. DURING THIS TIME THERE WAS NO ECG ON THE MX40 SCREEN VISIBLE WHILE THERE WAS A PATIENT CONNECTED WITH THE LEAD SET. THERE ALSO WAS NO INOP LEAD OFF" ON THE SCREEN. THE MX40 WAS STILL CONNECTED TO THE CENTRAL STATION AS THE PATIENT NAME AND BED LABEL WERE STILL VISIBLE ON THE MX40. ON THE CENTRAL STATION THERE WAS AN INOP VISIBLE "ECG/ARRH ALARM OFF" | 1
5331526 | A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER CONTACTED PHILIPS TO REPORT AN ISSUE WITH THE CALCULATIONS ON THE SYSTEM. THERE WAS NO PATIENT HARM.><. | 1
4993977 | THE STX PADS IN COMPLAINT WILL NOT BE RETURNED FOR INVESTIGATION. THEREFORE, A PHYSICAL INVESTIGATION WILL NOT BE PERFORMED. IT SHOULD BE NOTED THAT UPON FURTHER INVESTIGATION BY THE WOUND CARE NURSE, IT WAS FOUND THAT THE PATIENT WAS NOT TURNED REGULARLY PER PATIENT PROTOCOL. PER THE IFU, SKIN CHECKS SHOULD BE CONDUCTED EVERY 20 MINUTES. SKIN INJURIES CAN OCCUR WITH ANY SURFACE COOLING DEVICES EVEN WHEN FOLLOWING PRODUCT INSTRUCTIONS FOR ROUTINE SKIN EVALUATIONS AND IN PATIENTS WITHOUT A PRIOR HISTORY OF SKIN PROBLEMS. EVEN THOUGH SKIN INJURY IS AN ANTICIPATED COMPLICATION FOR SURFACE-COOLING, THE RISK CAN BE MITIGATED BY FOLLOWING IFU.><A (B)(6) MALE PATIENT WEIGHING (B)(6) HAD A CARDIAC ARREST AT WORK AND WAS SHOCKED IN THE FIELD. PATIENT WAS UNCONSCIOUS WHEN HE PRESENTED TO THE EMERGENCY DEPARTMENT (ED). PATIENT HAD THE FOLLOWING CLINICAL SYMPTOMS: DIABETES. ON (B)(6) 2015, PATIENT WAS STARTED ON THE INNERCOOL STX SURFACE COOLING PER HOSPITAL PROTOCOL FOR THERAPEUTIC HYPOTHERMIA POST CARDIAC ARREST. THERE WERE NO ADJUNCTIVE TEMPERATURE MANAGEMENT OR RELATIVE PROCEDURES PERFORMED ON THE PATIENT. ON (B)(6) 2015, THE WOUND CARE NURSE WAS ASKED TO CONDUCT A CONSULT OF THE PATIENT WHERE SHE FOUND BLISTERS ON THE PATIENT'S BACKSIDE IN THE SAME PATTERN AS THE STX PADS. UPON FURTHER INVESTIGATION BY THE WOUND CARE NURSE, IT WAS FOUND THAT THE PATIENT WAS NOT TURNED REGULARLY PER PATIENT PROTOCOL. PATIENT WAS ON 2 VASOPRESSORS (NEOSYNEPHRINE AND ANOTHER UNKNOWN TYPE) AND DIPROVAN WHEN THE REPORTED BLISTERS WERE NOTED. THE PATIENT DID NOT REQUIRE INTERVENTION OR TREATMENT FOR THE BLISTERS. PER THE WOUND CARE NURSE, THE BLISTERS WERE INTACT AND VERY MINOR. PATIENT LATER EXPIRED ON THE SAME DAY. PER THE TREATING PHYSICIAN'S OPINION, THE CAUSE OF DEATH WAS CARDIAC ARREST. HOWEVER, THE OFFICIAL CAUSE OF DEATH IS UNKNOWN. IN THE PHYSICIAN'S OPINION, THE REPORTED BLISTERS WERE ATTRIBUTABLE TO THE ZOLL DEVICE BECAUSE THEY HAD THE SHAPE OF THE PADS ON THE PATIENT'S SKIN AND THE BLISTERS WERE IN THE EXACT SHAPE OF THE INSIDE OF THE PADS. IN THE PHYSICIAN'S OPINION, THE PATIENT'S DEATH WAS NOT RELATED TO THE BLISTERS OR TO THE ZOLL DEVICE BECAUSE THE BLISTERS WERE MINOR. THE PHYSICIAN STATED THAT THERE WAS NO PUSS OR INFECTION. NO FURTHER INFORMATION WAS PROVIDED. | 0
4992877 | AT APPROXIMATELY 11:57 AM ALARMS STARTING SOUNDING. THE PATIENT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PATIENT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATION ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL PATIENT MONITOR. AFTER ABOUT 15 MINUTES IT WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PATIENT MONITOR WAS REMOVED FROM THE NETWORK AND THE PATIENT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
5051280 | CALLER REPORTED THE ACCU-CHEK CONNECT APP ADVISED AN INCORRECT BOLUS AMOUNT. CALLER STATED SHE ADMINISTERED THE ADVISED BOLUS OF 5 UNITS OF INSULIN AT 8:55AM AND AT 11:20AM SHE EXPERIENCED HYPOGLYCEMIA; SHE WAS ABLE TO SELF-TREAT. NO ADVERSE EVENT REPORTED. NO PRODUCT WAS REQUESTED TO BE RETURNED.><THE EVENT OCCURRED IN (B)(6). WHILE THIS PRODUCT IS NOT SOLD IN THE UNITED STATES, IT IS LIKE OR SIMILAR TO A PRODUCT MARKETED IN THE UNITED STATES. DEVICE WAS NOT REQUESTED TO BE RETURNED. | 0
4794446 | WHILE PRIMING THE WHITE BLOOD CELL SET ON THE COPE APHERESIS MACHINE, THE MACHINE ALARMED PLASMA PUMP ERROR" - DIRECTIONS WERE FOLLOWED FOR TROUBLESHOOTING ALARMS 3 SEPARATE TIMES. ALARMS COULD NOT BE PASSED. TUBING REMOVED NEW TUBING SET RETRIEVED USED AND PRIMED/PROCEDURE DONE WITHOUT INCIDENT OF ALARMS." | 0
4797714 | ROBOTIC MARYLAND BIPOLAR CABLE BROKE DURING USE. NO PT HARM AND VISUAL INSPECTION INDICATED ALL WAS INTACT. | 0
5167878 | .><PATIENT WITH SEVERE CARDIAC AMYLOIDOSIS UNDERWENT STEM CELL MOBILIZATION FOR PERIPHERAL BLOOD PROGENITOR CELL COLLECTION IN PREPARATION FOR AUTOLOGOUS STEM CELL TRANSPLANTATION. THE COLLECTION PROCEDURE WAS PERFORMED WITHOUT DEVIATIONS OR UNEXPECTED EVENTS; ON ARRIVAL HE COMPLAINED OF NAUSEA AND RECEIVED ONDANSETRON FOR THIS; HIS BLOOD PRESSURE DECREASED MODESTLY DURING THE PROCEDURE BUT REMAINED WITHIN THE PATIENT'S SELF-REPORTED USUAL RANGE; HE WAS EVALUATED BY THE MEDICAL DIRECTOR AND ADDITIONAL FLUID REPLACEMENT WAS INITIATED; HE REMAINED ASYMPTOMATIC WITH STABLE BLOOD PRESSURE. APPROXIMATELY TWO HOURS INTO THE COLLECTION PROCEDURE THE PATIENT ACUTELY DECOMPENSATED: HE SAT UP, VOMITED AND WENT INTO CARDIOPULMONARY ARREST. HE RECEIVED IMMEDIATE CPR, THE CODE TEAM ARRIVED AND WAS ABLE TO RESUSCITATE THE PATIENT FROM PEA (PULSELESS ELECTRICAL ACTIVITY) WITH BRADYCARDIA AND HE WAS TRANSFERRED TO THE MICU. SHORTLY AFTER TRANSFER, HE DEVELOPED BRADYCARDIA AND SUBSEQUENT CARDIOPULMONARY ARREST AGAIN. HE WAS NOT ABLE TO BE RESUSCITATED. THE CAUSE OF DEATH WAS ATTRIBUTED TO CARDIAC AMYLOIDOSIS. | 0
4384857 | DURING A ROBOTIC PROSTATECTOMY, TROUBLE WITH THE INSTRUMENT SWAP OUT.PHYSICIAN REQUESTED A NEW PAIR OF MONOPOLAR SCISSORS, STAFF TRIED TO PUT A NEW PAIR OF SCISSORS IN AND THE SCISSORS WOULD NOT ADVANCE ENOUGH FOR PHYSICIAN TO TAKE OVER AND CONTINUE THE SURGERY, WE UNDOCKED THE ARMS PUT A NEW METAL TROCAR, TESTED THE TROCAR TO MAKE SURE IT WAS ROUND AND NOT BENT, WE TOOK THE FACE PLATE OFF AND ON TO SEE IF THAT WAS THE ISSUE, FINALLY AFTER THE THIRD TIP COVER FOR THIS INSTRUMENT WE WERE ABLE TO ADVANCE THE INSTRUMENT. IT APPEARS THE TIP COVERS WERE THICKER THAN OTHER TIP COVER OF SAME LOT NUMBER. | 0
4729747 | DURING ROBOTIC LAPAROSCOPIC RESECTION, THE HOOK BROKE OFF. | 0
4473000 | STEEL CABLES COMING OUT OF INSTRUMENT WHILE THE PHYSICIAN WAS WORKING.MANUFACTURER RESPONSE FOR DRIVER NEEDLE MEGA SUT CUT, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================TOOK A DESCRIPTION OF THE PROBLEM, ISSUED RGA#. | 0
4459185 | DURING THE OPERATIVE PROCEDURE, THE HARMONIC INSERT SHEARS WAS BEING USED TO DISSECT THE UTERUS. THE TIP OF THE SHEARS BROKE OFF AND WAS LOCATED WITHIN THE PATIENT AND RETRIEVED. | 0
4403776 | DURING PROCEDURE, VESSEL SEALER WAS BEING USED. THE VESSEL SEALER REPORTEDLY REGISTERED A RECOVERABLE FAULT MESSAGE IMMEDIATELY FOLLOWED BY ARTERIAL BLEEDING. CASE WAS CONVERTED TO AN OPEN CASE. BLEEDING WAS CONTROLLED. | 0
4426250 | MY HUSBAND USED THE ALERE PT/INR HOME MONITORING KIT FOR 2 YEARS. WE CHECKED HIS INR EVERY 10 DAYS TO 2 WEEKS. HE WAS CONSISTENTLY IN THE THERAPEUTIC RANGE FOR INR, ACCORDING TO THE MACHINE. WE LAST MEASURED HIS INR ON (B)(6) 2014, AND IT WAS 2.5. ON (B)(6) 2014, HE HAD A MASSIVE STROKE DUE TO A CLOT IN HIS LEFT MIDDLE CEREBRAL ARTERY. HIS INR WAS 2.1 IN THE EMERGENCY DEPARTMENT. HE HAD THE CLOT REMOVED BY A NEURORADIOLOGIST AT (B)(6) WITHIN 90 MINUTES OF SYMPTOM ONSET (SOMNOLENCE, GLOBAL APHASIA, RIGHT HEMIPLEGIA). HE RECOVERED SPEECH BUT CONTINUES TO HAVE RECEPTIVE APHASIA AND SEVERE GAIT DEVIATIONS AND INSTABILITY. IN (B)(6) OF 2014, WE RECEIVED A LETTER FROM ALERE TITLED URGENT: MEDICAL DEVICE CORRECTION. THE LETTER STATED THAT THERE HAD BEEN REPORTS OF RECORDINGS WITH THE ALERE MACHINE THAT WERE LOWER THAN THE ACTUAL INR, SO THERE WAS A RISK OF BLEEDING. WE WANTED TO REPORT MY HUSBAND'S EPISODE, IN CASE THERE HAD BEEN ANY INCIDENT OF THE ALERE MACHINE RECORDINGS BEING HIGHER THAN THE ACTUAL INR, WHICH MAY CONTRIBUTE TO CLOT FORMATION. WE DID NOT HAVE A PROBLEM WITH THE DEVICE, WE JUST WANTED TO REPORT THIS EPISODE, SINCE WE RECEIVED THE LETTER FROM ALERE. | 1
4478145 | WHILE SUTURING MESH, A CABLE FROM THE ENDOWRIST MEGA SUTURECUT NEEDLE DRIVER CAME DISLODGED FROM ITS PULLEY AND STOPPED FUNCTIONING. AFTER INSPECTION, NOTHING APPARENT WAS MISSING OR BROKEN. | 0
4351286 | THE INVESTIGATION INTO THIS ISSUE IS ON-GOING. THE CONCLUSION FROM THE INVESTIGATION WILL BE PROVIDED THROUGH A FOLLOW-UP REPORT. (B)(4).><DATE OF REPORT 02/10/2015. DATE RECEIVED BY MANUFACTURER 02/10/2015. FOLLOW UP REPORT 1. ADDITIONAL INFORMATION / DEVICE EVALUATION. DEVICE EVALUATED BY MANUFACTURER YES. (B)(4). A CUSTOMER FILED A COMPLAINT STATING THAT 12 (B)(6), GENERATED WITH THE COBAS AMPLIPREP / COBAS TAQMAN 48 SYSTEM, WERE MIS-ASSIGNED TO WRONG PATIENTS. THE 12 RESULTS WERE IDENTICAL TO 12 RESULTS PREVIOUSLY RUN WITH DIFFERENT SPECIMENS ON THE SAME THERMAL-CYCLER OF THE COBAS TAQMAN 48 (CTM 48) ANALYZER. THE INVESTIGATION DETERMINED THAT THE CUSTOMER'S PHYSICAL NETWORK SETUP WAS INCORRECT AND THE HARDWARE WAS FAULTY. BOTH THE LABORATORY LOCAL AREA NETWORK (LAN) AND THE INSTRUMENT LAN WERE CONNECTED TO THE SAME NETWORK HUB AND THE HUB WAS MOST LIKELY DEFECTIVE, WHICH CAUSED FREQUENT DISCONNECTIONS OF BOTH CTM 48 ANALYZERS THAT WERE CONNECTED. THE DISCONNECTION EVENTS WERE VISIBLE IN THE RUN LOG OF THE ANALYZERS. AFTER CORRECTING THE NETWORK LAYOUT AND INSTALLATION OF A SWITCH TO REPLACE THE FAULTY HUB, NO FURTHER LAN DISCONNECTION EVENTS OCCURRED. THE NETWORK MIS-CONFIGURATION AND FAULTY HARDWARE AT THIS CUSTOMER SITE CAUSED THE INITIAL RAW DATA DELETION ON THE CTM 48 ANALYZER, REQUESTED BY THE AMPLILINK SOFTWARE, TO FAIL FOR THERMAL CYCLER B. WHEN THE NEW RUN WAS PERFORMED ON THERMAL CYCLER B, RAW DATA WAS DELETED BEFORE THE RUN STARTED, BUT THE DELETION WAS SLIGHTLY DELAYED. A SOFTWARE TIMING ISSUE ALLOWED FOR AN EXTREMELY LIMITED TIMEFRAME WHERE OLD RAW DATA THAT STILL EXISTED ON THERMAL CYCLER B WERE ABLE TO BE USED BY AMPLILINK. NEW RAW DATA GENERATED FOR THE CURRENT RUN ON THERMAL CYCLER B WAS NOT CONSIDERED BY AMPLILINK SINCE RAW DATA FOR ALL CYCLES HAD ALREADY BEEN RECEIVED. AMPLILINK CALCULATED RESULTS FOR THE RUN ON THERMAL CYCLER B USING INCORRECT RAW DATA FROM THE PREVIOUS RUN, WHICH WERE INCORRECT. AFTER THE NETWORK SETUP WAS CORRECTED AND THE POTENTIALLY FAULTY HUB REPLACED, FAILURE OF THE INITIAL RAW DATA DELETION EVENT DID NOT RECUR. THE ISSUE WOULD BE EXTREMELY UNLIKELY TO HAVE OCCURRED OR TO RECUR IF NOT FOR THE IMPROPER LAN CONFIGURATION AND THE FAULTY HUB HARDWARE. THE ISSUE WOULD REQUIRE AN EXTREMELY UNLIKELY PREREQUISITE SEQUENCE OF EVENTS AND TIMING TO OCCUR. THIS WAS THE FIRST OCCURRENCE OF THIS ISSUE REPORTED IN OVER TWELVE YEARS SINCE THE RELEASE OF THE AMPLILINK SOFTWARE. THERE WAS NO HARM REPORTED FOR THIS COMPLAINT. (B)(4)><A CUSTOMER SITE IN (B)(6) FILED A COMPLAINT ALLEGING THAT WRONG RESULTS WERE ASSIGNED TO PATIENT SAMPLES WHEN USING THE COBAS AMPLIPREP / COBAS TAQMAN 48 (CAP/CTM) SYSTEM WITH AMPLILINK SOFTWARE VERSION 3.3.5, RUNNING THE CAP/CTM (B)(4) TEST, V2 . RESULTS WERE ASSIGNED TO 12 SPECIMENS ON A SAMPLE RACK THAT ARE IDENTICAL TO THE FIRST 12 RESULTS ON AN UNRELATED RACK WHICH WAS PREVIOUSLY RUN WITH DIFFERENT SPECIMENS. INCORRECT BUT BELIEVABLE RESULTS WERE GENERATED FOR 11 SAMPLES. THE CUSTOMER STATED THAT (B)(6) TREATMENT FOR ONE PATIENT WAS NOT STARTED DUE TO THE ORIGINAL (B)(6) RESULT REPORTED ON (B)(6) 2014 OF 23 IU/ML. WHEN REPEATED ON (B)(6) 2014, THE (B)(6) RESULT WAS 580,000 IU/ML. | 1
4372875 | ON (B)(6) 2014 CARDINAL SCALE MFG WAS MADE AWARE OF LITIGATION FOR AN ALLEGED INCIDENT THAT INVOLVED THE USE OF A DETECTO SCALE. IN THIS LITIGATION IT IS CLAIMED THAT A HEALTHCARE PROFESSIONAL COMMANDED THE PT TO STAND ON SCALE. AT THIS TIME THE PT WAS STRUCK BY A FOLD DOWN SEAT FROM THE SCALE. PT IS CLAIMING THAT THIS CAUSED FURTHER DAMAGE TO KNEE THAT HAD BEEN SURGICALLY RECONSTRUCTED DAYS EARLIER.><CARDINAL IS IN PROCESS OF MAKING CONTACT WITH THE FACILITY TO EVALUATE THE DEVICE IN QUESTION. WILL SUBMIT FOLLOW UP REPORT ONCE EVAL IS COMPLETED. THE MODEL NUMBER AND SERIAL NUMBER OF SCALE IN THIS REPORT IS BASED ON DESCRIPTION OF SCALE IN PAPERWORK AND RECORDS OF SCALES SOLD TO THIS LOCATION. | 0
4768867 | RADIOGRAPHIC EQUIPMENT EVENT: UNDER TABLE DETECTOR ASSEMBLY-HINGED MOUNTING BRACKET FAILED. SPECIFICS. THE MOUNTING BRACKET THAT CONNECTS THE TABLE BASED TO THE IMAGING DETECTOR ASSEMBLY FAILED CAUSING THE ENTIRE ASSEMBLY TO FALL TO THE FLOOR. | 0
4538361 | THE CABLE BROKE INSIDE THE INSTRUMENT.MANUFACTURER RESPONSE FOR ROBOTIC MEGA SUTURE CUT, (BRAND NOT PROVIDED) (PER SITE REPORTER)======================TOOK DESCRIPTION OF WHAT HAPPENED AND ISSUED RGA#. | 0
5027338 | .><VESSEL SEALER STOPPED FUNCTIONING AS DESIGNED WITH THE BLADE FAILING TO FIRE AND CUT THE INTENDED TISSUE. A 2ND VESSEL SEALER WAS OPENED AND THE PROCEDURE WAS COMPLETED WITH ONLY A MINOR PAUSE AS THE INSTRUMENTS WERE SWAPPED. NO PATIENT HARM RESULTED. | 0
5177208 | ICU NURSE REPORTED THE FOLLOWING: 0300: THE CENTRAL MONITORING SCREEN WENT BLACK WITH NO DISPLAY. IT STILL HAD POWER AND, WHILE TROUBLESHOOTING THE SYSTEM, A SCREEN APPEARED STATING NO SIGNAL INPUT. NURSE CHECKED THE INCOMING CABLES AND THERE WERE NO DISCONNECTED CABLES. THE ONLY THING THAT WAS SHOWING WAS THE ROOM NUMBERS ON THE MONITOR SCREEN. AT 0305: NURSING SUPERVISOR ARRIVED IN UNIT AND STATED THE MED/SUR MONITOR ALSO LOST ITS DISPLAY. AT 0310: PHILIPS TECH SUPPORT NUMBER CALLED. PHILIPS OPERATOR STATED THE TECH WILL BE CALLING SHORTLY. AT 0330: PHILIPS TECHNICIAN RETURNED CALL, WAS UPDATED BY ICU NURSE AND REQUESTED THAT A BIO-MED PERSON LOOK AT THE SYSTEM BELIEVING IT WAS PROBABLY A PROBLEM WITH THE COMPUTER TOWER SERVICING THE MONITOR OR A CABLE FAILURE IN THE SYSTEM. AT 0340: NURSING SUPERVISOR PAGED INFORMATION SYSTEMS (IS). THEY CALLED RIGHT BACK AND ADVISED NURSING TO CALL ANOTHER IS TECHNICIAN BECAUSE HE WAS THE BEST OPTION FOR FIXING THIS PROBLEM. THE IS TECHNICIAN WAS PAGED BY THE NURSING SUPERVISOR. AT 0345: IS CALLED AND ASKED THE NURSE TO ATTEMPT TO RESOLVE THE PROBLEM BY CHECKING VARIOUS CABLES AND UNPLUGGING SOME POWER CORDS TO NO AVAIL. IS RELAYED THAT HE WILL BE COMING TO THE HOSPITAL. AT 0425: IS ARRIVED IN THE ICU AND BEGAN TO TROUBLESHOOT THE SYSTEM. AT 0433: CENTRAL MONITOR BACK ON LINE AND FUNCTIONING AS IT SHOULD. PROBLEM: IS FOUND A POWER SUPPLY MODULE BEHIND THE MONITOR THAT LOOKED SUSPICIOUS WITH MELTED AND DEFORMED HOUSING. IS TOOK THE EQUIPMENT INFORMATION AND CASE NUMBER SAYING HE WAS GOING TO CALL THE PHILIPS TECHNICIAN AND ADVISE HER OF WHAT HE FOUND. THE MED/SUR MONITOR CAME BACK ON LINE AT THE SAME TIME THE ICU MONITOR DID. A NEW POWER SOURCE WAS ORDER TO BE DELIVERED OVERNIGHT. IN THE MEANTIME, IS PROVIDED A TEMPORARY POWER SOURCE. DISCUSSION- PLEASE NOTE: THIS IS ONE OF THREE SIMILAR FAILURES OF MONITORS SECONDARY TO FAILURE OF THE SAME MODEL POWER SUPPLY DURING THE PAST 8 WEEKS. POWER SUPPLIES ARE ORIGINAL POWER SUPPLIES PURCHASED FROM PHILIPS AS PART OF THE MONITOR INSTALLATION. TWO OF THESE FAILURES HAPPENED TODAY. NO ADVERSE PATIENT EVENT AROSE. HOWEVER, LOSS OF PATIENT MONITORING ENDANGERED PATIENT SAFETY. THE POWER SUPPLY OVERHEATING TO THE EXTENT OF MELTING AND CHARRING THE PLASTIC HOUSING AND THE LABEL SUGGESTS A FIRE HAZARD THAT COULD ENDANGER THE SAFETY OF MULTIPLE PATIENTS, HEALTH CARE WORKERS AND VISITORS. PHOTOS OF THE DEFECTIVE POWER SUPPLY WILL BE SENT VIA E-MAIL. THERE IS NO SUGGESTION OF POWER PROBLEMS IN THE BUILDING. NO RECORDED BLACK OUTS OR BROWN-OUT" (POWER DROPS) DURING THIS TIME FRAME. PHILIPS' INITIAL RESPONSE WAS TO ATTRIBUTE THIS TO COMMERCIAL POWER PROBLEMS. THESE ARE THE ONLY THREE DEVICES IN THE BUILDING EXPERIENCING PROBLEMS - WITH THOUSANDS OF OTHER ELECTRICAL / ELECTRONIC DEVICES IN USE. THESE POWER SUPPLIES ARE NOT PHILIPS INTERNALLY MANUFACTURED PRODUCT. THEY ARE PRODUCED BY GLOBTEK | 1
4482884 | VESSEL SEALER QUIT WORKING IN THE MIDDLE OF THE CASE. THE STAFF TRIED TO RESET THE INSTRUMENT, BUT THE SURGEON'S CONSOLE IDENTIFIED IT AS AN ERROR, AND THEREFORE COULD NO LONGER USE THE VESSEL SEALER. | 0
4546581 | ><BATTERY FAILURE DURING USE IN A SURGERY CASE ON THREE DIFFERENT CASES. COMES UP WITH FAILURE ON DISPLAY, BUT BATTERY INDICATES 100% CHARGED. THESE ARE ON MONITORS THAT ARE LESS THAN 8 MONTHS OLD. | 0
4908643 | WHEN SEARCHING FOR RAZADYNE IN OUR CPOE SYSTEM USING RAZ AS THE SEARCH TERM, IT WAS THE 79TH OPTION IN THE DROP DOWN. WHEN SEARCHING RAZ IN LEXICOMP IT IS THE FIRST RESULT RETURNED. THE MOST RELIABLE AND RECOGNIZED DRUG INFORMATION SEARCH ENGINES (LEXICOMP, EPOCRATES, ETC) RETURN RESULTS THAT START WITH THE SEARCH TERM SELECTED AND FOR GOOD REASON. LISTING RESULTS THAT INCLUDE ANY LETTER COMBINATION INSIDE A DRUG NAME IS GOOD FOR A GAME OF (B)(6), BUT BAD FOR MED ORDER ENTRY. (B)(6). THIS DOCUMENT WAS CREATED AS AN INTEGRAL PART OF THE QUALITY ASSURANCE PROGRAM AT THIS COMMAND, AND AS SUCH IS A PRIVILIGED DOCUMENT WHICH IS PROTECTED FROM UNAUTHORIZED DISCLOSURES, DISCUSSION, OR REPRODUCTION BY TITLE 10, U.S. CODE SECTION 1102 (1968). (B)(6). | 1
4552718 | WHEN A MICROTUBE IMAGE DOES NOT EXACTLY FIT THE DEFINED PARAMETERS OF THE READING SOFTWARE FOR NEGATIVE (NO AGGLUTINATION) OR POSITIVE (AGGLUTINATION OF 1+" | 0
4800655 | IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A CARTO 3 SYSTEM AND THE VISITAG'S DISAPPEARED. ERROR 19 (FAILURE IN THE ECG1 CARD) WAS SEEN RANDOMLY DURING THE PROCEDURE, AFTER RF APPLICATION WAS STARTED. THE USER COULD NOT TAKE POINTS, BUT VISITAG'S WERE STILL APPEARING AS DESIRED. THEY REBOOTED THE PATIENT INTERFACE UNIT (PIU) AND THE ERROR CLEARED. HOWEVER, MOST OF THE VISITAG'S DISAPPEARED ONCE THE SYSTEM RE-INITIALIZED AND RESPIRATORY GATING WAS RESET. CHANGING THE SETTINGS UNDER THE VISITAG TOOLS RESULTED IN THE LOSS OF MORE OF THE TAGS TAKEN PRIOR TO THE REBOOT. THE PHYSICIAN DID NOT WANT TO PERFORM MORE TROUBLE SHOOTING AT THIS TIME. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THE BIT ERROR ISSUE WAS ASSESSED AS NOT REPORTABLE. THE MOST LIKELY CONSEQUENCE IS AN INTRAPROCEDURAL DELAY . THE POTENTIAL RISK THAT IT COULD CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH IS REMOTE. THIS EVENT IS BEING REPORTED BECAUSE THE VISITAG¿S DISAPPEARED AND SHOULD REAPPEAR AFTER REINITIALIZING.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4).><EVALUATION SUMMARY: (B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT A PAROXYSMAL ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A CARTO 3 SYSTEM AND SEVERAL ISSUES WERE SEEN DURING THIS PROCEDURE. A CURRENT LEAK¿ ERROR WAS SEEN ON THE CARTO 3 SYSTEM. AFTER EXTENSIVE TROUBLE SHOOTING | 1
5239406 | (B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT A RIGHT SIDED ATRIAL FLUTTER PROCEDURE WITH A CARTO 3 SYSTEM AND TWO LASSO® 2515 NAV ECO VARIABLE CATHETERS A PROCEDURE CANCELLATION OCCURRED BECAUSE THE PHYSICIAN NO LONGER TRUSTED THE CARTO 3 SYSTEM. BASED ON THE BWI REPRESENTATIVES REPORT, THE REBOOT OF THE PIU RESOLVED THE ISSUE. FIELD SERVICE ENGINEER (FSE) WENT ON-SITE TO EVALUATE THE SYSTEM. FSE COMPLETED FULL FUNCTIONAL TEST OF THE SYSTEM. SYSTEM PASSED ALL TESTS. ISSUE WAS NOT DUPLICATED. FSE RECOMMENDED USING NEW CATHETERS AND CATHETER CABLES AS OPPOSED TO REPROCESSED CATHETERS. SYSTEM IS READY FOR USE. DEVICE HISTORY RECORD REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. CONCOMITANT PRODUCTS: EZ STEER THERMOCOOL SF NAV CATHETER (MODEL# D-1313-04-S, LOT# 17271444L). C3 INTERFACE CABLE ¿ THERAPEUTIC (MODEL# D-1286-03-S LOT# 105039. (B)(4).><IT WAS REPORTED THAT A PATIENT UNDERWENT A RIGHT SIDED ATRIAL FLUTTER PROCEDURE WITH A CARTO 3 SYSTEM AND TWO LASSO 2515 NAV ECO VARIABLE CATHETERS A PROCEDURE CANCELLATION OCCURRED. DURING THE PROCEDURE, THE LASSO CATHETER DISPLAYED A MAGNETIC SENSOR ERROR. THE CATHETER WAS CHANGED AND THE SECOND LASSO CATHETER LOOP WOULD NOT DISPLAY THEREFORE THE CATHETER COULD NOT BE USED FOR FAST ANATOMICAL MAPPING. ERROR 371 MAPPING ELECTRODES NOT AVAILABLE" WAS SEEN. MORE THAN 30 MINUTES OF TROUBLESHOOTING WAS DONE | 1
5086371 | THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4).><IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A CARTO 3 SYSTEM AND A PROCEDURE DELAY OCCURRED. THE CARTO 3 SYSTEM WAS DISPLAYING ERROR 17 LOCATION PAD CURRENTS UNSTABLE. THEY REBOOTED THE PATIENT INTERFACE UNIT (PIU) WITHOUT RESOLUTION. THEY TURNED OFF THE PIU AND DISCONNECTED AND RECONNECTED THE LOCATION PAD (LP) EXTENSION CABLE FROM THE BACK OF THE PIU AND THE SAME ON THE LP END OF THE EXTENSION CABLE. THEY POWERED ON THE PIU AND ERROR 17 RETURNED. THEY THEN TURNED OFF THE PIU AND CONNECTED THE LP DIRECTLY TO THE PIU. WHEN THE PIU BOOTED UP, THE ERROR 17 RETURNED. THEY REBOOTED THE PIU AND THE WORKSTATION. THE PIU WAS CYCLING FROM 1, 2, AND THEN BACK TO 1. THE PATIENT WAS UNDER GENERAL ANESTHESIA FOR APPROXIMATELY 30 MINUTES. A TRANSSEPTAL PUNCTURE HAD BEEN PERFORMED. THE USE OF THE CARTO 3 SYSTEM WAS ABORTED AS THE CARTO 3 SYSTEM ISSUES COULD NOT BE RESOLVED. HOWEVER, THE PROCEDURE WAS CONTINUED WITH THE USE OF CRYO. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCE. THE PROCEDURE DELAY HAS BEEN ASSESSED AS A REPORTABLE MALFUNCTION BASED ON THE FACTS THAT THE PATIENT WAS UNDER GENERAL ANESTHESIA, THERE WAS AS TRANSSEPTAL PUNCTURE, AND THERE WAS ASSESSED PHYSICIAN RISK. ON ATTEMPTING TO CLARIFY THE PHYSICIAN RISK OPINION, INFORMATION WAS PROVIDED WHICH INDICATES THAT THE ASSESSED RISK MAY BE DUE TO A CUSTOMER PREFERENCE REGARDING THE LOCATION PAD. AS SUCH, THIS ASSESSMENT IS CONSERVATIVE.><(B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION (AFIB) PROCEDURE WITH A CARTO 3 SYSTEM AND A PROCEDURE DELAY OCCURRED. THE CARTO 3 SYSTEM WAS DISPLAYING ERROR 17 LOCATION PAD CURRENTS UNSTABLE. THEY REBOOTED THE PATIENT INTERFACE UNIT (PIU) WITHOUT RESOLUTION. THEY TURNED OFF THE PIU AND DISCONNECTED AND RECONNECTED THE LOCATION PAD (LP) EXTENSION CABLE FROM THE BACK OF THE PIU AND THE SAME ON THE LP END OF THE EXTENSION CABLE. THEY POWERED ON THE PIU AND ERROR 17 RETURNED. THEY THEN TURNED OFF THE PIU AND CONNECTED THE LP DIRECTLY TO THE PIU. WHEN THE PIU BOOTED UP, THE ERROR 17 RETURNED. THEY REBOOTED THE PIU AND THE WORKSTATION. THE PIU WAS CYCLING FROM 1, 2, AND THEN BACK TO 1. THE USE OF THE CARTO 3 SYSTEM WAS ABORTED AS THE CARTO 3 SYSTEM ISSUES COULD NOT BE RESOLVED. HOWEVER, THE PROCEDURE WAS CONTINUED WITH THE USE OF CRYO. THE BIOSENSE WEBSTER FIELD SERVICE ENGINEER ARRIVED TO THE ACCOUNT AND WAS ABLE TO DUPLICATE THE REPORTED COMPLAINT. THE BIOSENSE WEBSTER FIELD SERVICE ENGINEER DISCOVERED THAT THE FAULTY LP EXTENSION CABLE CAUSED THE REPORTED ISSUE. THE BIOSENSE WEBSTER FIELD SERVICE ENGINEER PLACED AN ORDER FOR THE LP EXTENSION CABLE. THE BIOSENSE WEBSTER FIELD SERVICE ENGINEER FOLLOWED UP ON THE ISSUE WITH THE BIOSENSE WEBSTER FIELD REPRESENTATIVE AND CONFIRMED THAT THE REPLACEMENT LP EXTENSION CABLE WAS DELIVERED TO THE ACCOUNT AND RESOLVED THE ISSUE. THE SYSTEM IS READY FOR USE. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. | 1
4402695 | EVALUATION SUMMARY: (B)(4). IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION PROCEDURE USING A CARTO 3 SYSTEM, A MAP SHIFT ISSUE AROSE DUE TO METAL INTERFERENCE. THE BWI FAILURE ANALYSIS LAB (FAL) ANALYZED THE COMPLAINT DEVICE. IT WAS DETERMINED THAT THE MAP SHIFT ISSUE OCCURRED DUE TO THE USER USING THE WRONG WORKFLOW. THE ISSUE COULD NOT BE DUPLICATED AFTER THE USER WAS ADVISED ON HOW TO PROPERLY USE THE SYSTEM. THE BWI SERVICE ENGINEERS ALSO DISCOVERED THAT THE METAL INTERFERENCE WAS CAUSED BY A NEAR BY XRAY MACHINE.><A DHR REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION PROCEDURE USING A CARTO 3 SYSTEM, A MAP SHIFT ISSUE AROSE. THE SKIN CONTACT PATCHES COULD NOT BE VISUALIZED CORRECTLY DURING THE PROCEDURE. ADDITIONAL INVESTIGATION DETERMINED THAT THE MAP SHIFT OCCURRED DUE TO METAL INTERFERENCE AND THERE WAS NO IMMEDIATE SYSTEM ERROR MESSAGE, MAKING THE EVENT MDR REPORTABLE. THE USER WAS ABLE TO COMPLETE THE PROCEDURE AND NO ADVERSE PATIENT CONSEQUENCES WERE REPORTED. THE HARDWARE INVESTIGATION HAS BEGUN, BUT IS NOT COMPLETED AT THIS TIME.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. CONTACT OFFICE AND MANUFACTURING SITE SHOULD REFLECT: (B)(6). CONCOMITANT PRODUCTS: UNK. (B)(4). | 1
5016650 | THE FDA SHUT DOWN THEIR EMDR SERVER ALONG AUGUST 6TH ¿ AUGUST 10TH FOR MAINTENANCE AND THE 3500A CODES WERE UPDATED. WHILE BWI COMPLAINTS SYSTEM IS UPDATED, THE MAPPING WILL BE DONE MANUALLY: (B)(4) IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION PROCEDURE WITH A CARTO® 3 SYSTEM AND AT THE MIDDLE OF THE CASE A MAP SHIFT WITHOUT ERROR MESSAGE WAS EXPERIENCED. THE ISSUE WAS RESOLVED BY CREATING A NEW MAP. FOLLOW UP INVESTIGATION WAS PERFORMED TO CLARIFY THE EVENT AND IT WAS CONFIRMED THAT THE ISSUE OCCURRED WHILE MAPPING AND THAT THE MAP SHIFTED ABOUT ONE OR TWO INCHES. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THIS EVENT IS BEING REPORTED BECAUSE SUCH MAP SHIFTS POTENTIALLY CAUSED BY A SYSTEM MALFUNCTION, COULD MISLEAD THE PHYSICIAN AND POSSESSES A POTENTIAL RISK TO THE PATIENT. THE FIELD SERVICE ENGINEER (FSE) CONTACTED BWI REPRESENTATIVE REGARDING THE ISSUE. IT WAS DISCOVERED THAT THE CUSTOMER USED RE-STERILIZED LASSO NAV CATHETER (REPROCESSOR: STRYKER). FSE WAS ALSO INFORMED THAT THE PATIENT WAS UNDER GENERAL ANESTHESIA AND THERE WAS NO CARDIOVERSION. FSE SUGGESTED SENDING DATA FILES TO DEVICE MANUFACTURER (HTC) FOR ANALYSIS. THE BWI REPRESENTATIVE STATED THAT THE SPECIFIC COMPLAINT WAS RELATED TO A RE-STERILIZED CATHETER AND NOT RELATED TO THE SYSTEM. THE BWI REPRESENTATIVE EXPLAINED THAT THE ISSUE HAS NOT HAPPENED SINCE ORIGINAL COMPLAINT WAS PLACED AND CUSTOMER ALSO HAS NOT USED THE RE-STERILIZED LASSO CATHETERS SINCE. NO SERVICE REQUESTED. SYSTEM IS OPERATIONAL. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. THE CUSTOMER COMPLAINT WAS NOT CONFIRMED. MANAGEMENT IS NOTIFIED OF FAILURE ANALYSIS THROUGH THE MONTHLY TRENDING REPORTS. NO SIGNIFICANT TRENDS HAVE BEEN IDENTIFIED AT THIS TIME; THEREFORE NO CAPA ACTIVITY IS REQUIRED.><A) THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED.(B)(4).><IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION PROCEDURE WITH A CARTO® 3 SYSTEM AND AT THE MIDDLE OF THE CASE A MAP SHIFT WITHOUT ERROR MESSAGE WAS EXPERIENCED. THE ISSUE WAS RESOLVED BY CREATING A NEW MAP. FOLLOW UP INVESTIGATION WAS PERFORMED TO CLARIFY THE EVENT AND IT WAS CONFIRMED THAT THE ISSUE OCCURRED WHILE MAPPING AND THAT THE MAP SHIFTED ABOUT ONE OR TWO INCHES. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THIS EVENT IS BEING REPORTED BECAUSE SUCH MAP SHIFTS POTENTIALLY CAUSED BY A SYSTEM MALFUNCTION, COULD MISLEAD THE PHYSICIAN AND POSSESSES A POTENTIAL RISK TO THE PATIENT. | 1
4670203 | THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4).><(B)(4). IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO® 3 SYSTEM AND A SYSTEM PERFORMANCE ISSUE OCCURRED WHEN TRYING TO ACQUIRE A POINT WITH THE FOOT PEDAL AFTER STARTING ABLATION, THE VISITAG POINT DOESN¿T SHOW WITH AN ABLATION TAG. THE ISSUE WAS REPORTED AFTER THE CASE AND NO PATIENT CONSEQUENCE WAS NOTED. THE VISITAG INVESTIGATION IS IN PROCESS TO DETERMINE PATIENT RISK WHEN NOT FUNCTIONING PROPERLY; HOWEVER THAT INVESTIGATION IS NOT COMPLETE. THEREFORE, IN TAKING A CONSERVATIVE APPROACH THIS EVENT HAS BEEN ASSESSED AS REPORTABLE SINCE THIS ISSUE MAY POTENTIALLY LEAD TO INADVERTENT PROLONGED ABLATION IN SOME AREAS WHICH MIGHT RESULT IN THE PATIENT INJURY. THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED. FSE CLARIFIED THAT THE ISSUE WAS REPORTED AFTER THE CASE SO NO LIVE TROUBLESHOOTING WAS PERFORMED. FSE ASKED THE CALLER CAS TO PROVIDE SOME MORE INFORMATION ON THE ISSUE. THE CAS CONTACTED FSE AFTER THE NEXT PROCEDURE AND CONFIRMED THAT THE SYSTEM IS WORKING PROPERLY AFTER REBOOT. THE ISSUE WAS NOT DUPLICATED. SYSTEM IS READY FOR USE. THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH CARTO 3 SYSTEM # 10244 WAS REVIEWED. THERE WAS NO ANY ADDITIONAL COMPLAINT OUT OF 46 ADDITIONAL REPORTED COMPLAINTS THAT MAY BE RELATED TO THE REPORTED ISSUE. A DHR REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO 3 SYSTEM AND A SYSTEM PERFORMANCE ISSUE OCCURRED WHEN TRYING TO ACQUIRE A POINT WITH THE FOOT PEDAL AFTER STARTING ABLATION, THE VISITAG POINT DOESN¿T SHOW WITH AN ABLATION TAG. THE ISSUE WAS REPORTED AFTER THE CASE AND NO PATIENT CONSEQUENCE WAS NOTED. THE VISITAG INVESTIGATION IS IN PROCESS TO DETERMINE PATIENT RISK WHEN NOT FUNCTIONING PROPERLY; HOWEVER THAT INVESTIGATION IS NOT COMPLETE. THEREFORE, IN TAKING A CONSERVATIVE APPROACH THIS EVENT HAS BEEN ASSESSED AS REPORTABLE SINCE THIS ISSUE MAY POTENTIALLY LEAD TO INADVERTENT PROLONGED ABLATION IN SOME AREAS WHICH MIGHT RESULT IN THE PATIENT INJURY. | 1
4813878 | THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. BWI CONCOMITANT PRODUCT USED: PRODUCT NAME: THERMOCOOL SMART TOUCH BI-DIRECTIONAL NAVIGATION CATHETER, US CATALOG #: D132705. (B)(4).><(B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION PROCEDURE WITH A CARTO® 3 SYSTEM AND THE CUSTOMER COULD NOT PACE. THE ECG CARD, EP OUT CARD AND DC/DC CARD WERE INITIALLY REPLACED TO DETERMINE IF THE PROBLEM WAS WITH THE CARTO SYSTEM OR THE ENVIRONMENT. REPLACING PARTS DID NOT RESOLVE THE ISSUE SO THE ORIGINAL PARTS WERE RE-INSTALLED. DURING INVESTIGATION THE SYSTEM PASSED FULL ATP. MEDIFIX CONFIRMED THAT THE SYSTEM IS OPERATIONAL AND THE ISSUE HASN¿T BEEN DUPLICATED. DEVICE HISTORY RECORD REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION PROCEDURE WITH A CARTO 3 SYSTEM AND CUSTOMER COULD NOT PACE. FOLLOW UP INVESTIGATION WAS PERFORMED TO CLARIFY EVENT AND IT WAS NOTICED THAT PATIENT WAS IN ATRIAL FIBRILLATION (PREEXISTING CONDITION). HOWEVER, DURING CAVOTRICUSPID ISTHMUS LINE PATIENT WENT TO JUNCTIONAL RHYTHM. THEREFORE, PACING WAS ATTEMPTED BUT NO CAPTURED WAS NOTED. ISSUE WAS TROUBLESHOT BY PACING WITH PACING LEADS TO STIM INPUTS CONNECTED ON PIU AND ALSO FROM STIM INPUTS TO DIRECT PACING PORTS, WHICH DID NOT RESOLVED THE PROBLEM. PACING WAS ACHIEVED BY BYPASSING CARTO SYSTEM. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THIS EVENT IS BEING REPORTED BECAUSE DIRECT PACING PORT FAILURE COULD POTENTIALLY DELAY PATIENT TREATMENT. | 1
5162915 | THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(6). (B)(4).><IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION PROCEDURE WITH A CARTO 3 SYSTEM AND A MAP SHIFT OCCURRED WITH NO WARNING. WHILE MAPPING WITH A LASSO CATHETER, THERE WERE INACCURACIES, HOWEVER THE CARTO 3 SYSTEM DID NOT INDICATE TO THE USER THAT A MAP SHIFT OCCURRED. THE PULMONARY VEIN ISOLATION PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. MULTIPLE ATTEMPTS HAVE BEEN MADE TO OBTAIN CLARIFICATION IN THIS COMPLAINT. HOWEVER, NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE. THEREFORE THIS EVENT IS BEING CONSERVATIVELY REPORTED SINCE THE SYSTEM DID NOT ALARM FOR A MAP SHIFT AS THIS COULD POTENTIALLY HARM THE PATIENT.><DUE TO THE AUGUST 2015 FDA MAINTENANCE WHERE THE 3500A CODES WERE UPDATED, THE 3500A CODES (EVALUATION CODES) WILL BE ADDED UNTIL THE BIOSENSE WEBSTER SYSTEM IS ALSO UPDATED. THEREFORE THE FOLLOWING CODES APPLY: (B)(4). MANUFACTURER'S REF. NO: (B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FIBRILLATION PROCEDURE WITH A CARTO 3 SYSTEM AND A MAP SHIFT OCCURRED WITH NO WARNING. BWI REPRESENTATIVE INFORMED THE FIELD SERVICE ENGINEER (FSE) THAT DURING THE PULMONARY VEIN ISOLATION PROCEDURE THE MAPPING CATHETER LASSO WAS SHOWN ON ONE POSITION ON THE CARTO 3 SYSTEM AND ON ANOTHER POSITION ON FLUOROSCOPY. SINCE THE PROCEDURE WAS FINISHED, TROUBLESHOOTING WAS NOT POSSIBLE TO PERFORM. FSE FOLLOWED UP THE ISSUE WITH BWI REPRESENTATIVE AND THE BWI REPRESENTATIVE INFORMED THAT HE HAD TWO CASE SUPPORTS IN THE LAB WITHOUT ANY PROBLEM. ISSUE WAS NOT DUPLICATED. SYSTEM IS OPERATIONAL. THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH THIS CARTO 3 SYSTEM # 11742 WAS REVIEWED. 1 OUT OF 18 ADDITIONAL REPORTED COMPLAINTS MAY BE RELATED TO THE REPORTED ISSUE. DEVICE HISTORY RECORD REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. | 1
4891568 | (B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA PROCEDURE WITH A CARTO 3 SYSTEM AND A MAP SHIFT WITH NO ERROR MESSAGE OCCURRED. FIELD SERVICE ENGINEER SPOKE TO THE BWI REPRESENTATIVE ABOUT REPORTED ISSUE. THE BWI REPRESENTATIVE EXPLAINED THAT THE COMPLAINT WAS CALLED IN FOR DOCUMENTATION ONLY. THE FACILITY DID NOT WANT ONSITE SERVICE AS THEY HAVE HAD SEVERAL CASES SINCE REPORTED ISSUE WITHOUT THE MAPPING ISSUES. THE ISSUE WAS NOT CONFIRMED. THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH THIS CARTO 3 SYSTEM WAS REVIEWED. OUT OF (B)(4) ADDITIONAL REPORTED COMPLAINTS THERE WERE (B)(4) ADDITIONAL COMPLAINTS THAT MAY BE RELATED TO THE REPORTED ISSUE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4)><IT WAS REPORTED THAT A PATIENT UNDERWENT AN IDIOPATHIC VENTRICULAR TACHYCARDIA PROCEDURE WITH A CARTO 3 SYSTEM AND A MAP SHIFT WITH NO ERROR MESSAGE OCCURRED. DURING MAPPING AND ABLATING, THE MAP HAD AN UNEXPLAINED SHIFT DURING THE CASE WITH NO ERRORS. METAL VALUES WERE LESS THAN TWO ON THE BACK PATCHES AND THE CHEST PATCHES READ 10.3 AND 10.9. NO DEFIBRILLATION WAS PERFORMED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY WITH THE ISSUE. THERE WAS NO PATIENT CONSEQUENCE. UPON REQUEST ADDITIONAL INFORMATION WAS RECEIVED ON THE EVENT. THE ABLATION CATHETER WAS COMPLETELY OUTSIDE THE FAST ANATOMICAL MAP. THE ORIGINAL LOCATION COMPLETELY SHIFTED. THE APPROXIMATE DIFFERENCE IN CATHETER LOCATION BEFORE AND AFTER MAP SHIFT WAS ABOUT ONE INCH. MAP SHIFTS WITH NO ERROR MESSAGE COULD POTENTIALLY CAUSE A RISK TO PATIENT. | 1
4936393 | IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION ABLATION PROCEDURE WITH A CARTO® SYSTEM AND EXPERIENCED A MAP SHIFT WITHOUT SYSTEM DISPLAYING AN ERROR MESSAGE. THIS WAS NOTICED SINCE THE RIGHT SUPERIOR PULMONARY VEIN WAS NOT AT THE SAME PLACE AS ON THE BEGINNING OF THE PROCEDURE. PHYSICIAN USED FLUOROSCOPY TO CONFIRM CATHETER'S POSITION. FURTHER INVESTIGATION WAS PERFORMED AND IT INFORMED THAT THE MAP SHIFT ISSUE WAS SEEN DURING MAPPING AND ABLATING, AS WELL AS THE MAP MOVEMENT WAS APPROXIMATELY FROM 5 TO 10 MILLIMETERS. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THIS EVENT IS BEING REPORTED BECAUSE SUCH MAP SHIFTS COULD POTENTIALLY MISLEAD PHYSICIAN AND CONTRIBUTE IN A POTENTIAL RISK TO PATIENT.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. PRODUCT NAME: PENTARAY NAV F CURVE US CATALOG #: UNKNOWN LOT #: UNKNOWN PRODUCT NAME: SMARTTOUCH DF CURVE US CATALOG #: UNKNOWN LOT #: UNKNOWN PRODUCT NAME: BW DECAPOLAR DF CURVE US CATALOG #: UNKNOWN LOT #: UNKNOWN (B)(4).><(B)(4). IT WAS REPORTED THAT A PATIENT UNDERWENT A PULMONARY VEIN ISOLATION ABLATION PROCEDURE WITH A CARTO® 3 SYSTEM AND EXPERIENCED A MAP SHIFT WITHOUT SYSTEM DISPLAYING AN ERROR MESSAGE. THIS WAS NOTICED SINCE THE RIGHT SUPERIOR PULMONARY VEIN WAS NOT AT THE SAME PLACE AS ON THE BEGINNING OF THE PROCEDURE. PHYSICIAN USED FLUOROSCOPY TO CONFIRM CATHETER¿S POSITION. FURTHER INVESTIGATION WAS PERFORMED AND IT INFORMED THAT THE MAP SHIFT ISSUE WAS SEEN DURING MAPPING AND ABLATING, AS WELL AS THE MAP MOVEMENT WAS APPROXIMATELY FROM 5 TO 10 MILLIMETERS. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THIS EVENT IS BEING REPORTED BECAUSE SUCH MAP SHIFTS COULD POTENTIALLY MISLEAD PHYSICIAN AND CONTRIBUTE IN A POTENTIAL RISK TO PATIENT. THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED. FSE CONFIRMED THAT NO SERVICE ACTIVITY WAS REQUIRED FOR THIS COMPLAINT. FSE CONTACTED REPS INVOLVED IN SUPPORTING THIS CUSTOMER, FOLLOWING UP THE ISSUE, AND WAS INFORMED THE NEXT TWO CASES REPORTED MAP SHIFT ISSUE WAS NOT REPEATED. SYSTEM IS OPERATIONAL. THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH CARTO 3 SYSTEM # (B)(4) WAS REVIEWED. THERE WAS NO ANY ADDITIONAL COMPLAINT OUT OF (B)(4) ADDITIONAL REPORTED COMPLAINTS THAT MAY BE RELATED TO THE REPORTED ISSUE. DHR REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. | 1
4648829 | PT PRESENTED TO ED ON (B)(6) 2014 WITH SEVERE SHORTNESS OF BREATH. PT WAS ADMITTED WITH CHF EXACERBATION AND HYPERTENSION ON (B)(6) 2014. PLACED ON CAPA AND GIVEN DIURETICS. PT WAS IMPROVING AND EXPECTED TO BE DISCHARGED ON (B)(6) 2014. DURING THE NIGHT THE CARDIAC RHYTHM CHANGED (BEGAN TO SLOW LEADING TO A CARDIOPULMONARY ARREST. RESUSCITATION WAS NOT SUCCESSFUL AND THE PT TIME OF DEATH WAS (B)(6) 2014 AT 00:33. | 0
5230409 | OMRON HEALTHCARE, INC. RECEIVED A LETTER REGARDING A CLAIM FROM (B)(6) ON (B)(6) 2015. IN THE LETTER, IT WAS REPORTED THAT THE INSURED SUSTAINED SEVERE DAMAGES TO THEIR HOME AS A RESULT OF THE FIRE. THE ENGINEER THAT WAS CALLED IN TO INVESTIGATE THE CAUSE OF FIRE IDENTIFIED THE CAUSE OF THE FIRE AS BEING THE OMRON BLOOD PRESSURE MONITOR WHICH WAS THE ONLY ITEM PLUGGED INTO THE ELECTRICAL OUTLET IN THE MASTER BEDROOM IN THE AREA OF THE ORIGIN OF THE FIRE. THE ENGINEER CONCLUDED: BASED UPON OUR INVESTIGATION TO DATE, WE HAVE FORMED THE FOLLOWING OPINIONS: THE AREA OF FIRE ORIGIN WAS AT THE OMRON BLOOD PRESSURE MONITOR LOCATED IN THE AREA OF FIRE ORIGIN; THE PROBABLE SOURCE OF HEAT FOR IGNITION WAS AN ELECTRICAL FAILURE OF THE POWER CORD FOR THE OMRON BLOOD PRESSURE MONITOR. ADDITIONAL EXAMINATION IS SUGGESTED TO VERIFY; THE PROBABLE FIRST FUEL IGNITED WAS THE POWER CORD INSULATION AND COMBUSTIBLES ON THE WOODEN DESK INCLUDING PAPER/DOCUMENTS; THE FIRE IS CLASSIFIED AS ACCIDENTAL AT THIS TIME BASED ON THE REMAINING PHYSICAL EVIDENCE AND REPORTED INFORMATION; NO OTHER COMPETENT SOURCES OF HEAT FOR IGNITION WERE IDENTIFIED IN THE AREA OF FIRE ORIGIN. ADDITIONAL DESTRUCTIVE EXAMINATION OF THE UNIT IS AN OPTION AT YOUR DISCRETION AFTER THE MANUFACTURER IS PLACED ON NOTICE. AT THIS TIME, OUR CLAIM IS STILL ONGOING. ADDITIONAL COMMUNICATION FROM THE (B)(6) REPRESENTATIVE (DATED 11/02/2015) STATED THEY ARE UNABLE TO PROVIDE COMPLETE ENGINEERING REPORT AT THIS TIME. AND, THEY WERE NOT IN A POSITION TO FORWARD THE EVIDENCE TO OMRON. THEY ALSO STATED, IF OMRON WISHES TO EXAMINE THE EVIDENCE, OMRON MAY HAVE OUR EXPERT CONTACT THEIR ENGINEER TO MAKE ARRANGEMENTS FOR JOINT EXAMINATION. | 0
5008666 | ><THE REPORTING FACILITY ALLEGED THAT THE ECHO IMAGES FOR PATIENT A WERE INADVERTENTLY SENT TO THE FOLDER OF PATIENT B, WHOSE FOLDER SUBSEQUENTLY CONTAINED THE WRONG ECHO IMAGES. THE ISSUE OCCURRED AT 4:20 PM. AS A RESULT OF THIS ISSUE, PATIENT B WAS TRANSFERRED TO THE CARDIAC CARE UNIT (CCU) AT APPROXIMATELY 6:30 PM, AFTER THE DEPARTMENT RECEIVED THE WRONG ECHO IMAGES. THE ISSUE WAS DISCOVERED WHEN THE REPORTING FACILITY STAFF REVIEWED PATIENT B'S RECORDS THAT EVENING. PATIENT A'S ECHO IMAGES HAVE NOT BEEN AFFECTED, SO THERE ARE NO CONSEQUENCES TO PATIENT A AS A RESULT OF THIS ISSUE AND THERE HAS BEEN NO REPORT OF INJURY OR HARM TO THE PATIENT B. | 1
4506215 | WHILE USING MEGASUTURE CUT NEEDLE DRIVERS TO SEW, SURGEON NOTICED A WIRE STICKING OUT OF THE INSTRUMENT. REMOVED INSTRUMENT FROM SURGICAL FIELD; NO HARM TO PATIENT. | 0
5041363 | THE LOGIQ E9 WAS BEING USED FOR AN OB/GYN STUDY WHERE MULTIPLE LENGTH/HEIGHT/WIDTH MEASUREMENTS WERE TAKEN TO CREATE A VOLUME CALCULATION FOR AN OVARY. THE USER REPORTED DELETING THE LENGTH/HEIGHT/WIDTH MEASUREMENT VALUES AND NOTICED THE VOLUME DID NOT DELETE AS WELL. THE USER THEN CREATED MORE LENGTH/HEIGHT/WIDTH MEASUREMENTS AND NOTICED THAT THE VOLUME KEPT INCREASING UNEXPECTEDLY. HOWEVER, THE INFORMATION WAS DISPLAYED CORRECTLY WITHIN THE WORKSHEET. THE ISSUE WAS OBVIOUS TO THE USER AND THERE WAS NO ADVERSE PATIENT IMPACT REPORTED. AN INVESTIGATION BEGAN AND IS STILL IN PROGRESS. UPON FURTHER REVIEW BY THE DESIGN TEAM, IT HAS BEEN DETERMINED THAT DELETED LENGTH/HEIGHT/WIDTH MEASUREMENTS USED TO GENERATE THE AVERAGE VOLUME CALCULATION MAY STILL BE INCORRECTLY USED AS AN INPUT VALUE BY THE SYSTEM FOR THE AVERAGE VOLUME CALCULATION. THE IMPACT OF THIS IS AS FOLLOWS. IF THE ABOVE SCENARIO EXISTS AND THE USER RE-OPENS A CLOSED STUDY, THERE MAY BE A POTENTIAL TO MISREAD THE WORKSHEET WHICH CONTAINS BOTH LENGTH/HEIGHT/WIDTH MEASUREMENTS AND THE CALCULATED VOLUMES, AND SELECT AN INCORRECT VOLUME FOR DIAGNOSIS BECAUSE THE SYSTEM HAS THE POTENTIAL TO INCORRECTLY STORE THE WRONG VOLUME CALCULATION PRIOR TO CLOSING THE ORIGINAL STUDY.><GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED AFTER THE INVESTIGATION IS COMPLETED. DATE OF DEVICE MANUFACTURE UNKNOWN.><GE HEALTHCARE¿S INVESTIGATION CONCLUDED THAT A SOFTWARE ISSUE CAUSED MEASURED VALUES FOR OB EXAM TYPES WHERE MULTIPLE MEASUREMENTS WERE TAKEN AND A SUBSET OF MEASUREMENTS ARE DELETED TO DISPLAY INCORRECTLY, THEREFORE CAUSING THE MISMATCH BETWEEN DISPLAYED MEASURED VALUES AND THE RESULTING CALCULATIONS; HOWEVER, THE CALCULATIONS ARE CORRECT IN EVERY CIRCUMSTANCE. THE CLINICAL EVALUATION OF THE ISSUE CONCLUDED THAT THE USER WOULD DETECT THE MISMATCH BETWEEN CALCULATED VALUES AND MEASURED VALUES AND WOULD CLEAR AND/OR DELETE ALL MEASUREMENTS, AND RE-TAKE THE MEASUREMENTS. ADDITIONALLY, IT IS CUSTOMARY FOR THE USER TO PERFORM ADDITIONAL DIAGNOSTIC IMAGING TO CONFIRM MEASUREMENTS AND CALCULATIONS. SHOULD THE ISSUE BE UNDETECTABLE, THERE COULD BE A POTENTIAL FOR A MINOR DELAY IN CARE IN THE CASE OF POLYCYSTIC OVARIAN DISEASE. NO FURTHER ACTION IS REQUIRED. | 1
4775011 | ><THE INITIAL REPORTER IS LOCATED OUTSIDE THE U.S., AND THEREFORE THIS INFORMATION IS NOT PROVIDED DUE TO COUNTRY PRIVACY LAWS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
5161871 | ABIOMED IMPELLA CIRCULATORY ASSIST DEVICE SHUT OFF WITHOUT WARNING/PRECIPITATING EVENT; DULL BEEP COULD BE HEARD BUT ONLY IF ACTUALLY IN ROOM AT BEDSIDE. NO ADVERSE PATIENT EVENT OCCURRED. A BIOMED REPRESENTATIVE WAS CONTACTED AND REMOVED DEVICE FROM SERVICE. | 0
5107971 | (B)(4).><INITIAL COMPLAINT FROM CUSTOMER WAS UNABLE TO IMPORT". UPON FURTHER TROUBLESHOOTING WITH CUSTOMER | 1
5337769 | .><THERE WERE THREE SEPARATE ISSUES WITH THIS MACHINE. THERE WAS A WARNING ON THE CARTO 3 ON THE PATCHES AND A SENSOR ERROR. NO TEMPERATURE BEING READ, SWITCHED OUT ABLATION CATHETER. ABLATION DATA NOT BEING CROSSED OVER TO CARTO GE CARDIOLAB. | 1
4608399 | ><IN THESE MONITORS, IF GRAPHIC TRENDS ARE DISPLAYED, THE GRAPH FAILS TO SHOW THE EVOLVING PATIENT COURSE AND ALLOW CLINICIANS TO BE MISLED AT TIMES OF PATIENT CHANGE. IT HAS BEEN OBSERVED THAT ONCE THE TRENDED DATA WINDOW IS OPENED, THE DATA DISPLAYED DOES NOT ALWAYS UPDATE AS TIME PROGRESS TO SHOW THE MOST RECENT DATA TREND. MANUFACTURER RESPONSE: AS OF THE FALL OF 2014, THE SOFTWARE FIX FOR THIS ISSUE IS NOT PLANNED FOR THE NEXT SOFTWARE RELEASE. | 1
4794457 | PATIENT ARRIVED FOR R/O LABOR AND DECREASED FETAL MOVEMENT. WHEN PATIENT WAS PUT ON THE CORO FOR FETAL TRACING, UNABLE TO DETECT HEART TONES, BUT COULD FEEL FETAL MOVEMENT. THERE WAS INTERMITTENT FETAL TRACING WITH MINIMAL VARIABILITY, WHICH WAS THEN UNDETECTABLE ON THE FETAL TRACING. VERIFIED FETAL HEARTBEAT AND BPP WITH ULTRASOUND, WHICH DID NOT CORRESPOND WITH CORO TRACING OF A FLAT LINE. TRANSFERRED PATIENT'S TRACING TO ROVER 1, WHICH CORRESPONDED WITH ULTRASOUND SHOWING MODERATE VARIABILITY. NOTE: TRANSDUCER INVOLVED IN THIS INCIDENT WAS NOT REMOVED FROM SERVICE AT THE TIME THE ISSUE OCCURRED. TRANSDUCER SERIAL NUMBER REDACTED MALFUNCTIONED ON THE FOLLOWING DAY, CREATING A SIMILAR EVENT WITH A DIFFERENT CORO MACHINE. | 1
5014149 | .><AN ERROR MESSAGE WAS DISPLAYED ON THE DA VINCI TOWER. THE VESSEL SEALER INSTRUMENT WAS NOT RECOGNIZED. THE INSTRUMENT DID NOT TOUCH THE PATIENT AND A NEW ONE WAS OPENED. | 0
5266164 | ON 2015 (B)(6) BIOPSY STAINED WITH HERCEPTEST LOT 20016932 AT (B)(6) GAVE RESULT HER2 SCORE 3+, STRONGLY POSITIVE. BASED UPON THE INITIAL HERCEPTEST RESULT, THE PATIENT WAS TREATED WITH HERCEPTIN. THE PATIENT DID NOT RESPOND TO TREATMENT AND ON 2015 (B)(6) PATIENT WAS RETESTED WITH SURGICAL TISSUE STAINED WITH HERCEPTEST LOT 20022527 AND RECEIVED HER2 NEGATIVE RESULTS. MOREOVER, ON 2015 (B)(6) SURGICAL TISSUE STAINED WITH ZYTOLIGHT SPEC HER2/CEN17 DUAL COLOR PROBE AT (B)(6) RESULT: NOT HER2 GENE AMPLIFIED. (DAKO COMPLAINT NUMBER (B)(4)). | 0
4974707 | AN OUTPATIENT WAS ORDERED TO HAVE AN ULTRASOUND GUIDED ASPIRATION OF A CHRONIC FLUID COLLECTION IN THE DEEP MUSCLES OF THE HIP AREA THAT WAS TRACKING VIA FISTULA CUTANEOUSLY. CULTURES AND CHEMISTRIES OF THE FLUID WERE ORDERED BY THE OUTPATIENT DOCTOR AND SENT BY THE RADIOLOGY TEAM TO THE LAB. TEN DAYS PASSED AND NONE OF THE RESULTS WERE RECEIVED BY THE ORDERING PHYSICIAN. IT WAS DISCOVERED THAT WHEN THE PT PRESENTED FOR F/U, THAT THE RESULTS WERE IN THE EHR MDDS DEVICE, BUT NO ONE KNEW. THE RADIOLOGIST WAS NOT EXPECTED TO SEARCH FOR THE RESULTS, NOR WERE THE RESULTS SENT TO THE RADIOLOGIST OF THE ORDERING PHYSICIAN. THE RESULTS WERE IN ONE OF THE SILOS IN THE EHR BUT THERE WAS NOT ANY WARNING, NOTIFICATION, OR RECONCILIATION THAT RESULTS WERE NOT RECEIVED BY THE ORDERING DOCTOR WHO DID NOT HAVE ACCESS TO THIS HOSPITAL'S EHR. THE ABSCESS GREW A VIRULENT ORGANISM. THE TREATMENT WAS DELAYED BY THE FAILURE OF THE DEVICE TO DELIVER THE RESULTS. | 1
4901363 | A ROBOTIC LARGE NEEDLE DRIVER JAWS WERE REPORTED TO NOT OPEN AND CLOSE PROPERLY DURING A CASE WITH USING DAVINCI ROBOTIC SURGICAL UNIT. THIS RESULTED IN THE NEEDLE BEING DROPPED WHILE INSIDE THE PATIENT. THE NEEDLE WAS FREE FLOATING IN THE PATIENT UNTIL IT WAS FOUND AND RECOVERED. UPON INVESTIGATION OF THE DEVICE CLINICAL ENGINEERING WAS ABLE TO MANUALLY MANIPULATE THE JAWS OF THE ROBOTIC ARM BY ROTATING THE COGS.======================MANUFACTURER RESPONSE FOR ROBOTIC LARGE NEEDLE DRIVER, DAVINCI ENDOWRIST INSTRUMENT (PER SITE REPORTER).======================WE CONTACT INTUITIVE SURGICAL QUARTERLY FOR REPORTS REGARDING OUR FAILED SURGICAL ARMS. THEY HAVE ALWAYS BEEN HELPFUL WITH THEIR FEEDBACK. | 0
4974708 | IHEALTH BG1 ALIGN MINI GLUCO-MONITORING SYSTEM BATTERY LIFE ONLY LASTS FOR TWO BLOOD SUGAR READINGS. TWO OF THESE DEVICES WERE PURCHASED FROM MULTIPLE (B)(6) LOCATIONS AND BOTH HAD THE SAME ISSUES. AT THIS RATE IT WOULD COST IN UPWARDS TO (B)(6) A DAY TO USE THIS PRODUCT DUE TO THE COST OF BATTERY REPLACEMENTS. AFTER A LITTLE RESEARCH, THE TESTS THAT WERE PERFORMED BY THE FDA STATES THAT THIS DEVICE LASTED FOR 200 READINGS ON A SINGLE BATTERY. THIS IS 100 TIMES MORE RELIABLE OF A PRODUCT THAN WE HAVE PURCHASED FROM OUR RETAILER. IT APPEARS THAT SOMETHING IS NOT WORKING CORRECTLY CAUSING BATTERY LIFE TO BE REDUCED DRAMATICALLY. I HAVE LISTED THE SERIAL NUMBER OF THE LAST DEVICE THAT WAS RETURNED TO (B)(6) FOR YOUR INFO. SERIAL NUMBER OF LAST DEVICE WE PURCHASED: (B)(4). | 0
4961012 | IT WAS REPORTED THAT: DESPITE THE UPPER SPO2 ALARM LIMIT WAS EXCEEDED | 1
4534574 | INITIALLY THERE WAS NO OUTPUT WHEN THE SURGEON ACTIVATED THE DEVICE. THEY DISCONNECTED AND RECONNECTED THE TOOL AND MACHINE CONNECTOR AND STILL NO OUTPUT WHEN ACTIVATED. THEY DISCONNECTED AND RECONNECTED A 2ND TIME AND WHEN THE SURGEON ACTIVATED THE ELECTROSURGICAL UNIT (ESU), THE CABLE BLEW APART AND THERE WAS A FIRE AT THE CONNECTOR. THE SURGEON WAS USING DAVINCI MONOPOLAR SCISSORS CONNECTED TO THE ESU VIA AN INTEGRA LIFE SCIENCES JARIT ACTIVE CABLE PART 600-290. INITIAL INSPECTION OF DEVICE SHOWS OBVIOUS BURNING OF THE CABLE AND CABLE IS IN 2 PIECES BUT NO OBVIOUS ISSUES WITH THE ESU DEVICE ITSELF. ESU DEVICE PASSED ALL BIOMED TESTING WITH NO FAILURES. THE MANUFACTURER REQUESTED THE UNIT FOR EVALUATION. | 0
4794456 | VESSEL SEALER DID NOT WORK DURING LAPAROSCOPIC PROCEDURE. | 0
4902287 | IRIS FIELD SERVICE ENGINEER EVALUATED THE INSTRUMENT AND CONFIRMED THE LEAK WAS COMING FROM THE A SMALL BRACK IN THE CGM TUBING. THE FSE REPLACED THE PIPETTE PARKING DEVICE TO RESOLVE THE LEAK. THE FSE RAN CONTROLS WHICH PASSED AND SYSTEM WAS OPERATIONAL. (B)(6).><THE CUSTOMER REPORTED A CONTAINED PROBE LEAK COMING FROM THE COLOR CLARITY AND SPECIFIC GRAVITY MODULE (CGM) TUBING NEAR THE SHORT SAMPLE SENSOR. THE CUSTOMER INDICATED QUALITY CONTROLS (QC) WERE NOT PASSING. THE CUSTOMER WAS WEARING PPE AT THE TIME OF DISCOVERY. THERE WAS NO EXPOSURE TO THE LEAK AND NO INJURIES. NO ERRONEOUS RESULTS WERE GENERATED OR REPORTED OUT OF THE LAB. | 0
4671125 | CUSTOMER REPORTED THEIR CONTROL IS FAILING FOR BILIRUBIN.><CUSTOMER REPORTED THE ICHEM VELOCITY FAILING CA CONTROL FOR BILIRUBIN. THERE WERE NO REPORTS OF PATIENT RESULTS BEING AFFECTED AND NO REPORTS OF CHANGE TO PATIENT MANAGEMENT AS A RESULT. AN IRIS FIELD SERVICE ENGINEER (FSE) OBSERVED THE PROBE WAS WORN AND THE ALIGNMENT OF THE PROBE WAS OFF. THE FSE REPLACED THE PROBE AND ADJUSTED THE ALIGNMENT. THE FSE RAN QC WHICH PASSED. SYSTEM WAS OPERATIONAL. | 0
5218655 | IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION AND FOUND THE TUBING BETWEEN THE SAMPLE FILTER AND PBV VALVE CRACKED DUE TO NORMAL WEAR. THE FSE REPLACED THE TUBING TO RESOLVE THE ISSUE. THE FSE RAN CONTROLS AND THEY PASSED. THE SYSTEM WAS OPERATIONAL. (B)(4).><THE CUSTOMER STATED THE TUBING BETWEEN THE SAMPLE FILTER AND THE PIPPETTER BYPASS VALVE (PBV) CRACKED AND WAS LEAKING. THE LEAK WAS CONTAINED TO THE INSTRUMENT AND THE CUSTOMER WAS WEARING THEIR PPE. THE CUSTOMER WAS NOT EXPOSED TO THE LEAK AND DID NOT SEEK MEDICAL ATTENTION. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB. | 0
5215477 | IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION. THE FSE FOUND AIR IN THE FLUIDICS FROM THE SYRINGE PUMP THROUGH TO THE COLOR CLARITY AND SPECIFIC GRAVITY MODULE (CGM). THE FSE REPLACED THE RINSE PUMP ASSEMBLY P/N: 700-7513 TO RESOLVE THE CONTROL FAILURE. THE FSE RERAN THE CONTROLS AND THE CONTROLS PASSED. THE SYSTEM WAS OPERATIONAL. (B)(4).><THE CUSTOMER REPORTED FAILING CA QUALITY CONTROLS FOR BILIRUBIN. THE CUSTOMER WAS GETTING A FALSE NEGATIVE RESULT AND TRIED USING FRESH CONTROLS, BUT STILLED FAILED. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB. | 0
5215475 | IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION. THE FSE DISCOVERED AIR IN THE FLUIDICS FROM THE RINSE PUMP ASSEMBLY TO THE SYRINGE. THE FSE REPLACED THE RINSE PUMP ASSEMBLY P/N: 700-7513 TO RESOLVE THE CONTROL FAILURE. THE FSE RERAN THE CONTROLS AND THE CONTROLS PASSED. THE SYSTEM WAS OPERATIONAL. (B)(4).><THE CUSTOMER REPORTED FAILING (B)(6) QUALITY CONTROLS FOR MULTIPLE ANALYTES. THE CUSTOMER PREFORMED A CLEANING PROCEDURE AND THEY WERE STILL FAILING CONTROLS. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB. | 0
5275348 | IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION. THE FSE CONFIRMED THE CA CONTROL FAILURE FOR BILIRUBIN. THE FSE REPLACED THE SYRINGE PUMP P/N: 700-7151S TO RESOLVE THE CUSTOMER ISSUE. THE FSE RERAN CONTROLS AND THE CONTROLS PASSED. (B)(4).><THE CUSTOMER REPORTED CA QUALITY CONTROLS FAILING FOR BILIRUBIN. THE CUSTOMER INDICATED THEY ARE GETTING FALSE NEGATIVE QUALITY CONTROL FAILURE. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB. | 0
4441596 | CUSTOMER REPORTED AN INTERMITTENT LACK OF SHORT SAMPLE ERROR FLAG WHEN EMPTY TUBES RAN ON IQ200.><CUSTOMER REPORTED INTERMITTENTLY SHORT SAMPLE DETECTOR FAILING TO DETECT SHORT SAMPLE WHEN EMPTY PATIENT SAMPLES WERE RUN. THERE WERE NO REPORTS OF CHANGE TO PATIENT MANAGEMENT AS A RESULT. | 0
5215450 | THE CUSTOMER CALLED IN TO REPORT A FOCUS FAILURE. THE CUSTOMER ALSO REPORTED THE POSITIVE CONTROL WOULD FAIL INTERMITTENTLY. THE CUSTOMER FLUSHED THE SAMPLE FILTER 3 TIMES AND THE CONTROLS STILL FAILED. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB.><IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION AND DISCOVERED THE VALVES WERE NOT OPENING WHEN FOCUS WAS RUN ON THE INSTRUMENT. THE FSE REPLACED THE EVACUATION BYPASS VALVE (EBV) P/N: 700-3880 AND THE DRAIN VALVE (DRV) P/N: 700-3882 TO RESOLVE THE ISSUE. THE FSE RERAN CONTROLS AND THE CONTROLS PASSED. THE SYSTEM WAS OPERATIONAL. (B)(4). | 0
5215459 | THE CUSTOMER CALLED IN TO REPORT A FOCUS FAILURE AND POSITIVE CONTROL FAILURE ON THEIR IQ200 INSTRUMENT. THE CUSTOMER WAS ABLE TO TROUBLESHOOT WITH CUSTOMER TECHNICAL SUPPORT (CTS) AND WAS INSTRUCTED TO REMOVE AND REPLACE THE SPECIMEN FILTER. THE CUSTOMER STILL FAILED CONTROLS. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB.><IRIS FIELD SERVICE ENGINEER WAS NOT SENT TO THE CUSTOMER LOCATION. ONSITE CUSTOMER BIOMEDICAL TECHNICIAN FOUND A SMALL LEAK IN THE TUBING BETWEEN THE FILTER AND THE PIPETTOR BYPASS VALVE (PBV). AFTER THE TUBING REPLACEMENT CALIBRATION FAILED. THE TECHNICIAN REPLACED THE EVACUATION BYPASS VALVE (EBV) P/N: 700-3880 TO RESOLVE THE ISSUE. THE FSE RERAN CONTROLS AND THE CONTROLS PASSED. THE SYSTEM WAS OPERATIONAL. (B)(4). | 0
5246937 | THE CUSTOMER REPORTED CA QUALITY CONTROLS FAILING FOR BILIRUBIN. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB.><IRIS FIELD SERVICE ENGINEER WAS SENT TO THE CUSTOMER LOCATION. THE FSE OBSERVED AIR IN THE FLUIDIC SYSTEM. THE FSE REPLACED THE FLUID BLOCK ASSEMBLY (FBA) P/N: 700-7509S) TO RESOLVE THE ISSUE. THE FSE RERAN CONTROLS AND THE CONTROLS PASSED. BEC INTERNAL IDENTIFIER FOR THIS REPORT IS (B)(6). | 0
5048106 | THE CUSTOMER REPORTED THEY ARE FAILING CA QUALITY CONTROL FOR BILIRUBIN. THE CUSTOMER OBSERVED THE BILIRUBIN PAD WAS NOT GETTING DOSED. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB.><IRIS FIELD SERVICE ENGINEER CONFIRMED THE CA CONTROL WAS FAILING FOR BILIRUBIN. THE FSE IDENTIFIED THE STRIP PROVIDER MODULE (SPM) WAS DELIVERING CHEMISTRY STRIPS ON TOP OF THE STRIP CONVEYOR SYSTEM (SCS) CAUSING THE PIPETTE TO RUN INTO THE STRIP. THE FSE ADJUSTED THE ALIGNMENT OF THE SPM TO RESOLVE THE ISSUE. THE FSE RERAN CONTROLS AND THE CONTROLS PASSED. BEC INTERNAL IDENTIFIER FOR THIS REPORT IS (B)(6). | 0
4997460 | IRIS FIELD SERVICE ENGINEER EVALUATED THE INSTRUMENT AND OBSERVED AIR IN TUBING BETWEEN THE COLOR CLARITY AND SPECIFIC GRAVITY MODULE (CGM) AND THE PROBE. THE FSE REPLACED THE TUBING BETWEEN THE CGM AND THE PORBE TO CORRECT THE ISSUE. THE FSE RAN CONTROLS WHICH PASSED AND SYSTEM WAS OPERATIONAL. (B)(4).><THE CUSTOMER REPORTED THEY ARE FAILING CA QUALITY CONTROL FOR BILIRUBIN. THERE WERE NO ERRONEOUS RESULTS GENERATED OR REPORTED OUT OF THE LAB. | 0
4798113 | THE CUSTOMER REPORTED THAT THE TUBE MOUNT (CONNECTED TO THE LOWER TELESCOPIC TUBE, CARRYING THE X-RAY TUBE, COLLIMATOR AND ALPHA MOTOR DRIVE) FELL DOWN BY ITS OWN. NOBODY WAS HURT.><THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA.><THE FOUR M8 FIXATION BOLTS OF THE TUBE MOUNT (CONNECTION TO THE LOWER TELESCOPIC TUBE, CARRYING THE X-RAY TUBE, COLLIMATOR AND ALPHA MOTOR DRIVE) WERE COMPLETELY LOOSE AND HAS CAME OFF OVER TIME RESULTED THAT THE TUBE ASSEMBLY FELL DOWN. INVESTIGATION SHOWED THAT THE THIRD PARTY INSTALLATION TEAM DID NOT APPLY A THREAD LOCKER (LOCTITE) AND DID NOT APPLY THE REQUIRED TORQUE OF 26NM TO THE BOLTS, WHICH BOTH ARE CLEARLY DESCRIBED IN THE INSTALLATION MANUAL AND WOULD ALSO NEED TO BE CONFIRMED ON A CHECKLIST BY THE INSTALLER. EACH OF THOSE WOULD HAVE PREVENTED THE BOLTS FROM COMING LOOSE. THE FIELD SERVICE ENGINEERS ADDED NEW M8 BOLTS, APPLIED LOCTITE 243 TO THE THREADS AND TIGHTENED THE BOLTS WITH A TORQUE OF 26NM ACCORDING TO THE SYSTEM MANUAL INSTALLATION INSTRUCTION. ALL SYSTEMS INSTALLED BY THE SAME TEAM OF THIRD COMPANY WERE INSPECTED WITH NO FAILURE FOUND. (B)(4). | 0
5254231 | (B)(4). SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><(B)(4). A BATCH REVIEW WAS CONDUCTED AND THERE WERE NO DEVIATIONS FOUND RELATED TO THIS REPORTED CONDITION DURING THE MANUFACTURE OF THIS LOT. THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><IT WAS REPORTED THAT A PATIENT HAD A LEAKING CASSETTE DURING PERITONEAL DIALYSIS THERAPY. THE PATIENT NOTICED FLUID ON THE FLOOR. THERE WAS NO PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. | 0
4466408 | WE USE ELEKTA'S MOSAIQ RADIATION ONCOLOGY RECORD AND VERIFY SYSTEM VERSION 2.41.01J0 IN OUR RADIATION ONCOLOGY CLINIC. WE RECENTLY FOUND IT NECESSARY TO RESEND REFERENCE CT AND STRUCTURE SET INFORMATION FOR AN IGRT PATIENT, WITH THE INTENT TO APPLY THE NEW REFERENCE DATA TO FUTURE IMAGING AND TREATMENT SESSIONS. TO OUR GREAT DISMAY, WE FOUND (AND AN ELEKTA REPRESENTATIVE CONFIRMED) THAT MOSAIQ ALSO APPLIES THE NEW REFERENCE DATA RETROACTIVELY TO PREVIOUS APPROVED CBCT IMAGES, SHOWING THE SAME SHIFTS WERE MADE BASED UPON DIFFERENT REFERENCE DATA. THIS EFFECTIVELY CHANGES THE RECORD OF WHAT WAS DONE AND WHAT WAS APPROVED, MAKING IT LOOK LIKE DIFFERENT (AND POSSIBLY COMPLETELY INAPPROPRIATE) LOCALIZATION DECISIONS WERE MADE. SINCE THIS HAPPENS RETROACTIVELY IT MAY NOT BE A PATIENT SAFETY CONCERN, BUT IT COULD BE A HUGE LIABILITY ISSUE FOR RADIATION ONCOLOGISTS AND RADIOTHERAPY CLINICS. IT IS DIFFICULT TO BELIEVE THE FDA WOULD GRANT APPROVAL TO A SYSTEM THAT ALLOWS RETROACTIVE CHANGES TO APPROVED RECORDS. | 1
4956410 | PATIENT WAS ISSUED A BOSCH A AND D MEDICAL DIGITAL BP MONITOR, MODEL NUMBER - UA-767PC, SERIAL NUMBER (B)(4) TO MONITOR HIS BP'S IN HIS HOME. THE PATIENT CALLED IN AND REPORTED THE UNIT WAS NOT FUNCTIONING AS IT APPARENTLY OVERHEATED AND MELTED TO THE DESK. THE PATIENT REPORTED HE WAS USING THE ISSUED SURGE PROTECTOR AND WAS ALSO USING A POWER CORD/AC ADAPTOR THAT WAS NOT ORDERED/ISSUED. THE DOOR TO THE BATTERY COMPARTMENT IS MELTED AND FUSED TO THE BATTERIES AND UNIT CASE. THE UNIT WAS SET UP AS A PERIPHERAL DEVICE FOR A HEALTH BUDDY, A AND D SMART CONNECTING BP MONITOR TO HEALTH BUDDY, BATTERIES INSTALLED, AND PATIENT ALSO HAD THE BP MONITOR PLUGGED INTO POWER. HE REPORTED TO THE CARE COORDINATOR THAT HE DID USE A SURGE PROTECTOR. | 0
4423357 | CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE (B)(4) LETTER. NOTE: THE DEVICE MANUFACTURER DATE FOR THE SUSPECT DEVICE IS UNKNOWN. THE DATE ENTERED IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 0
4839079 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><THIS IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATA AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC006 LETTER. THE DEVICE MANUFACTURER DATE FOR THE SUSPECT DEVICE IS UNKNOWN. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 0
4438539 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><(B)(6). THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE DEVICE MANUFACTURER DATE FOR THE SUSPECT DEVICE IS UNKNOWN. THE DATE ENTERED IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 0
5088649 | THE CUSTOMER OBSERVED THAT THE SAMPLE TUBE DID NOT HAVE THE CORRECT AMOUNT OF REAGENT WHEN THE SAMPLE WAS PREP ON THE TQ PREP INSTRUMENT. THE SAMPLE LINE HAD ASPIRATED AIR SINCE THE REAGENT BOTTLES WERE EMPTY. THERE WAS A FAILURE OF THE SENSORS. NO ERROR MESSAGE ALERTED THE CUSTOMER THAT THE REAGENT BOTTLES WERE EMPTY SINCE THE SENSORS HAD FAILED. THE FSE REPLACED BOTH REAGENT SENSORS TO RESOLVE THE ISSUE.(B)(4).><THE FSE (FIELD SERVICE ENGINEER) REPORTED THAT THE REAGENT SENSORS (REAGENT A AND REAGENT B SENSORS) IN A COULTER TQ-PREP WORKSTATION FAILED AND DID NOT GENERATE THE LOW REAGENTS MESSAGES. THE CUSTOMER PERFORMS THE DAILY INSTRUMENT CHECK AS PER INSTRUCTIONS FOR USE AND THE CUSTOMER VISUALLY CONFIRMS THE TUBES AFTER PREPARATION, WHICH ALERTED THE CUSTOMER TO THE ISSUE. ERRONEOUS RESULTS WERE NOT GENERATED AND THERE WAS NO CHANGE OR EFFECT TO PATIENT TREATMENT IN CONNECTION WITH THIS EVENT. | 0
4583046 | THE CARBOPLATIN DOSE DELIVERED TO THE PT WAS BASED ON THE CREATININE VALUE MANUALLY ENTERED AND APPROVED BY ARIA BY THE PHYSICIAN INSTEAD OF THE VALUE LATER RECEIVED BY ARIA FROM THE LAB INTERFACE. THE PHARMACIST WAS NOT AWARE OF THE INTERFACED LAB VALUES FOR CREATININE, BECAUSE THE FLOWSHEET BY DESIGN HAD DISPLAYED THE LAB VALUE WITH THE MOST RECENT COLLECTED DATE AND TIME, WHICH IN THIS CASE WAS NOT THE MANUALLY ENTERED VALUE. BASED ON THE PHARMACIST'S COMMENTS DURING AN INTERVIEW ON (B)(6) 2015, THE PHARMACIST WAS NOT AWARE OF HOW TO DIFFERENTIATE MANUALLY ENTERED RESULTS FROM LAB INTERFACED RESULTS AND MIGHT HAVE RESPONDED DIFFERENTLY BASED ON KNOWLEDGE OF THE SOURCE OF THE CREATININE VALUE. NEVERTHELESS, THE PHARMACIST APPROVED THE DISPENSING OF CARBOPLATIN ACCORDING TO THE PHYSICIAN'S MANUALLY ENTERED ORDER. THE PT THEREFORE RECEIVED A HIGHER DOSE OF CARBOPLATIN THAN WHAT WOULD HAVE BEEN CALCULATED USING THE INTERFACED LAB VALUE. THE COMPLAINT DESCRIPTION INCLUDED A COMMENT FROM CUSTOMER SUPPORT INDICATING THE FLOWSHEET DID NOT HAVE A GREEN BOX NEXT TO THE CREATININE TO INDICATE MULTIPLE VALUES. THIS ISSUE, PREVIOUSLY IDENTIFIED IN DISCREPANCY REPORT (B)(4), WAS ANALYZED AND DETERMINED NOT TO BE A CONTRIBUTING FACTOR IN THIS INCIDENT BECAUSE THE PHARMACIST DID NOT KNOW HOW TO EFFECTIVELY USE THE FLOWSHEET. DURING THE INTERVIEW WITH THE PHARMACIST, HER ONLY CONCERN WAS HOW TO DIFFERENTIATE MANUALLY ENTERED RESULTS FROM LAB INTERFACE RESULTS. IN ANALYZING THE COLLECTED DATE AND TIME, THE FLOWSHEET BY DESIGN DISPLAYED THE PHYSICIAN'S MANUALLY ENTERED ESTIMATED CREATININE VALUE AS THE MOST RECENT VALUE BECAUSE THE PHYSICIAN USED THE CURRENT DATE AND TIME AS THE COLLECTED DATE AND TIME FOR THE ESTIMATED CREATININE VALUE. THE FIRST ROOT CAUSE IS THAT THE PHYSICIAN APPROVED THE DRUG DOSING BASED ON ENTERING AN ESTIMATED CREATININE VALUE INSTEAD OF WAITING AND APPROVING THE DRUG DOSING BASED ON THE SAME DAY LAB VALUES. THE SECOND ROOT CAUSE IS THAT THE PHARMACIST WHO RE-CHECKS THE DRUG ORDERS FOR CORRECTNESS DID NOT KNOW HOW TO DIFFERENTIATE MANUALLY ENTERED RESULTS FROM LAB INTERFACED RESULTS. (B)(6). | 1
4785230 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) HAD AN INTERMITTENT DISPLAY SETTING FAILED" MESSAGE. HE TREIDD TO FIX IT BY REINSTALLING IN THE DRIVER AND ADJUSTING THE RESOLUTION. HOWEVER | 1
4385452 | THE CUSTOMER STATED THAT THE KEYBOARD, VIDEO AND MOUSE (KVM) STATION FOR THE CENTRAL MONITORING SYSTEM (CNS) WAS NOT WORKING. THE CUSTOMER TRIED POWER CYCLING THE KVM AND ALSO TRIED A DIFFERENT USB MOUSE AND KEYBOARD. THE KVM IS STILL NOT FUNCTIONAL. MFR REF #: 8030229-2014-00186. | 1
4752871 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) WAS FREEZING MULTIPLE TIMES A DAY. REFERENCE MFR # 8030229-2015-00068. | 1
4991348 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) WAS MONITORING PTS, BUT THE SECONDARY MONITOR WAS NOT DISPLAYING THE INFO. (B)(4). | 1
4536553 | THE CUSTOMER STATES THAT A PT EXPIRED WHILE BEING MONITORED ON THE BEDSIDE VITAL SIGN MONITOR (BSM-6500) AND THAT THE SYSTEM WAS TESTED AND PERFORMED AS EXPECTED AND THAT THERE IS NO EVIDENCE THAT THE MONITOR MALFUNCTIONED. THE CUSTOMER WOULD LIKE US TO REVIEW THE EVENT LOGS TO SEE IF THE ALARMS WERE SILENCED. MFR REF# 8030229-2015-00023. | 1
4385978 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) FROZE. ALL WAVEFORMS AND ETC. ARE LIKE STILL PICTURES. | 1
4785260 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) DID NOT ALARM FOR A PATIENT WHO HAD A 6 BEAT RUN OF VENTRICULAR TACHYCARDIA. REFERENCE MFR # 8030229-2015-00096. | 1
4386899 | THE CUSTOMER STATED THAT THE PATIENT EXPIRED WHILE THE ALARMS HAD BEEN SHUT OFF ON THE CENTRAL MONITORING SYSTEM (CNS) DUE TO EXCESSIVE ALARM FATIGUE (EXCESSIVE FALSE ALARMS). THEY WOULD LIKE US TO IMMEDIATELY INTERVENE TO DETERMINE IF THE ISSUE IS PRODUCT RELATED. THEY FEEL THAT THE NURSES HAVE BEEN WELL TRAINED ON SKIN PREP AND LEAD PLACEMENT, ALARM MODIFICATION AND MONITORING SYSTEM SETTINGS. THEY FEEL THAT THERE IS A TRUE PROBLEM WITH THE SETTINGS OR THE EQUIPMENT ITSELF. REF MFR # 8030229-2014-00191. | 1
4802830 | CUSTOMER STATED THAT WHEN THEY TURN ON THE Z-920PA TELEMETRY TRANSMITTER, ALL OTHER RECEIVERS THAT ARE MONITORED FROM THE SAME ORG RECEIVER SHOW ARTIFACT. MFR REF # 8030229-2015-00111. | 1
4792971 | THE CUSTOMER STATED THAT UPS IS SHUTTING DOWN FOR THE CENTRAL MONITORING SYSTEM (CNS). | 1
4668487 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) DISPLAY SCREEN WENT BLACK, SO THEY REPLACED IT WITH ANOTHER NON NIHON KOHDEN DISPLAY. THE REPLACED WORKED FOR A WHILE BUT LATER A WINDOWS ERROR APPEARED ERROR INTEL MATRIX CONSOL FAIL - SMART EVENT." REF MFR.# 8030229-2015-00046." | 1
4391023 | THE CUSTOMER STATED THAT THE ZM-920PA TELEMETRY TRANSMITTER FELT IT WAS GETTING WARM. THE BIOMED CAN SEE THAT IN THE BATTERY COMPARTMENT THERE IS A YELLOWISH COLOR ON THE RIGHT SIDE AT THE NEGATIVE MARK. NO PT INJURY. MFR REF # 8030229-2014-00122. | 1
4873865 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) SCREEN LOCKED UP. THEY ARE UNABLE TO MAKE CHANGES TO THE CNS. | 1
4385449 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) INTERMITTENTLY GOES BLANK. MFR REF #: 8030229-2014-00184. | 1
4991354 | THE CUSTOMER REPORTED THAT THE RIGHT DISPLAY OF THE DUAL DISPLAY SET UP OF THE CENTRAL MONITORING SYSTEM (CNS) IS INTERMITTENTLY TURNING ON AND OFF. THIS IS THE 2ND OCCURRENCE. (B)(4). | 1
5005245 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) TURNED OFF BY ITSELF AND WILL NOT POWER BACK ON. REFERENCE MFR # 8030229-2015-00211. | 1
4390981 | THE CUSTOMER REPORTED THAT THE BEDSIDE VITAL SIGN MONITOR READS INACCURATE CO2 VALUES AND THE UNIT IS UNABLE TO PERFORM GAS CALIBRATION PROCEDURE. REF MFR REPORT 8030229-2014-00112. | 1
4793004 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) DISPLAY IS COMPLETELY BLACK WITHOUT ANY POWER INDICATORS. | 1
4384964 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) INTERMITTENTLY FREEZES UP. THE CUSTOMER REPLACED THE HDD LESS THAN A YEAR AGO. REF MFR REPORT #: 8030229-2014-000176. | 1
4994716 | THE CUSTOMER REPORTED THAT THERE WAS A POWER OUTAGE AND INDICATED THAT THE CENTRAL MONITORING SYSTEM (CNS) SPONTANEOUSLY WENT TO A COMPLETE BLUESCREEN. THEY TRIED TO RESTART THE UNIT WITH THE POWER BUTTON, BUT THE UNIT WOULD NOT REBOOT. THE CUSTOMER INDICATED THAT THE UPS DID NOT FAIL AND THE BATTERIES WERE REPLACED IN (B)(6) 2015. | 1
5334190 | UNABLE TO DOWNLOAD PILLCAM FILE, SCREEN INDICATED FILE WAS CORRUPT. THE RECORDER DECK MALFUNCTIONED. | 1
4673553 | OR STAFF REPORTED FAILURE OF DEVICE. A NEW ONE WAS OPENED AND IT FAILED TO OPERATE AS WELL. LOT # M10141211 AND M10150120. | 0
5155661 | THIS WAS A CATHETER CONNECTING ISSUE. WHEN THE CATHETER WAS CONNECTED TO THE SYSTEM, THERE WAS NO IMAGE ON THE DISPLAY. THE CATHETER WAS REMOVED, AND A NEW CATHETER WAS USED. SECOND CATHETER FUNCTIONED AND PROCEDURE COMPLETED. MANUFACTURER RESPONSE FOR DIAGNOSTIC ULTRASOUND CATHETER, SOUNDSTAR ECO (PER SITE REPORTER): CLASS II RECALL.><. | 1
4546534 | ><BATTERY FAILURE DURING USE IN A SURGERY CASE ON THREE DIFFERENT CASES. COMES UP WITH FAILURE ON DISPLAY, BUT BATTERY INDICATES 100% CHARGED. THERE ARE ON MONITORS THAT ARE LESS THAN 8 MONTHS OLD. BATTERY SHOWS DATE CODE OF 04/2014. | 0

## 2016
5406460 | background: we use abacus software to mix our tpn's, and we also carry clinimix-e5/20 premade tpn/s, 2000 ml size. for proper and consistent labeling, we run all of our tpn's through abacus, including clinimix-e orders. we also add lipids to all of our tpns to make them 3-in-1, a concentration of 125 ml lipids per liter tpn. this means that the 2000 ml clinimix bags, we add 250 ml lipids for all clinimix tpn orders. problem: an order was received yesterday to add 10 units of regular insulin per liter of tpn". the tpn was running at 80 ml per hour. the patient happened to be receiving a clinimix-e 5/20 tpn a 2000 ml bag to which we add 250 ml of lipids to make a lipid concentration of 125 ml lipids per liter tpn. it was discovered that the abacus software calculates the amount of insulin incorrectly. for an insulin concentration of 10 units per liter tpn 22.5 units of insulin is the correct amount that should be added (2.25 liters of tpn order for 10 units per liter). however the abacus software calculates the insulin based on 2000 ml the size of the premixed bag instead of the size of the bag after we add the lipids. this creates an error of 12.5%. apparently this is a known problem that baxter hasn't fixed. medication not administered to or used by the patient. (b)(6)." | 1
6017475 | it was reported that a tpn order, created using abacus v3.1 order entry software, included a higher than intended dose of potassium phosphate. this tpn order was compounded, released, and infused on a patient, who experienced ventricular instability" and "coded four times" before being resuscitated and airlifted to another hospital for care. the patient is reported to be in stable condition. follow-up with the reporting facility revealed lab tests confirmed the patient had high potassium levels. tpn therapy was discontinued and dextrose and heparin were ordered and administered to stabilize the patient. the facility reviewed the tpn order mixcheck report (a mixcheck report provides details on orders compounded on the exactamix compounder such as ingredients ingredient amounts overall bag weight etc.) and identified that an incorrect concentration of potassium phosphate was listed on the mixcheck report and that the potassium phosphate was listed as a manual addition to the tpn order. further investigation by the facility found that an incorrect tpn formula template had been selected and used to create the tpn order in the abacus v3.1 order entry software. the template included a diluted form of potassium phosphate (0.01429mmol/ml) not currently in use by the facility; the facility currently utilizes a higher concentration form of potassium phosphate (3mmol/ml). this wasn't caught by the pharmacists creating or reviewing the tpn order nor was it caught when the potassium phosphate was manually added to the tpn bag. this resulted in the 3mmol/ml concentration of potassium phosphate getting added at the volume intended for the 0.01429mmol/ml concentration. baxter technical support guided the facility in making the tpn formula template with the 0.01429mmol/ml concentration of potassium phosphate unavailable for use to prevent this issue from occurring in the future. no additional information was provided.><(b)(4). patient weight is (b)(6). device not returned to manufacturer. a review of the abacus and compounder database and event logs in addition to a review of the bag label physician's order and the mixcheck report found: the formulary and templates were configured correctly and abacus functioned as designed. during order entry the pharmacist who entered the order into abacus selected the order template from the list of applicable templates. however the template chosen was unintended for this patient at this time. the template chosen was created previously for use during a time when there was a shortage of potassium phosphate and included a less-concentrated form (0.01429 mmol/ml) no longer in use at the facility. neither the template nor the ingredient was marked unavailable in abacus after the shortage was resolved. additionally the following was found: abacus order comparison feature displayed an indication that the template was different than the one used the day before. abacus calculated the correct volume of potassium phosphate based on the order form and the ingredient concentration defined. a second pharmacist approved the order within abacus (with the different template and less-concentrated potassium phosphate). upon order approval abacus produced a file and printed a formula label that clearly lists the potassium phosphate ingredient and its 0.01429 mmol/ml concentration. both pharmacists initialed the formula label to indicate review and approval. the order was compounded using an exactamix compounder. because the less-concentrated form of potassium phosphate was not attached to the compounder the exactamix mixcheck report listed it to be manually added. the mixcheck report clearly lists the potassium phosphate ingredient and its concentration of 0.01429 mmol/ml. the pharmacist manually injected a higher concentrated potassium phosphate (3 mmol/ml) into the bag. both pharmacists initialed the mixcheck report to indicate review and approval of the completed solution. based on this investigation the cause was determined to be use error. the abacus v3.1 user guide states: serious harm or death may occur if an adequate review isn't completed. the abacus software changes an order to complete only after an authorized user completely finishes and approves an order. a thorough review of all data on each tab in the order entry process will minimize the risk of medication error. while the exactamix 2400 operator manual states: it is important to print a mixcheck report for every order then have a cosigner (pharmacist) view and approve the entire report especially the formula name; expected weight measured weight and difference; manual additions; and details. should additional relevant information become available a follow-up repot will be submitted." | 1
5795224 | (b)(6) pts (26 pediatric, 21 adult) received x 4 magnesium sulfate doses compounded in their tpn because of a new diluted magnesium sulfate file added to our baxter abacus tpn software unintended to go live but was created for initial testing. the diluted file was created on (b)(6), and the compound used that file for 24 hr tpn bags pumped on (b)(6) resulting in x 4 doses above orders. the software allowed the same ndc to be entered for the current mag sulfate file as well as the new diluted mag sulfate file. the software also defaults to have the unavailable" option box unchecked. the pharmacist adding the new file did not realize that this boxed needed to be checked to prevent the file from potentially going live. we in contact with baxter regarding the event. we have also contacted (b)(4) to report the event and forwarded software screenshots." | 1
5946166 | ><permanent foreign body. healthtronics amica microwave 16 g probe utilized for liver ablation. within 10 seconds of inserting probe, the rib was encountered and probe pulled back. however, tip was noted to be dislodged. | 0
5946165 | ><healthtronics amica microwave 6 g probe utilized for liver ablation. within 10 seconds of inserting probe, the rib was encountered and probe pulled back. however, tip was noted to be dislodged. | 0
5806719 | trouble shooting with technical support confirmed that the display was turning off. likely potential causes for no display are blown display fuse, blown capacitor, blown fuse to power supply, or display power supply malfunction. welch allyn replaced the display to the customer. no further investigation will be conducted.><welch allyn received a report from a welch allyn customer stating that their acuity display is blanking in and out. a display that shuts off during use could result in the loss of centralized patient monitoring. there was no death or injury associated with this complaint. the customer did not provide any patient information. this report was filed in our complaint handling system as complaint (b)(4). | 1
5705332 | ><during surgery, the provider noticed the readings for co2 and o2 were far below normal. ventilator volume looked normal and there were no error messages displayed on the equipment. the gas analyzer module was tested with calibration gas and had readings that were within normal tolerances for the cal gas. just prior to swapping out the entire machine, it was decided to replace the gas sampling line. this immediately corrected the problem. there was delay of about 20 minutes in the procedure while troubleshooting. the suspect sampling line was visually inspected, but no defects were noted. pressure testing of the sampling line indicated a significant leak, which would be consistent with the problem in the room. to find the source of the leak, the line was being prepared to be submerged in water to look for bubbles. at this time the fitting at one end fell apart, so the leak became obvious. packaging from the defective line was not retained, so the lot number is unknown. another line from the same supply bin had lot# 155042. | 0
5920937 | ><ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 0
5830461 | ><the hospital reported that, during the preoperative checkout, the unit had a pressure spike. there was no report of patient involvement. | 0
5764721 | ><the hospital reported that, at the start of a case when switching from manual mode to ventilator mode, the ventilator did not start. the clinician manipulated the bag/vent switch and eventually the ventilator started. there was no reported patient injury. | 0
5447820 | ><the hospital reported that, during a case, the unit had a ventilator failure. the clinician reportedly switched to manual mode to maintain ventilation of the patient. there was no reported patient injury. | 0
5953012 | ><the hospital reported that, during pre-use checkout, the adjustable pressure limiting (apl) valve stuck in the closed position. there was no reported patient involvement. | 0
5563494 | ><the hospital reported that, during a procedure, the unit alarmed for a ventilator failure. the clinician reportedly switched to manual mode, cycled power, switched back to mechanical ventilation and was able to continue the case without further complaint. there was no reported patient injury. | 0
6188871 | ><the hospital reported that, the unit is leaking. there was no reported patient involvement. | 0
6181254 | ><ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. no report of patient involvement. device evaluation anticipated, but not yet begun. | 0
5866932 | ><no report of patient involvement. the initial reporter is located outside the u.s., and therefore this information is not provided due to country privacy laws. ge healthcare's investigation into the reported occurrence is ongoing. a follow-up report will be issued when the investigation has been completed. device evaluation anticipated, but not yet begun. | 0
5616227 | ><machine passed fda checkout. following induction and successful intubation, unable to ventilate patient. several error messages received: vt compensation locked; inspiration stopped and cannot drive bellows. machine removed from service and replaced with a properly functioning unit so that the surgical procedure could be completed. upon examination by clinical engineering, unit was found to have a defective gas inlet valve. | 0
5476561 | a lumbar laminectomy was performed on the pt in the operating room. as is standard for these cases, the pt was moved from the supine to prone position after induction and intubation. the blood pressure cuff and ekg leads were removed from the pt to facilitate moving from supine to prone position on the operating room bed. the ekg leads and blood pressure cuff were then placed on the pt again; 20-25 minutes later, it was noted that the blood pressure cuff was no longer in 'automatic' mode, and blood pressure was not being measured during that time. the issue is that the blood pressure cuff automatically turns off of 'automatic' mode and returns to 'manual' mode. there is no alarm to alert the anesthesia provider that blood pressure measurements are not being taken. the word 'manual' does appear in the blood pressure box on the monitor, however, it is small and difficult to notice, especially with the hectic nature of the operating room. this seems to be an ongoing problem and an issue that occurs frequently to anesthesia providers. in certain cases, this could lead to prolonged undetected hypertension/hypotension that could prove harmful and even deadly for a certain subset of pts. dates of use: 3 hours. diagnosis or reason for use: used to monitor vitals for pts undergoing surgery. | 1
6114572 | patient care was not compromised and was treated appropriately. pending return.><patient was asked to come in for re-biopsy due to inconclusive result and no sample left for re-test. | 0
5871605 | circadiance, llc received customer complaint allegation on (b)(6) 2016 detailing a product issue associated with a smartmonitor 2 device. the customer reported that the device failed to alarm during the functional self-test. no resulting effect on a patient or potential user has been reported.><the complaint issue alleged by the customer was not able to be confirmed because the device has not been sent to circadiance for evaluation. the customer stated that the device will not be sent in for service and requested that the complaint record be closed. the smartmonitor 2 device has provisions to allow the user to detect failure of the audible alarm. specifically, the users are instructed that the audible alarm will annunciate every time that the monitor is turned on allowing the user to verify alarm functionality prior to placing the device into service. additionally, users are instructed to perform a full device checkout by following the published smartmonitor 2 checkout procedure between uses on different patients or at least annually if the monitor has not been in use. the steps of the checkout procedure fully verify all of the alarm functions of the smartmonitor 2. these instructions and features minimize the risk that an audible alarm failure can result in patient injury. the user of this equipment is responsible for reading, understanding, and following the warning and caution statements throughout the manual. no additional information is available at this time. if additional information becomes available or the device is sent to circadiance for evaluation, a follow-up report will be filed. | 1
5439907 | the reported event was confirmed. initial visual inspection revealed no damages. the inability to cool was confirmed by reviewing the case data and the longer than normal cool time was due to a bad mixing pump. replaced the mixing pump with high pump hours and low pump speed and high speed because the pump motor was not keeping steady speed of the motor; the motor was going bad so it was replaced. the unit passed acats testing with burn in, electrical safety testing and a 3 day run period with bypass loop attached. a coin cell leakage test was performed and the unit passed. the coin cell battery was replaced for routine maintenance. the control panel passed final testing. software loaded. the device was returned to the customer. the device history record was reviewed and found nothing that could have caused or contributed to the reported event. the instructions for use state the following: the arctic sun? temperature management system is intended for monitoring and controlling patient temperature. warnings and cautions: warnings: do not use the arctic sun? temperature management system in the presence of flammable agents because an explosion and/or fire may result. do not use high frequency surgical instruments or endocardial catheters while the arctic sun? temperature management system is in use. there is a risk of electrical shock and hazardous moving parts. there are no user serviceable parts inside. do not remove covers. refer servicing to qualified personnel. power cord has a hospital grade plug. grounding reliability can only be achieved when connected to an equivalent receptacle marked ?hospital use? or ?hospital grade?. when using the arctic sun? temperature management system note that all other thermal conductive systems such as water blankets and water gels in use while warming or cooling with the arctic sun? temperature management system may actually alter or interfere with patient temperature control. do not place arcticgel? pads over transdermal medication patches as warming can increase drug delivery resulting in possible harm to the patient. cautions: this product is to be used by or under the supervision of trained qualified medical personnel. federal law (usa) restricts this device to sale by or on the order of a physician. use only distilled or sterile water. the use of other fluids will damage the arctic sun? temperature management system. when moving the arctic sun? temperature management system always use the handle to lift the controller over an obstacle to avoid over balancing. the patient?s bed surface should be located between 30 and 60 inches (75 cm and 150 cm) above the floor to ensure proper flow and minimize risk of leaks. the clinician is responsible to determine the appropriateness of custom parameters. when the system is powered off all changes to parameters will revert to the default unless the new settings have been saved as new defaults in the advanced setup screen. for small patients (=30 kg) it is recommended to use the following settings: water temperature high limit =40??c (104??f); water temperature low limit =10??c (64.4 ??f); control strategy =2. the operator must continuously monitor patient temperature when using manual control and adjust the temperature of the water flowing through the pads accordingly. patient temperature will not be controlled by the arctic sun? temperature management system in manual control. due to the system?s high efficiency manual control is not recommended for long duration use. the operator is advised to use the automatic therapy modes (e.g. control patient cool patient rewarm patient) for automatic patient temperature monitoring and control. the arctic sun? temperature management system will monitor and control patient core temperature based on the temperature probe attached to the system. the clinician is responsible for correctly placing the temperature probe and verifying the accuracy and placement of the patient probe at the start of the procedure. medivance recommends measuring patient temperature from a second site to verify patient temperature. medivance recommends the use of a second patient temperature probe connected to the arctic sun? temperature management system temperature 2 input as it provides continuous monitoring and safety alarm features. alternatively patient temperature may be verified periodically with independent instrumentation. the displayed temperature graph is for general information purposes only and is not intended to replace standard medical record documentation for use in therapy decisions. patient temperature will not be controlled and alarms are not enabled in stop mode. patient temperature may increase or decrease with the arctic sun? temperature management system in stop mode. carefully observe the system for air leaks before and during use. if the pads fail to prime or a significant continuous air leak is observed in the pad return line check connections. if needed replace the leaking pad. leakage may result in lower flow rates and potentially decrease the performance of the system. the arctic sun? temperature management system is for use only with the arcticgel? pads. the arcticgel? pads are only for use with the arctic sun? temperature management systems. the arcticgel? pads are non-sterile for single patient use. do not reprocess or sterilize. if used in a sterile environment pads should be placed according to the physician?s request either prior to the sterile preparation or sterile draping. arcticgel? pads should not be placed on a sterile field. use pads immediately after opening. do not store pads once the kit has been opened. do not place arcticgel? pads on skin that has signs of ulceration burns hives or rash. while there are no known allergies to hydrogel materials caution should be exercised with any patient who has a history of skin allergies or sensitivities. do not allow circulating water to contaminate the sterile field when patient lines are disconnected. the water content of the hydrogel affects the pad?s adhesion to the skin and conductivity and therefore the efficiency of controlling patient temperature. periodically check that pads remain moist and adherent. replace pads when the hydrogel no longer uniformly adheres to the skin. replacing pads at least every 5 days is recommended. do not puncture the arcticgel? pads with sharp objects. punctures will result in air entering the fluid pathway and may reduce performance. if accessible examine the patient?s skin under the arcticgel? pads often especially those at higher risk of skin injury. skin injury may occur as a cumulative result of pressure time and temperature. do not place bean bag or other firm positioning devices under the arcticgel? pads. do not place positioning devices under the pad manifolds or patient lines. the rate of temperature change and potentially the final achievable patient temperature is affected by many factors. treatment application monitoring and results are the responsibility of the attending physician. if the patient does not reach target temperature in a reasonable time or the patient is not able to be maintained at the target temperature the skin may be exposed to low or high water temperatures for an extended period of time which may increase the risk for skin injury. ensure that pad sizing / coverage and custom parameter settings are correct for the patient and treatment goals water flow is greater than or equal to 2.3 liters per minute and the patient temperature probe is in the correct place. for patient cooling ensure environmental factors such as excessively hot rooms heat lamps and heated nebulizers are eliminated and patient shivering is controlled. otherwise consider increasing minimum water temperature modifying target temperature to an attainable setting or discontinuing treatment. for patient warming consider decreasing maximum water temperature modifying target temperature to an attainable setting or discontinuing treatment. due to underlying medical or physiological conditions some patients are more susceptible to skin damage from pressure and heat or cold. patients at risk include those with poor tissue perfusion or poor skin integrity due to diabetes peripheral vascular disease poor nutritional status steroid use or high dose vasopressor therapy. if warranted use pressure relieving or pressure reducing devices under the patient to protect from skin injury. do not allow urine antibacterial solutions or other agents to pool underneath the arcticgel? pads. urine and antibacterial agents can absorb into the pad hydrogel and cause chemical injury and loss of pad adhesion. replace pads immediately if these fluids come into contact with the hydrogel. do not place arcticgel? pads over an electrosurgical grounding pad. the combination of heat sources may result in skin burns. if needed place defibrillation pads between the arcticgel? pads and the patient?s skin. carefully remove arcticgel? pads from the patient?s skin at the completion of use. discard used arcticgel? pads in accordance with hospital procedures for medical waste. the usb data port is to be used only with a standalone usb flash drive. do not connect to another mains powered device during patient treatment. users should not use cleaning or decontamination methods different from those recommended by the manufacturer without first checking with the manufacturer that the proposed methods will not damage the equipment. do not use bleach (sodium hypochlorite) as it may damage the system. medivance will not be responsible for patient safety or equipment performance if the procedures to operate maintain modify or service the medivance arctic sun? temperature management system are other than those specified by medivance. anyone performing the procedures must be appropriately trained and qualified." (b)(4). the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><it was reported that the device was not cooling the water. therapy was interrupted and the device was replaced with another arctic sun.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard." | 0
5442225 | the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><case data was received. the reported issue was confirmed as a use of device issue. the clinical representative reviewed the casa data and along with knowledge of renal therapy being delivered, the ability to not rewarm is confirmed. arctic sun was functioning properly, but couldn't rewarm patient due to conflicting temperature inputs to the patient with continuous renal replacement therapy (crrt). the root cause for the reported issue that the device couldn't rewarm the patient is due to competing therapy with crrt. device remained in service. the device history record was reviewed and found nothing that could have caused or contributed to the reported event. the instructions for use state the following: the arctic sun? temperature management system is intended for monitoring and controlling patient temperature. warnings and cautions: warnings: do not use the arctic sun? temperature management system in the presence of flammable agents because an explosion and/or fire may result; do not use high frequency surgical instruments or endocardial catheters while the arctic sun? temperature management system is in use; there is a risk of electrical shock and hazardous moving parts. there are no user serviceable parts inside. do not remove covers. refer servicing to qualified personnel; power cord has a hospital grade plug. grounding reliability can only be achieved when connected to an equivalent receptacle marked ?hospital use? or ?hospital grade?; when using the arctic sun? temperature management system note that all other thermal conductive systems such as water blankets and water gels in use while warming or cooling with the arctic sun? temperature management system may actually alter or interfere with patient temperature control; do not place arcticgel? pads over transdermal medication patches as warming can increase drug delivery resulting in possible harm to the patient. cautions: this product is to be used by or under the supervision of trained qualified medical personnel; federal law (usa) restricts this device to sale by or on the order of a physician; use only distilled or sterile water. the use of other fluids will damage the arctic sun? temperature management system; when moving the arctic sun? temperature management system always use the handle to lift the controller over an obstacle to avoid over balancing; the patient bed surface should be located between 30 and 60 inches (75 cm and 150 cm) above the floor to ensure proper flow and minimize risk of leaks; the clinician is responsible to determine the appropriateness of custom parameters. when the system is powered off all changes to parameters will revert to the default unless the new settings have been saved as new defaults in the advanced setup screen. for small patients (=30 kg) it is recommended to use the following settings: water temperature high limit =40??c (104??f); water temperature low limit =10??c (64.4 ??f); control strategy =2; the operator must continuously monitor patient temperature when using manual control and adjust the temperature of the water flowing through the pads accordingly. patient temperature will not be controlled by the arctic sun? temperature management system in manual control; due to the systems high efficiency manual control is not recommended for long duration use. the operator is advised to use the automatic therapy modes (e.g. control patient cool patient rewarm patient) for automatic patient temperature monitoring and control; the arctic sun? temperature management system will monitor and control patient core temperature based on the temperature probe attached to the system. the clinician is responsible for correctly placing the temperature probe and verifying the accuracy and placement of the patient probe at the start of the procedure; (b)(4) recommends measuring patient temperature from a second site to verify patient temperature. (b)(4) recommends the use of a second patient temperature probe connected to the arctic sun? temperature management system temperature 2 input as it provides continuous monitoring and safety alarm features. alternatively patient temperature may be verified periodically with independent instrumentation; the displayed temperature graph is for general information purposes only and is not intended to replace standard medical record documentation for use in therapy decisions; patient temperature will not be controlled and alarms are not enabled in stop mode. patient temperature may increase or decrease with the arctic sun? temperature management system in stop mode; carefully observe the system for air leaks before and during use. if the pads fail to prime or a significant continuous air leak is observed in the pad return line check connections. if needed replace the leaking pad. leakage may result in lower flow rates and potentially decrease the performance of the system; the arctic sun? temperature management system is for use only with the arcticgel? pads; the arcticgel? pads are only for use with the arctic sun? temperature management systems; the arcticgel? pads are non-sterile for single patient use. do not reprocess or sterilize. if used in a sterile environment pads should be placed according to the physician request either prior to the sterile preparation or sterile draping. arcticgel? pads should not be placed on a sterile field; use pads immediately after opening. do not store pads once the kit has been opened; do not place arcticgel? pads on skin that has signs of ulceration burns hives or rash; while there are no known allergies to hydrogel materials caution should be exercised with any patient who has a history of skin allergies or sensitivities; do not allow circulating water to contaminate the sterile field when patient lines are disconnected; the water content of the hydrogel affects the pad adhesion to the skin and conductivity and therefore the efficiency of controlling patient temperature. periodically check that pads remain moist and adherent. replace pads when the hydrogel no longer uniformly adheres to the skin. replacing pads at least every 5 days is recommended; do not puncture the arcticgel? pads with sharp objects. punctures will result in air entering the fluid pathway and may reduce performance; if accessible examine the patient skin under the arcticgel? pads often especially those at higher risk of skin injury. skin injury may occur as a cumulative result of pressure time and temperature. do not place bean bag or other firm positioning devices under the arcticgel? pads. do not place positioning devices under the pad manifolds or patient lines; the rate of temperature change and potentially the final achievable patient temperature is affected by many factors. treatment application monitoring and results are the responsibility of the attending physician. if the patient does not reach target temperature in a reasonable time or the patient is not able to be maintained at the target temperature the skin may be exposed to low or high water temperatures for an extended period of time which may increase the risk for skin injury. ensure that pad sizing / coverage and custom parameter settings are correct for the patient and treatment goals water flow is greater than or equal to 2.3 liters per minute and the patient temperature probe is in the correct place. for patient cooling ensure environmental factors such as excessively hot rooms heat lamps and heated nebulizers are eliminated and patient shivering is controlled. otherwise consider increasing minimum water temperature modifying target temperature to an attainable setting or discontinuing treatment. for patient warming consider decreasing maximum water temperature modifying target temperature to an attainable setting or discontinuing treatment; due to underlying medical or physiological conditions some patients are more susceptible to skin damage from pressure and heat or cold. patients at risk include those with poor tissue perfusion or poor skin integrity due to diabetes peripheral vascular disease poor nutritional status steroid use or high dose vasopressor therapy. if warranted use pressure relieving or pressure reducing devices under the patient to protect from skin injury; do not allow urine antibacterial solutions or other agents to pool underneath the arcticgel? pads. urine and antibacterial agents can absorb into the pad hydrogel and cause chemical injury and loss of pad adhesion. replace pads immediately if these fluids come into contact with the hydrogel; do not place arcticgel? pads over an electrosurgical grounding pad. the combination of heat sources may result in skin burns; if needed place defibrillation pads between the arcticgel? pads and the patient?s skin; carefully remove arcticgel? pads from the patient?s skin at the completion of use. discard used arcticgel? pads in accordance with hospital procedures for medical waste; the usb data port is to be used only with a standalone usb flash drive. do not connect to another mains powered device during patient treatment; users should not use cleaning or decontamination methods different from those recommended by the manufacturer without first checking with the manufacturer that the proposed methods will not damage the equipment. do not use bleach (sodium hypochlorite) as it may damage the system; (b)(4) will not be responsible for patient safety or equipment performance if the procedures to operate maintain modify or service the (b)(4) arctic sun? temperature management system are other than those specified by (b)(4). anyone performing the procedures must be appropriately trained and qualified." (b)(4). the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard. case data was received.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed.><it was reported that the device was not heating the patient. therapy was interrupted in order to replace the device with another device.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard." | 0
5435611 | the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><it was reported that the device was unable to reach target temperature of 33c. the patient was taken off of therapy because the target temperature was not reached.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><case data was received. the reported issue was confirmed. the mixing pump was found to be the root cause of the inability to cool. the mixing pump failed to circulate water from the main circulation tank to the chiller tank to be cooled. this can be caused by either the pump head or pump motor. further evaluation of the pump is not possible as no sample will be returned for evaluation. this failure resulted in the water level not rising in the chiller tank to properly overflow chilled water back into the main circulation tank. this failure of the mixing pump causes the reduced water temperature control and prevents chilled water from being provided to the patient therapy pads for proper patient temperature management. the mixing pump was replaced by palex, spain. the device passed their testing and was released back to use. the device history record was reviewed and found nothing that could have caused or contributed to the reported event. the instructions for use state the following: the arctic sun? temperature management system is intended for monitoring and controlling patient temperature. warnings and cautions - warnings: do not use the arctic sun? temperature management system in the presence of flammable agents because an explosion and/or fire may result. do not use high frequency surgical instruments or endocardial catheters while the arctic sun? temperature management system is in use. there is a risk of electrical shock and hazardous moving parts. there are no user serviceable parts inside. do not remove covers. refer servicing to qualified personnel. power cord has a hospital grade plug. grounding reliability can only be achieved when connected to an equivalent receptacle marked ?hospital use? or ?hospital grade?. when using the arctic sun? temperature management system note that all other thermal conductive systems such as water blankets and water gels in use while warming or cooling with the arctic sun? temperature management system may actually alter or interfere with patient temperature control. do not place arcticgel? pads over transdermal medication patches as warming can increase drug delivery resulting in possible harm to the patient. cautions: this product is to be used by or under the supervision of trained qualified medical personnel. federal law (usa) restricts this device to sale by or on the order of a physician. use only distilled or sterile water. the use of other fluids will damage the arctic sun? temperature management system. when moving the arctic sun? temperature management system always use the handle to lift the controller over an obstacle to avoid over balancing. the patient bed surface should be located between 30 and 60 inches (75 cm and 150 cm) above the floor to ensure proper flow and minimize risk of leaks. the clinician is responsible to determine the appropriateness of custom parameters. when the system is powered off all changes to parameters will revert to the default unless the new settings have been saved as new defaults in the advanced setup screen. for small patients (=(b)(6) kg) it is recommended to use the following settings: water temperature high limit =40??c (104??f); water temperature low limit =10??c (64.4 ??f); control strategy =2. the operator must continuously monitor patient temperature when using manual control and adjust the temperature of the water flowing through the pads accordingly. patient temperature will not be controlled by the arctic sun? temperature management system in manual control. due to the systems high efficiency manual control is not recommended for long duration use. the operator is advised to use the automatic therapy modes (e.g. control patient cool patient rewarm patient) for automatic patient temperature monitoring and control. the arctic sun? temperature management system will monitor and control patient core temperature based on the temperature probe attached to the system. the clinician is responsible for correctly placing the temperature probe and verifying the accuracy and placement of the patient probe at the start of the procedure. medivance recommends measuring patient temperature from a second site to verify patient temperature. medivance recommends the use of a second patient temperature probe connected to the arctic sun? temperature management system temperature 2 input as it provides continuous monitoring and safety alarm features. alternatively patient temperature may be verified periodically with independent instrumentation. the displayed temperature graph is for general information purposes only and is not intended to replace standard medical record documentation for use in therapy decisions. patient temperature will not be controlled and alarms are not enabled in stop mode. patient temperature may increase or decrease with the arctic sun? temperature management system in stop mode. carefully observe the system for air leaks before and during use. if the pads fail to prime or a significant continuous air leak is observed in the pad return line check connections. if needed replace the leaking pad. leakage may result in lower flow rates and potentially decrease the performance of the system. the arctic sun? temperature management system is for use only with the arcticgel? pads. the arcticgel? pads are only for use with the arctic sun? temperature management systems. the arcticgel? pads are non-sterile for single patient use. do not reprocess or sterilize. if used in a sterile environment pads should be placed according to the physician request either prior to the sterile preparation or sterile draping. arcticgel? pads should not be placed on a sterile field. use pads immediately after opening. do not store pads once the kit has been opened. do not place arcticgel? pads on skin that has signs of ulceration burns hives or rash. while there are no known allergies to hydrogel materials caution should be exercised with any patient who has a history of skin allergies or sensitivities. do not allow circulating water to contaminate the sterile field when patient lines are disconnected. the water content of the hydrogel affects the pad adhesion to the skin and conductivity and therefore the efficiency of controlling patient temperature. periodically check that pads remain moist and adherent. replace pads when the hydrogel no longer uniformly adheres to the skin. replacing pads at least every 5 days is recommended. do not puncture the arcticgel? pads with sharp objects. punctures will result in air entering the fluid pathway and may reduce performance. if accessible examine the patient skin under the arcticgel? pads often especially those at higher risk of skin injury. skin injury may occur as a cumulative result of pressure time and temperature. do not place bean bag or other firm positioning devices under the arcticgel? pads. do not place positioning devices under the pad manifolds or patient lines. the rate of temperature change and potentially the final achievable patient temperature is affected by many factors. treatment application monitoring and results are the responsibility of the attending physician. if the patient does not reach target temperature in a reasonable time or the patient is not able to be maintained at the target temperature the skin may be exposed to low or high water temperatures for an extended period of time which may increase the risk for skin injury. ensure that pad sizing / coverage and custom parameter settings are correct for the patient and treatment goals water flow is greater than or equal to 2.3 liters per minute and the patient temperature probe is in the correct place. for patient cooling ensure environmental factors such as excessively hot rooms heat lamps and heated nebulizers are eliminated and patient shivering is controlled. otherwise consider increasing minimum water temperature modifying target temperature to an attainable setting or discontinuing treatment. for patient warming consider decreasing maximum water temperature modifying target temperature to an attainable setting or discontinuing treatment. due to underlying medical or physiological conditions some patients are more susceptible to skin damage from pressure and heat or cold. patients at risk include those with poor tissue perfusion or poor skin integrity due to diabetes peripheral vascular disease poor nutritional status steroid use or high dose vasopressor therapy. if warranted use pressure relieving or pressure reducing devices under the patient to protect from skin injury. do not allow urine antibacterial solutions or other agents to pool underneath the arcticgel? pads. urine and antibacterial agents can absorb into the pad hydrogel and cause chemical injury and loss of pad adhesion. replace pads immediately if these fluids come into contact with the hydrogel. do not place arcticgel? pads over an electrosurgical grounding pad. the combination of heat sources may result in skin burns. if needed place defibrillation pads between the arcticgel? pads and the patient?s skin. carefully remove arcticgel? pads from the patient?s skin at the completion of use. discard used arcticgel? pads in accordance with hospital procedures for medical waste. the usb data port is to be used only with a stand alone usb flash drive. do not connect to another mains powered device during patient treatment. users should not use cleaning or decontamination methods different from those recommended by the manufacturer without first checking with the manufacturer that the proposed methods will not damage the equipment. do not use bleach (sodium hypochlorite) as it may damage the system. medivance will not be responsible for patient safety or equipment performance if the procedures to operate maintain modify or service the medivance arctic sun? temperature management system are other than those specified by medivance. anyone performing the procedures must be appropriately trained and qualified." (b)(4). case data received." | 0
5441038 | the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard.><the investigation is still in progress. once the investigation is complete, a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard. sample not available.><it was reported that the device was not cooling the patient. therapy was interrupted and in order to replace the device with another arctic sun.><the reported issue was confirmed. no sample was returned for evaluation; however, further investigation by the biomedical technician isolated the issue to a bad mixing pump. the mixing pump was replaced and the device passed final testing, calibration and electrical safety testing at medtech. the device was placed back in use. the device history record was reviewed and found nothing that could have caused or contributed to the reported event. the instructions for use state the following: the arctic sun? temperature management system is intended for monitoring and controlling patient temperature. warnings and cautions: warnings: do not use the arctic sun? temperature management system in the presence of flammable agents because an explosion and/or fire may result; do not use high frequency surgical instruments or endocardial catheters while the arctic sun? temperature management system is in use; there is a risk of electrical shock and hazardous moving parts. there are no user serviceable parts inside. do not remove covers. refer servicing to qualified personnel; power cord has a hospital grade plug. grounding reliability can only be achieved when connected to an equivalent receptacle marked ?hospital use? or ?hospital grade?; when using the arctic sun? temperature management system note that all other thermal conductive systems such as water blankets and water gels in use while warming or cooling with the arctic sun? temperature management system may actually alter or interfere with patient temperature control; do not place arcticgel? pads over transdermal medication patches as warming can increase drug delivery resulting in possible harm to the patient. cautions: this product is to be used by or under the supervision of trained qualified medical personnel; federal law (usa) restricts this device to sale by or on the order of a physician; use only distilled or sterile water. the use of other fluids will damage the arctic sun? temperature management system; when moving the arctic sun? temperature management system always use the handle to lift the controller over an obstacle to avoid over balancing; the patient?s bed surface should be located between 30 and 60 inches (75 cm and 150 cm) above the floor to ensure proper flow and minimize risk of leaks; the clinician is responsible to determine the appropriateness of custom parameters. when the system is powered off all changes to parameters will revert to the default unless the new settings have been saved as new defaults in the advanced setup screen. for small patients (=30 kg) it is recommended to use the following settings: water temperature high limit =40??c (104??f); water temperature low limit =10??c (64.4 ??f); control strategy =2; the operator must continuously monitor patient temperature when using manual control and adjust the temperature of the water flowing through the pads accordingly. patient temperature will not be controlled by the arctic sun? temperature management system in manual control; due to the system?s high efficiency manual control is not recommended for long duration use. the operator is advised to use the automatic therapy modes (e.g. control patient cool patient rewarm patient) for automatic patient temperature monitoring and control; the arctic sun? temperature management system will monitor and control patient core temperature based on the temperature probe attached to the system. the clinician is responsible for correctly placing the temperature probe and verifying the accuracy and placement of the patient probe at the start of the procedure; (b)(4) recommends measuring patient temperature from a second site to verify patient temperature. (b)(4) recommends the use of a second patient temperature probe connected to the arctic sun? temperature management system temperature 2 input as it provides continuous monitoring and safety alarm features. alternatively patient temperature may be verified periodically with independent instrumentation; the displayed temperature graph is for general information purposes only and is not intended to replace standard medical record documentation for use in therapy decisions; patient temperature will not be controlled and alarms are not enabled in stop mode. patient temperature may increase or decrease with the arctic sun? temperature management system in stop mode; carefully observe the system for air leaks before and during use. if the pads fail to prime or a significant continuous air leak is observed in the pad return line check connections. if needed replace the leaking pad. leakage may result in lower flow rates and potentially decrease the performance of the system; the arctic sun? temperature management system is for use only with the arcticgel? pads; the arcticgel? pads are only for use with the arctic sun? temperature management systems; the arcticgel? pads are non-sterile for single patient use. do not reprocess or sterilize. if used in a sterile environment pads should be placed according to the physician?s request either prior to the sterile preparation or sterile draping. arcticgel? pads should not be placed on a sterile field; use pads immediately after opening. do not store pads once the kit has been opened; do not place arcticgel? pads on skin that has signs of ulceration burns hives or rash; while there are no known allergies to hydrogel materials caution should be exercised with any patient who has a history of skin allergies or sensitivities; do not allow circulating water to contaminate the sterile field when patient lines are disconnected; the water content of the hydrogel affects the pad?s adhesion to the skin and conductivity and therefore the efficiency of controlling patient temperature. periodically check that pads remain moist and adherent. replace pads when the hydrogel no longer uniformly adheres to the skin. replacing pads at least every 5 days is recommended; do not puncture the arcticgel? pads with sharp objects. punctures will result in air entering the fluid pathway and may reduce performance; if accessible examine the patient?s skin under the arcticgel? pads often especially those at higher risk of skin injury. skin injury may occur as a cumulative result of pressure time and temperature. do not place bean bag or other firm positioning devices under the arcticgel? pads. do not place positioning devices under the pad manifolds or patient lines; the rate of temperature change and potentially the final achievable patient temperature is affected by many factors. treatment application monitoring and results are the responsibility of the attending physician. if the patient does not reach target temperature in a reasonable time or the patient is not able to be maintained at the target temperature the skin may be exposed to low or high water temperatures for an extended period of time which may increase the risk for skin injury. ensure that pad sizing / coverage and custom parameter settings are correct for the patient and treatment goals water flow is greater than or equal to 2.3 liters per minute and the patient temperature probe is in the correct place. for patient cooling ensure environmental factors such as excessively hot rooms heat lamps and heated nebulizers are eliminated and patient shivering is controlled. otherwise consider increasing minimum water temperature modifying target temperature to an attainable setting or discontinuing treatment. for patient warming consider decreasing maximum water temperature modifying target temperature to an attainable setting or discontinuing treatment; due to underlying medical or physiological conditions some patients are more susceptible to skin damage from pressure and heat or cold. patients at risk include those with poor tissue perfusion or poor skin integrity due to diabetes peripheral vascular disease poor nutritional status steroid use or high dose vasopressor therapy. if warranted use pressure relieving or pressure reducing devices under the patient to protect from skin injury; do not allow urine antibacterial solutions or other agents to pool underneath the arcticgel? pads. urine and antibacterial agents can absorb into the pad hydrogel and cause chemical injury and loss of pad adhesion. replace pads immediately if these fluids come into contact with the hydrogel; do not place arcticgel? pads over an electrosurgical grounding pad. the combination of heat sources may result in skin burns; if needed place defibrillation pads between the arcticgel? pads and the patient?s skin; carefully remove arcticgel? pads from the patient?s skin at the completion of use. discard used arcticgel? pads in accordance with hospital procedures for medical waste; the usb data port is to be used only with a standalone usb flash drive. do not connect to another mains powered device during patient treatment; users should not use cleaning or decontamination methods different from those recommended by the manufacturer without first checking with the manufacturer that the proposed methods will not damage the equipment. do not use bleach (sodium hypochlorite) as it may damage the system; (b)(4) will not be responsible for patient safety or equipment performance if the procedures to operate maintain modify or service the (b)(4) arctic sun? temperature management system are other than those specified by (b)(4). anyone performing the procedures must be appropriately trained and qualified." (b)(4). the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard. the device was evaluated by a biomedical engineer at (b)(4).><the investigation is still in progress. once the investigation is complete a supplemental report will be filed. the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information the complainant / reporter was unable or unwilling to provide any further patient product or procedural details to bard.><it was reported that the patient was cooled for normothermia." | 0
5688075 | this is our final report on this incident.><this is our initial report on this incident.><the customer reported that a patient sample yielded a false negative test result using biotestcell 3 on one tango optimo (tango #1) at his site. a new draw of the same patient was tested on a second tango optimo on his site (tango #2) and yielded a positive result. an anti-jk(b) was identified. the customer retested the first patient sample on the second tango ( #2) and got a positive result. the customer advised to send in the patient sample and the supposedly defective product biotestcell 3 for investigational testing. our quality control laboratory is still waiting for the material. the affected tango optimo (#1) was checked by our field service engineers. evaluation of the data is still ongoing.><the customer reported that a patient sample yielded a false negative test result when tested with biotestcell 3 on one tango optimo (tango #1) at his site. a new draw of the same patient was tested on a second tango optimo on his site (tango #2) and yielded a positive result. an anti-jk(b) was identified. the customer retested the first patient sample on the second tango ( #2) and got a positive result. the customer did neither return the patient sample that had caused the false negative test result nor the supposedly defective product. therefore our quality control laboratory tested their retained sample of biotestcell 3 with different samples and controls on tango optimo. all positive and negative reactions were correct. we did not observe any false negative reaction. testing by our quality control laboratory confirmed that the allegedly defective lot of biotestcell 3 functions correctly. a review of the batch record documentation showed no irregularities which might have negatively affected the quality of the allegedly defective lot. the affected tango optimo (tango #1) was inspected by our field service engineer. one syringe was replaced and the result camera was replaced, too. a post service performance procedure was performed and the instrument found to be working again within its specification. | 0
6025351 | this is our final report on this incident.><a false negative ssc ii control - which is a diluted anti-d, was reported on a training instrument while teaching new customers. when this actual control was run on a second tango infinity, it yielded a positive result. furthermore the control was also retested on the first tango infinity which had previously given a negative test result and now yielded positive results. the customer returned the supposedly defective product and the affected reagents such as the reagent red blood cells for the antibody screening test and the anti-human globulin anti-igg solidscreen ii. the supposedly defective product as well as the reagents returned by the customer were tested with various samples and controls and yielded correct positive and negative results. we did not observe any false negative results. testing by our quality control laboratory confirmed that the allegedly defective lot of solidscreen ii control functions correctly. a review of the batch record documentation showed no irregularities which might have negatively affected the quality of the allegedly defective lot. the affected tango infinity was inspected by one of our field service engineers. the washing process was found to be fully functional. the analysis of the log files is still ongoing.><a false negative ssc ii control - which is a diluted anti-d - was reported on a training instrument while teaching new customers. when this actual control was run on a second tango infinity, it yielded a positive result. furthermore the control was also retested on the first tango infinity which had previously given a negative test result and now yielded positive results. the customer returned the supposedly defective product and the affected reagents such as the reagent red blood cells for the antibody screening test and the anti-human globulin anti-igg solidscreen ii. the supposedly defective product as well as the reagents returned by the customer were tested with various samples and controls and yielded correct positive and negative results. we did not observe any false negative results. testing by our quality control laboratory confirmed that the allegedly defective lot of solidscreen ii control functions correctly. a review of the batch record documentation showed no irregularities which might have negatively affected the quality of the allegedly defective lot. the affected tango infinity was inspected by one of our field service engineers and no indication for an instrument and software malfunction could be identified on current data. the ssc ii control failed only once and the re-testing on the same instrument revealed correct results. the failed control sample was not flagged, because the washing error message (flag 15) did not occur on this sample well. the instruments washing process was confirmed to operate within specification.><this is our initial report on this incident. | 0
5629661 | the customer reported that a patient sample was processed on erytype s abd+rev. a1, b on tango optimo and the instrument called the patient o rhd neg. the images looked off - dark brown and grainy. the sample was processed on the customer's second tango optimo and this instrument called the patient o rhd pos. the reactions were very clear and correct. the customer did neither return the supposedly defective product erytype s abd+rev. a1,b fjor investigational testing nor the patient sample that had caused the false negative test result. therefore our quality control laboratory tested their retained erytype s abd+rev.a1,b sample with different donor samples. all positive and negative reactions were correct. we did neither observe any false negative reaction nor dark brown and grainy images. all images were clear and showed the expected red color of red blood cells. testing by our quality control laboratory confirmed that the allegedly defective lot of erytype s abd+rev.a1,b functions correctly. a review of the batch record documentation showed no irregularities which might have negatively affected the quality of the allegedly defective lot. the affected tango optimo was inspected by our field service engineer. we are waiting for the results of this investigation.><this is our initial report on this incident.><the customer reported that a patient sample was processed on erytype s abd+rev. a1, b on tango optimo and the instrument called the patient o rhd neg. the images looked off - dark brown and grainy. the sample was processed on the customer's second tango optimo and this instrument called the patient o rhd pos. the reactions were very clear and correct. the customer did neither return the supposedly defective product erytype s abd+rev. a1,b for investigational testing nor the patient sample that had caused the false negative test result. therefore our quality control laboratory tested their retained erytype s abd+rev.a1,b sample with different donor samples. all positive and negative reactions were correct. we did neither observe any false negative reaction nor dark brown and grainy images. all images were clear and showed the expected red color of red blood cells. testing by our quality control laboratory confirmed that the allegedly defective lot of erytype s abd+rev.a1,b functions correctly. a review of the batch record documentation showed no irregularities which might have negatively affected the quality of the allegedly defective lot. the affected tango optimo was inspected by our field service engineer. relevant components were replaced, inspected, cleaned, and lubed as required during maintenance. the system was also checked for any leaks or visible errors, none found. a performance verification was ran and the system found to be performing to bio-rad specifications. the tango optimo was returned to the customer in full operation.><this is our final report on this incident. | 0
6068871 | this is our initial report on this incident.><the customer stated that the tango infinity reported false negative results with for dat positive control several times. testing on another tango infinity instrument at the same site revealed 2+ results. the customer said to send in the sample that had caused a false negative test result and also the supposedly defective product for investigational testing. our quality control laboratory is still waiting for these samples. the affected instrument was thoroughly inspected by our field service engineers. root cause analysis is still ongoing. | 0
6120620 | on (b)(6) 2016, a customer site reported an unexpectedly negative antibody screen when tested on a galileo echo instrument on (b)(6) 2016.><immucor technical support used a remote electronic connection method to assess the instrument test well image in question on (b)(6) 2016, and that test well image appeared visually positive. | 0
5394040 | a review of the image result files showed that cell 2 of the initial antibody screening assay resulted as negative, but visually appeared weakly positive. customers were notified of technical communication cc-09-042-02 which advises customers to visually inspect all negative antibody screening and identification results on the echo prior to release of results.><on (b)(6) 2015, the customer reported that an unexpected negative result was obtained with a sample known to contain anti-e. | 0
6100534 | an immucor field service engineer (fse) visited the customer site on 26oct2016 to assess the testing instrument in question, and determined that the syringes were dirty. consequently, the fse removed and cleaned the syringes.><on (b)(4) 2016, a customer reported an abo mistype when testing blood samples on a galileo neo on (b)(6) 2016. | 0
6166099 | an immucor field service engineer (fse) visited the customer site on 01dec2016 to assess the testing instrument in question. the fse determined that the instrument was operating as expected.><on (b)(6) 2016, an unexpectedly negative antibody screen was documented, when tested on a galileo echo instrument. | 0
5739882 | on (b)(6) 2016 a customer reported unexpected negative results on the galileo echo instrument. customer states that patient had transfusion reaction which prompted them to do a transfusion reaction work up. they tested both the pre- and post-transfusion samples and both resulted as negative on echo serial number (b)(4).><review of instrument results files: sample id (b)(6) (pre transfusion sample) echo reported negative for all screen cells. screen cells 1 and 3 are visually negative. screen cell 2 is visually weakly positive. sample id (b)(6) (post transfusion sample) echo reported negative for all screen cells. screen cells 1 and 3 are visually negative. screen cell 2 is visually weakly positive. a service call was made. the instrument was tested and is operating within specifications with no errors observed. customer was informed that negative wells that visually appear positive are addressed in (b)(4). customer advised to perform a visual verification of negative reactions before final release of those well results. | 0
5894638 | immucor technical support used a remote electronic access method to assess the testing well images in question on (b)(6) 2016. the test well in question which was unexpectedly negative from the instrument in question was visually positive. this was well number two (2) of the antibody screen, which was known to be e antigen positive.><on (b)(6) 2016, an immucor technical employee at a customer site reported an unexpected negative antibody screen when tested on a galileo echo instrument. | 0
6100526 | an immucor field service engineer (fse) visited the customer site on 26oct2016 to assess the testing instrument in question, and determined that the syringes were dirty. consequently, the fse removed and cleaned the syringes.><on 24oct2016, a customer reported an abo mistype when testing blood samples on a galileo neo on (b)(6) 2016. | 0
5926597 | on (b)(6) 2016, a (b)(6) customer site reported an unexpectedly negative antibody identification when tested on a galileo echo instrument, but it was only on 15aug2016 that sufficient information was obtained to determine event reportability.><by the time that adequate information was obtained on 15aug2016, the antibody identification panel test microstrips were already previously expired on 26jul2016. therefore no evaluation was possible related to this event. however, the lot of microstrips used on the galileo echo was previously already evaluated for k antigen status, which at that time performed as expected. | 0
5695725 | on (b)(6) 2016, a customer reported unexpected negative reactions on the 3_cell assay on the galileo echo during correlation studies.><a review of the images from m01332: batch 22641 (r725 and 221612) tested on wednesday, may 4, 2016 11:47, sample w14795 krohne 753794, all 3 cells were negative, cell 2 was visually weakly positive, cell 2 is e positive, controls worked as expected. technical communication cc 09-042-02 states that the echo may generate a negative result with capture-r plates, where upon subsequent visual inspection the cell appears weak positive or equivocal. customer was advised to perform a visual verification of negative reactions before final release of those well results. | 0
5864857 | on (b)(6) 2016, a customer reported an unexpected abo blood grouping mistype when testing a blood sample on a galileo, when tested on (b)(6) 2016.><immucor technical support assessed the instrument test well images in question using faxed copies of reports received (b)(6) 2016. an immucor field service engineer (fse) visited the customer site to assess the testing instrument in question on 27jul2016 and changed the reagent probe because of the presence of chips on the probe tip. the immucor laboratory tested retention product on (b)(6) 2016 which performed as expected. the immucor laboratory also tested the returned blood sample from the customers site on (b)(6) 2016. the immucor laboratory was unable to reproduce the customer's observations. the donor red blood cells tested negative with anti-b, as expected, and the full abo typing yielded ntd because of the lack of reaction with b reverse typing reagent red blood cells. | 0
5840316 | (b)(6) customer did not provide any additional information or documented instrument access. no blood sample or testing product was returned to the immucor laboratory for any testing.><on (b)(6) 2016, a customer site in europe ((b)(6)) reported an unexpected negative antibody screen when tested on a galileo echo instrument. | 0
5587356 | review of instrument images: batch (b)(4) screen, cells 1 and 3 are e+(homozygous). cell 1: negative, visually positive. cell 3: negative, visually positive. customer was referred to cc 09-042-02 which states that the echo may generate a negative result with capture-r plates, where upon subsequent visual inspection the cell appears weak positive or equivocal. customer was recommended to perform a visual verification of negative reactions before final release of those well results.><on (b)(6) 2016, a customer reported unexpected negative reactivity with capture-r ready-screen (3) (crrs 3) on the galileo echo instrument when testing a patient sample having a known anti-e. no adverse event occurred as a result of the unexpected negative screen. | 0
5739877 | review of results files for batch 22102. lots in use: r734/221628. sample id (b)(6) - echo reported all screen cells as negative. wells are visually positive for screen cells 1 and 2. screen cell 3 is visually negative. control well resulted as expected. customer was informed that negative wells that visually appear positive are addressed in cc-09-042-02. customer advised to perform a visual verification of negative reactions before final release of those well results.><on 24may2016 a customer reported unexpected negative results when testing a patient sample using capture-r ready-screen 3 (crrs 3) on the galileo echo instrument. the customer reports results are visually positive despite echo reporting them as negative. customer reports patient has an anti-d. | 0
5823262 | on (b)(6) 2016, a customer reported unexpected negative reactivity for a patient sample ((b)(6)) on a crossmatch assay on the galileo echo. the sample had a previous history of having an anti-c, anti-e and anti-jka. the donor is c=e=jka+. initial testing reported the sample out as incompatible as expected with 1+ reactivity. however, repeat testing for this sample with the same donor reported out as compatible. no adverse event occurred as a result of the unexpected result.><immucor technical support used a remote electronic access method to assess the testing well images in question: the camera report appeared optimal. batch 860 and 861 were run without error. batch 860 reported the testing with sample id: (b)(6) and donor id: (b)(6) as incompatible with 1+ reactivity as expected. the well displayed slight red cell adherence. the control well reported 4+ reactivity and displayed strong red cell adherence also as expected. batch 861 reported the testing with sample id: (b)(6) and donor id: (b)(6) as compatible. the test well reported a reaction strength of 3 and visually appeared negative. the control well reported 4+ reactivity and displayed strong red cell adherence as expected. the same reagent lots were used for testing both batches. other donors samples tested at the same time with these batches reported out as expected. based on the data provided, a reagent related problem did not appear likely as the testing reported out as expected during the initial run. a sample related problem appears likely. the customer was informed of the results but would like service come out. | 0
5695060 | a review of the image result files showed that reactions appeared as reported by the instrument. a service call was made and the instrument was found to be performing within specifications. customer does not know if the sample contained passively administered anti-d.><on 04may2016 immucor became aware of an unexpected negative result on the galileo instrument when performing a 2 cell screen on a patient sample. the sample contained an anti-d. | 0
5882428 | haemonetics received a complaint on 07/22/2016 stating that during a routine plasmapheresis procedure, there was bowl noise during the middle of the 3rd return. a centrifuge underspeed error was received and plastic shavings were observed on the top of the bowl. the donor did not get all rbcs returned, less than 200ml were in the bowl. after being disconnected from the procedure, the donor began experiencing symptoms of chest pain, trouble speaking, weakness and dizziness. the donor was diaphoretic, had slurred speech and was warm. vitals were taken and ems was called. the donor was transported to the local emergency room. outcome is unknown. the center has no additional information and the donor has not returned to donate since this incident. the bowl was returned for evaluation and the center kept the device in service.><the customer returned the bowl used in the procedure for evaluation. the pcs 2 plasma collection system was not evaluated as the customer continued to utilize the device in subsequent procedures with no issues. the bowl was visually evaluated and found the white ceramic seal was cracked. the crack was classified as a hairline crack which has been deemed to have a higher potential to cause functional issues. in this instance, the header shield and upper assembly was frozen upon receipt and there were plastic shavings found on bottom center register of bowl, which holds the bowl in place in the centrifuge. if the upper assembly is to freeze up during use, it is possible that the centrifuge chuck will continue to spin with the bowl not spinning, creating noise and shaving portion of the plastic bowl loose. with centrifuge chuck spinning the shavings will work their way to the top of the centrifuge well. the shavings were not found inside the bowl body and would not have the potential to enter the fluid pathway of the disposable set. | 0
5690864 | the night shift at a hospital clinical laboratory reported that the sta compact blood coagulation analyzer was not functioning as expected. the equipment was taken out of service until a company representative was able to perform a repair. the field service engineer found a very small crack in a valve which was replaced (suspect part was returned to company for evaluation). the facility reported that there was no injury to any patients or hospital personnel. the valve from the device has been checked by the manufacturer. the presence of a small crack in the valve was confirmed. this component does not have a history of cracking in this manner. efforts were made, using the subject valve and similar valves with various defects, to mimic the hospital's malfunction experience. because the results could not be duplicated, it is not completely certain that this malfunction and this cracked component are directly linked. this event is considered an isolated occurrence. | 0
5917336 | the ortho vision analyzer has functioned as per design. the root-cause of the discrepant negative grading could not be confirmed, however it could not be excluded that it is sample-related, the anti-cw (rh8) and non-specific antibodies being weak and/or at detection limit of the reagents and techniques used associated to donor sample preparation. no product failure is identified. a biased result was reported to physician. the patient was not harmed. (b)(4).><complaint reporter is reporting one discrepant negative grading for a crossmatch reaction in indirect antiglobulin test (iat) for one patient sample and one donor unit using ortho biovue&#59411;system igg cassette lot igc644a and 0.8% ortho red cell diluent lot rcd793a in conjunction with their ortho vision analyzer. the customer stated that the discrepant negative reaction was observed in column 1 of cassette id (b)(6) versus a repeat testing performed unintentionally for the same patient and donor using the same reagents where a positive reaction (0.5+ reaction strength) was observed for column 1 of cassette id (b)(6). the customer stated that images of these cassette ids are looking similar and therefore they expected to have the ortho vision analyzer grading them the same way (0.5+ reaction strength). the customer stated that antibody identification allowed to identify an anti-cw(rh8) antibody and a very weak non-specific antibody reacting with cell 2 of the identification panel for this patient. no further detail was provided. the customer stated that they consider both antibodies as not clinically significant and that therefore donor units do not require being antigen negative. the customer stated further that donor units were chosen on basis of patient's phenotype. no further detail was provided. the customer stated further the donor unit involved in this incident was not transfused to this patient. event date: (b)(6) 2016. reported on: (b)(6) 2016. the complaint reporter said that no patient was harmed as a result of the reported event. | 0
6142030 | a customer complained after observing discrepant abo reverse grouping results for one patient that were uploaded to their datavue middleware and not blocked. complainant/complaint reporter: mrs. (b)(6) ? laboratory technician. event date: (b)(6) 2016. reported on: (b)(6) 2016 by mrs. (b)(6) to the helpdesk . software version: 3.6.0. reagent information: ortho biovue system reverse cassette lot rdc018f exp. 11 august 2017. affirmagen 4 lot 4a929z exp. 06 december 2016. patient information: (b)(6). the customer said that their ortho vision biovue analyzer is connected to their laboratory information system (lis) through their datavue middleware. the customer said that, on (b)(6) 2016, they had tested a patient sample for abo forward and reverse grouping using ortho biovue abdd/k grouping cassette, ortho biovue system reverse cassette and affirmagen 4 in conjunction with their ortho vision biovue analyzer and that they had obtained: - for abo forward grouping, a positive reaction (with 4+ reaction strength) with the biovue anti-a reagent and a negative reaction with the biovue anti-b reagent. - for abo reverse grouping, negative reactions with all four cells of the red cell reagent. the customer said that an inconclusive abo grouping result was sent to their datavue middleware and stayed in the pending review/acceptance queue. the customer said that, on the same day, as they were troubleshooting the discrepant negative reactions obtained in abo reverse grouping, they made a new order request from their datavue to test the same patient sample for abo reverse grouping using ortho biovue system reverse cassette and affirmagen 4 in conjunction with the same analyzer and that they had obtained negative reactions with all four cells of the red cell reagent. the customer said that a ab group for abo reverse grouping result was sent to their datavue middleware and stayed in the pending review/acceptance queue. the customer said that, as they are suspecting the patient to be immunocompromised, they had repeated the abo reverse grouping test in tube method using the same red cell reagent with a longer incubation time at room temperature and that they had obtained a positive reaction with the b cell of the red cell reagent, confirming the group a of the patient. the customer said that no discrepant result was reported to the physician and that the patient was not harmed as a result of this event.><discrepant abo reverse grouping results for one patient. the root cause of the discrepant abo grouping result not being blocked by the analyzer is user related, the user sending orders for abo reverse grouping only, preventing the analyzer from comparing to an abo forward grouping. the root cause of the discrepant negative abo reverse grouping is sample-related, the antibodies being weak and/or at detection limit of the technique used. no product failure was identified. no biased result was reported. no patient was harmed. | 1
6141771 | discrepant negative grading in antibody identification test for one patient sample. the root cause could not be determined, although it cannot be excluded the negative reactions obtained are sample-related, the antibodies being weak and or at around the detection of the reagents and techniques used. no general product failure is identified. no biased result was reported to a physician. the patient was not harmed.><complaint reporter is reporting a discrepant negative grading in antibody identification for a patient sample in indirect antiglobulin test (iat) using biovue technique on their ortho vision biovue analyzer. complainant/complaint reporter: mrs. (b)(6)- laboratory technologist event date: (b)(6) november 2016. reported on: 08 november 2016 by mrs. (b)(6) to the helpdesk. software version: 3.6.0. reagents: ortho biovue system igg cassette lot igc013f exp. 04 december 2016. ortho biovue system neutral cassette lot and expiry date not provided. the 0.8% resolve panel c lot 8rc516 exp. 29 november 2016. patient information: sample id (b)(6). no further detail was provided. the customer said that, on (b)(6) 2016, they had tested sample id (b)(6) for antibody identification in iat technique using ortho biovue system igg cassette and 0.8% resolve panel c lot 8rc516 in conjunction with their ortho vision biovue analyzer and that they had obtained negative reactions with the eleven cells of the red cell reagent. the customer said that, when reviewing the reactions manually they would have graded the reaction obtained with cell 1 as weak positive (with 0.5+ reaction strength). the customer said that, on the same day, they had tested the same sample for antibody identification in enzyme technique using ortho biovue system neutral cassette and the same lot of red cell reagent in conjunction with the same analyzer and that they had obtained a positive reaction (with 2+ reaction strength) with cell 1 and negative reactions with the other cells of the red cell reagent. the customer said that they could identify an anti-cw(rh8) antibody. the customer said that no biased result was reported to the physician. the customer said that the patient was not harmed as a result of the reported event. | 0
5994042 | customer contacting cts to report 19 events of potential misread 1+ reactions associated to patient antibody screens as part of an internal validation study being performed in support of an ongoing audit investigation being conducted by the (b)(6). the validation testing was conducted with known previously identified antibodies initially tested on the provue, manually reviewed off of the service rack and then repeated by the manual gel method. eleven (11) out of the 19 events are reportable. eight (8) out of 11.><on 9-15-2016 an ocd field engineer (fe) arrived at the customer site and replaced the visor as it was worn, fe cleaned the led arrays and performed the gripper to the incubator adjustment. fe also verified centrifuge speed and timing which were in specification per procedure. replaced the card and barcode illuminators and readjusted the camera per procedure. the provue's final brightness is acceptable at 121, range is 101-128. reference image inspected and was acceptable per procedure. customer ran qc and accepted analyzer performance. no incorrect or erroneous results were reported as a result of this incident. there was no risk present at the time of the incident. there was no harm to any patient. the investigation determined that the most likely root cause of this event could potentially be a combination between the instrument and the weak antibody samples used during testing. | 0
5457754 | customer reporting false negative antibody screen for one sample. possible anti-d and anti-c. patient was transfused on (b)(6) 2015. ( # of units not provided) according to customer, sample in question was tested on provue on (b)(6) 2016 and reported as positive ( cell #1 =1+). during antibody id, abs screen was repeated using manual gel method, and 2+ positive was observed in cell #1 and 1+ positive in cell #2. abs screen was repeated on provue and gave negative reaction. customer is concerned with differences in abs reactions between provue and manual gel.><on (b)(4) 2016 an ocd field engineer (fe) arrived at the customer site and noticed a pinch/dent on teflon tubing during trouble-shooting. fe replaced teflon tubing, checked syringe, probe and wash station, and ran rinse/priming test through. customer passed qc and accepted service of the instrument. repairs have returned this instrument to expected operation. the investigation determined that the most likely root cause of this event was instrument related. no incorrect or erroneous results were reported as a result of this incident. there was no risk present at the time of the incident. there was no harm to any patient. | 0
5466226 | customer technical support personnel discussed the possible reason for the negative reaction in the anti-a and anti-d microwell; is due to the depth of where the provue aspirates red cells from the samples compared to where the user aspirates cells for manual gel samples along with the theory that donor transfused red cells are denser and would settle to the bottom of patient samples after centrifugation. while the operator would typically aspirate cells more towards the top of the tube collecting autologus red cells. (feb 15, 2016 1 unit of blood group o neg transfused to the patient). communication (b)(4) which discusses the possible reason of the false negative results. customer indicates their understanding of the potential cause of the reported concern and indicates their confidence that the root cause was sample related and not any defect in the gel card or the provue operations. customer is satisfied with the performance of the provue and requires no additional follow up. the investigation determined that the most likely root cause of this event was sample related due to the recent transfusion.><ocd rep contacting cts to report false negative reactions to the forward anti-a and anti-d microwell to the mts abd/rev gel card on the provue to a patient who consistently shows reactivity in these wells by the manual gel method. | 0
5446804 | on (b)(6) 2016 an ocd field engineer (fe) arrived at the customer site, inspected images from the log and found no cells appearing in the well. reader camera adjustment value is bright cam 0: 119. fe cleaned diffusion plate and visor, performed reader camera alignments and optics (focus and brightness) adjustments. post adjustments, bright cam0 value is: 119. fe created a new reference image and confirmed image is good. customer ran controls successfully. customer ran the same patient that resulted in false negative and results were again negative on the provue. fe inspected the gel card and it is clearly neg. fe returned to the customer site on (b)(6) 2016, performed the cavro volume verification 50-25-10ul volumes verification passed for all three volumes. fe confirmed reagent spinner disks work, measured incubator temperature for 37 degrees. it give 36.5. pressure at 2.6 psi and vacuum at 196 inh2o, all good, probe height, all good. cts stated that after all the successful verifications made on analyzer, provue works as intended. manual and provue method are a little different as igg cards are already at 37 in provue vs room temp for manual (cold agglutination possibilities). centrifugation speed is also different. a customer letter will be issued to discuss testing in provue vs. manual. the root-cause could not be confirmed although the most probable root cause is associated with an anti-s antibody being weak and/or at the detection limit of the technique and reagents used. no incorrect or erroneous results were reported as a result of this incident. there was no risk present at the time of the incident. there was no harm to any patient.><ocd rep contacting cts on behalf of the customer who has reported false negative reactions to a patient's antibody screen tested on the provue but was later identified to have an anti-s which reacted by manual gel . two samples from the same patient were tested on provue on (b)(6) 2016. 1 of 2 | 0
5695239 | on 5-4-2016 an ocd field engineer (fe) arrived at the customer site and verified the gripper and camera adjustments. the fe also performed a health check on the fluidics. the customer ran controls and accepted all results confirming them to all be in range. repairs have returned this instrument to expected operation.><the customer reported that the ortho provue missed an antibody that was detected in manual gel. no incorrect results were reported. incident 1 of 2. (b)(4). | 0
5446819 | ocd rep contacting cts to report a false negative result in the anti-d micro well to the mts abd/rev gel card to a patient that has had a previous rh pos typing. customer reports the patient sample was tested on the provue to on (b)(6) 2016 and a blood group of o neg was issued based on the negative result obtained in the anti-d well. the patient in question received 1 unit of rh negative pack cells on (b)(6) 2016 and another on (b)(6) 2016. customer states that they have rerun both manually (tubes) and manual gel and they got a 4+ reaction. customer notes a small haze that they feel the camera should have caught.><on (b)(6) 2016 an ocd field engineer (fe) arrived at the customer site and discussed the issue with customer. fe sent result image to cts who confirmed that the image is negative. fe did not make any changes to the provue as no malfunction was identified. fe ran wad diagnostic successfully. customer ran qc, results were acceptable. customer reran patient sample on provue and received dual population which was explained by cts as: the more you test a sample, the more of the transfused cells will be used up at the bottom of the tube. if the sample tube is being tested multiple times then the transfused cells will be removed and all will be left is patient's own cells. the investigation determined that the most likely root cause of this event was sample related due to the recent transfusion. | 0
5477703 | ge healthcare's investigation is ongoing. a supplemental report will be submitted when the investigation has been completed.><it was reported that a primary lesion marker for a breast ultrasound scan allegedly moved posterior unexpectedly on a subsequent reading of the study. in this particular case, the reading physician was the same physician who performed the study and therefore the issue was obvious to the user. it is unknown whether or not the reading device was manufactured by ge healthcare. there was no report of patient injury. | 0
5607995 | while abbott field service engineer (fse) was on site one of the lab employees got water/bleach solution in her eye. the employee immediately rinsed the eye with water and with eye solution. the employee then went to the emergency department where they continued flushing the eye. the employee was then given eye drops to continue using. this is considered an adverse event for any treatment of an eye for injury or exposure to biohazardous / potentially biohazardous material beyond flushing of the eye with water.><the device evaluation was reassessed and concluded that a malfunction occurred, therefore the device was not performing as intended.><(b)(4). a laboratory employee was splashed in the eye with a water/bleach solution while assisting an abbott field service representative (fsr) performing troubleshooting / maintenance on architect c1600375. the employee was not using any ppe (personal protection equipment) such as goggles or a face shield. a review of the architect c1600375 service history found no similar issues or related complaints. a historical review of complaints did not identify any other complaints for similar incidents. a review of product labeling shows adequate precautionary warnings and information necessary to ensure safety and the safe use of the architect c1600375, including use of personal protective equipment (ppe). based on this investigation, neither a malfunction nor deficiency was identified for architect c1600375. | 0
5471773 | in addition to the on-site investigation performed by field service, the evaluation included reviewing the architect c4000 analyzer serial (b)(4) instrument logs and service history, historical data and associated labeling. the customer replaced the 1 ml syringe which had evidence of leaking without full resolution. field service (fs) identified a misaligned/loose peristaltic head tubing as the likely cause. fs replaced the peristaltic head tubing which resolved the issue. a review of the analyzer service history identified no additional contributing factors on or around the date of the complaint. a review of the 12 month search for similar complaints identified no adverse trends of either the 1 ml syringe or the peristaltic head tubing related to the current complaint issue involving discrepant erratic results. a review of archietct c4000 tracking and trending data revealed no systemic issues or adverse trends related to issue described in this complaint. the architect c4000 erratic result rate is within acceptable limits with no trends identified. the architect system operations manual provide information with regard to specimen handling, maintenance, and troubleshooting the reported issue. the issue was resolved by replacing the used parts in accordance with current product labeling. based on the evaluation performed neither a deficiency nor a malfunction were identified.><the account generated elevated magnesium processed on architect c4000 analyzer for 2 patients. patient 1 generated architect magnesium of 5.8 meq/l that repeated 2.4, 2.2, 2.1 meq/l. patient 2, sample id (b)(6) generated elevated architect magnesium of 12.1 meq/l that repeated 1.6, 1.7, 1.6 meq/l. the account uses a normal reference range of 1.3 to 2.1 meq/l. no impact to patient management was reported. no specific patient information was provided.><patient 3 generated architect magnesium of 4.9 meq/l that repeated 1.7, 1.7, 1.6 meq/l. no impact to patient management was reported. no specific patient information was provided.><(b)(). (b)(4). a followup report will be submitted when the evaluation is complete. evaluation in process.><the device evaluation was reassessed and concluded that a malfunction occurred, therefore, the device was not performing as intended at the customer site. however, no systemic issue and/or product deficiency was identified. | 0
5912780 | the device evaluation was reassessed and concluded that a malfunction occurred, therefore, the device was not performing as intended at the customer site. however, no systemic issue and/or product deficiency was identified.><multiple patients: id (b)(6) is (b)(6) year old male and id (b)(6) is (b)(6) year old male. an evaluation is in process. a followup report will be submitted when the evaluation is complete. evaluation is in process.><the account generated false elevated magnesium on 2 patient samples that repeated in the normal range when processed on the architect c16000 analyzer. id (b)(6) generated magnesium of 2.50 but repeated 0.74 mmol/l. id (b)(6) generated magnesium 2.55 but repeated 0.86 mmol/l. the account uses a normal reference range of 0.7 to 1.1 mmol/l. no impact to patient management was reported.><the abbott field service representative (fsr) was dispatched to troubleshoot the issue. when observing subsequent instrument runs, the fsr noted erratic mixer vibration. the fsr replaced the r1 b and r2 b mixers to troubleshoot the issue and assigned the mixers as likely cause for the elevated magnesium issue. there were no additional discrepant results reported on the architect c16000 analyzer after the mixers were replaced. an instrument service history review revealed no additional erratic or discrepant results reported on architect c16000 analyzer, serial (b)(4), nor did it identify any service or complaint issues that may have contributed to this issue. the calculated erratic rates for the architect c16000 systems were below the upper control limit. additional metrics review found no systemic issues or adverse trends for the issue associated with erratic/discrepant results. a review of historical data for the part associated with this complaint revealed no adverse trend or systemic issues associated with the mixer. a review of the product labeling concluded that the issue is sufficiently addressed. based on the available information and evaluation the architect c16000 analyzer is performing acceptably. | 0
5913522 | an evaluation is in process. a follow-up report will be submitted when the evaluation is complete.><an evaluation was performed by reviewing the c802159 service history, complaint text, other customer complaints for similar issues, and a review of labeling. an abbott field service representative (fsr) visited the account and indicated the customer performed quarterly maintenance to include replacing the worn out o-ring, sample/wash syringe. the complaint text stated that cuvette nozzle #1 was clogged with paper and could be related to the result issue. the customer cleaned the clogged nozzle and performed the as needed procedure 2134 (change water bath). a review of the architect c802159 service history identified no contributing factors on or around the date of the complaint. a search for similar complaints did not identify any trend of the o-ring, sample/wash syringe ((b)(4)), the nozzle #1, #4, and #5 (rohs) (pn 7-93256-02), or the architect c8000 related to the complaint issue of discrepant results. the architect system operations manual addresses component replacement and troubleshooting for falsely elevated results. the corrective actions include but are not limited to: examine samples for fibrin or other large particles, rerun the sample, replace the syringe o-ring and seal tips, perform monthly maintenance procedure 6018 clean cuvette washer nozzles, and contact your area customer support. based on this investigation no malfunction or deficiency of the o-ring, sample/wash syringe ((b)(4)), the nozzle #1, #4, and #5 (pn 7-93256-02), or the architect c8000 were identified.><the device evaluation was reassessed and concluded that a malfunction occurred, therefore, the device was not performing as intended.><the customer stated that the architect analyzer generated falsely increased calcium results on one patient. the results provided were: sid (b)(6) on (b)(6) 2016, initial = 2.03mmol/l / repeat = 5.32 / third run = 2.03mmol/l (normal range 2.10-2.55mmol/l). there was no reported impact to patient management. there was no additional patient information provided. | 0
6110835 | investigation of the customer issue included a review of the complaint text, a search for similar complaints, a review of labeling, and instrument service. no returns were made available from the customer site for this evaluation. no adverse trend was identified for the customer issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation, the analyzer performed as intended and no product deficiency was identified. this issue was previously reported under mdr number 1628664-2016-00227 under a different suspect device.><the device evaluation was reassessed and concluded that a malfunction occurred, the device did not perform as intended. however, no systematic issue or product deficiency was identified.><the customer observed false elevated magnesium patient results generated using the architect c8000 analyzer. the following data was provided. the customer uses normal range 0.7 to 1.0 mmol/l. sid (b)(4) initial greater than 3.9, repeat 0.65 results other analyzer (serial number (b)(4) 0.6. no impact to patient management was reported. | 0
5591745 | (b)(4). an evaluation is in process. a follow-up report will be submitted when the evaluation is complete.><concomitant medical products and therapy dates added the lot number 151023062 of the ict module. an evaluation was performed by reviewing the complaint text, other customer complaints for similar issues, a review of labeling, and a review of historical records. the architect serial number (sn) c401617 instrument logs confirm the falsely depressed potassium, sodium, and chloride issue. the history log shows ict module 151023062, expiration date 23jul2016, was installed on 13jan2016 and was within the use warranty period of 3 months after installation. the ict module performance between 29mar2016 and 01apr2016 was stable and within acceptable limits. a review of service history for architect sn (b)(4) did not reveal any possible causes and no additional reports for erratic or discrepant assay results have been received. a review of all complaints for ict module lot 151023 did not identify any other complaints for similar incidents. a review of all ict module complaints did not identify any related issues or trends. a review of complaints for architect c4000 systems revealed no issues or trends related to the current complaint. a review of product labeling shows adequate information with regard to maintenance, component replacement, specimen collection and handling, result interpretation, and troubleshooting for the reported issue. based on this investigation, neither a malfunction nor deficiency was identified for architect c401617 or ict module 151023062.><the device evaluation was reassessed and concluded that a malfunction occurred, therefore the device was not performing as intended.><the customer stated that the architect analyzer generated a falsely decreased potassium (k) result on one patient. the results provided were: (b)(6) - initial results k=1.6mmol/l / k=4.0 (within customer's normal range). there was no additional patient information provided. there was no reported impact to patient management. | 0
5608547 | an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. (b)(4).><abbott personnel advised the customer to replace and calibrate the reagent probe 1 on the architect c8000 analyzer. subsequent instrument operations and test results were acceptable. the part was worn from normal use. a review of complaint tracking and trending metrics was performed and identified no adverse trends in conjunction with the complaint issue currently under evaluation. the architect systems operation manual provides information to address the current customer issue. based on the available information from the customer site and from the results of this evaluation, there is no evidence to reasonably suggest a product malfunction occurred. the issue was addressed through standard troubleshooting procedures.><the customer reports the following potassium assay results generated on an architect c8000 analyzer: patient a: initial result of 10.0 mmol/l; sample retested ten times with a mean of 4.0 mmol/l and a %cv = 3.22. patient b: initial result of 8.0 mmol/l; sample retested ten times with a mean of 4.5 mmol/l and a %cv of 7.5. no suspect results were reported from the lab with no reported patient impact.><the device evaluation was reassessed and concluded that a malfunction occurred; therefore, the device was not performing as intended and the evaluation codes were corrected. | 0
5658006 | further investigation of the customer issue included a review of the complaint text, a search for similar complaints, review of instrument logs, a review of labeling and instrument service. no adverse trend was identified for the customer's issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures by replacing the used misaligned and loose mixer. based on all available information and abbott diagnostics' complaint investigation, the analyzer performed as intended and no product deficiency was identified.><the device evaluation was reassessed and concluded that a malfunction of the mixer was identified, therefore, the device was not performing as intended, however, no systemic issue or product deficiency was identified. the issue was resolved by replacing the used misaligned/loose mixer.><the customer observed falsely elevated creatinine results on the architect c16000 analyzer. the following data was provided: (b)(6) initial 1.55 mg/dl, repeat 2.61 mg/dl. there was no impact to patient management reported.><an evaluation is in process. a follow up report will be submitted when the evaluation is complete. (b)(4). an evaluation is in process. | 0
5907160 | manufacturer name, city and state, corrected.><the device evaluation was reassessed and concluded that a malfunction occurred; therefore, the device was not performing as intended .><during troubleshooting, the customer observed that the mixer paddles (list 09d59-02) were bent. the mixer paddles were not able to enter the cuvettes for mixing. the paddles were replaced. a review was conducted of the architect c16000 instrument logs and did not find any cause for the bent paddles. the reaction graph for the initial low result does not show an increase in absorbance after the addition of the r2 reagent indicating the lack of mixing due to the bent paddles. after the paddles were replaced, control samples were tested and met specifications. subsequent instrument operations were acceptable. a review of the architect c16000 analyzer's (serial number (b)(4)) service history identified no additional contributing factors on or around the date of the issue. no additional result issues have been reported since replacing the bent mixing paddles. a review of complaint tracking and trending metrics was performed and identified no adverse trends in conjunction with the complaint issue currently under evaluation. the architect system operations manual provides information to address the current customer issue. based on the available information from the customer site and from the results of this evaluation, there is no evidence to reasonably suggest a product malfunction occurred. the customer resolved the issue by performing troubleshooting and replacing the bent mixers in accordance with current product labeling.><the customer reports that one patient sample generated an initial clin chem magnesium assay result of 0.29 mmol/l / 0.71 mg/dl on an architect c16000 analyzer (the typical normal reference range for this assay is 0.66-1.07 mmol/l or 1.6-2.6 mg/dl). the patient was administered magnesium based on this decreased result. a subsequent blood sample generated a magnesium result of 1.57 mmol/l / 3.82 mg/dl. the initial sample was then retested and generated a result of 0.69 mmol/l / 1.7 mg/dl. no further patient information was made available. there was no impact to patient management reported. | 0
5698313 | customer reports that a mini drawer on a pyxis anesthesia system 4000 did not function properly causing medications in other drawers to be unavailable. the customer was able to retrieve mediations from a nearby device. no patient harm reported as a result of this incident.><carefusion received a medwatch (sus mw5061518) on may 9, 2016 from (b)(6) of (b)(6). the report date on the medwatch lists (b)(6) 2016. on (b)(6) 2016 (b)(6) from (b)(6) hospital in (b)(6) opened a case with carefusion documenting the same incident. a carefusion field service technician went on site attempted to replicate the issue without success. the mini drawer functioned correctly. neither report alleges patient harm. | 1
5878646 | a facility reported that they were not connecting the medivators hookup air/water channel connector to the endoscope when using a dsd 201 automated endoscope reprocessor, thus flushing, high level disinfection, and/or rinsing of this channel was potentially not achieved. there is potential for patient cross contamination or chemical exposure due to improper reprocessing of endoscopes. medivators sales representative was informed that this facility was instructed by an endochoice sales representative, reprocess the endoscope without the air/water connection of the hookup attached because they do not need to attach this channel." this facility reached out to medivators for support. medivators sales representative informed them that there was no way to validate flow to this channel without appropriate connection. visual inspection of all connections is required per aer/hookup ifu. since the initial contact they have ordered a new hookup and they are reprocessing with all channels appropriately connected to the aer and endoscope. to date there have been no reports of illness of injury to patients or handlers. this complaint will continue to be monitored within the medivaotrs complaint system.><a facility reported that they were not connecting the medivators hookup air/water channel connector to the endoscope when using a dsd 201 automated endoscope reprocessor thus flushing high level disinfection and/or rinsing of this channel was potentially not achieved. there is potential for patient cross contamination or chemical exposure." | 0
5651052 | the case states that the facility is not appropriately pre-cleaning and manual cleaning endoscopes prior to reprocessing in the medivators dsd edge automated endoscope reprocessor.><medivators clinical specialist (cs) visited this facility to perform an inservice on the operation of the dsd edge machine. the cs reported the following observations: in review of sgna guidelines, the facility has never manually cleaned a scope in a sink of h2o and detergent in 12 years. it was observed that the facility used same water all day to leak test scopes. they reported they had never flushed a scope in 12 years. facility reusing olympus maj-855 auxiliary water tube without reprocessing between patients. there is potential for patient cross contamination according to the dsd edge ifu, endoscopes must be manually cleaned prior to reprocessing in aer. medivators cs recommended that they familiarize themselves with society guidelines for reprocessing as well as directed them to review their scope manufacturers manual cleaning procedures and reach out to their scope manufacturer representative for additional guidance. the importance of appropriate manual cleaning prior to placing in the aer for hld was stressed. to date, there is no reported patient illness or injury. this complaint will continue to be monitored within medivators complaint system. | 0
6071238 | the facility reported that they did not have a 0.2 micron absolute bacteria retentive water filter as required in the water filtration system for a medivators cer automated endoscope reprocessor. there is potential for contaminated water being used in the reprocessing of endoscopes, thus potential for waterborne bacteria to be left on the scope during the rinsing cycle. this could result in patient infection. medivators technical service has been in contact with this facility and they have been informed of the cer filtration requirements. medivators has shipped them the filtration system for their cer machines for them to install. to date, there have been no reports of patient illness or injury. this complaint will continue to be maintained within the medivators complaint system.><the case states that the facility did not have the 0.2 micron absolute bacteria retentive water filter as required in the water filtration system for a medivators cer automated endoscope reprocessor. there is potential for contaminated water used in the reprocessing of endoscopes and potential for waterborne bacteria to be left on the scope during the rinsing cycle, which could potentially affect the patient. | 0
5745787 | first of 12 cases of isolation of candida lipolytica from bronchoscopically obtained respiratory (n=6 via olympus bf1th190) or mediastinal lymph node needle aspirates (n=6 via olympus ebus bfu180f) from 12/30/2015 to 05/23/2016 to date. eight of the 12 bronchoscopes, including the first 7 cases, were obtained from bronchoscopes that had been reprocessed on the same olympus oer-pro endoscope reprocessor, one of 5 such devices in use at our facility. no evidence of clinical infection associated with any of these (b)(6) cultures. no other (b)(6) clinical cultures, including those obtained endoscopically, no (b)(6) surveillance cultures from selected olympus bf1th190 therapeutic bronchoscopes and olympus bfu180f ebus scopes after reprocessing or from monthly surveillance cultures of olympus duodenoscopes or linear array ultrasound endoscopes. | 0
5949028 | near the end of (b)(6), we inadvertently discovered that the server/software designed to capture the reprocessing data of our medivators does not function consistently on two of our machines at (b)(6) medical center (b)(6). as a result we are missing reprocessing data on an undetermined number of dates, times and pieces of equipment which data we are required to maintain. we are concerned that the vendor was aware of the potential issue and failed to disclose this to us prior to purchasing the server and software package to go with the automated endoscope reprocessor. when contacted to come and help with the issue the vendor disclosed that they had experienced similar issues and suggested that we obtain the software upgrade as soon as its available. it was also newly suggested that the medivator machines be rebooted weekly, an action that can only be accomplished by is upon request from the department each week. | 0
5588462 | (b)(4). hazy smoke with a burning odor was observed from an infrequently used tdxflx during assay calibration with the access door open. a review of service history for the tdxflx, serial number (b)(4), did not identify any issues that may have contributed to the current complaint. a review for similar complaints against the tdxflx analyzer found no similar issues or trends identified. a review of the tdx/flx system operations manual gives adequate instructions for shut down of the analyzer and also advises keeping all access doors closed to prevent damage to the air heater and the photomultiplier tube while running. the ambassador configured the system 2 parameters to disallow assay runs with the access door open. additionally, a review of the 2016 ul certification memo contains adequate information noting abbott diagnostic equipment and accessories are certified to the appropriate safety standards, and adequate protection is provided for the operator against spread of fire from the equipment. based on this investigation no systemic deficiency or malfunction was identified for the tdxflx (list number 04a24-96), serial number (b)(4).><the device evaluation was reassessed and concluded that a malfunction occurred, therefore the device was not performing as intended, and conclusions code was corrected from (b)(4).><the customer observed smoke coming from the sample carousel area on the tdxflx analyzer. the customer stated that the smoke dissipated quickly and a burning odor was present. the customer stated that the analyzer is not used frequently. no flames were observed and no injury was reported. | 0
5952658 | an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. an evaluation is in process.><an abbott field service engineer was dispatched to the account and found a visibly damaged charred component on the vpm board (vacuum pressure supply flow part 8960122101) on the cell-dyn ruby analyzer. the vpm board contained a burnt area at the joint of the cr1 cathode and j1, pin1 (positive side of the solenoid). the issue was limited to this part and did not spread to other parts of the equipment. the cell-dyn ruby was repaired and the analyzer was returned to normal operation. investigation of the customer issue included a review of the complaint text, a search for similar complaints, a review of labeling, a review of instrument service, and examination of the returned part. returns were made available from the customer site for this evaluation. abbott engineers examined the damaged board. the burnt area was due to overheating of the cr2 (1n4003, flyback diode) that was due to arcing and/or voltage spike imposed on cr2. the arcing and/or voltage spike could have been due to either the solenoid cable that connects to j1 being removed while q1 (switching power transistor) was turn on, and the solenoid coil was energized or the solenoid valve was power cycled by the flow script as normal and the cr2 failed to absorb the stored energy dissipated from the coil while q1 was turn off. the driver circuitry was examined and no other issues were found on the same printed circuit board. cell-dyn ruby analyzer serial number 34641bg service history was reviewed and no contributing factors were found. no subsequent reports of smoke/burn issues were identified. historical complaint data was reviewed and no adverse trend was identified. product labeling was reviewed and found to be adequate. the issue was resolved through normal troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation for the cell-dyn ruby analyzer, no product deficiency was identified.><the customer reported that the cell-dyn ruby analyzer did not operate. the customer heard a sound and they smelled a burnt odor. error code 2077 (flow script timeout) was observed. the screen initially displayed, then shut down. the customer disconnected the analyzer from the power outlet. during repair by abbott service fse personnel, a visibly damaged charred component on the vpm board was observed (vacuum pressure supply flow part (b)(4)). no fire or injury was reported. | 0
5617376 | (b)(4). an evaluation is in process. a followup report will be submitted when the evaluation is complete.><a review of the complaint information showed the instrument was in ready mode, and the customer stated that she replaced the waste container without waiting for the system initiated message (sim), waste container full". additionally personal protective equipment (ppe) was not worn by the operator. a review of the product labeling concluded that the issue is sufficiently addressed; specifically the cell-dyn sapphire operations manual under: principles of operation under halt behavior associated with sims hazards under operator responsibility hazards under biological and chemical hazards troubleshooting and diagnostics under 0425 waste full. a review of historical data including complaint history review service history review and tracking and trending showed no adverse trends or product issues. this event was determined to be due to use error. no product deficiency was identified for the cell-dyn sapphire.><an operator was replacing the waste container on the cell-dyn sapphire when after disconnecting the tube the cell-dyn sapphire started an automatic flush and splashed in her eyes. the operator was diagnosed with microburns and chemical keratitis with treatment of eye drops and intraocular antibiotic ointment. the operator was not wearing personal protective equipment." | 0
5466301 | the customer stated that the cell dyn sapphire analyzer generated a decreased hemoglobin (hgb) compared to a blood gas analyzer. the results provided were: sapphire = 7.1 / blood gas analyzer 10. there was no reported impact to patient management. there was no additional patient information provided .><(b)(4). a follow-up report will be submitted when the evaluation is complete.><an evaluation was performed by reviewing the complaint text, other customer complaints for similar issues, accuracy testing, and a review of labeling. the customer stated that the cell-dyn sapphire analyzer, serial number (b)(4), generated decreased hemoglobin results compared to a blood gas analyzer. a review of customer complaints received for this issue did not identify any trends or increased complaint activity. a review of the submitted data and information showed dilution dependent parameters, like rbc (1.9 10e6/ul) and mcv (115 fl) are exactly the same which shows, the instrument is performing well. there is a difference in technology between a blood gas counter and the cell-dyn sapphire analyzer. in this case, the difference between the cell-dyn sapphire hgb result and the hgb of the blood gas analyzer has no major impact on medical intervention since both results were below the hgb normal range. a review of the cell-dyn sapphire operator's manual was performed and found to adequately address the issue. based on available data and information, the cell-dyn sapphire, list number 08h00-01, serial number (b)(4), is performing acceptably with no product deficiency identified. | 0
6099054 | the customer observed smoke coming from the top left of the cell-dyn sapphire analyzer. the customer returned the motion control power supply, part number 8405700302 to aid in the investigation. a review of complaints for this issue did not identify any trends or complaints associated with the part number or the instrument base. examination of the returned part found no specific or visible burnt area in the electrical components inside the power supply; however, there was an excessive amount of dust/dirt inside the power supply. a review of labeling shows the cell-dyn sapphire operator's manual addresses power specifications, electrical hazards, and troubleshooting of the issue. based on this investigation no product deficiency was identified for the cell-dyn sapphire, serial number (b)(4), or the motion control power supply assembly, part number 8405700302.><the customer observed smoke coming from the top left of the cell-dyn sapphire analyzer. the instrument was powered down and disconnected from the power source. there was no personal injury, impact to patient management, or facility damage reported.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. (b)(4). | 0
5677537 | evaluation of the customer issue included a review of the complaint text, a search for similar complaints, review of instrument logs, a review of labeling and instrument service. no returns were made available from the customer site for this evaluation. no adverse trend was identified for the customer's issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation, no product deficiency was identified.><(b)(4). an evaluation is in process. a follow-up report will be submitted when the evaluation is complete.><the customer observed falsely elevated white blood cell (wbc) results generated using the cell-dyn ruby analyzer. the following data was provided (10e3/ul). (b)(6): initial 42.4, repeat 23.6, 20.6. no impact to patient management was reported. | 0
5822684 | the customer results were reviewed and suggest that there was an analytical issue with the hgb flow pathway of the instrument. there are several factors that could affect the hemoglobin (hgb) results and these factors could not be eliminated. first, blockage in the hgb flow system due to salt build up. second, leaks in lines around hgb flow cell. third, a dirty flow cell due to organic build up. field service replaced the bubble mix tubing due to salt build up resolved the complaint issue. the cell-dyn 1800 system operator's manual provides troubleshooting information to assist in problem identification, isolation, and corrective action. the complaint incident was an instrument specific issue; therefore, a product deficiency for the cell-dyn 1800 analyzer and the silicon tubing was not identified.><the account generated a falsely depressed hemoglobin of 4.9 g/dl processed on the cell-dyn 1800 on a patient sample that repeated 11.0 g/dl. no impact to patient management was reported. no specific patient information was provided.><(b)(4). a followup report will be submitted when the evaluation is complete. | 0
5474863 | the customer observed falsely depressed wbc results on the cell-dyn sapphire analyzer. the following data was provided: wbc initial 0.016, repeat 5.94 k/ul . 2nd patient wbc initial 0.009, repeat 4.80 k/ul . there was no impact to patient management reported.><(b)(4). further investigation of the customer issue included a review of the complaint text, a search for similar complaints, review of instrument logs, a review of labeling and instrument service. no adverse trend was identified for the customer's issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures; replacements of the worn out mix motor sensor assembly, leaking wash probe block, and preventive maintenance resolved the complaint issue. based on all available information and abbott diagnostics' complaint investigation no product deficiency was identified. | 0
5677573 | the customer observed low hemoglobin (hgb) results for two patients while using the cell-dyn emerald analyzer. the customer normal range is 12.0 to 16.0 g/dl. patient 1: sid (b)(6): initial 5.6, repeat 5.6. repeat (no sid) at hospital 12.7. previous result, 11.9, from (b)(6) 2015 was also provided. the patient was sent to the hospital after the initially low values, however, no treatment was given. patient 2: no sid: initial 11.5. repeat (no sid) at hospital 7.4. the patient was sent to the hospital after the initial result was obtained, however, no treatment was given. no impact to patient management was reported.><the customer suspected a preanalytical issue, such as sample handling and collection, may have occurred during collection of these samples. initial and repeat results were generated with different sample draws. the diagnosis and treatment of the patients was not provided in the complaint information. therefore, a definitive cause for the different hemoglobin results could not be determined. the true hemoglobin value for these specific patient samples could not be verified. investigation of the customer issue included a review of the complaint text, a search for similar complaints, and a review of labeling. no returns were made available from the customer site for this evaluation. no adverse trend was identified for the customer's issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures. based on all available information and abbott diagnostics' complaint investigation, the analyzer performed as intended and no product deficiency was identified.><(b)(4). an evaluation is in process. a follow-up report will be submitted when the evaluation is complete.><the device evaluation was reassessed and concluded that a malfunction occurred, the device did not perform as intended. however, no systematic issue or product deficiency was identified.><the customer observed low hemoglobin (hgb) results for two patients while using the cell-dyn emerald analyzer. the customer normal range is 12.0 to 16 g/dl. patient 1: (b)(6): initial 5.6, repeat 5.6, repeat (no sid) at hospital 12.7, previous result, 11.9, from (b)(6) 2015 was also provided. the patient was sent to the hospital after the initially low values, however no treatment was given. patient 2: no sid: initial 11.5, repeat (no sid) at hospital 7.4, the patient was sent to the hospital after the initial result was obtained, however no information about treatment was provided. no impact to patient management was reported. | 0
5884483 | it is suspected that a fire may have originated inside the analyzer after a sysmex (b)(4) representative observed wires that appeared to be burnt. the analyzer is being sent to the manufacturer for further investigation to determine the root cause of the event.><a user of an xn-10 analyzer in china reported a fire starting inside the analyzer. the user noticed the fire immediately and used a fire extinguisher to put it out. no one was injured or required treatment during this event. this event is being reported because of the potential for a serious injury if this incident were to reoccur. the investigation is ongoing. a root cause has not been determined.><final investigation determined that the master valves (mv) that control the movement of reagents through the system developed a leak over time. the design of the seal within the master valves may not have allowed for sufficient control of fluids within the valve. the protective sheet intended to cover electronic components was misaligned, allowing salty reagent solutions to reach electronic components. root cause was a sequence of events that caused a salt bridge at the solenoid valve (sv) connector that allowed current to continue flowing, causing heat and ignition. the solenoid valve wiring bundle was routed from above, and lay on top of the protective sheet, so leaking reagent traveled down the wiring to the solenoid connector. this was the ignition point. a field corrective action will be performed on all xn-series analyzers. service engineers will be visiting each site to align the protective sheet, reroute the wiring bundle under the sheet, clear any salt deposits from electronic parts if present, and replace any leaking master valve seals. additionally, any master valves that may be susceptible for leakage will be replaced. | 0
5696219 | the xn series administrator's guide (ag), chapter 5 - appendix, section 5.2 transportation controller connections (ct-90), subsection 5.2.1 explains the flow of data of the system. the first step is to determine the conveying destination. the rack containing samples is placed in the start yard (st-40) where the barcode terminal (bt-40) reads the sample/ rack label and sends the sample position information to the transportation controller (ct-90). the ct-90 sends the sample position information to the host computer. the host computer sends the sample analysis order by rack to the transportation controller. the bt-40 reads the barcode label at the carry-out position and sends the rack number to the ct-90. the ct-90 identifies the rack and determines the conveying destination. " "the second step is the analysis of the sample. the ct-90 sends an order to the xn conveyor (cv-50) for a rack to be conveyed. the cv-50 sends the order for the rack to be conveyed to an analyzer. the analyzer identifies the sample from the order and the host computer is queried for the sample information. the host computer sends the sample analysis order. the analyzer performs analysis based on the received order and sends the analysis data to the host computer and the analysis result to the cv-50. this is repeated until there are no more samples to be analyzed. the cv-50 sends the analysis result to the ct-90. the ct-90 sends a query for additional orders." summary: barcodes read: patient b - barcode correctly read at 20:25:39 placed in rack 30 tube position 1 downloaded patient b demographics - per analyzer host log. patient a - barcode correctly read at 20:25:59 placed in rack 30 tube position 1 was not logged in to the lis so no patient demographics were downloaded - per analyzer host log. results generated: results from patient a were generated at 20:26:57 and run in rack 30 tube position 9 - per analyzer printout. results from patient b were generated at 20:31:45 and run in rack 30 tube position 1 - per analyzer printout. rerun information: patient a sample was repeated at 21:00:38 - a sample mixing error occured - per analyzer printout note: these results were erroneous and reported. patient a sample was rerun the morning of (b)(6) 2016 - data was not provided. per user results were compatible with first run at 20:25:59. the discrepancy in the order of barcode read location in the sample rack and order of report generation is still under investigation. confirmation is needed to rule out operator error or a communication issues between the analyzer barcode terminal and laboratory information system (lis). the analyzer runs tubes in the order that they are placed in the sample rack. patient b results generated 5 minutes after patient a indicating that this specimen was placed further down the rack. the host log indicates other patient samples in positions 2 - 8. at this time we cannot rule out an operator error such as moving a sample tube after being read by the bt-40 or manually changing the demographics. confirmation of which patient required a corrected report has also been requested.><sample id (b)(6) were run on the xn-10 analyzer. barcodes were read correctly. the same patient demographics downloaded to both samples. the user reported that the sid belonging to patient b was applied to patient a's sample. they verified the patient information in their laboratory information system (lis) was correct. the user stated that the demographic information downloaded was the correct demographics for patient b. review of the host log indicates that when patient a was first run it was unregistered and no order with demographics were assigned to the sample. this sample was run first. an analyzer printout of the results was provided with patient demographics belonging to patient b. it is unknown how the incorrect demographics became associated with patient a's sid. patient b was run next and results were reported. patient a was run again in the manual mode and a mixing error occured. the incorrect demographics from patient b were once again assigned to patient a and these results were reported. a corrected report was later issued. it is unknown at this time which patient results were erroneously reported. this incident is under investigation. a follow-up report will be sent when the investigation is complete. this incident is reported because a communication issue or operator error can not be ruled out as a cause of the incorrect demographics. no harm to the patient was incurred.><further investigation determined that the root cause was user error. when the barcode is read at the bt-40 the laboratory information system (lis) provides the test order and patient demographics for the sample. the user incorrectly moved a sample tube to a different position in the rack after the rack passed through the barcode terminal (bt-40) of the conveyer system. the problem was compounded by the operator not registering the sample in their host computer so no order information was available for the sample tube that was present on the host computer at the time of analysis. the barcodes were read correctly at both the analyzer and the barcode terminal. the xn software will be improved to change the communication between the bt-40 and xn analyzer so when the xn-10 reads an unexpected barcode number for a sample that was not read by the bt-40 it will not associate any patient demographic information to the sample. there is a software version update to mitigate operator errors due to moving tubes incorrectly or placing racks at the analyzer that have not been read by the bt-40. it was determined there is minimal risk because results sent to the lis are accurate for the barcodes that are read at the analyzer even though the display on the xn and any printouts would show incorrect demographics. xn-9000 users were provided a product notification reminding them that sample tubes should not be moved replaced or swapped from their position after leaving the bt-40." | 0
6143391 | with further investigation, it was determined the root cause was due to user error. when the barcode is read at the bt-40 the laboratory information system (lis) provides the test order and patient demographics for the sample. the operator did not follow proper sample processing instructions and placed a rack at an analyzer to be processed without allowing the barcode to be read by the barcode terminal first. the xn software has been improved to change the communication between the bt-40 and xn analyzer, so when the xn-10 reads an unexpected barcode number for a sample that was not read by the bt-40, it will not associate any patient demographic information to the sample. there is a software version update to mitigate operator errors due to moving tubes incorrectly or placing racks at the analyzer that have not been read by the bt-40. it was determined there is minimal risk because results sent to the lis are accurate for the barcodes that are read at the analyzer even though the display on the xn and any printouts would show incorrect demographics.? xn-9000 users were provided a product notification reminding them that sample tubes should not be moved, replaced or swapped from their position after leaving the bt-40.><a correction to the xn-10 software was developed to enhance communication between the bt-40 and xn analyzer, so when the xn-10 reads an unexpected barcode number for a sample that was not read by the bt-40, it will not associate any patient demographic information to the sample. a field corrective action will be performed to implement the software.><on (b)(6) 2016 at 07:18 am, the xn-10 with bar code terminal (bt-40) in an xn-9000 configuration identified quality control (qc) level 1, 2 and 3 results with the patient demographics from 3 different patients. the operator stated they were processing patient samples and qc at the same time. data demonstrates the barcode terminal (bt-40) read the patient sample barcodes in the rack (b)(4) at 07:13. when the barcode is read at the bt-40 the lab information system (lis) provides the test order and patient demographics for the sample. the operator placed the qc rack at the analyzer sampler area, bypassing the bt-40, so qc barcodes were not read first by the bt-40 before analysis. the qc rack proceeded to the xn-10 sampling area, and the bar code for each sample is read. the xn-10 read the bar code of the first qc sample in the qc rack at 07:16. the barcode numbers for the qc were read accurately; however, the qc results were assigned the patient demographics that belonged to three patients samples that were read in rack (b)(4) by the bt-40. the operator incorrectly processed the qc samples, but the xn-10 incorrectly applied demographics that did not belong to the qc barcode that was read when the qc sample was analyzed. this event occurred with qc samples and no harm to patients were incurred; however, if this issue were to reoccur the xn-10 could apply incorrect patient demographics to patient samples on the xn analyzer display. the results sent to the lis are accurate for the barcodes that are read. | 0
5921103 | the customer observed falsely decreased intact parathyroid hormone results generated using the architect i2000sr analyzer for one patient. the following data was provided (pg/ml). initial (undiluted) greater than 2500 (flagged result), diluted and repeated 4000. repeated using another analyzer (undiluted) 2300 (not flagged). no impact to patient management was reported.><an abbott field service engineer was dispatched to the account. the fse inspected the architect i2000sr analyzer and replaced the stat, r1, and r2 probes and the associated tubings; cleaned the wash zone manifolds; tightened tubing connections; and performed diagnostic checks to verify system performance. the analyzer was returned to normal operation. additional testing could not be performed after the analyzer was serviced as the sample was depleted. evaluation of the customer issue included a review of the complaint text, a search for similar complaints, a review of labeling, and a review of instrument service. no returns were made available from the customer site for this evaluation. no adverse trend was identified for the customer issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation, the analyzer performed as intended and no product deficiency was identified.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. an evaluation is in process. | 0
5573688 | (b)(4). a mandatory technical service bulletin (tsb) was issued on all impacted i2000 sr instruments. the mandatory tsb instructs field service to replace the incorrect waste system, vacuum system, and some buffer system tubings.><field service replaced the incorrect tubing on the architect i2000sr analyzer per the mandatory technical service bulletin (tsb 116-132a). there was no impact to patient results. | 0
5763091 | the device evaluation was reassessed and concluded that a malfunction occurred, the device did not perform as intended. however, no systematic issue or product deficiency was identified. conclusion code in evaluation codes was corrected.><the customer observed a burning odor on the right side of the architect i2000sr analyzer in the reaction vessel (rv) area and observed visible smoke from the back of the analyzer. the analyzer was powered off. no fire or injury were reported.><an abbott field service engineer was dispatched to the account and inspected the architect i2000sr analyzer. the fse determined the middle of the three card cage fans [circulator fan repair kits (part number 7-77147-02)] were not rotating. after cleaning and inspecting the fans, no damage was observed. the analyzer was powered on and initialized to running status. the fse did not observe any additional smoke. no analyzer parts were replaced, as the analyzer was prepared for de-installation due to the laboratory closing. investigation of the customer issue included a review of the complaint text, a search for similar complaints, a review of labeling, and a review of instrument service. architect i2000sr analyzer serial number (b)(4) service history was reviewed and no contributing factors were found. field service previously replaced two of the three worn circulator fan repair kits (part number 7-77147-02) for this analyzer on 05/27/2016 because the customer reported a burning odor and noise from the right side of analyzer, however at that time smoke was not observed and no damaged parts or evidence of fire were found during the replacement. returns were not made available for this evaluation and no parts were replaced related to the current issue. no other smoke/burn issues were identified. historical complaint data was reviewed and no adverse trend was identified for circulator fan repair kit (part number 7-77147-02) or the analyzer. product labeling was reviewed and found to be adequate. the issue was resolved through normal troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation for the architect i2000sr analyzer, no product deficiency and no malfunction were identified.><(b)(4). an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. an evaluation is in process. | 0
6016012 | the device evaluation was reassessed and concluded that a malfunction occurred; therefore, the device was not performing as intended and the evaluation codes were corrected.><the customer reports the following architect cyclosporine assay result generated on an architect i1000sr analyzer on (b)(6) 2016: (b)(6)= 140.5 ng/ml. the following day, the customer reported numerous aspiration errors ((ec 3350: the test cannot be processed, aspiration error for (pipetting arm) at (manager sample)) being generated on the analyzer with the failure of the daily maintenance procedure. various hardware parts were replaced. the sample was then retested on (b)(6) 2016 and generated a result of 297.4 ng/ml. controls were within specifications on all runs. no suspect results had been reported from the lab. there is no impact to patient management reported.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete.><abbott personnel at the customer site observed that the daily maintenance failed on the architect i1000sr analyzer, serial number (b)(4), with aspiration errors and that on the day of testing error code 3350 [test cannot be processed, aspiration error for (pipetting arm) at (manager sample)] was generated on the system. while troubleshooting the issue abbott personnel replaced the probe, the pressure monitor sensor and the probe tuing. subsequent instrument operations were acceptable. no returns were made available from the customer site. a review of complaint tracking and trending metrics was performed and identified no adverse trends in conjunction with the complaint issue currently under evaluation. the architect system operations manual and the architect cyclosporine assay package insert contain information to address the current customer issue. based on the available information from the customer site and from the results of this evaluation, there is no evidence to reasonably suggest a product malfunction occurred. sample integrity cannot be ruled out as a possible cause. the issue was addressed through standard troubleshooting procedures. | 0
5455706 | donor ids generated (b)(6) prism (b)(6) results but (b)(6) confirmatory results. additional testing for (b)(6) was (b)(6). no impact to patient management was reported. no specific patient information was provided.><evaluation of complaint data for the products and likely cause lots did not identify elevated complaint activity. label reviews were performed and found them to adequately address the customer's issue. additionally, review of the device history records for the likely cause lots did not reveal any issues related to the customer's observations. a review of the prism metrics field data for lot 57108m500 was also performed and found the (b)(6) (rr) rates to be 0.11% and 0.09%, which is less than the abbott prism hcv package insert upper 95% confidence intervals of 0.38% and 0.37%, respectively. additionally, field data was reviewed by customer site and the ir and rr rates for lot 57108m500 in (b)(6) 2016 and (b)(6) 2016 were 0.07% and 0.06%, and 0.05% and 0.04%, respectively. finally, the prism metrics field data from december 2015 (earliest manufacture date of the prism reaction tray likely cause lots 57127m500 and 57128m500) was also reviewed and the initial and (b)(6) rates for the three months were 0.10%/0.11%/0.11% and 0.09%/0.10%/0.09%, respectively. as a result, all of the reactive rates under review were within package insert specifications; therefore, the product is meeting labeling claims for clinical specificity. there is no malfunction or product deficiency.><(b)(6). (b)(4). a followup report will be submitted when the evaluation is complete. evaluation in process. | 0
5746223 | field service replaced the incorrect tubing on the architect i2000sr analyzer per the mandatory technical service bulletin. there was no impact to patient results.><(b)(4). a mandatory technical service bulletin (tsb) was issued on all impacted i2000 sr instruments. the mandatory tsb instructs field service to replace the incorrect waste system, vacuum system, and some buffer system tubings. | 0
5838137 | a mandatory technical service bulletin (tsb) was issued on all impacted i2000 sr instruments. the mandatory tsb instructs field service to replace the incorrect waste system, vacuum system, and some buffer system tubings.><field service replaced the incorrect tubing on the architect i2000sr analyzer per the mandatory technical service bulletin. there was no impact to patient results. | 0
5628676 | field service replaced the incorrect tubing on the architect i2000sr analyzer per the mandatory technical service bulletin. there was no impact to patient results.><(b)(4). a mandatory technical service bulletin (tsb) was issued on all impacted i2000 sr instruments. the mandatory tsb instructs field service to replace the incorrect waste system, vacuum system, and some buffer system tubings. | 0
5370602 | an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. (b)(4).><the customer reports imprecision of results for some assays run on an architect i2000sr analyzer. controls have remained within specifications. the following examples were provided: patient (b)(6) (female): tsh = 0.006 uiu/ml; retest = 2.1203 uiu/ml. patient (b)(6): tsh = 0.0008 uiu/ml; retest = 1.0281 uiu/ml. there is no impact to patient management reported.><an abbott field service engineer visited the customer site and found bubbles in the wash buffer line. the lines were reconnected to the reservoir sensor. afterwards, the customer also replaced the buffer filter as it was obstructed. subsequent instrument operations and test results were acceptable. there were no returns made available from the customer site. a review of complaint tracking and trending metrics was performed and identified no adverse trends in conjunction with the complaint issue currently under evaluation. the architect system operation manual and the architect tsh assay package insert provide information to address the current customer issue. a single definitive cause of the issue was not identified as a clogged/obstructed buffer filter was replaced and the buffer manifold/reservoir tubing was adjusted to resolve the erratic/discrepant results. based on the available information from the customer site and from the results of this evaluation, there is no evidence to reasonably suggest a product malfunction occurred. the issue was addressed through standard troubleshooting procedures. | 0
5455241 | patient identifier multiple ids consists of (b)(6) and (b)(6). (b)(4). an evaluation is in process. a followup report will be submitted when the evaluation is complete. evaluation in process.><evaluation of complaint data for the products and likely cause lots did not identify elevated complaint activity. label reviews were performed and found them to adequately address the customer's issue. additionally, review of the device history records for the likely cause lots did not reveal any issues related to the customer's observations. a review of the prism metrics field data for lot 57108m500 was also performed and found the initial reactive (ir) and repeat reactive (rr) rates to be 0.11% and 0.09%, which is less than the abbott prism hcv package insert upper 95% confidence intervals of 0.38% and 0.37%, respectively. additionally, field data was reviewed by customer site and the ir and rr rates for lot 57108m500 in (b)(6) 2016 were 0.07% and 0.06%, and 0.05% and 0.04%, respectively. finally, the prism metrics field data from (b)(6) 2015 (earliest manufacture date of the prism reaction tray likely cause lots 57127m500 and 57128m500) was also reviewed and the initial and repeat reactive rates for the three months were 0.10%/0.11%/0.11% and 0.09%/0.10%/0.09%, respectively. as a result, all of the reactive rates under review were within package insert specifications; therefore, the product is meeting labeling claims for clinical specificity. there is no malfunction or product deficiency.><donor ids (b)(6) and (b)(6) generated reactive (b)(6) results but (b)(6) confirmatory results. no impact to patient management was reported. no specific patient information was provided. | 0
6203382 | the customer observed sparks coming from inside the architect i2000sr analyzer. no injuries occurred.><further investigation of the customer issue included a review of the complaint text, a search for similar complaints, review of instrument logs, a review of labeling and instrument service. no adverse trend was identified for the customer's issue. labeling was reviewed and found to be adequate. the issue was resolved through standard troubleshooting procedures. an abbott service representative identified a loose trigger/pre-trigger manifold drain tubing that was not inserted in baseplate drain. fluid accumulated in the power supply drip pan and then overflowed onto the artesyn/astec power supply quad module board (part number (b)(4)) resulting in a short circuit to the board. the replacement of the power supply quad module board by your abbott service representative returned the analyzer to normal function. based on all available information and abbott diagnostics' complaint investigation no product deficiency was identified. | 0
5495807 | (b)(4). photographs provided by the customer illustrated evidence of burn spots on the board. additionally, prism systems integration engineers performed a visual inspection of the returned board. the inspection showed that a capacitor had failed. the photos included in their review show evidence of burn spots related to the failed capacitor. this appears to be the most likely cause of the board failure. evaluation of the capacitor specifications determined it was the correct type for the application. a malfunction was identified as the abbott prism transport interface pcb failed and subsequent part replacement resolved the issue, which reasonably suggests the part did not perform per specification. a review of product labeling was performed and was found to sufficiently address hazards with respect to this event. additionally, the abbott prism system service and support manual was found to provide adequate instruction for use of the abbott prism transport interface pcb. a product deficiency was not identified as review of historical complaint data for the product did not identify increased complaint activity related to the complaint issue. no additional complaints were found for the complaint issue.><the customer observed a transport timeout error on the prism analyzer, serial number (b)(4). service was dispatched to resolve the issue. service identified damage on the transport interface printed circuit board (pcb). both sides of the board showed burn marks. no smoke or fire was observed. no injury occurred. | 0
5749054 | the customer reports a falsely decreased architect stat high sensitive troponin-i assay result generated on an architect i1000sr analyzer. an initial result of 0.0 ng/l was generated and reported from the lab. the sample was then tested on another architect isystem in the lab and generated a result of 49.9 ng/l. a corrected report was issued. a field service call was initiated. there is no impact to patient management reported.><an abbott field service engineer (fse) visited the customer site to inspect the architect i1000sr analyzer. the fse found and replaced a leaking trigger manifold kit valve (list 7-77612-03) and one leaking trigger no. 4 two-way bypass valve (list 7-200607-01). the parts were deemed worn from normal use. subsequent instrument operations were acceptable. a review of complaint tracking and trending metrics was performed and identified no adverse trends in conjunction with the complaint issue currently under evaluation. the architect system operations manual and the architect stat high sensitive troponin-i assay package insert contain information to address the current customer issue. based on the available information from the customer site and from the results of this evaluation, there is no evidence to reasonably suggest a product malfunction occurred. the issue was addressed through standard troubleshooting procedures.><the device evaluation was reassessed and concluded that a malfunction occurred; therefore, the device was not performing as intended and the evaluation codes were corrected.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. (b)(4). | 0
5458344 | (b)(4). a followup report will be submitted when the evaluation is complete. evaluation in process.><(b)(6) generated reactive prism hcv results but a negative confirmatory result. additional testing for id (b)(6) was nat negative. no impact to patient management was reported. no specific patient information was provided.><evaluation of complaint data for the products and likely cause lots did not identify elevated complaint activity. label reviews were performed and found them to adequately address the customer's issue. additionally, review of the device history records for the likely cause lots did not reveal any issues related to the customer's observations. a review of the prism metrics field data for lot 57108m500 was also performed and found the (b)(6) rates to be (b)(6), which is less than the abbott prism (b)(4) package insert upper 95% confidence intervals of (b)(6). additionally, field data was reviewed by customer site and the ir and rr rates for lot 57108m500 in january 2016 and february 2016 were (b)(6). finally, the prism metrics field data from december 2015 (earliest manufacture date of the prism reaction tray likely cause lots 57127m500 and 57128m500) was also reviewed and the initial and (b)(6) rates for the three months were (b)(6). as a result, all of the (b)(6) rates under review were within package insert specifications; therefore, the product is meeting labeling claims for clinical specificity. there is no malfunction or product deficiency. | 0
5573715 | field service replaced the incorrect tubing on the architect i2000sr analyzer per the mandatory technical service bulletin (tsb 116-132a). there was no impact to patient results.><(b)(4). a mandatory technical service bulletin (tsb) was issued on all impacted i2000 sr instruments. the mandatory tsb instructs field service to replace the incorrect waste system, vacuum system, and some buffer system tubings. | 0
6107261 | the customer observed smoke from the architect i2000sr analyzer system control center (scc). the customer stated that they smelled smoke, and when they opened the door where the scc is located, they saw visible smoke coming from the rear of the scc. no fire or other visible damage was seen. the analyzer and the scc were powered off. no fire or injury was reported.><an abbott field service representative (fsr) was dispatched to the account and found the failed component was the scc f+power supply (part number 7-206072-01) within the scc computer on the architect i2000sr analyzer. the smoke from the analyzer was limited to this part and did not spread to other parts of the equipment. the scc computer (part number 07d01-06) includes scc f+power supply (part number 7-206072-01). the fsr replaced the scc computer. the analyzer was returned to normal operation. evaluation of the customer issue included a review of the complaint text, a search for similar complaints, a review of labeling, and a review of instrument service. returned part was made available from the customer site for this evaluation. inspection of the scc f+power supply (part number 7-206072-01) confirmed burn damage of a component on the circuit board inside the power supply. the damage was limited to this part and did not spread to other parts of the equipment. architect i2000sr analyzer serial number (b)(4) service history was reviewed and no contributing factors were found. no subsequent reports of smoke/burn issues were identified after the scc computer was replaced. product labeling was reviewed and found to be adequate. the issue was resolved through normal troubleshooting procedures. based on all available information and abbott diagnostics complaint investigation for the architect i2000sr analyzer, no product deficiency was identified.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. an evaluation is in process. | 0
5458686 | evaluation of complaint data for the products and likely cause lots did not identify elevated complaint activity. label reviews were performed and found them to adequately address the customer's issue. additionally, review of the device history records for the likely cause lots did not reveal any issues related to the customer's observations. a review of the prism metrics field data for lot 57108m500 was also performed and found the initial reactive (ir) and repeat reactive (rr) rates to be 0.11% and 0.09%, which is less than the abbott prism hcv package insert upper 95% confidence intervals of 0.38% and 0.37%, respectively. additionally, field data was reviewed by customer site and the ir and rr rates for lot 57108m500 in (b)(6) 2016 were 0.07% and 0.06%, and 0.05% and 0.04%, respectively. finally, the prism metrics field data from (b)(6) 2015 (earliest manufacture date of the prism reaction tray likely cause lots 57127m500 and 57128m500) was also reviewed and the initial and repeat reactive rates for the three months were 0.10%/0.11%/0.11% and 0.09%/0.10%/0.09%, respectively. as a result, all of the reactive rates under review were within package insert specifications; therefore, the product is meeting labeling claims for clinical specificity. there is no malfunction or product deficiency.><(b)(4). a followup report will be submitted when the evaluation is complete. evaluation in process.><(b)(6) generated reactive prism (b)(6) but negative confirmatory results. additional testing for id (b)(6) was nat negative. no impact to patient management was reported. no specific patient information was provided. | 0
5463602 | possible false negative. per hologic's cytology applications specialist (cas) a customer's lab personnel reported a case with observed abnormal cells that were not located in the 22 fovs selected by the imager. cas reviewed slide on the review scope as an unknown mixed with other slides. case presented had no triggers in the 22 fovs to prompt an autoscan. full review of slide during qc showed abnormal cells outside. case was picked up on qc. there were two groups of hsil cells on the slide, with no abnormal cells in the 22 fov. the customer would like to submit the slide for internal review though the cas explained they would not receive a report. this is a reportable event in the us as this is a potential delay in diagnosis. | 0
5957849 | the customer reported a bad scatter pattern when samples were prepared on their coulter tq-prep workstation. erroneous results were not generated and there was no change or effect to patient treatment in connection with this event. customer stopped using the instrument as soon as the problem was detected.><the customer reported a bad scatter pattern when samples were prepared on their tq prep instrument. the fse (field service engineer) observed a failure of the sensors. no error message alerted the customer that the reagent bottles were empty since the sensors had failed. the fse replaced the reagent a and reagent b level sense assemblies to resolve the issue. repairs were verified per service procedures. the beckman coulter internal identifier for this event is (b)(4). | 0
5743361 | the customer stated that the mpa system aliquoted samples that had not been centrifuged. these samples were sent to another instrument for testing. the customer stated this occurred twice, beginning around one week ago on (b)(6) 2016. there were several error messages on the system at the time of the event, including a rack identifier error. the customer stated that there were erroneous results for an unspecified number of samples tested for multiple assays. of the results provided, one patient sample had an erroneous troponin i (tni) result that is reportable as a malfunction. this sample initially resulted as 0.83 ng/ml and this value was reported outside of the laboratory to the doctor. the lab caught the result within a few minutes and called back to the doctor before any patient was affected. the sample was then centrifuged and repeated, resulting as < .3 ng/ml. the repeat result was believed to be correct. there was no adverse event. the troponin i reagent lot number was 188618, with an expiration date of january 2017. the field service representative could not determine a cause. the system was found to be in operation. he states that the problem may be related to an operator handling issue.><.><investigations have determined that the affected sample was processed 2 times and each time, it was processed in a rack that had been defined to skip the centrifuge. no system malfunction was found. | 0
5533007 | the customer stated that they had an issue on the aliquoter of the cobas p612 system where a small water bubble would form at the end of the pipettor needle after rinsing. aliquots of one patient sample were created by the cobas p612 system and this sample was analyzed on an immulite 2000xp analyzer and a siemens zehntower xp analyzer. the affected sample had erroneous results that were reported outside of the laboratory for testosterone and human sex hormone-binding globulin (shbg). the doctor asked for a repeat of the primary tube of the sample and the primary tube was repeated. refer to the attachment for the initial and repeat patient results. the listed patient results on the left are the repeat results. the initial results can be seen in the text below the listed patient results. it was asked, but it is not known if the patient was adversely affected. no adverse events were alleged. the field service engineer changed tubing on the cobas p612 system and after this, the aliquoter was said to be working fine. no water bubble was seen after rinsing. the customer later called back to state that they observed an issue with a new tip on the cobas p612 aliquoter, which was waiting to aspirate primary samples. they noticed that there was some quantity of the water in the tip. the customer threw the tip away and did not perform any aliquots. the field service engineer changed the tubing again on the cobas p612 system. the customer called back again on 03/22/2016 to state that the cobas p612 aliquoter is diluting material again. no water drop was observed on the needle. it was noted that during aliquoting, an air/water bubble was aspirated. the field service engineer replaced the syringe on the cobas p612 on 03/23/2016, but the issue with the air bubble remained. he changed a metal part on the pipetting unit and found that tubings were damaged after dismantling them. it was stated that the instrument was working fine after replacing the metal part.><investigations have determined that the customer was not using the recommended cleaning agent. the customer uses merck extran ap21 and the recommended cleaning agent is extran ap 18. it was determined that the system was working according to specifications.><this event occurred in (b)(6). component code - device subassembly = metal part on the pipetting unit"." | 0
5976433 | the complaint states that sample presentation and sample queue errors (spe/sqe) were generated on the accelerator aps, causing discrepant results to be generated. spe/sqes occur when a gate malfunction or a misalignment of the gate and/or track at the aspiration point fails to hold the sample in position until all aspirations are complete. field service replaced and aligned the gate, part number 8-205912-01, and adjusted other components to address the issue. system logs showed that the system recognized and generated the appropriate error messages for the affected samples. samples with sqes do not generate results. the complaint notes that none of the suspect results were released by the system, but results for sample id (b)(6) were manually released by the lab tech; corrected results were reported after the sample was rerun. a review of tickets for part number 8-205912-01, did not identify any other complaints or trends involving the gate. a review of current labeling provides troubleshooting assistance with probable causes and corrective actions to take when errors are generated. all results are to be deleted for the sample generating the spe. the middleware and/or host lis is to be configured to prevent the release of erroneous results and the sample is to be rerun for all ordered tests. samples with sqes must have current orders canceled and reordered on a new sample id to be rerun. based on the available information, a deficiency was not identified. a malfunctioning gate failed to hold samples in the sampling position until aspirations were complete. the system properly recognized and generated the appropriate error messages.><an evaluation is in process. a follow-up report will be submitted when the evaluation is complete. section a patient information: patient identifier, age, sex . sid (b)(6), male, (b)(6) sid (b)(6), male, (b)(6).><the customer observed falsely decreased tbili results on two patients. the results provided were: sid (b)(6) initial run on 06sep2016 =0.29mg/dl / repeat on different analyzer = 0.62mg/dl; sid (b)(6) initial run on 09sep2016 = 0.37 / repeat on different analyzer = 0.63. there was no reported impact to patient management. | 0
5419704 | the sp-10 automated slide preparation unit is used in the clinical laboratory to stain patient slides. an external container with distilled water is connected to the analyzer. the container is not pre-filled. distilled water is added to the container by the laboratory associates. the laboratory has the option to determine an appropriate fill level based on their test volume and a volume of distilled-water that allows for ease in handling the container. the caller reported that the container is usually not filled completely with distilled water but this one time a laboratory associate had filled a container completely and another laboratory associate suffered a lumbar strain after lifting it. the treatment provided was used to treat symptoms associated with a lumbar strain. the treatment was not given to prevent permanent damage to a body structure. there was no report of permanent back injury. the laboratory associate was off work after the event on (b)(6) 2015 until (b)(6) 2015, and returned to modified duty until (b)(6) 2015. the laboratory associate was diagnosed with a lumbar strain on (b)(6) 2015 and the following treatment plan was prescribed; apply ice to low back for 20 minutes every hour for 2 days, then heat, stretch low back every hour, reduce flexeril to 10mg every 8 hours, ibuprofen 800mg every 8 hours, and reduce percocet 5/325 to bedtime as needed for pain. it is unknown if the laboratory associate was taking any medication prior to the incident. it is unknown how long the medication was used. this event was initially reported to (b)(4) on (b)(6) 2015 by (b)(6). | 0
5583428 | an mdr follow-up report will be submitted after the elevated complaint investigation concludes.><elevated complaint investigation (b)(4) for mdr 3005248192-2016-00002 follow-up report 1 included an existing data review, quality data review (capa/nonconformance review), product/system/instrument evaluation, and complaint history review. the information in this complaint identified that dryer assembly (part 30-144163ro) failures did occur however the overall product design requirements continue to be met. the overall product performed as designed, per safety standard iec 61010-1, and continues to perform as expected when this failure mode occurs, as there is no evidence that the single fault condition temperature exceeded 105??c on any touchable surface. based on the investigation results, no product deficiency was identified.><it was identified on april 20, 2017, that upon submission of the initial report, date of this report incorrectly stated 03/17/2016 when it should have stated 04/15/2016.><the abbott vp 2000 processor is a device designed to automate and standardize slide specimen processing and routine slide staining for the laboratory. customer reported that when slides were removed from the drying station on the vp2000 instrument the slide rack was noticeably hotter than usual. there was no report of injury. the abbott field service engi | 0
6077193 | leica biosystems received a complaint that tissue samples exhibited sub-optimal processing and may be undiagnosable. on (b)(6) 2016, leica biosystems received confirmation that seven (7) cases were undiagnosable and would require rebiopsy. patient identifier information was provided for all cases and separate reports have also been submitted for all cases. (patient identifier) is (b)(6) report numbers for all cases are as follows: 1423337-2016-00017, 1423337-2016-00018, 1423337-2016-00019, 1423337-2016-00020, 1423337-2016-00021, 1423337-2016-00022 and 1423337-2016-00023. please refer to these reports for specific details of the other patients involved. | 0
5521516 | on (b)(6) 2016, leica biosystems received information from the risk management office representative that a re-biopsy had been conducted for a patient from whom a bone specimen was not diagnosable. | 0
6142817 | on 08 november 2016, following a complaint that tissue was not processing, a leica field support specialist (fss) attended the laboratory to obtain additional information and provide applications support. the fss was notified by the complainant that five (5) tissue samples were undiagnosable and were recommended for re-biopsy. liver biopsy samples derived from three (3) male patients and one (1) female patient and one (1) prostate biopsy sample were the undiagnosed samples and recommended for re-biopsy. on (b)(6) 2016, the fss confirmed that no further information would be provided for the five (5) patients for whom re-biopsy was recommended. only the patient gender was provided. refer to importer report numbers 1423337-2016-00025, 1423337-2016-00026, 1423337-2016-00027, 1423337-2016-00028, 1423337-2016-00029 for all patients involved. | 0
6126707 | leica biosystems received a complaint regarding poor processing of tissue samples from a processing run in retort a, which completed on (b)(6) 2016. the complainant reported that 150 blocks of tissue may have been adversely affected; the tissue samples smelt of alcohol and had not been processed. on 31 october 2016, leica biosystems received information from a leica field support specialist (fss) that tissue in five (5) blocks from three (3) cases was not diagnosable and re-biopsy had been advised (this information was received by leica biosystems (b)(4) on 01 november 2016). the fss also reported that: ..the customer agrees that they think an incorrect reagent was put on the instrument." an identifier the age or date birth and the gender of each patient was requested. on 07 november 2016 leica biosystems (b)(4) received information that re-biopsy of two (2) patients was required. the date of birth and gender of each patient were also provided. refer to mfr. report# 8020030-2015-00076 for specific details of the other patient involved.><investigation of this complaint found that the reagent in bottles 4 (ethanol) and 13 (xylene) and the wax in wax chamber 4 had been replaced in the period prior to commencement of the "routine overnight 45" protocol started in retort a at 16:57pm on (b)(6) 2016 from which sub-optimal tissue processing was reported by the complainant. the reagent in bottles 4 (ethanol) and 13 (xylene) was used for the final dehydration and clearing steps respectively of the "routine overnight 45" protocol started in retort a at 16:57pm on (b)(6) 2016. the instrument was found to have functioned as designed during execution of this protocol. based on information reported by a leica field support specialist that: "..the customer agrees that they think an incorrect reagent was put on the instrument" determined that the root cause of the sub-optimal tissue processing reported by the complainant was a use error which occurred prior to commencement of the "routine overnight 45" protocol started in retort a at 16:57pm on (b)(6) 2016. as the reagents on the instrument at the time of the event had been discarded prior to attendance at the laboratory by a leica representative it was not possible to either measure the ethanol concentration in bottles designed for this reagent or to perform chemical analysis of the reagent in bottles designated for xylene. as a consequence it is not possible to establish whether the use error occurred during replacement of the reagent in either bottle 4 (ethanol) or 13 (xylene). the manufacturer instructions for manual reagent replacement are detailed in section 5.4.4 of the leica peloris/peloris ll user manual; and contain the following warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss."><on 26 january 2017 leica biosystems received the following information in written communication from the quality manager and head of department histopathology at the institution: "...we would like to express our disappointment on the reported findings particularly in regards to the quote "...the customer agrees that they think an incorrect reagent was put on the instrument' we do not recall at any occasion we made such a statement." the written communication also included the following new information: "tissue processing issue was first reported on (b)(6) 2016 friday when in the morning the samples coming out of wax smelt of alcohol. the same day a few cases processed on (b)(6) were also returned from pathologist with processing problems. this has been caught during qa reporting." the clarifying information provided by the institution in the written communication indicates that a use error did not occur prior to commencement of the "routine overnight 45" protocol started in retort a at 16:57pm on (b)(6) 2016 from which sub-optimal tissue processing was reported in the complaint received by leica biosystems on (b)(6) 2016; and as a consequence the root cause of the sub-optimal tissue processing reported from the "routine overnight 45" protocol started in retort a at 16:57pm on (b)(6) 2016 could not be determined from the information available. investigation of the additional information provided regarding "processing problems" involving cases processed on (b)(6) 2016 found that two (2) processing runs either started or completed on this date. specifically the "routine overnight 45" protocol which started in retort a at 16:35pm on (b)(6) 2016 and completed at 07:00am on (b)(6) 2016; and the "routine overnight 45" protocol which started in retort a at 16:59pm on (b)(6) 2016 and completed on (b)(6) 2016. evaluation of the instrument logs found that the instrument operated within specification during execution of these two (2) protocols; and there was no evidence of a user interaction with the instrument that would have adversely impacted the quality of the tissue processing from these two (2) protocols. the root cause of the "processing problems" reported from "routine overnight 45" protocols which started in retort a at 16:35pm on (b)(6) 2016 and 16:59pm on (b)(6) 2016 could not be determined from the information available." | 0
6137375 | investigation of this complaint found that a use error occurred at 16:40pm on (b)(6) 2016, when a user affirmed in the instrument software that the reagent concentration in bottle 8 (absolute ethanol) was to be set to default value of 100%. the actual ethanol concentration in bottle 8 (absolute ethanol) remained unchanged at 80.2%, because bottle 8 (absolute ethanol) had not been removed from the instrument for sufficient time to replace the reagent. as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled; the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence, bottle 8 (absolute ethanol) was used for the final dehydration step of both the skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the tissue samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used to re-process these samples. use of the cleaning protocol to re-process samples is a use error as detailed in section 3.2 of the (b)(4) peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal tissue processing reported from both the "bx and large runs" protocols executed on (b)(6) 2016 was a use error. specifically a user failed to complete manual replacement of the reagent in bottle 8 (absolute ethanol) in accordance with the manufacturer instructions detailed in the (b)(4) peloris/peloris ll user manual. the (b)(4) peloris/peloris ll user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss." the tissue samples would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used for re-processing of the affected samples. use of the cleaning protocol to re-process samples is also a use error as detailed in section 3.2 of the (b)(4) peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal processing reported from a processing run(s) executed on both (b)(6) 2016 could not be determined from the information available.><(b)(4) received a complaint that tissue from a protocol run in retort b on the evening of (b)(6) 2016 was "not processing" and "pathologist says alcohol in xylene." the complainant also advised that the reagent in bottles 3 6 7 and 8 had been changed; "xylenes look good" and "waxes were not checked." on (b)(6) 2016 a (b)(4) field support specialist (fss) attended the laboratory in order to obtain further details of the circumstances involved in the complaint and to provide applications support. the complainant advised the fss that: "the first problem with processing occurred on (b)(6) 2016 on both the bx and large runs. the tissue came off both of these tissue runs seemed raw and not processed." the complainant also advised that the affected tissue samples from these protocols had been re-processed by "first running tissue through the cleaning cycle and then re-running the tissue program." the complainant reported that the following samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and/or the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 were undiagnosable: liver biopsy samples derived from three (3) male patients and one (1) female patient and one (1) prostate biopsy sample. the fss also documented that re-biopsy had been recommended for the patients for whom tissue was undiagnosable. the complainant also reported that intermittent processing issues which were described by the pathologists as an artefact caused by "alcohol in the xylene on the tissue processor" had been specifically reported by the pathologists on both (b)(6) 2016. on 10 november 2016 the fss confirmed that no further information would be provided regarding the five (5) patients for whom re-biopsy was recommended. as a consequence neither an identifier nor the age/date of birth is available for the five (5) patients for whom re-biopsy was recommended. refer to mfr. report #8020030-2016-00083; 8020030-2016-00085; 8020030-2016-00086; and #8020030-2016-00087 for the other patients involved." | 0
6197275 | leica biosystems received a complaint that biopsy samples from a protocol run in retort b were under-processed. on 14 december 2016, leica biosystems received information that re-biopsy of two (2) male patients will be required. it was also documented that the institution was not prepared to provide any further patient details. investigation of this complaint by leica biosystems is in progress. refer to mfr. report# 8020030-2016-00091 for details of the other patient involved.><investigation of this complaint found that a use error occurred at 21:13pm on (b)(6) 2016 when a user affirmed in the instrument software that the reagent concentration in bottle 5 (absolute ethanol) was to be set to default value of 100%. the actual ethanol concentration in bottle 5 (absolute ethanol) remained unchanged at 78.9%, because bottle 5 (absolute ethanol) had not been removed from the instrument for sufficient time to replace the reagent. as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled; the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence, the reagent in bottle 5 (absolute ethanol) was used for the final dehydration step of both the skin/gi bxs" protocol started in retort b at 21:13pm on (b)(6) 2016 from which sub-optimal tissue processing was reported by the complainant. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. investigation of this complaint found that the instrument operated within specification during execution of the "skin/ gi bxs" protocol started in retort b at 21:13pm on (b)(6) 2016. the root cause of the sub-optimal tissue processing reported was a use error which occurred at 21:13pm on (b)(6) 2016. specifically a user failed to complete manual replacement of the reagent in bottle 5 (absolute ethanol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual which contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss." user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss." the tissue samples would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used for re-processing of the affected samples. use of the cleaning protocol to re-process samples is also a use error as detailed in section 3.2 of the leica peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue."" | 0
6071752 | leica biosystems received a complaint that tissue in 45 cassettes, which had been processed overnight on friday through to saturday in retort a and b, was very burnt" and dehydrated. the complainant advised that the reagent in each of two (2) bottles was replaced with either 70% alcohol and 90% alcohol on saturday; the quality of tissue processing from a validation run comprising a single block which was performed following replacement of the reagent in the two (2) bottles was satisfactory and a power outage in the building "later on saturday" necessitated re-start of the instrument as the ups was unable to support the instrument for the duration of the power outage. on 15 september 2016 leica biosystems received information that tissue samples exhibiting sub-optimal processing would be un-diagnosable. the number of undiagnosable samples was not provided. on 06 october 2016 leica biosystems received further information that seven (7) cases were un-diagnosable and would require re-biopsy. an identifier the age and the gender for each patient requiring re-biopsy was provided. (b)(6). refer to #8020030-2016-00061; #8020030-2016-00063; #8020030-2016-00064; #8020030-2016-00065; #8020030-2016-00066; and #8020030-2016-00067 for specific details of the other patients involved.><(b)(4). investigation of this complaint found that the instrument operated within specification during execution of both the "overnight breast biopsy" protocol started in retort a at 17:48pm on (b)(6) 2016 and the "gi/needle core overnight" protocol started in retort b at 17:49pm on (b)(6) 2016 from which sub-optimal tissue processing was reported. it was determined that the requirement to re-start the instrument because the ups did not support the instrument for the duration of the power outage which the complainant reported had occurred "later on saturday" (b)(6) 2016 did not either cause or contribute to the sub-optimal tissue processing reported because the processing runs from which sub-optimal tissue processing was reported completed at 03:45am on (b)(6) 2016. an error recorded on two (2) occasions at 17:46pm on (b)(6) 2016 indicates that the protocol being scheduled by the user could not be run because none of the bottles designated for alcohol contained reagent at a concentration either at or above the minimum required for use as the final reagent. the following information was also displayed: "cannot run protocol; final concentration threshold for alcohol exceeded. replace least pure alcohol station bottle 3." an error code recorded at 17:47pm on (b)(6) 2016 indicates that bottle 3 (alcohol) was not in contact with the corresponding sensor for 3 seconds which is not sufficient time to complete manual replacement of the reagent; and the properties of the corresponding reagent station were reset at 17:47pm on (b)(6) 2016. resetting a reagent station sets the concentration to the default value configured in the reagent types definition unless an alternative value is entered into the instrument software by the user; and resets the number of cassettes processed and the number of cycles and days to zero. the properties of the reagent in bottle 3 prior to this user action were: alcohol concentration = 55.2% cycles = 58 cassettes = 1296 days = 22. at 17:47pm on (b)(6) 2016 a user affirmed in the instrument software that the default concentration of 100% was to be set for bottle 3 (alcohol). as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled the reagent station with the lowest (in -threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type. reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type; and reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence bottle 3 (alcohol) was used for the final dehydration step in "overnight breast biopsy" protocol started in retort a at 17:48pm on (b)(6) 2016 and the "gi/needle core overnight" protocol started in retort b at 17:49pm on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the root cause of the sub-optimal tissue processing reported was a use error which occurred prior to commencement of the protocols from which sub-optimal tissue processing was reported. specifically a user failed to complete manual replacement of the reagent in bottle 3 (alcohol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual. the leica peloris/peloris ll user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss."" | 0
5346021 | bottle 8 (ethanol) was removed from the instrument for approximately four (4) minutes at 09:25am on (b)(4) 2015, which is sufficient time to complete manual replacement of the reagent: and the station properties for bottle 8 (ethanol) were reset at 09:30am on (b)(4) 2015. resetting a reagent station sets the concentration to the default value configured in the reagent types definitions unless an alternative value is entered into the instrument software by the user; and resets the number of cassettes processed and the number of cycles and days to zero. based on information provided by the complainant, it was determined that the sub-optimal tissue processing reported was derived from the dpc 3 hr" protocol started in retort a at 16:34pm on (b)(4) 2015 which completed at 05:00am on (b)(4) 2015; the "dpc 10 hr" protocol started in retort b at 16:43pm on (b)(4) 2015 which completed at 05:00am on (b)(4) 2015; the "dpc 3 hr" protocol started in retort a at 07:36am on (b)(4) 2015 which completed at 10:40am on (b)(4) 2015 and the "dpc 3 hr" protocol started in retort b at 08:04am on (b)(4) 2015 which completed at 11:08am on (b)(4) 2015. these protocols comprised five (5); 86; 30 and six (6) cassettes respectively based on data entered into the instrument software by the user. the reagent in bottle 8 (ethanol) was used for the first time in the final dehydration step of the "dpc 10 hr" protocol started in retort b at 16:43pm on (b)(4) 2015. bottle 8 (ethanol) was subsequently used for the final dehydration step of the "dpc 3 hr" protocol started in retort a at 16:34pm on (b)(4) 2015; the "dpc 3 hr" protocol started in retort a at 07:36am on (b)(4) 2015; and the "dpc 3 hr" protocol started in retort a at 08:04am on (b)(4) 2015. although the ethanol concentration calculated by the instrument software for bottle 8 (ethanol) was 99.9% the ethanol concentration measured by the hydrometer following completion of the "dpc 3 hr" protocol started in retort b at 08:04am on (b)(4) 2015 which completed at 11:08am on (b)(4) 2015 was 90%. this finding indicates that a use error occurred during the replacement of the reagent in bottle 8 which was performed on (b)(4) 2015 prior to commencement of the four (4) protocols from which sub-optimal tissue processing was reported by the complainant. the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled. the reagent station with the lowest (in -threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type. reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence bottle 8 (ethanol) was used for the final dehydration step of the "dpc 3 hr" protocol started in retort a at 16:34pm on (b)(4) 2015; the "dpc 10 hr" protocol started in retort b at 16:43pm on (b)(4) 2015; the "dpc 3 hr" protocol started in retort a at 07:36am on (b)(4) 2015; and the "dpc 3 hr" protocol started in retort b at 08:04am on (b)(4) 2015 from which sub-optimal tissue processing was reported by the complainant. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. investigation of this complaint found that the instrument operated within specification during execution of the "dpc 3 hr" protocol started in retort a at 16:34pm on (b)(4) 2015; the "dpc 10 hr" protocol started in retort b at 16:43pm on (b)(4) 2015; the "dpc 3 hr" protocol started in retort a at 07:36am on (b)(4) 2015; and the "dpc 3 hr" protocol started in retort b at 08:04am on (b)(4) 2015 from which sub-optimal tissue processing was reported by the complainant. the root cause of the sub-optimal tissue processing reported was a use error which occurred when the reagent in bottle 8 (ethanol) was replaced prior to commencement of the protocols from which sub-optimal tissue processing was reported by the complainant.><leica biosystems received a complaint regarding sub-optimal tissue processing from four (4) processing runs which completed on (b)(6) 2015. the complainant stated that "some tissue cut and sent to pathologist some tissue on second peloris for reprocessing. fatty tissue and large excisions more affected". on (b)(4) 2015 a leica field support specialist (fss) provided telephone support to the laboratory in order to assist with the circumstances involved in this complaint and to provide applications support. the laboratory provided the ethanol concentration in bottles 3-11 inclusive which had been measured using a hydrometer. the variance between the directly measured ethanol concentration and that calculated by the instrument software which is based on data entered by the user was acceptable in all instances except for bottle 8. the measured ethanol concentration in bottle 8 was 90% and the calculated concentration was 99.9%. the laboratory replaced the reagent in bottles 8 (ethanol) 12 (xylene); 13 (xylene) and 14 (xylene) and the wax in the four (4) wax baths. on (b)(4) 2015 leica biosystems melbourne received information from the laboratory confirming that tissue sample from one (1) patient was not diagnosable. a patient identifier the age/date of birth and the gender of the patient requiring re-biopsy was requested. on (b)(4) 2015 leica biosystems melbourne received further information from the laboratory indicating that re-biopsy of the one (1) patient from whom tissue was not diagnosable had not been performed; and that "the physician was notified but will decide whether or not to re-biopsy upon follow up visit with the patient". the gender and age of the patient were provided. on 06 january 2015 leica biosystems melbourne received the following information from the laboratory "as far as we know a new biopsy was not performed on this patient"." | 0
6197460 | bottle 7 (ethanol) was not in contact with the corresponding sensor for 30 seconds from 08:43am on (b)(6) 2016, which is not sufficient time to replace the reagent; and the properties of reagent station were reset. although the user affirmed in the instrument software that the reagent concentration in bottle 7 (ethanol) was to be set to 100% at 08:44am on (b)(6), the actual concentration of reagent in bottle 7 (ethanol) remained unchanged at 56.9% because the bottle had not been removed from the instrument for sufficient time to replace the reagent. as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled, the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence, bottle 7 (ethanol) was used for the final dehydration step of the routine 8hr" protocols started in retort b at 14:09pm on (b)(6) 2016 and in retort a at 02:01am on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using reagent at a concentration less than the minimum required for the final dehydration/clearing step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. investigation of this complaint found that the instrument operated within specification during execution of the routine 8hr" protocols started in retort b at 14:09pm on (b)(6) 2016 and in retort a at 02:01am on (b)(6) 2016. the root cause of the sub-optimal tissue processing reported was a use error which occurred at 08:44am on (b)(6) 2016. specifically a user failed to complete manual replacement of the reagent in bottle 7 (ethanol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual which contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss."><leica biosystems received a complaint that tissue was "very mushy" following processing. the complainant advised that the tissue samples exhibiting sub-optimal processing had been re-processed on another tissue processor located in the laboratory; a validation processing run(s) had been completed; all tissue samples exhibiting sub-optimal processing were diagnosable and the instrument had been used by the laboratory on both (b)(6) 2016 without reported incident." | 0
6071760 | leica biosystems received a complaint that tissue in 45 cassettes, which had been processed overnight on friday through to saturday in retort a and b, was very burnt" and dehydrated. the complainant advised that the reagent in each of two (2) bottles was replaced with either 70% alcohol and 90% alcohol on saturday; the quality of tissue processing from a validation run comprising a single block which was performed following replacement of the reagent in the two (2) bottles was satisfactory and a power outage in the building "later on saturday" necessitated re-start of the instrument as the ups was unable to support the instrument for the duration of the power outage. on 15 september 2016 leica biosystems received information that tissue samples exhibiting sub-optimal processing would be un-diagnosable. the number of undiagnosable samples was not provided. on 06 october 2016 leica biosystems received further information that seven (7) cases were un-diagnosable and would require re-biopsy. an identifier the age and the gender for each patient requiring re-biopsy was provided. (b)(6). refer to #8020030-2016-00061; #8020030-2016-00062; #8020030-2016-00063; #8020030-2016-00064; #8020030-2016-00066; and #8020030-2016-00067 for specific details of the other patients involved.><(b)(4). investigation of this complaint found that the instrument operated within specification during execution of both the "overnight breast biopsy" protocol started in retort a at 17:48pm on (b)(6) 2016 and the "gi/needle core overnight" protocol started in retort b at 17:49pm on (b)(6) 2016 from which sub-optimal tissue processing was reported. it was determined that the requirement to re-start the instrument because the ups did not support the instrument for the duration of the power outage which the complainant reported had occurred "later on saturday" (b)(6) 2016 did not either cause or contribute to the sub-optimal tissue processing reported because the processing runs from which sub-optimal tissue processing was reported completed at 03:45am on (b)(6) 2016. an error recorded on two (2) occasions at 17:46pm on (b)(6) 2016 indicates that the protocol being scheduled by the user could not be run because none of the bottles designated for alcohol contained reagent at a concentration either at or above the minimum required for use as the final reagent. the following information was also displayed: "cannot run protocol; final concentration threshold for alcohol exceeded. replace least pure alcohol station bottle 3." an error code recorded at 17:47pm on (b)(6) 2016 indicates that bottle 3 (alcohol) was not in contact with the corresponding sensor for 3 seconds which is not sufficient time to complete manual replacement of the reagent; and the properties of the corresponding reagent station were reset at 17:47pm on (b)(6) 2016. resetting a reagent station sets the concentration to the default value configured in the reagent types definition unless an alternative value is entered into the instrument software by the user; and resets the number of cassettes processed and the number of cycles and days to zero. the properties of the reagent in bottle 3 prior to this user action were: alcohol concentration = 55.2% cycles = 58 cassettes = 1296 days = 22. at 17:47pm on (b)(6) 2016 a user affirmed in the instrument software that the default concentration of 100% was to be set for bottle 3 (alcohol). as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled the reagent station with the lowest (in -threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type. reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type; and reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence bottle 3 (alcohol) was used for the final dehydration step in "overnight breast biopsy" protocol started in retort a at 17:48pm on (b)(6) 2016 and the "gi/needle core overnight" protocol started in retort b at 17:49pm on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the root cause of the sub-optimal tissue processing reported was a use error which occurred prior to commencement of the protocols from which sub-optimal tissue processing was reported. specifically a user failed to complete manual replacement of the reagent in bottle 3 (alcohol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual. the leica peloris/peloris ll user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss."" | 0
6137378 | investigation of this complaint found that a use error occurred at 16:40pm on (b)(6) 2016, when a user affirmed in the instrument software that the reagent concentration in bottle 8 (absolute ethanol) was to be set to default value of 100%. the actual ethanol concentration in bottle 8 (absolute ethanol) remained unchanged at 80.2%, because bottle 8 (absolute ethanol) had not been removed from the instrument for sufficient time to replace the reagent. as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled; the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence, bottle 8 (absolute ethanol) was used for the final dehydration step of both the skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the tissue samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used to re-process these samples. use of the cleaning protocol to re-process samples is a use error as detailed in section 3.2 of the leica peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal tissue processing reported from both the "bx and large runs" protocols executed on (b)(6) 2016 was a use error. specifically a user failed to complete manual replacement of the reagent in bottle 8 (absolute ethanol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual. the leica peloris/peloris ll user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss." the tissue samples would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used for re-processing of the affected samples. use of the cleaning protocol to re-process samples is also a use error as detailed in section 3.2 of the leica peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal processing reported from a processing run(s) executed on both (b)(6) 2016 could not be determined from the information available.><leica biosystems received a complaint that tissue from a protocol run in retort b on the evening of (b)(6) 2016 was "not processing" and "pathologist says alcohol in xylene." the complainant also advised that the reagent in bottles 3 6 7 and 8 had been changed; "xylenes look good" and "waxes were not checked." on (b)(6) 2016 a leica field support specialist (fss) attended the laboratory in order to obtain further details of the circumstances involved in the complaint and to provide applications support. the complainant advised the fss that: "the first problem with processing occurred on (b)(6) 2016 on both the bx and large runs. the tissue came off both of these tissue runs seemed raw and not processed." the complainant also advised that the affected tissue samples from these protocols had been re-processed by "first running tissue through the cleaning cycle and then re-running the tissue program." the complainant reported that the following samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and/or the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 were undiagnosable: liver biopsy samples derived from three (3) male patients and one (1) female patient and one (1) prostate biopsy sample. the fss also documented that re-biopsy had been recommended for the patients for whom tissue was undiagnosable. the complainant also reported that intermittent processing issues which were described by the pathologists as an artefact caused by "alcohol in the xylene on the tissue processor" had been specifically reported by the pathologists on both (b)(6) 2016. on (b)(6) 2016 the fss confirmed that no further information would be provided regarding the five (5) patients for whom re-biopsy was recommended. as a consequence neither an identifier nor the age/date of birth is available for the five (5) patients for whom re-biopsy was recommended. refer to mfr. report #8020030-2016-00083; 8020030-2016-00084; 8020030-2016-00085; and #8020030-2016-00087 for the other patients involved." | 0
6137373 | leica biosystems received a complaint that tissue from a protocol run in retort b on the evening of (b)(6) 2016 was not processing" and "pathologist says alcohol in xylene." the complainant also advised that the reagent in bottles 3 6 7 and 8 had been changed; "xylenes look good" and "waxes were not checked." on 08 november 2016 a leica field support specialist (fss) attended the laboratory in order to obtain further details of the circumstances involved in the complaint and to provide applications support. the complainant advised the fss that: "the first problem with processing occurred on (b)(6) 2016 on both the bx and large runs. the tissue came off both of these tissue runs seemed raw and not processed." the complainant also advised that the affected tissue samples from these protocols had been re-processed by "first running tissue through the cleaning cycle and then re-running the tissue program." the complainant reported that the following samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and/or the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 were undiagnosable: liver biopsy samples derived from three (3) male patients and one (1) female patient and one (1) prostate biopsy sample. the fss also documented that re-biopsy had been recommended for the patients for whom tissue was undiagnosable. the complainant also reported that intermittent processing issues which were described by the pathologists as an artefact caused by "alcohol in the xylene on the tissue processor" had been specifically reported by the pathologists on both (b)(6) 2016. on 10 november 2016 the fss confirmed that no further information would be provided regarding the five (5) patients for whom re-biopsy was recommended. as a consequence neither an identifier nor the age/date of birth is available for the five (5) patients for whom re-biopsy was recommended. refer to mfr. report #8020030-2016-00083; 8020030-2016-00084; 8020030-2016-00086; and #8020030-2016-00087 for the other patients involved.><investigation of this complaint found that a use error occurred at 16:40pm on (b)(6) 2016 when a user affirmed in the instrument software that the reagent concentration in bottle 8 (absolute ethanol) was to be set to default value of 100%. the actual ethanol concentration in bottle 8 (absolute ethanol) remained unchanged at 80.2% because bottle 8 (absolute ethanol) had not been removed from the instrument for sufficient time to replace the reagent. as the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled; the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence bottle 8 (absolute ethanol) was used for the final dehydration step of both the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016. the minimum final reagent concentration required for ethanol is 98%. the consequences of using ethanol at a concentration less than the minimum required for the final dehydration step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the tissue samples from the "skin/gi bxs" protocol started in retort a at 16:46pm on (b)(6) 2016 and the "routine 6hr" protocol started in retort b at 16:47pm on (b)(6) 2016 would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used to re-process these samples. use of the cleaning protocol to re-process samples is a use error as detailed in section 3.2 of the leica peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal tissue processing reported from both the "bx and large runs" protocols executed on (b)(6) 2016 was a use error. specifically a user failed to complete manual replacement of the reagent in bottle 8 (absolute ethanol) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual. the leica peloris/peloris ll user manual contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss." the tissue samples would have been further adversely impacted by the reported use of the cleaning protocol as the first step in the regimen used for re-processing of the affected samples. use of the cleaning protocol to re-process samples is also a use error as detailed in section 3.2 of the leica peloris/peloris ll user manual which contains the following specific warning: "do not use cleaning protocols for reprocessing as the dry step will damage tissue." the root cause of the sub-optimal processing reported from a processing run(s) executed on both (b)(6) 2016 could not be determined from the information available." | 0
5731224 | leica biosystems received a complaint that tissue samples processed in both retort a and b of peloris tissue processor serial number (b)(4) were over-processed"; and as a consequence the instrument had temporarily been taken out of service. the complainant also reported that a wax seal was torn. on-site assessment of the operation and function of the instrument was conducted by a (b)(6) field service engineer (fse) on (b)(6) 2016. the fse noted that reagents on the instrument at the time of the complaint had been replaced on (b)(6) 2016. on (b)(6) 2016 the fse was advised by the complainant that the quality of tissue processing from a validation run(s) performed subsequent to fse attendance at the laboratory using non-diagnostic tissue samples was satisfactory. on (b)(6) 2016 information provided by the complainant to a (b)(6) field support specialist indicated that tissue from one (1) case involved in this event was not diagnosable and re-biopsy of the patient had been recommended. however (b)(6) biosystems (b)(6) received information on (b)(6) 2016 that re-biopsy of the patient was not required because the pathologist was able to make a diagnosis after the tissue was re-cut.><a code recorded at 19:09pm and 19:10pm on (b)(6) 2016 indicates that the protocol being scheduled by the user could not be run because the final cassettes processed threshold for ipa had been exceeded. the following information was displayed on both occasions: "final cassettes processed threshold for ipa exceeded; processing quality may be compromised - 1 run(s) remaining. replace least pure ipa station bottle 9." a code recorded at 19:13pm on (b)(6) 2016 indicates that bottle 9 (ipa) was not in contact with the corresponding sensor for three (3) seconds which is less than the minimum period configured of 20 seconds; and the station was reset to 100%. resetting a reagent station sets the concentration to the default value configured in the reagent types definition unless an alternative value is entered into the instrument software by the user; and resets the number of cassettes processed and the number of cycles and days to zero. the properties of the reagent in bottle 9 prior to this action were: ipa concentration=92.5% cycles=61 cassettes=4433 and days=36. the reagent in bottle 9 (ipa) was used for the first time in the final dehydration/clearing step of the "factory 2 hr xylene free reduced heat" protocol started in retort a at 19:14pm on (b)(6) 2016; and was subsequently also used for the final dehydration/clearing step of the "factory 2 hr xylene free reduced heat" protocols started in retort a at 18:02pm on (b)(6) 2016; in retort b at 19:47pm on (b)(6) 2016; in retort a at 15:08pm on (b)(6) 2016; in retort a at 18:19pm on (b)(6) 2016 and in retort b at 19:09pm on (b)(6) 2016. it was determined that the quality of tissue processing from the "factory 2 hr xylene free reduced heat" protocols started in retort a at 19:14pm on (b)(6) 2016; in retort a at 18:02pm on (b)(6) 2016; in retort b at 19:47pm on (b)(6) 2016; in retort a at 15:08pm on (b)(6) 2016; in retort a at 18:19pm on (b)(6) 2016; and in retort b at 19:09pm on (b)(6) 2016 would have been adversely affected by the incorrect user action of resetting the station properties for bottle 9 (ipa) at 19:13pm on (b)(6) 2016 without replacing the reagent. although the user affirmed in the instrument software that the reagent concentration in bottle 9 (ipa) was to be set to default value of 100% at 19:13pm on (b)(6) 2016 the actual concentration of reagent in bottle 9 (ipa) remained unchanged at 92.5% because the bottle had not been removed from the instrument for sufficient time to replace the reagent. the instrument software uses reagent concentration to select reagent stations when a protocol is scheduled. the reagent station with the lowest (in-threshold) concentration of a reagent group or type is selected for the first step using that reagent group or type; and reagent stations of increasing concentration are used for the succeeding processing steps of the reagent group or type. reagent with the highest concentration is always used for the final processing step of a reagent group or type before changing to another reagent group or type. as a consequence bottle 9 (ipa) was used in the final dehydration/clearing step of the "factory 2 hr xylene free reduced heat" protocols started in retort a at 19:14pm on (b)(6) 2016; in retort a at 18:02pm on (b)(6) 2016; in retort b at 19:47pm on (b)(6) 2016; in retort a at 15:08pm on (b)(6) 2016; in retort a at 18:19pm on (b)(6) 2016 and in retort b at 19:09pm on (b)(6) 2016. the minimum final reagent concentration required for isopropanol is 95%. the consequences of using reagent at a concentration less than the minimum required for the final dehydration/clearing step in a protocol is re-introduction of water into the tissue which cannot be displaced in subsequent processing steps; and contamination of reagents used in the subsequent processing steps ultimately resulting in sub-optimal tissue processing. the root cause of the sub-optimal tissue processing reported was a use error which occurred at 19:13pm on (b)(6) 2016. specifically a user failed to complete manual replacement of the reagent in bottle 9 (ipa) in accordance with the manufacturer instructions detailed in the leica peloris/peloris ll user manual; which contains the following specific warning: "always change reagents when prompted. always update station details correctly - never update the details without replacing the reagent. failure to follow these directives can lead to tissue damage or loss."" | 0
5857709 | customer reported that while troubleshooting focus, positive control failures, a leak was observed on the top of the silver panel and also dripping from probe and from blue valve (pipette bypass valve-pbv) of their iq 200 system. the customer was wearing personal protective equipment (ppe) consisting of a lab coat, gloves and goggles. there was no exposure to the leaking hazardous fluid. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><device manufacturing date has been revised to reflect correct manufacturing date. bec internal identifier for this report (b)(6).><the field service engineer (fse) was at the customer site on (b)(6) 2016. the fse found that the source of the leak to be the tubing that goes from the specimen filter to the probe bypass valve (pbv). the fse replaced the affected tubing to resolve the issue and was able to run controls successfully. bec internal identifier for this report is (b)(6). | 0
5498833 | the customer reported ca control failed false negative for bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site and replaced the syringe pump to resolve the issue. the repairs were verified per established service procedures. bec internal identifier for this report is (b)(6). | 0
5897089 | the field service engineer (fse) was at the customer site and observed that the cause of the issue was the drain valve (drv) was sticking and putting back pressure on the flow cell. the fse replaced the drv and that resolved the reported issue. the fse successfully ran calibration, autofocus and controls. bec internal identifier for this report is (b)(6).><the customer reported positive control failing low on their iq200 select urine microscopy instrument. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. | 0
5717508 | the customer reported that there was a focus failure and positive control failure on their iq 200 system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site and found the cannula pump was faulty. the fse replaced the pump to resolve the reported issue the repairs were verified per established service procedures. (b)(4). | 0
5792437 | the customer reported cb control failed false negative for urobilinogen on their ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site on 06/22/2016. the fse observed that the issue was at the syringe pump. the fse relapsed the syringe pump and qc was no longer failing. the repairs were verified per established service procedures. bec internal identifier for this report is (b)(6). related event: 2023446-2016-00307, bec internal identifier (b)(6). | 0
5504687 | the customer reported positive control failure on their iq 200 system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site and found that the problem was caused by a damaged the pbv (pipette bypass valve). the fse replaced the pbv. the system was operational. the repairs were verified per established service procedures. bec internal identifier for this report is (b)(6). | 0
5574813 | the customer stated they are failing ca for bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site on 03/18/2016. the fse found that the rinse pump was producing air bubbles and required replacement. the fse replaced the rinse pump to resolve the issue. the repairs were verified per established service procedures. (b)(4). | 0
5455211 | the customer reported they are failing ca controls for bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated or reported out of the lab and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) observed that the syringe pump was weak, causing inconsistent sample to be dispensed onto the strip pad without a short sample error message or flag. the fse replaced the sample syringe and was able to successfully run controls.. the repairs were verified per established service procedures. (b)(4). | 0
5549350 | the field service engineer (fse) was at the customer site on (b)(6) 2016. the fse began to run qc and noticed the tubing on the probe was not attached properly causing air in the lines. he reattached the tubing and all qc passed.. the repairs were verified per established service procedures. bec internal identifier for this report is (b)(6).><the customer stated that there were (b)(6) urine bilirubin and (b)(6) blood results on control sample run on the ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. | 0
5401358 | the customer reported they are failing positive controls and focus on their iq 200 system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event.><iris field service engineer (fse) was at the customer site on 01/04/2016. the fse found that the evacuation bypass valve (ebv) was not functioning correctly. the fse replaced the ebv resolving the issue. the system was operational. bec internal identifier for this report is (b)(6). | 0
6019491 | the customer reported that ca and cb controls were failing on their ichem velocity automated urine chemistry system. erroneous patient results were not reported out and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site and observed that there was a cracked tubing at the color gravity module (cgm) allowing air bubbles in the sample line. the fse replaced the tubing at the cgm and was able to run the instrument without further issues. the repairs were verified per established service procedures. bec internal identifier for this report is (b)(4). | 0
5887572 | the field service engineer (fse) was at the customer site and observed that the cause for the reported event was a clogged flowcell. the fse flushed the flow cell, performed a manual focus of the instrument and was able to run control successfully. (b)(4).><the customer reported that there were erroneous result reported by their iq 200 system for three patients and focus failed when they attempted to run controls. although erroneous results were reported out for the three patients, the results were corrected and resent. there was no change or effect to patient treatment. | 0
5946211 | the customer reported that there was a contained fluid leak that consisted of a few drops of possible wash solution and urine on their ichem velocity automated urine chemistry system. the customer also reported that they are unable to get qc to run. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. the customer was wearing personal protective equipment (ppe), there was no direct exposure of mucous membranes. there was no injury or medical attention required.><the field service engineer (fse) was at the customer site and inspected the connectors for the color gravity module (cgm) tubing. the tubing on the right side of the cgm was loose. the fse tightened the connectors and there were no more leaks. the fse also switched the settings to primary (iris fully automated), entered the correct lot numbers for the strips and the controls, and re-ran qc with no issues.. the repairs were verified per established service procedures. (b)(4). | 0
6017908 | the customer reported that their ichem velocity automated urine chemistry system was leaking from the color gravity module (cgm). the leak consisted of only a few drops onto the strip provider module (spm). the customer was wearing personal protection equipment (ppe) consisting of a lab coat and gloves at the time of discovery. medical attention was not sought, there was no direct exposure to open wounds or mucous membranes. erroneous patient results were not reported out and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site and observed that the color gravity module (cgm) required replacement. the fse also found that the syringe pump assembly was also malfunctioning and not allowing for the completion of the cgm calibration. he replaced the cgm assembly, it's respective tubing, and the syringe pump assembly and was able to run the instrument without further issues. the repairs were verified per established service procedures. (b)(4). | 0
5401645 | the field service engineer (fse) instructed the customer via telephone to troubleshoot on (b)(6) 2016. the fse guided the customer with the troubleshooting process and the customer identified that the syringe was not completely tightened and the pipette was dripping. the customer tightened the syringe barrel and this resolved the instrument issue. the repairs were verified per established laboratory procedures. (b)(4).><the customer reported ca is failing bilirubin & ph and cb failing blood and specific gravity (sg) on the ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. | 0
5668921 | the customer reported that their iq 200 system was blowing urine out of the tube when aspirating. according to the customer, when the instrument does the air mix of the sample, some of the sample is forced out of the tube and onto the stm (sample transport mechanism). erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. the customer was wearing personal protective equipment (ppe) at the time of discovery. there was no reported exposure to wounds or mucous membranes.><the field service engineer (fse) was at the customer site on 04//29/2016 and found that the pressure regulator malfunctioned. he replaced the air charge valve (acv) air charge and pressure transducer.. the repairs were verified per established service procedures. (b)(4). | 0
5499266 | the field service engineer (fse) was at the customer site on 02/16/2016 and determined the evacuation bypass valve (ebv) was faulty. the fse replaced the ebv to resolve the reported issue. the repairs were verified per established service procedures. (b)(4).><the customer reported body fluid (bf) control failure on their iq200 system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. | 0
5448573 | the customer reported ca is failing bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated or reported out of the lab and there was no change or effect to patient treatment in connection to the event.><the field service engineer (fse) was at the customer site on (b)(4) 2016. the fse replaced the syringe pump assembly and the rinse pump to resolve the reported problem. the repairs were verified per established service procedures. (b)(4). | 0
6096739 | the customer reported that bf control level 1 and level 2 failed for rbc (low) after replacing the filter, lyse and diluent on their iq 200 system. there were no erroneous patient results generated. there was no change to patient treatment.><the field service engineer (fse) was at the customer site and observed that the evacuation bypass valve (ebv) was causing the control failures. the fse replaced the ebv to address the reported issue. the repairs were verified as per service procedures. bec internal identifier for this report is (b)(6). | 0
6068431 | the field service engineer (fse) was at the customer site and observed that there was a leaking tubing at the end of the color gravity module (cgm) assembly. the fse replaced the tubing at the cgm and confirmed there was no more leaking. the repairs were verified per established service procedures. (b)(4).><the customer reported a contained leak on their ichem velocity automated urine chemistry system. the customer also stated they were experiencing chem na and short sample errors on all patient samples and controls. the customer was wearing personal protective equipment (ppe) at the time the leak was observed. there were no reports of exposure to open wound or mucous membranes. no one sought medical attention. erroneous patient results were not reported out and there was no change or effect to patient treatment in connection to the event. | 0
5458475 | the field service engineer (fse) was at the customer site on (b)(4) 2016. the fse then replaced the syringe pump assembly and was able to run controls successfully. the repairs were verified per established service procedures. (b)(4).><the customer reported they are failing ca controls for bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated or reported out of the lab and there was no change or effect to patient treatment in connection to the event. | 0
5475459 | on 02/08/2016 the field service engineer (fse) observed air bubbles in the rinse pump and the color gravity module (cgm) tubing. per the fse, the air bubbles were being introduced into the rinse pump and the cgm tubing by the old iwash tubing. the fse replaced the iwash tubing to resolve the reported issue. the repairs were verified per established service procedures. (b)(4).><the customer reported ca is failing bilirubin and cb is failing urobilinogen & leukocytes on their ichem velocity automated urine chemistry system. erroneous patient results were not generated or reported out of the lab and there was no change or effect to patient treatment in connection to the event. | 0
5417019 | the field service engineer (fse) was at the customer site on (b)(4) 2016. the fse observed that there were bubbles in rinse line. the fse replaced the rinse pump and rinse valve to resolve the reported problem. the repairs were verified per established service procedures. (b)(4).><the customer reported ca failing bilirubin on their ichem velocity automated urine chemistry system. erroneous patient results were not generated and there was no change or effect to patient treatment in connection to the event. | 0
5787264 | this event occurred in (b)(6).><during further investigation, the device was disassembled. there were no signs of melting on the printer cover; however, the surrounding plastic of the printer roller was deformed. the strip holder of the device was contaminated. traces of urine were found inside the device and the printed circuit board (pcb) was visibly damaged by it. the thermoprinter's temperature controlling chip was damaged along with other parts of the pcb. the liquid also caused the device to short circuit. liquid (urine or a disinfectant) got into the device. it destroyed the pcb and most of it's components. because the thermoprinter's temperature controlling chip was damaged, the printer overheated and the plastic melted. this issue is not related to a product quality issue. the most likely root cause is a user handling issue. instructions on performing tests correctly and the appropriate cleaning of the device are addressed in product labeling.><the customer had returned their urisys 1100 analyzer because the printer paper burnt and the meter did not turn on any longer. the customer indicated that when trying to print from the instrument, the paper wouldn't feed through the printer roller and it became burnt. it is not known if the paper was being inserted correctly or not. the printer paper was burnt but no flame was produced. there was a piece of heavily" burned paper that has since been discarded. no adverse event occurred. it was not necessary to evacuate the laboratory. during the preliminary investigation with the returned analyzer and power adapter the instrument would not turn on. a power adapter used at the investigation site was used with the customer's analyzer and the instrument still would not turn on. a retention meter was plugged in with the customer returned power adapter and the retention meter turned on indicating there was no issue with the customer's returned adapter. the customer's returned analyzer makes a faint buzzing noise when turned on. there are also visible burn marks on the printer roller and melted plastic and burn marks near the printer area of the instrument." | 1
5375631 | the customer found that the pads were coming aff the strips. a new set of strips was provided to the customer. (b)(4).><the customer reported the pads were falling off the strips of the ichem velocity instrument. erroneous patient results were not generated and there was no change or effect to patient treatment in connection with this event. | 0
5375629 | the field service engineer (fse) evaluated the instrument. the fse found that the syringe pump was damaged. the fse replaced syringe assembly, which repaired the failing controls. the repairs were verified by the fse. (b)(4).><the customer reported the ichem velocity instrument was failing bilirubin controls. erroneous patient results were not generated and there was no change or affect to patient treatment in connection with this event. | 0
5354169 | the customer reported the pads of the ichem velocity instrument were falling off the strips. erroneous patient results were not generated and there was no change or affect to patient treatment in connection with this event.><the customer found the pads were falling off the strip provider module. the customer ran controls after the incident and verified that the controls were within specifications. (b)(4). | 0
5378590 | the customer reported the ichem velocity instrument was failing controls for multiple parameters. erroneous patient results were not generated and there was no change or affect to patient treatment in connection with this event.><the field service engineer (fse) evaluated the instrument. the fse found that the tubing to the color and gravity module was loose at the fitting. the fse replaced the tubing, which repaired the failing controls. the repairs were verified by the fse. (b)(4). | 0
5352771 | correction: the date of this report and the date received by the manufacturer were changed from 08/25/2015 to 12/22/2015.><the customer reported the positive control of the iq200 elite instrument was running low. there were no erroneous patient results generated or reported out of the lab.><the field service engineer (fse) evaluated the instrument. the fse replaced the drain valve, the pipette bypass valve, the rinse/waste station and the sample probe. the fse re-reran controls and the controls passed. the system was operational. (b)(4). | 0
5529758 | (b)(4). a postage paid label has been sent to retrieve the unit for further investigation. the u.s importer is requesting manufacturer of the device to further investigate this incident. the instruction manual for the home unit (model hem-741creln4) has following cautions for the consumers. consult your physician before using the device for any of the following conditions: if you have had a mastectomy. people with severe blood flow problems or blood disorders as cuff inflation can cause bruising". device not returned.><consumer had initially reported on (b)(6) 2016 to the retailer indicating his bp monitor kept inflating until 300 and it left bruises on the arm. he had to go to the hospital because he was on blood thinners and something happened to his veins. he went back to retailer to get his money back and retailer would not accept it. the retailer contacted the importer/distributor and requested to contact consumer. consumer was contacted (on (b)(4) 2016) by the importer/distributor to obtain additional information. consumer reported he has a bruise on the upper arm. the unit was in inflation mode and it had pumped up all the way and it would not stop pumping and they finally had to unplug the cuff for the unit to make it stop putting up. he tried to return the bp monitor to retailer. he is on blood thinners and it has created a bruise and he did go to the er to have the bruise checked out but he was not admitted and they did not change his meds and they did not give him anything for it. the cuff was snugged. he is only user and is supposed to use it three times a day. he uses his left arm. the cuff was on his bare arm not over a shirt. he was sitting on a hard chair with feet on floor. the cuff inflated up to 300. his average reading is 185/123. he has never had this issue before. the left arm and the bruise were dark and it covered the whole upper arm. consumer indicated once again he is on blood thinner. it has been three days for the bruise. consumer was advised to stop using the unit and was asked to return the unit further investigation via postage paid label that was sent to him. during follow-up call with quality analyst (qa) on (b)(6) 2016 consumer stated he purchased the unit last friday and on saturday he noticed his bruise. consumer stated he is on blood thinners. the bruise was dark and he went to the er to have it checked out. the er took him off of his blood thinners and gave him medication to prevent clotting and for the pain. consumer stated he was not admitted. he still has the bruise but it is small. he uses the unit 3 times a day. consumer has not replaced the cuff. consumer reported he is not sure if he will return the unit as he is speaking to his lawyer regarding next steps. qa informed consumer a postage paid label has been sent and requested to return the unit for further investigation and replacement unit will be sent to him. consumer sent an email on (b)(6) 2016 indicating he was called yesterday about this defective product. he was told he was going to receive a shipping label via email to return the unit. after further discussion with family he is going to contact an attorney about injuries he received while using this item. he is still out the cost of purchasing the bp200 plus the cost of an er room visit and the cost of replacing the blood pressure monitor. not to mention his time lost and pain suffered. consumer called back again on (b)(6) 2016 and indicated the he was using left arm unit caused the bruise which turned into a blood clot and was admitted over the weekend. he stopped using product before he called omron to report problem. he is taking product to his lawyers. consumer was advised a postage paid label has been sent for retrieval of the unit. quality analyst called again consumer on (b)(4) 2016 to obtain additional information. consumer stated he purchased the unit on (b)(6) 2016 per his doctor's recommendation. consumer reported that saturday (b)(6) 2016 he went to the hospital and was admitted for a blood clot. consumer stated they ran too many tests for him to remember specifically which tests were executed. he saw his doctor yesterday (b)(6) 2016 where the doctor evaluated him put him on new medication and adjusted the dosage of his current medication. consumer stated all of the medication he is taking is a result of the blood clot and his high blood pressure/hypertension. consumer stated he is on about 4-5 different medications but would not provide their names. consumer stated he has been on blood thinners since (b)(6) 2015.><(b)(4). verbal requests were made to retrieve the unit. in addition postage paid labels were sent to retrieve the unit back for inspection. however the end-user consumer did not return the home blood pressure monitor for inspection. the device was not received for evaluation; therefore a device analysis could not be completed. the manufacturer reviewed the device history record qa test data and risk analysis complaint history for the model number and similar models and complaint records for similar issues. no issue/problem was noted during data reviewed by the manufacturer. all risk mitigations warnings and cautions are still correct and in place. no further investigation required." | 0
5725505 | the aerobic blood culture bottle (bact alert) that was collected was defective. the bottle was lacking the carbon dioxide sensor, so it could not be used on the analyzer. blood culture were then reordered and obtained within 45 minutes. no adverse outcome. report and photos of bottle with missing sensor were sent to biomerieux 06/09/2016. | 0
6169966 | further review of the complaint determined this event is not reportable as a malfunction. the customer indicated the bottles screened positive and were smeared/subcultured. the package insert provides instructions for the test procedure (reference 9309027 a - en - 2014-12). test procedure section 4 states all bottles designated positive should be smeared and subcultured for acid-fast bacilli." "if the acid-fast smear is positive proceed with the mycobacteria specific identification procedures employed by your institution. if the concentrated smear is negative for acid-fast bacilli a gram stain should be performed. if both the acid-fast smear and the gram stain are negative indicating a possible false positive the bottle should be loaded back into the instrument until growth of subculture or redesignation as positive or negative. cultures which were initially determined false positive and were redesignated positive should be smeared and subcultured." as the customer followed the package insert which directs all bottles designated positive to be smeared and subcultured there was no patient impact. it has been determined that this issue should it reoccur will not likely lead to death or serious adverse event as the package insert clearly defines the test procedure for bottles that screen positive.><a customer in the (b)(6) notified biom?rieux of false positive results associated with the bact/alert? mp culture bottles. the customer reported the samples mostly sputum pleural fluid and bronchial washes are broken down with mucolyse for 15 minutes. then the samples are decontaminated with naoh 4% vortex mixed for 20-30 seconds every five (5) minutes for 30 minutes mixed with a phosphate buffer (checked with a ph color indictor in the neutralization buffer) centrifuged on 3000g for 15 minutes and then inoculated with mas into mp bottles. the customer indicated that when the bottles signal positive they get a zn stain which indicates no growth. the bottles are then reloaded into the bact/alert instrument. the customer reported there has been no patient impact and incorrect results were not reported to the wards; however a delay in processing may have occurred with some samples due to this issue. an internal biom?rieux investigation will be initiated." | 0
6129009 | a customer in (b)(6) notified biom?rieux of a problem with duplicate bar codes associated with bact/alert? sa bottles. the customer reported that in the emergency room of the facility, they identified duplicate bar codes for the bact/alert? sa bottles and that their erp system would not allow the user to enter a generic barcode. as a result, they were unable to enter the blood culture samples or create records in the laboratory information system (lis). the customer reported having to send samples to another facility where a workaround in the erp system is being performed; however, this has the potential to cause result delays. an internal biom?rieux investigation was performed related to this issue. the bact/alert? bottle label supplier (ccl) was notified that biom?rieux had received a complaint for a duplicate bottle id number. ccl indicated the numbers were generated by a biomerieux-supplied software. the numbers are 100% scanned by ccl for no read, no scan, poor quality, and to ensure there are no duplicated bottle id numbers within the set. as part of the ccl investigation, the numbers reported as duplicate were compared to a database of bottle id numbers ccl had generated prior to biom?rieux supplying the bottle id numbers. the reported bottle id numbers were confirmed to have been previously issued. ccl concluded the root cause was that biom?rieux had supplied bottle id numbers identical to numbers ccl had historically generated. a retain label analysis" by biom?rieux confirmed ccl's conclusion on september 12 2016. further analysis of equipment configurations concluded "bact/alert? connected to observa?" or "bact/alert? connected to bact/view" could allow the result for a previous patient to be assigned to a new patient. however investigation determined it is not likely that bact/alert? connected to observa? will impact results since the workflow delays linking the culture bottle to the patient record. in each case the correct result will be assigned to the new patient record upon completion of the bottle test. in addition clinical environment mitigations decrease the likelihood of an erroneous result going unnoticed. clsi guidance m47-a principles and procedures for blood cultures regarding the required number of blood culture bottles indicates "a single culture bottle is clearly inadequate ..... multiple sets would help a clinician to distinguish .... ." as this guidance is expected to be followed by laboratories other bottles within the set may be positive as well and it is unlikely that a doctor would stop empirical therapy because of one negative blood culture bottle result reported from the lis. the expected preliminary result would most likely be negative-to-date rather than the negative that is potentially reported due to this issue. cumitech also states "it can be concluded from available data that two to four blood cultures are necessary to optimize detection of bacteremia and fungemia." cumitech 1c blood culture iv states that prior to leaving the patient's bedside the following details are to be included on culture bottle label(s): two unique identifiers (e.g. name birthdate social security number etc.); collector's name employee number and/or employee code; collection site (e.g. vein arm leg etc.); the amount of blood collected if different from the culture bottle requirements). if the bottle is called positive the technologist will unload it to subculture and gram stain. proper bottle labeling by the user makes it likely that the negative result initially reported would be detected. additionally if the bottle is true positive the physician is immediately notified and the result is updated in the database management system (i.e. observa?) or laboratory information system (lis). therefore it is highly likely that the false initial report would be detected and the positive result will be reported for the correct patient. health hazard assessment (hha) by medical doctor indicated the issue is not expected to have a negative impact on patient health due to both product mitigations as well as mitigations that occur in the context of the clinical environment. even though there is no necessary corrective action needed in the field this event is being reported under 21 cfr 803 medical device reporting as a reportable malfunction. a corrective and preventive action (capa) investigation has been initiated to ensure this type of labeling issue does not recur.><biom?rieux conducted an internal investigation: the investigation examined the bact/alert manufacturing directions including the quality control release testing documentation and all results were within specification. quality assurance subsequently released the lots for distribution to the field. the customer lis is the root cause of this investigation. since all bottle ids are retained in memory indefinitely and bottle id numbers are finite a duplication of bottle id will eventually occur. the current workflow of using a generic bottle barcode label on bottles recognized as duplicate should be continued. biom?rieux recommends a purge of bottle ids from the lis database with more than a predetermined number of days of being unloaded based on customer needs. as interim mitigation (b)(4) cases of bact/alert sa product were shipped. these bottle ids were evaluated against the customer lis to ensure no previously provided bottle ids were repeated. the complaint was received from one customer at two different locations for alleged duplicate bottle ids. the root cause was determined to be limitation created by the customer lis setup. the bottles performed as intended no irregularities were identified during the production process and the lot itself was not adversely impacted." | 1
6181870 | a customer in the (b)(6) notified biom?rieux of a false negative result associated with the bact/alert? bpa culture bottle. the customer reported platelets were screened using the bact/alert? system and recorded negative with the bact/alert? bpa and bpn culture bottles after seven (7) days of incubation and on terminal subculture. the platelets were hla matched for a particular patient and were made into two (2) doses (after bacteriology testing). the first was transfused without event but before the second was due to be transfused (2 days later), a platelet clot was noticed in the pack. upon retesting, the pack was found to contain staphylococcus aureus after only four (4) hours incubation on the bact/alert?. an internal biom?rieux investigation will be initiated.><an internal biom?rieux investigation was conducted. the customer's initial testing performed before the platelets were split into two packs showed a negative result . the first pack was administered with no issues to the patient. the second pack after two days had a visible clot and when tested was positive for staphylococcus aureus. the second pack with the visible clot is indicative of contamination introduced when the platelets were split into two packs. the bact/alert? bpa eu instructions for use were reviewed and it provides sufficient caution to the user on proper handling technique. great care must be taken to prevent contamination of the platelet sample during inoculation into the culture bottles. contamination could lead to a specimen being determined positive when a clinically relevant isolate is not actually present in the donated platelet unit. the investigation examined the bact/alert? manufacturing directions, including the quality control release testing documentation, and all results were within specification. the investigation determined the most likely root cause is contamination during the division of the platelet sample. the bact/alert? bpa eu bottles performed as expected during initial use. | 0
5745553 | a customer in the united states notified biomerieux of discrepant results associated with bact/alert&#59395;control module pkgd 3d (ref 210147). the customer reported the bact/alert&#59397;erroneously produced false positive results. as a result, one (1) patient, seen in the emergency room, was admitted and treated with antibiotics after the results were communicated to the physician. no additional information regarding the outcome of the patient was provided by the customer. an internal biomerieux investigation will be initiated.><this report was initially submitted following notification that a customer in the united states reported false positive results in association with the bact/alert 3d control module. the results were reported to the physician and used in treatment decisions. biom?rieux investigation was conducted. based on review of the log files associated with corresponding bottles, the false positives were caused by loading multiple bottles in the same rack and temperature effects due to unloading. instructions for use indicate that large numbers of bottles loaded or unloaded at the same time, or in the same area, can cause heat loss within the racks and can trigger acceleration or rate algorithms to erroneously flag positive. the customer is instructed to keep the loading and unloading times to two minutes and 15 minutes, respectively, and to load bottles into various racks during a single loading session. in addition, the user manual supplement - 514777-1en1 - bactlink external specification for bactalert systems states, because the bact/alert systems are strictly for screening positive results are always reported as preliminary results to the lis. positive results require laboratory personnel to perform gram stains (or other smear examination) and subcultures to confirm the presence of microorganisms. positive results should never be reported as final by the lis until after confirmation has been performed. bact/alert systems report unloaded negative results as final results." the investigation concluded the user acted outside the instructions for use." | 0
5953777 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. based on the initial report of bacillus, the physician started the patient on antibiotics and removed the patient's picc line. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
6073629 | a customer in the united states contacted biomerieux to report bacillus contamination associated with a bact/alert fa plus culture bottle. the culture bottle was inoculated in the emergency room (er). the corresponding box of bact/alert fa plus bottles was inspected and one empty bottle was found; this empty bottle had a hole in the bottom. this bottle was discarded. the customer then examined the contaminated bottle more closely and he observed that the label looked like it had gotten wet because the printing was slightly smeared and the paper looked slightly bubbled. he surmised that the bottle with the hole in it leaked media onto the contaminated bottle. he swabbed the label of the contaminated bottle and the subculture grew bacillus sp. the customer was not willing to provide photos of the bottle label; although he was willing to allow a biomerieux representative visit the lab. the customer stated that, to his knowledge, the patient was not harmed or required modified treatment due to the contamination issue. the customer stated they followed the collection instructions in their local sop. each bottle should have an alcohol pad placed over it until it is inoculated. review of the local sop by biomerieux personnel indicated the procedure does not state to wipe the bottle top and wait for the disinfectant to dry. it only states to remove flip cap and put an alcohol prep on top of each bottle. it should also state to wipe the bottle top and allow the alcohol to dry. the wiping action removes spores, and the drying action kills bacteria. biomerieux investigation will be initiated.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root cause. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and re-emphasizes the preparation protocol described in the package insert. | 0
5951238 | a customer reported a discrepant result when using the bact/alert fa plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient is not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. the customer could not identify the exact lot involved within the complaint; however, was able to narrow it down to either lot 3046221 (exp. 01-mar-2017) or lot 3046537 (exp. 14-apr-2017). biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5951225 | a customer reported a discrepant result when using the bact/alert? fa plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient is not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. the customer could not identify the exact lot involved within the complaint; however, was able to narrow it down to either lot 3046221 (exp. 01-mar-2017) or lot 3046537 (exp. 14-apr-2017). biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5940853 | a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product. | 0
5940811 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product. | 0
5940795 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product. | 0
5951255 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert? fn plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient is not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. the customer could not identify the exact lot involved within the complaint; however, was able to narrow it down to either lot 3046214 (exp. 26-feb-2017) or lot 3046250 (exp. 08-mar-2017). biomerieux has initiated an investigation into this issue. | 0
5951270 | a customer reported a discrepant result when using the bact/alert fn plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient is not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. the customer could not identify the exact lot involved within the complaint; however, was able to narrow it down to either lot 3046214 (exp. 26-feb-2017) or lot 3046250 (exp. 08-mar-2017). biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
6010899 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert? fn plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><adverse event and life-threatening should not be selected in the previous supplement 1. | 0
5952934 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert fn plus culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6002941 | adverse event and life-threatening should not be selected in the previous supplement 1.><a customer reported a discrepant result when using the bact/alert fn plus culture bottle. the culture bottle flagged (b)(6) and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5941003 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product.><a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant (bacillus spp.); however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6012156 | a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5954148 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5954195 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert. | 0
5954185 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
5940943 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product.><a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant (bacillus spp.); however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
5954161 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6016571 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product.><a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant (bacillus spp.); however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6012122 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root cause. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and re-emphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
5940900 | a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant (bacillus spp.); however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product. | 0
5954145 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6100197 | a customer in the united states notified biom?rieux of a complaint regarding a warning message on the myla? server indicated as an error code 6 instrument change involving four (4) bact/alert? sn culture bottles for four (4) patients. the bottles were loaded into the system on (b)(6) 2016. the customer described their system configuration as including bact/alert? 3d (version: b.40 rel 4) and myla? (version: 3.3.2-5). the customer reported that they did not relocate any bottles and stated they performed gram stain and subculture on two (2) of the four (4) bottles. the customer was directed to gram stain and subculture all four (4) bottles. the customer indicated one bottle as being false positive flagged". the customer did not report adverse events or negative patient impact associated with this issue. a biom?rieux investigation was performed. as the myla? server performed as intended by preventing an update to the database for this patient record and instructed the customer regarding culture bottle handling there was no myla? malfunction. biom?rieux opened an investigation to determine root cause of the trigger of the myla? alarm. the bact/alert? bottle label supplier (b)(4) was notified that biom?rieux had received a complaint for a duplicate bottle id number. (b)(4) indicated the numbers were generated by a biomerieux-supplied software. the numbers are 100% scanned by (b)(4) for no read no scan poor quality and to ensure there are no duplicated bottle id numbers within the set. as part of the (b)(4) investigation the numbers reported as duplicate were compared to a database of bottle id numbers (b)(4) had generated prior to biom?rieux supplying the bottle id numbers. the reported bottle id numbers were confirmed to have been previously issued. (b)(4) concluded the root cause was that biom?rieux had supplied bottle id numbers identical to numbers (b)(4) had historically generated. a "retain label analysis" by biom?rieux confirmed (b)(4) conclusion on september 12 2016. further analysis of equipment configurations concluded "bact/alert? connected to observa?" or "bact/alert? connected to bact/view" could allow the result for a previous patient to be assigned to a new patient. however investigation determined it is not likely that bact/alert? connected to observa? will impact results since the workflow delays linking the culture bottle to the patient record. in each case the correct result will be assigned to the new patient record upon completion of the bottle test. in addition clinical environment mitigations decrease the likelihood of an erroneous result going unnoticed. ?? clsi guidance m47-a principles and procedures for blood cultures regarding the required number of blood culture bottles indicates "a single culture bottle is clearly inadequate ..... multiple sets would help a clinician to distinguish .... ." as this guidance is expected to be followed by laboratories other bottles within the set may be positive as well and it is unlikely that a doctor would stop empirical therapy because of one negative blood culture bottle result reported from the lis. the expected preliminary result would most likely be negative-to-date rather than the negative that is potentially reported due to this issue. (b)(6) also states "it can be concluded from available data that two to four blood cultures are necessary to optimize detection of bacteremia and fungemia." ?? (b)(6) 1c blood culture iv states that prior to leaving the patient's bedside the following details are to be included on culture bottle label(s): two unique identifiers (e.g. name birthdate social security number etc.). collector's name employee number and/or employee code. collection site (e.g. vein arm leg etc.). the amount of blood collected if different from the culture bottle requirements). ?? if the bottle is called positive the technologist will unload it to subculture and gram stain. proper bottle labeling by the user makes it likely that the negative result initially reported would be detected. additionally if the bottle is true positive the physician is immediately notified and the result is updated in the database management system (i.e. observa?) or laboratory information system (lis). therefore it is highly likely that the false initial report would be detected and the positive result will be reported for the correct patient. health hazard assessment (hha) by medical doctor indicated the issue is not expected to have a negative impact on patient health due to both product mitigations as well as mitigations that occur in the context of the clinical environment. even though there is no necessary corrective action needed in the field this event is being reported under 21 cfr 803 medical device reporting as a reportable malfunction.><device not returned to manufacturer." | 0
6012135 | a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal a bacillus species organism; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
6016614 | an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. the root cause is not within the control of biom?rieux. an info bulletin will be issued to the field that describes the investigation results and reemphasizes the preparation protocol described in the package insert.><a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant (bacillus spp.); however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue. | 0
5940889 | a customer reported a discrepant result when using the bact/alert? culture bottle. the culture bottle flagged positive and was sub-cultured to reveal an environmental contaminant; however, the healthcare professional determined the patient was not infected by that organism. the customer reported there were no signs of contamination, yellow-colored sensors or turbid media present prior to inoculation. when specifically asked, the customer indicated that no death, injury or mistreatment was associated with this issue. biomerieux has initiated an investigation into this issue.><an internal biom?rieux investigation concluded the following: the organisms are not present within the bottle when received at the sites, but rather is introduced to the bottle during inoculation at the testing site. the broad range of species and strains recovered indicate the contaminants do not stem from a single source. testing practices or environmental conditions at the testing site are the root causes. remediation of the root cause is not within the control of biom?rieux. as a preventive action, an info bulletin will be issued to the field that describes preventive actions in order to improve methods of testing and associated environmental controls at test sites to reduce the likelihood of inoculated bottle contamination. the proposed corrective action does not introduce any new risks associated with the use of bact/alert product. | 0
5563018 | ><in (b)(6) 2015, the blood bank encountered an unanticipated problem with scanning expiration dates and unit numbers after implementation of a cerner upgrade resulting in truncated or scrambled numbers. while some of these issues were resolved, the problem remains for (b)(6) numbers and the unit numbers. this occurs randomly and across all pc's and bar-code readers. this can unnecessarily disrupt the provision of care to patients and result in the wasting of blood product. we do not know if this problem can or will extend to blood type bar coding, which may have significant impact for the patient. we have notified cerner of the issues, but have been unable to reproduce the errors. a software update by cerner is available that should address the unit numbers, but not the (b)(4) number problem. on 3/9/2016 the software version was updated to the (b)(6) program on the blood bank computers ((b)(6).) on 3/12/2016 the bar-code scanners were recalibrated to read slower. errors still occurred. | 1
6099215 | when changing drainage bag on evd, the luer lock cracked and was unable to be removed properly. the entire drainage system had to be replaced by 2 rns. this is the third failure in the past 2 months. no harm to pt. | 0
6030169 | (b)(4). qc investigation findings: the samples for mdr 1060680-2016-00031 (bed sensor pad) and mdr 1060680-2016-00032 (bed sensor pad) were returned together, as they were used together when the reported issue occurred. the end user had not dated the pad with the start date/end dated to track the pads use. all of the test/trials referenced in the complaint were performed on the monitor/pad/batteries as returned by the customer, there were no alterations of any kind to the batteries or setting on the alarm. quality control connected the devices and began system checks as instructed by the setup information on the sensor pad. the unit was tested twenty-five+ times by quality control and the operations manager, each and every time weight was removed from the pad, the alarm sounded. when weight was reapplied the alarm reset. the pad is printed with setup instructions, maintenance information, trouble shooting and warnings. there were no issues/errors defects found. root cause: the investigator was unable to determine root cause as there was no defect/error found. the unit operated as intended. corrections: replacement product was requested/sent. corrective action: there is no corrective action required at this time, the returned sample /unit operated as intended. preventive action: there is no preventive action required at this time, there were no vendor issues found. no further information is available at this time. we will provide follow up report if additional information becomes available.><a bed alarm was connected to the corresponding sensor pad. patient was laying on bed pad. it was plugged into the deroyal alarm. the patient got up and fell. the alarm never sounded. it was plugged in. customer did not believe the patient hit the reset button either. no patient injury was reported. | 1
6075178 | ><on (b)(6) 2016, pt had a 7.1 second gap without a qrs beat and per ge asystole arrhythmia definition, this falls outside their 6 to 7 seconds with a hr=0 with no qrs beat detected within the prior 2 seconds. the ge equipment took too long during this incident to trigger the asystole alarm. | 1
6075227 | ><on (b)(6) 2016, pt had a 6.3 second gap without a qrs beat and per ge asystole arrhythmia definition, this falls outside their 6 to 7 seconds with an hr=0 with no qrs beat defected within the prior 2 seconds. the ge equipment took too long during this incident to trigger the asystole alarm. | 1
6055315 | it was initially identified in the investigation that the prohibiting factors for 'lowvol' were not in place on the user table setup. the customer's initial mitigation after the issue was discovered was to add a prohibiting factor to the 'lowvol' failure codes so that any subsequent donations recorded with this failure code will allow the defined draw failure prohibiting factor to post to those failed donations. the investigation is not complete and no root cause of the issue has been identified. a supplemental report will be submitted upon conclusion of the investigation. software issue, no device available.><the customer reported on (b)(6) 2016 that an untested donation rbc component was able to be labeled based on the donor record's historic blood type without the eldorado donor application requiring an override. the donor had been collected off-line (i.e. on a paper medical history form) and there was a problem with the donor's initial phlebotomy needle stick. the staff member entered the 'lowvol' draw failure code on the medical history form for that initial inadequate donation draw. subsequently, a second phlebotomy stick was successful and a full donor unit was collected and the draw information was entered on the same paper medical history form that had been used for the initial needle sick. however, the initial stick's 'lowvol' draw failure code was not removed (in error) from. when the draw information for the successful second collection attempt was entered into the eldorado donor application, the data entry user entered the 'lowvol' draw failure code and met the successful full donor unit. the 'lowvol' failure code was not mapped in the user table setup to post a prohibiting factor nor was it defined to require unit testing or draw information. because there was no prohibiting factor defined to post, and this failure code was defined to not require unit testing, there were no conditions associated to the component manufacture record to prevent the untested donation to be identified. this donor had a historical blood type on record and the application allowed the rbc component to be labeled based on the donor's historic blood type when no current test results had been received for that donation. the customer found the issue just prior to shipping this untested rbc component to their end user. the customer subsequently provided the information that of the original four cases, the initial case was caught and stopped at shipping. the other three cases all had components distributed and transfused, however the customer verified testing from donations prior to and following the problem units and all routine donor testing is acceptable. no adverse transfusion reactions were reported. | 0
6104457 | the deficient code was introduced with the release of rsa v1.0.12.0 on 05/22/2016. a complaint pertaining to this system issue was received on (b)(6) 2016. an initial risk assessment and investigation was performed and staff incorrectly classified this complaint to a peripheral device used with the software. upon further review on 10/21/2016, it was identified that this complaint was incorrectly assessed and was reclassified as a malfunction. a capa has been initiated to further investigate the mishandling of this complaint. additionally, a review has been conducted on all open complaints to ensure no other complaints have been similarly mishandled. review of the application code responsible for the errant functionality found the design was missing necessary coding to create the particular deferral when triggered by the first level answer (direct answer to a parent question). this is a result of incorrect design methods for handling the two dhq records created during this process. no blood products or donors were affected as a result of the software design issue. a communication was provided to all users of the failure scenario. a software modification was performed, verified and validated. the scope of testing included the software modification itself, and all potentially impacted areas of the software, and other regression testing to ensure that the fix was effective and did not adversely affect the software. the software modification was implemented on 10/26/2016. notification of the correction was provided subsequent to implementation. currently, the becs is used only by a single corporate entity. as such, all users affected by the issue are currently working with the updated software application.><on 10/21/2016, oneblood, inc. identified that specific user circumstances could allow a potentially ineligible donor to be accepted despite answering no" to donor history questionnaire (dhq) question #1: "are you feeling healthy and well today". when a donor answers an electronic dhq administered via a tablet computer the dhq is submitted for review and follow up with the potential donor by a oneblood health history screener. after review with the donor the dhq is submitted the staff member will investigate the answer and not allow the donor to donate and the system applies any applicable deferrals based on dhq answers. for dhqs completed via a tablet testing has confirmed that the intended one-day deferral that should be placed on the donor is not being applied resulting in a potentially ineligible donor to be accepted despite answering the dhq in a deferrable way. this failure scenario has been confirmed to only occur if the dhq response to question #1 is answered by the donor using a tablet. all other methods of dhq completion have been confirmed to not represent this risk. additionally the failure would only manifest if staff incorrectly allow progression of the donation after health screening or the donor attempts to donate again on the same day. review of production data found that no inappropriate donations have occurred as a result of this potential failure." | 1
5469941 | i became sick from side effect of meds, gained weight, shortness of breath, waking up in nightmare, dry cough, weakness, assault due to sleeplike state/numbness, unusual bleeding. i came to the emergency room at the (b)(6) complaining of stomach pain, got shot in stomach and i have not been right since, i cry a lot, feel movement inside of me, and get sick a lot, can't hold much food down. this was in (b)(6) 2015 or (b)(6) 2016, (b)(6) records. tetra abdominal binder to help with bloating. had trouble with equipment use, and when i would finish how long it took my blood pressure to be stable. well i started having side effects and flashbacks after using some of my meds, nightmares, feeling like someone is taking advantage of me when i sleep, i am so used to working and living around other soldiers, loss of interest in work, i had no sex drive after use of some meds and experiencing ptsd, mst and depression, anxious. reaction from stress test got sick and felt nauseous. i sought medical attention at the emergency room for chest pain and pain in my stomach. the attending physician told me that i had some form of physical injury to my face and chest and that i "suffered from negative thoughts and had reoccurring thoughts. the attending physician also stated that i suffered from ptsd which was not being treated. i received several "blood test i have requested, and i am still awaiting the receipt of the result(s) of those test. i am questioning how the drug cocaine" was found in my urine as i am not now nor have i ever been a drug user. i believe that i could have possibly been the victim of a "date rape" drug and i requested that a "rape test/examination" be performed and my request was not honored." | 0
6115007 | .><the 10/17/18 tele bp 8 am 233/98 tele rn assesses, no symptoms. repeat at 11 am 204/83, she denied symptoms but dr was notified.. when tele rn called to assess again, pt told her that t k from dr's office told her to take an extra dose of clonidine on sunday evening. pt also stated that nurse was there today for snv and did a manual bp which was 146/82 versus tele bp at same time 204/83. tele rn sent an email to the rn who did the visit to confirm the discrepancy between the 2 bp units. tele rn informed philips and ordered new bp unit.><philips q&r has investigated the reported failure of tele health unit reading 204/83 vs 146/82 by manual reading. this complaint was not confirmed. device was returned to the 3rd party logistics supplier (b)(4) on october 30, 2016. the blood pressure (bp) cuff device passed all performance specifications for bp readings. investigation showed that device failed leakage test but this failure would not contribute to the failure of a high blood pressure reading. q&r will continue to monitor for trends. | 0
5350694 | the product should have been shipped to university of maryland medical but was actually shipped to incorrectly to the home of the integra neurospecialist assigned to that area. the integra neurospecialist took the device to the hospital the following day. oxygen monitoring for this pt was delayed by 24 hours due to the incorrect shipment of device. no pt injury was reported due to this delay. the pt is in stable condition.><the device is not expected to be returned. this incident has been determined to be reportable since pt monitoring and possible treatment may have been delayed due to late shipment. | 0
6025507 | .><when ecg leads were removed from the patient the monitor went into asystole. no patient harm, nurse was in the room. manufacturer response for pt. monitor, philips (per site reporter): troubleshooting info/checked with clinical support for possible input. | 1
6030469 | .><wrong vitals (bp) were being transmitted to emr. spoke with philips, they had staff discharge old patient from monitor and cycle power, problem was then resolved. manufacturer response for pt monitor, mp30 (per site reporter): service instructions - issue resolved. | 1
5536781 | the customer reported the emergency stop function on their nordictrack x9i treadmill causes the jlab software to shut off. when the software shuts off so does the ecg and breath by breath monitoring. the customer did not clarify if there was patient involvement or not, it is currently unknown.><this mdr is being submitted due to a retrospective complaint review performed by carefusion. the carefusion identification file is (b)(4). the device was evaluated at the customer's facility by a carefusion field service engineer (fse). he was unable to duplicate the reported event. as a correction, the fse installed the hotfix software and did a system calibration. carefusion continues to track and trend any incident related to this issue. (b)(4).><the customer reported the emergency stop function on their nordictrack x9i treadmill causes the (b)(4) software to shut off. when the software shuts off so does the ecg and breath by breath monitoring. there was no known patient involvement.><(b)(4). | 1
6198459 | on (b)(6) 2016, (b)(6) was notified by our business partner (b)(6) diagnostics that their customer ((b)(6) health system in (b)(6)) wrote a rule that fired incorrectly and erroneous results were sent to the lis system. a specimen was run, and the initial result was greater than the reportable range of the instrument, with no dilution. the results were rejected by instrument manager and the specimen was automatically diluted by the instrument and rerun. the rerun results were again above the range and rejected by instrument manager. the specimen was diluted again automatically by the instrument and run again. the results of this third run were still above the reportable range for that analyte. the results of the third run should have been held within instrument manager for further action by the lab, but instead the rules rounded the result up to greater than the reportable range" of the analyzer and released them to the lis. this was discovered later by the lab when a final diluted result came off the analyzer and it was indeed still above the reference range for the test but not above the reportable range for the analyzer. this issue was caused by a user error of instrument manager. the rule to evaluate the results when they are diluted was written incorrectly and the validation performed by the site did not defect this issue. the issue was caught before the pt was treated so there was no harm to the pt specified. the rule in question contains a value list with virtually the entire test menu on the list which means the rule can fire on more than just the one test in this scenario. the rule has been in place for over one year. due to the inability to positively verify that no other pt results were affected during the time the rule was in place this report is being filed. (b)(6) has not received reports of harm but is reporting because pts could have possibly been harmed." | 1
6069088 | on september 29, 2016 data innovation was notified by abbott diagnostics that one of their customer ((b)(6)) sent them a complaint for a potential reportable event ((b)(4)) because results that should be reported as (b)(6) were reported as (b)(6) on (b)(6) 2016. a specimen is run and result is (b)(6) causing duplicate tests to be ordered. a rule was written to process the duplicate tests, however, the rule fires prematurely because there is no test code in the sel event and the results are evaluated with no test results. customer indicated that (b)(6) results are being reported as (b)(6) as a result of this rule firing. this was due to a user error of instrument manager. the rule was written incorrectly and did not contain 'is numeric' and 'test resulted' as it should have, and the validation performed by the site did not detect this issue. site reported the issue and indicated they discovered in the lab so there no pt harm for this incident. the issue is being reported because erroneous rules have been in place since august 2015, so it is possible that (b)(6) test results should have been reported as (b)(6) were reported as (b)(6). data innovations has not received reports of harm but reporting because pts could have possibly been harmed. | 1
5442355 | importing error code mapping into instrument manager (im) v 8.12.10 and higher may clear another configuration's error code mapping. this would cause results that should be flagged from that configuration to have no error code on results, and it could have an effect on rules that evaluate the error code(s). if customers have a connection that is using configuration number 105, and they are using error code mapping on that configuration; then importing error codes on any other configuration will clear the lis error code mapping for configuration number 105. this may lead to invalid results being sent to the lis.><(b)(4), reported to data innovations late on friday january 15, 2016 in incident (b)(4) that the error codes were all unmapped on the n-mod-ppe configuration. the error codes missing were initially reported after the customer was reviewing results in the lis, and looked into a result that was questionable and found no error flag attached to the result in the lis. when data innovations became aware of the incident on january 18, 2016, investigation into this issue started by reviewing system logs, reports, and configurations. data innovations' research found that tests resulted on the n-mod-ppe connection before (b)(6) 2016 13:36:12 had error codes processed correctly. however, after this point (for 79 hours) all error codes received from the n-mod-ppe connection were suppressed. the system log showed that a user had imported an error code mapping file into another configuration at this time. the configuration provided by the customer had no error code mapping in it as if someone had cleared the mapping. data innovations was able to recreate the issue and determined that after importing error code mapping, all mapping for configuration number 105 was cleared. a review of the software code found an issue with the import error code mapping feature that caused all mapping on configuration number 105 to be cleared when importing the file. | 1
5926719 | during the investigation the central receiving table showed multiple updates to patient demographics that have been done both through the electronic kipu interface and manually by the client. there was no product malfunction, the issue was caused by the client. there are no reports of adverse health consequences due to this issue.><starlims clinical solution is a software solution for managing research and diagnostic laboratory data and processes and is intended to be used by trained healthcare and research laboratory professionals. the application provides single and multiple site facilities the ability to manage patient and sample registration, pre-analytical processing, manual entry of information and collection of data from ivd analyzers and/or non-diagnostics instruments, data validation (technical and/or clinical) and result reporting. a customer notified abbott informatics that patient demographic data had been overwritten with different patient demographic data. more specifically, the patient's last name, first name, date of birth, and gender were changed and the date of birth was set as the current date. | 1
5817179 | starlims public health solution is intended for use by public health laboratory professionals and it is designed to receive and process samples from clinical, culture, environmental, food and other ad-hoc or planned sample submissions. the application provides single and multiple site facilities the ability to store, retrieve, process and manage the data associated with specimen testing including but not limited to patient demographics, tests ordered, test results, test result interpretation, quality control results, and result codes (flags) that may be associated with the specimen from its entry into the laboratory workflow until the report (including patient report) is generated and released. sample information may also be retained and used for epidemiological studies or submission to adjunct agencies such as governmental public health organizations. when the software is deployed, configurable settings are programmed based on the customer preferences. for instance, each laboratory enters their specific population based reference ranges that are used to evaluate if a patient's test results are normal, below or above the reference range and require further action such as repeat testing. each laboratory enters their specific reporting rules for reporting samples through various electronic interfaces. the (b)(6) public health lab notified abbott informatics of an issue with the script code used to generate the hl7 lab result message due to the 2 of 3 rule for the (b)(6) testing protocol. the rule is the following: (b)(6). during the approval process, it was discovered that the initial result should have been (b)(6). the supervisor sent the sample back for retest, and resulted as (b)(6). the sample now has (b)(6) result line items in the results table, but the hl7 lab result message contained a repeatedly (b)(6) lab result. the public health clinic called the patient into the clinic, and provided the patient with a paper lab result report that stated the patient was (b)(6).><test configuration was reviewed and an issue has been found on executing the rule that computes the final result. if the first result is (b)(4) the software re-triggers the test in duplicate. then the 2/3 rule is applied. the error is in the test configuration of the algorithm of the 2/3 rule when the initial result is changed from (b)(6). rules are created by the customer during the implementation. there are no reports of adverse health consequences due to this issue. | 1
5517933 | the age or date of birth and weight were not supplied. based on the information available, there is no evidence of a product or device malfunction. the information reasonably suggests a user error occurred since appropriate ppe were not in place while the bec employee was servicing and troubleshooting the device. furthermore; if troubleshooting is being performed with open doors and covers, patient samples should be removed. otherwise, covers must be closed during normal operation. as per automate 2500 family instructions for use (reference pn b50901) and service manual (revision 3) safety instructions: close covers during normal operation: while the system is operating, do not touch any moving robot. make sure that all covers and doors are closed during operation. avoid contact with sample material: do not touch patient samples, disposable tips, removed caps, any machine components that come in contact with the sample material, or waste liquid with your bare hands. always wear gloves and other appropriate protective gear to protect yourself from infection. before troubleshooting the system with open doors and covers, remove patient samples.><a beckman coulter (bec) technical support employee, reported whilst replacing a drawer unit on an automate 2500 sample processing system serial number (b)(4) at a customer's site, suffered cuts to his hands due to sharp edges. the bec employee reported appropriate personal protective equipment (ppe) such as gloves were not in use. afterwards the system started operating and tried to place uncapped samples into the racks that were not present as they were removed by another operator at the customer's site. following, the bec technical support employee attempted to prevent spillage of samples and reached out to catch the tubes, spilling serum onto hands that had previously been cut. the bec employee halted the system immediately and was directed to the accident and emergency (a&e) unit in at the customer's site hospital. the a&e department contacted the bec employee's general physician (gp); however were unable to determine record of last hepatitis vaccination. the a&e department administered a booster shot for hepatitis, and gave direction for the gp through a letter requesting follow up jabs at 3 months and 6 months. additionally two blood samples were collected at the a&e department around 4 hours after the incident for additional base line testing. the bec technical support employee confirmed following up with the gp on next day of friday (b)(6) 2016 and discussion of the incident in details. to date, no further effect was reported to bec in association with this event. | 0
6157605 | ><new ge b850 monitor installed in operating room was reported by crna that the pa and cnp waveform displays appeared incorrect. pressure values did not match expected with insertion of pa catheter. difficult advancement of pa catheter. discovered later that the pa catheter was knotted and unable to removed until the second day after it had become free of the knot or loop on its own. the pa catheter was successfully removed on (b)(6) 2016. in discussion with crna and anesthesiologist about the pa and cvp waveform displays, they reported that the 2 waveforms alternated between the 2 rows of the waveform displays in the monitor screen on their own. this was discovered on (b)(6) 2016 after second event with same ge 850 monitor in same operating room was reported to leaders. the monitor and all accessory equipment were taken out of service at that time. | 1
6157606 | ><new ge b850 monitor installed in operating room was reported by anesthesiologist that the pa catheter waveform was displayed on the cvp waveform line and cvp was on the pa waveform line. the anesthesiologist and crna reported having to correct the labeling several times as it continued to transpose on the monitor. while the pt was coming off of cardiopulmonary bypass, they were unable to get readings due to continued problem, delaying the pt on bypass about 10 minutes. the monitor and all accessory equipment were taken out of service after this event. | 1
5846089 | ><took system 9 seconds to alarm asystole after last qrs beat. there was a 3 second pause arrhythmia preceding the last qrs beat. pt's heart rate had been slowing down, should not have taken so long to determine asystole. this is another example to show why the current ek-pro asystole criteria shouldn't include the hr due to its lag time. the pt had a seizure 2 seconds after the cscs (carescape central station) alarmed asystole. | 0
5492310 | ><patient taken to cardiac cath lab after having two separate occurences of cardiac arrest. the diagnostic portion of the procedure was completed when patient had cardiac arrest for the third time. cpr was initiated. the fluoroscopy failed when physician attempted to continue with the final attempt at intervention. | 0
5677472 | the customer reported that the monitor on the cns (central monitoring station) has a buzzing sound and will not stay powered on. when the monitor is on there are lines displayed on the screen.><the customer reported that the monitor on the cns (central monitoring station) has a buzzing sound and will not stay powered on. when the monitor is on there are lines displayed on the screen. customer switched out monitor with a spare to resolve the issue. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned. | 1
5781092 | manufacturer narrative: the customer reported that the cns primary hard drive failed. the customer sent in the defective hard drive to be re-imaged. hard drive has been reimaged and has had software ver 01-96 installed. reimafed hard drive sent back to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the cns primary hard drive failed. the customer is sending in the hard drive to be re-imaged. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that the cns (central nursing station) primary hard drive failed. | 1
5496683 | the biomedical engineer reports there was a receiving data interrupted" error on the cns (central monitoring system). the biomedical engineer restarted the device which back filled the old patient data. no patient data was lost. patients were being monitored on the bedside monitor. no patient harm was reported. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the biomedical engineer reports there was a "receiving data interrupted" error on the cns (central monitoring system). the biomedical engineer restarted the device which back filled the old patient data. no patient data was lost. patients were being monitored on the bedside monitor. no patient harm was reported." | 1
5450578 | the customer reports that the cns (central monitoring system) software locks up/freezes and was not displaying patient vitals, but data was passing through to their emr system. the bme (biomedical engineer) performed a reboot of the system to restore functionality, but the cns will freeze again after a few hours. had the bme check both the hard drives which were indicating good/green, however nihon kohden technical support believes the issue may be the hard drives. gave the customer the part number for a backup hard drive to order in case the issue persists. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><manufacturer narrative: the customer reports that the cns (central monitoring system) software locks up/freezes and was not displaying patient vitals, but data was passing through to their emr system. the bme (biomedical engineer) performed a reboot of the system to restore functionality, but the cns will freeze again after a few hours. had the bme check both the hard drives which were indicating good/green, however nihon kohden technical support believes the issue may be the hard drives. gave the customer the part number for a backup hard drive to order in case the issue persists. due to the age of this complaint additional information necessary to conduct an investigation is not readily available. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reports that the cns (central monitoring system) software locks up/freezes and was not displaying patient vitals, but data was passing through to their emr system. | 1
5892655 | the customer reported that the central nurse's station (cns) is rebooting on its own. the error log indicates that the application cannot locate a file, which then causes the central nurse's station (cns) application to restart. a reconfigured hard disk drive (hdd) was sent to the customer. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that the central nurse's station (cns) is rebooting on its own. the error log indicates that the application cannot locate a file, which then causes the central nurse's station (cns) application to restart. a reconfigured hard disk drive (hdd) was sent to the customer. the device was never sent in for evaluation. due to the age of this complaint additional information necessary to conduct an investigation is not readily available. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the customer reported that the central nurse's station (cns) is rebooting on its own. the error log indicates that the application cannot locate a file, which then causes the central nurse's station (cns) application to restart. | 1
5905556 | the customer reported that after the patient was removed and then reattached to the bedside monitor (bsm), the monitor alarmed for a low heart rate but did not alarm for a 15 beats per minute run of vtach.><the customer reported that after the patient was removed and then reattached to the bedside monitor (bsm), the monitor alarmed for a low heart rate but did not alarm for a 15 beats per minute run of vtach. after changing the leads and relearning the dominant qrs, the monitor worked correctly, alarming for the vtach dysrhythmia. once the electrodes were reapplied, the issue was resolved. this was determined to be a use error situation. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 if additional information becomes available. device was not returned. | 1
6038744 | the customer reports that the display on the cns (central monitoring system) is black and not working. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><manufacturer narrative: the customer reports that the display on the cns (central monitoring system) is black and not working. a new unit was sent to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reports that the display on the cns (central monitoring system) is black and not working. | 1
5874285 | the customer reported that they have had multiple power outages and now the cns (central monitoring system) will not boot up. it is on a boot screen and will not change. the ups worked normally. the customer stated it had been 4 or 5 years since the hdd was replaced. he is aware the facility should change the hdd's every 2 years. the bme ordered a hdd for this cns. the bme put a spare cns into place in the meantime. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 if additional information becomes available.><the customer reported that they have had multiple power outages and now the cns (central monitoring system) will not boot up. it is on a boot screen and will not change. the ups worked normally. | 1
5668767 | the customer reports that the vga cable was unplugged on the cns (central monitoring system) by accident and when he plugged it back in the cns would not display the patients on the monitors. after restarting the cns 3 times the device booted back up and showed 32 patients on a single screen instead of their dual screen setup. the bme went into the system setup and switched the monitoring screens to dual display and the device was still not displaying properly. shortly after, the cns gave 3 beeping sounds and restarted randomly. no patient harm was reported during this incident. the device was returned to nihon kohden and evaluated. the device boots up but does not start the cns application. the hard drives were reimaged with software version 02-53 and it was configured as a 32 bed system which corrected the problem. all steps on the maintenance check sheet of the service manual were performed and an extended test was performed with a bedside monitor, recorder and printer. the repaired device was returned to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reports that the vga cable was unplugged on the cns (central monitoring system) by accident and when he plugged it back in the cns would not display the patients on the monitors. after restarting the cns 3 times the device booted back up and showed 32 patients on a single screen instead of their dual screen setup. the bme went into the system setup and switched the monitoring screens to dual display and the device was still not displaying properly. shortly after, the cns gave 3 beeping sounds and restarted randomly. no patient harm was reported during this incident. | 1
5708008 | the customer reported that a patient had a vtach event and that the cns (central monitoring system) did not alarm. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that a patient had a vtach event and that the cns (central monitoring system) did not alarm. | 1
5935047 | the customer reports intermittent communication loss on the cns (central network station). the communication loss went from occurring once a day for about a month to up to 3 hours at a time. the customer restarts the cns to resolve the issue. the customer has been using a known working cns in place of the defective cns. the device was never sent in for evaluation. due to the age of this complaint additional information necessary to conduct an investigation is not readily available. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the customer reports intermittent communication loss on the cns (central network station). the communication loss went from occurring once a day for about a month to up to 3 hours at a time. the customer restarts the cns to resolve the issue. the customer has been using a known working cns in place of the defective cns. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reports intermittent communication loss on the cns (central network station). the communication loss went from occurring once a day for about a month to up to 3 hours at a time. the customer restarts the cns to resolve the issue. the customer has been using a known working cns in place of the defective cns. | 1
5345972 | the customer reported that the cns (central monitoring system) went into communication loss several times over the weekend.><the customer reported that the cns (central monitoring system) went into communication loss several times over the weekend. nihon (b)(4) continues to investigate the reported event. nihon (b)(4) will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available. | 1
5375442 | the device related to this complaint was returned and evaluated. the customer complaint was confirmed and the cause of the malfunction was identified. the hard drive was replaced and the repaired device was returned after we received a po from the customer.><the customer reported that the cns (central monitoring system) froze up when it was rebooted due to a ups power outage. the device was returned to nihon kohden, evaluated, and the reported issue was confirmed. the hard drive was replaced which resolved the issue reported by the customer. contacted the customer regarding this unit on (b)(6) 2016. currently waiting to hear from the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 when additional information becomes available.><the customer reported that the cns (central monitoring system) froze up when it was rebooted due to a ups power outage. | 1
5356528 | the customer reported that the cns (central monitoring system) shows multiple beds in communication loss for 12 seconds every 3-4 minutes. the customer reports that the switch shows amber lights. the customer was asked to reboot the cns which fixed the issue. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if new information is obtained. device not returned to manufacturer.><the customer reported that the cns (central monitoring system) shows multiple beds in communication loss for 12 seconds every 3-4 minutes. the customer reports that the switch shows amber lights. the customer was asked to reboot the cns which fixed the issue. | 1
5459480 | manufacturer narrative: the customer reported that the cns (central monitoring system) periodically freezes. they can only get about 15 minutes of use before it freezes. the device has been taken out of clinical use, and no patient harm was reported. the device was never sent in for evaluation as the customer was sent a replacement hard drive. the defective hard drive was destroyed per hospital policy. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the customer reported that the cns (central monitoring system) periodically freezes. they can only get about 15 minutes of use before it freezes. the device has been taken out of clinical use, and no patient harm was reported.><the customer reported that the cns (central monitoring system) periodically freezes. they can only get about 15 minutes of use before it freezes. the device has been taken out of clinical use, and no patient harm was reported. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available. | 1
5540399 | the customer reported that the cns (central monitoring system) is having intermittent lock ups. the biomedical engineer said he was going to review the alarm history for the patients admitted during the reported issues and call back, but has not contacted nihon kohden tech support since the initial report date. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that the cns (central monitoring system) is having intermittent lock ups. | 1
6032291 | manufacturer narrative: the biomed reported that the central nurse's station (cns) had a hdd failure message after a power cycle. after troubleshooting the issue, it was suggested that the customer send the device in for repair due to a corrupt file. the device was never sent in for evaluation. due to the age of this complaint additional information necessary to conduct an investigation is not readily available. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the biomed reported that the central nurse's station (cns) had a hdd failure message after a power cycle. after troubleshooting the issue, it was suggested that the customer send the device in for repair due to a corrupt file. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the biomed reported that the central nurse's station (cns) had a hdd failure message after a power cycle. | 1
5927396 | the customer reported that the central nurse's station (cns) shut down during an accidental network loop that happened while construction work was going on at the hospital. this caused a lot of bedsides to alarm and go into communication loss and some were still in comm loss after the loop was fixed.><the customer reported that the central nurse's station (cns) shut down during an accidental network loop that happened while construction work was going on at the hospital. this caused a lot of bedsides to alarm and go into communication loss and some were still in comm loss after the loop was fixed. they were advised to power cycle every dropped bedside to bring them back into communication, which resolved the issue. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available. | 1
5780600 | manufacturer narrative: the customer reported that the cns (central nurse's station) is not exhibiting sound after a power outage. they rebooted and the problem persists. customer called stating that after plugging in external speakers the problem persists. device was returned and evaluated. alarm sound and sync sound were working fine with direct connection of sound cable from cns to touch monitor. alarm sound was set to iec type. all steps were completed on the maintenance check sheet of service manual and extended testing with the recorder, printer, and bedside monitor were performed. the issue could not be duplicated. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the cns (central nurse's station) is not exhibiting sound after a power outage. they rebooted and the problem persists. customer called stating that after plugging in external speakers the problem persists. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that the cns (central nurse's station) is not exhibiting sound after a power outage. they rebooted and the problem persists. customer called stating that after plugging in external speakers the problem persists. | 1
5828407 | the customer called stating the org is not communicating with the cns (central monitoring system).><the customer called stating the org is not communicating with the cns (central network). there was a large power outage recently and the network has been in communication loss ever since. customer was asked to go to the it closet where the nihon kohden switching and patch panels are located and found that the switch was not powered on. customer was unable to power on the switch with a different cable and outlet. customer had a spare netgear switch which she tried and that fixed the issue. customer was provided part number for purchase of a new switch. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned. | 1
5874866 | the customer reported that the cns (central network station) freezes every 10-15 minutes. the device was returned to nihon kohden, evaluated, and the reported issue was confirmed after 4 days of extended testing. nihon kohden continues to investigate the reported event. currently waiting for customer to approve repair costs. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><the customer reported that the cns (central network station) freezes every 10-15 minutes. | 1
5719070 | manufacturer narrative: the customer reports that the cns (central monitoring system) malfunctioned and was frozen on the main screen. he power cycled the device and it came back on with an error message, but could not recall what it said. once he was able to power on the cns, none of the bedside monitors were being monitored. the product involved in this event has not been returned to date to allow for an analysis to be performed. the customer replaced the hard drives to resolve the issue. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned.><the customer reports that the cns (central monitoring system) malfunctioned and was frozen on the main screen. he power cycled the device and it came back on with an error message, but could not recall what it said. once he was able to power on the cns, none of the bedside monitors were being monitored. the product involved in this event has not been returned to date to allow for an analysis to be performed. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 if the product is returned for evaluation or new information is obtained.><the customer reports that the cns (central monitoring system) malfunctioned and was frozen on the main screen. he power cycled the device and it came back on with an error message, but could not recall what it said. once he was able to power on the cns, none of the bedside monitors were being monitored. | 1
5655430 | the biomedical engineer reports that the cns (central monitoring system) is experiencing communication loss with multiple bedside monitors.><the biomedical engineer reports that the cns (central monitoring system) is experiencing communication loss with multiple bedside monitors. biomedical engineer was advised to try resetting the network interface cards on the affected bedside monitors which did not resolve the issue. the biomedical engineer was asked to try a different network interface card and report if that fixed the issue. on (b)(4) 2016 the biomedical engineer was contacted and reported that there was a bad switch that had caused the problem. the biomedical engineer changed out the switch with a spare that he had which fixed the issue. the biomedical engineer purchased a spare for the future as well. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 if additional information becomes available. | 1
5660845 | field service found a burned r100 component on the main electronics board, part number 8-35003665-01. the board was replaced to resolve the issue. the accelerator aps operations manual and the centrifuge operations manual were reviewed. troubleshooting and safety information is provided. customer complaint data was reviewed and an increase in failures for the main electronics board, part 8-35003665-01, triggered an alert which suggests an adverse trend. based on this information, a product deficiency was identified and further investigation is being performed. a follow-up report will be submitted at the completion of the investigation.><the abbott field service engineer observed error code 10 [lid not locked, replace main board]. when the fse replaced the board they noticed damage to the component. there was no fire or smoke seen, and no burning smell associated with the issue. there was no impact to patient management, user safety, or facility damage reported.><a product deficiency was identified due to an increase in complaints for failures of the r100 component of the main electronics board (parts 8-35003665-01 and 8-206623-01). the supplier, inpeco, determined that the issue is from the r100 component of the hettich centrifuge board which was too small to withstand a current spike. hettich released an enhanced main board revision 06 with a different r100 resistor. it is considered acceptable to continue using the automation system because the issue does not impact operator nor patient safety. the accelerator aps operations manual and the centrifuge operations manual were reviewed and show that troubleshooting and safety information is provided. there have been no reported injuries for this issue. if the electrical components of the main electronics board fail, it will shut down the centrifuge module. if the centrifuge shuts down, the operator will receive an error message. the frequency of this hazard was assessed and it was determined that though there was a trend in complaints, the overall frequency of low had not changed from the predicted frequency as outlined in the risk management file. | 0
5449080 | the customer stated that the centrifuge produced a burning smell but no visible smoke or fire. the abbott field service engineer found that the component overheated and was charred. there was no impact to patient management, user safety, or facility damage reported.><a product deficiency was identified due to an increase in complaints for failures of the r100 component of the main electronics board (parts 8-35003665-01 and 8-206623-01). the supplier, (b)(4), determined that the issue is from the r100 component of the hettich centrifuge board which was too small to withstand a current spike. hettich released an enhanced main board revision 06 with a different r100 resistor. it is considered acceptable to continue using the automation system because the issue does not impact operator nor patient safety. the accelerator aps operations manual and the centrifuge operations manual were reviewed and show that troubleshooting and safety information is provided. there have been no reported injuries for this issue. if the electrical components of the main electronics board fail, it will shut down the centrifuge module. if the centrifuge shuts down, the operator will receive an error message. the frequency of this hazard was assessed and it was determined that though there was a trend in complaints, the overall frequency of low had not changed from the predicted frequency as outlined in the risk management file.><(b)(4). field service found a burned r100 component on the main electronics board, part number 8-35003665-01. the board was replaced to resolve the issue. the accelerator aps operations manual and the centrifuge operations manual were reviewed. troubleshooting and safety information is provided. customer complaint data was reviewed and an increase in failures for the main electronics board, part 8-35003665-01, triggered and alert which suggests an adverse trend. based on this information, a product deficiency was identified and further investigation is being performed. a follow-up report will be submitted at the completion of the investigation. | 0
5728699 | description: a prescription was written for norco 325/10mg tablet, but what was dispensed was hydrocodone 5/325. the way the drug is noted in the prescriber's software (b)(4) lists the ingredients in a manner which is backwards from the conventional listing where the hydrocodone dose is listed before the acetaminophen dose. for example, (b)(4) denotes the different hydrocodone products as: norco 325/5mg; norco 325/7.5mg; and norco 325/10mg. we have had several instances where we nearly made an error, and now one confirmed error. in reading the prescription, it is easy to recognize the 325mg and easily overlook the other number, or to expected to see the dose we usually use. there are many instances where the electronically generated prescriptions are misleading. i have two examples: guanfacine 2mg tab 24hr, it is easy to miss the fact that this is the long-acting form of intuniv, not the regular guanfacine 2mg tablet. the same thing is a problem with metoprolol 50mg 24h instead of saying metoprolol succinate" 50mg. medication administered to or used by the pt: no. (b)(4)." | 1
6108901 | pt was discharged from our children's hospital with a prescription for acetaminophen 60 mg. the product selected was 160mg/5ml and showed a dose of 60mg or 1.88 ml which is what the pt was receiving in house. we use (b)(4) at our facility. the discharge medication reconciliation was done through (b)(4) as is our process, but the dose of the drug does not transmit to the pt summary that is provided to the pt. only the product and frequency transmit. the dose does transmit to any prescriptions but since this was an over the counter medication, the parents were using the pt summary to know what to give their child. it appeared on the pt summary that the pt should receive acetaminophen 160mg/5ml as their dose. the parents gave the pt this much higher dose of acetaminophen after discharge and the pt was then readmitted to the emergency department with signs and symptoms of tylenol toxicity and received acetylcysteine for treatment of overdose. this error was presented at our facility safety operations council and our info systems department escalated this concern to (b)(4) to address since this potentially impacts all users of (b)(4) who prescribe over the counter products upon discharge. we have not received a response from them regarding a resolution for this issue. (b)(4). severity: error resulted in treatment or intervention; temporary pt harm. | 1
5846304 | a bayer service representative was on site to re-install a certo wireless system in a new mri suite. when the representative picked up the certo box and turned around, the unit was pulled out of his hands and into the magnet bore. the magnet was subsequently ramped down and the certo box was removed from the scanner without further issue. no injury or adverse event was reported.><bayer service was dispatched to re-install a certo system into the customer's new mri suite. as the service engineer was unfamiliar with the specific location of the fringe fields of this scanner and knew there were limitations regarding the placement of this device, he contacted the scanner oem (siemens) for a drawing which identified the fringe fields. the service engineer identified the location of the 30 gauss line in the room at the end of the patient table (approximately 6 feet from the bore at the end of the table) which is within the recommended fringe field strength for the installation of this component. however, the room configuration was such that the location at the point where the engineer was attempting to install the certo module was in a fringe field higher than is recommended in the installation and service manual, which is no closer than 50 gauss, for this product. as a result, when the engineer picked up the module and turned toward the magnet, the module was drawn into the magnet bore. in order to remove the certo scan room module from the magnet, the scanner had to be ramped down. the certo module was then removed from the scanner without further issue. bayer product analysis received and examined the returned certo module. visual examination confirmed evidence of physical damage consistent with the reported event. the mounting flanges were bent in an upward direction from the mounting surface as well as the disengagement of the power supply for the ethernet port from the velcro strapping. our investigation determined that the reported problem occurred as a result of human error of the service engineer upon installation. as the certo module contained ferromagnetic material, it was drawn into the magnet bore when the service engineer moved it within the 50 gauss line. the certo installation and service manual contains the following warning: ?? warning: mr conditional. contains ferromagnetic material. keep device outside the 50 gauss line, or 10ft (3.1m) from the magnet bore. device may be attracted to the magnet bore and could result in serious injury or death.><this investigation remains in progress. once the investigation is completed, a follow-up report will be submitted. | 0
5476570 | the clinical decision support device mitigating drug therapy for all pts has poorly usable functionality such that it pressures the user and defaults to heparin therapy for protection from thrombosis, but fails to consider the frequent situations when the pt is already on an anticoagulant such as warfarin or noacs. it requires a significant amount of clinician time to click the way out of the morass of the flawed decision tree to cancel heparin. it is sufficiently onerous that it is much easier to simply allow the pt to remain over anticoagulated, as was the case in this pt. this puts most pts at higher risk of bleeding events. | 1
6145664 | .><the customer reported that they received a speaker error message. the speaker would not produce sound. this occurred during preventative maintenance. there was no patient involvement.><(b)(6). | 0
5562681 | .><the customer reported that the equipment did not alarm when saturation fell below the alarm limit. a nursing mother noticed that her baby turned blue. the nurse reported that the alarm was set at 85% and the neonate's spo2 fell to 40%. there was no treatment or patient harm. | 1
6157710 | a follow-up report will be submitted upon completion of the investigation.><it was reported by nursing to the biomed department that on (b)(6) 2016, the sp02 did not alarm. the monitor was quarantined until (b)(6) 2016. the customer has requested the logs reviewed. the device was in clinical use at the time the issue was discovered. no adverse event or patient harm was reported.><. | 1
5903861 | .><the customer reported a speaker malfunction error". the device was in clinical use at the time the issue was discovered. there was no adverse event or patient harm reported." | 1
5946105 | evaluation of the provided data by the philips research and development expert responsible for the sp02 measurement confirmed that the issue is not associated with a product failure. the issue is consistent with normal device behavior which the user misunderstood. the patient had a desat alarm event which was silenced. the subsequent low sp02 alarm was provided because the desat alarm criteria was no longer met but the patient's sp02 was still below the low sp02 limit thereby prompting the yellow low alarm to be presented. when the desat alarm is active, the low sp02 alarm remains in the background but not displayed until such time as the desat condition ends. if the sp02 value is rising but remains below the low limit a low sp02 alarm is then presented and the alarm text will contain the maximum exceeded value or lowest value seen during the measurement interval in the enhanced alarm text. the actual current value for the sp02 measurement is available/displayed on the monitor screen.><the customer reported that the neonate's saturation dropped to 22 and there was only a yellow spo2 low alarm. the desat limit was set to 75. the baby was manually ventilated. further information about the patient's condition is currently not available.><a follow-up report will be submitted once the investigation is complete. | 1
5834689 | .><the customer reported that they had a speaker malfunction" error; they were unsure if the speaker produced sound at the time of the event. at the time of the alleged malfunction the device was being used for clinical monitoring. there was no adverse event or patient harm reported." | 1
5738846 | the customer reported that a yellow non invasive blood pressure (nbp) alarm occurred at the mp50 bedside monitor but did not appear at the information center. no patient harm was reported.><. | 1
6157899 | the customer reported that on (b)(6) at 11pm the baby was found vital signs absent. a code was called. baby was resuscitated and died 48 hours after transfer. the staff believe that the monitor did not alarm.><a follow up report will be submitted once the investigation is complete. | 1
5507786 | the customer reported an error message: loudspeaker error". additional information received indicated that the loudspeaker was not producing sound. the device was not in clinical use at the time the issue was discovered. there was no adverse event or patient harm reported.><." | 1
6079426 | (b)(6)><the customer reported that the speaker failed; there was no sound coming from the speaker. | 1
5701798 | ><the patient was wrapped in a blanket on the ct table for the scan procedure; a patient strap was not used to secure the patient's arm, despite the fact that the patient was comatose. during the scan, the patient's left arm slid down from his chest and his hand hit the table top cover when the cradle was moved back to a home position. a force of approx. 200n was applied to his arm, which resulted in a fracture of the humerus. the operator was reportedly not aware that the injury had occurred since the patient was wrapped in a blanket. the conclusion of the investigation is that improper positioning/monitoring of the patient caused the injury. | 0
5832004 | orthalign conducted a conference call with the reporting surgeon, dr. (b)(6), on 06/30/2016, regarding the reported issue. the following points were noted: dr. (b)(6) stated that the patient was a normal sized woman. there were no special circumstances or surgical complications. the patient had standard osteoarthritis, without dysplasia. bone quality was normal. the orthalign jig was securely fixed during the procedure; dr. (b)(6) had two assistants helping him with the procedure; they removed the jig" (pelvic base bracket and probe) after navigating and impacting the cup placement. the jig was removed while they were putting in retractors and preparing the femur; the pins were also removed at this time using a (b)(4) (b)(4) system 6 surgical drill; he suspects the pin broke while it was being removed. this was not noticed at the time; the pin tip was visible on the first x-ray of the patient's hip taken at the 6-week follow-up visit. after seeing the tip at the edge of the x-ray dr. (b)(6) took a second x-ray to include the full hips. the 2nd x-ray showed the full pin section in the right ilium; the patient presented as asymptomatic and "doing great" at the 6 week follow-up visit; dr. (b)(4) explained the x-ray and retained pin to the patient. he made the patient aware of the metal in her body for potential mri procedures. he recommended leaving the pin in since there were no symptoms or adverse reactions to it. the patient elected to leave the pin in and was "fine with it" and "accepting of the situation"; dr. (b)(6) also elected to do nothing additional due to the retained pin. he scheduled a routine 1-year follow-up with the patient and intends to x-ray the pin location again at that time; dr. (b)(6) stated that he sees a drill bit pin or similar fragment left in the bone of one-out-of-two hundred to one-out-one-hundred patients (from prior surgical procedures). foreign object breakage and retention in a patient's bone is an existing occurrence in orthopedic surgery. he recently saw a patient who had been previously treated in (b)(6) and had a drill bit piece retained in his ilium. the patient was asymptomatic and "fine with it."; we informed dr. (b)(5) that the pin material was electropolished 17-4 ph h900 stainless steel and that it is biocompatible but not validated for long term implantation. 17-4 stainless is a common widely used material in orthopedic surgical instruments. dr. (b)(6) noted that other retained objects (like the pin and drill bit fragments mentioned above) were in the same category possibly of the same material; dr. (b)(6) stated that the risk from the retained pin is a potential foreign body reaction "just like with any implant." there is sometimes a risk of migration with retained objects in the body but he thought there "limited odds" of that happening in this case given its location in the ilium and the nature of the screw thread fixation; we enquired if dr. (b)(6) recommended or made any technique changes after this incident. he replied that he had discussed the incident and the importance of inspecting pins at the time of removal with his surgical assistants. the orthalign plus hipalign instrument sets (403007) were returned to orthalign for evaluation. no issues were discovered with any of the instruments in the sets. all included parts met specifications. the pins were not returned with the sets. they were most likely discarded by the hospital after use. so the pin involved in this incident was not available for further evaluation. the lot history including inspection records and material certifications for the pins supplied with the sets to the user were evaluated. no issues or discrepancies were discovered.><a surgeon reported the tip of a fixation pin retained in the pelvis of a patient following a tha procedure navigated with the orthalign plus system. there was no detection of pin breakage or complications at the time of surgery. an x-ray during the patient's six week follow-up visit showed a pin section lodged in the right iliac wing. the patient did not have any complications or symptoms related to the retained pin section at the time of the follow-up visit. the surgeon and patient elected to leave the retained pin section in place. no adverse reactions or complications were expected." | 0
5707990 | during a pulmonary vein isolation procedure, a patient burn occurred. following the successful ablation, a burn was noted on the left upper back of the patient. the burn was slightly larger than a quarter, described as round and red with a white skin tone around the burn, which was slightly deep. the burn was noted at the corner of the ensite navx patch proximal to the spine and the dispersive pad; however, no apparent patch overlap was noted. the patient's skin was not hairy or diaphoretic during the ablation procedure and no skin preparation was performed due to this. a wound specialist was consulted, unknown treatment was administered, and the patient has since been discharged. there were no performance issues with any sjm device during the procedure.><(b)(4). the results of the investigation are inconclusive since the device was not returned for analysis. the device history record for the above-referenced product was unable to be reviewed since the batch number is unknown. based on the information received, the cause of the reported incident could not be conclusively determined. | 0
6149729 | it was reported that during use the vigileo monitor had a burned smell; however, the monitor worked fine. no additional information was available. no patient compromise was reported. no other system-related devices were reported as suspect. no patient demographics were able to be obtained.><examination of the returned vigileo monitor confirmed the customer?s complaint. the monitor was plugged in and after a few minutes a smell was detected. evaluation testing of the ''c480'' connector component from the main board was found to have smoke damage. the main board was replaced to restore the monitor to functionality. review of the device history record supported that there were no non-conformances noted for any reason. complaint histories for all reported events are reviewed against trending control limits on a monthly basis, and any excursions above the control limits are assessed and documented as part of this monthly review. no corrective or preventative actions are required at this time. | 1
5721003 | it was reported that the vigileo monitor showed high values when compared with the values provided by the echocardiography. the co (cardiac output) value displayed was approximately 9.0 l, the sv (stroke volume) was 155 ml, the patients pulse rate was 60 and the blood pressure was low. an error message of ?sv high? was displayed, however, there were no error messages or alarms related to the co values. the expected values were 3 l for the co and 60 ml for the sv. no additional information was able to be obtained. there was no allegation of patient compromise and no other system related devices were identified as suspect><the device was expected to be returned for evaluation; however, new information was received that the device is not available for evaluation. although the device was not returned the device history record was reviewed and it supports that there were no non-conformances noted for any reason. no further actions are required as this time. edwards will continue to review and monitor all events. if action is required, an appropriate investigation will be performed.><the device is expected to be evaluated; however, at the time of this report, the device has not been returned. a supplemental report will be submitted to communicate the results of the complaint investigation results. | 1
6019311 | the device is expected to be evaluated; however, at the time of this report, the device has not been returned. a supplemental report will be submitted to communicate the results of the complaint investigation results.><the review of the device history record supports that there were no non-conformances noted for any reason. evaluation testing supports that the cable functions as expected. no physical damage was found in the visual inspection. it is unknown whether procedural processes or a specific circumstance played a role in the customer?s experience, as the fault could not be replicated and no other issues were detected. invasive procedures involve some patient risks. with any hemodynamic monitoring, pressure readings can change quickly and dramatically. pressure readings should correlate with the patient?s clinical manifestations. complaint histories for all reported events are reviewed against trending control limits on a monthly basis, and any excursions above the control limits are assessed and documented as part of this monthly review. no corrective or preventative actions are required at this time. the flotrac sensor was also submitted and the mdr number is 2015691-2016-03140.><it was reported, that during use of an ev1000 monitor, the values displayed on the screen were frozen for eight hours. the stroke volume was displayed and the colored targets remained visible, indicating that the values were live data. the patient's vital signs continued to record and monitor and the map (mean arterial pressure) was displayed on the patient's bedside monitor. the customer tried to resolve the issue, however, error message databox disconnected" was observed and the databox remained connected at all times. an edwards representative was informed of the issue and went to the hospital to try and troubleshoot. when the representative attempted to do an engineering download error message "unable to download data" was observed. there was no allegation of patient injury." | 1
5355966 | the device is expected to be returned; however, it was not returned a the time of this report. a follow-up submission will communicate the results of the device history record review, the date of manufacture and expiry date, the evaluation results and conclusion.><new information was received which changed the description summary for this event. new information supports that the device alarmed and the initial report of inaccurate values was not correct. the new information supports that the issue was related to an event involving an apc09 cable. an error message was displayed on the vigileo monitor screen, prompting the customer to trouble-shoot and exchange the apc09 cable to resolve the error message and the issue. there was no allegation of any patient compromise. the complaint file has been updated to reflect the new information and amend the suspect device from a monitor to an apco9 cable.><it was reported that the arterial pressure values displayed in the zeroing screen on the vigileo monitor went up and down from 50 mmhg to 120 mmhg; thus the customer did not find the values credible and the monitoring was stopped. the values monitored were co (cardiac output), sv (stroke volume), svv (stroke volume variation), ci (cardiac index) and svi (stroke volume index). the co value was around 5 l/min, although the expected value was less than 4 l/min. no error or alert messages were noted as displayed on the monitor. there was no allegation of patient compromise and no additional system-related devices were identified as suspect. the monitor is available for return and evaluation. | 1
5796888 | the device is expected to be evaluated; however, at the time of this report, the device has not been returned. a supplemental report will be submitted to communicate the results of the complaint investigation results.><it was reported, that during the preparation of the set for the procedure, the ev1000 power supply cable was damaged beyond repair due to a possible fluid ingress. no sparks or smoke was reported but there was evidence of melted plastic on the power supply cable. there was no report of patient or hospital staff compromise and no other system-related devices were reported as suspect.><one ev1000 power adapter cable was received for evaluation and the device history record was unable to be performed as the serial number is unknown. examination of the returned cable confirmed the customer's complaint it was observed that the power cord connections and the power supply leads were burnt due to fluid penetrating the inside of the power port connection. it was not possible to connect the power cord to the power supply for further testing due to the damage sustained. a corrective and preventative action has been generated to address this issue. no further actions will be performed at this time. edwards will continue to review and monitor all events. if action is required an appropriate investigation will be performed. | 1
5846338 | .><while the patient was being monitored the screen froze and the patient was un-monitored for several hours. this is the third ev1000 in the last year to have similar problems. manufacturer response for non invasive cardiac function monitor, ev1000 (per site reporter): will replace the defective unit. | 1
5737710 | a spark came from an icu room with a patient using ev1000 when the power cord was attached to battery pack. patient was moved to another room and a loud explosion along with smoke billowed out. fire from ev1000 and power surge protector. iv fluids above may have dripped down into electrical unit. rca done found all our ev1000 units had been installed by the manufacturer rep inverted posing risk of recurrence.><. | 1
5716777 | at end of procedure, the surgeon noticed that the continuous cardiac output (cco) monitor was not functioning properly. the surgical staff, biomed, and unit management troubleshooted the problem for over 2 hours before a reading could be obtained. the manufacturer was contacted and provided direction over the phone during the course of the problem. cco monitor was changed x3, cco cable was changed x5, and swan ganz catheter was changed x2. the problem was only resolved when the patient was moved to a separate or. manufacturer, biomed, and unit director unable to determine cause of monitor failure. of note: the first two swan ganz catheters used were of the same lot number. the third catheter that was used successfully had different lot number. cco reading confirmed and surgery was concluded.><. | 1
5350104 | both catheters would not communicate with computer. both seemed to be having sensor issues with the catheters. both catheters were not showing any signals when connected to the cable, and they were removed at the time. a third catheter was functioned fine and had signals on the display. the procedure was completed smoothly and no injury to the patient. we have reported numerous of these catheters manufacturer response for bidirectional navigation catheter, nav st thermacool smarttouch (per site reporter): this is the 7th such catheter we have reported about this device. there has been no response from mfr.><. | 1
5614290 | during procedure, catheter showed artifact in the intra cardiac ecg. replaced cable connecting catheter to carto unit with no resolution. doctor and vendor representative in room determined product was defective.><. | 1
6002807 | the procedure had to be cancelled because sensis equipment in cardiac cath lab was not working properly. patient was ready in pre op with iv in place when it was determined that we were not going to be able to do the procedure. the sensis system was upgraded on with new software in early december. this ep case was the first case after upgrading the system. it was not working properly and siemens was contacted and their support team also had difficulty fixing the system. the applications rep was present trying to fix it, and found the problem with the system. the system problem was fixed but we still have problems with the sensis system for eps procedures and we are working with siemens to found a solution to our problems and potential need for a different ep system may be required.><. | 1
6211010 | it was reported that a patient underwent an atrial flutter (afl) procedure with a carto 3 system. a map shift with no error message was reported with the use of the carto 3 system during the procedure. there was no movement of the patient or patches and there was no metal interference. it was stated that the catheter seemed to have dropped into a pouch. no troubleshooting could be performed. the procedure was completed with no patient consequences. multiple attempts have been made to obtain clarification to this complaint. however, no further information has been made available. this map shift with no error message was assessed as a reportable malfunction as such map shifts could potentially be caused by a system malfunction and there would be a potential risk to the patient.><manufacturer's ref. (b)(4), it was reported that a patient underwent an atrial flutter (afl) procedure with a carto 3 system. a map shift with no error message was reported with the use of the carto 3 system during the procedure. there was no movement of the patient or patches and there was no metal interference. it was stated that the catheter seemed to have dropped into a pouch. no troubleshooting could be performed. the procedure was completed with no patient consequences. the biosense webster field service engineer arrived on site for troubleshooting and repair. there was noticeable damage to the location pad cable. exposed wires were found. the biosense webster field service engineer was unable to initialize with the current location pad cable. the biosense webster field service engineer confirmed that the reported issue was due to the damaged location pad cable. upon repair, the biosense webster field service engineer disposed of the location pad cable, as there was clear visible damage that the cable was torn away from base and this was the root of the reported issue. full atp/preventative maintenance was successfully completed. issue is resolved. system is ready for use. the device history record (dhr) review was performed by the manufacturer and no anomalies were noted in manufacturing or servicing of this equipment.><the hardware investigation has begun but it has not been completed at this time. when the investigational analysis has been completed, a supplemental 3500a report will be submitted. if additional information is received regarding this event, a supplemental 3500a report will be submitted to the fda. (b)(4). | 1
5737287 | (b)(4). it was reported that a map shift occurred during a procedure. user couldn't find the right position of the vein and she had to use fluoroscopy to know the exact position of the catheter. case completed with no patient consequence. issue was related to system behavior. as describe in ifu (v4 section on chapter 2) when the relative position between the patches remains the same but the position of the heart relative to the back patches has changed the system will not give a warning and an incorrect map (map shift) might be generated. the sedation to the patient during the case was causing a movement of the hearth related to the back patches and this was clearly evident by the analysis of the patches position through the extended features. avoiding sedation peaks during the case solved the issue. the issue did not occured the following cases. system was ready for use. the history of customer complaints associated with carto 3 system # 50119 was reviewed. two (2) out of 25 additional reported complaints may be related to the reported issue. dhr review was performed by the manufacturer and no anomalies were noted in manufacturing or servicing of this equipment.><it was reported that a map shift occured during the procedure. the procedure was completed without patient consequence using fluoroscopy to visualize the catheter. no other information was available. this is information by itself is not assessable for reportability. additional information received on 05/26/2016 clarifies the circumstances in which the map shift occurred, and quantifies the amount of catheter displacement caused by the shift: the map shift was approximately 6-12 mm and carto did not display any error messages. cardioversion had not been performed, and no patient movement had been noted at the time the map shift was discovered. this issue is mdr reportable as such map shifts could potentially be the result of system malfunction, and there is a potential risk to patient. this clarification of the initial event description provides us the information necessary to classify this event as possibly reportable, and thus the alert date has been changed 5/26/2016. | 1
5919219 | manufacturer's reference number: (b)(4) it was reported that a nurse fell and broke her elbow secondary to a screw on the floor. the detached plastic screw from the carto door locker indirectly contributed to the event. the door locker does not impact the carto system?s functionality, safety or performance. no complaints related to the cart door locker have been reported since the carto system was released to the market. also, the manufacturer reported they have not received any complaints related to the locker. a device history record (dhr) review was performed by the manufacturer, and no anomalies were noted in the manufacturing or servicing of this equipment.><the hardware investigation has begun but it has not been completed at this time. when the investigational analysis has been completed, a supplemental 3500a report will be submitted. (b)(4).><it was reported that a nurse fell and broke her elbow secondary to a screw on the floor. an image of the screw was evaluated by an engineer, who believed it to be a part from the cart where the carto system is stored. follow-up information has been requested. upon receipt of the information, the case will be re-evaluated and a supplemental report will be submitted. since this adverse event might result in permanent impairment of a body function or permanent damage to a body structure, it is to be considered serious and mdr reportable. | 0
5548254 | it was reported that a patient underwent a paroxysmal atrial fibrillation procedure with a carto 3 system and a map shift without error message occurred. after performing circumferential ablation of the right pulmonary veins and posterior wall of the left posterior veins, a substantial shift in the location of the ridge between the left atrial appendage and anterior left veins was observed. while using a smarttouch catheter, it was noticed that the amplitude and direction of the force did not match the fast anatomical mapping (fam) map. a new fam map was created which revealed a shift in the location of the ridge on the old map. on the new map, when the catheter was along the ridge, it projected into the left veins on the original map. the new map appeared to be rotated anteriorly versus the old map. there was no cardioversion performed and the patient did not make any noticeable movements. the new map was used to finish ablation. the procedure was completed with no patient consequence. this event is mdr reportable because map shifts with no error message could potentially cause a potential risk to the patient.><(b)(4). it was reported that a patient underwent a paroxysmal atrial fibrillation procedure with a carto? 3 system and a map shift without error message occurred. a bwi representative reported to the field service engineer (fse), that the case data was uploaded to the server, for analysis by device manufacturer. fse confirmed that the system is fully functional. device manufacture and research and development investigated the issue. this was a user error issue. the device manufacture analysis: it looks as there was an electromagnetic noise (metal, probably the fluoroscopy) issue at this site. there was a change in back patch metal values which does not reach level cause warning. this change in metal causes back patch position change (due to the metal effect on position calculation). this change is in non-uniform manner which does not preserve their inter patch distance, for example, did not move as a rigid body. this caused a map-shift as there is no available bcs compensation at such case. the device history record review was performed by the manufacturer and no anomalies were noted in manufacturing or servicing of this equipment.><the hardware investigation has begun but it has not been completed at this time. when the investigational analysis has been completed, a supplemental 3500a report will be submitted. manufacturer's ref. no: (b)(4). | 1
5632139 | it was reported that a patient underwent an atrioventricular nodal reentrant tachycardia (avnrt) procedure with a carto 3 system. after the first radio frequency application, a map shift occurred and no error message displayed. the catheter appeared to be outside the existing fast anatomical mapping shell area. there was no cardioversion performed. also, the patient did not move. there was no active troubleshooting performed. the procedure was completed successfully with no patient consequence. the reported map shift is assessed as a reportable malfunction. it could be caused potentially by a system malfunction and therefore is a potential risk to the patient.><due to the august 2015 fda maintenance where the 3500a codes were updated, the 3500a codes (evaluation codes) will be added until the biosense webster system is also updated. (b)(4). (b)(4). it was reported that a patient underwent an atrioventricular nodal reentrant tachycardia (avnrt) procedure with a carto 3 system. after the first radio frequency application, a mapshift occurred and no error message displayed. no active troubleshooting could be performed, the case completed successfully. the biosense webster field service engineer contacted the biosense webster field representative regarding the issue. the customer has performed cases since the initial issue and there has been no reoccurrence of the issue. the biosense webster field service engineer informed that the system is operational. the device history record (dhr) review was performed by the manufacturer and no anomalies were noted in manufacturing or servicing of this equipment.><the hardware investigation has begun but it has not been completed at this time. when the investigational analysis has been completed, a supplemental 3500a report will be submitted. manufacturer's ref. no: (b)(4). | 1
5496643 | .><all mac-lab screens went dark (two in the back control area, one in the procedure room itself). | 1
5414003 | the bwi failure analysis lab received the device for evaluation. the analysis has begun but is not completed at this time. when the investigational analysis has been completed, a supplemental 3500a report will be submitted. (b)(4)><it was reported that a patient underwent an atrial fibrillation procedure with a 20 pole eco cable and a noise issue occurred. when the lasso eco was connected to the carto 3 system there was noise on all of the electrograms. the catheter cable was changed which did not resolve the issue. the eco dongle cable was changed which resolved the issue and the procedure was continued. the procedure was completed with no patient consequence. multiple attempts have been made to obtain clarification to this complaint. however, no further information has been made available. since it is unknown if noise occurred on all the channels, including the 12 leads of body surface electrocardiogram (ecg) and all intracardiac ecgs on the carto and recording system, and it is also unknown if there was any ecg signal available for the physician to monitor the patient's heart rhythm, bwi conservatively reports this event.><(b)(4). it was reported that a patient underwent an atrial fibrillation procedure with a 20 pole eco cable and a noise issue occurred. the returned device was visually inspected and one pin was missing on p1 connector side. then a carto 3 test was performed with the cable and a known good catheter and no malfunctions were observed. no errors or noise were displayed. all the cables are inspected for visual damages before packaging. on line inspections are in place to prevent this type of damage/defect from leaving the facility. the device history record (dhr) was reviewed and no anomalies were found related to this complaint. in addition, the dhr review verifies that the device was manufactured in accordance with documented specification and procedures. catheter failed during the visual inspection. however based on the available analysis finding results, the failure mode reported by the customer does not appear to be caused by any internal bwi processes; since cable pass all the functional tests. the root cause of the broken pin cannot be determined. | 1
5837126 | poor signal during procedure. when the catheter was connected, the image on the display was poor. the catheter was removed, and a new catheter was used. the new catheter worked fine with clear image. no injury to the patient.><. | 1
5390311 | .><completed case on the host" x-per station fellow went to review waveforms on another x-per station and waveforms would not load. fellow had to use the "host" x-per system in order to be able to review the waveforms." | 1
6105472 | cath lab reports that the mac lab monitoring equipment stopped working. staff unable to record bp, pulse oximetry, hr and arterial line. staff report mac lab is losing communication during patient acquisition.><. | 1
5797764 | .><the patient was prepared for an atrial fib ablation. upon testing the patch system prior to the procedure, an elevated impedance was noted and the patch was changed. the procedure proceeded with no further problems noted. there was no patient harm noted. | 0
5339670 | the merge hemodynamic recording system is used to monitor and record o2 saturation while performing cardiac catheterization procedures. it was reported that the spo2 value display was freezing on the hemo monitor. if pulse rate is measured from the spo2 values, then pulse rate display also freezes, which causes the hemo monitor to display stale data><upon retrospective review, this issue was determined to be reportable as an mdr. | 1
5339668 | upon retrospective review, this issue was determined to be reportable as an mdr.><during cath lab procedures while using the phasein etco2 option for merge hemo, the system could occasionally become non-responsive to user actions (freezes up), the record button might become grayed out and unable to start /stop recording, or multiple short recordings might be displayed versus the typical long, continuous recording. a customer reported that they observed a display of rapid repeated recordings after disconnecting the etco2 module. | 1
5437917 | merge hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. on (b)(6) 2014, the site reported that there was a hemomonitor error while a patient was on the table. they lost vitals on win7. there was no direct harm to the patient due to this issue. the hemo pc was swapped and the system sent for evaluation by engineering. this error may impact the physician's ability to continue vitals monitoring during cath procedures. (b)(4).><upon retrospective review, this issue was determined to be reportable as an mdr. | 1
5437923 | upon retrospective review, this issue was determined to be reportable as an mdr.><merge hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. on (b)(6) 2014, hemomonitor.exe was reported to have stopped working unexpectedly. the error was reported as: software_hemo-6092_lab2 hemomonitor.exe. merge assisted wih re-imaging hemo pcs. this error impacts the physician's ability to continue vitals monitoring during cath procedures. | 1
6174109 | siemens is conducting a thorough investigation of the reported events. as this event is under investigation, a root cause has not yet been determined. a supplement report will be filed upon completion of the investigation.><it was reported to siemens that a malfunction occurred while operating the axiom sensis system. the customer reported that a patient was ready in pre op with an iv in place when it was determined that the system in the cardiac cath lab was not working properly. the procedure was cancelled and we are unaware of any impact to the state of health of the patient involved. | 1
6131721 | (b)(4)><(b)(4) date sent: 1/3/2017. per service manual operational and diagnostic analysis does confirm the pump door does not stay latched due to cracked latch.><it was reported that there is an issue with the door latch on the device. the pump door will not stay closed. there was no patient involvement. | 0
6188550 | at the beginning of the surgery, surgeon tried to install the mayfield adapter but it resisted to tightening movements. as it was difficult to immobilize the patient, surgeon decided to use a hammer to tighten the part.><device has not been made available yet. | 0
5881063 | (b)(4).><patient contacted dexcom on 07/24/2016 to report that the receiver displayed a hardware error on (b)(6) 2016. at the time of contact, no injury or medical intervention was reported. no additional event or patient information is available. | 1
6181632 | the reported complaint was not confirmed. during laboratory analysis, the product surveillance technician (pst) removed the three wires from the plastic connector housing for inspection. the plastic connector was purposely damaged during removal in order to preserve the condition of the crimped terminals. per drawing, both red wires measured approximately 1k measure of resistance (ohm) to the white wire. actual measurement taken at 23.5?? celsius was 1.09k ohm (typical). magnified visual inspection of each crimp proves all three were properly assembled and crimped with no intermittency of contact. no further testin"was performed. if additional information becomes available on this complaint that would alter the facts and/or conclusion, a supplemental report will be filed accordingly side b had the issue and the function stopped for a short period of time but the error cleared and both sides were able to be used to complete the procedure. no change out was needed. the case was completed successfully without delay and without associated blood loss. there was no harm observed.><(b)(4). the field service representative (fsr) tested the unit and could not verify any errors. he checked the unit and found no evidence for loose connections and corrosion. the fsr replaced the resistance temperature detector (rtd) as a precaution. the unit operated according to manufacturer specifications and was returned to clinical use. as per perfusionist they no longer have the problem since the sensor was replaced. the suspect part was returned to the manufacturer. this complaint is related to (b)(4) / medwatch #1828100-2016-00798. if additional information becomes available on this complaint that would alter the facts and/or conclusion a follow-up emdr will be filed accordingly." | 0
5881848 | orders are created on the cpoe machine and sent electronically to the intended recipients, including nurses, ancillary services, labs, and pharmacies. it has been a widespread and formidable problem that these digital orders do not reach the recipient and are not carried out in a timely manner. to make these matters worse, there is not any easy mechanism on these computerized medical devices to reconcile orders with their execution and/or results. this is true for all vendors of these devices. in this case, because of a decrement in hemoglobin level, in a critically ill patient with multi system failure, a stool guaiac test was ordered. the patient was being treated with prophylaxis for deep vein thrombosis, which was continued. it was not until 5 days later when the patient was acutely bleeding that the test was done. the patient went into shock requiring multiple interventions. this case and the genre of problem it represents, is disturbing on many levels, including the fact that the fda does not provide oversight of cpoe systems of devices as it does for all other apparatus intended for use in the diagnosis or cure of disease. | 1
5528946 | with the work flow controlled by the electronic care record system, vital sign data is out of sight, tucked away is a silo that requires multiple clicks and time to access. consequently, the key data of vital signs is out of sight, and thus, out of mind. no one knows the pt data, yet medications are administered resulting in complications because of the inappropriate (for the pt's vital signs) administration of therapies. in this particular case, the pt was hypotensive and received add'l blood pressure lowering medication resulting in critical condition. | 1
5525718 | stat arterial blood gas ordered pre-op in pt or parenteral narcotic who was in delirium. the test was not done. the pt underwent an operation. after the operation, the pt became unresponsive and was found to be in respiratory failure requiring critical care. the order was not carried out by the ancillary service department. no one on the operating team saw the order such that it would be done prior to the anesthesia (not the anesthesiologist, surgeon, nurse, no one.) the recipient was not in receipt of the order, operating room, the recipient did not timely see it. communication has been disrupted with the deployment of these care controlling devices. this is a widespread problem. | 1
5470041 | datascope cs300 intra-aortic balloon pump therapy was initiated in the cath lab on an acute mi patient. after transport to icu, the console repeatedly displayed the alarm autofill failure." troubleshooting did not solve the alarm and the second iabp console was retrieved and replaced the first. no further "autofill failure" alarms on the replacement console. no apparent patient harm." | 0
6080963 | ><problems with high hounsfield units (hu#s) were identified shortly after the ge discovery hd750ct was placed into service on 10/30/2014. initial medical physics testing on 10/30/2014 revealed a high hu bias on the hu # of water (subsequent annual medical physics testing in 2015 and 2016 verify the consistently high hu# of water in adult abdomens). in february 2015, ge healthcare was first notified of instances where renal cysts had elevated hu#s (compared to the same exams on the same pt on different ct scanners). a ct scan of the chest/abdomen/pelvis was performed on (b)(6) 2016 on the ge discovery hd750ct at the (b)(6). lesions along the posterior aspect of the right lobe of liver measured 30 hy (hounsfield units). these were interpreted as metastatic disease. the pt had another ct chest/abdomen/pelvis done on (b)(6) 2016 on our philips ct scanner. the same lesions measured 14 hu. these are in consistent with cysts, not metastatic disease. this was later confirmed to be true with a follow up mri. in this event, the falsely elevated attenuation measurement on the ge ct scanner caused significant misinterpretation. multiple requests have been made to ge healthcare (i.e., service requests and direct request to ge health senior director and senior engineering for physics and image quality); however, high water hu# artifacts persist. therefore, (b)(6) staff radiologists do not have clinical confidence in the ge discovery ct750 hd scanner for diagnostic abdomen/pelvis studies. the majority of ct scans are abdomen/pelvis diagnostic studies; therefore, utilization of the ge discovery ct750 hd is severely limited. | 0
5431165 | ><patient was having a composite filing done of tooth #3 when she began to experience a burning sensation inside and outside the mouth while the staff was using the curing light. | 0
5532837 | (b)(4).><it has been reported that the device is not functioning properly.><(b)(4) | 1
5440767 | the manufacturer received the bipap for investigation and there were no operational issues noted. there was no evidence of a failure of the device that would have resulted in patient data from a previous patient being included in another patient's data. the root cause of the event is due to the operator of the device not clearing the memory for the bipap and associated modem and sd card. the modem management user guide (pn 1109520) instructs the user, be sure to clear the memory of any previously used therapy devices modems smartcards or sd cards prior to reissuing them to other patients. consult the documentation included with the therapy device or modem for instructions on clearing its memory. use the data card utilities to clear the memory of smartcards and sd cards. the data card server must be installed and running on your computer in order to access data card utilities. the data card server utility for installation can be found by clicking on the encoreanywhere pre-requisites page hyperlink on the login page for encoreanywhere." there was no patient involvement and no risk of harm or injury if the event should recur. based on the information available the manufacturer concludes no further action is necessary.><the manufacturer was made aware of an allegation by a customer that encore anywhere (ea) software was responsible for data from one patient being included in another patient's data when downloaded from an a30 bi-level positive airway pressure (bipap) device. there was no direct patient involvement and no patient harm or injury was reported. the alleged incident occurred on (b)(6) 2015 but was not communicated to the manufacturer until 01/21/2016. the manufacturer has requested return of the a30 bipap sn (b)(4) material number (b)(4) for investigation in this alleged event. the manufacturer is currently investigating this allegation. preliminary assessment of the downloaded data indicates that the customer did not correctly erase the memory of the therapy device after the device was removed from the first patient thus allowing additional data from the first patient to be included in the second patient's data. the manufacturer's investigation on-going. a follow up report will be submitted when the manufacturer has completed the investigation. note that this software is associated with these 510(k) numbers: k993307 k012323 k061034 and k121623." | 1
5962724 | it was reported that there was data missing in this customer's t:connect data management system account. reportedly, the customer's bolus deliveries were not logged into their account roughly from (b)(6) 2016. during troubleshooting with customer software support, it was determined that the t:connect bolus graph indicated that the customer only had one bolus delivered during the week; however, the customer reported that they were using the pump the entire time.><.><follow up on 08/30/2016 confirmed that the customer was using the pump and delivering boluses regularly during this time. as the customer did not have the pump during the time of the call, no troubleshooting was able to be performed. although requested, no additional information was provided regarding the reported issue. no product was returned for evaluation. should new relevant information become available, a supplemental report will be submitted. device not returned. | 1
5940041 | no product was returned for evaluation. should new relevant information become available, a supplemental report will be submitted.><it was reported that data in the customer's t:connect data management system account was inaccurate. reportedly, the profile account displayed excessive bolus deliveries. during troubleshooting with tandem's customer technical support (cts), the contact reported that the customer did not input these deliveries. cts verified in the engineering logs that all bolus deliveries were delivered, and that the data reflected in the t:connect account is accurate. reportedly, there was no impact to the customer's blood glucose level. multiple attempts were made by cts to obtain additional information; however, no additional information regarding this event was provided. | 1
5621334 | customer reported receiving an e6 message on the display of their adc blood glucose meter and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event.><this is a known malfunction with the precision link software that can lead to incorrect trending results. this occurs when results, obtained on a meter with incorrect data and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc (b)(4) letter. the device manufacturer date for the suspect device is unknown. the date entered is the date abbott diabetes care became aware of the event. all pertinent information available to abbott diabetes care has been submitted. | 1
5790967 | this is a known malfunction with the precision link software that can lead to incorrect trending results. this occurs when results, obtained on a meter with incorrect data and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec006 letter. the device manufacturer date for the suspect device is unknown. the date entered in section h4 is the date abbott diabetes care became aware of the event. all pertinent information available to abbott diabetes care has been submitted.><customer reported that her daughter received an e6 message on the display of her adc blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
5982027 | the customer contacted a siemens technical service center (tsc) specialist. tsc reviewed the data provided. the tsc found that the centralink data management system omitted a test. tsc verified the centralink configuration and no errors were found. tsc modified the frequency of the omit_command_2 task from every one minute to every five minutes. a siemens headquarters support center (hsc) specialist reviewed the data supplied. hsc stated the system was configured correctly. all hold tests were setup for both vista instruments as per the vista with aliquot support documentation. hsc stated that the hold tests were reflexed appropriately by the sample tube re-run button. hsc found that the hold tests were omitted by centralink before the customer had the ability to setup the vancomycin test for a manual re-run. hsc was able to duplicate the event. the cause of the discordant, vancomycin results was due to centralink running a batch command to omit the hold tests". the instrument is performing according to specifications. no further evaluation of the device is required.><discordant vancomycin results were obtained on two dimension vista instruments that were linked to a centralink data management system. the customer initially received an above assay range result on their dimension vista 2 instrument (>50). the customer followed proper protocol to repeat the sample using a "sample tube rerun". the customer then performed a manual 1:6 dilution and set the sample to run in centralink. the customer then manually programmed the sample to run on their dimension vista 1 instrument. a falsely elevated result of >250 obtained from dimension vista 2 populated the current result column in centralink followed by a falsely low result of 9.8 obtained from dimension vista 1. the depressed result (9.8) was not released to the physician(s) as the operator who performed the dilution knew the result should be higher. it is unknown if a result was released to the physician(s). it was determined that the "hold" tests function was properly triggered when the customer used the "sample tube rerun". however the "hold" tests were omitted by a "batch" (not by user) function/mode. because the "hold" tests were omitted centralink sent the order again to dimension vista 2 instrument which caused the result to be run neat. when the repeat result from the dimension vista 2 instrument transmitted to centralink as >50 a dilution factor of 6 was applied on a neat sample thus obtaining an elevated result of >250. minutes later the low result of 9.8 was received from dimension vista 1 but the dilution factor had already been applied to the result of >50 thus obtaining a falsely low result. there are no reports of patient intervention or adverse health consequences due to the discordant results." | 1
6207650 | the customer contacted a siemens customer care center (ccc) specialist. review of the sample showed that it was clotted. the ccc reviewed the instrument data and found that the instrument had a normal setting of -1 for alkaline phosphatase. the ccc reviewed the normal ranges and found the instrument was not set up to hold a result of <10. the customer added 10 as the low limit on normal evaluation orders. the ccc testing the instrument and a result of <10 was held in review. the cause of the depressed, less than ten alkaline phosphatase result being released is due to human error, a setting was not set. the instrument is performing according to specifications. no further evaluation of the device is required.><customer reported that a depressed, less than ten (<10) alkaline phosphatase result was not held in review and was auto-validated by their centralink data management system. the initial result was released to the physician(s), which was questioned. the patient was redrawn and tested, resulting within the expected range and clinical picture of the patient. a corrected report was issued to the physician(s). there are no known reports of patient intervention or adverse health consequences due to the depressed, less than ten alkaline phosphatase result being released. | 1
5998099 | siemens healthcare determined that the cause of the discordant results was due to an equation error typed into a new centralink installation at the customer site. the equation was incorrectly configured to multiply the relevant ratio by 1000 instead of 100. a siemens technical application specialist corrected the equation. the instrument is performing according to specifications following the specialist's actions. no further evaluation of the device is required.><discordant results were obtained with the microalbumin to urine creatinine ratio calculation performed by the centralink data management system. the discordant results were falsely elevated by a factor of 10, and were reported to the physicians. the reported results likely influenced clinicians to prescribe treatment, such as blood pressure medication. specific details were not provided by the customer. the discordant results were obtained as a result of an equation error in the centralink data management system. once the equation was corrected and verified, 186 corrected reports were issued to the physicians. it is unknown if there were adverse health consequences due to the discordant results.><the initial mdr 2432235-2016-00590 was filed on (b)(6) 2016. additional information (b)(6) 2016: a siemens headquarter support center (hsc) specialist investigated the event. the customer's old centralink was being replaced with a new server, and rules and configurations were manually copied by siemens to the new server from a customer maintained workbook. on (b)(6) 2016, the customer emailed the workbook to a siemens automation informatics specialist for use in their new centralink. hsc reviewed system files and logs received from the customer. hsc verified that configurations and rules were entered into the new system by siemens support per the settings documented in the customer workbook. upon review by siemens hsc specialist it was determined that the customer had changed the units for reporting the microalbumin method but did not update the calculation on the spreadsheet. the default calculation in the customer workbook for microalbumin was in mg/dl (x1000 calculation), not mg/l (x100 calculation). in conclusion, the error was introduced into the customer maintained workbook by the customer. the instrument is performing according to specifications. no further evaluation of the device is required. | 1
5886886 | the (b)(6) public health newborn screening laboratory uses the (b)(6) specimen gate screening center (sgsc) software (a class i medical device) to store, retrieve, and process the data associated with specimen testing including but not limited to patient demographics, tests ordered, test results, test result determination (interpretation), quality control results, and result codes (flags) that may be associated with the specimen from its entry into the laboratory workflow until patient report is generated and released. when sgsc software is installed, configurable settings are programmed based on the customer preferences. for instance, each laboratory enters their specific population based reference ranges that are used to evaluate if a patient's test results are normal, below or above the reference range and require further action such as repeat testing. this laboratory specific logic automates the flow of the specimen through the laboratory. the customer defines the nomenclature used for their result codes (flags), and when the software encounters the predefined result code the software processes the specimen according to the logic associated with the result code. for instance in (b)(6) the result code tbc" means "to be confirmed". the software will apply this result code to an initial analytical result that is outside of the reference range. the software then based on the tbc result code flags the specimen as one which requires repeat testing to be performed. to perform the repeat testing new punches of the dried blood spot specimen are required; the software identifies this to the customer. the software is configured during installation based on the procedural workflow of the laboratory. when a laboratory technician deviates from their routine procedural workflow the software may not be able to recognize the deviation and account/correct for the event if the specific procedural deviation wasn't defined in the software during installation. in this specific scenario a newborn sample was entered into sgsc by the laboratory samples were punched from the dried blood spots and testing was completed for the (b)(6) defined disorder panel. one of the tests run by the laboratory referenced as apga generates results for both the galt and biotinidase markers. when the apga galt marker result is abnormal a tgal screen is then ordered. the tgal test profile defined by (b)(6) includes running tests for tgal and galt markers. when the apga biotinidase result is abnormal the apga assay is identified for repeat testing which would include repeat testing of the btd and galt markers. in this specific scenario both galt and biotinidase results were abnormal in the original apga assay and the laboratory repeated the btd marker and within the tgal profile ran both tgal and galt. upon apga repeat the biotinidase was normal but galt remained abnormal. upon tgal testing the tgal was normal but galt remained abnormal. the report should have resulted in an "abnormal" determination (galt repeat abnormal tgal repeat normal) however on (b)(6) 2016 the reported result determination was "normal" for both galt and tgal. the correct disorder comment was included on the report by the laboratory: "test for galt enzyme defect is positive which indicates possible galactosemia. if not previously contacted by the emory newborn screening follow-up program please contact them at (b)(6)." a corrected report was released two days later on (b)(6) 2016. due to the abnormal screen report a repeat specimen was received for testing and was tested on (b)(6). the repeat specimen tested normal for galt. in the reporting phase sgsc evaluates result codes associated with the analytical result against a library of defined result codes which directs the software on which analytical result determination to print in the report. during sample punching an issue occurred which led to one of the confirmation apga assay to be identified as initial producing tbc result codes. in the routine procedure workflow of the laboratory the tbc result code (indicating initial test phase) should not reach the reporting phase and therefore is not defined in the result code library for reporting. in this situation encountering an undefined result code in the reporting phase sgsc is configured to use a default determination which is "normal". in the event there is more than one assay for a single marker such as in this case for galt the order in which the confirmation samples for the tests are punched dictates which result is printed in the report. the last assay run is reported by sgsc. in this case it was the apga galt result reported on the original report. the laboratory chose to report the galt result from the tgal assay instead of the apga assay when the corrected report was created. both test results were abnormal for galt.><on (b)(6) 2015 (b)(6) initiated recall z-0166-2016 for specimen gate screening center. (b)(6) was issued the recall letter on (b)(6) and the signed response form was received the same day. (b)(6)has been working with the (b)(6) newborn screening lab to identify the result code classifications impacted by the software anomaly. since the recall issuance(b)(6) newborn screening lab have held routine meetings to identify and resolve the result code classifications. implementation of the final corrective action is targeted to be complete by (b)(6) 2016." | 1
6136094 | a (b)(6) female pt was being discharged from our hospital with prescription for aspirin 40.5 mg chewable tablet (1/2 of 881 mg chewable tablet). we use cerner at our facility. the discharge medication reconciliation was done through cerner as is our process, but the dose of the drug does not transmit to the pt summary that is provided to the pt. only the product and frequency transmit. the dose does transmit to any prescriptions but since that was an over the counter medication, the parents were using the pt summary to know what to give their child. it appeared on the pt summary that the pt should receive the entire 81 mg tablet as their dose. this was caught by the physician because this particular physician had been involved in another error because of this same issue. the family was educated on the correct dose to administer to their child, but the pt could have received a 50% increase in the amount of aspirin they were supposed to receive. medication administered to or used by the pt: yes. where did the error occur: hospital. (b)(6). | 1
5761049 | during surgery, the sdc video capture device was used to save numerous photos of the surgery. photos are used to document procedure and become part of the patient's medical record. photos may be needed to get reimbursement for some procedures. at the end of surgery, it was discovered that most of the photos were not captured. we have been having an abnormally high failure rate of the dvi video input boards on these units. in one area, 7 out of 8 have failed. there have been several documented delays during surgery caused by failures of the recording device. not related to the dvi input, we have also had problems with the software becoming corrupted if the unit loses ac power, like during the emergency generator test. in most cases this can be corrected by reloading the software. software version is 1.3.1. we believe the version hasn't changed since we bought these. we think this is mainly a hardware issue with the dvi input board. manufacturer response for video capture device, (brand not provided) (per site reporter): the dvi input board failed and needs to be replaced.><. | 1
5856194 | ><patient was positioned for a laparoscopic roux-en-y gastric bypass procedure. during the procedure the patient had slid down the operating room table approximately 9 inches. the patient was positioned and secured as usually required for the procedure. a foot board was not in place for this procedure. pt. on the operating room table in reverse trendelenburg position. | 0
5856210 | ><patient was positioned for a laparoscopic gastric bypass roux-en-y procedure. during the procedure the patient had slid down the operating room table approximately 2 inches. the patient was positioned and secured as usually required for this procedure. a foot board was not in place for this procedure. pt on the operating room table in reverse trendelenburg position. | 0
5952897 | it was unknown if the initial reporter sent report to the fda.><it was reported that the patient alleges the bolus recommendations provided by the software application are not accurate. no adverse event reported. no product will be returned for evaluation and backup files are unavailable. | 1
6137318 | it was unknown if the initial reporter sent report to the fda.><it was reported that the patient alleges the bolus recommendations provided by the connect app ((b)(6)) are not accurate. the patient had low blood glucose levels from the incorrect bolus advice. no adverse event reported. the device is not expected to be returned for product evaluation. | 1
6099009 | while checking blood pressure with automatic machine, numbers rapidly escalated high and cuff tightened down causing pain to patient's right arm. | 1
5986816 | equipment failure during 3d mammogram.><. | 0
6085003 | (b)(6) center had a patient on the table for a dcs implant and had failed to obtain an implant for the surgery. (b)(6), the sales rep in that area, did not have any inventory on hand. (b)(6) contacted (b)(6) whom was able to provide an implant for the surgery. the item they provided was lot # 107903, part number 17111. this item was packaged in july of 2009 and expired in july of 2016. (b)(6) was contacted by biopro in early 2016 to send that inventory back, (b)(6) indicated that they had no inventory of that lot number and biopro was unable to retrieve it for repackaging. (b)(6) kept expired inventory, (b)(6) received the inventory, which was labeled, and never checked the expiration date on the label prior to implantation. this lot # did indeed expire in july of 2016; the job packet with the label is attached to the complaint. upon review of the real time aging study for the tyvek pouches it was noted that although biopro has established the shelf life for the pouches at 7 years, many of the pieces that were tested to determine this were actually aged to 7 years and 7 months. the complete real time aging study is located in regulatory affairs. the aging study was conducted in october of 2003 and the samples tested were packed from march, 1996 through january, 1998. the expiration date is clearly marked on the dcs label, but it is small. the current version of the dcs label is larger and the expiration date is more pronounced.><(b)(6) center did not verify the expiration date of the implant in question prior to implanting it. biopro attempted to retrieve this specific item prior to expiration and was told that (b)(6) did not have it. (b)(6) did indeed have it and provided it to (b)(6) center and it was implanted after the date of expiration. the expiration date was clearly visible on the label. the following letter was prepared and provided to (b)(6): re: digital compression screw, item number 17111, lot # 107903, expiration date 07-2016. at the time of implantation in (b)(6) 2016, the label on this component indicated that the part was beyond its expiration date. this expiration date is in relation to the sterile packaging, not the actual component. the shelf life for product utilizing tyvek pouches like the dcs is packaged in has been established by biopro to be 7 years from date of packaging. this time was established in 2003 as a result of a real time aging test conducted by an outside contractor. the pieces tested were at that time up to 7 years and 7 months post packaging, which exceeds the 7 year and 3 month age of the piece in question. test results of all components of the real time aging test were reviewed and all pieces tested met or exceeded the acceptance criteria to be considered sterile. based on the testing conducted and the age of the piece implanted i am certain that the device in question maintained its sterility and there is no risk to the patient. device is implanted and not available. | 0
5612222 | dental x-ray unit tube head leaking insulating oil. this is the 2nd dental unit of same manufacturer that has same problem within 1 1/2 years. same manufacturer and lot. leaking coolant oil from x-ray tube head on two gendex expert dc dental x-ray units. this oil surrounds the x-ray tube (heats up) to keep it within an operating range. the gasket / seal failed. i do not know how much oil is contained in the unit, nor do i know how much oil leaked out of the 2 units. thus, i had to discontinue use for fear it would expose the patients to overdosed radiation. dates of use: (b)(6) 2008. reason for use: dental pathology/x-ray. | 0
6055753 | .><transfer ribbon started smoking in the robot. | 1
5350469 | upon return of the faulty transformer, the investigation will begin to determine cause of fire and transform failure. on completion of investigation, the conclusion will be made available in a follow-up report.><while in service mode customer heard a noise from behind the facial wall. upon investigation, the customer discovered transformer t13 smoking and a small flame on the filter board. | 0
5986603 | acquired x-ray images display sideways. this has been reported to the mfr (ziehm imaging, inc.) as of (b)(6) 2015 and there has been no clinically viable permanent solution to address this problem. from clinical perspective, this can be pt safety issue as images are mislabeled and can be misinterpreted. we have samples of images that have been sent to the mfr with no response on when this will be fixed. instead workarounds were proposed rather than solving the underlying product problem. these workarounds take more time than is necessary and prolong the pt's medical procedure. | 1
6092403 | this involves an error related to dispill multidose software. the program keeps a manually entered medication profile for each patient who receives the dispill package. when trying to print a dispill card for one patient, then pharmacist erroneously changed the name of the patient, though they thought they were searching for the patient. a different patient's profile had been left up on the screen, which allowed the pharmacist to alter the name. we reached out to dispill to see if there could be an auto-log out put in place, so if a patient was left up for a particular amount of time, the system would log out. we were told that dispill would not do this however. we do have other catches internally in place to make sure that these types of errors don't occur, however the pharmacist did not follow those policies. medication not administered to or used by the patient. relevant material not provided. (b)(6). | 1
6031967 | customer questioning 3 runs of tox panel on triage meter, each time run, different analytes tested positive. customer stated a urine sample from one patient was tested 3 times on 2 triage meters and both meters gave the same results each time, but the results themselves were different for each run. the 1st run- 7 analytes tested positive. the 2nd run (time not provided)- just thc positive. the 3rd run (time not provided)- just thc and bar positive. patient was visiting the hospital for a potential medication overdose. customer was unable to recall patient data due to no patient identifier given on initial report. caller unable to gather information on which specific analytes were positive on the first tests run. at the time reported, no confirmatory testing was performed. although requested no confirmatory results were reported. no reported adverse patient sequela. no further information provided.><investigation conclusion: investigation pending.><investigation/conclusion: received 1 false negative bar result when testing retain lot w61921 with positive tox calibrator 1.5x. review of the device scan showed an abnormally thin and long peak at the bar position, this is the root cause of the false negative result. no issues observed with tox calibrator 0.5x testing. no sample was returned to product support, further investigation cannot be pursued. reviewed the batch record for lot w61921. lot met all final release specifications. | 0
5386541 | on (b)(6) 2016, abbott point of care (apoc) was contacted by a customer regarding an analyzer that was placed in the downloader/recharger (drc) and became hot. the customer states that an analyzer was put into the drc and got hot. the customer removed the analyzer, swapped drcx and the analyzer no longer got hot. the analyzer was using rechargeable batteries at the time of the event. the product (drc) was replaced at no charge and returning for investigation. per i-stat system manual: art: 714368-00k, rev. date: 02-aug-12: the downloader/recharger can recharge a rechargeable battery in the analyzer. if the analyzer contains a rechargeable battery, the battery begins recharging automatically as soon as the analyzer is placed in the downloader/recharger. the downloader/recharger also has a compartment for recharging a rechargeable battery outside the analyzer. placing an analyzer in a downloader/recharger will automatically initiate recharging of the rechargeable battery. the indicator light on top of the downloader/recharger will be green (trickle charge), red (fast charge), or blinking red (fast charge pending) when an analyzer with a rechargeable battery is placed in the downloader/recharger. as well, placing a rechargeable battery into the recharging compartment will automatically initiate trickle recharging. the indicator light near the recharging compartment will be green when a rechargeable battery is placed in the compartment. there are no injuries associated with this event. at this time there is no evidence of a confirmed malfunction and product was replaced at no charge.><apoc incident # (b)(4). the investigation was completed on 04/06/2016. the customer reported that an analyzer was warm to touch when placed onto downloader recharger (drc) s/n (b)(4), and the drc was hot to touch. drc s/n (b)(4) has not been returned to flextronics; therefore, no evaluation could be performed.><(b)(4).><n/a. | 0
5815080 | it was reported that a patient has received a dynesys implant on an unknown date and had to be revised on (b)(6) 2016 due to breakage of the cord. it was also stated the following the removal was from an anterior vertebral body tethering case an off-label use of the dynesys system the recommended technique was not followed during the initial surgery."><the manufacturer did not receive devices x-rays or other source documents for review. as no lot numbers were provided for the devices the device history records could not be reviewed. a cause for this specific event cannot be ascertained from the information provided should additional information become available and an investigation result be available that changes this assessment an amended medical device report will be submitted. zimmer's reference number of this file is (b)(4).><updates: initial reporter age/date of birth weight describe event or problem model #/lot # implant date device available for evaluation? date received by mfr type of reports if follow-up what type? device evaluated by mfr? evaluation codes additional mfr narrative corrections: sex adverse event and/or product problem event problem codes. the manufacturer received the device for investigation on october 5 2016. the investigation is pending. a cause for this specific event cannot be ascertained from the information provided should additional information become available and an investigation result be available that changes this assessment an amended medical device report will be submitted. (b)(4).><updates: date received by mfr type of reports if follow-up what type? evaluation codes additional mfr narrative. trend analysis: no trend identified. the dhr check could not be performed for the cord as the lot number was not available. the product experience report (per) states that a removal was performed of the dynesys top loading system because the cord broke between two screws. the returning explanted materials form lists ruptured vertebral body tether of the thoracolumbar spine as reason for explant. the revision procedure is described as removal of a left vertebral body tether at t11-l3. the form states (b)(6) 2013 as the original surgery date and (b)(6) 2016 as the explant date. an overview sketch for the complete instrumentation of this patient was attached which indicates that a surgery was performed on (b)(6) 2013 at t5-t10 on the right side and at t11-l3 on the left side. some staples are also mentioned and it is noted that the instrumentation on the left side is ?out? on (b)(6) 2016. the discrepancies of the dates are only one day and do not affect the investigation. it is assumed that the dates on the overview sketch are correct. visual examination: the retrieved products were implanted at the levels t11-l3 but were not marked with the anatomical location from which they were removed. however with the information at hand it was possible to reconstruct some of the screws locations. the pedicle screws show none to very slight remains of bone attachments. on all of the screw heads damages from the revision surgery can be observed in form of slight scratches slight deformations and indents. the set screws show none to only very slight damage on the inner hexagon most likely from either the implantation surgery and / or the revision surgery; in general all of the set screws look inconspicuous. three pieces of a cord implanted at the levels t11-l3 a long and two short pieces with unknown lot number were received. all of the cord pieces are slightly yellowish discolored and tissue attachments can partially be observed. the long piece shows the green thread marking of the working zone of the cord at one end which should not be implanted. the cord end shows a clear cut which is melted and merged. the other end is slightly bulged and the fibers of the outer layer seem to be frayed. three contact areas of the screws are clearly visible on the cord piece. the outer cord layer is slightly damaged in the contact areas where the tip of the set screw was placed. on the opposite side of two of the contact areas a dark spot can be observed which correlates with the cannulation of the cannulated pedicle screws. the second cord piece shows one contact area where the outer layer is slightly damaged. both ends exhibit a slightly bulged end with frayed and / or separated fibers from the outer cord layer. one cord end of the third cord piece shows a clear cut which is melted and merged. the other end exhibits a slightly bulged end with frayed and / or separated fibers. one contact area of the screws is visible right next to the slightly bulged end. no other damaged areas e.g. worn areas could be identified on the cord pieces. based on the appearance of the cord ends the locations of the contact areas and the assumption that the clean cut ends are the ends of the entire implanted cord it was tried to reconstruct the cord. the five unknown staples show damage in form of nicks and scratches. the inner thread is mostly damaged and some of them exhibit bone ingrowth. one prong of a staple is fractured. conclusion: the dynesys top loading system used for vertebral body tether of the thoracolumbar spine was revised after approximately 2 years and 7 months in vivo due to a ruptured cord. the patient?s (b)(6) and can be classified according to the bmi scale as obese class i (moderately obese) ((b)(6)). contraindications of the dynesys spinal system and the zimmer dto implant are similar to other commercially available posterior spinal fixation systems including obesity. five unknown staples were received. with the available information at hand it stays unknown for what indication or where they were implanted. the cord was implanted at the levels t11-l3 and it was tried to reconstruct the cord. with the available information it stays unknown which cord end ended at t11 respectively at l3. therefore the anatomical location from where the individual pieces were removed stays unclear. two of the cord ends show a clear cut which are melted and merged while the fibers of the outer layer seem to be frayed on the remaining cord ends. based on the reconstructed cord two possible locations of a rupture could be identified. at those two locations the cord ends have similar bulged and frayed appearance that does not allow a clear identification of a rupture of the cord. the per states a cord rupture in between two screws while the returning explanted materials form lists ruptured vertebral body tether. therefore it is unclear if the cord possibly ruptured in one or two locations. the appearance of the cord ends and the information at hand are not sufficient to conclude if and where a rupture of the cord might have occurred. the reason for a possible cord rupture stays unknown. from the available information it can be concluded that the surgeon performed a vertebral body tethering (vbt) for idiopathic scoliosis as described by (b)(6) et al.. this is off label use of the dynesys top loading system. as the surgical intentions of vbt differ from those of the dynesys system?s intended uses it is also unknown if the cord tensioning during the implantation surgery might have influenced the cord rupture or if the cord tensioned over the time in vivo due to growing of the patient might have been a contributing factor. in conclusion it is hypothesized that the following possible factors might have contributed to the cord rupture either alone or in combination: - cord tensioning during the implantation surgery - cord tensioned over the time in vivo due to growing of the patient - the system being in an unknown/untested loading environment - the patient?s overweight. however an exact root cause could not be identified. the need for corrective actions is not indicated and zimmer (b)(4) considers this case as closed. (b)(4).><it has now been reported that a patient was implanted a dynesys lis stabilizing cord 200 on (b)(6) 2013 and had to be revised on (b)(6) 2016 due to breakage of the cord. it was also stated the following "the removal was from an anterior vertebral body tethering case an off-label use of the dynesys system the recommended technique was not followed during the initial surgery."" | 0
5705315 | e-prescriptions were sent using aria ehr software produced by (b)(4). they were never received by (b)(6) pharmacy on (b)(6). ehr vendor informed, but insists that prescriptions were sent. pharmacy insists they never received them. | 1
5537447 | at our institution, no one sits in front of the ehr mdds terminal surveying the data of each patient as it gets deposited into the ehr mdds. additionally, there is not any alarm or notification that new results arrived, nor is there a notice of any abnormal results. in this case, the serum sodium was dangerously elevated and was in the mdds, but because of the time of arrival and absence of notice, no one knew, resulting in further increase before intervention was ordered. the patient suffered altered level of consciousness and the hospitalization was prolonged. | 1
5690723 | the ehr mdds component fails to provide adequate notice that new lab results have been deposited; and more dangerous, there is zero notice that the silently deposited results have life threatening implications. in this case of atrial fibrillation being treated with heparin and being converted to coumadin (receiving both drugs), the hemoglobin was 7.0 g/dl on a 6 am blood draw. the result was in the silo of the ehr, but no one saw it because they all rounded earlier. it was more than 12 hours later that it was seen by a health care professional, during which time full heparin infusion continued. the heparin was stopped at the time the results were seen but the coumadin had already been administered. the next morning, the hemoglobin was less than 6 g/dl, the patient was bleeding from the gi tract, and there was incipient shock requiring icu transfer. the device was defective in that it failed to alert of new results, which resulted in delay of appropriate care. | 1
5481905 | (b)(4). date of event since (b)(6) 2015. as the reported event was a configuration issue and resolved by over the phone by baxter technical services, no product was returned for evaluation. the serial number was not provided; therefore, a manufacturing review could not be performed. a review of the mix check report indicated that the bags were within range after the configuration changes. should additional relevant information become available, a follow-up report will be submitted.><it was reported that an over-delivery of dextrose was observed during compounding. customer discarded the bags that were out of range. during troubleshooting, baxter technical support determined that the customer had been using the incorrect inlet and flow factor configuration for dextrose 70%. this issue was resolved by baxter technical support by having the customer update their formulary to correct the inlets and the associated flow factors. there was no patient injury or adverse events reported. no additional information was provided. | 1
5458687 | the subject device is not returned to olympus yet. olympus will investigate the subject device to determine the cause of this phenomenon when olympus receives it. there were no further details provided. if significant additional information is received, this report will be supplemented.><the monitor image disappeared suddenly 2 hours after the user facility started the laparoscopic distal gastrectomy. after the phenomenon occurred, the user facility turned off and on the subject device, and the monitor image was displayed properly. the user facility completed the procedure using the subject device. there was no patient injury in this event.><the subjected device was returned to olympus medical systems corp.(omsc) for evaluation. omsc tried to reproduce the disappearance of the endoscopic image, but this phenomenon was not reproduced. omsc evaluated the subject device, and found the following. there was no irregularity found in the subject device. the subject device has the error log that the communication error between the subject device and a camera-head occurred. there were not any scratches, distortions and foreign materials on the video connector socket of the subject device. there were not any foreign materials in the subject device. there was no abnormality found at the solder on the circuit board of the subject device.. there was no abnormality found at the communication state. there was no abnormality found at the internal harness connection of the subject device. olympus checked the device history record of the subject device, and there was no irregularity found. the exact cause of the reported event could not be conclusively determined, because this phenomenon was not reproduced. omsc surmised that the cause of this event was the following in theory. there were any foreign materials (e.g. dust, medicinal solution) on the electrical contacts inside the camera control unit's connectors. the communication between the subject device and a camera-head was blocked by these foreign materials. so the communication error between the subject device and a camera-head occurred. as a result, the endoscopic image disappeared. the subject device and/or any combination device were moved by the facility. at that time the stress was applied at the connection part of the subject device and a camera-head. so the connection state became erratic. as a result, the communication error of the subject device and a camera-head occurred and the endoscopic image disappeared. a member of the facility made physical contact with the subject device and/or any combination device when a member moved. at that time the stress was applied at the connection part of the subject device and a camera-head. so the connection state became erratic. as a result, the communication error of the subject device and a camera-head occurred and the endoscopic image disappeared. the otv-s400 instruction manuals state the corresponding method when there is an abnormality in the endoscopic image. there were no further details provided. if significant additional information is received, this report will be supplemented. this report is being submitted as a medical device report in an abundance of caution. | 1
5481863 | oca undertook a retrospective review of its mdr files for the period of january 2005 to april 2015. based upon this review, we are submitting this mdr to separately account for each of the three patients involved in this event. the hospital reported that the light source was used in three different procedures over a period of 7-8 days in which there were allegations of three colon perforations. one of the three patients was known to have cancer, which may have compromised the intestinal walls. the other two patients were perforated in the cecum area. hospital staff described the perforations as minimal". the hospital has been contacted for additional information regarding this report and according to the hospital staff; they do not suspect the device to have caused the patient's outcome as there was no problem found with the device during an evaluation conducted following the reported events. the same clinicians were said to have been involved in all three procedures and hospital staff have acknowledged that these reports may be technique related. please cross reference mfr. report numbers: 8010047-2007-00109 and 2951238-2016-00146 to account for the three patients as referenced in the original report. the following report will be supplemented to cross reference the two associated complaints: (b)(4).><the light source was not returned to olympus for investigation and the hospital has not provided further details regarding the reported events. the user facility's biomedical engineer reported no problems were found with the light source after testing it with their video processor and colonoscopes. the cause of the reported events cannot be determined. user technique and the patients pre-existing conditions cannot be ruled-out as contributing factors to the reported events." | 0
5773482 | based on customer complaints, it was discovered that the ph of the target retrieval solution measured 4-5 instead of the correct ph of 9.1. the customer complaint was related to no staining. | 0
5535111 | my health care provider uses epic software to send prescriptions to the pharmacy. the provider submitted 9 vials of novalog for 90 days, but what was received at the pharmacy was 6 vials of humalog for 60 days. i have had this occur twice in the past 18 months with a different health care provider and different pharmacy. something is wrong with the file once it leaves through epic software's ehr gateway and is received by the pharmacies. i would have received the wrong insulin and would not have had enough insulin for the 90 days that is required by my health insurance. i believe epic software is regulated as a medical device because it provides drop downs for medications and their uses. if not it should be regulated the same as other medical device software. i did not pick up the prescription. i had the doctor send it again to third pharmacy. i have not attempted to fill yet. | 1
5350910 | ><as we are validating the build of erxs within epic, we found an interesting bit of information. for ambulatory scripts for insulin u-500, it seems that epic's foundation system" (the system that has recommendations/pre-built files on how to build drugs) has a build that is unique for u-500 which is not aligned with medication safety best practices. epi's foundation system recommends to convert u-500 doses so that one may use a u-100 syringe. we are addressing this issue with our ehr team endocrinology service and diabetes nurse educations but felt that it is necessary to share since many institutions are/have transitioning/transitioned to epic. one plan is to link an order for (b)(6) syringes so both of these get ordered together. there are other challenges we expect from the diabetes nurse educators but we will address them once we meet this week. of note this is erx live in the 2015 version so it would be ver obvious to epic who identified it since very few institutions are going live with this version. coincidentally there are tow hospital systems in xxxxxxxx xx that are going live with epic 2015. patient counseling provided: unknown. (b)(4)." | 1
5353662 | description: during an investigation for a tpn compounding error, a potential contributing cause was that the 2 labels generated from the baxter exactamix /abacus software do not list ingredients in the same order and the naming of the ingredients are not consistent. the label that gets affixed to the tpn bag is designed to match the prescriber's tpn order (e.g., 1st ingredient is dextrose, 2nd is amino acids; salts are listed by their full name - sodium chloride"). the second label lists the "solution formula". the ingredients on the 2nd label do not match the order of the tpn bag label. also salts are listed by the chemical name (e.g. "na chloride"). this makes cross referencing the 2 labels during the checking process difficult. we contacted baxter to see if the order of the ingredients on the "solution formula" label could be changed to match the tpn label. they said this is not something that can easily be changed and it could be a major change that would cost us a lot of money. they also said that the name of the ingredients ("na chloride") cannot be changed. baxter should not make it costly / difficult to change the order of ingredients to match the tpn label. no one orders tpn ingredients in the manner they currently have been listed (na chloride 1st dextrose and amino acids are in the middle of the list). they should also use normal naming conventions for ordering salts (sodium chloride potassium chloride - not na chloride and k chloride). (b)(4)." | 1
5413568 | ge corometrics 250cx series maternal/fetal monitor produced inconsistent coincidence monitoring during high risk labor. the equipment failure may have been a contributory factor in the subsequent fetal demise. used during labor until emergent c-section. | 1
5446457 | pt in room and the fetal monitor is not feeding data to the central monitor. cables checked and cuts in and out in room but doesn't feed to central desk.><. | 1
5375044 | per sales rep fms duo has no flow. surgery type: hip; patient harm: none; patient delay - 60 minutes. pump started working after the delay and the case was completed with the pump. our sales rep confirmed via phone on (b)(6) 2016 that the patient was under anesthesia for approximately 45 minutes during the delay in procedure. he stated there were no reported adverse patient consequences.><if information is obtained that was not available for the initial medwatch, a follow-up medwatch will be filed as appropriate. the complaint device was returned and an evaluation was performed. the evaluation revealed the irrigation roller pump head was loose and pushed back against the front panel causing it to not rotate; therefore the unit had no flow. a review into the depuy synthes mitek complaints system revealed one other dissimilar complaint for this eight year old reusable pump. the device was repaired and has passed all functional and safety tests, and is fully operational. at this point in time, no corrective action is required and no further action is warranted. however, depuy synthes mitek will continue to track any related complaints within this device family as a means of monitoring the extent with which this complaint is observed in the field. (b)(4). | 0
6132407 | a (b)(6) male underwent an elective cystoscopy with transurethral resection of left posterior lateral bladder tumor and resection of bladder neck papillary tumor on (b)(6) 2016. the procedure was done for recurrent bladder cancer. the fluid warmer (thermedx fluidsmart system) was being used for the first time in the hospital's operating room and was operated by the vendor during the procedure. air bubbles were noted on the cystoscopy screen during the procedure and the pressure was increased from 60 mmhg to 70 mmhg to help eliminate the bubbles. a total of 12,210 ml of ns was pumped in during the procedure. the pt became bradycardic and hypotensive during the procedure and was treated with ephedrine with adequate response. upon arrival in the recovery room, the pt awoke in extreme pain and was noted to have a rigid, distended abdomen. additionally, he had a declining respiratory status which required reintubation. a stat abdominal ct revealed the presence of free fluid and free air in the abdomen and pelvis secondary to perforation in the left anterior aspect of the bladder. | 0
5956385 | it was reported that an acquisition work station radiation shield spontaneously shattered on a mobile trailer after it had been moved. we were informed there were no injuries. | 0
5813512 | on (b)(6) 2016, a patient's breast screening was conducted using philips microdose si (model: l50) system. the acquired image was reviewed by radiologist on a siemens pacs reviewing workstation. radiologist reviewed the image and no clinical finding is reported. the patient felt something in the breast and came back for consultation and then the patient underwent diagnostic scanning on another microdose si (model: l50 system) on (b)(6) 2016 and radiologist reported clinical finding in the breast image. | 1
5699563 | the technologist was positioning the patient when they heard a loud pop sound and the back of the c-arm started sparking. no one was injured.><the sites bio-med called in an electrical issue. our field engineer initially couldn't reproduce the sparking but after moving the c-arm all the way down and bringing it back up it sparked. the wire on the vta board had some grease on it and once that was cleaned he noticed a nick in the insulation. this wire runs behind the vta board and sits very close to the lead screw. this was causing the sparking when the c-arm was moved. he replaced the vta board. | 0
5758435 | the compression motor brake was slipping. the dealer field engineer replaced this part.><after an exposure the compression paddle was released automatically, the patient stepped back out of compression and the compression paddle came back down on its own. the compression paddle won't drive up or down unless they are doing it manually. no one was hurt. | 0
5706260 | ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. patient age and weight could not be obtained to date. report source other: user report mw5061306. device evaluation anticipated, but not yet begun.><ge healthcare (gehc) received a user medwatch report on may 10 2016 where a patient reported that in (b)(6) 2013, she had a mammogram done because she felt a lump in her left breast during monthly self-check - diagnosis was no cancer found. the site lost its acr accreditation due to suboptimal images and in 2015, patient received a letter from the facility stating fda had concerns with results of mammograms done during certain time period due to suboptimal imaging and her mammogram fell into that category. there was no allegation of ge healthcare mammography unit malfunction. in 2015 she had repeat mammogram done at different facility and results stated to see a doctor and get ultrasound. in 2016, she got ultrasound and upon doctor reviewing ultrasound she received a biopsy the next week. two week post biopsy was told that she has ductual carcinoma in situ in left breast and date 2 invasive ductual carcinoma in right breast. she met with surgeon who recommended a double mastectomy with chemotherapy for 8 months. based on this reported information, it has been determined on (b)(6) 2016 that the mammography system may have caused or contributed to the reported event which was initially assessed as a serious injury.><ge healthcare received an fda medwatch report on may 10, 2016, where a patient reported that in (b)(6) 2013, she had a mammogram because she felt a lump in her left breast during monthly self-check. diagnosis was no cancer found. in 2015, she received a letter from the (b)(6) center stating fda had concerns with results of mammograms done during certain time period and her mammogram fell into that category. in 2015 she had a repeat mammogram done at a different facility and results stated to see a doctor and get an ultrasound. in 2016, she got an ultrasound and upon doctor reviewing the ultrasound, she received a biopsy the next week. two weeks post biopsy, she was told that she has ductal carcinoma in situ in left breast and date 2 invasive ductal carcinoma in right breast. she met with surgeon who recommended a double mastectomy with chemotherapy for 8 months. | 0
5881811 | .><scanned in a patient to mammography unit and another patient populated the screen. | 1
5382443 | ><the anesthesiologist was asked to anesthetize patient. during intubation the patient became hypoxic. ventilation was difficult and there was on going bradycardia and hypertension. it was noted that the co2 waveform was flat and not correlating with ventilation breathes or compressions. severe ventilation effort was noted and correct ett placement was verified with fluoroscopy. the required ventilation effort caused the ventilation bag to rupture. during replacement of the ventilation bag other circuit connections were checked and a rubber test stopper was found occluding the output port of the anesthesia machine. the stopper was removed and the patient was reintubated. ventilation normalized and skin color improved. | 0
5491665 | ><anesthesia machine was malfunctioning. after operating room case is finished, the start/end case button would freeze and anesthesia tech was unable to reset machine for next case. the only way to reset machine was to turn machine off with knob. | 0
5608892 | an (b)(6) woman in a wheel chair was rolled into the room for an exam. the technologist placed the patient in the wheel chair underneath the gc85a table for an upper extremity exam. the x-ray technologist (user) pressed the park button by mistake causing the table to move down onto the patient. the force of the table caused bruising and a puncture wound to the patient's legs requiring medical attention.><this was caused due to careless use by the operator and the device was not the contributory cause of the incident. retraining has been provided to the user. the park" button since this occurrence has been reprogrammed to prevent any table movement to the patient support. customers narrative indicates that they appear satisfied with this solution. future design enhancements to the benefit of patient safety are being reviewed and discussed." | 0
5832729 | ><i had back surgery and the doctor saw it on x-ray. the filter has an arm that's came off and is just floating around. when to 2 doctors and they told me it cannot be removed and just don't get a blood clot. can you help me? | 0
5619454 | lis malfunction - two techs were working on the same result at the same time. sid urine creat was first resulted by the analyzer as >346.0. one tech put the sample back on the analyzer for a rerun. the sample was now resulted and sent to the lis as 417.0. one tech was looking at an unrefresh screen and saw >346.0. that teach cleared it out (incorrectly causing it to read 0.00.) the other tech approved what she was seeing on her screen, the 417.0 result. the pt's report reads <1.2. (there is a rule in orchard to report a 0.0 result as <1.2). how does the report show <1.2 with no initials when hs resulted it as 417.0? | 1
6190452 | the device was not returned and the lot number is unknown; therefore, a device analysis could not be completed. should additional relevant information become available, a supplemental report will be submitted.><should additional relevant information become available, a supplemental report will be submitted.><it was reported that a homechoice cassette was leaking from where the plastic connector and the tubing meet. this occurred priming of peritoneal dialysis therapy. the patient was able to continue their therapy using all new supplies. there was no patient injury or medical intervention associated with this event. no additional information is available. | 0
6194769 | (b)(4). since the lot number was not provided, this information cannot be determined.><according to the reporter, the support was broken on the scale of the device. this was noticed while setting up for the procedure. | 0
5832763 | the telecare management program is a service provided to qualifying individuals with a diagnosis of chf, cad, diabetes, and htn. this service is provided at no cost to the individual and is covered by their health plan in order to decrease health care costs and undesirable healthcare utilizations for a high risk population. this program provides in-home biometric monitoring equipment to qualifying participants. the equipment uses bluetooth technology to transmit biometric readings to a web portal that is monitored daily. biometric readings that are out of range will trigger an alert in the web portal. an alert will be triggered for blood pressure readings that meet the following criteria: systolic blood pressure greater than 18 0" or less than 89 diastolic blood pressure greater than 100 or less than 54 and heart rate greater than 100 or less than 49. the biometric equipment is provided by an equipment vendor that is responsible for but not limited to installation of the equipment instruction on using the equipment and troubleshooting/servicing the equipment. the blood pressure monitor provided by the equipment vendor is manufactured by ideal life ( model no: bpm 0001). the registered nurse that monitor the biometric readings in the web portal have reported numerous suspicious blood pressure readings with concerns regarding the accuracy of the blood pressure monitor. program participants and their provider(s) (i.e. primary care physicians cardiologists etc) also voiced concerns to the registered nurses involved with this program regarding the accuracy of the blood pressure monitor. research was conducted based on the information reported to me and i gathered data regarding the claims. on 04/07/2016 i submitted the data that i found to the equipment vendor to conduct further research. approximately 353 out of 441 participants in this program are receiving blood pressure monitoring using an ideal life blood pressure monitor. the telecare management pressure monitoring using an ideal life blood pressure monitor. the telecare management program requested that the equipment vendor replace all of the ideal life blood pressure monitors with a monitor made by a different manufacturer until further research was conducted. approximately 43 participants were identified with suspicious blood pressure readings and this was based on the research that i provided to the vendor as well as a query conducted by the vendor. the query conducted by the equipment vendor looked for participants with the following criteria: diastolic blood pressure readings over 105; participant reporting as asymptomatic; participant not showing any signs of heart rate and/or systolic blood pressure elevation. on (b)(6) 2016 a registered nurse with a caseload of approximately 50 participants reported that she noticed an increase in blood pressure alerts from participants that rarely had blood pressure alerts in the past. many participants in the telecare management program have taken their ideal life blood pressure monitor to their physician's office to compare with the blood pressure monitor at the office. the diastolic blood pressure readings at the physician offices were 30-40 points lower on average when compared with the ideal life blood pressure monitor. one participant reported the following data after a follow-up appointment with his doctor: arm ideal life bp monitor doctor's office bp monitor right 151/105 136/78 left 139/101 146/88. the registered nurse that works with this program participant spoke with a nurse at the doctor's office. the nurse states "we believe the equipment is providing falsely elevated readings and request that the blood pressure monitor be serviced and/or replaced." the participants blood pressure monitor was replaced approximately 3-4 times with ideal life blood pressure monitors but the issue did not resolve. this participant was installed with a blood pressure monitor from a different manufacturer at the end of (b)(6) 2016. since receiving the new blood pressure monitor." | 0
5660584 | it is indicated that the device will not be returned for evaluation. a review of the batch history, historical trending, and similar complaint trending review for the product family will be conducted. if there is any further relevant information from that review, a supplemental medwatch will be filed. (b)(4).><same case as: 2134265-2016-03990 and 2134265-2016-04245. it was reported that automatic pullback failure occurred. a 100v ilab ultrasound imaging system was used in conjunction with an opticross imaging catheter and a pullback sled. during procedure, it was noticed that sometimes automatic pull back did not work and the pullback gear of the motor drive unit (mdu) 5 plus was damaged. furthermore, the lock part of reusable sled sometimes floated. no patient complications were reported. | 0
5562730 | ><during the procedure the x-ray equipment began leaking large amounts of water on the floor, creating a puddle under operating electrical equipment. source of the leak found to be the chiller"." | 0
5566856 | ><the giraffe looked like it had crumbs down in the moulding and we were suspicious of some black crumbs" that looked like they were in a clump. the giraffe was pulled from service and sent to biomed. we noticed the new giraffes have a new moulding on their canopy that would not allow these "crumbs" to collect. biomed opened the giraffe moulding and called infection prevention to take a look. there was a dark brown slime between the outside moulding and glass. we had our microbiology department culture for mold and bacteria and nothing grew. we double-checked the recall system and didn't see anything regarding this style of canopy moulding but still wondered why there was a design change. there were a total of 6 giraffes that were in danger of having this same type of moulding so all units were pulled and checked by biomed. of these 6 giraffes 4 units were serviced for this issue." | 0
5864830 | ><latch broken on device causing isolette door not to close properly. | 0
5959484 | (b)(4).><according to the reporter, a patient had a pillcam procedure that has possibly retained in the stomach. an xray was performed about two weeks later and the xray technician said it either looks like a piece of the capsule or a surgical clip was inside still inside. follow up information from the account confirmed that the patient had a clip placed in the colon at the exact spot that showed up in the xray. the radiologist confirmed the clip as well. the patient did not have any symptoms but the patient did not see the capsule expelled and wanted to see if it was retained. no adverse events to the patient was reported. | 0
6080980 | i am a type 1 diabetic. i wear an omni pod to inject my insulin. omni pod changed their packaging for the pods that i use. since then, i can only wear the pod for 24-36 hours at a time. the pods are to be used for 72 hours. i itch so bad that i have to change them really often. my skin actually has scar marks from the new pods. i still have a box of the old packaged pods and can still wear them for 72 hours at a time. when i remove the new pods i have hives underneath them. the red sores that i'm left with last a week, then for months after you can still see where it was. they, at omni-corp, says that i am not alone but they don't have an answer. these pods cost my ins (b)(6) every 90 days, and they are made in (b)(4). i am worried because these pods have a cannula that is inserted into my skin. don't know if i may be in contact with a product that may cause other problems down the line. | 0
5597759 | patient was in pacu and philips monitor showed an ecg flatline 3 times. connections were checked and leadwires were connected correctly. two ecg electrode pads were replaced and normal sinus rhythm returned. | 1
5762254 | a follow up report will be submitted once the investigation is complete.><the customer reported that the speaker malfunctions intermittently. the device was received by a field service engineer, who verified that there was no sound from the speaker. the device was in clinical use at time the issue was discovered. there was no adverse event or patient harm reported.><. | 1
6151807 | a follow up report will be submitted once the investigation is complete.><the customer reported that the wave is not showing on the bedside monitor screen". a patient death was reported. the device was used for monitoring at the time of the alleged malfunction." | 1
5483291 | a follow up report will be submitted once the investigation is complete.><.><the customer reported that the monitor fell on the patient's shoulder and they got bruised and a slight cut.><it was reported that the patient had just been discharged and the customer's anesthetic technician was cleaning the equipment. as he was cleaning it, the monitor and arm fell and caught him on the wrist; he had a small cut. there were two staff present and they were able to catch it before it fell on the floor. | 0
6089624 | pt purchased a halloween-type contact lens online and suffered an injury because of it. the pt only wore the contact lens for 5-10 minutes before removing it from his left eye (didn't wear the right one as the left one was hurting so bad). when i saw him, nearly his entire left eye corneal epithelium was missing. it was almost like he had undergone prk refractive surgery (where the epithelium is intentionally removed). the pt was in pain and had significantly reduced vision from that eye. a bandage contact lens was applied to aid in the healing, and a 4th generation fluoroquinolone (moxifloxacin) was prescribed as well as a topical nsaid (ketorolac). the pt developed an elevated intraocular pressure and was treated with anti-glaucoma medication. i think it is likely that the contact lens packaging solution caused a toxic reaction with the pt's eye. therapy start date: (b)(6) 2016. the product is over-the-counter. | 0
5913357 | ><about two months ago, upon completion of a successful cardiac catheterization, images from the catheterization failed to send to the synapse archive. ge field engineer on site and attempted to delete images to increase the storage space. it was unsuccessful. the system was rebooted and the images from the just completed case could not be reclaimed. the hard drive was replaced about six days after the event but was not functional until two days after replacement. manufacturer response for ge innova igs520, (brand not provided) (per site reporter): ge service program manager on site, along with field engineer, investigated the event. the removed hard drive was sent to france for engineers to try and retrieve the lost images from the case in question. in the process of removing the hard drive the chassis shorted out and needed to be replaced. the chassis was returned to gehc harvested parts depot as damaged and per the service manager, has likely been destroyed. per the service manager, ge engineering confirmed that the disc full message is due to a known software logic error which is intermittently displayed on some systems running the igs innova m3b software, when dosemapp and fluoro store options are also enabled. recommended that if the message appears, a member of the staff check to ensure there is adequate disc storage space. | 1
5675311 | needless connector on a central line was found to be broken into two pieces. there was a sharp green portion that had detached from the clear cap. the maintenance iv fluids and vasopressor was infusing into the pillow. the patient's port to the central venous line was open to the air. | 0
6063716 | invision-plus iv connector made by rymed technologies broke with the threads cracked off and stuck in the tubing connector. the patient noticed blood on his shirt and realized there was blood coming out of his central line and medication dripping from the line that was not connected into his central line anymore. he took pictures and sent to the sp. he also saved the broken connector and tubing. he is currently out of town so he doesn't have the lot number, but said he will call on monday with it. we have sent him a different brand connector to try because he is afraid to use this brand. he stated he had minor side effects. i did not ask him if it is ok if the manufacturer contacts him. i can ask when he calls with the lot number. he willing to mail the broken connector and tubing. dates of use: from (b)(6) 2016 to currently. diagnosis or reason for use: pah. | 0
6061038 | pt was sitting in church this weekend when her invision plus iv connector broke. blood started draining into her hickman. pt panicked and called 911. she was taken to (b)(6) hosp because she was told to go there if there ever was any issue with veletri or her line. she states that the emergency room was not familiar with veletri. pt kept telling emergency room staff they need to restart infusion asap because of the short half life of veletri. they told her to calm down and she looked ok, so nothing to worry about. emergency room had no tubing available, so they had to call the cardiac floor to get tubing brought down. pt said it took over an hour and a half before everything was taken care of and restarted. pt is now doing ok. i reviewed with pt that if this ever happens again, she should clamp her hickman to stop the blood flow while she sets everything back up to resume infusion. i advised pt that she can always call the pharmacy to assist with any infusion issues and we have an on-call pharmacist available 24 hours a day. hosp can also call pharmacy to review everything if they are not familiar with veletri infusion. dates of use: (b)(6) 2014 to present. diagnosis or reason for use: pah. | 0
5784891 | rymed 460207 invision-plus catheter extension set and 415303 end cap became separated when patient got up to go to the bathroom and did not pull iv pole in with him. doctor called for nurse when he went into the pt room as the patient had blood pouring out of antecubital vein. copious amt of blood on patient, floor and pt bed. nurse ran in grabbed a towel and place pressure on iv catheter causing an exposure of blood to patient. this is a frequent occurrence at (b)(6) and has been reported to manufacturer distributor and hospital with minimal changes. product still remains. many blood exposures still occurring. previously it was a daily occurrence. dates of use: (b)(6) 2014 - (b)(6) 2016. diagnosis or reason for use: used on all iv infusions. event abated after use stopped or dose reduced: no. event reappeared after reintroduction: yes. | 0
6112097 | the customer reported that there was an adverse patient event and the nurse staff alleged that the central station monitor did not alarm. the date of the adverse event was (b)(6) 2016 between 9:55 and 10:03pm for bed 918b. the customer was also concerned the monitor was reading artifact as a rhythm and the staff thought that was why they did not notice the patient condition. investigation is required going forward to review the piic alarm logs and the patient strips to ascertain whether the philips device may have caused or contributed to the adverse event. the patient was being monitored at the central station information center. the patient was using a dual chamber pacemaker as seen from the patient strips. the patient expired.><a follow up report will be submitted once the investigation is complete.><there was no malfunction. the biomedical engineer wanted all information from the logs available from 8pm to 10pm for bed 918b (tel30) on (b)(6) 2016. per patient safety officer, the patient was using a dual chamber pacemaker and had been experiencing abdominal pain. the issue was that there was a delay in recognizing the patient?s condition and in providing treatment. per patient safety officer the patient coded at 22:00pm and expired at 22:22pm. the piic alarm logs and system logs were sent to philips r&d for analysis. the logs show that the piic device did alarm as a red v-tach alarm for the time in question (21:59:49.843 (b)(6) 2016) and supports that the device was monitoring and annunciating alarms per specifications. there were no patient alarms prior to 21:59 on that day ((b)(6) 2016). the logs taken from the device had wrapped, meaning that due to the passage of time since the event, not all relevant log data was still present, therefore r&d could not give more information on the device. however, r&d stated that the patient admit state didn't change, they stayed in the same bed assignment, and data was being collected for the bed for hours before the event. based on the information provided there is no data to support a malfunction. | 1
5553735 | ><i had a mammogram done at (b)(6) center on date above at (b)(6). i had mammogram because i felt a lump in my left breast during monthly self check. diagnosis was no cancer found. (b)(6) 2015 received letter from (b)(6) stating fda had concerns with results of mammograms done during certain time period and mine fell into that category. in (b)(6) 2015 i had repeat mammogram done at a different facility. results stated to see a doctor and get ultrasound. (b)(6) 2016 got ultrasound and upon doctor reviewing ultrasound received a biopsy the following week. two week post biopsy was told that i have ductal carcinoma in situ in left breast and date 2 invasive ductal carcinoma in right breast. (b)(6) met with surgeon who recommended a double mastectomy with chemotherapy for 8 months. | 0
5890853 | ><the preliminary analysis and root cause determination indicate that the device operated as expected. the investigation is ongoing and mckesson will follow-up with the fda with any new information learned. | 0
5353652 | ><description: there are two different info of cyclosporine marketed in the us. cyclosporine and cyclosporine modified. they are not generically interchangeable. we had a mix up in our pharmacy that was caught as we were putting the product on our shelf that we order from (b)(6). we believe the less than optimum description in mckesson connect contributed to this. as you can see from the copy of the mfr's box this product ((b)(6)) in attachment the product name is cyclosporine capsules usp modified". this is true also for all the other products listed that are actually "modified" formulation. (b)(4)." | 0
5479532 | blood cultures were ordered as part of a fever work up. the results were randomly and intermittently electronically transmitted to the ehr data storage library, but there was not any warning or notification that a positive blood culture was reported. more than 24 hours elapsed before any clinician knew and antibiotics started. the delay, commonly caused by silent delivery of life critical results, was life threatening in this case. | 1
5495524 | there was a deep vein thrombosis afflicting this pt, for which he was being treated with coumadin. an inr was deposited into the electronic care record device without notice or warning. it was elevated and required that coumadin not be administered, but, no one saw the result that was silent in the coag silo. consequently, the usual dose of coumadin was administered that resulted in an tnr that was markedly elevated, that required immediate reversal. there was a 5 gram decreased in hemoglobin that threatened the life and well being of this pt. it is a widespread problem in ehr driven care that disease critical results come back silently, and sometimes such is deadly. oversight and correction of this defect is indicated. | 1
5719804 | the pt was seen in the post op anesthesia care area (pacu). post op electronic orders were not released while pt was in pacu, and then pt was transferred to med surg. after transfer the pacu orders were released. the usual method was bypassed, and the post op orders did not auto d/c upon transfer. this left the medication and dosages active on the mar for the post op pt, which presented a potential risk for giving a medication not ordered, and in dosages not appropriate for the pt's location and level of care. this has occurred on a number of occasions and staff are being educated to the problem and potential risks. medication not administered to or used by the pt. relevant materials not provided. (b)(4). | 1
5719844 | (b)(4). pt was admitted to our hospital on xx/xx/2016 for induction therapy for newly diagnosed secondary acute leukemia. she was still in house on xx/xx/2016. our computer system is set so that all medications will auto-discontinue after 30 days. among other issues, she was poorly controlled hypertension and was taking multiple antihypertensives. one of these is clonidine 0.3 mg tid. on xx/xx/2016, the clonidine was discontinued and not reordered. she also was taking amlodipine 10 mg daily, which appears to also have been discontinued at the same time. pt's blood pressure was elevated overnight starting approx 12 hours after clonidine stopped (up to 185 sbp). unclear if providers or rn's realized clonidine and amlodipine were stopped. clonidine restarted about 24 hours after stopped. amlodipine not restarted but nifedipine sustained released started also about 24 hours after the amlodipine was stopped. pt did develop bilateral scleral hemorrhages during this time frame - potentially from the hypertension. her platelet count is approximately 50 at this time. we are struggling with this automatic state date - this is a new functionality for us. we integrated into a larger health care system in xxx 2015 - using the same ehr platform. we did not have a automatic stop date prior to the integration so are not really looking for this info - pharmacists don't use the functionality where the list of to be expiring" medications show up. we would like to not have the 30 days stop date but get pushback from corporate compliance due to their interpretation of the cms 482.25 (b)(5). from what we can tell both (b)(4) are not in favor of automatic stop dates. we would love some advice from you on this matter or you bring it to a larger forum of discussion. medication not administered to or used by the pt." | 1
6157633 | the cpoe system, epic, is one of the largest in the united states. epic has an inherent problem: orders for medication show up in the medication administration charting (mar) section in an erroneous, and, i believe, confusing and possible dangerous manner. the mar, and the label that issues from the printer in epic, confuses product strength with dose. for example, a 3 mg daily dose of sirolimus appears as sirolimus tablet 3 mg; dose 3 mg oral daily. sirolimus is not marketed in a 3 mg tablet. similarly any dose of a product that does not exactly match the tablet strength appears in the same confusing pattern: e.g. a dose of mesalamine capsules 1600 mg appears as "mesalamine capsule 1600 mg dose 1600 mg etc." mesalamine is not marketed as a 1600 mg capsule. the above should read correctly in the mar (and on labels) as: sirolimus tablet (or tablets) 1 mg dose 3 mg and mesalamine capsule (or capsules) 400 mg dose 1600 mg etc. this inherent software error has widespread distribution i believe and should be rectified. my attempts to point out this problem have been ignored locally. severity: circumstances or events have capacity to cause error. (b)(6)." | 1
5719849 | the patient was seen in the post op anesthesia care area (pacu) post op electronic orders were not released while the patient was in pacu, and then patient was transferred to med surg. after transfer the pacu orders were released. the usual method was bypassed, and the postop orders did not auto d/c upon transfer. this left the medication and dosages active on the mar for the postop patient, which presented a potential risk for giving a medication not ordered, and in dosages not appropriate for the patient's location and level of care. this has occurred on a number of occasions and staff are being educated to the problem and potential risks. medication not administered to or used by the patient. relevant materials provided: (b)(6). | 1
6040502 | .><omnicell would not dispense norco in ed. this item is available in multiple omnicells. to address immediate patient issue, product could be obtained from other omnicell. the next day, am pharmacy checked and found no problem with hydrocodone. no hydrocodone was dispensed that day. that day, i verified the following: hydrocodone was cycle counted - count was correct. coil containing hydrocodone was inspected and found to be intact. hydrocodone was destocked and restocked and coil functioned as expected. | 1
5684655 | .><the pentax monitor not working during the egd procedure. the procedure was finished, but unable to take pictures. there was no harm to the patient. pentax tech was notified via phone and said there might be something wrong with the capture card or the s video cable and biomed should handle it. manufacturer response for pentax endo pro, endopro (per site reporter): they understand the problem and know that this is an issue. the only solution is to reseat the video card. | 1
5478588 | the customer reported that incorrect susceptibility results were reported out of the lab for an escherichia coli isolate with multi-drug resistance. the customer stated they reuse specimen numbers and discovered that the (b)(4) report for a recent specimen tested in (b)(6) was reported incorrectly. the isolate was an esbl positive e. coli from a urine source with multiple drug resistance; the drugs were reported out as susceptible. customer reported it appeared that the ((b)(6) hospital system) had results from the same specimen number tested in 2015 - same panel type, same id, but esbl negative. the customer stated the patient was treated with ciprofloxacin however when the corrected report was released, the patient was put on a different drug, details unknown. customer stated the results were finalized on (b)(6) 2016 and the corrected report was issued on (b)(6) 2016. patient treatment was impacted in connection with this event.><review of the submitted data revealed specimen numbers stored in (b)(4), m145f and m145g, was likely downloaded from the (b)(4). the f" tag was applied to 2015 and "g" to 2016. both specimens had collect dates that corresponded to download and run dates. device configuration indicated the active specimen date range was set to 180 days. microbiology reports for each specimen indicate the correct ast results were stored in (b)(4) for each specimen. the customer further reported that they manually transmit all specimens and use the "active specimens only" and "test group status date = 0" settings. customer stated they did not visually confirm the results before transmitting. customer stated that both specimens were downloaded from the (b)(4) so each should have had a tag applied for the year making them unique. the failure mode for the incorrect susceptibility reporting was likely user error. since the customer reuses specimen numbers it is likely that the previous year specimen was inadvertently selected by the user during manual transmission of the results. there is insufficient evidence to suggest an instrument software or panel malfunction. data provided confirmed the correct results were stored in (b)(4) for each specimen and the software functioned as intended; results reported by (b)(4) were based on customer input. (b)(4)." | 1
5903950 | (b)(4). no malfunction of the device was reported to maquet. death of a patient was reported to maquet. the death of the patient is not related to the use of the maquet product however since a maquet product was involved we decided to report this case to the competent authorities. a maquet-service-technician was on site and confirmed that the death of the patient is not related to the use of the maquet product. maquet (b)(4) provides product failure investigation, analysis and resolution for the device described in this report.><patient death was reported to maquet. the death of the patient is not related to the use of the maquet product however since a maquet product was involved we decided to report this case to the competent authorities. manufacturer reference # : (b)(4). | 0
5413703 | on (b)(6) 2016, the customer site in (b)(6) contacted agfa to move/delete bookmarks in report after information was not being transferred to the his system by oru_outbound interface. a complaint was registered and was determined, that this information did not transfer because it was written after the report end bookmark, and then filtered out by the interface. there have been no reports of injury when using the reporting system at this customer site. this issue is a known issue to agfa. a problem record was created on (b)(6) 2012 to address the issue via (b)(4). in 2014, an important information letter was distributed to global customers with both impax and impax ris to inform them of the issue. product corrections were issued for both the impax (pacs) product and the impax ris product as addressing this issue required changes to both products. impax 6.5.5 su1 was released on march 20, 2014 and introduced the new implementation of protected sections in impax reporting. this fix was ported forward to all newer versions of impax. impax ris 6.0.0 was released on february 07, 2014 and introduced protected sections of the report template. this site is currently running versions below these corrections. impax pacs 6.5.3 su3 impax ris 5.8.1 hf1 agfa's long term solution has been recommended based on the important information letter distributed in 2014 and we continue to work with the site to find an acceptable workaround until the permanent solution can be applied to their system. the customer needs to plan improvement of the template outlay with existing bookmarks in current versions (site is a multi-site solution where upgrades will take time to plan and implement).><. | 1
6165304 | on (b)(6) 2016 it was reported by users at (b)(6) that earlysense device provided a bed exit alert for a patient that left bed. the earlysense central display station (cds), also provided the bed exit alert at the same time at the nursing station. alert was confirmed at the bedside and at the cds by healthcare providers. no injuries were reported, however, the customer reported that the external wireless communication devices ((b)(6)) did not receive the alerts on all three tech phones. in conclusion: the earlysense bedside unit and central display station were performing as expected according to their specifications. however, the gateway software was not supporting the delivery of the earlysense alerts to the external mobile alerting devices that belong to the hospital ((b)(6)). correction provided to the site. updated gateway software was installed in collaboration with the hospital it that recognizes when the gateway application is down' (watchdog) and restarts the gateway so the external alerting devices are supported." | 1
5429299 | fujifilm medical systems u.s.a., inc. (fmsu) recently completed a retrospective review of complaint files from 2013 through 2015 and has determined that this complaint may meet the threshold for reporting because images were reported as lost. fmsu is therefore reporting it at this time. the site (b)(6) reports that they have lost two images during operation of a clustered (b)(4) console running v6.2. these images were taken between 1:00 and 1:10 pm on (b)(6) 2013. these images had become inaccessible and the (b)(4) was then rebooted. rebooting deletes the images from the database. once the reboot process was complete the images were no longer present on the (b)(4) console. the facility did not provide any information indicating this as an adverse event, i.e., that the patient was re-imaged. fujifilm therefore is considering this as a possible malfunction.><fmsu submitted this mdr (along with three other mdrs) on october 6, 2015. unfortunately, the submission failed and we did not realize it until last week. we are therefore submitting it today. the following is the acknowledgment message received on october 6, 2015: (b)(4) has received your submission." | 1
5829949 | rn accidently used the connector to the lp evd drain instead of iv tubing connector to infuse dilaudid for pain. pt remained stable and was observed in the icu overnight, no adverse outcome. | 0
5737773 | a mobile cardiac telemetry device reported 15 hours of additional data after it was disconnected, turned off and the battery removed. i started monitoring at 2 pm, (b)(6) 2016 and finished about 7:30 pm the next day, (b)(6) 2016 for a total of just over 29 hours. the official 40 page study report states that i had a total of 16552 pvcs which may be a serious medical problem. the report further states that the recording time was 48 hours and that the total time analyzed was 44 hours. in fact, specific patient data is reported for 15 hours after the device was completely disabled, electrodes removed, sensor turned off and batteries removed, and the associated smart phone turned off. this calls into question the entire study and whether another patient's data was totally or partially reported. of note, approximately 25 hours after the study terminated, saturday at 8:30 pm ((b)(6) 2016), i received a call from the monitoring company which requested i turn the smart phone back on, so that additional data could be retrieved, some problem had occurred during the 29 hours the device was connected. the company refuses to talk with me ((b)(6)) or my physician to explain this serious discrepancy. it appears that i may require another holter monitor to resolve this disturbing issue. | 1
5644645 | a scottcare customer, calling on the behalf of their client, contacted scottcare on (b)(6) 2016 stating that a li-ion battery for the telesentry mobile cardiac telemetry monitor had exploded" while charging at a patient's home. the customer clearly stated that no one was injured by the battery's event. the battery was returned by the patient to our customer on (b)(6) 2016 and ultimately returned to scottcare on (b)(6) 2016. scottcare was able to see the battery had a catastrophic event with evidence of charring on the battery materials and some residue on the plastic battery shell. scottcare's investigation of the shell did not uncover any evidence of explosion melting of the plastic or disfigurement one would expect with an "explosion" but rather something akin to smoldering as evidenced by the residue on the battery shell. scottcare was not able to determine if the battery had been stored properly by the customer and the patient or had been exposed to any adverse environmental conditions such as excessive moisture or heat. the device was sent to the battery manufacturer (batteryspace.com/aa portable power) for further investigation.><scottcare has received a single reported incident of this nature for this specific model battery since the battery was commercially released in 4/28/2010. approximately (b)(4)batteries have been produced and shipped to customers to date which provides an approximate (b)(4) failure rate. the failure occurred during the charging of the battery which can only be charged while detached from the telesentry cardiac monitor. furthermore the telesentry mct monitor is shipped with two batteries each charged and used on alternate days when provided to the patient. as such we do not foresee an immediate threat of injury or death for this battery. thus this will likely be a class ii recall requiring relabeling with "date of manufacture" and "use by date." we will also subsequently evaluate sourcing a new battery cell and redesigning the plastic compartment that houses the battery cell. scottcare plans to take the following actions by march 15 2016: include "date of manufacture" and "use by date" (set at two years from the date of manufacture) on all new batteries for future sales. send a letter to all customers who have received these batteries in the past with the following corrective actions: batteries older than two years: inform the customer of the potential battery issue and ensure they are aware of our 2 year use guidline in our manual. if the customer is no longer using the battery advise the customer to properly dispose the battery per scottcare's telesentry ts01 operator manual (pd-100975). batteries less than two years: inform the customer of the potential battery issue and ask them to return the battery for relabeling with proper "date of manufacture"" and "use by date.". if the customer is no longer using the battery advise the customer to properly dispose the battery per scottcare's telesentry ts01 operator manual (pd-100975). concurrently scottcare will evaluate and test an alternative cell which complies with the iec62133 standard and assess a redesign of the cell compartment as recommended by the battery manufacturer. scottcare plans to complete these actions by april 30 2016." | 1
5764587 | a lithium ion battery used for our portable computers sparked, overheated and melted. an entire floor of our hospital had to be evacuated in order to protect our patients from smoke inhalation. | 1
5522299 | ><during the surgical procedure, anesthesia was unable to monitor/measure/determine and amount of anesthetic agent being delivered. the nurse anesthetist believed that the ge aestive 5 gas delivery unit was faulty and immediately contacted the biomedical engineering staff to urgently evaluate the machine. despite the failure, there was no impact to the patient. biomedical inspected the equipment and determined that the agent failure issue was related to the ge datex carescape b850 monitor and the ge datex m-caio parameter module. the agent id failure message occurred on these two units. biomed swapped in a spare m-caio module and temporarily resolved this issue. the same message occurred on the monitor with this new module later in the same day. at that point, the biomed tech shut off the monitor cpu to reset the device. the monitor and module came back from rest and functioned as normal without recurrence of the failure message. biomedical engineering technicians reported they have seen a fair amount of calls indicating the agent id failure" message . most often corrected by shutting off the cpu and rebooting it if you will. it's like the cpu gets locked up and produces this particular error message. there were no reports of patient awareness. the patient is a minor. documentation indicated that the patient received general anesthesia and recovered satisfactorily from the anesthesia. the anesthesiologist documented there were no apparent anesthesia-related complications. b850 cpu runs continuously. no manufacturer recommendations to reboot power." | 1
6105545 | during cardiac ablation the recording system was showing a signal even when not connected to catheter. this happened on two other patients mrn (b)(6) on (b)(6) 2016 and mrn (b)(6) on (b)(6) 2016. clinical engineering and st. jude were involved in trouble shooting this machine. | 1
5541501 | mms model m3001a for the intellivue mp 70 monitor fell out of brackets. | 0
6105629 | during conduction of a bioequivalence study of fexofenadine/pseudoephedrine 60/120 mg (on (b)(6) 2016), an abon (biopharm) multi-drug 5 panel screen test (coc-thc-amp-mop-tra) for urine sample was used prior to conducting the study to screen for drug abuse in healthy male volunteers recruited and included in the study the result was negative for the 5 drugs (coc-thc-amp-mop-tra) before dosing the volunteers then after dosing the volunteers another random screen test was conducted for a volunteer at the 8 hour interval post dosing and the test showed a positive for (tra) and negative for the rest (coc-thc-amp-mop) although the volunteer is healthy and not abusing (tra) and showed a negative result for the test predosing so an investigation was carried out on another new three healthy male volunteers other than those included in the study where a predose abon (biopharm) multi-drug 5 panel screen test was carried out and showed negative results for the 5 drugs (coc-thc-amp-mop-tra) then a single dose of fexofenadine/pseudoephedrine 60/120 mg was administered by the three volunteers and a urine sample was collected 2 and 4 hours postdose and an abon (biopharm) multi-drug 5 panel screen test was carried out. the results showed positive for (tra) and negative for (coc-thc-amp-mop-tra) although they are not abusing (tra) and showed a negative result for the test predosing so according to this result a cross reaction of pseudoephedrine is suspected which may have shown the false positive results for (tra) post dosing. dates of use: (b)(6) 2016. diagnosis or reason for use: bioequivalence study. is the product over-the-counter: yes. event abated after use stopped or dose reduced: yes. event reappeared after reintroduction: yes. | 0
5969051 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><spacelabs received a report that on (b)(6) 2016 a telemetry central monitor powered off by itself. the customer removed the product from service. no injury was reported.><the device was sent to spacelabs healthcare for further analysis. the reported problem was verified. the power switch and nvram were replaced. the repaired unit passed all functional tests and was restored to service. this report is complete and this particular issue is considered closed. | 1
6002993 | i had a loop recording device implanted under my skin in my left chest area in order to detect any episodes of atrial fibrillation that may occur while i am asleep. i was given a medtronic mycarelink patient monitor model 24950 to keep by my bedside to receive and forward any a-fib episodes. the 2 instruction manuals that came with the device did not adequately describe how the device worked. the first manual indicated the various displays that the monitor would display during operation with the final picture in the manual showing a large green check mark which i would expect to see if the monitor was functioning properly. the 2nd more detailed manual also showed the various displays the monitor would display during operating with an additional display after the green check mark. this display depicted the status of the internal battery, a picture of a radio antenna with signal strength and a green check mark plus the current date. it also described 3 led lights on the monitor (green, blue and amber) indicating proper operation. in reality the way the monitor works is that after about 5 minutes the display is blank as well as the led lights are turned off. i called medtronic and they stated the blank display indicated everything is operating correctly and there is nothing to worry about. i have 2 problems with this operation. first, nowhere in either manual does it indicate the display will go blank. it just shows the final display inferring that is what the patient will see. second, a blank display indicating proper operation is counterintuitive. i am a retired electrical engineer and have worked with hundreds of shipboard and aircraft displays for the us navy and air force and can guarantee that neither the navy or air force would not accept any display that does not have a positive indication of proper operation. the blank display gives me zero confidence the system is doing anything. i may as well have a brick sitting on my end table. | 0
5760702 | the customer reported that the hospital network was being flooded by one of the nk servers. because this caused the internet to become unusable,the customer closed the hospital side port to stop the multicast from flooding the network. as a result, all of the rns (remote network systems ) in the facility were showing communication loss. also, all of the devices that were physically at the milford campus, but remotely monitored at the dover campus, were going in and out of comm loss. the reported event was evaluated by the nihon kohden technical support group. nihon kohden's netkonnect system was sending multicast out through the hospital side's network. the service was stopped and restarted, and as a result, the services and the multicast began going out the correct port. the system was back up and running. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the hospital network was being flooded by one of the nk servers. because this caused the internet to become unusable,the customer closed the hospital side port to stop the multicast from flooding the network. as a result, all of the rns (remote network systems ) in the facility were showing communication loss. also, all of the devices that were physically at the milford campus, but remotely monitored at the dover campus, were going in and out of comm loss. | 1
5688302 | i was required to take a pre-employment drug screen, a non-dot 9 panel screen via (b)(6). the results came back with a positive indication of marijuana use despite no use by me. four days before the test, i oral surgery and was prescribed oxycodone for pain relief which i took many in the interval between the surgery and the pre-employment test. no indication for opiates was returned despite my heavy consumption of oxycodone due to pain, the last dose the evening before the test. when i asked the tester when they called, why there was no indication for opiates due to my substantial oxycodone use, the answer was we don't test for that". amazing - a non-dot 9 panel test requires opiate testing! since i am not a marijuana user and this test which should have indicated opiates the test is totally inconsistent and should be retested as either chain of custody was violated the sample was contaminated or the device or process used to do drug testing is flawed. i previously worked for an (b)(6) company ((b)(6)) where i developed software for a blood-gas analyzer - a device and process very similar to urine drug testing. control solutions are used to calibrate the device for correct operation as well as periodically verify that the device is working properly. the only recourse the testing company ((b)(6)) provides is to retest the sample which is pointless if any of these conditions are true. the fda requires accuracy and repeatability of measurements so i have no doubt that a flawed process will yield the same flawed results. i had my oral surgery at (b)(6) on the (b)(6) 2016 and when i checked with my oral surgeon he said he surprised if there was not a positive indication for opiates as such my urine sample was a control solution for opiates which the device failed to detect - how can there be any confidence in any measurement taken? it should have been detected for 4-6 days after use. this area of drug screen testing is sorely in need of government regulations to protect the employment applicant. as it stands now we have no rights against false testimony like this. at least a retest a few days later would be a start. as i said we have no rights against false testimony like this. please start regulating these drug screening so this does not happen to anyone else. i lost a (b)(6) job because of this!" | 0
5496126 | ><manufacturer narrative: ge healthcare performed a checkout of the device and found the screw (m16") stud and screw holes of the three-leg jig tool had a thread galling issue making it difficult to rotate the handle. the field engineers applied a large force to the handle which caused two other screws ("m4") on the handle to break. the m16 screw stock has been checked and no issues were found. it was confirmed that other units shipped to the field do not have the issue. service procedures includes proper instructions for using the tool. no further actions are planned at this time." | 0
6171440 | .><on december 22, 2016 agfa has implemented a correction via agfa (b)(4) ((b)(4)) and an urgent safety notice was sent to affected consignees for this issue. on january 26, 2017, a mandatory service bulletin was released to provide documentation on how to deploy the correction for this issue. any additional reporting for this correction activity will be reported to fda via 21 cfr 806 reporting via agfa id - (b)(4) ((b)(4)). there has been no reported patient harm for this event.><we are submitting this medical device report today december 13, 2016. on november 22, 2016, the customer reported that while using an nx workstation with software version nx 3.0.8950 software and selecting the affected patient/exam from closed exams, initially the wrong image was linked to the exam and appeared. after a short time the wrong image was replaced by the correct image, however the wrong image was used for transmitting to pacs. to correct this issue, the user backed up all logs and rebooted the system. the correct image was then transferred to pacs. during the investigation by agfa november 24, 2016 - december 1, 2016, the root cause was identified as a memory leak which was introduced in the image processing software in nx 3.0.8950. the user received two warning messages of potential malfunctioning prior to the issue. the user did not follow the warning instructions and this resulted in the image of a patient being archived in another patient exam record. a restart of the system at that time would have prevented the malfunction from happening. as a preventive action, agfa has recommended the following to the customer: restart the system on a more regular basis. a daily restart would certainly be sufficient to prevent memory issues. the system is configured to display images by default in print mode although print automation is not active. reconfiguring the system to display images in softcopy mode further reduces the impact of the memory leak on memory consumption. agfa solution development is underway and a correction will be implemented for this issue. the memory leak is resolved in nx 3.0.9000. a supplemental report will be provided. there has been no reported patient harm for this event. | 1
6215319 | the patient underwent surgical procedure of radical prostatectomy with regional lymphadenectomy" by two urologists. the surgeon documented during the procedure that olympus thunderbeat was used for hemostasis after using blunt and sharp dissection on the nodal tissue from the external iliac vein and freeing lateral rings. during the procedure it was reported that a probe area on thunderbeat machine showed damage error that required opening up a 2nd (tb-0920oc) to finish the case. no injury or further complications were reported for the patient." | 0
6215324 | ><patient found on floor, back portal of isolette found open. | 0
5367822 | the omnipod failed to deliver insulin. | 0
6056960 | i use the omnipod insulating pump system. the omnipod has an adhesive that allows the insulin pump pod stick to my skin. i have been using the omnipod system for 8 plus years but have recently developed an allergic reaction to what appears to be the adhesive. the reaction consists of blisters, sores and a terribly itchy rash. the omnipod is supposed to last 3 days prior to changing the site. however because of the allergic reaction i am having to change omnipod sites after 1-2 days of placement of the omnipod. based on diabetes message boards it appears that the manufacturer of the omnipod system may have changed the adhesive which is causing my allergic reaction and many other's reactions as well. i am getting very little sleep, high blood sugars and am having to take antibiotics and a topical steroid cream to treat the allergic reaction areas. | 0
5929494 | i used the omnipod insulin pump. there is an adhesive that is used to attach the pod to various places on your body. my last shipment came about 6 weeks ago, and at that time i started having reactions to the site, redness, itching, raised skin, blistering with weeping. when i called omnipod, they sent some replacements pods and when i started using them, the reactions got worse. when i called omnipod again to report the problem for the second time, they were not helpful at all and said there was no problem with their product and they had not changed anything. even though the problem started with the first pod out of the new shipment. their solution is to put other products, other adhesives on top of the product i'm having trouble with, it's getting to the point where i can't use them anymore. | 0
6037595 | (b)(4). consumer did not provide model number and therefore, actual place of manufacture could not be determined. the u.s importer is requesting manufacture of the device to further investigate this incident. a root cause has not been determined. it has not been confirmed that the blood pressure monitor provided an inaccurate result or if the device caused or contributed to the reported incident. however, due to the need for medical intervention and in an abundance of caution, this medwatch is being filed.><(b)(4). additional information added to device evaluated by mfr?, evaluation codes and additional mfr narrative. multiple attempts were made to reach the consumer to obtain additional information, however were unable to reach consumer. in addition, a postage paid label was sent to retrieve the unit back for inspection. the device has not been returned as of 12/1/2016. the model number and lot number were not provided; therefore, a review of the manufacturing records could not be completed. without return of the product, the device manufacturer is unable to perform a complete investigation into the root cause of the reported event. at this time, the manufacturer cannot confirm the reported issue. the device manufacturer reviewed the risk analysis for major wrist blood pressure monitors (bpm). events/problems relating to inaccurate readings have been analyzed. risk mitigations, warnings and cautions are documented in the risk files as applicable. review of complaint history showed there was a reportable event for wrist bpm. additionally, the unit was not returned for the previous reportable event therefore, a problem with device could not be confirmed at that time. the device was not received for evaluation; consequently, a device analysis could not be completed. should additional relevant information become available, a supplemental report will be submitted.><a letter was received on (b)(6) 2016 addressed to omron board of directors, trustees, and associates of omron. in the letter sent by the consumer the following comments were stated. yesterday, her husband returned home from a hospital stay from a near stroke experience. he had been feeling ill, but mostly just feeling nauseated. nausea is not enough to go to the doctor so they looked for other vital signs and symptoms. they checked his blood pressure multiple times on the wrist device with completely normal readings. she insisted that he go to the hospital, in spite of the normal bp (blood pressure) readings. they arrived at the hospital and his blood pressure was 260/196- well in to stroke range. the hospital doctor informed them that in no way, shape, or form, was he anywhere close to being normal. they also informed that his readings have not been normal for some times. the hospital also said that whatever device consumer used to obtain these readings was inaccurate and was almost the cause of her husband's death. they also said that not only did that device not give accurate readings, but the lack of those readings almost killed her husband! consumer wrote she understands that omron product does not replace the role of family doctor, nor did they expect it to. they used it to supplement doctor's visits and to monitor home health to stay ahead of any abnormalities. furthermore, consumer wrote the wrist bp monitor was more than 20-40 points off from an accurate reading, and it was not giving readings anywhere close to indicating a problem. not only did the device fail in its ability to give an accurate reading, but it isn't giving a reading at all. they were told not to worry about the nausea because without high bp readings he would be fine. they monitored his bp four-five times a day on omron's device. the bp machine showed perfect bp over two weeks. consumer wrote that omron almost stole life of her husband. she has problem with omron selling products and claimed she will be pursuing legal action and going to the national press to ensure omron is no longer allowed to endanger the lives of family. she not only would want her money back for the machine but wants appropriate damage compensation including medical bills and pain and suffering. quality analyst (qa) made multiple attempts (on (b)(6) 2016) to reach consumer to obtain additional information. no answer. qa left message with consumer requesting a call back. in addition, consumer was sent a written request requesting consumer to call to provide additional information. no additional information received. | 0
5905420 | the icup" point of care test device is marketed such that the result is supposed to not be subjective and/or require any interpretation of the results. however with this particular device the result appears as a "red line". however this is not made clear enough in the... the drug screen results were misinterpreted given the fact that this device is 510(k) cleared but still requires interpretation of the result which defeats the entire purpose. there are untrained unqualified individuals interpreting complex toxicology results and interpreting these themselves. this resulted in 2 erroneously reported drug test results. this then led to 2 incarcerations which are now on my record the results were leaked to my 2 professional licensing boards who used these results to investigate me suspend each license for 18 months and then permanently revoke both of these - my livelihood is taken away. further each result was used to substantiate a criminal charge. in total i spent 45 days incarcerated have 2 criminal charges lost my 2 professional licenses and am living off (b)(6). i have not worked in 3 years and this is the fault of this device which is not marketed properly and should not be waived! there is a problem with this particular drug screen device in that it does not make clear the cross reactivity of a specific product-nasal decongestant nasal inhaler. the product monograph is written in 6 point font of smaller and is almost impossible to see visually. they do have listed l-ephedrine on the non-cross reactivity list. the first issue is that the drug has been renamed a number of years ago - at least 10 years - and is now named l-levmetamfetamine. this is the first point of confusion. second it is listed as non-cross reactive when in fact this is not the case and it does cross react with this drug product when used at the recommended dosage. there is no information provided on this by the manufacturer which is severely lacking and deceptive to the public. incarceration revocation of professional licenses substantiation of criminal charges false reporting of a relapse." | 0
5437788 | merge unity pacs is a medical image and information management system that allows viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. customer reported that a good numbers of reports have been missing headers. reports are for distribution to referring physicians. if the patient demographics is missing and the report is distributed via fax or printer, the report could not be associated to any patient since no patient demographics are available. this event did not result in a serious injury to the patient. (b)(4).><upon retrospective review, this issue was determined to be an mdr. | 1
5416623 | on (b)(6) 2015, we upgraded our telemetry system from philips classic to philips picc ix on one of our nursing units. on (b)(6) 2015 we experienced a central station going completely black rendering the staff unable to see a monitor rhythm. the system rebooted itself. philips was notified and responded. since (b)(6) 2015, there have been 7 instances of this happening with the most recent incident (b)(6) 2016. the reboots can take from 5-30 minutes leaving our patients unmonitored. philips has held conference calls and have attended meetings with no resolution. last week, philips informed us that the same problem was identified in 5 other hospitals since (b)(6) 2015. our patients are at risk, we need a resolution to this issue. another product that is not working correctly is philips mx40 telemetry receivers. | 1
5766357 | (b)(4) reports that the dm3 monitor (91330) failed to alarm for low nibp on patient that coded. patient admit for cerebral infarction and heart disease. patient had multiple stents. new stent @ (b)(6) on (b)(6) 2016. admit to (b)(6) 2016 at 02:30. patient code on (b)(6) 2016 approx. 09:10 in neuro unit on dm3. patient code on (b)(6) 2016 approx. 10:34 in icu. patient passed on (b)(6) 2016 at 15:20. hosp. risk manager still collecting data. dm3 was removed from patient room by (b)(4) at request of nursing for testing. monitor was turned off and defaulted back to spot mode. (b)(6).><. | 1
5832256 | edan instruments investigated that the low values were primarily due to inaccurate calibration performed at the factory. notification was sent to the distributors of the products, and corrections includes replacement of co2 module. the affected unit was tested for the non-stable readings during usage. during testing, the unit worked as intended with stable values. investigation will continue to determine if device malfunction occurred, or if there was user error during this first usage of the device.><customer reported an issue with the incorrect co2 reading during dental anesthesia usage. co2 readings were both low and not stable, occasionally going to zero. due to the low values, the patient was transferred a local hospital to be weaned off of anesthesia. no sequelae were observed and there were no adverse events for the patient. | 0
5693119 | a follow up report will be submitted once the investigation is complete.><.><the customer reports that on (b)(6) 2016, a mp70 patient monitor from intensive care unit of (b)(6) hospital fell off from the rack, there was no patient on the bed and did not fell on any patient or health professional. it has been confirmed there was no health impact at any time neither to a patient nor a bystander due to the damaged mount with a monitor falling off.><the customer reports that the mp60 patient monitor fell off from the hanging tower's rack on (b)(6). no person was injured. there was no report of any adverse patient impact resulting from the reported problem. | 0
5959808 | the customer reported that the audio monitor does not work. the device was not in clinical use at time the issue was discovered. there was no adverse event or patient harm reported.><. | 1
5780418 | .><removed c50097-printed circuit board.><the customer reported that the mx700 monitor show message speaker malfunction. no patient or user harm was reported. | 1
5717715 | the customer reported that the mx500 monitor shows a message of speaker malfunction. the device was used for monitoring at the time of the alleged malfunction. no patient or user harm was reported.><. | 1
5580667 | a follow up report will be submitted upon completion of the investigation.><the philips field service engineer (fse) went to the customer site. it was established that the patient in bed (b)(4) went through a run of premature ventricular contractions (pvc) on (b)(6)-2016 around 10:14 a.m. and the intellivue mp5 patient monitor did not alarm for this condition. the patient was having chest pains and was subsequently taken to the cath lab for emergency treatment. the patient survived the incident and was later discharged from the ward. no emergent or adverse impact was reported.><. | 1
5550330 | a philips fse tested the device and found the device working to specifications. the user confirmed that during the asystole, the electrodes were detached from the patient and the device provided an inop alarm for this condition. the fse trained the customer on the device operation. the complaint is considered as user related and no malfunction of the device. the reported issue is readily detected and poses minimal risk to patient safety. the device provided an inop message (and alarm tone) to alert users to a problem. device labeling (instructions for use) describes alarm behavior. no further investigation or action is warranted.><the customer reported that the monitor does not sound an alarm in case of asystole. the device was in use during monitoring a patient. the patient received an asystole and the monitor did not sound for that condition.><a follow up report will be submitted once the investigation is complete. | 1
5412425 | .><the customer reported there were no audible alarms and no qrs tone on the mx450 monitor. the device was not in clinical use at the time of the event. no report of patient or user harm. | 1
5376171 | .><the customer reported that he is not get any tones from device during power on test. at the time of the alleged malfunction, the device was being used for clinical monitoring. no patient harm was reported. | 1
5390254 | .><the customer reported a speaker malfunction alarm on the display. at the time of the alleged malfunction, the device was being used for clinical monitoring. no patient harm was reported | 1
5813537 | the customer reported that on (b)(6) 2016 at 10:56:08a. the patient in room ((b)(6)) ecg hr dropped and rhythm showed asystole event but the monitor did not trigger the asystole alarm event until 10:58:18a. the caregiver claims mp70 asystole alarm was delayed up to 2 minutes after event. the patient had to have chest compression as a result of the event but expired.><the complaint is considered as user related and not a malfunction of the device. logs were sent for review and show that the alarms were paused for 3 minutes at 10:55:08. at 10:58:13 the logs show that the device provided an asystole alarm at 10:58:13 as configured instead of the expected time of 10:56. the device is designed to have the ability to configure pause alarms at intervals of one minute, two minutes, three minutes, or infinitely. device labeling adequately describes alarm behavior when alarms are paused noting that the monitor displays the message alarms paused and the red alarms paused lamp on the monitor front panel is lit. | 1
6177493 | a follow up report will be submitted upon completion of the investigation.><the customer reported a loudspeaker error. there was no reported patient involvement. | 1
6029626 | the customer reported that the system had limited use". the loudspeaker of the intellivue mp2 patient monitor was malfunctioning. the device was not in clinical use at the time of the alleged malfunction.><a follow up report will be submitted upon completion of the investigation.><." | 1
5694133 | the customer reported that the system was not alarming. further information received stated that on (b)(6) 2016, the rn went into room 4.412 on 4 south to deliver care at 15:36. at this time, the rn noted the patient not breathing; on assessment patient noted to have no pulse. cardiac arrest called, immediately started acls, and team arrived to help. the patient died.><there was no malfunction of the device. a review of the provided alarm logs found that the monitor had alarmed, during the time of the reported event, and had been silenced. a philips field service engineer (fse) went on site, and found the device to be in use (it has been used on other patients, since the initial report) and working properly. the fse was able to verify operation of the monitor where red and yellow alarms were simulated and both visual and audible alarms were verified. the device remains at the customer site, and no subsequent calls have been logged for this device. no further investigation is warranted at this time. | 1
5903417 | a follow up report will be submitted once the investigation is complete.><the report of the alarms not going off was not supported by the monitoring data. there was no malfunction. a review of the alarm logs also showed several alarms at the time of the reported event, including instances of the alarms being silenced. philips cannot determine whether the delay in treating this patient was causal or contributory to the death. the device remains at the customer site, and no subsequent calls have been logged for this device/issue. no further investigation is warranted at this time.><the customer reported that the monitor did not trigger an alarm between 1:30-1:45am on (b)(6) 2016. the patient was connected to a monitor. only ekg and af were controlled, and blood pressure and spo2 were not measured. no alarm was generated regarding the personnel. no ecg alarms were transmitted to the monitoring center. the attending clinician was aware of the patient's condition, when the issue occurred. it was noted that there was a delay in treatment, due to this issue. the patient died on (b)(6) 2016. the customer suspected a stroke, patient died after coronary and not executed cpr.><the customer reported that the monitor did not trigger an alarm between 1:30-1:45am on (b)(6) 2016. the patient was connected to a monitor. only ekg and af were controlled, and blood pressure and spo2 were not measured. no alarm was generated regarding the personnel. no ecg alarms were transmitted to the monitoring center. the patient died on (b)(6) 2016. the customer suspected a stroke, patient died after coronary and not executed cpr.><based on the information contained in the alarm logs, the arrhythmia alarms were turned off the day prior to the reported event; the arrhythmia alarms were turned back on after the incident. there was no product malfunction. a review of the alarm logs confirmed the status of the arrhythmia alarms being off during the time of the reported event. philips cannot determine whether the delay in treating this patient was causal or contributory to the death. the device remains at the customer site, and no subsequent calls have been logged for this device/issue. the ifu for this device states warning selecting pulse as the active alarm source for hr/pulse switches off the arrhythmia alarms listed in the section ?ecg and arrhythmia alarm overview? on page 165 including asystole vfib and vtach alarms and the heart rate alarms. this is indicated by the message ecg/arrh alarmsoff (unless this has been configured off for your monitor) and the crossed out alarm symbol beside the ecg heart rate numeric. the message ecg/arrh alarmsoff can be configured off or to switch to a yellow (medium severity) inop after a fixed number of hours. high and low pulse rate and extreme bradycardia and extreme tachycardia alarms from pulse are active." no further investigation is warranted at this time." | 1
5971289 | the customer reported a loudspeaker failure. at the time of the alleged malfunction, the device was being used for clinical monitoring. no patient harm was reported.><. | 1
5554430 | .><the customer states that the unit is giving a speaker malfunction error message. the unit is not making a sound, and when an alarm happened it does not alarm. at the time of the alleged malfunction, the device was not being used for clinical monitoring. | 1
5586490 | .><the customer reported that the device indicates a loudspeaker error and no alarm sound can be heard during turning the device on and during operation. at the time of the alleged malfunction, the device was being used for clinical monitoring. no patient harm was reported. | 1
5535934 | the customer reported a speaker malfunction. at the time of the alleged malfunction, the device was not being used for clinical monitoring. no patient harm was reported.><. | 1
5551007 | the investigation revealed that the device worked as expected and no malfunction of the device could be identified. post incident testing of the monitor found alarms working per specification. log files show several alarms for a vent fib/tach alarm at 16:08 which were immediately silenced by the user. no further investigation or action is warranted.><the customer reported that on (b)(6) 2016 around 16:08 for bed (b)(4), a red alarm was generated but the audio for the alarm was not heard. the patient's scream alerted the nurses to the patient's v-tach event. | 1
6100895 | it was reported that the non invasive blood pressure hose used with the passport 12m patient monitor was leaking and the monitor presented an erroneous value. the hospital staff treated the patient under the impression that these were accurate readings and administered medication. it was only after the patient did not respond to the medication that the blood pressure readings were found to be reporting lower than actual.><(b)(4).><five defective hoses were returned for evaluation. it was determined that the leak was caused by the deterioration of the inner o-ring of the nibp hose, due to the use of non-recommended cleaning agent by the customer. the customer was provided a list of recommended cleaning agents, also located in the operator's manual. mindray ds, usa, inc. ((b)(4)).><it was reported that the non invasive blood pressure hose used with the passport 12m patient monitor was leaking and the monitor presented an erroneous value. the hospital staff treated the patient under the impression that these were accurate readings and administered medication. it was only after the patient did not respond to the medication that the blood pressure readings were found to be reporting lower than actual. | 1
5744800 | the customer reported no temperature alarm. there was no report of a death or serious injury, nor was there a report of any adverse impact to any user or patient.><.><. | 1
6072489 | .><the customer stated that the device gives audio error, sound not working. there was no report of a death or serious injury, nor was there a report of any adverse impact to any user or patient. the device was in use at the time of the reported issue.><a follow-up report will be submitted once the investigation is complete.><the customer stated that the device gives audio error, sound not working. there was no reported patient incident injury. | 1
6136997 | the customer stated that the gasket has become unseated in the lead set. there was no reported patient incident injury.><.><phone number removed from grid - failing electronic submission. a follow-up report will be submitted upon completion of the investigation. | 0
6127717 | customer reports respiration rate values are too high.no patient injury/ harm was reported.><a follow-up report will be submitted upon completion of the investigation.><. | 1
6124806 | the investigation was started but has not yet been concluded. the investigation result will be reported in a follow up-report.><it was reported that the dr?ger monitor broke away from base plate. it was just sitting on the base of the machine and was able to be pushed off, simply by touching the dials. the top of the dock where the monitor locks into position has lifted away from the dock assembly. there was no patient involvement and there was no injury reported. | 0
6135258 | i was informed that a serious arrythmia alarm (ssystoly) occurred with a patient. the telemetry system did not sound an alarm, the visible alarm was present and the duplicate display on another floor had an audible alarm, when the alarm was not cleared in a timely manner the other floor called 2a and informed them. the alarm was false and the patient was not in distress. that day, we had been testing alarm integration into vocera (a voice communication system for staff. i was called, responded within 15 minutes, and rebooted the central station pc. tested and noted alarm sounds were working. the next day i informed (b)(4) (patient monitoring vendor) and vocera of the problem. i was told in the afternoon that the cause had been identified and resolved (conflict between 2 systems caused failure). i'm not sure i understand that and have asked for a written explanation of the problem. the same issue occurred a few days earlier when we were testing the system on the third floor. i did not put the failure and testing together that day and rebooting resolved that issue. during testing on both system the audible alarm worked fine, when the testing/training room was configured back to normal is apparently when the failure begins. the system is not up now and will be adequately tested before being released back to service for any further work on the system.><. | 1
5642870 | .><the fetal heart monitor ultrasound probe burned patient. there was another similar event with the same type probe that occurred recently as well. manufacturer response for fetal heart rate ultrasound probe, fhr monitoring probe (per site reporter): service rep making arrangements for inspection. manufacturer response for avalon fm50 fhm, avalon fm50 (per site reporter): service rep will be making arrangements to have it inspection. | 0
5642869 | fhm probe burned patient's abdomen. manufacturer response for avalon fm50, avalon fm50 (per site reporter): service rep making arrangements to have device inspected. manufacturer response for fetal heat rate ultrasound probe, fhr ultrasound probe (per site reporter): service rep making arrangements for return and inspection.><. | 0
6064812 | a patient's caregiver reported an event of peritonitis during a call to technical support. upon follow up, the patient's peritoneal dialysis nurse stated that a laboratory test of the patient's dialysis effluent revealed the presence of staph epidermis bacteria, evidence of touch contamination. the patient was administered vancomycin. the patient was not hospitalized and last reported to be asymptomatic.><(b)(4)- a follow up mdr will be submitted following the manufacturer's evaluation.><a visual inspection of the returned cycler exterior showed no sign of physical damage. a simulated treatment was performed and completed without any failures or problems. no fluid leaks in the test cassette during the test treatment. the valve actuation test passed. the system air leak test passed. the patient sensor calibration check passed. there were no discrepancies encountered in the internal visual inspection of the cycler. an investigation of the device manufacturing records was conducted by the manufacturer. there were no deviations or non-conformances during the manufacturing process. in addition, the device manufacturing review confirmed the labeling, material, and process controls were within specification. a clinical investigation concluded that there was no information provided that indicated a causal relationship between the use of fresenius peritoneal dialysis products and the peritonitis. the most common cause of peritonitis is a breach of technique or biofilm in the peritoneal dialysis catheter.><the peritoneal dialysis patient reported having an infection and seeing fibrin within the fluid. the peritoneal dialysis nurse confirmed that the peritoneal dialysis patient was diagnosed with touch contamination peritonitis on (b)(6) 2016. no further information is available. | 0
5825825 | a dialysis patient reported they developed peritonitis on (b)(6) 2016. follow-up with the patient's peritoneal dialysis registered nurse (pdrn) confirmed the report of peritonitis. the patient was not hospitalized. the patient was treated with vancomycin (dosage unknown) and later transitioned to gentamycin (dosage unknown.) the pdrn reported the patient was asymptomatic and has remained on ccpd (continuous cycler assisted peritoneal dialysis) therapy throughout the episode and afterward.><the liberty cycler was not repaired or replaced in association with this event. an investigation of the device manufacturing records was conducted by the manufacturer. there were no deviations or non-conformances during the manufacturing process. in addition, the device record review confirmed the labeling, material, and process controls were within specification.><(b)(4) a supplemental mdr will be submitted when plant investigation is complete. medical records have been requested. | 0
5596664 | medical records were received from the patient's treatment facility. a peritoneal dialysis (pd) patient presented to their pd clinic with slightly swollen ankles, abdominal pain that was tender to touch and cloudy effluent on (b)(6) 2016 (not (b)(6) 2016 as previously reported). results of (b)(6) 2016 laboratory test of body fluid revealed gram positive cocci, confirming bacterial peritonitis. the pd fluid showed few white blood cells and the isolated organism was staphylococcus epidermidis. the patient was diagnosed with touch contamination peritonitis and was treated with fortaz (ceftazidime) 1 gram and vancomycin 1 gram via ip route. on (b)(6) 2016, the patient completed his prescribed antibiotic therapy. the patient's pd catheter was removed on (b)(6) 2016; the patient's modality switched to in-center hemodialysis.><the actual device was returned to the manufacturer for physical evaluation. a visual inspection of the returned cycler showed no exterior damage. the heater tray/scale was not obstructed. the load cell value and verification was not within tolerance. after calibrating the scale the load cell value and verification was within tolerance. physical testing was performed and passed. a simulated use test was performed and passed. there were no discrepancies encountered in the internal inspection of the cycler. an investigation of the device manufacturing records was conducted by the manufacturer. there were no deviations or non-conformances during the manufacturing process. in addition, the device record review confirmed the labeling, material, and process controls were within specification.><(b)(4)><a peritoneal dialysis (pd) patient's caregiver called technical services and reported the patient was recently getting over an infection and is experiencing abdominal pain. follow-up with the patient's peritoneal dialysis registered nurse (pdrn) reported the patient developed peritonitis due to a breach in sterility. the patient was treated with an unspecified medication and the peritonitis resolved.><there was no documentation in the medical record supporting a possible association between the liberty cycler and the event of peritonitis. however, there is a probable association between breaking the sterile pathway of pd therapy and peritonitis. | 0
5515912 | patient had bravo capsule placed endoscopically. when receiver was returned, no info was downloaded to receiver. pt had chest pain, capsule was removed endoscopically using snare cautery, capsule was supposed to fall off after 48-72 hours. | 1
5694336 | philips lifelink gosafe medical alert system. communicator does not perform as documented (quick setup guide, p/n 0940833, rev. 01) during setup. setup guide reads turn on the home communicator using the switch on the back." "press the flashing orange message button when you are ready to continue." "plug the power cord into an easily accessible outlet. choose a power outlet that can not be turned off by a wall switch. press the flashing orange message button when you are ready to continue." problem: flashing orange message button does not come on after turning communicator on but before plugging in. when communicator is plugged in the button does flash. the voice guidance that launches when the flashing button is pressed (after the communicator is plugged in) says "now plug me in." confusing to elderly customers because machine does not perform as the setup guide says it should. call to technical support. agent says that the communicator must be plugged in before the message button light will flash. told him that this is not what the instructions say. he replies that instructions are wrong." | 1
5986831 | please refer to the initial report.><the investigation was started but has not yet been concluded. the investigation result will be reported in the follow up report.><screenshots from a technical problem report confirmed that the etco2 values were reading higher than what was expected. the m540 was displaying units of mmhg on the screen and was expecting the values from the micro stream pod in these same units. it was found that the etco2 pod default units were set to %volume, and therefore the etco2 measurement values are not accurate. a field safety corrective action has been opened and a safety notice has been sent out to current users of iacs vg5 for a mandatory downgrade to vg4. the field safety corrective action released for this issue is not applicable for the devices in the us market where the affected software version is not used.><it was reported that, etco2 values from iacs microstream were consistently and significantly higher than pco2 from blood gas machine. the etco2 and pco2 was compared on 5 patients. the etco2 was approximately 10 mmhg higher in 3 patients and 15 and 20 mmhg higher in two patients. there is no patient injury reported. | 1
5984627 | it was reported that, after installation 9:30pm (b)(6), all functions of iacs are fine, however, the next morning the clinical specialist noticed that whilst the m540 was acoustically alarming (which it shouldn't in an iacs configuration), and the cockpit was not annunciating an alarm. further communication with customer indicated the m540 intermittently had no audible alarm. there is no injury reported.><the iacs logs from the event were reviewed by engineering. a video capturing the issue was provided. after reviewing the provided video and corresponding logs, it could be confirmed that the pulse tones and audible alarms were not working during the event after being in standby. the no audio issue at both the cockpit and m540 was reproduced during the investigation. it has been determined that there is a bug in the microsoft windows audio service subsystem that causes a loss of all audio. a field safety corrective action has been released to current users of iacs vg5 for a mandatory downgrade to vg4. the field safety corrective action released for this issue is not applicable for the devices in the us market where the affected software version is not used.><please refer to the initial report.><the investigation was started but has not yet been concluded. the investigation result will be reported in the follow up report. | 1
5895926 | .><rn received report from nicu rn. it was reported at that time that the baby had a little mark on his left foot that seemed to be an 02 sat probe burn (redness). rn noted it in charting that day as pt was transferred. noted that mark was first charted in nicu notes on the previous day at 09:00. | 0
5667884 | the investigation was started but has not yet been concluded. the investigation result will be provided in the follow up-report.><it was reported that on (b)(6) 2016 at 12:29 a covidien adult single-use spo2 sensor failed without an alarm being generated by the delta xl monitor. no measurements and waveforms were displayed, and the red light of the sensor did not glow. after replacement the measurement worked again. it was reported that the event led to deterioration in health of the patient. no serious injury has been reported.><the bedside monitor diagnostic logs, the infinity central station (ics) logs, and the spo2 screenshot from the ics was received for evaluation. the diagnostic logs show the monitor logged a nellcor defective sensor error at the time of the event 11:14:06 on (b)(6) 2016. this indicates that the delta monitor detected a defective nellcor sensor. the bedside alarm history log was not received to verify the alarm condition however the ics log at 11:14:39 on (b)(6) 2016 indicates there was user interaction at the central station pressing alarm stop and alarm silence. this means the bedside monitor alarmed when the reported issue occurred and also alerted staff at the central station by relaying an alarm to take action. by design the delta displays blank" in the box until it can find the patient baseline and then calculate the numerical value which can take up to 15 seconds. if the device is unable to calculate the waveform data then an error ?***? notification will alert the user of the issue and a banner will be displayed at the bedside monitor with the error. the reproducibility testing using the same setup as the customer showed the device behaved as intended to annunciate alarms. it's reported after replacing the defective sensor at customer side the device works properly again. there is no error found regarding alarming function.><please refer to the initial report." | 1
5682756 | ics logs were submitted for analysis. the logs of m300 cited in this complaint was not provided as it could not be determined which specific unit was exhibiting the cited issue. it was also communicated that a new patient was admitted on the m300, and the records that covered the date of the cited issue were not available for the m300. in the ics clinical logs where a discharge would be evident, there were no discharges identified. as a result a hardware analysis would need to be conducted, however, the m300 was not available for the root cause investigation. the exact time of event was not communicated however for the identified ip address (11 192.168.72.110) the ics showed 3 bed no longer being monitored" on the ics for the date cited ((b)(6) 2016). a precise root cause as to why the issue would occur was not able to be determined as requested logs and material is not available. by design if the m300 turns off or discharges an audio and visual offline/discharge message will be displayed at the ics as well as the m300. in this case the ics logs showed numerous offline messages indicating that the bed specific to this complaint was not being monitored.><please refer to the initial report.><the investigation was stated but has not yet been concluded. the investigation result will be reported in the follow up report.><it was reported that during the operation of the m300 transmitter (m300) repeatedly disappears from the screen ( discharge itself) and then must be admitted again. it could be named only on transmitter (no. 9) with this error. no patient injury reported." | 1
6090775 | there was no alarm to indicate the device was too close to the magnetic field. there was no reported patient harm.><further clarification reveals the issue was no audible alarm of the magnetic field was provided by the mri monitor. the issue was discovered during planned maintenance activities. the alarm card ((b)(4) was returned to ge healthcare engineering for investigation in which the cause was determined to be a single component failure: the sonitron buzzer in the alarm card was not working. the alarm leds and the circuits of the alarm card were functioning correctly. the magnetic field alarm functionality is as follows: when the magnetic field where the mri monitor is located gets stronger than 150 g yellow light is lit in the high magnetic field alarm led on the front panel of the shield. when the magnetic field gets stronger than 250 g the led starts blinking red and an audible alarm is triggered. when the monitor is turned on the high magnetic field alarm alarms once."><incident is from 2016. specific date is not known at this time. procode is not known at this time. the 510k# is not known at this time. date of device manufacture is not known at this time." | 1
5438252 | the customer reported a failure of the defib sync port on the dash 4000 patient monitor. not in patient use when failure occurred.><ge healthcare's investigation is ongoing. a follow-up report will be submitted when the evaluation is complete. device evaluation anticipated, but not yet begun.><the defective main board is no longer available to be returned for investigation. the customer biomed cleaned the defib sync port however, it did not resolve the issue. the biomed then replaced the failed dash central processing unit (cpu) board, which includes the defib sync port. this resolved the issue. the issue with the dash defib sync port was caused by failure of the dash cpu board. | 0
5485557 | a philips field service engineer (fse) was dispatched to the customer site to evaluate the device and wireless ecg module. per the fse, the hospital biomed asked if he could do a preventive maintenance on the device and wireless ecg module. according to the fse, he found nothing out of calibration and the device and wireless ecg module performed properly. according to the customer, the wireless ecg module was located in between the patient and the gurney. there were no melting or burn marks noted on wireless ecg module and the same wireless ecg module has continued to be used without issue. following the incident, the biomed tested the wireless ecg module and was not able to find any problems with it. additional information regarding the incident was requested, but additional information was not received. based on the available information, it is possible that the injury was caused by skin to skin contact between the patient?s thumb and hip. when this occurs the patient?s body becomes a conductive path by the creation of a loop. this skin to skin contact can be reduced by placing foam pads or other nonconductive material between the patient?s hand and leg, if necessary. the wireless ecg module and patient monitoring device remain at the customer site and both continue to be used without further incident. a search found no further related calls. the absence of further calls supports that the reported problem has not recurred. the customer claimed that a patient was burned on his hip and finger. the customer believes the wireless ecg module is what burned the patient. hospital clinical engineering/biomed staff reported that they did not think the patient was treated. it was reported that the patient was immediately discharged and when the customer followed up a few days later, the patient said issue had cleared. testing completed by both the hospital?s biomed and a philips fse found that both the device and wireless ecg module performed properly. based on the available information, it is possible that the injury was caused by skin to skin contact between the patient?s thumb and hip. when this occurs the patient?s body becomes a conductive path by the creation of a loop. this skin to skin contact can be reduced by placing foam pads or other nonconductive material between the patient?s hand and leg, if necessary.><the customer claimed that a patient was burned on his hip and finger. the customer believes the ecg module is what burned the patient. it is not known at this time if or how the patient was treated for the injury.><a follow up report will be submitted after philips obtains more information concerning this event. | 0
5977812 | the customer reported that the device was not producing any alarms and that the alarms appear to be muted. technical support (ts) checked volume settings and found them to be appropriate. there was no reported patient impact.><a follow up report will be submitted once the investigation is complete.><.><the customer reported that the device was not producing any alarms and that the alarms appear to be muted. it was later clarified that both the alarm sound and all other sounds (touch and qrs tone) were not working. technical support (ts) checked volume settings and found them to be appropriate. the biomed stated he tried turning the ip5 off and on but still no sound. there was no report of a death or serious injury, nor was there a report of any adverse impact to any user or patient. the device was being used for patient monitoring when the problem occurred. | 1
5434056 | per depot repair, the tram 451n central processing unit (cpu) board was discarded and is not available for further investigation. the tram sends alarm information in parameter data packets for visual display and audio processing to the host monitor (a solar 8000 in this case). the tram software determines when a parameter is in alarm based on parameter limit settings stored on the tram main cpu board memory components. failure of tram main cpu board memory components due to electrical, thermal or mechanic stress over time, could result in the tram not providing alarm information to the host monitor. based on the information available, the root cause of the tram not producing audible alarms on the solar monitor was determined to be caused by failure of the tram main cpu. the tram main cpu board was replaced.><the customer reported the tram did not produce visual or audio alarms. the tram was not in use when the issue was noted. no patient involvement.><there was no patient involvement. date of event is not known. ge healthcare's investigation is ongoing. a follow-up report will be submitted when the investigation is complete. device evaluation anticipated, but not yet begun. | 1
5851400 | problem description: following a client complaint, intelerad has identified a software problem in the image fusion module which may display an incorrect calculation of standard uptake value (suv) when loading the presentation state of image for a previously calculated (and saved) region of interest (roi). this problem could lead radiologists to base their diagnosis/report on incorrect anatomical measurements leading to incorrect medical treatment. no patient incident has been reported at this moment - this problem is documented as a precaution. investigation details: one customer complaint ((b)(4)) bug identified: (b)(4). this is a bug due to a missing coordinate transformation, which manifests itself only when the ct axial spacing is not 1mm. most if not all of our testing data to that point had been from ge which had 1mm spacing, so this bug was not seen, until we had a pet/mr dataset with axial spacing not equal to 1mm. even then, the bug as stated was that the max pixel was not in the cubical voi region. we did not note that the max pixel was incorrect for the given region. the bug only applies to restored presentation states; any modification of a voi read from presentation states causes max suv to be calculated correctly the incorrect max suv is that of an adjacent region, so is random, and more than likely not to correspond to the actual region of interest, therefore the error should be noticeable. the displayed suv at the cursor location is correct, contradicting the incorrect max suv. the go to max" associated with the voi will triangulate out of the voi box indicating that something is wrong. all of the max suvs in the presentation states will be incorrect as above increasing the likelihood that the error will be noticed the fix ((b)(4)) is available in inteleviewer 4-11-1 and later versions. the investigation was conducted on the client's systems and was reproduced on intelerad's systems (inteleviewer version 4-101-1-p180). assessment of risk to user: description of hazardous situation in use of the device: incorrect calculation of suv when loading presentation state of image for a previously calculated (and saved) roi. this may lead to incorrect diagnosis. this problem is currently documented as a known risk: risk #2.3 in risk management summary (eng-007 issue 21). the mitigation in place reduces the risk for the general case of suv measurements but not for datasets such as the one that caused the incident nor the specific use case of displaying a measurement when reloading a presentation state. the risk can be further reduced through additional design factors taking into account specific datasets and the case of reloading the presentation state. probability of risk: before mitigation: occasional (3). after mitigation: remote (2). severity of risk: before mitigation: injury requiring treatment (3) (in the worst case). after mitigation: injury requiring treatment (3) (in the worst case) . detectability (user awareness. if device failure occurs is it easily recognized by user): the reported incorrect suv is more or less random so does not correspond to the associated volume. detectability increases with the number of measurements since the lack of correspondence between incorrect values and associated volumes becomes more obvious. can be recognized in the normal case by trained users. risk vs. benefit comparison (to help determine the recall strategy and type of the recall): reloading a saved presentation state is a basic function of the device. the risk is small compared to the benefit as it is a significant time saver and normally reduces the risk or errors of repeating measurements by hands. how often does the problem occur? how often can the problem cause direct or indirect harm to the patient? only caused by data where the axial spacing is not 1mm. it may cause harm if it goes undetected and there is a major discrepancy between the real value and what is displayed and there is no other means to confirm the diagnosis before the patient is treated. the device has been in use for two years with only a single reported incident with no indication of real harm. describe the factors that may contribute to product risk to the patient problem occurs for patient high/low resolution ct or mr scans where axial resolution is != 1 and where patient measurements are stored for later reload. factors that may contribute to product risk (i.e. device design manufacturing problems or use error): as above. risk assessment summary: function: display previously saved suv measurement. failure mode/defect: incorrect suv displayed. local effect: incorrect measurement leading to incorrect diagnosis if undetected. system effect: none. adverse health consequences: incorrect diagnosis leading to incorrect treatment if undetected resulting in harm from incorrect treatment or missing/delayed treatment. severity: 3 injury requiring treatment. probability: 3 occasional. risk rating: medium. corrective action: removing the affected version will prevent users from encountering the problem. proposed recall action: intelerad will: identify all clients affected i.e. all clients with access to the image fusion module of intelepacs - 31 august 2016. inform all concerned clients of the issue (recall letter) - 31 august 2016. provide all clients that do not already have a fixed version of the device with a software update - 30 september 2016. estimated completion date of the proposed recall action is end of september 2016." | 1
5809958 | the customer reported that the stat priority icon was not displayed in the study list when the study was opened. no patient harm was reported due to this issue.><mckesson completed an investigation and concluded that the root cause of the problem was a software defect. mckesson found that the stat priority alert icon may not be displayed in the list, for any user, when all stat studies are opened at the same time and there are no high" priority studies in the unreported study list. under these conditions the stat priority button will not display the correct number of stat studies. mckesson will work with affected customers to apply a software update to prevent the recurrence of the problem." | 1
5437884 | upon retrospective review, this issue was determined to be reportable as an mdr.><an iconnect enterprise archive customer reported that reports are not matching/reconciling after name change. customer further explained if a trauma patient comes in and patient information is not available immediately, patient is stored under doe patient. when the report is being transcribed as trauma and patient has been discharged, the report type does not match the study after the name change. (b)(4). | 1
5480535 | the customer reported that the incorrect patient is displayed when clicking directly on a patient name in the dexis (patient) administration screen. no injuries have been reported.><the customer used a scenario which was not indicated in the dexis user manual. a software evaluation was performed on dexis 10.1.2. evaluation revealed that the reported issue is only observed when multiple instances of dexis (patient) administration are active, which is not considered to be a normal scenario for the use of dexis. the issue can be resolved by reloading the patient list in the active dexis administration window. the user is able to determine the correct patient image and no injuries have been reported. this concludes our investigation. | 1
5475692 | merge cardio is a system intended to be used to acquire, store, print, transfer, and archive clinical information including images, hemodynamic studies and reports, measurements (via import from dicom structured reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. it was reported that the report is missing a piece of information that it says should be there. specifically, the report says to look at 3rd lesion intervention, but in that information was not available to the user. technical support discussed this with the customer and explained what to do to get the 3rd intervention section to display. steps (including a screen shot) were provided on how to do this (e.g., procedure narrative>interventions and select the intervention number). there was no report of patient injury. (b)(4).><upon retrospective review, this issue was determined to be reportable as an mdr. | 1
5763242 | merge healthcare is further investigating the allegation from the customer to determine if any corrections or corrective actions are necessary.><merge unity pacs a medical image and information management system that is used for viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. on june 1, 2016 customer reported the distance indicator between image slices are not reflecting what is displayed on the modality. it is intermittent and is not presently known if it is limited to one modality or vendor. merge healthcare was notified by the customer that they were informed by american college of radiology (acr) that their facility did not pass accreditation due to spacing listing on images not consistent throughout the series. initial evaluation of images shows the sagittal images within a specific series should be spaced 3.9mm apart. the modality shows the consistent spacing, however, the unity pacs viewer shows a 4.0mm spacing between 2 of the images. due to the potential for the inconsistency in spacing measurements that is displayed on some images, there may be confusion which could lead to a potential delay in diagnosis or treatment for a patient. there is no indication that any patients were harmed as a result of the displayed inconsistency in slice measurements. there was no reported adverse event to a patient. (b)(4). | 1
5475697 | upon retrospective review, this issue was determined to be reportable as an mdr.><merge cardio is a system intended to be used to acquire, store, print, transfer, and archive clinical information including images, hemodynamic studies and reports, measurements (via import from dicom structured reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. it was reported that studies did not contain hemo data. it was further reported that multiple studies showed this problem. if there is an automatic association of images and hemo data, then it shows on the webviewer and the workstation that there is hemo data for the study. if there is a manual association of the two, it says there is no hemo data for the study. even when the report is started, and even if it says there is hemo data, there is no hemo data in the report. it was determined by customer support that there weren't any hemo transfers in place. the issue was resolved and data are populating the report. there were no reports of patient injury. (b)(4). | 1
5623275 | merge healthcare is continuing to evaluate this allegation to determine if any corrections or corrective actions are required. depending on the results, a supplemental report may be submitted.><additional information received from the customer confirmed that there were no known adverse events or harm to a patient. the assignable cause of the failure was not confirmed due to the logs from the time period indicated being no longer available. this issue will continue to be monitored. a review of the complaint history indicates there is no significant trend. trending will continue to be monitored. revised information contained in the supplemental report includes the following: on 6/3/2016: new information received by manufacturer.><merge unity pacs a medical image and information management system that is used for viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. on 04/1/2016, merge was notified that on an unknown date, studies were marked read and the reports were approved, however, no dictation or reports were found. merge support was able to recover all but 2 reports, which had to re-read and re-dictated. however, a cause of the issue could not be determined since the reports were older, and no logs could be reviewed. there was no reported adverse event to a patient. however, a missing report has the potential to delay patient treatment and/or diagnosis, especially if the physician or health care professional is not aware that the report is missing. further information is needed to clarify this event. if additional information is received, a supplemental report will be submitted. (b)(4). | 1
5867550 | after further investigation by merge support, the cause of the missing report could not be determined as the log files were truncated, however the most probable cause was an issue with uploading the report to the server.><merge unity pacs a medical image and information management system that is used for viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. on (b)(6) 2016, merge was contacted by a customer indicating that they were unable to find a report for an exam that was saved under pending approval status after dictation. the merge support technician was not able to find the report or the word document file for the report. due to not being able to find the report, re-dictation was necessary. there was no reported adverse event to a patient. however, not being able to find a report that then needs to be re-dictated, there is a potential to delay patient treatment and/or diagnosis. (b)(4). | 1
5521088 | upon retrospective review, this issue was determined to be reportable as an mdr.><officepacs is used for the manipulation and displaying of medical images. officepacs is designed and marketed for soft copy reading and storage of studies produced by digital modalities. on (b)(6) 2013 a customer reported missing images from an old study. merge customer support conducted a search of the server and was unable to located the missing studies. this may potentially result in delay in patient care. there was no report of adverse patient impact. (b)(4). | 1
5467141 | due to the potential for multiple customers to have the same reported problem, a recall has been initiated.><merge cardio is an integrated cardiovascular information system classified as a picture archiving and communication system (pacs). on (b)(6) 2016 this customer reported to merge healthcare that a patient report was confirmed (signed) with the incorrect patient data. based on this allegation from that customer, merge healthcare investigated the issue and confirmed that there was a malfunction of the medical device. this malfunction if not detected, could potentially lead to an incorrect diagnosis and/or treatment of a patient based on the inaccurate information in the report. merge healthcare was able to query potentially affected customer's databases on (b)(6) 2016 to identify if there were any suspect reports stored there. the query detected a suspect report for this customer and they have been made aware of the issue. (b)(4). | 1
5468289 | due to the potential for multiple customers to have the same reported problem, a recall has been initiated.><merge cardio is an integrated cardiovascular information system classified as a picture archiving and communication system (pacs). on (b)(6) 2016, a customer reported to merge healthcare that a patient report was confirmed (signed) with the incorrect patient data (see original report 2183926-2016-00441). based on this allegation from that customer, merge healthcare investigated the issue and confirmed that there was a malfunction of the medical device. this malfunction if not detected, could potentially lead to an incorrect diagnosis and/or treatment of a patient based on the inaccurate information in the report. merge healthcare was able to query potentially affected customer's databases on (b)(6) 2016 to identify if there were any suspect reports stored there. the query detected a suspect report for this customer and they have been made aware of the issue. (b)(4). | 1
5467888 | merge cardio is an integrated cardiovascular information system classified as a picture archiving and communication system (pacs). on (b)(6) 2016, a customer reported to merge healthcare that a patient report was confirmed (signed) with the incorrect patient data (see original report 2183926-2016-00441). based on this allegation from that customer, merge healthcare investigated the issue and confirmed that there was a malfunction of the medical device. this malfunction if not detected, could potentially lead to an incorrect diagnosis and/or treatment of a patient based on the inaccurate information in the report. merge healthcare was able to query potentially affected customer's databases on (b)(4) 2016 to identify if there were any suspect reports stored there. the query detected a suspect report for this customer and they have been made aware of the issue. (b)(4).><due to the potential for multiple customers to have the same reported problem, a recall has been initiated. | 1
5467898 | due to the potential for multiple customers to have the same reported problem, a recall has been initiated.><merge cardio is an integrated cardiovascular information system classified as a picture archiving and communication system (pacs). on (b)(6) 2016, a customer reported to merge healthcare that a patient report was confirmed (signed) with the incorrect patient data (see original report 2183926-2016-00441). based on this allegation from that customer, merge healthcare investigated the issue and confirmed that there was a malfunction of the medical device. this malfunction if not detected, could potentially lead to an incorrect diagnosis and/or treatment of a patient based on the inaccurate information in the report. merge healthcare was able to query potentially affected customer's databases on (b)(6) 2016 to identify if there were any suspect reports stored there. the query detected a suspect report for this customer and they have been made aware of the issue. (b)(4). | 1
5731380 | merge cardio is a system intended to be used to acquire, store, print, transfer, and archive clinical information including images, hemodynamic studies and reports, measurements (via import from dicom structured reporting, text files or optical character recognition of measurements captured on images) and cardiology signal (waveform) data. on (b)(6) 2016, a customer contacted merge healthcare and stated that a diagnostic measurements taken by their modality was importing into the incorrect node within the clinical reporting tool. due to an incorrect values displaying in the diagnostic report, there is a potential for incorrect treatment of a patient that could result in harm. reference complaint number (b)(4).><the customer's reported problem is currently under investigation by merge healthcare. for this reason, conclusions code was used. when more information becomes available, a supplemental report will be submitted. | 1
5528660 | merge pacs is a picture archiving communication system that is intended to create and display two-dimensional and three-dimensional images of anatomy from a series of digitally acquired images. pacs is designed and marketed for soft copy reading, communication, and storage of studies produced by digital modalities. the customer called into merge support in (b)(6) 2016 because they found that some outside studies that they were bringing into their system were conflicting and merging to some current studies on their system. merge support recognized the issue and informed client they needed to create a dicom header modification rule to prefix the mrns that are coming from that facility. this would have fixed the issue. client stated that they would work with merge implementation staff to resolve the issue. the customer reported that there had been cases where rad had read images of patient a thinking they were patient b and the wrong report was attached to the wrong patient. the site was able to catch these very quickly and correct the issue. there is no information suggesting that a patient has been adversely affected by this. to make sure the incorrect merging does not happen, site is manually qcing everything that comes into the system. however, there is a risk that one could be missed or that the study was read after hours and was not qc'ed. again, there is no evidence of a specific patient being adversely effected by this. (b)(4).><the customer is working with interface and modality work list project with implementation team and full implementation of the devices is not complete. the customer elected to began reading studies for outside sites before work was done and is doing so without utilizing workflows recommended by merge heathcare. specifically use of an ipid or a prefix to ensure each study is considered unique due to multiple individualized fields. | 1
5770930 | all three images imported into merge pacs were timestamped for 10 a.m. none of these images were found caught in the system which would prevent the customer from viewing all three images at the time of review. after further review, it was found that the customer did not scroll down to view the third image and third image was later found on second review of images by a radiologist. it was this subsequent review where it was discovered that the third image revealed the patient had a fracture.><merge pacs is a picture archiving communication system that is intended to create and display two-dimensional and three dimensional images of anatomy from a series of digitally acquired images. pacs is designed and marketed for soft copy reading, communication, and storage of studies produced by digital modalities. on (b)(6) 2016 a customer contacted merge healthcare and alleged that an incident had occurred where the third image on a report was not reviewed by the radiologist. the third image showed a fracture that initially went undetected due to not being reviewed. this resulted in a delay in diagnosis and/or treatment for the patient. currently there is no indication of an adverse event to the patient. (b)(4). | 1
5583324 | merge healthcare is further investigating the customer's allegation and will determine if any corrections or corrective actions are required.><merge pacs is a picture archiving communication system that is intended to create and display two-dimensional and three-dimensional images of anatomy from a series of digitally acquired images. pacs is designed and marketed for soft copy reading, communication, and storage of studies produced by digital modalities. on (b)(6) 2016 a customer reported that while viewing images of clavicle they observed that the dicom tags were identifying the images as being the patient's left side regardless of whether it was an image of the patient's left or right clavicle. as the customer also places markers to identify the patient's side the issue is easily noticed. however, if the customer did not place the marker the dicom tags incorrectly tagging the images has the potential to lead to incorrect treatment or misdiagnosis. there was no indication from the customer that any patients were harmed due to this complaint. (b)(4).><the investigation performed found that the merge pacs software was functioning as designed. the dicom tag was coming over from the 3rd party modality incorrectly. the customer is working to resolve the issue with their modality vendor. the information is available to the customer on their modality. | 1
5758289 | annual preventative maintenance was performed, by the manufacturer, in late march 2016. reagent valve seals were replaced but not tightened to manufacturer specifications. over a period of three weeks, the reagent valve seals loosened during the normal operating runs. the loosening of the valve seals prevented reagents from completely returning to their original supply containers. this resulted in a small quantity of reagent to be held within a manifold line where it was subsequently transferred to the next adjacent reagent. during the evening run approximately a month after pm, the cross contamination of reagents was significant enough to damage approximately one-third of the patient specimens processed that evening, roughly 47 out of 140 total patient specimens. this problem was noticed on the following morning during tissue embedding and microtomy when tissue was discovered to be brittle. the pathologists had difficulty interpreting some of the h&e prepared slides and required steps to be taken to try and salvage the specimens in order to render a diagnosis. as a result of this instrument malfunction the pathologists were unable to provide a diagnosis on two patients, both of which underwent extreme procedures to obtain the specimens for diagnosis. to reiterate: approximately 47 patient specimens had suboptimal tissue processing, attempts to salvage the specimens were not successful, two patient cases had specimens processed so poorly that no diagnosis could be made. surgical procedures may need to be repeated on these two patients. manufacturer response for histology tissue processor, logos (per site reporter): manufacturer inspected the device and identified a problem of loose and leaking reagent valve seals resulting from inadequate tightening during the previous maintenance visit. device was adjusted to correct operation. manufacturer will inspect the device again in 30 days to confirm that device is still operating properly.><. | 0
6065834 | written event details were obtained from the customer but were conflicting and also indicated that there may have been more than one (1) patient affected by the problem. multiple attempts have been made by merge healthcare to reach the customer for clarification; however, the attempts have not been successful to date. for this reason, conclusion code \ (conclusion not yet available-evaluation in progress) was used. when more information becomes available, a supplemental report will be submitted.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemo monitor unexpectedly shutdown because of a faulty ups (uninterrupted power supply) during a stemi procedure and resulted in a 10-15 minute delay while the hemo monitor was rebooted. with merge hemo not presenting physiological data during treatment, there is a potential for delay in treatment that could result in harm to the patient. however, the customer reported that the procedure was completed successfully. (b)(4). | 1
5536909 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the nibp and spo2 were not functioning properly during a case that resulted in the patient coding and respiratory failure. the o2 level remained at 100%. information obtained from the customer revealed that the involved patient was considered high risk and had been previously diagnosed with several co-morbidity diseases. however, the patient presented stable vital signs so sedation was administered. shortly thereafter, the patient coded and required intubation. the patient was administered medication and was revived. due to the inability to gain vascular access, the procedure was stopped. a less-invasive procedure will be scheduled in the future. (b)(4).><remote diagnostic testing was performed on the customer's hemo software. it was found that the pdm firmware was version 2.52.02 and the site was using nelcor, so the firmware was upgraded to 2.52.04. the customer tested the software and stated that the pleth waveform and a valid reading were displayed and verified that the software functioned correctly. the customer confirmed that the site's other lab already had the correct version of firmware installed. the system was returned to service at the customer site. | 1
5686001 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that moisture appeared to be in the pb-1000. after direct communication with the customer additional information was obtained on (b)(6) 2016 that stated additional problems were experienced where artifacts" showed in a pediatric patient's ecg and full disclosure was also affected. the "artifacts" issue was not initially reported to merge healthcare. with merge hemo not functioning as intended during treatment there is a potential for incorrect treatment that results in harm to the patient. the customer stated that they were unsure of the outcome of this procedure but felt that the reported problem could affect the patient's care. (b)(4).><the complainant device has not yet been returned to merge healthcare for evaluation to determine the source and potential hazard of the reported moisture. the cause of artifacts appearing in the patient's ecg has not yet been determined; however the customer reported that the ground cable was replaced. when more information becomes available a supplemental report will be submitted." | 1
5527250 | based upon analysis of the customer's software data logs, a usb recovery thumb drive for a win 7 hemo-monitor t3600 with k600 video card was sent to the customer to upgrade their software. merge healthcare is continuing to work with the site to upgrade the affected devices.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the site was seeing diagonal lines across the waveform display that caused a 15 minute delay during an active case. additional information obtained from the customer revealed that the patient had not yet been sedated; the unwanted lines were intermittent and were present when played back in full disclosure mode. the device failed to perform an essential function which may lead to delay in diagnosis and/or treatment of patients; however, the customer reported that the procedure was completed successfully once the hemo system was rebooted. | 1
5437926 | merge hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. merge healthcare received reports of hemomonitor application crashes with the following error - hemomonitor.exe has stopped working unexpectedly" being displayed on the desktop. this error does not restart the computer since no keyboard and mouse are available the customers have had to restart via the front panel power button. this event did not result in a serious injury to the patient. this error impacts the physician's ability to continue vitals monitoring during cath procedures. (b)(4).><upon retrospective review this issue was determined to be reportable as an mdr." | 1
5728238 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that issues were experienced with a third party ffr (fractional flow reserve) device and its interaction with merge hemo. the available information indicates that a reboot of the merge hemo application occurred during an active case that may have resulted in a brief lapse of patient monitoring and may have caused the procedure to be delayed for a short period time; however, several attempts have been made to obtain additional event information from the customer with no success. therefore, merge healthcare is conservatively reporting this event. with merge hemo not presenting physiological data during treatment, there is a potential for a delay in care that results in harm to the patient. however, the customer did not report or allege that any serious injury or death occurred as a result of the reported problem. (b)(4).><the customer's reported problem is currently under investigation by merge healthcare (hemo-8794). for this reason, conclusions code (conclusion not yet available-evaluation in progress) was used. when more information becomes available, a supplemental report will be submitted. | 1
5792130 | the customer's reported problem is currently under investigation by merge healthcare. for this reason, conclusions code (conclusion not yet available-evaluation in progress) was used. when more information becomes available, a supplemental report will be submitted.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemo monitor turned black. information obtained from the customer revealed that the problem occurred in the middle of a case and the hemo monitor was rebooted. it was further stated that the patient did not have to be moved to another lab to complete the procedure. with merge hemo not presenting physiological data during treatment, there is a potential for a delay in care that results in harm to the patient. however, the customer stated that the procedure was completed successfully once the hemo monitor was rebooted. (b)(4). | 1
5543003 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that problems were experienced with the ekg output to a volcano device. information obtained from the customer revealed that the problem occurred while a transplant patient was sedated, and the procedure was delayed 20 minutes. it was also reported that the heart catheterization procedure was completed successfully; however, a flow wire study was attempted but cancelled due to the problem. the flow wire study will be completed at the patient's next visit - usually an annual visit unless problems present before then. it was further reported that the problem was not caused by merge hemo but rather a cabling issue with an external system. (b)(4).><based upon the available information, there is no indication that the reported problem was due to merge healthcare's device malfunctioning. the customer performed simulation testing and found that the problem was due to a cabling issue with an external system. (b)(4). | 1
6153720 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that during a stemi procedure (st elevation myocardial infarction) the pdm (patient data module) disconnected from the hemo monitor. subsequently, the patient was moved to another lab onsite after sedation and active monitoring had been initiated. with merge hemo not capturing physiological data, there is a potential for incorrect treatment that could result in harm to the patient. the customer reported that procedure was completed successfully once the patient was moved. (b)(4).><merge technical support remotely accessed the customer's server and found that the pdm (patient data module) was not connected. the customer was instructed to reconnect the cables to the device. device labeling, hemo v9.40 user manual, addresses such an occurrence with statements such as, inspect overall physical condition of the system components peripherals and interconnecting cables. perform any corrective actions required."" | 1
6095057 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that problems were experienced with the hemo monitor in the beginning of a procedure resulting in the hemo monitor being rebooted. once the reboot was complete, an error message was displayed that prevented the patient's vitals from being monitored. an external monitoring device was used and manual charting was completed. the customer did not report the problem to merge healthcare when it occurred but rather disclosed the occurrence on (b)(6) 2016 when information was obtained for another reported problem. the customer did not know the exact date the problem occurred but stated that it happened about 2 weeks ago." the date of (b)(6) 2016 has been used based on the approximate timeframe. with merge hemo not capturing physiological data there is a potential for delay in treatment that could cause harm to the patient. however the customer reported that the procedure was completed successfully using external monitoring. reference complaint number (b)(4).><the customer's reported problem is currently under investigation by merge healthcare. for this reason (b)(4) (conclusion not yet available-evaluation in progress) was used. when more information becomes available a supplemental report will be submitted. hardware not yet returned." | 1
5655739 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemomonitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemomonitor pc went blank during an active case and the pdm (patient data module) alarmed. this caused the procedure to be delayed 15 minutes. with merge hemo not presenting physiological data during treatment, there is a potential for a delay in care that results in harm to the patient; however, the customer reported that the patient was not harmed. the procedure was continued once the system was rebooted indicating that it was completed successfully. it was further reported that the lab was operational the remainder of the day. reference complaint number: (b)(4).><the reported problem is currently under investigation. a follow-up report will be submitted when more information becomes available. | 1
5998948 | information obtained from the customer revealed that the problem was resolved by the site due to the urgency of the patient's condition. the problem was found to be caused by a loose connection to the pdm (patient data module) at the trunk cable. the customer stated that once the cable was securely fastened, the problem was correct. site biomed replaced the trunk cable the following day. no further instances have occurred. device labeling, user manual, addresses the potential for such an occurrence in the general equipment care section with statements such as, inspect overall physical condition of the system components peripherals and interconnecting cables. perform any corrective actions required." for this reason conclusions code 18 (failure to follow instructions) was used. further there was no indication from the available information that the problem was caused from device malfunction and/or product defect.><merge hemodynamics monitors measures and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016 a customer reported to merge healthcare that patient vitals were not displaying on the hemo monitor. the problem occurred during a stemi (st elevation myocardial infarction) procedure and resulted in a 20 minute delay. subsequently the hemo monitor and client pc were rebooted by the user during the procedure. with merge hemo not capturing physiological data there is a potential for incorrect treatment that could cause harm to the patient. the customer reported that the procedure was completed successfully once the hemo monitor was rebooted. (b)(4)." | 1
6004869 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that problems were experienced with the patient data module (pdm) resulting in no output to the hemo monitor after sedation and active monitoring were initiated. external equipment was used to complete the procedure; however, the patient was moved to another lab, so that pressure values and ffr (fractional flow reserve) could be obtained. with merge hemo not capturing physiological data, there is a potential for incorrect treatment that could cause harm to the patient. the customer reported that the procedure was completed successfully using external equipment. (b)(4).><the customer reported that the procedure was completed successfully using external equipment that included a philips portable ekg machine." the customer elected to close the lab where there were alleged problems for troubleshooting activities. during this time there were no cancelled or rescheduled procedures. the customer's reported problem is currently being evaluated by merge healthcare. when more information becomes available a supplemental report will be submitted." | 1
5774351 | an internal investigation found that the issue is likely related to the customer running on the (b)(4) operating system. (b)(4) was sent to the customer on 6/07/2016. this will enable the customer to upgrade their current software. however, the site has indicated that they do not plan on upgrading until july 2016.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that unexpected lines appeared in the top corner of a patient's waveform and went diagonally across to the bottom following the ao mean. information obtained from the customer revealed that the patient's numbers and data were accurate; however, the lines appeared in full disclosure. with merge hemo displaying unexpected lines while capturing and recording physiological data, there is a potential for a delay or an incorrect treatment that may result in harm to the patient. however, the customer reported that the procedure was completed successfully, there was no delay, and no patient harm occurred. (b)(4). | 1
5536910 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemo application automatically rebooted during an acute intervention case for a heart attack patient that resulted in the procedure being delayed 30 minutes. the problem has been attributed to a sfc / scannow" system error. with merge hemo not presenting physiological data during treatment there is a potential for a delay in care that results in harm to the patient. however it was indicated that the procedure was completed successfully once the system was rebooted. reference complaint number: (b)(4).><remote diagnostic testing was performed on the customer's hemo software and a "sfc / scannow" error was found. the hemo pc was re-imaged to version 10.0 and subsequent testing confirmed that the hemo application between the client pc and hemo computer found no issues. the equipment was returned to service at the customer site." | 1
5740782 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that tabs were locked out on a patient's study and the medical staff could not edit them. these tabs included patient/study, hemodynamics, and equipment. the customer further reported that when the tabs are locked out, it prevents the medical staff from confirming the patient's study. depending on which tab is locked, it could prevent a calculation from computing correctly if the patient's data, such as height and weight, are in a locked state. it could also prevent the patient study from promptly being confirmed. this could pose a potential risk to the patient caused by a delay in treatment. however, it was confirmed that merge healthcare quickly resolved the problem by resetting the database. the customer did not allege or indicate that the patient was at risk of serious injury. to date, merge healthcare has not received allegations of a serious injury because of this reported problem. (b)(4).><the customer's reported problem was internally investigated by merge healthcare (b)(4). merge healthcare has identified this as an issue and has executed a recall to address this as a potential risk due to delay in treatment. this phantom locking issue has been corrected in hemo 10.0 software release. | 1
5976144 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that a memory error occurred on the site's server that caused the lab to be down for a week" and subsequently caused procedures to be rescheduled during that time. with the merge hemo server not functioning as intended there is a potential for delay in treatment that could result in harm to the patient. (b)(4).><merge technical support assisted the customer with preliminary troubleshooting and found that the server would not boot-up into "bios" or diagnostics. the site was then instructed to ship the faulty server to merge healthcare for repair. upon evaluating the faulty server a bad memory stick of ram was found and subsequently replaced along with the addition of 1.5 gb memory. the server passed all testing and was returned to the customer. once re-installed the server would boot-up and the proper light indicators were lit on the hard drive and led's to the mother board and power supplies; however there was no video or mouse/keyboard control. technical support instructed the user to re-seat the cpu (central processing unit). the customer then confirmed that the server was functioning as intended. (b)(4)." | 1
5750132 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on may 28, 2016, a customer reported to merge healthcare that the pressure recording was not functioning as expected during a patient study. after the user reset the hemo monitoring computer, pressures and recording of physiological data functioned as expected. information obtained from the customer revealed that there was a delay of ~5-10 minutes while the system was rebooted. when the hemo monitor is rebooted during a procedure, the patient's data is not captured. with merge hemo not capturing physiological data, there is a potential for incorrect treatment that results in harm to the patient. however, the customer reported that the procedure was completed successfully once the hemo monitor was rebooted. (b)(4).><the customer's reported problem is currently under investigation by merge healthcare. when more information becomes available, a supplemental report will be submitted. | 1
5450584 | merge hemo displays, measures, and records physiological data from a patient undergoing a cardiac catheterization procedure. during a cardiac catheterization procedure, all three monitor screens went black and the patient data module (pdm) emitted a beeping tone. the site rebooted the hemo pc multiple times in order to load the (b)(4) operating system. once the user was able to log into windows, merge healthcare's support technician remotely connected to the hemo pc and created a shortcut in the startup for the hemo monitor application in order to get the merge hemo application functioning again. once that was complete, the procedure resumed and was completed. there was a patient on the table and the procedure was delayed by 35 minutes. with merge hemo not presenting physiologic data during a cardiac catheterization procedure, there is a potential for a delay in care that results in harm to the patient. the customer did confirm that the procedure completed successfully and there was no harm to the patient. following this event, the customer requested that the operating system and merge hemo application be reinstalled on the pc. a technician was sent on site on (b)(4) 2016 to reinstall the operating system and merge hemo application. (b)(4).><the patient that was on the table when this event occurred was not harmed, however, additional sedation had to be administered due to the delay. the patient's demographs is unknown.><merge hemo displays, measures, and records physiological data from a patient undergoing a cardiac catheterization procedure. during a cardiac catheterization procedure, all three monitor screens went black and the patient data module (pdm) emitted a beeping tone. the site rebooted the hemo pc multiple times in order to load the windows 7 operating system. once the user was able to log into windows, merge healthcare's support technician remotely connected to the hemo pc and created a shortcut in the startup for the hemo monitor application in order to get the merge hemo application functioning again. once that was complete, the procedure resumed and was completed. there was a patient on the table and the procedure was delayed by 35 minutes. with merge hemo not presenting physiologic data during a cardiac catherization procedure, there is a potential for a delay in care that results in harm to the patient. the customer did confirm that the procedure completed successfully and there was no harm to the patient. following this event, the customer requested that the operating system and merge hemo application be reinstalled on the pc. a technician was sent on site on 1/22/2016 to reinstall the operating system and merge hemo application. | 1
5437755 | merge hemo monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. merge healthcare received reports of hemomonitor application crashes with the following error - hemomonitor.exe has stopped working unexpectedly" being displayed on the desktop. this error does not restart the computer since no keyboard and mouse are available the customers have had to restart via the front panel power button. this event did not result in a serious injury to the patient. this error may impact the physician's ability to continue vitals monitoring during cath procedures. (b)(4)><upon retrospective review this issue was determined to be reportable as an mdr." | 1
5700119 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to themerge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemo monitor locked up" during a case resulting in a 10 minute delay in treatment. with merge hemo not presenting physiological data during treatment there is a potential for a delay in care that results in harm to the patient. however the customer reported that the procedure was completed successfully once the hemo monitor was rebooted. (b)(4).><the customer's reported problem is currently under investigation by merge healthcare ((b)(4)). for this reason conclusions code (conclusion not yet available-evaluation in progress) was used. when more information becomes available a supplemental report will be submitted." | 1
6126698 | merge healthcare support determined that the pdm (patient data module) required replacement. a new pdm was shipped to the customer on 10/25/2016, and installed on (b)(6) 2016. the defective pdm was shipped to the 3rd party manufacturer (schiller) for investigation, and they found no issues with the pressure ports.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the hemodynamics hemo monitor pc. on october 25, 2016, a customer reported to merge healthcare that they experienced problems obtaining a patient's invasive pressure, and they were unable to zero or equalize pressure channel 1. the staff rescheduled the following day. with merge hemo not presenting physiological data during treatment, there is a potential for a delay in care that results in harm to the patient. however, the customer reported that the procedure was completed successfully, and there was no harm to the patient. (b)(4). | 1
6095050 | merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that problems were experienced in the middle of a procedure after active monitoring and sedation had been initiated. the hemo monitor display turned black so the patient was connected to an external monitoring device and manual charting was completed. the hemo monitor was subsequently rebooted successfully. with merge hemo not capturing physiological data, there is a potential for delay in treatment that could cause harm to the patient. however, the customer reported that the procedure was completed successfully using external monitoring. (b)(4).><the customer's reported problem is currently under investigation by merge healthcare. when more information becomes available, a supplemental report will be submitted. | 1
5942097 | merge technical support provided assistance for the customer to troubleshoot the reported problem. a loose vga splitter was found, so the user reconnected it and confirmed that both hemo displays functioned correctly. another staff member found that the vga cable was bad so replacement hardware was shipped to the customer on 10aug2016. the faulty vga cable has not yet been returned to merge healthcare for evaluation. once the device is received and evaluated, a supplemental report will be submitted. when more information becomes available, a supplemental report will be submitted.><merge hemodynamics monitors, measures, and records physiological data from a human patient undergoing a cardiac catheterization procedure. the system comprises the patient data module and the merge hemodynamics hemo monitor pc. the two units are connected via a serial interface. all vital parameters and evaluations are registered and calculated in the patient data module. this data is then transmitted to the merge hemodynamics hemo monitor pc via the serial interface. all data can be shown and monitored on the merge hemodynamics hemo monitor pc. on (b)(6) 2016, a customer reported to merge healthcare that the hemo monitor froze during a procedure and the patient was moved to another lab. with merge hemo not presenting physiological data during treatment, there is a potential for incorrect treatment that results in harm to the patient. however, the customer reported that the procedure was completed successfully once the patient was moved. (b)(4). | 1
5804318 | it was reported that the physician was experiencing communication issues with the programming tablet. a company representative then performed troubleshooting. it was confirmed that the tablet was not plugged into the wall. it was verified that the port on the tablet was not loose and there were no issues noted with the wand. the serial cable was then unplugged and plugged back in however the issue was not resolved until the serial cable was replaced. the physician was then provided with a new serial cable and the faulty serial cable was disposed of. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection.><. | 0
5401156 | the physician stated that he did receive the box which included the replacement product; however, he had not opened it. additionally, a returned product kit was sent to the physician so that he could returned the complaint handheld programming device to the manufacturer for analysis. attempts were made to have the product returned; however, they have been unsuccessful to date.><product analysis on the returned handheld programming computer was completed. analysis showed the allegation of the handheld computer not holding a charge was not verified. no anomalies associated with the handheld computer's performance were noted during testing using an ac adapter or the main battery with a full charge. the handheld computer performed according to functional specifications. product analysis on the returned software was completed. analysis was performed on the returned flashcard and no anomalies associated with the flashcard software or databases were identified. the flashcard and software performed according to functional specifications.><.><the handheld programming computer was received with it's associated software. device analysis is expected but has not been completed to date.><the psychiatrist reported his handheld programming computer would no longer hold a charge. it was noted the device would light up when plugged into the wall, but it would immediately shut off after being unplugged from the wall. no additional relevant information has been received to date. | 0
5618331 | (b)(4).><the cause for the reported allegation is associated with two disconnected wire connections in the returned serial cable. once the wires were soldered onto the serial cable pcb, no further anomalies were identified. note it's okay to use similar language to the pa results for these events.><it was reported that the manufacturer's consultant's tablet's programming cable was not working. the issue was confirmed to be with the cable because other backup cables worked. the serial cable was received for analysis. however analysis has not been completed to date." | 0
5957461 | (b)(4).><the tablet was received on 09/28/2016. analysis was approved on 10/17/2016. during the analysis it was verified that the usb port was damaged due to mechanical stress. no other anomalies were identified.><it was reported that a physician unplugged the usb serial cable from their tablet, and the usb port on the side of the tablet came out. there was no known mishandling of the tablet, and no patients were affected. the tablet has not been received to date. | 0
5853068 | additional information from a nurse indicated that the event occurred on (b)(6) 2016. the suspected usb cable was returned to the manufacturer on 08/19/2016. analysis is underway but it has not been completed to date.><(b)(4).><it was reported that the medical professional was having communication problems with her programming system. it was reported that she could not interrogate generators when using the programming system. troubleshooting was performed and the connection of the usb cable to the tablet was suspected to be at fault. review of manufacturing records confirmed that the tablet passed all functional tests prior to distribution. the return of the suspect usb cable is expected but it has not been received to date.><an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with two disconnected wire connections in the returned serial cable. once the wires were soldered onto the serial cable pcb, no further anomalies were identified. | 0
5526000 | .><it was reported that the physician's tablet serial cable was faulty. the physician was provided a new serial cable and the issue resolved. the faulty serial cable was received for analysis. analysis is underway, but has not been completed to date.><analysis of the serial cable was completed on 04/06/2016. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. | 0
5515349 | an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with two disconnected wire connections in the returned serial cable. once the wires were soldered onto the serial cable pcb, no further anomalies were identified.><(b)(4).><it was reported that the physician's wand was not working. via troubleshooting, it was identified that the problem was isolated to the tablet usb cable. the cable was received by the manufacturer for analysis. however, analysis has not been completed to date. | 0
6121090 | suspect device software version: (b)(4).><information was received indicating that the malfunction with the hand held device was suspected to just be due to the battery of the hand held being depleted. the sales representative up was able to troubleshoot the hand held though with his demo generator and stated that it worked fine. he also learned that the issue was that the hand held would not turn on which is why the nurse had reported it. after he spoke with the physician he stated that she just thinks that the hand held was not powering on because of a depleted battery. the hand held was not plugged in when she went to use it prior to troubleshooting by the representative, so she plugged it in, the battery charged and when troubleshooting was performed on the demo the device worked fine. the patient who was mentioned to be involved was never unable to be interrogated as she did not show for an appointment.><.><it was reported on (b)(6) 2016 that the physician?s hand held is not functioning. no troubleshooting has been performed and no additional information has been received to date. | 0
5426909 | it was reported than a medical professional could not interrogate generators when using her programming system. troubleshooting was performed and the connection of the usb cable to the tablet was suspected to be at fault. review of manufacturing records confirmed that the tablet passed all functional tests prior to distribution. the return of the suspect usb cable is expected but it has not been received to date.><the suspected usb cable was returned to the manufacturer on (b)(6) 2016. analysis is underway but it has not been completed to date.><.><an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. | 0
5502341 | (b)(4).><it was reported that the site's tablet was missing its power button, presumed lost after several months in use. the power button has been replaced and the device is able to power on as expected. review of the tablet computer's device history records confirmed all quality tests were passed prior to distribution. no additional pertinent information has been received to date.><the suspect tablet was received by the manufacturer, and product analysis was subsequently completed. the failure for the tablet to power on was confirmed, which was determined to be due to the missing power button. when the push button switch located where the power button was previously attached was pressed, the tablet turned on and completed all functional testing with no observed anomalies. product analysis concluded that the missing power button appears to be related to retention tabs and/or user error, which allowed accidental removal of the component.><it was reported that the tablet?s power button was found to be missing again. the tablet has not been received by the manufacturer to date. no additional pertinent information has been received to date.><(b)(4). | 0
5913279 | it was reported that the customer cannot interrogate or performs system diagnostic test, an error message immediately received indicating that no connection can be established. the physician replaced the 9v alkaline wand battery without success. company representative visited the customer for a surgery support and to complete the troubleshooting as the customer has only one programming system. usb cable replacement was requested as vns implant and follow-up on other patients was planned. further follow up indicated that the issue occurred again on other patient's generator. the troubleshooting was completed by the company representative as planned, using different usb cables on two programming tablets and wands. the result showed a faulty usb white cable. the customer received a replacement usb cable from company representative and the issue was resolved.><. | 0
5437201 | product analysis was completed on the returned handheld computer and flashcard on 05/26/2016. attempts were made to power on the handheld using the returned power supply, but these were unsuccessful. visual inspection of the handheld identified that the battery cover was bent due to a swollen main battery. as a result, the battery cover was unable to make good contact with the battery latch switch that was on the main board. this resulted in the handheld receiving a false open battery latch condition. after installation of a modified main battery cover, the handheld device was powered on using the returned ac power supply with no anomalies. no further anomalies were identified. no anomalies associated with the battery latch were identified. an analysis was performed on the returned flashcard. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications.><additional information was received that the physician's office was still able to use the handheld device despite the impact of the allegedly faulty battery latch on device performance. the device has still not been returned to the manufacturer to date. no additional pertinent information has been received to date.><.><the physician's handheld computer was returned to the manufacturer and is undergoing product analysis. no additional pertinent information has been received to date.><it was reported that the battery latch was not locking properly on the physician's handheld computer, leading to computer resets. the product has not been returned for analysis to date. no additional pertinent information has been received to date. | 0
6050205 | analysis of the returned flashcard was completed and no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications. an analysis of the returned handheld computer was completed and the reported allegation was not verified. no anomalies associated with the handheld were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications.><it was reported that the medical professional was having difficulties to interrogate generators when using his handheld computer. it was reported that the device which is very old, does not communicate anymore with generators. review of manufacturing records confirmed that the programming computer passed all functional tests prior to distribution. the suspected programming computer was returned to the manufacturer on 10/13/2016. analysis is underway but it has not been completed to date.><. | 0
5811838 | it was reported that a physician's tablet was acting funny". clarification was provided stating that the screen was having difficulty responding. a full reset was performed and the problem persisted. there were no signs of physical damage on the tablet. the physician's office did not know of any damage occurring on the tablet and the whole screen was reportedly affected. the tablet had been shipped to the physician on 01/14/2016. no patient's therapy was affected by this event.><(b)(4).><it was reported that the tablet would not be returned because the screen had been functioning properly. the physician was provided a replacement and the manufacturer's sales rep opted to keep the tablet for back up use. no further relevant information has been received to date." | 0
5697699 | .><it was reported by the nurse practitioner that she had an extra programming wand from a former physician. the extra programming wand was reported to not be working and that it had not affected any patient, but she would like to have a new wand. the programming wand, along with the serial cable, were returned for product analysis. it was found there were no anomalies with the returned programming wand; however, it was found that the serial cable had malfunctioned. the cause of the malfunction was due to a disconnected wire connection. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. | 0
5400189 | analysis was completed for the returned tablet on (b)(6) 2016. during the analysis it was verified that the tablet was unable to power on. the cause for the anomaly is associated with a defective main board. once the main board was replaced, no further anomalies were identified.><it was reported by a physician's office that one of their programming tablets was unable to turn on. it was reported that the issue started a week prior and there was no indication of mishandling. the tablet was charging for some time and would not turn on after that. a company representative went to visit the office and confirmed the tablets were unable to charge after 45 minutes of charging. the company representative indicated that the battery power icon was not lit and the charger was confirmed to be working by using it on the company representative's tablet. the cable connections were checked but the indicator lights did not power on. the power buttons were verified to be in the correct orientation. no additional information has been received to-date. the programming tablet has been received by the manufacturer for analysis, which is underway but has not been completed to-date.><. | 0
5663050 | (b)(4).><it was reported a company representative was unable to interrogate his demo generator with his tablet. he got an error establishing communication" message prior to interrogating the device. the tablet was not plugged into the wall. the wand was confirmed to have enough battery life. the serial cable was replaced and the communication issue resolved. the faulty serial cable was received by the manufacturer. analysis is expected but has not been completed to date.><product analysis on the returned serial cable was completed and found the cause of the reported communication issues was associated with a disconnected wire connection. once the wire was soldered onto the serial cable pcb no further anomalies were identified." | 0
5781789 | it was reported that during an implant surgery, there was a communication issue from the programming system. they could not interrogate a generator. troubleshooting was performed and the connection of the usb cable to the tablet was suspected to be at fault. when it was replaced by another one from an alternative programming system, the issue was solved. the suspected usb cable was returned to the manufacturer on 06/27/2016. analysis is underway but it has not been completed to date. review of manufacturing records confirmed that the tablet passed all functional tests prior to distribution.><(b)(4).><an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. | 0
6046846 | .><it was reported that replacing the serial adapter cable paired with the handheld computer had resolved the communication issues. the suspect cable has not been received by the manufacturer to date.><the suspect serial cable was received by the manufacturer for product analysis. product analysis was completed on the returned serial cable. visual inspection initially identified no anomalies. it was identified that the serial cable was unable to establish communication using a known good wand and handheld. the cause for the communication difficulties was associated with a broken wire connection on the handheld plug printed circuit board. this cause was most likely due to excessive wear. once the wire was soldered back onto the plug pcb, the serial cable was able to establish communication between the handheld and wand for all remaining functional tests. no further anomalies were identified.><it was reported that a physician's programming system had been unable to communicate with their patient's device for some time. no patients were affected as the site had a backup system. a company representative performed trouble-shooting. the handheld was confirmed to not being plugged into the wall at the time of use. a known good wand did not work with the physician's handheld system. therefore, it was identified either the handheld or serial cable adapter as the source of the problem it was noted that the connection of the adapter to the hhd is loose, but it is unknown if this is the cause. no additional relevant information has been received to date. the suspect product has not been received by the manufacturer to date. | 0
5361035 | .><analysis was competed on the returned handheld. no anomalies associated with the handheld were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications. analysis was competed on the returned flashcard. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications.><it was reported that a physician's programmer was not holding a charge. the programmer and software have been received and analysis is pending, but has not been completed to-date. no additional relevant information has been received to-date.><date of report, corrected data: the date that the report was received was inadvertently provided incorrectly as (b)(6) 2016 on the initial report. | 0
5665116 | .><it was reported that the physician's jornada handheld was not working. the vns representative had it plugged in for a while and it would not hold any charge. an analysis was performed on the handheld. during the analysis, it was identified that the power supply was damaged. because of the damage, the power supply was unable to power the returned handheld. no further anomalies were identified during the analysis. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. | 0
5811815 | .><it was reported that a physician?s handheld was not working. the handheld had a note that said ?not working (b)(6) 2015.? the physician could not provide any specifics on what was meant by the handheld not working. the programming wand was verified to work with the new programming tablet. there was no notification of patients? therapies being affected. it was confirmed that the physician had an additional programming computer, so no patients were affected. the physician?s handheld and programming software were both returned for product analysis to be performed. no anomalies were identified with the handheld during testing, and no anomalies were identified with the flashcard software or database during analysis. both the handheld and the programming software performed according to functional specifications. the applicable serial cable and ac adapter were not returned for analysis. no further relevant information has been received to date. | 0
5389129 | on (b)(6) 2015 it was reported that a physician's handheld would no longer hold a charge. the handheld was returned for product analysis on (b)(6) 2016. product analysis is still underway and has not yet been completed.><product analysis was completed on the handheld and flashcard on 2/2/2016. no anomalies associated with the handheld performance were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications.><. | 0
5903472 | it was reported that the physician was having communication issues with his programming system. a company representative performed troubleshooting which included confirming that the 9v battery in the wand was not depleted and that all of the cables were properly connected. the company representative exchanged various parts of the physician's system with parts from the representative's system. the issue persisted until the serial cable was replaced. the physician was provided a new serial cable and the faulty serial cable was disposed of. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. it was confirmed that no patient were affected by this issue.><. | 0
5396282 | (b)(4).><an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified.><the vns representative reported that prior to a new implant surgery, she was unable to interrogate the new generator with the surgeon?s tablet programming system. she kept receiving the ?retry? message. she then used her wand and serial adaptor cable with the physician?s tablet, and it worked fine. the representative therefore came to the conclusion that the physician?s serial adaptor cable must be defective and also she noticed that it was loose at the usb connection. a replacement cables were provided. the suspect usb serial adapter cable was received by the manufacturer for analysis. however, analysis has not been completed to date. | 0
5399578 | .><during the patient's consult, it was determined that the patient's generator was functioning properly. the issue was narrowed down to the serial cable between the physician's wand and the tablet. once the serial cable was replaced, the issue resolved. the serial cable was received on 01/28/2016. analysis confirmed that the cause for the communication issue was associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. this information was previously reported in mfr. report # 1644487-2016-00298, prior to determining that the cause of the communication issue was the serial cable.><it was reported that a physician was unable to get a reading from a patient's vns device. the patient had been punched hard in the vns when she got in between two people. the physician's notes stated that impaired function was noted" regarding the patient's vns. the physician did not get an error message and was completely unable to communicate with the patient's device. the physician's programming system had worked properly since the patient's device was unable to be interrogated so the programming system was not suspected as the cause of the issue. the physician referred the patient for replacement. no surgical intervention has occurred to date." | 0
5702659 | analysis of the returned tablet was completed and a cause for the reported failure to power on or off" allegation was identified. during the analysis it was identified that the tablet could not charge the main battery. the cause for the identified anomaly is associated with a loose battery cable to the motherboard. once the cable was reseated no further anomalies were identified during the analysis.><.><it was reported that demo tablet could not be turned on despite being charged. the tablet could switched on when it plugged into the wall. the review of manufacturing records confirmed all tests passed for the device prior to distribution. the return of the suspect tablet is expected but it has not been received to date.><the suspected programming tablet was returned to the manufacturer on 06/09/2016. the analysis is underway but it has not been completed to date." | 0
5978651 | it was reported that there were communication issues encountered with the physician's programming system. the error received was error establishing communication." troubleshooting was attempted including repositioning the wand reseating the serial cable and restarting the tablet but the issue persisted. the system's white serial adapter cable was switched with a known working cable and the issue resolved. no patients were affected. the serial cable does not require return as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection.><udi of suspect device: (b)(4)." | 0
6059763 | it was reported that physician experienced technical difficulties with the programming system. further information was received that the handheld battery did not keep the charge, besides, the handheld serial cable was suspected to be faulty. the programming wand was checked, battery changed and the wand confirmed to be working fine. additional information was received the programming computer remains plugged into the wall when not in use because the battery doesn't hold it. it was reported that the physician sometimes has to use the programming computer plugged into the mains. review of manufacturing records confirmed all tests passed for the device prior to distribution.><.><an analysis was performed on the returned flashcard and no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications. an analysis was performed on the returned handheld and the reported allegation of the handheld not keeping charge was not verified. no anomalies associated with the handheld performance were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications . | 0
5513544 | .><on 4/20/2016 the handheld and flashcard were returned for product analysis. the flashcard and software performed according to functional specifications. product analysis on the handheld revealed that the handheld was unable to advance past the screen alignment utility. the cause for the display anomaly is associated with resistance values being higher than expected in the touch screen circuitry. following replacement with a known good screen, full product functionality was restored. no further anomalies associated with the handheld performance were identified during the analysis.><on (b)(6) 2016 it was reported that the handheld is stuck on the alignment screen. the nurse has tapped the crosshairs over 20 times and it is not advancing. it was reported that the handheld would be returned for product analysis but it has not been received to date. | 0
6029169 | the tablet was received, and analysis was approved. no anomalies associated with the tablet performance were noted during testing using the ac adapter or the main battery with a full charge.><(b)(4).><it was reported that a company representative's tablet was having trouble doing heartbeat detection during at least three surgeries. she was able to interrogate and do diagnostics in each case without issue, but she could not do the heartbeat verification because it would just keep going and going without any heartbeat shown. she had not tried doing heartbeat verification outside of the operating room. in each case, she switched out her tablet for the surgeon's but kept the same serial cable and wand, and there were no issues with the heartbeat detection with the surgeon's tablet. further information was received indicating that she had tried moving to different areas of the operating room, but the heartbeat verification did not work. she did report that the serial cable felt loose inside the usb port of the tablet. the company representative attended another surgery on (b)(6) 2016, and she was unable to communicate with the patient's generator at all. she was holding the serial cable in the tablet tightly and moved the tablet away from the operating table as far as possible, but she was still unable to communicate with the generator. once the patient was in post-op, she was able to communicate with the patient's generator and perform heartbeat verification, using the same programming system, without any problems. during another surgery around (b)(6) 2016, the company representative was unable to communicate with another generator. she then switched out her serial cable and wand with known good ones, and she still was unable to communicate with the generator. then, after the surgery was complete, her programming system was able to communicate with the patient's generator in post-op without issue. a new tablet was then provided to the company representative because the tablet had been identified as the most likely cause of the communication issues. the faulty tablet has not been received to date. | 0
5905229 | (b)(4).><it was reported (b)(6) 2016 that a company representative was attempting to interrogate a patient's device but received a communication error when trying to program. he changed out the serial cable and was able to program and interrogate the patient's device. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. additional relevant information has not been received to-date. | 0
5603951 | it was reported that a physician has communication issue with two different patient's generators, a communication error message showed on the screen that request to keep the wand on the generator, this was done but without success. the 9v battery was replaced and checked ok but without resolving the issue. the usb white cable was suspected as the faulty product part. a replacement was requested and sent to the customer. the review of the manufacturing records confirmed that the motion tablet passed all functional tests prior to distribution. follow up indicated that the customer received the new usb cable. the physician stated that the new usb cable works fine. the return of the suspect usb cable is expected but it has not been received to date. no additional relevant information has been received to date.><further follow up indicated that the facility discarded the suspected device; therefore, no analysis can be performed.><(b)(4). | 0
5660623 | (b)(4).><it was reported that a provider at (b)(6) hospital was having difficulty interrogating a device. through troubleshooting, the difficulty was identified to be the result of the usb serial cable that was not working. when the serial cable was replaced, the programming system was reported to be working properly. additional information was received that no patients were affected as a different programming system was able to be used. the suspect serial cable was reported to have been discarded. | 0
5666173 | the suspected usb cable was returned to the manufacturer on 05/23/2016. an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified.><(b)(4).><it was reported that a medical professional could not interrogate a generator when using the tablet. troubleshooting was performed and the connection of the usb cable to the tablet was suspected to be at fault. it was reported that the patient's generator was interrogated and programmed successfully with a handheld computer. the return of the suspect usb cable is expected but it has not been received to date. | 0
6064930 | it was reported by the company representative that she had to replace the physician's white serial cable on (b)(6) 2016. the company representative performed the appropriate troubleshooting which isolated the physician's tablet issue to the serial cable. once the serial cable was replaced, the issue resolved. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection.><. | 0
5873653 | a physician reported her programming system would not turn on. the handheld was reportedly charged for a few hours and would not turn on at all. no patient therapy was reportedly affected. there was no physical damage noted on the handheld. the handheld was returned to the manufacturer for review. product analysis has not been completed to date.><.><the handheld and software/flashcard were returned for analysis. the handheld was received with the backup battery status indicating a low battery, which is an expected event. the handheld was charged using the returned ac power cable and verified to fully charge. the handheld was used continuously for over an hour using only the main battery after being charged. no anomalies were identified with the handheld or the returned ac adapter cable and the handheld was verified to perform according to functional specifications. an analysis was performed on the returned flashcard/software and confirmed that the flashcard and software also performed according to functional specifications. no anomalies were identified during analysis of the software. | 0
5459094 | it was reported that the physician?s tablet would not interrogate the patient?s generator. the tablet had been successfully used on other generators prior to this event. troubleshooting was attempted where the tablets, wands, and serial adapter cables were interchanged between two systems, the wand batteries were replaced, the tablets were unplugged from the wall to minimize interference, and the wands were rotated. the programming system would not still interrogate the patient?s generator. it was later reported that the neurologist found the issue to be with the two serial adapter cables. one suspect serial cable adapter is captured in this medwatch report, while the other cable was reported in medwatch report 1644487-2016-00340. follow up confirmed the programming system was working as expected with the new adapter cable, and that the patient involved was successfully interrogated with another system. the adapter was reported to have been discarded by the site. attempts for additional pertinent information have been unsuccessful to date.><. | 0
5669885 | (b)(4).><it was reported that during the initial dosing appointment the physician experienced issues interrogating the patient's newly implanted generator. a company representative visited the next day and performed troubleshooting. the issue was isolated to the serial cable. the physician was provided a new cable at this time and the programming system functioned as intended. the patient returned to the physician's office and was successfully interrogated and programmed without incident. during the troubleshooting session an additional cable was found to be malfunctioning. both serial cables were returned for product analysis however they were returned in the same packaging so it is unclear which was the original cable and which was the additional cable. product analysis confirmed that the serial cable performed to specification. no anomalies were seen with it's performance. however the other cable also underwent product analysis and it was found that this cable had a disconnected wire connection within the returned serial cable. once the wire was soldered to the pcb the serial cable then function as intended. the second cable was reported in mfr. report # 1644487-2016-01084. | 0
5697425 | it was reported that a vns patient was seen in a dosing appointment, and the physician assistant's programming system would not function as intended. troubleshooting was performed where it was determined that the usb to db9 serial data cable was at fault, but the exact troubleshooting performed was not clearly indicated. it was not stated whether or not the patient's therapy could not be adjusted as a result of the product issue. the suspect cable was reportedly discarded after the issue was identified. no additional pertinent information has been received to date.><suspect device udi: (b)(4).><it was reported that the programming system has been successfully used since the replacement of the suspect usb to db9 adapter cable. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. no additional pertinent information has been received to date. | 0
5355187 | the suspect device was received on 01/20/2016. an analysis was performed on the returned handheld and the reported allegation of hhd not holding charge was not verified. no anomalies associated with the handheld performance were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications . no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications.><.><it was reported that the physician's handheld device was not holding charge. the device was reported to be functional for about 30 minutes without being plugged in. when it is plugged in, there is interference during the programming due to electromagnetic interference (emi) and therefore has to be unplugged. the handheld is expected to be returned but has not been received to date. | 0
5883747 | it was reported that a programming system at the site was showing an error during attempted communication with a demo generator. the serial adapter cable was switched out with a known working cable and the error message resolved. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. no additional pertinent information has been received to date.><. | 0
6096348 | (b)(4).><it was reported by a company representative that she was unable to interrogate a new generator prior to a surgical case. she used the physician's tablet to successfully complete the case. upon inspection, she noticed that the usb port on her tablet was damaged. the tablet was received for analysis and was completed 11/09/2016. during the analysis, it was verified that the usb port was damaged. although the usb port was damaged, the pins were not bent. as a result, the damaged usb port was able to be used during testing. no other anomalies were identified. additional relevant information has not been received to-date. | 0
5801833 | the suspected handheld computer and flashcard were returned to the manufacturer for analysis. an analysis was performed on the returned flashcard and no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications. an analysis was performed on the returned handheld and the reported battery failure will not hold charge" allegation could not be verified. the handheld was received without a main battery. during the analysis no anomalies associated with the handheld performance were noted during testing using the ac adapter or a known good main battery with a full charge. the handheld performed according to functional specifications.><it was reported that a customer has an issue with the programming system. it was reported that the handheld computer will not hold a charge. it was also reported that they could not turn that device off. review of manufacturing records confirmed all tests passed for the device prior to distribution. the return of the suspected handheld computer is expected but it has not been received to date.><.><." | 0
5567561 | (b)(4).><it was reported on (b)(6) 2016 that the physician was unable to interrogate a patient's m106 device. the physician stated that he could palpate the generator so the issue was not suspected to be location or depth of generator. the wand battery was checked and the light stated on for more than 25 seconds. the tablet was not plugged into the wall. the usb connection was checked and fit snuggly. he tried interrogating while holding the usb plugged in as well. the wand was 3-4 feet away from the tablet and he tried rotating the wand handle 45 degrees. he also tried to interrogate the patient in the reception room in case of electro-magnetic interference. the patient was recently implanted with a m106 in (b)(6) 2015. the physician stated he last successfully interrogated a patient about 3 weeks ago. troubleshooting was performed on (b)(6) 2016 and it was found that the wand to usb connection adapter was the reason for the failure to interrogate. once the adapter cable was replaced the device worked correctly. the patient's generator was also able to be interrogated successfully after the cable was replaced. the cable was thrown away and i snot available for analysis. | 0
5865230 | (b)(4).><it was reported that the programming system was inconsistently communicating with generators. patient therapy was not affected by these intermittent issues. a company representative performed troubleshooting and determined that the serial cable was loose. the serial cable was replaced with a new cable. the programming system then interrogated and performed as normal. the faulty serial cable was discarded following the troubleshooting. therefore product analysis cannot be completed. | 0
5771811 | (b)(4).><it was reported that a physician's programming system was experiencing communication issues. the battery on the programming wand was verified as adequate and the connections were verified as secure. when the serial cable was changed out, the communication issues stopped. the physician had another programming system so no patients were affected. the faulty cable was cut to prevent inadvertent use of a known defective cable. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. | 0
6002052 | on (b)(6) 2016, a physician's tablet was not charging despite being connected to a good power source. when plugged in, the tablet's charging indicator did not light up. the power cable of the programming system was able to charge other tablets, isolating the tablet as the issue. the tablet still did not work when plugged in and trauma to the device was denied. no additional relevant information has been received to date. the tablet has not been received by the manufacturer to date.><suspect device udi number: (b)(4).><the physician's tablet and ac adapter was received for product analysis. it was found that the tablet's main battery was depleted; however, after the tablet was charged for at least three hours with the provided ac adapter, the tablet stayed powered on continuously for an hour or more. no anomalies were observed. both the amber and green battery indicator was observed while the tablet was charging. no additional relevant information has been received to date. | 0
5341422 | on (b)(6) 2016 product analysis was completed on the handheld and flashcard. no anomalies associated with the handheld were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications. additionally the handheld was received with an x50 serial cable. the reason for the serial cable return is unknown. an analysis was performed on the returned serial cable and no anomalies associated with its performance were noted during testing using a known good x50 handheld. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications.><it was reported that the physician's handheld does not turn on and when it did it did not hold a charge. the handheld, flashcard, and power cable were returned for product analysis on (b)(4) 2015. product analysis is still underway and has not yet been completed.><. | 0
5762441 | (b)(4)><it was reported that the physician was unable to interrogate patient's generator, but that a backup programming system was available so no patient was affected. a company representative performed troubleshooting on the physician's device by swapping a known good wand and serial cable with the physician's wand and serial cable on the known good tablet; however, a demo generator was unable to be interrogated. the physician's tablet was then used with a known good serial cable and wand and the generator was able to be interrogated. the physician's tablet was used with the physician's serial cable and a known working wand and the generator could not be interrogated. the physician's wand was then used with the physician's tablet and a new serial cable which resolved the issue. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. | 0
5541499 | .><the operating room specialist reported that during a surgery case that there were communication issues. the case was ultimately able to be completed with no issues. she confirmed that when she switched the tablets and used her wand and serial cable, the issue was resolved. the connections and 9v battery were said to have been checked and switching the tablets. later, she reported that despite trying to identify the cause, she was not able to replicate the communication issues. it was thought that it was likely emi since she checked the connections and wand battery prior to the call. the operating room specialist went to the hospital again on (b)(6) 2016 and found that the programming system was initially was working properly. the operating room specialist was using the physician's serial cable on her programming system, which initially was working properly, but then stopped working. she was able to confirm that the physician's serial cable was having the issue, and was thought to be an intermittent issue. after replacement of the serial cable, the or specialist was able to successfully interrogate with her tablet, indicating the issue was related to the serial cable. the suspect serial cable was received by the manufacturer for analysis. however, analysis has not been completed to date.><the initial report inadvertently reported the conclusion of the troubleshooting incorrectly. the issue was isolated to the company representative's cable, rather than the physician's.><based on the troubleshooting performed, the or specialist was able to confirm that her serial cable was having the issue, and was thought to be an intermittent issue. an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified. | 0
5791165 | .><it was reported that the handheld programmer presented with a clear memory message". a company representative will perform troubleshooting on the handheld programmer; however this has not occurred to date. no additional relevant information has been received to date.><a company representative went to the physician's office and performed troubleshooting. it was reported that the staff could not remember what they were doing with the handheld programmer when it presented with the "clear memory" message. the office reported that they do not have many vns patients and do not use the programming system very often. they did not report that any patients were affected by the issue. during troubleshooting the company representative performed a hard reset on the handheld programmer and it started up as normal. a vns patient was then brought into the clinic and the generator was successfully interrogated using the re-started handheld programmer." | 0
6140943 | per a livanova representative, the hand held computer was received with the battery cover from the physician, but the representative did not send the cover in for analysis. the battery on the hand held device was inserted backwards upon receipt from the physician by the livanova representative. it is believed that the battery being inserted backwards is likely the reason the hand held computer would not power on.><the physician gave a livanova therapeutic consultant a hand held programming device (hhd) stating that it was not working. the therapeutic consultant was not able to get the hhd to turn on, even after charging it over night. the hhd was received on 11/17/2016. analysis is underway, but hasn't been completed to date. no additional relevant information has been received to date.><it was discovered that a company therapeutic consultant returned the handheld programming computer without a battery cover. an analysis was performed on the returned handheld programming computer. during the analysis it was identified that the handheld was returned without a battery cover. as a result, the handheld would not power on. also, during the analysis it was identified that the main battery was installed backwards. as a result, the handheld was unable to charge and receive power from the battery. once the battery was removed and installed correctly, no further anomalies were identified during the analysis. analysis was performed on the flashcard software. no anomalies associated with flashcard software or databases were identified during the flashcard analysis. the flashcard and software performed according to functional specifications. no additional relevant information has been received to date.><.><previous supplemental mdr #1 inadvertently did not include information received by a company representative who stated that the battery cover was received with the hand held programming computer from the physician, but was not sent in for analysis, and the battery was backwards upon receipt from the physician. | 0
5845991 | .><an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified.><it was reported that the programming system is not working due to faulty usb cable. a replacement usb cable was used and the programming system functioned fine. review of manufacturing records confirmed all tests passed for the device prior to distribution. the suspect usb cable was received by the manufacturer. analysis of the device is underway, but it has not been completed to date. | 0
5434085 | .><it was reported the physician was having issues with his vns programming system and it was not working. attempts for additional information have been unsuccessful to date. | 0
5404675 | (b)(4).><the usb serial cable was received for analysis. an analysis was performed on the returned usb to db9 cable and the reported allegation was verified. the cause for the reported allegation is associated with a disconnected wire connection in the returned serial cable. once the wire was soldered onto the serial cable pcb, no further anomalies were identified.><the technical support specialist reported on (b)(6) 2016 that the a new m106 generator was unable to be interrogated during replacement surgery. the data received light was not illuminating during the interrogation. the company or specialist was able to program patient initials with his wand and tablet through the packaging before it entered the sterile field. once in the sterile field and after the new leads were placed, the m106 was unable to establish communications after about 15 attempts. the specialist removed the usb cord, reinserted, and waited 15 seconds before reinterrogating again. she then disconnected and reconnected the wand and usb cable which continued to not resolve the issue. she then replaced the 9v battery and tried it all again. she held down reset buttons and the wand light illuminated for 20 seconds. the tablet was turned off and restarted. the hospital's wand with the specialist's usb cable was then used and the same troubleshooting steps were performed but communication was still unsuccessful. the surgeon decided to close the patent with the new leads and generator. however, diagnostics and heartbeat detection were unable to be tested as a result of the communication issues. the issue was suspected to be related to the tablet usb cable or tablet connection. a replacement usb cable was provided to the specialist. once the cable was replaced, the programming system has been functioning since the case on (b)(6) 2016. it was reported that potential sources of emi were turned off as much as possible in the or to mitigate emi communication issues. communication with the generator was attempt in different areas of the sterile field to try and move away from sources of emi. the usb serial cable was received for analysis. however, analysis has not been completed to date. | 0
5793680 | .><further follow up with company representative confirmed that the programming system works fine with the new usb cable (black).><it was reported that a medical professional was having issues with his programming system. when using the tablet and a wand to interrogate a device, the communication fails in 70% of the cases. it was reported that the issue occurred on more than one patient. the wand battery was replaced with a new one without resolving the issue. the physician reported that no interference is suspected to be the cause of those communication issues. the connection of the usb cable to the tablet was suspected to be at fault. a new usb cable was sent to the physician as replacement. review of manufacturing records confirmed that the tablet passed all functional tests prior to distribution. additional information was received from the physician, indicating that they tried with the new usb cable but the issues persisted. in around 50% of the communication attempts they fail. it was reported that the communication is established most of the time after many attempts. no additional information was provided to date.><device evaluated by mfr; corrected data: the previously submitted mdr inadvertently provided the wrong suspect device for the event. | 0
5579736 | it was reported that the patient experienced erratic stimulation. patient's off time was recently programmed to 5 minutes but the patient was able to feel stimulation every 3 minutes and sometimes not feel stimulation at all. patient's spoke was concerned of increased seizures as a result. at that time, neurologist's programming system was reported to be not working. attempts for additional relevant information regarding the programming issue were unsuccessful.><. | 0
5888445 | it was reported that customer had communication issue during vns implant. the programming system was unplugged from the outlet wall, the connection cables were checked and were ok, the 9 v battery was replaced but the issue persisted. the programming system was not able to communicate with other generator as well. an usb white cable issue was suspected, therefore a replacement was requested. follow up with the physician indicated that the new usb cable was received and the communication issues were resolved. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection.><. | 0
5873751 | (b)(4).><it was reported that the site was not able to use their programmer due to an allegedly malfunctioning serial adapter cable. follow up troubleshooting was performed. the battery in the programming wand was checked, connections were confirmed to be firmly seated, and parts of the programming system were swapped with known working components. the problem was narrowed down to the serial adapter cable, and after switching that device with a known working cable, the programming system successfully interrogated a demo generator. the serial cable does not require return, as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection. | 0
5885916 | (b)(4).><a physician was unable to interrogate a specific patient. the physician's device was later able to interrogate other patients on the same day. the patient reportedly could still feel stimulation, so no device issue was expected with the patient's generator. the manufacturer's sales representative performed troubleshooting on the programming system. the sales representative was unable to interrogate known good demo generators with the programming system. when the tablet was used with a known good serial adapter cable, interrogation was successful. when the physician's serial cable was used with a known good tablet and wand, an error message stating there is an error establishing communication with the generator. please try repositioning the programming wand." was noted. the troubleshooting confirmed that the issue was related to the serial adapter cable. the serial cable was subsequently discarded. the serial cable does not require return as the failure mode is understood to be a failure of the serial cable associated with a disconnected wire connection." | 0
5867478 | review of the patient's in-house programming history identified a likely programming anomaly sometime between (b)(6) 2012 and (b)(6) 2013 that turned the off time from 5min to 60sec unintentionally. there were no noted diagnostic tests between those dates. a final interrogation was performed on (b)(6) 2012 which confirmed the intended settings. upon initial interrogation on (b)(6) 2013, the off time was at 60min.><. | 0
5593317 | during review of the in-house programming history database, it was observed that a faulted system diagnostic test occurred on (b)(6) 2014. a faulted system diagnostics occurred and a final interrogation was not performed. on the next recorded visit dated (b)(6) 2014, the device was interrogated and the was at inefficacious settings. during the visit, the output current and the magnet output current were changed, however the 'off' time remained at 60 min off. the 'off' time was changed to efficacious settings on (b)(6) 2014.><. | 0
6115096 | .><programming data was reviewed, and the first diagnostic test was not in the history. however, the most recent diagnostics from the previous office visit were the same as what the nurse reported the diagnostics that didn't show up in the programming history. based on the available history, it appears that a second programming system may have been used, or there is a possibility that the physician did not run diagnostics and transposed the most recent diagnostic results from the previous visit into the notes.><serial # and udi #, corrected data: the initial report inadvertently listed the incorrect product information for the suspect device. additional manufacturer narrative, corrected data: the initial report inadvertently listed the incorrect software version of the suspect device. the correct software version is 11.0.4.><it was reported that a nurse's programming tablet history was not showing a diagnostic test performed on a patient's vns system. the patient's device was able to be programmed and interrogated. the physician received an error message while trying to perform diagnostics, but she was able to adjust the usb serial cable and perform diagnostics successfully. the nurse reported that she was unable to see the first diagnostic test in the programming history on the tablet upon later review. however, diagnostic tests performed later that same day were present in the programming history on the tablet. no further relevant information has been received to date. | 0
5439596 | during review of the in-house programming/diagnostic history database, it was observed that a partial programming event occurred on (b)(6) 2008 during which the device's output current was reprogrammed from 0.5ma to 0.0ma. subsequent interrogations and a programming change are seen in the programming history and the output current was not changed from 0.0ma indicating that the provider may have intended to change the output to 0.0ma on (b)(6) 2008. the magnet output current was reprogrammed to 0.0ma on (b)(6) 2009 for an unknown reason. attempts to identify the operation of the handheld programmer on (b)(6) 2008 have been unsuccessful. the operator of the handheld programmer on (b)(6) 2009 is no longer at the facility. the patient's current providers did not have information regarding the partial programming event or the magnet disablement.><. | 0
5701554 | on 5/17/16 during review of the patient's programming history, it was found that on (b)(6) 2015, the generator was interrogated and a system diagnostics was performed which changed the patient's settings. the patient's device was re-interrogated but only the normal mode output current was corrected. it was also observed that on (b)(6) 2013 it appears that there was a settings change from an interruption in a system diagnostics test but that there is missing data in the programming history. the patient left the visit at faulted system diagnostics test settings.><. | 0
5513078 | a computer software error was observed from a review of the manufacturer's programming history database. on (b)(6) 2014 a system diagnostic was performed, however a final interrogation was not performed. on the next recorded visit dated (b)(6) 2014 the first interrogation showed the settings were indicative of an interrupted communication during the previous diagnostic that was performed. additional relevant information has not been received to-date.><follow-up from the treating physician revealed that at the patient's last checkup the settings were turned on as intended.><. | 0
5807233 | during review of the in-house programming/diagnostic history database, it was observed that on office visit on (b)(6) 2014 the patient's settings were different than what were programmed at the previous office visit on (b)(6) 2014. the settings found were indicative of a faulted diagnostic test; however, review of system diagnostic testing from office visit on (b)(6) 2014 was within normal limits. the device was interrogated prior to the patient leaving the office on (b)(6) 2014 as recommended by device manufacturer to ensure the device is at the correct settings; therefore, the change in settings occurred after the (b)(6) 2014 visit. the physician indicated that the patient was not seen in between these two dates; however, it is unknown if the patient was seen by another physician during this time. no patient adverse events were reported. no additional relevant information has been received to date.><. | 0
5492011 | on (b)(6) 2016, the vns patient's wife reported that the patient's device off time was changing to 60 minutes and had to be restarted three times in the last two weeks." the patient subsequently began experiencing an increase in depression so the physician prescribed six-weeks of tms treatment. additional information was received stating that the physician tested the patient's device on (b)(6) 2016 and diagnostic results showed normal device function. the next day the patient reported that he was no longer able to perceive device stimulation on-times. follow-up revealed that the patient underwent tms treatment on (b)(6) 2016. the patient's device was tested following the procedure and both normal mode and system diagnostic results showed normal device function. the company representative attended the patient's appointment on (b)(6) 2016 and reported that the device was programmed to 1.25/20/500/30/5.00. the device was tested and again and system diagnostic results again showed normal device function. follow-up with the physician revealed that the patient's device was found to be programmed to a 60 minute off time when the patient returned to the clinic for tms treatment.><." | 0
5375832 | .><during an internal review of programming and diagnostic history, it was identified, during an office visit on (b)(6) 2014, that an interrupted system diagnostic test occurred that caused an unintended change in device settings. it's unknown when that change in device settings occurred. the settings were corrected on (b)(6) 2014 when the event was found. diagnostics history was reviewed and no record was found of an interrupted system diagnostic test that might have caused the change in device settings. no patient adverse events were reported. | 0
5366189 | analysis of the returned flashcard / software was completed and the reported allegation was verified. during the analysis it was identified that the 2577 installation folder was missing. as a result, the vns software could not be installed into the handheld using the returned flashcard. additionally, it appears that the contents of the vns81.arm.cab (located in the 2577 folder) had been extracted to the root directory of the flashcard. in addition to the .cab files, (b)(4). the origin of the files is unknown. no other anomalies were identified. analysis of the returned handheld computer was completed and no anomalies associated with the handheld performance were noted during testing using the ac adapter or the main battery with a full charge. the handheld performed according to functional specifications.><it was reported that a medical professional was having difficulties with her programming system. the flashcard of a handheld computer was no longer supporting the program. it was reported that the handheld computer could not perform any task with its flashcard. by using others flashcards, it was functioning perfectly. troubleshooting and hard reset were performed without success. review of manufacturing records confirmed that the handheld computer passed all functional tests prior to distribution. the handheld computer and the flashcard were returned to the manufacturer on (b)(6) 2016. analysis are underway but they have not been completed to date.><. | 0
5515264 | (b)(4).><it was identified via manufacturer's review of the in-house programming history that a programming anomaly occurred on (b)(6) 2015. on (b)(6) 2015, the device was interrogated, a system diagnostic was performed, and a final interrogation was not performed. on the next recorded visit dated on (b)(6) 2015, the device was interrogated and the settings were found to be at unintended parameters indicative of a faulted system diagnostic. the settings were not corrected on (b)(6) 2015. | 0
5505045 | a nurse reported on (b)(6) 2016 that the physician's handheld experienced a frozen screen after interrogating a patient's device. the company representative reported that she was at the office on (b)(6) 2016, tested the handheld by interrogating with two different generators and it worked fine. she was performing training with a physician. however, the reporter called again on (b)(6) 2016 reporting that the physician had to go through many steps" to properly interrogate the patient who had been seen in clinic lately but the patient was eventually interrogated. she describes the issue as "being able to interrogate device but it freezes." she did not have a specific message or screen which was being displayed. she said the physician told her "it should show treatment parameters but it does not" and was not sure if she was providing us with the right information. since then however this device is still locking up upon interrogation. she requested that the company representative re-visit and acknowledged that the programmer worked properly during her visit on (b)(6) 2016. based on the company representative's conversation with the site the issue was rising from interrogating a patient directly from the parameter screen of a different patient rather than exiting to the main menu. the company representative went to the site and found the programming settings for the first patient that was seen and therefore she believes that it was a cross-programming event between the two generators. if they were interrogating a patient directly from the parameter screen from a different patient (rather than exiting to the main menu) this could result in a cross-programming event if the model of the generators was the same (and were m100-m102/m102r). no additional relevant information has been received to date.><.><.><date received by manufacturer (mo/day/yr) corrected data: supplemental mdr 1 inadvertently had the wrong date received date. correct date for supplemental mdr 1 is 05/13/2016." | 0
5710346 | the pyxis es platform was implemented in our 6th hospital health system last year. the application is unstable and has resulted in hours/days of extended downtime, removing the electronic safety net that our clinicians rely on that scans patient and specific medication. this is a serious safety issue. we have learned that other health systems have had the same problems and carefusion is yet to give this sufficient attention. | 1
5597092 | problem: no patient involvement, problem identified during internal testing. potential hazard(s): when using enhanced ct format dicom data, there may be: incorrect orientation of displayed image data indicating incorrect patient side (left, right, superior, inferior, anterior or posterior) (sidedness). inaccurate measurement.><no specific medical device failure. no patients involved. issue identified during echopixel internal testing. health hazard evaluation completed by quality management.. all current users of product have been contacted, provided information on the issue. all current users have indicated that none use the stated enhanced multi-frame ct formatted dicom data, therefore no risk to patients corrective action process in place. no specific device failure involved. | 1
5437766 | upon retrospective review, this issue was determined to be an mdr.><merge unity pacs is a medical image and information management system that allows viewing, selection, processing, printing, telecommunications, and media interchange of medical images from a variety of diagnostic imaging systems. diagnostic images services reported that montage was displaying an incorrect tomo slice. the first tomo image added to the montage is changed when adding a second slice. (b)(6) radiology also reported when adding multiple tomo images to montage, the images selected are not the ones that insert to the montage. a physician encountered this issue when presenting to other physicians. the montage that was created with tomo images did not have the intended tomo images. defect was found in unity pacs 11.0 code. code was not handling montages correctly when closing exams. (b)(4) | 1
5985838 | ><report # 2915056-2016-00020 is a literature report from (b)(6) concerning iodine-125 (i-125) seed migration without associated adverse event (ae). this report involves a (b)(6) year-old male (4 of 6 patients) implanted with 86 i-125 free seeds (oncoseed) for the treatment of stage t1-t2 prostate cancer. concurrent medical conditions/past medical history included stage t1-t2 prostate cancer. concomitant medications were not reported. on an unknown date, the patient underwent a trans perineal interstitial prostate brachytherapy treatment of stage t1-t2 prostate cancer with oncoseed at a total dose ranging from 110 to 145 gy; lot number was not provided. one month after brachytherapy, an abdominal ct did not detect seed migration. however, one year after brachytherapy, abdominal ct revealed seeds lodged in the right atrium and right inguinal canal. treatment details were not provided. there was no associated adverse event (ae). the outcome had not resolved. article summary: seed migration is a well-recognized event that occurs after trans perineal interstitial prostate brachytherapy, and it is observed more often with loose seeds than with linked seeds. it is well known that the most frequent site of seed migration is the chest. seed migration can result in a decreased dose being administered to the prostate and has serious effects on various organs, such as the heart. it is also associated with the development of secondary malignancies and inflammation. the study was conducted to evaluate the incidence and predictors of seed migration after trans perineal interstitial prostate brachytherapy using i-125 free seeds. from mar-2007 to mar-2011, 121 patients with stage t1-t2 prostate cancer underwent trans perineal interstitial prostate brachytherapy alone, or in combination with hormal" therapy or external beam radiotherapy. all patients received oncoseed. pre-planning was performed 3 weeks prior to implantation and the implants were inserted using the standard parallel needle insertion technique. all patients underwent a series of radiographs [chest radiography kidney-ureter-bladder (kub) radiography and a ct scan] to assess whether seed migration had occurred on postoperative days 1 and 30 and at 12 months. results: seed migration occurred in 31 (25.6%) of 121 patients. in 16 (51.6%) of the 31 patients only a single seed migrated. in the remaining 15 (48.4%) patients two to five seeds migrated. a total of 51 of 7 883 (0.65%) implanted seeds migrated. migration was detected on postoperative day 1 in 16 patients day 30 in 13 patients and at 12 months in 4 patients (migration occurred at different times in 2 patients). the migrated seeds were found in the lungs pelvis heart mediastinum kidney inguinal canal liver and sacrum. the number of needles was a statistically significant factor in seed migration. there was no significant difference in post-planning dose delivered to the prostate between the patients that displayed seed migration and those that did not. conclusion: the authors reported that seed migration occurred in 25.6% of patients and it was detected within 1 month of the insertion procedure in many cases (88% of patients). the number of migrated seeds ranged from one to five (mean 1.6). seeds that become lodged in the pelvis on day 1 might re-migrate to other organs. no difference was noted in the quality of dosimetry. according to the authors although follow-up may have been inadequate to determine whether toxicity will result in vessels secondary to seed embolization no adverse effects associated with seed migration were detected. the number of needles was a significant predictor of seed migration. the authors concluded that the results of this study suggest the number of needles might be used to predict the risk of seed migration. miyazawa k matoba m minato h et al. seed migration after transperineal interstitial prostate brachytherapy with i-125 free seeds: analysis of its incidence and risk factors. japan radiological society. (2012); 30: 635-641." | 0
5985801 | ><device not returned to mfr. | 0
5920023 | the customer reported that the rns (remote network station) screen was dark this morning and now it does not appear to be on. if the screen is powered off, then back on, it will come on for a couple of seconds.><the customer reported that the rns (remote network station) screen was dark this morning and now it does not appear to be on. if the screen is powered off, then back on, it will come on for a couple of seconds. the device was returned to nihon kohden, evaluated, and the reported issue was confirmed. the device was returned to the vendor. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. | 1
5556826 | the customer reported that the rns (remote network station) shutdown by itself and won't turn back on.><the customer reported that the rns (remote network station) shutdown by itself and won't turn back on. customer tried different power cords and flipping the power switch on the back and front which did not resolve the issue. the device was returned to nihon kohden, however it has not yet been evaluated. the unit was evaluated and the reported problem was duplicated with main switch in on position and known good power cord and ac socket. the customer was sent an exchange unit. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the rns (remote network station) shutdown by itself and won't turn back on. customer tried different power cords and flipping the power switch on the back and front which did not resolve the issue. the device was returned to nihon kohden, however it has not yet been evaluated. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 when additional information becomes available. | 1
5432813 | the customer reported that he cannot hear any sound from the rns (remote network system or remote monitoring system). while on the phone with nihon kohden, the customer tried power cycling, reinstalling the driver, and reinstalling the ports. the customer was then asked to try to hear the sound through an external speaker. customer was told to call in if that did not work and a repair/exchange would be provided. no other information is available at this time. the device involved in this event has not been returned to date to allow for an evaluation / analysis to be performed. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 when additional information becomes available.><the customer reported that he cannot hear any sound from the rns (remote network system or remote monitoring system). while on the phone with nihon kohden, the customer tried power cycling, reinstalling the driver, and reinstalling the ports. the customer was then asked to try to hear the sound through an external speaker. customer was told to call in if that did not work and a repair/exchange would be provided. no other information is available at this time.><manufacturer narrative: the customer reported that he cannot hear any sound from the rns (remote network system or remote monitoring system). while on the phone with nihon kohden, the customer tried power cycling, reinstalling the driver, and reinstalling the ports. the customer was then asked to try to hear the sound through an external speaker. customer was told to call in if that did not work and a repair/exchange would be provided. the device involved in this event was not returned to allow for an evaluation/ analysis to be performed. due to the age of this complaint additional information necessary to conduct an investigation is not readily available. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned. | 1
5339721 | the customer reported that the monitor on the rns (remote network system or remote monitoring system) has a dimmed backlight and you can barely see the display.><nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 when additional information becomes available. | 1
5852534 | the customer reported that the rns (remote network station) boots up to dos mode and cannot be started to normal mode.><the customer reported that the rns (remote network station) boots up to dos mode and cannot be started to normal mode. the device was returned to nihon kohden, evaluated, and the reported issue was not confirmed. completed all steps on the maintenance check sheet of the user manual and performed extended test with cns and bedside monitor. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. | 1
5340984 | ><while patient was on ventilator the ventilator unit will not charge battery for emergency power failure. patient was transferred to another ventilator. vent was sent in for repair. this ventilator was at the ge repair center for almost 2 months and out of service. pulled out of service. | 0
5896010 | .><during the robotic procedures the cables on the instruments break. this has been an issue over this past year. in the past few months we have had an increase in the amount of breakages. this has occurred across many different procedures and many different surgeons. (each of the instruments involved were at various stages of their 10 usages per device) | 0
5366191 | i just want my personal experience with this company on the record. you are free to use it anyway you wish including my name and complaint. to put things into perspective i need to go back a bit. for personal health purposes my doctor recommended a humidifier for my bedroom. i purchased a holmes humidifier and was pleased with the product until it broke down in less than a year. i wrote (b)(6). after several e-mail exchanges they acknowledged their faulty product by sending me a replacement bionaire humidifier a month or so after my initial complaint. in the fall of this year ((b)(6)) i again contacted (b)(6) (via their bionaire customer service) and reported that the replacement they had sent me less than a year after my first humidifier broke down, was causing problems which i identified as fan or electric motor. i had obvious concerns with a possible electrical fire, i sent them all the info they requested and that it could glean from the product including model, serial, electric prong numbers. i stressed that it was the replacement that they had sent me that broke down and was causing fire concerns to me. after more delays and more questions, they abruptly said we will not replace our product." i contacted them several more times including asking to speak with a supervisor. the supervisor stated that their product warranty was for one year on the original product. they further mentioned that i would need to provide a sales receipt. i pointed out that it was the product they had sent as a replacement that was now broken. obviously i would not have a receipt for the replacement. but their records should show when they had sent the replacement. i felt sure it had been less than a year. the supervisor replied that (b)(6) policy applied only to the original purchase. end of story. i can't afford to get into a legal battle with a multi-national corporation. but i hope others will benefit from my personal story. sorry for the long account of my problems. but it looks like the big corp win again. purchase date: (b)(6) 2015; this date is an estimate. (b)(6)." | 0
5563024 | ><while oral care was being rendered to the patient using a toothette suction swab, the patient bit down on the plastic part of the toothette and broke it in half, retaining that piece in his mouth. numerous, various attempts to remove the (3.5 cm) distal end of the toothette were unsuccessful. the patient had increased dyspnea and a moist cough, but no stridor. the md examined the patient and via direct laryngoscopy visualized the sponge stick piece in the lower hypopharynx overlying the vocal cords, partially obstructing the airway. the piece was removed intact using mcgill forceps and the oropharynx was suctioned thoroughly. cxr post-event confirmed removal of the foreign body. | 0
5998182 | the customer of a versacell x3 sample management system reported that three (3) sample tubes were dropped while moving between the track and the versacell x3 sample management system. the customer stated that some sample fluid (urine sample) was spilled. there was no cross contamination of samples. the customer was unable to determine if the sample tubes were going to the track or coming from the track. it is unknown if the samples had been aspirated or not. there are no known reports of patient intervention or adverse health consequences due to the dropped sample tubes.><a siemens customer care center (ccc) specialist was contacted by the customer. the ccc confirmed the customer cleaned the fluid with proper personal protective equipment (ppe). a siemens customer service engineer (cse) specialist was dispatched to the customer's site. the cse found the gripper was not in alignment with the tubes on the track. the cse realigned the instrument to the track and verified the proper alignment. the cse readjusted the placement with an immulite instrument and verified the tubes were properly picked up. the cause of the sample tubes being dropped was from a misaligned gripper. the gripper was misaligned due to being bumped into or moved out of position. the instrument is performing according to specifications. no further evaluation of the device is required.><the initial mdr 2247117-2016-00064 was filed on october 04, 2016. additional information (10/26/2016): headquarters support center (hsc) reviewed the event. hsc concluded the patient samples results were obtained and the tubes were dropped going back to the track after being tested. the instrument is performing according to specifications. no further evaluation of the device is required. | 0
5960605 | the reported issue has been identified as a software problem. spacelabs has initiated a field corrective action and notified the fda (b)(4) district office of this action on august 25th, 2016. (recall report number: 3010157426-08/25/16-003-c) we also began notifying customers of this activity with a letter dated august 25th, 2016. this investigation is considered complete and the issue closed.><spacelabs received a report that on (b)(6) 2016 at 18:51 a telemetry patient experienced an episode of asystole that did not generate an audible or flashing visual alarm at the central station. the alarm was captured in the retrospective database. no one was injured as a result of this event. | 1
5958279 | spacelabs received a report that on (b)(6) 2016 at 00:39 am, a telemetry patient experienced an episode of psvt that did not generate an audible or flashing visual alarm at the central station. the alarm was captured in the retrospective database. no one was injured as a result of this event.><spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><the investigator is unable to confirm whether an appropriate audible alarm notification occurred at the reported event time through the database. per the complaint description, the appropriate visual alarm notifications were provided. audible tones were confirmed to have generated appropriately during the same time period on other telemetry channels. we know of no reason that audible alarm tone would not have been present at the time of the complaint episode as these indicators operated according to specifications when tested by a spacelabs? field service engineer. this investigation is considered complete and the issue closed. | 1
5972189 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><spacelabs received a report that on (b)(6) 2016 the xhibit central monitor system displayed an offline" message in the patient zones for telemetry beds. no injury was reported as a result of this event." | 1
6155748 | spacelabs has launched an investigation and will file a supplemental report once the investigation is complete.><spacelabs received a report that on (b)(6) 2016 at 7:50pm a telemetry waveform displayed as a dotted waveform then dropped off the display of the xhibit central station. | 1
5870758 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><the reported issue has been identified as a software problem. spacelabs has initiated a field corrective action and notified the fda (b)(4) district office of this action on august 25th, 2016. (recall report number: 3010157426-08/25/16-003-c) we also began notifying customers of this activity with a letter dated august 25th, 2016. this investigation is considered complete and the issue closed.><spacelabs received a report that on (b)(6) 2016 telemetry waveforms disappeared off the central monitor. the issue was resolved by restarting the system. no injury was reported as a result of this. | 1
5430837 | spacelabs has launched an investigation into this event and will file a supplement report once the investigation is complete.><the issue has been traced to the texas instrument (ti) msp430 microcontroller used on the model 96280 telemetry receiver quad receiver card (qrc) pcba that has an internal hardware defect that affects the proper clocking of digital data communications on some qrc units. the presence of an ?xtr offline? message reflects an extremely high number of the digital data interruptions rapidly occurring. in order to correct this issue, a new version of the ti msp430 is being implemented that includes a hardware correction for the clocking defect. spacelabs considers this medwatch closed. further actions and information will be documented under the spacelabs field corrective action process.><spacelabs received a report that on (b)(6) 2016 four telemetry beds went offline from the xhibit central monitor model 96102. no injury was reported as a result of this event. this condition was quickly remedied by resetting the monitor. | 1
5881075 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><the reported issue has been identified as a software problem. spacelabs has initiated a field corrective action and notified the fda (b)(4) district office of this action on august 25th, 2016. (recall report number: 3010157426-08/25/16-003-c) we also began notifying customers of this activity with a letter dated (b)(6) 2016. this investigation is considered complete and the issue closed.><spacelabs received a report that on (b)(6) 2016 the xhibit central monitor system on med west displayed offline" and "not available" messages. no injury was reported as a result of this event." | 1
5909482 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><audible and visual alarm data was collected from the central station and analyzed for evidence of alarm indicators during episodes of bradycardia. findings from this data indicate that an audible alarm tone was initiated and that the waveform zone flashed for episodes of bradycardia. thus, per episodes collected during investigation, appropriate audible and visual alarm notifications were provided and the device performed to specifications. we know of no reason that alarm notifications would not have been present at the time of the complaint episode as these indicators operated according to specifications when tested by a spacelabs? field service engineer. this investigation is considered complete and the issue closed.><spacelabs received a report that on (b)(6) 2016 a telemetry patient experienced an episode of bradycardia that did not generate an audible or flashing visual alarm at the central station. the alarm was captured in the retrospective database. no one was injured as a result of this event. | 1
5917724 | spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><spacelabs received a report that on (b)(6) 2016 the xhibit central monitor lost connection to the telemetry receiver and all patients dropped off the telemetry central. the issue was resolved by restarting the system. no injury was reported as a result of this event.><the reported issue has been identified as a software problem. spacelabs has initiated a field corrective action and notified the fda (b)(4) district office of this action on august 25th, 2016. (recall report number: 3010157426-08/25/16-003-c) we also began notifying customers of this activity with a letter dated august 25th, 2016. this investigation is considered complete and the issue closed. | 1
5919761 | spacelabs received a report that on (b)(6) 2016 the xhibit central monitor wouldn't stop rebooting. no injury was reported as a result of this event.><onsite investigation by a spacelabs field service engineer (fse) confirmed a video splitter attached to the xhibit was causing the re-booting. the splitter was replaced. the unit passed all functional tests and was restored to service. this report is considered final and the issue closed. | 1
5870759 | spacelabs received a report that on (b)(6) 2016 telemetry waveforms disappeared off the central monitor. the issue was resolved by restarting the system. no injury was reported as a result of this event. this was a second occurrence for this central station on the same date.><spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete. | 1
6037503 | spacelabs received a report that on (b)(6) 2016 the xhibit central station kept rebooting. no injury was reported as a result of this event.><spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><patient monitoring resumed by disconnecting the network cable, rebooting the central monitor, then readmitting the telemetry patients to the device. onsite investigation by a spacelabs field service engineer (fse) found that the problem had resolved, and could not be reproduced. there have been no further reports of recurrence for the past two months. this investigation is considered complete and the issue closed. | 1
5932722 | spacelabs received a report that on (b)(6) 2016 the xhibit central station on tower 2 was in use, then began continuously restarting with no intervention from the user. no injury was reported as a result of this event.><spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete. | 1
5897978 | the reported issue has been identified as a software problem. spacelabs has initiated a field corrective action and notified the fda (b)(4) district office of this action on august 25th, 2016. (recall report number: 3010157426-08/25/16-003-c) we also began notifying customers of this activity with a letter dated august 25th, 2016. this investigation is considered complete and the issue closed.><spacelabs received a report on july 18, 2016 that the xhibit central monitor system showed all beds offline" messages at 2:30am (b)(6) 2016. the issue was resolved by system restart which did restore operation. no injury was reported as a result of this event.><spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete." | 1
5662246 | tubing removed from packaging, a section of the tubing was visualized to be kinked. tubing not used due to the kinking.><. | 0
5800205 | a pediatric patient in (b)(6) was undergoing therapeutic plasma exchange treatment (tpe) for guillain-barre syndrome . the tpe set was connected via an extracorporeal membrane oxygenation (ecmo) circuit with smith stopcocks. at the completion of the tpe treatment, kelly clamps were used to disconnect the tpe access line from the ecmo circuit and the male luer broke off in the stopcock. the blood in the extracorporeal circuit was not returned to the patient resulting in an estimated blood loss of 71 ml. | 0
5641499 | haemonetics received a complaint on (b)(6) 2016 for a report of anticoagulant depletion during a plasmapheresis. there was no report of any donor injury or reaction.><this device was not returned for evaluation. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace. (b)(4). | 0
5353813 | haemonetics received a complaint on (b)(6) 2015 for a report of anticoagulant depletion during a plasmapheresis. there was no report of any donor injury or reaction.><this device was not returned for evaluation. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace. haemonetics filed a medical device safety alert and is awaiting the fda correction/removal reporting number. not returned. | 0
5782517 | haemonetics received a complaint on (b)(6) 2016 for a report of anticoagulant depletion during a plasmapheresis. there was no report of any donor injury or reaction.><haemonetics has requested that the rotors be returned for evaluation. the rotors have not been returned. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace. it was confirmed with the customer that since receiving the haemonetics medical device safety alert they have been using the approved cleaning solutions on the pump rotors. not returned. | 0
5412917 | haemonetics received a complaint on (b)(6) 2016 for a report of anticoagulant depletion during a plasmapheresis. there was no report of any donor injury or reaction.><the haemonetics field service engineer (fse) went to the center to evaluate the device as it was also due for preventative maintenance. the issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. the fse completed full diagnostic testing on the device. during the testing, 2 ac depletions were observed. the rotors were checked with durometer testing and they passed the specification, however, the fse observed the rotors to be sticking. the fse replaced the rotors and the device ran as expected. haemonetics filed a medical device safety alert and is awaiting the fda correction/removal reporting number. | 0
5538910 | haemonetics received a complaint on (b)(6) 2016 for a report of anticoagulant depletion during a plasmapheresis. there was no report of any donor injury or reaction.><this device was not returned for evaluation. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace. haemonetics filed a medical device safety alert and is awaiting the fda correction/removal reporting number. | 0
6078775 | haemonetics has requested that the rotors be returned for evaluation, however, the customer has not responded to date. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace.><haemonetics received a complaint on 10/14/2016 for a report of anticoagulnat (ac) depletion noticed at end of procedure, no donor reaction. | 0
6078790 | haemonetics has requested that the rotors be returned for evaluation, however, the customer states the rotors are not available to return. this issue of anticoagulant (ac) depletion has been investigated under a corrective action. the results of that investigation determined the likely cause was the harsh cleaning solution used to clean the pump rollers at the customer site can cause damage to the pump rollers which may lead to a device malfunction. new rotors were sent to the customer for the on-site technician to replace. haemonetics confirmed that the customer is aware of the haemonetics medical device safety alert regarding the proper cleaning solutions to use to prevent the early degradation of the rotors. not returned.><haemonetics received a complaint on 10/17/2016 for a report of anticoagulant (ac) depletion noticed at end of procedure, no donor reaction. | 0
5980600 | ><an (b)(6) male admitted with leg weakness and fall, no loss of consciousness or head trauma on (b)(6) 2016. pt had a recent history of tachy-brady syndrome and was being followed by cardiology with plan for a permanent pacemaker placement. pt had a ct of his chest/abdomen and pelvis showing widespread metastatic disease. mri of the thoracic spine was ordered to rule out spine metastasis in the setting of leg weakness. pt did have a significant past history of peripheral vascular disease with history of abdominal aortic aneurysm repair, with left leg bypass graft in addition to aneurysm surgery in 1982 or 1983 at (b)(6) hospital. however, there was no readily available info regarding clipping of aneurysm. pt was alert and oriented and had the mri check list presented to him twice, once on the nursing unit and then in the mri suite and pt denied having any metal in his body. during the mri, pt became unresponsive and had a code stroke called, repeat ct head scan showed a large intracranial bleed. findings discussed with family and decision was made for comfort care and pt expired on (b)(6) 2016. final autopsy results received on (b)(6) 2016 with cause of death attributed to the rupture of the aneurysm of the right posterior communicating artery. the previously placed aneurysm clips were present: one sundt-kees clip at the point to tri-furcation of the right mca and one mcfadden clip at the a1 segment of the left aca. there was no evidence of vascular tear or damage at or adjacent to the clips, or evidence of displacement of the clips as both clips were present at their original location as referenced on the cranial ct's. | 0
5681325 | .><mgs-80 gram stainer showing potential carryover/cross contamination issue. the instrument has been placed out of service; investigation into gram stain and culture results is ongoing. | 0
5537445 | at our institution, no one is sitting in front of the ehr mdds terminal monitoring the incoming results of every patient, as we do with heart telemetry monitors. therefore, results come back to the mdds and no one knows that they are there. there is not an alarm or warning; and certainly zero notification of abnormal results. this includes data as vital signs, which have been relegated to a less than vital status by the flawed or absent alarm system in the ehr. in this case, as in many others, vital signs were obtained by machine prompted by techs and these machines send the data to the ehr. the blood pressure was dangerously low, yet the nurse administered three medications that lower blood pressure. the blood pressure decreased further requiring intervention and a increased length of stay. there have been other incidents of disease critical vital signs being deposited in the ehr without the awareness of the clinicians, including a case of what turned out to be (b)(6), when the patient was discharged despite a markedly elevated temperature (and ultimately died, and caused (b)(6) in additional primary and derivative expenses). | 1
5683641 | during a total hip arthroplasty, it was discovered the orthosize v1.2.6 software templated a size 7 stem. the surgeon broached to the appropriate size and implanted a size 10 stem. there was no patient injury or delay in procedure.><current information is insufficient to permit a conclusion as to the cause of the event. evaluation of software found no evidence of product non-conformance. digital implant image comparison testing was conducted and all images were accurate. analysis of patient x-rays found the incorrect protocol was used, resulting in the incorrect template sizes. there are warnings in the package insert that state that this type of event can occur: a trained medical professional determines final implant type and size intraoperatively." "orthosize software does not determine the final size and type of hardware to be implanted."" | 1
5341146 | it was reported that during an initial total hip arthroplasty procedure on an unknown date the orthosize v1.2.6 software did not correctly template the stem. the surgeon broached and attempted to implant the stem, but it did not fit. the stem was removed and the surgeon continuously broached until the appropriate size was reached.><this follow-up report is being filed to relay additional information, which was unknown at the time of the initial medwatch. evaluation of software found no evidence of product non-conformance. digital implant image comparison testing was conducted and all images were accurate. a conclusive root cause of the event could not be determined.><current information is insufficient to permit a conclusion as to the cause of the event. there are warnings in the package insert that state that this type of event can occur: a trained medical professional determines final implant type and size intraoperatively." "orthosize software does not determine the final size and type of hardware to be implanted."><this follow-up report is being filed to relay additional information which was unknown at the time of the initial medwatch.><it was reported that during an initial total hip arthroplasty procedure on (b)(6) 2015 the orthosize v1.2.6 software templated a size 8 stem. the surgeon broached to size 8 and attempted to implant the stem; however it did not fit. the stem was removed and the surgeon continuously broached until the appropriate size was reached. a size 12 stem was implanted." | 1
5683639 | current information is insufficient to permit a conclusion as to the cause of the event. evaluation of software found no evidence of product non-conformance. digital implant image comparison testing was conducted and all images were accurate. a conclusive root cause of the event could not be determined. there are warnings in the package insert that state that this type of event can occur: a trained medical professional determines final implant type and size intraoperatively." "orthosize software does not determine the final size and type of hardware to be implanted."><during a total hip arthroplasty it was discovered that the orthosize v1.2.6 software templated a stem 3 to 8 sizes smaller than the stem size that was implanted. the surgeon broached to the appropriate size and implanted a size 11 stem. there was no patient injury or delay in procedure." | 1
5750278 | this report documents information related to an issue identified with functionality included in cerner's millennium powerchart ltc nursing documentation. the issue occurs when the nursing documentation downtime mar report is used to administer medications to patients when cerner millennium powerchart ltc is unavailable. in cerner millennium, when the user prints the downtime mar report, suspended medications are included on the report but do not include a note of the suspended" status. patient care could be adversely affected as clinicians may administer suspended medications in error. cerner has not received communication on any adverse patient events as a result of this issue.><cerner distributed a priority review flash notification flash16-0291-0 on june 15 2016 to all potentially impacted client sites. the software notification includes a description of the issue and notice that a software modification is being developed to address the issue for all sites that could potentially be impacted. cerner distributed an updated flash notification flash16-0291-1 on november 4 2016. the software notification includes a description of the issue and notice that a software modification has been developed and is available to address the issue for all sites that could be potentially impacted. cerner corporation considers the issue to be resolved and no further narrative is required for follow-up.><cerner distributed a priority review flash notification (b)(4) on june 15 2016 to all potentially impacted client sites. the software notification includes a description of the issue and notice that a software modification is being developed to address the issue for all sites that could potentially be impacted. ? cerner corporation will provide a follow-up report when the software modification is available.><this report documents information related to an issue identified with functionality included in cerner's millennium powerchart ltc nursing documentation. the issue occurs when the nursing documentation downtime mar report is used to administer medications to patients when cerner millennium powerchart ltc is unavailable. in cerner millennium when the user prints the downtime mar report suspended medications are included on the report but do not include a note of the "suspended" status. patient care could be adversely affected as clinicians may administer suspended medications in error. cerner has not received communication on any adverse patient events as a result of this issue." | 1
5579500 | this report documents information related to an issue identified in cerner's healthelife. this solution is a patient portal that provides patients with a view of clinical information from their cerner millennium emr (electronic medical record), including a list of their prescribed medications. the issue occurs when the prescription is filled by a pharmacy using cerner's retail medication manager, and the pharmacy changes the strength, dose or frequency of the ordered medication and adds signature line instructions to describe these changes. the updated signature line instructions are displayed in the portal, conflicting with the dose and frequency of the original prescription which remain unchanged in the portal. patient care could be adversely affected in the event the patient mistakenly uses the portal as a source of dosage instructions, as the instructions displayed are not for the strength, dose and/or frequency displayed in the portal. this may lead to the patient taking an incorrect amount of the medication resulting in overdose or underdose. cerner has not received communication on any adverse patient events as a result of this issue.><cerner distributed a priority review flash notification (b)(4) on april 8, 2016 to all potentially impacted client sites. the software notification includes a description of the issue and notice that a software modification has been developed to address the issue for all sites that could potentially be impacted. cerner corporation considers this issue to be resolved and no further narrative is required for follow-up. | 1
5483227 | linac treatment 40 ap without shift back to iso center (1/2 treatment shift). dose of 170 c6y versus 175 c6y. linac displayed 2 untreated fields when treatment had been completed. radiation therapy chest on line did not display plan that was filmed and treated in the patient history. varian help desk notified, after 1 hour search treatment of 2 fillers were located in a backup file on the linac computers. treatment and plans were manually imported to rt chart. six chemotherapy treatments completed (b)(6) 2005. | 0
5343898 | in 1999, clarify case was submitted to the american red cross (arc) to report an issue with group operations. when group operations was used to add a donation property(gatx) to a series of whole blood numbers (wbns), the property should have automatically been applied to every component for that donation. instead, group operation did not add the property to all components. upon investigation, it was discovered that group operations will only add donation properties to in-house components (e.g. not shipped). for example, assume a wbn is collected and red cell and platelet components are produced. the components are tested and determined to be suitable for release and the platelets are shipped to a customer, while the red blood cells remain in house. subsequently, a (b)(6) property (unsuitable for release) is added to the wbn through group operations. if the platelets are then returned, they would not receive the (b)(6) property and would be allowed to be shipped. in summary, if the shipped component is returned to inventory, there is no computer mechanism to prevent re-issuing of the component. due to the fact that configurable donation properties that are designed to prevent release would not have been added to shipped products, american national red cross (arc) believe this should have been reported. donation properties should be added to every component for the donation, regardless of whether a component has bee shipped.><arc sent a blood services letters (bsls) ((b)(4)) to inform regions of the issue and instructed the regions not to use group operations to add donation properties to any donations that had shipped components. the bsl was issued in may 1999. in addition a record review was performed to ensure that no erroneous releases were identified. the result of this review found that no erroneous releases occurred and no units were required to be recalled. the root cause of this defect is in the original software received from (b)(4) in 1996. the design and functionality did not include adding a donation property to a component that had been shipped when applied using group operations. a change request ((b)(4)) was developed to modify group operations functionality to add donation properties to all components for a donation, including shipped components. this change was delivered by minor software patch, s1364 to all regions by september 8, 1999. | 0
6139589 | the reported issue has been determined to potentially affect picis perioperative and critical care versions 8.6 service pack 4 and 8.6 service pack 5 when both preop manager and or manager are utilized together. a software solution has been created and will be made available to potentially affected clients and included in a future 8.6 service pack. additionally, the issue is experienced when an unusual client workflow is followed. according to clinical representative feedback, it is not typical for clients to enter precautions into picis preop manager ahead of a booking being created in picis or manager. the company reviewed related user documentation and determined the user documentation may be enhanced with a cautionary statement to ensure customers are aware of the need to create a booking in picis or manager prior to the precautions being entered in picis preop manager. user documentation will be enhanced with the release of a future 8.6 service pack.><picis clinical solutions, inc. (picis) has received and investigated a customer report of an interoperability issue within the picis perioperative and critical care software application. customer reported that a precaution entered into preop manager did not flow to the or manager application as expected. the investigation into the reported incident concluded that the root cause resulted in the software not capturing the precaution in or manager that was entered into preop manager. the investigation has identified an issue where the picis intermessaging service (pims) channels create excessive threads and adversely impact the picis connectivity manager (pcm). as a result, messages may not be processed as expected (the message that links the preadmission patient and the patient from the booking may not be processed). our findings revealed the unique situation reported by the complainant occurs when clients follow a particular workflow. after internal investigation with clinical resources, it was determined that the workflow utilized by this reporting client was not common. reporting of this condition from clients is rare (only one reported client complaint to date after more than ten years on the market), as the issue is only experienced with unique conditions and workflows. a software solution has been designed to address this issue. the user documentation will be enhanced to include a cautionary statement regarding the proper data workflow. event description customer report: this patient had a slate alert (osam) entered in pom (picis preop manager) by the pss rn yesterday (b)(6) 2016 but did not appear on the slate today (b)(6) 2016 as seen below. the nurse has now entered the alert in orm (picis or manager) and this carried over to the slate. we need to ensure this does not happen for safety reasons as we were under the understanding that we are to enter slate alerts in pom (picis preop manager) and these will carry forward to the slate."" | 1
6041018 | a morbidly obese patient was having a fibrillation or pulmonary vein isolation procedure in ep lab. the x-ray table is from general electric medical systems with system id innova. the doctor had difficulty exchanging the sheath in the right groin and moved the table forward to look at the groin under fluoro. when the table was moved, there was a loud noise and the upper part of the table dropped and the base of the table lifted up from the floor. the procedure was aborted and the patient was transferred to a stretcher.><. | 0
5384877 | at about 1350, digital image (di) techs reported smoke coming from the di film reader. manager went to investigate and noted smoke billowing out of the machine. smoke was filling the hallways with a pungent smell of electrical fire. 911 called and the fire alarm was pulled, on-site safety officer notified. fire company arrived and used industrial fans to blow smoke out of the ed. no patients or employees harmed. fire dept. cleared the scene at 1405. machine removed from service. no visible fire.><. | 0
5485538 | nicu staff smelled smoke in the area of the pacs cr reader. staff noted heavy smoke coming from the reader.><. | 0
6144295 | the customer complained that a female operator has got a fractured nose after unintended release of the stretch grip swivel lock (spline mechanism) during stretch grip positioning for the patient.><(b)(4). the investigation is still ongoing on this event. when the investigation is completed a follow-up report will be sent to the fda . | 0
5363844 | (b)(4).the icp express was returned for evaluation. it was found that the enclosure was dented at several edges and the rear connector hex screws were damaged. the lcd was not functional, the internal transducer cable was not functional, and the battery was weak and would not hold a charge. the supplier replaced the lcd, internal transducer cable, hex screws and battery. the supplier was not able to conclusively determine the direct cause of failure. however, the device was manufactured on 27 june 2013 and is therefore outside of warranty. trends will be monitored for this or similar complaints. at present, we consider this complaint to be closed.><(b)(4). upon completion of the investigation a follow up report will be filed.><unit error '-99' ? error with grey cable unit could not be zeroed. on (b)(6) 2016 per affiliate: there was no injury to the patient reported. please find attached an updated case form with serial number for the icp express. the icp express cable is also available for return. on (b)(6) 2016 per affiliate: the product specialist noted in the report that same like device was used to manage the problem during the procedure. a total of 35 mins is the time the rep indicated the delay for."" | 0
5719947 | dignishield stool management system device was used on pt. pt had a colonoscopy with biopsy and control procedure performed on (b)(6) 2016 which showed areas of necrosis from th is device. | 0
5719948 | dignishield stool management system device was used on pt. pt had a colonoscopy procedure performed on (b)(6) 2016 which showed areas of necrosis from this device. | 0
5354930 | ><report received of a user error resulting in suction swab disengagement; labeling instructions regarding use of bite block were not followed. the reporter provided the following information. on (b)(6) 2015, a nurse was providing oral care to a confused and combative patient. the patient reportedly bit off the tip of the suction swab. upon being bitten, the end piece of the suction swab, including a portion of the plastic straw, broke off in the patient's mouth. the nurse was unable to open the patient's mouth to retrieve the suction swab. the broken piece of the suction swab reportedly partially occluded the patient's airway. the patient was sedated and the broken piece was visualized and removed with forceps. the patient's lip and throat were cut and bleeding and the patient remained hospitalized unrelated to the incident. reporter stated, unrelated to the incident, the patient was intubated sometime after the event due to the patient's reported medical history. instructions for use state: use a bite block when performing oral care on patients with altered levels of consciousness or those who cannot comprehend commands." reporter provided that while the patient was reportedly confused and combative a bite block was not in use at the time of the incident. although requested no additional information was available." | 0
6142011 | the customer contacted the siemens customer care center (ccc). ccc connected remotely and checked a patient sample. ccc discovered that data was not purged after 30 days as configured. a siemens headquarter support center (hsc) specialist connected remotely and performed a data base clean up, purge and defragmentation. hsc checked the database, and no issues were noted. the cause of the database not purging as configured is unknown. the instrument is performing according to specifications. no further evaluation of the device is required.><the (b)(4) lab data manager was not purging data as configured, causing results to stay in the system past the time period configured. the customer reuses barcodes every 90 days, and discovered that old results were being confused with new orders and the same barcodes were associated with different patients. there were no incorrect results or data reported out by the customer. there are no known reports of patient intervention or adverse health consequences due to the (b)(4) not purging data as configured. | 1
5735992 | ><patient information could not be obtained due to country privacy laws. the initial reporter is located outside the u.s., and therefore this information is not provided due to country privacy laws. manufacture date is unknown. ge healthcare's investigation is ongoing. a follow-up report will be submitted when the investigation is complete. device evaluation anticipated, but not yet begun. | 0
5737782 | ><the stereotactic imaging on the cds was not working. required a total of 3 discs with imaging to be retrieved, while the patient was under anesthesia (approx. 14 hours). this is what the patient safety report stated: pa was attempting to bring up the stereotactic image on disc on the stryker unit prior to dr beginning surgery. the disc would not function or open the image. the patient was under anesthesia at this point without available image. pa called the file room. the cd burner was not working. he was connected to medical record department who burnt a disc for him. once back in the operating room, it was realized that the second disc had the incorrect image burnt on it. while pa was retrieving the 3rd disc from medical records, the stryker rep made available a second stryker unit. prior to pa returning to operating room with a third disc, the first disc was able to be utilized to open the image on the second stryker unit. dr was able to begin the surgery. it is reported that once surgery began, the patient had been under anesthesia without any surgical intervention for approximately one hour. it was later learned that the disc drive in the first stryker unit was dysfunctional. it is possible that if the disc drive were to have worked in the first stryker unit, that the first disc may have opened the image. overall, three concerns: the non functioning disc drive in the stryker unit. the inability to begin surgical intervention while the patient is under anesthesia. an inability to utilize the pac system to open image, requiring the image to be burnt on a disc (acknowledging this is not a new process). it is reported that it is not uncommon that the initial disc brought to the operating room is not functional for various reasons and a second disc has to be retrieved during or cases. the stryker imaging rep was present for the case. | 0
5490749 | the site indicated that the incident unit will not be returning to the manufacturer for evaluation. if additional information pertinent to the incident presents itself, a follow-up report will be submitted. (b)(4).><patient had a patency capsule procedure. she has a history of crohn's disease plus a couple of surgeries. she ingested the agile capsule and did not see it pass. she had a x-ray and they found the capsule (husband is not sure the location). the plan is the patient is to go to the hospital and ingest contrast material. the husband did not know if or what procedure would be done. he said they would wait for the results of the contrast ingestion to make a decision. the patient continues to have normal bowel movements but is currently in a lot of pain. two weeks later, the patient's husband reported the barium enema with contrast was done, capsule was not seen, and there was no obstruction. a possible 15mm kidney stone was visualized. next step is to meet with surgeon regarding possible kidney stone. | 0
5387506 | the capsule was placed endoscopically and data recorder was attached. after capsule was removed from the patient and connected to the computer, there was no video recorded.><. | 1
6010685 | (b)(4).><according to the reporter, the patient ingested a capsule and it got stuck in the esophagus. the patient previously had a gastric-band. the customer will perform an egd and try to remove the capsule also it location to see if it has moved. the account confirmed that there was no harm to the patient. a repeat procedure was not performed. follow up information revealed that the treating doctor removed the first capsule endoscopically and tried to place a second capsule at the same time. he was unable to place the second capsule also due to the gastric banding. post-endoscope, the patient recovered, did not require any additional medical intervention and was discharged to home doing well. no other known adverse events were reported. | 0
5997958 | according to the reporter, the patient swallowed a capsule but the capsule failed to transmit. the device operator had to endoscopically retrieve the capsule and then have the patient take a different capsule. the second capsule was blinking and was paired the entire time.><(b)(4). (b)(6). | 0
5921923 | (b)(4). the sample has been requested but to date, the incident sample has not been received for evaluation. if the sample is received or if additional information pertinent to the incident is obtained, a follow-up report will be submitted.><customer reported patient who ingested a pillcam capsule and within hours reported to having contracted an allergic reaction extending from the torso out to the arms. the allergic reaction has been resolved. | 0
6100411 | according to the reporter, a procedure was ongoing when the capsule was not found anymore. they took capsule out of stomach and it couldn't be found by the physician anymore. they used a 2nd capsule and that capsule worked fine.><(b)(4). | 0
5429902 | .><patient had capsule endoscopy. patient swallowed pill cam and had sensor belt applied. no video was recorded. | 1
6063550 | ><ge healthcare has received notice of a lawsuit, alleging that the corometrics monitor mistakenly monitored and recorded the maternal heart rate at times during the labor including the second stage. as a result the doctors and nurses were misled into thinking the fetus was fine when it was not and needed to be delivered by c-section." the baby reportedly "suffered brain cell death and brain damage due to hypoxic ischemic encephalopathy"." | 1
6088848 | unable to print strips in icu for patients interpretation. one patient had to temporarily be hooked up to crashcart defibrillator/monitor due to the patient being unstable. no identified harm to patients.><. | 1
6088924 | there were two central monitor banks blank. eight patients were non-monitored for a period of time; lost partial viewing capacity. internal facility biomed staff addressed. also, five minutes of storage capacity lost. there was no apparent harm to patients.><. | 1
6213299 | .><while doctor was using the fenestrated bipolar forceps, it was noticed to have a crack in the plastic/rubber near the head of the instrument. it was immediately removed from service and replaced with a new instrument. | 0
5537655 | the incident is similar to an fda alert described as in certain circumstances involving select davinci si single-site instruments the jaws of grip actuated instruments may become fixed in a closed position." in our specific case the physician was using a da vinci robot mono-polar curved scissors (ref: 420179). the scissors worked for half the case and then stopped opening. the doctor tried manually opening them from the dials in the back with no success.><." | 0
6145123 | .><during the procedure, the robot had an unrecoverable fault. lost the use of arm number one and had to finish case with three arms. | 0
6010536 | .><dr noticed (while watching his recording of the procedure he performed a few days prior) that a cable on the tip of the maryland bipolar forceps popped and a tiny piece of the plastic insulation broke off into the patient while he was using the instrument in a normal fashion. at the time, this was not noticed and the piece was left behind inside of the patient. | 0
5642877 | during a robotic hysterectomy procedure a blade broke off of the harmonic ace shears. the surgeon was notified. a search made for broken blade was found.><. | 0
6033516 | .><the blade would not advance and stopped cutting. | 0
5510055 | the vessel sealer was not properly cauterizing the tissue.><. | 0
5976514 | .><the device had six lives left ,yet was unable to cut the suture. | 0
6033548 | .><the wire broke with eight lives still remaining with instrument. | 0
6088776 | .><during a robotic-assisted laparoscopic total hysterectomy with bilateral salpingo-oophorectomy with lymph node dissection, the doctor and operating room staff noticed that while she was using the xi maryland bipolar forceps, the pulley cable on the tip of the instrument popped. the instrument was immediately removed from inside of the patient's body and inspected by the scrub nurse. it was noticed that a piece of insulation that had been underneath cable was missing. the scrub nurse informed the doctor right away. the doctor actively looked into the patient's cavity and recovered the broken piece. the instrument and insulation piece were passed off of the sterile field. | 0
6092184 | during robotic surgery, the permanent cautery spatula was being used inside the patient's abdomen when the scrub tech saw on the video monitor that it looked like a piece of plastic coating was flaking off the instrument. the surgeon used a robotic prograsp instrument to retrieve the flake from the cautery spatula. the plastic piece and the defective instrument were then withdrawn from the patient's abdomen. the plastic flake never fell off the instrument into the patient. a new cautery spatula was used for the rest of the case.><. | 0
6213328 | .><while docking, the robotic bipolar cord was attached to bipolar maryland instrument and began firing with no one at the robotic counsel. the cord was unplugged. we tried the instrument again and it did the same thing. the bipolar maryland and cord was replaced. | 0
5989733 | the hook portion of the davinci instrument separated from the instrument. pieces were removed from the field. update to company notification to follow.><. | 0
5737959 | .><while using da vinci xi stapler after first load was fired, we noticed a tear at the top of the sheath. after taking a closer look, a piece was missing, which had broken off in the patient. the piece was removed from the patient and a new sheath was put on the stapler. | 0
5619298 | .><during normal use of the intuitive davinci vessel sealer, the instrument would no longer articulate. upon inspection of the instrument, it was discovered that one of the cables had broke. instrument taken out and replaced. | 0
5821774 | .><during the procedure, the robot alarmed and sent an error message that arm number three was disabled. the representative from the company was in the room and was unable to fix the problem, so arm number three had to be disabled for the rest of the procedure. the procedure could not continue robotically and had to proceed to laparoscopic, and finally an open procedure. there was not injury to the patient. | 0
5636574 | upon inspection after the case, staff noticed the tip cover was cut, which is something that protects the patient from electrocautery burns. looking at the tip cover, it looks like it was burned, which caused the cut. the tip cover also has two burn holes on it. when looking at the scissors, burn marks match up with the burns on the tip cover. staff did not notice anything during case. these results were seen after the completion of the procedure.><. | 0
5768626 | .><at the start of the case the harmonic ace curved shear was attached to the robot. the instrument was plugged into the machine and activated. the machine alerted to check instrument." the instrument was removed from the robot cleaned checked and placed back in the robotic arm. the instrument was re-activated and at that time the harmonic ace broke and a piece of metal fell off. this metal piece was retrieved and removed from the patient." | 0
5384865 | .><blade became exposed, causing error and causing robot to terminate function ability. manufacturer response for vessel sealer, intuitive (per site reporter): pending return authorization. currently being shipped back to vendor. | 0
5878262 | .><in the middle of surgery, the vessel sealer alarmed and stopped working. it was removed and replaced. failed self-check" appeared on robot monitor. manufacturer response for vessel sealer endo wrist intuitive (per site reporter): replacement/ credit." | 0
5878263 | blade exposed. stopped working. manufacturer response for vessel sealer endo wrist, intuitive (per site reporter): credit.><. | 0
5878261 | .><in the middle of surgery, the vessel sealer alarmed and stopped working. it was removed and replaced. failed self-check" appeared on robot monitor. manufacturer response for vessel sealer endo wrist intuitive (per site reporter): credit." | 0
5440694 | a small piece of the 8mm permanent cautery hook belonging to the davinci xi robot fell off of the instrument. it was retrieved and removed from patient's abdominal cavity. the broken hook was removed from the surgical field immediately. the davinci representative was in the room at time of incident.><. | 0
6043595 | .><during the procedure, the grey piece of the tip cover accessory that is used with the endowrist monopolar curved scissor to insulate the tip, kept sliding off the scissors. this piece was removed and replaced with a new tip cover accessory. upon removal of the original cover, it was noted that the clear tip of the grey piece was cracked. while using the monopolar curved scissor with the second grey piece, it was noted that this second piece completely slid off the scissor and was in the patient's abdomen. this grey piece was then retrieved from the abdomen by the surgeon. a new monopolar curved scissor and tip accessory was obtained and they work appropriately. acm was notified and original scissor and two malfunctioning tip cover accessories were bagged and give to her. gray pieces bagged and labeled and given to acm. spd notified, they keep records of defective equipment. to be returned to company for review. team followed policy to ensure no retained surgical items. | 0
5784823 | worked a few times and then stopped working. manufacturer response for harmonic, (brand not provided) (per site reporter): in process of return and feedback.><. | 0
5999535 | vessel sealer was being used and an error occurred. the vessel sealer was removed and inspected. blade was found to be exposed.><. | 0
5398144 | when the vessel sealer was inserted into the cannula, it would not work. the arm went yellow and the light on the vessel sealer box went red. we opened another and it worked perfectly.><. | 0
5398143 | .><surgeon attempted to use the vessel sealer to cut a suture but the blade would not engage. it was tried a few times and still did not operate properly. a new item was opened and worked fine. | 0
5764610 | monopolar curved scissor that has ten uses came apart while in the abdomen of the patient.><. | 0
5938777 | .><small <2mm plastic piece of instrument became dislodged from instrument. item seen and removed by surgeon. there was no harm to patient. | 0
5748101 | .><harmonic scalpel instrument appeared to have damage to teflon coating. this resulted in the instrument sticking to the tissue. | 0
6156683 | a beckman coulter field service engineer (fse) evaluated the olympus pk7300 automated microplate system. the fse found clot buildup in the lower cup of the wash station drain line. the fse cleaned out the wash station drain line to resolve the issue.><the customer reported an uncontained deionized water leak, involving the beckman coulter pk7300 automated microplate system. approximately one (1) gallon of deionized water leaked underneath the instrument. the customer was wearing personal protective equipment consisting of a laboratory coat, and gloves at the time of the event. the customer slipped and fell as a result of the uncontained water leak and went to the clinic with a swollen knee. the customer was sent home. | 0
6188627 | ><ge innova fluoroscopy machine completely shut down twice during the procedure. attempted multiple times to reboot the machine but efforts were unsuccessful. patient was not harmed. machine malfunction delayed coronary angiogram procedure. | 1
6102885 | the pt was on tele. when we attempted to complete tele strip the wave review button was not on screen at the front nurses station, but it was present at the back nurses station. we were unable to appropriately complete tele strip because recorded information not present.><. | 1
5909169 | on (b)(6) 2016, 3m learned that during the use of a 3m true definition scanner, an (b)(6) year-old female patient experienced an electrical sensation/shock when the scanning wand touched her teeth. this event occurred in italy and during the scan in which the electrical sensation/shock occurred, the dentist reported that the true definition scanner was plugged into an extension cord.><(b)(4) received an electrical safety inspection report from the dentist which stated that the dental chair was found to have no functional defects or deficiencies that were safety-related. no further complaints have been reported by the dental office since they discontinued use of the extension cord with the true definition scanner.><appropriate electrical safety measures (isolation transformer) are present in the true definition scanner system. the 3m true definition scanner user manual and instructions for use (ifu) (available in english and italian) contain the following: warning. to reduce the risks associated with hazardous voltage and fire - use only a properly grounded power outlet; do not use extension cords or multiple portable power socket outlets.' a third party site assessment is planned to take place after the (b)(6) holiday ((b)(6) 2016). 3m may also be provided with a copy of the latest yearly power socket assessment. any additional information that becomes available will be submitted by 3m in a follow-up report." | 0
6196476 | on (b)(6) 2016, 3m was notified that a male subject experienced an electrical sensation during a demonstration scanning procedure using the 3m true definition scanner. the subject felt tingling sensation on lower tooth #2 and #3. scanning was immediately stopped. no patient injury occurred.><the true definition scanner system involved in this event was sitting on the work surface charging in an outlet without using an isolation transformer (iso puck). the ifu states that a qualified (ul listed) stand-alone isolation transformer is provided together with a small power supply component for the table computer. these must be used together. the 3m? true definition scanner mobile edition is intended to operate routinely while connected to power. users may elect to equip multiple locations with the 3m supplied or recommended isolation transformer and power supply component to simplify transport. in this case it is expected these parts are incorporated into the operatories in a semi-permanent fashion." 3m has requested that the entire true definition unit be returned to 3m for testing to ensure the product is operating in specification. 3m is still waiting to receive this unit back. it is anticipated that an electrical site assessment will be conducted at the uk dental office. if any further information becomes available 3m will file a follow-up report." | 0
6033483 | .><the blade would not advance and stopped cutting. | 0
5387519 | faulty instrument; would not operate.><. | 0
5614313 | when the vessel sealer was removed from the patient, the surgical technologist could not remove it from its obturator. it was observed that tip of vessel sealer had bent and would no longer fit through the obturator. vessel sealer removed from operative field and given to materials management. no harm to patient.><. | 0
5387521 | .><pins are loose at the top of the instrument. | 0
5387522 | joints were too stiff to use product.><. | 0
5387523 | .><forceps will not close properly. | 0
5387524 | .><the jaws are catching and getting stuck. they are cutting suture before the surgeon is ready for the suture to be cut. | 0
6002852 | .><a 1 mm piece of plastic insulation material broke off a da vinci curved bipolar dissector. it has been brought to my attention that this is not the first occurrence of this nature with da vinci instrumentation. | 0
5959509 | .><in the course of the procedure, the surgeon saw a piece of plastic in the patient's pelvic cavity. the plastic was approximately 5 mm long and 1mm wide. the piece appeared to have broken off from the fenestrated bipolar instrument of the davinci robot while the instrument was in use in the procedure. another member of the surgical team was able to remove the plastic piece using an endoscopic grasper. the or staff contacted the davinci rep who was in house who responded and inspected the equipment. | 0
6002766 | during the robotic low anterior resection, the robotic stapler was loaded on to the arm and the system went into a non-recoverable fault and had to be restarted for the surgery to continue. the patient sustained no injuries during the rebooting of the system.><. | 0
5694258 | there was an error message on vessel sealer machine stating not placed in socket...using old software" no power on end of device while being used. turned off the machine and back on after few seconds device worked for 4-5 minutes then did not function with error message on vision cart stating "blades exposed"><." | 0
5387503 | device stopped working; yellow light came on.><. | 0
5681367 | an elderly male patient presented with malignant neoplasm of the prostate. he was scheduled for a laparoscopic robotic-assisted radical prostatectomy. vessel sealer alarmed because the blade was exposed. another vessel sealer was used and also did not work. slow and patient dissection with the bipolar maryland dissection on the left side and the unipolar scissors on the right side were used. the procedures were finished with no further complication or patient harm.><. | 0
5758187 | .><during a robotic gastric sleeve surgery, noted a foreign body inside abdominal cavity of o -ring". the item was removed from abdomen -- noted it probably came from the 8 mm cannula seal on device. a portion of cannula seal pulled apart from the seal. this was identified by sales rep in room from davinci. this is a first time incident. these seals are used on every robot case and this has never happened before. the instrument being inserted was the correct size for which the seal was designed." | 0
5623590 | .><while using the robotic prograsp, the surgeon couldn't open or close the jaws. upon inspection, the prograsp had a wheel off the tracks of the mouth wheel. | 0
5624467 | .><during a davinci tlh, the physician was trying to use the vessel sealer and it would cauterize but not cut. a new vessel sealer was opened and it worked without difficulty. no harm to patient. not reprocessed. | 0
6154197 | .><procedural note states: when it came time to use the stapler to fire across the sigmoid i found that the stapler would not clamp using a blue load despite multiple attempts. i converted to a green load; however there were multiple issues with the stapler not clamping the tissue properly. a different green load was tried and the stapler was able to clamp down the tissue and indicated that there were no issues. when the stapler fired however there was another malfunction with it cutting. it only partially transected the colon. at this point i decided to complete the transection with regular laparoscopic staplers ... i scrubbed and went to the table and completed the transection laparoscopically with an echelon stapler."" | 0
5500267 | .><blade exposure error came up on vessel sealer screen during tlh. old vessel sealer will be returned to intuitive. a new vessel sealer opened and used for remainder of procedure. no harm done to patient. | 0
5453517 | .><the device was not working as intended. | 0
5995332 | .><the instruments rubbed up against each other inside the patient belly. surgeon noticed a white flake powder on patient intestines then noticed white rub marks on both instruments. surgeon was concerned but all of this was at the end of the case as the procedure finished and they undocked the robot. | 0
5603046 | .><during robotic sigmoidectomy, the physician began dissecting with the vessel sealer and proceeded to cut the mesentery which had a vessel attached to it. as the physician was cutting, the vessel sealer did not divide and seal the cut he made and the patient began to bleed. attempts were made to stop the bleeding robotically/laparoscopically, eventually a mid line incision was performed to stop the bleeding. device was removed from the field and sequestered for analysis with manufacturer. | 0
5976605 | intuitive stapler malfunctioned during use, causing the stapler to become stuck mid-fire. device ref# (b)(4), lot# s10160330 280, 50 firing capacity w/ 2 firings used. the device was removed using the intuitive removal kit. the device was damaged during removal, with no apparent harm to patient. a green da vinci si endowrist 45 stapler reload was in use. ref# (b)(4), lot# ds6142516, expiration 6-2016. chart reviewed: procedural note states: the anterior peritoneal reflection was incised entering into the plane just between the mesorectal fascia and denonvilliers fascia. this plane was further developed and extended distally down to the level of the levators as well. the complete mesorectal excision was completed at this point including circumferentially down to the level of the levators. a robotic stapler device was then brought into the field and attempted to transect the distal rectum just above the levators however this device misfired and an echelon 60mm stapler with a green load was used to transect the rectum with 3 firings. the left colon was then again retracted laterally and the mesentery was divided just proximal to the ima pedicle to the level of the bowel wall circumferentially along the descending sigmoid junction. indocyanine green was instilled intravenously and used to perform fluorescence angiography of the descending colon which revealed adequate blood flow at a planned site of proximal transection."><." | 0
5757881 | .><while stapling across fissure, the stapler with blue load stapled both sides of tissue but failed to cut. | 0
6137769 | .><surgeon was performing an incisional hernia repair robotic assisted using the robotic large suture holder. the cable on the robotic large suture holder broke with no injury to the patient. the instrument was removed and replaced to continue with surgery. the lives on the instrument were noted and a broken tag was applied. risk was called. the instrument was placed in the risk bin at the or front desk. sterile processing department (spd) was called to inform them of the instrument damage. | 0
5856255 | pk endowrist broke in patient. the pk was removed and replaced. the broken pk was inspected and is intact. no patient harm, the procedures were completed without further incident.><. | 0
5583936 | during robotic assisted sigmoid colectomy, the surgeon fired the device while grasping tissue to surgically cut it, however the device did not retract to cut. no patient harm.><. | 0
6002806 | .><it was brought to my attention that in late (b)(6), during a robotic assisted laparoscopic sleeve gastrectomy the robotic stapler was loaded to the arm and it instantly went in to a non-recoverable fault. the system was powered down and restarted but again went in to non-recoverable fault requiring the system to be restarted. the stapler was again reloaded and for the third time the system went into non recoverable mode. the manual stapler was used to complete the procedure. no injuries were sustained by the patient. | 0
5702715 | .><while surgeon was attempting to suture the cuff closed, the mega needle driver stopped opening on both sides. it only opened on one side. pulled instrument out of the davinci machine and reset it hoping it would work on both sides, but it didn't. manually it seemed to work ok when moving the discs on the back. however it didn't work right during the case. another mega needle driver was obtained to finish the case. | 0
5684567 | scissor tip on monopolar scissors for the da vinci robot did not prevent thermal spread to adjacent tissue. during case dr. noted a small liver burn where the scissors were touching the liver while he was dissecting the adrenal gland off the right kidney. the instrument scissor tip cover was changed and no more thermal spread was noted. it was a small burn, dr. looked at it and noted no bleeding. it was called in to intuitive surgical and a report was filed. the product and original wrapper were sent to intuitive to be looked at.><. | 0
5856294 | .><during the procedure the device stopped running. the davinci arm blinked yellow and a message stated blade exposed". it was removed we tried to reload but were unable to do so. we opened a new device no harm to patient. robot assisted total laparoscopic hysterectomy bilateral salpingo-oophorectomy possible staging. disposable unit was sequestered and a new unit was opened and used without any delay to the procedure; defective unit to be sent to manufacturer for investigation. or management tracking with manufacturer." | 0
5852967 | .><during the robotic assisted right lower lobectomy, the dr fired the stapler across bronchus and the stapler fired but did not cut across the bronchus as it should. we had to open a manual stapler and load to finish the resection. the rest of the procedure progressed uneventfully. stapler fired but did not resect bronchus as expected. | 0
5462743 | robotic endostapler malfunctioned while attached endoscopically to tissue. intuitive dvstst called; the surgeon to use staple tool removal kit to release instrument. staple tool removal kit used to release the instrument. device safely removed from the patient.><. | 0
5793550 | .><during robotic right colon resection, the stapler would not turn into the correct position for the surgeon to line up anastomosis to fire stapler. surgeon hand sewed the anastomosis. the reusable robotic stapler device was saved and sent to equipment manager to return to the manufacturer. | 0
5566710 | .><tip of davinci cautery hook has a hard plastic coating that chipped off when instruments collided in abdomen. the material itself is not radiopaque so x-ray would not be able to locate foreign body if still in abdomen. also the piece is so small it could have easily been suctioned up during the irrigation process. instrument has been removed from service and will be held by myself in case any further investigation is needed and will later be returned to intuitive surgical for further study and credit. | 0
5392790 | while the doctor was performing a robotic supracervical hysterectomy using pk forceps, during the procedure, he lost control of the instrument jams. the instrument was removed and inspected and loose wire was discovered. the instrument was replaced with a new one.><. | 0
5846394 | the physician was performing a robotic right lower lobectomy using davinci xi. during the stapling portion of the procedure, while using the robotic stapler, there was a stapler misfire where only part of the stapler fired. physician asked for the endo gia stapler to finish stapling the bronchus. a davinci rep was present during incident. the stapler was removed from service and a manual stapler used to complete the stapling. the patient has sustained no obvious injuries and remainder of procedure was uneventful. the stapler still has 25 of 50 lives remaining. lot # for endowrist stapler 45 reload is m10151217 exp:12/31/2017.><. | 0
5496634 | .><during robotic hysterectomy, a disposable vessel sealer failed to work. no harm or consequences. | 0
5839528 | .><during a procedure a permanent cautery hook was loaded on to the robot arm and inserted into abdomen. while the surgeon was working a small piece of the tan insulation fell off the tip of the instrument into the patient abdomen. this was instantly retrieved by the surgeon and removed from abdominal cavity and instrument removed from field. there were no injuries sustained by the patient. both the instrument and the dislodged insulation were secured by the cst. | 0
6144561 | .><during robotic thoracoscopy the stapler misfired, the tissue was not fully cut and stapled. a new stapler was used for the remainder of the surgery. no patient harm. | 0
5989899 | .><during a robotic assisted sigmoid colectomy, the speaker on the monitor facing the patient that is used to communicate with the console was cutting out sporadically. the surgeon noticed that the coagulation device was not sealing and told the physician assistant not to cut the tissue. this was not heard by the staff, including the physician assistant, who then divided the tissue resulting in a small amount of rapid blood loss that was controlled quickly. total blood loss was 50ml. | 0
6085332 | .><surgeon was performing a robotic assisted abdominal perineal resection and using the robotic grasping retractor when the cable on it broke. there was no injury to the pt. instrument removed from the sterile field and replaced. | 0
5771705 | the vessel sealer gave an error; the blade came out and would not go back into the device. opened a new device and it worked properly. no harm to patient.><. | 0
5818620 | .><while in use, the blade on the snap-fit endowrist detached inside of the patient. it was located by use of c-arm imaging in the left para-colic gutter and was removed. on post op day 3 patient found to have pneumoperitoneum and ileus. repeat surgery resulted in partial colon resection and end sigmoid colostomy. | 0
5793645 | .><a laparoscopic harmonic scalpel tip broke off during the case, it was retrieved with no harm to patient. the distal end of white insert was noted to be dangling. the instrument was removed and saved but the tip which was a few millimeters in size was misplaced but i did secure the rest of the instrument for pick up by risk. the patient did not suffer any harm and surgery progressed as expected. | 0
5390310 | .><a portion of this case was a transoral robotic surgery (tors) using the davinci robot. the 8mm permanent cautery spatula was used and it is thought that the entire bovie piece was conducting; however, it was not noticed until the end of that portion of the case that a burn was found in the patient's posterior right cheek area. the instrument was plugged into the triad bovie machine and the setting was 25/25. | 0
5440727 | .><da vinci sheath tip was burned away on the proximal part of the shaft. this was noticed when instruments were being cleared off the field. doctor noted no harm to patient; however, surgeon will monitor patient. | 0
5496627 | .><failed self check, did not work. the staff had to get another one to complete the case. | 0
5842740 | .><vessel sealer malfuntioned after 5 minutes of use. | 0
5903130 | .><the vessel sealer took an excessively long time to coagulate. we have encountered similar problems in the past. the device was removed from the field and another sealer was used. | 0
5619291 | the jaws of the pk were staying open. a piece of the wire that controls the gears for the pk was sticking out of the instrument.><. | 0
5750555 | .><while doing a procedure using the robot xi, the stapler was introduced it fired once. when the stapler was reintroduced an error occurred with the which lead to a call to servicing where we were instructed to remove all robotic and laparoscopic instruments and shut down the entire system and reboot. after rebooting the system, the same error occurred when using the stapler necessitating a second shut down of the system. we then had to abort the robotic portion of the procedure and open different instruments leading to increased patient cost. we were able to complete the procedure without injury to patient. note that the intuitive representative was present during the robotic portion of the procedure. follow up: intuitive had installed a software update on the xi robot on 05/04/2016. system was tested and verified as ready for use by intuitive. it was after this update that we had the fault. the intuitive repair technician came and replaced part 653140-07. we have not had any further problems since the part was replaced. | 0
5392805 | .><doctor was performing total laparoscopic hysterectomy single site. arm 1 -fenestrated bipolar forceps, arm 2 - permanent cautery hook and camera arm in the middle. camera arm collided with arm one cannula causing cannula mount to open. as a result, the fenestrated bipolar forceps bent. the forceps bent more than usual. | 0
5735248 | a 2nd 45mm stapler load was being fired and the blade would not retract, therefore the stapler would not open. the davinci system faulted and the alarm went off. the rep was immediately called while the scrub tech used the davinci emergency key to unlock" or "open" the endowrist stapler arm. this instrument was then able to be removed and the alarm stopped.><." | 0
5549794 | .><vessel sealer was used and then alarmed that the blade was exposed. manufacturer response for robotic vessel sealer, (brand not provided) (per site reporter): the product was returned to the manufacturer for evaluation.><.><vessel sealer was used and then alarmed that the blade was exposed. manufacturer response for robotic vessel sealer, (brand not provided) (per site reporter): the product was returned to the manufacturer for evaluation. | 0
6016500 | during a surgical procedure the prograsp forceps broke and a wire was hanging out of the instrument.><. | 0
5446453 | xi davinci fenestrated bipolar missing 1mmx1mmx3mm plastic piece of insulation on tip of instrument - unknown if lost in pt or before inserted in pt. in this case, there were no visible pieces of the plastic covering seen in patient. the missing portion was noted after it was removed since the jaws of the instrument would not operate properly. due to the small size; however, it could not be determined if it was removed throughout the course of normal irrigation and suctioning during the surgery.><. | 0
5602938 | .><robotic needle driver would not open and close. | 0
5426850 | .><the vessel sealer worked at beginning of procedure then it would not cut. had to open a new device. new vessel sealer functioned properly. no harm to patient. | 0
5426851 | .><vessel sealer was opened and used for davinci tlh. device was not cutting and cauterizing appropriately. not reprocessed. new vessel sealer was opened and functioned properly during case. no harm to patient. | 0
6069006 | .><during a thorascopy procedure, two small pieces of material were found in the patient's thoracic cavity. the spatula was examined and an area of coating was found missing near the metal tip. the device and the loose pieces of material were removed from the field. | 0
6050335 | during the case, the mega needle driver cable broke. no harm to pt. instrument had five lives left on it.><. | 0
5466175 | during robotic assisted inguinal hernia repair, the cadiere forceps wire broke while in use in the patient's abdomen. no injury to pt. instrument replaced and procedure continued and completed.><. | 0
5728655 | .><the harmonic tip broke off during use in the patient's abdomen. the piece was recovered. | 0
5829117 | .><while using the robotic micro scissor with cautery, physician noticed an arc come through the end of the instrument. patient experienced a burn on the colon. instrument removed and replaced. burn noted on tip cover. | 0
6085505 | a piece of the maryland bipolar forcep for the davinci xi robot broke off while in use. piece fell into surgical site, but was successfully retrieved. a piece of the yellow protective plastic also broke off and fell inside the patient - that small piece was also recovered. instrument has been taken out of service. manufacturer response for robotic forcep, maryland bipolar forceps (per site reporter). replacement will be provided.><. | 0
5774739 | .><intuitive vessel sealer #1 showed error at the beginning of case; had used 2-3 times. vessel sealer #2 would seal but not cut. both devices used with davinci machine. intuitive hooked up to allow company to satellite in; called, and they stated no error shown from their end but they would document the call and concern. | 0
5867967 | .><while using the endowrist 45 stapler, after several times of the instrument doing its clamp and unclamp function prior to stapling, it went into fault mode-stating on the monitor that it failed to unclamp. the stapler 45 emergency unclamp kit was used per directions on the monitor. the instrument jaws opened and it was removed from the patient without incident. the instrument had 1 life left. davinci stat" was called and rep notified." | 0
5528573 | .><during a robotic left inguinal hernia repair, the mega suture needle driver cable broke. no patient harm. instrument was removed and replaced in order to finish the case. | 0
6016540 | harmonic scalpel tip broke while performing robotic hysterectomy; tip was removed from the patient.><. | 0
5426845 | vessel sealer was working. during the procedure it quit and an error came up on screen of machine. it stated not compatible with software. not reprocessed. opened new vessel sealer. turned machine off and back on and new vessel sealer functioned properly. no harm to patient.><. | 0
5811612 | vessel sealer malfunctioned during robotic case. alarm was set off that the instrument was not working. upon inspection, noted blade unable to retract back to its normal position when used to cut tissue. no parts of instrument were left inside patient. instrument is intact of its parts. replaced with a new vessel sealer. manufacturer response for endowrist one single use instrument vessel sealer (si), endowrist one single use instrument vessel sealer (per site reporter): numerous reports re these devices but no response from mfg.><. | 0
5528561 | malfunctioned, gave recoverable fault.><. | 0
5789668 | vessel sealer was in use and working without incident. it was removed so that another instrument could be used. the vessel sealer was reinserted and a recoverable fault error message alarmed upon insertion. the fault was cleared, vessel sealer machine was rebooted and device reinserted. the fault occurred again so another vessel sealer was opened and worked with out further incident. there surgeon was present and there was no patient harm and there was only about a two minute delay in the procedure.><. | 0
5344434 | telemetry monitor for pt was on pt, when pt complained that monitor hurt and was hot. after removing monitor which was very hot where batteries are placed, both the inside and outside of monitor was hot. plastic covering on battery was melted and very hot and smelled like burning plastic. | 1
5646247 | while on telemetry, the pt was found unconscious and in cardiac arrest. visible or audible alarms not noted at the telemetry system's central station. the telemetry system's central station was under observation by a monitor technician. when a code was called and the pt was disconnected from the telemetry pack and connected to an external ecg display, the initial rhythm wa thought by the medical staff to be ventricular tachycardia or asystole. resuscitation was not successful. prior to the event, the pt-specific parameters set on the telemetry system on arrhythmia detection, as will as visible and audible alarms, included: ventricular tachycardia - rate 165 bpm and sever bradycardia - rate 45 bpm. several days after the event, an attempt was made by the hospital and the manufacturer to retrieve the telemetry system's waveform and event reports for this pt's event. the records for this pt were no longer stored in the telemetry system's memory. several days after the event, the telemetry system was tested by the hospital bioengineer. using a simulator test of sinus rhythm at rate of 30 bpm, the telemetry central station did produce auditory and visual alarms for a severe bradycardia" event." | 1
5905957 | gasket came loose in the part of the lead set that connects to the telemetry pack. this caused failure of the pt's ekg signal to transmit to the central station. telemetry pack worked fine with a new lead set. | 0
6022158 | spacelabs received a report on (b)(6) 2016 at 7:30am that an episode of bradycardia did not alarm at the telemetry central station. no injury was reported as a result of this event.><onsite testing of the involved devices conducted by a spacelabs field service engineer confirmed that the equipment performed to specifications. the testing was witnessed by a facility staff member. spacelabs has launched an investigation into this event and will file a supplemental report once the investigation is complete.><the patient's historical waveforms and trend information were reviewed by a spacelabs lead software engineer. findings show that at approximately 5:38am, the low hr alarm limit had been lowered to 40 bpm. a heart rate of 56 bpm would not have generated an alarm due to the patient-specific alarm limits set by the clinical staff. the device performed to specifications. this report is considered final and the issue is closed. | 1
5440640 | approximately 35 telemetry transmitters have heated enough for nursing to have concerns for patient burns. upon investigation by biomed, some units found to have heat damage to the battery and casing. these events started approximately in (b)(6) 2015 at which time the manufacturer was notified. we are aware of one employee that has been burned when removing the batteries. reason for use: cardiac condition. | 0
5513425 | telemetry box became very hot, nearly burning the patient. the nurse was able to discontinue the use of this telemetry box prior to the patient being injured. through our investigation, we have learned these telemetry boxes have a tendency to become very hot so much they begin to smoke. dates of use: (b)(6) 2016. event abated after use stopped or dose reduced: yes. | 0
6186515 | on (b)(6) 2016, (b)(6) hospital reported that they have had an incident of skin issues with the stx pads that occurred last week. the application was for post-cardiac arrest with a target temperature of 33c. an unresponsive (b)(6) male with no known past medical history (that they were aware) was admitted after ventricular fibrillation arrest in car, girlfriend called emergency medical services (ems). patient was resuscitated and taken to (b)(6 hospital. on (b)(6) 2016 at 1800, the patient was placed on stx surface wraps for application of hypothermia post-cardiac arrest; target temperature of 33c; bath temperature not documented throughout treatment. the patient was on propofol and was not given any neuromuscular-blocking drugs (nmbs). the patient was on levophed (norepinephrine) drip (b)(6): nursing staff titrating drip; max was 5mcg/min. it was documented that there was no skin breakdown for the remainder of (b)(6) 2016 until 16:00. on (b)(6) 2016 at 16:00, the patient was taken off of stx system/stx pads and intravascular catheter was placed (clinical contact is unable to provide an explanation as to why it took them so long to place the catheter and/or why it was chosen to do ivtm versus surface at that point); specific catheter placed was not documented.- on (b)(6) 2016 at 18:00, the day shift registered nurse documented that the skin had no breakdown. on (b)(6) 2016 at 20:00: the night shift registered nurse noted blisters on back". patient was seen by wound care registered nurse on (b)(6) 2016: "multiple open blisters mid to upper back consistent with 2nd degree burn"; recommended medi honey and hcs patches for treatment. wound care registered nurse continued to see patient throughout the rest of his hospital stay; blisters did resolve and no further treatment was required. patient was transferred out of hospital to rehab center with cpc of 1-2.><the stx pads in complaint will not be returned for investigation. therefore a physical investigation will not be performed. note: skin injuries can occur with any surface cooling devices even when following product instructions for routine skin evaluations and in patients without a prior history of skin problems. even though skin injury is an anticipated complication for surface-cooling the risk can be mitigated by following ifu. an international guideline published by national pressure ulcer advisory panel (npuap) and european pressure ulcer advisory panel (epuap) categorizes pressure ulcers into several stages of severity but common strategies include repositioning the patient on a recurrent basis to avoid long exposure to pressure points (i e. alternating patient positions between the patient's left-side right-side and center every 2 hours) performing comprehensive skin assessments frequently documenting the risk assessments on a prediction scale (e.g. braden scale) and communicating regularly regarding the status of skin monitoring." | 0
6205091 | ><transvaginal probe came apart during exam. parts were contained in condom cover. no harm to patient and no retention of any piece. | 0
5996507 | ><the patient was supine on the operating room table secured with safety straps at the hip and bilateral lower extremity. the patient was placed in slight reverse trendelenburg and tilted towards both surgeon and scrub tech. the patient slid to side onto both surgeon and scrub tech. patient pushed back onto the bed and transfer pad removed immediately. patient assessed and procedure resumed. | 0
5386110 | during review of run data files (rdfs) for an unrelated customer request, terumo bct found an incident of donor data entry error. in the rdf, it was noted that the weight of a donor had been entered incorrectly. no medical intervention was required for this event. actual weight: (b)(6) full donor (b)(6).><investigation: per the rdf analysis, the procedure that was run with incorrect donor weight resulted in ~17%tbv being collected. the ac infusion rate was calculated to be ~1.32ml/min/ltbv, which is > the trima maximum of 1.09ml/min/l tbv. no adverse events occurred during or after the donation and the donor has donated since this incident. a one year review of the device's service history was performed with no issues noted related to the reported condition. the terumo bct clinical specialist provided feedback to the customer regarding the implications of incorrect data entry of donor information and how it can affect available procedures. corrective action: an internal capa has been initiated to address data entry errors. root cause: the root cause of the incorrect donor weight was an user interface issue. | 1
5796124 | the customer reported that the collection set was failing the pressure test during set up. the customer returned the set for evaluation. upon evaluation, it was noted that the clamp was not fully occluding the tubing, allowing air into the sample bag, which has the potential to be returned to the donor. there was no donor connected at the time of the pressure test alarms. patient (donor) information is not available at this time.><no donor was connected at the time of the event. no patient (donor) information is reasonably known.><investigation: an unused set was returned for investigation. the set was visually inspected for kinks, occlusions, missing parts, mis-assembly and leaks which may have contributed to the pressure test alarm, none were found. the set was loaded onto the trima machine in the lab with the clamps in the position as received by the customer (white pinch clamp on the sample bag line closed). the trima machine alarmed, pressure test error". it was noted that the sample bag was full of air. closer inspection of the clamping on the inlet coil showed that the white pinch clamp going to the sample bag was not fully occluding the tubing as the tube was off-center within the clamp. the clamp was opened the air from the sample bag was released and the clamp was re-closed. the pressure test was restarted and the set passed the pressure test and made it to the next process stage. investigation is in process. a follow-up report will be provided.><investigation: a review of the device history record (dhr) for this unit showed no irregularities during manufacturing that were relevant to this issue. the run data file (rdf) was analyzed for this event. root cause: review of the returned set suggests that the white clamp did not fully occlude the tubing line. the signals in the run data file indicate that the disposable test took longer than expected during the disposable test substates and the trima accel system operated as intended by displaying a pressure test error alert. while this does not indicate a conclusive cause for air in the sample pouch or return line the system running the pumps longer in order to achieve the specific criteria can indicate that the tubing set may not have been clamped or loaded correctly." | 0
5661026 | investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher-than-expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected wbc content in the platelet product could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell (rwbc) testing, therefore no patient information is reasonably known at the time of the event. donor unit #(b)(4). the disposable kit is not available for return, because it was discarded by the customer. | 0
5393006 | investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file (rdf) was analyzed for this event. root cause: a definitive root cause for the observed leuko reduction failure remains undetermined at this time. the signals in the run data file indicate that the higher than expected wbc content in the platelet product was likely a result of an escape of wbcs from the leuko reduction system chamber during a portion of the procedure. based on the available information, it cannot be ruled out that the higher than expected wbc content in the platelet product could be donor related.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual wbc testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(4) the disposable kit is not available for return because it was discarded by the customer. | 0
5853274 | the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual wbc testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(4). the disposable set is not available for return, because it was discarded by the customer.><investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file (rdf) was analyzed for this event. root cause: a definitive root cause for the observed leuko reduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher-than-expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected wbc content in the platelet product could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product. | 0
5890274 | investigation is in progress. a follow-up report will be provided.><the customer reported using saline solution instead of anticoagulant citrate dextrose solution (acda) during a platelet collection procedure. at 42 minutes into the procedure, it was noticed the saline solution bag had been spiked instead of the acda bag. the collection was stopped immediately and no rinseback was given. the donor was kept at the collection site and the medical director was contacted. the medical director gave instructions to the donor on what signs and symptoms the donor should look for, and to call 911 if any occurred. the donor was given fluids and snacks at the collection site. the customer stated the donor is fine and no medical intervention was required for this event. during a follow-up with the donor from the customer later that day, the donor stated he felt well. the disposable set is not available for return, because it was discarded by the customer.><investigation: the disposable set was unavailable for return. a review of the procedure confirms that the run duration was 42 minutes and was not ended prior to first return cycle. a review of the device history record (dhr) for this unit showed no irregularities during manufacturing that were relevant to this issue. root cause: based on the customer's report, the root cause for the incident is that the user spiked saline instead of acda. correction: terumo bct has followed up with this customer to provide feedback on the reported condition. | 0
5733447 | investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. root cause: this disposable set was unavailable for specific root cause analysis. a definitive root cause for the observed leukoreduction failure remains undetermined at this time. possible root causes were provided in the initial report for this event.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual wbc testing, therefore no patient information is reasonably known at the time of the event. donor unit #(b)(6). the disposable set is not available for return, because it was discarded by the customer.><investigation: the run data file (rdf) was analyzed for this event. the signals in the run data file indicate that the higher than expected wbc content in the platelet product was likely a result of an escape of wbcs from the leukoreduction system (lrs) chamber at the end of the procedure. based on the available information, it cannot be ruled out that the reported results could be donor related. investigation is in progress. a follow-up report will be provided. | 0
5792095 | the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell (rwbc) testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(6) the disposable kit is not available for return because it was discarded by the customer.><investigation: the collection set was unavailable for return and evaluation. the run dta file(rdf) was analyzed for this event. the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. root cause: the signals in the run data file indicate that the plasma line was likely occluded during a portion of the procedure. if the plasma line does not contribute properly to the platelet pump it is possible that the flow through the lrs chamber could be higher than expected by the system. this could allow some wbcs to escape and contribute to the higher than expected wbc content reported for this collection. orientation of the hex in the hex holder may contribute to the above. possible causes for the plasma line occlusion include but are not limited to: pinching of the tubing due to the hex positioning as mentioned above, a significant kink in the plasma line tubing in the centrifuge. | 0
5511136 | full reporter's email address: (b)(6): the customer stated they have completed their own internal investigation, and the issue was determined to be donor related and they are now flagging all products from this donor for product white blood cell residual testing. investigation is in process. a follow-up report will be provided.><the customer reported an elevated white blood cell (wbc) count in the collected platelet product. there was no donor or transfusion recipient at the time of the residual wbc testing, therefore, no patient information is reasonably known. the disposable kit is not available for return, because it was discarded by the customer.><investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher than expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trimaaccel system operated as intended. based on the available information, it cannot be ruled out that the reported results could be donor related. | 0
5980978 | investigation: the run data file (rdf) was analyzed for this event. the analysis of the run data file did not find a conclusive cause for the higher-than-expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected wbc content in the platelet product could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product. investigation is in process. a follow-up report will be provided.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual wbc testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(6) the trima platelet collection set is not available for return because it was discarded by the customer.><the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher-than-expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accelsystem operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected wbc content in the platelet product could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product. | 0
5442544 | the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell testing, therefore no patient information is reasonably known at the time of the event. donor unit (b)(6) the disposable kit is not available for return, because it was discarded by the customer.><investigation: the device history records were reviewed for this event. there were no events noted in the device history record that would have contributed to the elevated white blood cell count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher-than-expected white blood cell content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected white blood cell content could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product. | 0
5458204 | investigation: the run data file (rdf) was analyzed for this event. the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. root cause: a definitive root cause for the observed leuko reduction failure remains undetermined at this time. the signals in the run data file are consistent with some possible clumping/aggregation during platelet collection that could have affected the separation of cells in the lrs chamber. based on the available information, it cannot be ruled out that the clumping/aggregation and resulting leuko reduction failure could be donor related.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(4). the disposable kit is not available for return because it was discarded by the customer. | 0
5669815 | investigation: the device history record (dhr) was reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated white blood cell (wbc) count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher-than-expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the higher-than-expected wbc content in the platelet product could be donor-related. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual wbc testing, therefore no patient information is reasonably known at the time of the event. donor unit #(b)(6) the disposable kit is not available for return, because it was discarded by the customer. | 0
5581217 | the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell (rwbc) testing, therefore no patient information is reasonably known at the time of the event. donor unit #: (b)(4). the disposable kit is not available for return, because it was discarded by the customer.><investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leuko reduction failure remains undetermined at this time. the analysis of the run data file did not find a conclusive cause for the higher than expected wbc content in the platelet product reported for this collection. no unusual process variable was identified and the signals in the run data file indicate that the trima accel system operated as intended. based on the available information, it cannot be ruled out that the reported results could be donor related. | 0
5662295 | the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell (rwbc) testing, therefore no patient information is reasonably known at the time of the event. donor unit #(b)(6) the disposable kit is not available for return, because it was discarded by the customer.><investigation: the run data file was analyzed for this event. after reviewing the signals in the run data file, it cannot be ruled out that the number of access alerts that occurred throughout the procedure disrupted the steady-state of the system and contributed to the elevated wbc content measured in the platelet product. it also cannot be ruled out that a sampling, calculation, or other process error could have contributed to the higher-than-expected wbc content in the platelet product. investigation is in progress. a follow-up report will be provided.><this report is being filed to provide additional information. investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. root cause: this disposable set was unavailable for specific root cause analysis. a definitive root cause for the observed leukoreduction failure remains undetermined at this time. possible root causes were provided in the initial report for this event. | 0
5685018 | investigation: the device history records (dhr) were reviewed for this lot. there were no events noted in the dhr that would have contributed to the elevated wbc count experienced by the customer. the run data file was analyzed for this event. root cause: a definitive root cause for the observed leukoreduction failure remains undetermined at this time. the signals in the run data file indicate that the higher than expected wbc content in the platelet product was likely a result of an escape of wbcs from the lrs chamber at the end of the procedure. based on the available information, it cannot be ruled out that the reported results could be donor related.><the customer would like the run data file investigated to determine a possible cause for the elevated white blood cell (wbc) content in the platelet product. there was not a transfusion recipient or patient involved at the time of the residual white blood cell (rwbc) testing, therefore no patient information is reasonably known at the time of the event. donor unit #(b)(6) the disposable set is not available for return, because it was discarded by the customer. | 0
6156900 | ><it was reported that three patients each were injured on separate dates. this report represents the second patient injury, patient b. the ric5-9w-rs ultrasound probe, along with a voluson p6, biopsy needle ebt 00230s-0 (elios bio tek), and a biopsy guide of unknown type were being used to assist an ovarian biopsy for harvesting oocytes when the ovary was punctured and/or torn with the biopsy needle. surgery, celiography, was required in order to stop the internal bleeding. the user reported that the virtual guide for the needle was offset and therefore needle placement was incorrect. | 0
6080674 | on 10/14/2016 - facility reported: at the time of placement, healthcare professional achieved max p wave and was confident that the picc was in the correct position (the depth marking on the picc was at 0cm). a confirmation chest x-ray at the time of placement was not done. two weeks later, a chest x-ray was taken and incidentally found that the picc was in the right atrium. the physician pulled the picc back 6cm and another chest x-ray confirmed that the picc was now in the correct position at the caj. facility reported that during initial placement, the nurse was confident that she was in the correct place but now it seems that she was 6cm too deep.><the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard. the following were reviewed as part of this investigation: patient severity, frequency analysis, applicable previous investigation(s), sample (if available), applicable fmea documents, labeling, and applicable manufacture records. based on a review of this information, the following was concluded: the complaint of a malpositioned picc was inconclusive as the exact circumstances in the clinical setting were unknown and no potential cause was discovered during evaluation of the returned equipment. evaluation of the unit was performed by the service facility, (b)(4). the scanner was evaluated by the service facility with an in-house sherlock and ecg simulator and the scanner was found to function normally. after manufacturing review of dhr and manufacturing records, there is no evidence that would support the event being a manufacturer-related failure. the reported issue could not be reproduced and the root cause is inconclusive as they could not duplicate the clinical setting.><the information provided by bard represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant / reporter was unable or unwilling to provide any further patient, product, or procedural details to bard. the device has not been returned to the manufacturer, at this time, for evaluation. a lot history review (lhr) review is not possible, as no manufacturing lot number has been provided by the complainant.><on 10/14/2016 - facility reported: at the time of placement, healthcare professional achieved max p wave and was confident that the picc was in the correct position (the depth marking on the picc was at 0cm). a confirmation chest x-ray at the time of placement was not done. two weeks later, a chest x-ray was taken and incidentally found that the picc was in the right atrium. the physician pulled the picc back 6cm and another chest x-ray confirmed that the picc was now in the correct position at the caj. facility reported that during initial placement, the nurse was confident that she was in the correct place but now it seems that she was 6cm too deep. additional information from facility stated: the external catheter measurement was at 49cm marking. the patient's rhythm during insertion was a-fib but clear p wave elevation with internal ecg insertion. the facility stated the negative deflection was noted when catheter hubbed. i like to see the deflection, especially on a-fib patients. the line was retracted back to 0cm. it was reported that the facility does not rely on 3cg with a-fib unless they can clearly see the p wave elevation. | 0
5624472 | ><cryoablation probe was not well insulated, creating a dime size thermal injury on right lower back. this defect does not show up on the shaft during the test as it took a few minutes to see it once the probe was removed and retested away from the patient. company notified and asked what steps should be taken so this doesn't happen again. photos of probe obtained. manufacturer response for cryoprobe, endocare tm 1.7mm long shaft renalcryotm cryoprobe oblong ice (per site reporter): replacing at no charge, following up on their end with an investigation of the product. | 0
5932051 | ><the shaft of the cryoablation probe should be insulated to prevent ice along the shaft of the needle that could cause the patient harm. in fact we have had 3 probes that have done this now, 2 of which caused a thermal injury to the skin with skin breakdown that was treated with silvadene cream and healed. manufacturer response for cryo probe, endocare tm 1.7mm round ice per cryo probe (per site reporter): replacing at no charge, following up on their end with an investigation of the product. | 0
6170075 | ><evaluation in progress. | 0
5547522 | ><during the procedure the perc-17 probe gave the patient a dime sized thermal injury to the skin. | 0
5936622 | ><probe did not freeze during pre-test. | 0
6198087 | ><device was tested with simulated use testing and found pressure relief valve spring degraded due to age. device was repaired and returned to the customer. | 0
6170060 | ><during test mode, there was some shaft frosting but the freeze zone" did not go up the shaft so this finding was not appreciated. two minutes into first ablation cycle it was noticed that the circle of skin around the probe was white and frozen; this was not the case for the other probe. the probe was shut off. due to concern about getting an inadequate lung ablation after the first cycle the probe was turned on at 30% and later to 50% with assiduous application of warm saline to the skin to keep it warm. the subcutaneous tissue felt hard like it was freezing even with reduced flow. on follow-up call the patient has developed a blister at the site of the probe with the frosted shaft." | 0
5623608 | ><cryoablation device was checked out and found to be fully functional before placing the patient under anesthesia for the ct guided procedure. after the patient was anesthetized and intubated, the device was found to be non-functional. the procedure could not be performed and needed to be rescheduled. manufacturer response for cryoablation device, cryocare (per site reporter): the manufacturer has evaluated the device and found that a printed circuit board had failed. the manufacturer repaired the device and has made it available to be returned to service. | 0
5988891 | this spontaneous report was received from (b)(6) (b)(4) on (b)(6) 2016. additional information received on 07-sep-2016 was combined with the initial. a consumer reported that on an unknown date in 2004 or 2006 after the female patient started the breathtek ubt kit (breathtek ubt kit), (unknown dose by mouth, one time dose), for helicobacter pylori urea breath test, she immediately experienced a severe allergic reaction which required her to go to the emergency room. while she was in the emergency room she was given an adrenaline shot;" after a few hours this "severe allergic reaction" resolved. this spontaneous report was received from (b)(6) (b)(4) on (b)(6) 2016. additional information received on 07-sep-2016 was combined with the initial. the patient's medical history included irritable bowel syndrome gluten sensitivity allergy to cow's milk allergy to grains allergy to peanuts and tree nuts allergy to whey protein allergy and other products. she was on "other unknown concomitant medications." diagnostic tests revealed the following results: unspecified test administered to determine lactose intolerance. she did not finish the breathtek test as she had the allergic reaction prior to the final sample collection. the consumer could not say for certain if the test had definitely caused her allergic reaction. her physician wants to administer a breathtek ubt kit in the near future. the outcome of severe allergic reaction was recovered/resolved. the consumer could not say for certain the previous testing had definitively caused her allergic reaction. otsuka causality assessment: the reported event allergic reaction is considered as related to breathtek ubt kit given the temporal association and positive dechallenge however the patient also had significant history of gluten sensitivity allergy to cow's milk allergy to grains allergy to peanuts and tree nuts allergy to whey protein allergy and other products which could have also contributed to the reported event. the information provided by the reporter is also not clear as it was reported the event happened in 2004 or 2006 and the consumer could not say for certain if the previous testing caused allergic reaction which precludes a meaningful medical assessment. otsuka does not consider the information contained in this individual case safety report to justify any change to the benefit-risk balance of breathtek ubt kit or to justify any safety measures." | 0
5341268 | ><pulled out of service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5341269 | ><ge ivent had a service error 705, and fan noise. ge required ventilator be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator started making a fan noise and had a service error code 705. pt was transferred to another ventilator. vent was picked up at pt's home and sent to ge for repair. ventilator had recently been repaired because the ventilator had a blank screen. | 0
5341256 | ><while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pulled out of service. pt was transferred to another ventilator. vent was sent in for repair. | 0
5341255 | ><while pt on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pulled out of service. pt was transferred to another ventilator. vent was sent in for repair. | 0
5341266 | ><while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pulled out of service. pt was transferred to another ventilator. vent was sent in for repair. | 0
5341261 | ><while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pulled out of service. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340952 | ><ge ivent had an error code 706 - pneumatic module issue. ge required ventilator be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, it started alarming with error code 706. pt was transferred to another ventilator. vent was sent in for repair on (b)(6) 2015 and it was returned after 4 months at the ge repair center, and out of service for our patients. vent had to be returned for more repairs because the ventilator is now showing a negative number of hours on the unit and battery light on ventilator is not working to show the status of the battery. ventilator was shipped back to ge on (b)(6) 2015. | 0
5340948 | ><ge ivent had an issue where the screen went blank. ge required ventilator be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340923 | ><ge ivent had an issue where the screen went blank. unable to monitor pts settings with a blank screen. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator screen blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340917 | ><ge ivent had an issue where the screen went blank. unable to monitor pts settings with a blank screen. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340914 | ><ge ivent had issue where the screen went blank. unable to monitor pts settings with a blank screen. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340908 | ><ge ivent had an error code 716 - fan malfunction. pt was required to be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, started making a grinding noise with error code 716. pt was transferred to another ventilator. vent was sent in for repair on (b)(6) 2015 and has not returned after 3 months at the ge repair center, and out of service for our pts. | 0
5340941 | ><ge ivent had an issue where the screen went blank. ge required ventilator be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5340937 | ><ge ivent had an issue where the screen went blank. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. ventilator was sent back and would not pass ovt. unit had to be sent back for more repairs. | 0
5340930 | ><ge ivent started making a grinding noise from blower motor. ge required ventilator bbe sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of service. while pt on ventilator, the ventilator started making a grinding sound and a whistling noise when cycling. pt was transferred to another ventilator. vent was picked up at pt's home and sent to ge for repair. ventilator was sent back to us and had an error code 706. unit had to be sent back for more repairs. this problem started on (b)(6) 2015 and had to be sent back to ge for repair three (3) times and the problem is not corrected. we just got this vent back from the ge repair center on (b)(6)2015, and the ventilator screens are not functional, and therefore cannot be placed safely on a pt. this ventilator has been out of service for over 10 months. | 0
5340929 | ><ge ivent had an issue where the screen went blank. ge required ventilator be sent in for repair. pt was required to be transferred to a back up ventilator and primary ventilator was pulled out of the service. while pt was on ventilator, the ventilator screen went blank, no settings could be seen to monitor pt. pt was transferred to another ventilator. vent was sent in for repair. | 0
5492266 | patient on telepak which may have been precipitated by lead" off problem with patient. upon review of waveforms in alarm review - did not match what reporter noted in 'real time'. patient required cpr/resusitation and was sent to the emergency to the cardiac cath lab for cardioversion - which was successful." | 0
5492265 | patient on telepak went into lethal arrhythmia where he fell in the bathroom. the system failed to alarm. the arrhythmia was unable to see the arrhythmia in real time on the central monitoring station. patient was resuscitated and sent to the cardiac cath lab and readmitted to cicu. the strip was later able to be retrieved which showed vtach/vfib (notable difference upon review in alarm review vs wave review. when vfib alarm event was revisited in alarm review wave forms would not record- showed single flat line. this occurred with another patient on (b)(6) 2005. | 1
5454767 | as part of a change of ehr at this institution, a new pacs system was introduced replacing one that was functioning well. none of the hospital it experts, nor experts from the vendor, checked the computers to determine if the imaging system's launcher could be loaded. as it turns out, innumerable computers do not have the storage to load the thin client launcher and thus, no radiograph images could be seen on these systems linked to the hospital's ehr. this has adverse impact especially at sites that have one or two computers causing delays in care and diagnosis. this is especially dangerous in pts with cardiopulmonary disease and heart failure or pneumonia. there has been no remedy put forth and no one has assumed leadership to solve the dysfunction. | 1
5459462 | the customer reported that the bsm (bedside monitor) lcd screen has gone out but is still responsive to the touchscreen inputs.><the customer reported that the bsm (bedside monitor) lcd screen has gone out but is still responsive to the touchscreen inputs. there were no patient harm. the patient was safely monitored from the cns (central monitoring system) until the monitor was swapped with another one. the device was returned to nihon kohden, evaluated, and the reported issue was confirmed. the device arrived with minor touchscreen damage. the inverter board was replaced to fix the issue and all functions were tested prior to shipment. the repaired device was returned to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. | 1
5496679 | the biomedical engineer reports that the staff noticed that the bedside monitor display showed all waveforms stop and restart during a case.><the biomedical engineer reports that the staff noticed that the bedside monitor display showed all waveforms stop and restart during a case. the staff power cycled the device and did a delete data" however the issue still remained. the device was switched out with a working device. by the time the biomedical engineer had a chance to test the device there were no issues. biomedical engineer was asked to check the connection between the input unit and the host monitor. the customer was instructed to contact nihon kohden after additional troubleshooting. the customer reported that they found no issues. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available. device not returned." | 1
5920036 | the customer reported that the bsm (bedside monitor) screen is black (blank) when it is turned on. the sounds can be heard in the background. the device was returned to nihon kohden, evaluated, and the reported issue was confirmed. the inverter board was replaced to resolve the issue. an evaluation also found a scratched touchscreen and a broken front enclosure. the unit was tested per the operator's/service manual and the results were recorded on the maintenance check sheet. the unit completed in 24 hours of extended testing and operates to manufacturer's specifications. the repaired device was returned to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 if additional information becomes available.><the customer reported that the bsm (bedside monitor) screen is black (blank) when it is turned on. the sounds can be heard in the background. | 1
5415349 | the customer reported that the display on the bsm (bedside monitor) is blank. the biomedical engineer switched the cables and the problem still persists.><manufacturer narrative: the customer reported that the display on the bsm (bedside monitor) is blank. the biomedical engineer switched the cables and the problem still persists. the customer was sent a replacement monitor. the unit was returned and evaluated. the reported problem was duplicated. unit has been scrapped, as it was not worth repairing. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the display on the bsm (bedside monitor) is blank. the biomedical engineer switched the cables and the problem still persists. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available. | 1
5743891 | the customer reports that the screen went black and there is no display. the unit functions with no display. the customer powered down and up, removed battery with the same result.><manufacturer narrative: the customer reports that the screen went black and there is no display. the unit functions with no display. the customer powered down and up, removed battery with the same result. the unit was evaluated and the customer's complaint of no display was confirmed. the lcd inverter was replaced to resolve the issue. the unit was tested and completed all steps in the maintenance check sheet per service manual. the device has been repaired and returned to the customer. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reports that the screen went black and there is no display. the unit functions with no display. the customer powered down and up, removed battery with the same result. nihon kohden continues to investigate the reported event. | 1
5909495 | the customer reported that the bedside monitor (bsm) shut down during a case. the unit will not fully boot up and only gets to the please wait" screen. the unit will not power off when the power button is pressed or even when the power cord is pulled from the wall outlet. nihon kohden continues to investigate the reported event. nihon kohden will submit a supplemental report in accordance with 21 cfr section 803.56 when additional information becomes available.><manufacturer narrative: the customer reported that the bedside monitor (bsm) shut down during a case. the unit will not fully boot up and only gets to the "please wait" screen. the unit will not power off when the power button is pressed or even when the power cord is pulled from the wall outlet. the unit has been cleaned and evaluated and the reported issue was duplicated. the customer has been sent an exchange unit. nihon kohden will submit a supplemental report in accordance with 21 cfr part 803.56 if additional information becomes available.><the customer reported that the bedside monitor (bsm) shut down during a case. the device will not fully boot up and only gets to the "please wait" screen. the device will not power off when the power button is pressed or even when the power cord is pulled from the wall outlet." | 1
5597181 | (b)(4). baxter received and evaluated the device. the device was found out of specification in relation to the reported heat damage, which was observed during physical inspection. evaluation observed damage to the front case from excessive heat due to radio printed circuit board (pcb) / flex assembly failure caused by fluid intrusion. this failure mode renders the unit unrepairable and therefore, no parts were replaced. (b)(4).><it was reported that a spectrum pump had burn/heat damage. it was also reported there was no patient involvement. | 0
5534514 | (b)(4). the device was not returned to baxter for evaluation. therefore, an evaluation could not be completed. if the device is returned, an evaluation will be completed and a supplemental report will be submitted.><it was reported that a spectrum pump powered off without user input during therapy (medication, programmed amount and delivery rate unknown). the customer also stated that the device was swapped out and that there was no patient injury. all the information regarding the patient and event reported to baxter by the customer has been reflected in this report. | 0
5586673 | (b)(4). baxter received and evaluated the device and found it was out of specification in relation to the heat damage which was observed during evaluation. visual inspection found heat damage to the positive battery contact pins on the rear case assembly and warping on the rear case in the area on the backflex assembly along with heat damage to positive battery contact pads of the wireless battery module caused by fluid intrusion. the wireless battery module over heating caused the plastic to melt on the rear case assembly on the pump and also due to a higher current draw also caused q5 on the backflex (pump) to overheat resulting in the rear case warping. the wireless battery module is not serviceable and requires replacing. (b)(4).><during baxter's evaluation of a spectrum pump, it was observed the wireless battery module had heat damage. it was also reported there was no patient involvement. | 0
5722919 | investigation based on the return of the workstation castor to the manufacturer. the workstation was manufactured in 2006. the most probable cause identified by the investigating engineer was that the castor washer inside diameter was not sufficient to accomodate the castor stem. the inside diameter of the washer was increased by 0.3mm following a design review in 2007><the olympus wm-np1 series mobile workstation is intended for use in medical facilities under the direction of a trained physician and has been designed to be used with a range of olympus equipment to facilitate gi endoscopy, endoscopy ultrasound, respiratory and surgical endoscopic procedures. keymed (medical & industrial equipment) ltd has been made aware of an event which took place in (b)(6). the threaded stud of the right rear castor of the workstation broke when the workstation was manoeuvred. the workstation was caught by healthcare facility staff. there is no report of injury to patient or users and this report is submitted in an abundance of caution.><device was manufactured in 2006. olympus (b)(4) plan to return the subject device to the manufacturer - keymed (medical & industrial equipment) ltd for investigation. keymed have requested a copy of the maintenance records from the healthcare facility and any service events at the olympus repair centre. device awaiting return of device. | 0
5535536 | omnicell in anesthesia area was set to reboot. began to emit a considerable amount of smoke. fire not seen but suspected. omnicell removed from operating room, no patient in room. technician removed ebox. | 1