## 2008
1186743 | A HEAD CT PERFORMED UNDER ADVERSE CONDITIONS (COMBATIVE PT, NEEDED TO BE RESTRAINED) ON THE SIEMENS CT SCANNER. THE IMAGE IN QUESTION WAS PERFORMED HELICALLY ON THE SIEMENS CT SCANNER WITH RECONSTRUCTION. IT WAS POST-PROCESSED IN THE OPPOSITE DIRECTION OF NORMAL RECONSTRUCTION (I.E. HEAD TO TOE RATHER THAN THE MORE COMMON PROTOCOL FROM TOE TO HEAD). THE IMAGES WERE SENT TO PACS IN THIS FORMAT AND PACS CORRECTLY RE-ORIENTED THE IMAGES FOR CORRECT VIEWING. GE BELIEVES THERE IS SOMETHING IN THE WAY THE GE WEB PRODUCT IS READING THIS STUDY AND MAY RANDOMLY DISPLAY SELECTED IMAGES FLIPPED IN THE MAIN WINDOW. GE HYPOTHESIZES THIS IS CAUSED BY AN UNKNOWN PRESENTATION STATE DISPLAY FLAG (COMES FROM SIEMENS CT SCANNER AND IS INHERENT IN THE IMPROPERLY DEVELOPED STUDY). STEPS HAVE BEEN PUT IN PLACE TO ENSURE PATIENT SAFETY WHILE THE INVESTIGATION CONTINUES (I.E. RECONSTRUCTIONS ONLY TO BE PERFORMED IN THE COMMON TOE TO HEAD DIRECTION). | 1
1034051 | DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AT THIS TIME. >< IT WAS REPORTED THAT THE TUBING CONNECTS TO THE VAMP WAS OBSERVED TO BE LEAKING BEFORE USE. REPORTEDLY, THE LEAKAGE WAS NOTED AT THE BONDED AREA OF THE STOPCOCK FOR THE VAMP. | 0
1000346 | PATIENT REPORTS SUSTAINING LINEAR SUPERFICIAL ABRASIONS ON BILATERAL ARMS THAT SHE CLAIMS WERE CAUSED BY NEW BLOOD PRESSURE CUFF. TO BE USED ON PATIENT MONITORS. | 0
1000042 | CUSTOMER REPORTED THAT G7 AUTOMATED HPLC ANALYZER REPORTED A HA1C RESULT OF 4.2%. RESULT WAS QUESTIONED BY PHYSICIAN; REPEAT RESULT WAS 6.5%. | 0
952468 | THE PT ADMITTED TO THE NEURO INTENSIVE CARE UNIT AFTER SUSTAINING A TRAUMATIC BRAIN INJURY. THE PT WAS PLACED ON THE ARCTIC SUN FOR FEVER CONTROL AND ICP MANAGEMENT. NO SKIN ISSUES NOTED DURING OR AT THE CONCLUSION OF THE ARCTIC SUN USE. ON THE FOLLOWING DAY, THE PT WAS NOTED TO HAVE CHANGES IN SKIN COLOR (REDDENED/PURPLE) ON THE RIGHT AND LEFT ABDOMEN AND THE TOP OF THE RIGHT THIGH. A SMALL AREA OF EPIDERMIS ON THE THIGH BLISTERED WHICH IS HEALING WITH CONSERVATIVE TREATMENT. THE TREATMENT EXTENDED THE PT'S HOSPITALIZATION. | 0
1014448 | DURING THE EARLY MORNING OF A SHIFT, THE CENTRAL MONITOR FOR A NURSING UNIT MALFUNCTIONED: SCREEN SHOWED NO COMMUNICATION IN VARIOUS ROOMS. CLINICAL ENGINEERING WAS UNABLE TO IDENTIFY CAUSE. | 1
1067720 | THE CUSTOMER STATED A PATIENT KNOWN TO BE HCV POSITIVE, GENERATED NONREACTIVE RESULTS ON THE AXSYM HCV V3.0 ASSAY. IN 2008, THE PATIENT YIELDED NONREACTIVE RESULTS OF 0.95 AND 0.94 S/CO ON THE AXSYM HCV V3.0 ASSAY. ON THREE MONTHS LATER, THE PATIENT GENERATED REACTIVE RESULTS OF 1.03 AND 1.14 S/CO. THE SAMPLE FROM MARCH WAS RETESTED AND GENERATED REACTIVE RESULTS OF 1.24 AND 1.08 S/CO. THE PATIENT WAS KNOWN TO BE HCV POSITIVE BASED ON TEST RESULTS FROM ANOTHER LABORATORY (METHOD UNKNOWN). NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1133447 | THE CUSTOMER REPORTED THAT A PATIENT WAS UNINTENTIONALLY REMOVED FROM A SLOT ON A CIC, THEN REAPPEARED APPROXIMATELY 10 MINUTES LATER. DURING THE 10 MINUTE TIME, A PATIENT HAD A HIGH HEART RATE AND THE CIC DID NOT ALARM AUDIBLY. NO PATIENT DEATH OR INJURY OCCURRED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
984133 | THE CUSTOMER REPORTED EXPERIENCING HOSPITAL-WIDE DROPOUT ON TELEMETRY AND HARDWIRE BEDS. PATIENT WAVEFORMS DISAPPEARED, AND PERIODS OF LOSS OF MONITORING OCCURRED FOR APPROXIMATELY 16 HOURS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
964616 | COMPLAINT WAS RECEIVED FROM OUTSIDE THE USA. DURING ROUTINE CLEANING OF THE SYSTEM, THE LABORATORY TECHNICIAN REPORTEDLY CUT THIS THUMB ON A METAL PLATE THAT IS EXPOSED TO BIOLOGICAL FLUIDS. NO INFORMATION HAS BEEN RECEIVED TO SUGGEST ANY ILL EFFECTS ASSOCIATED WITH THE CUT. ACCORDINGLY, THIS EVENT IS BEING FILED IN A EXCESS OF CAUTION. >< INVESTIGATION OF THE AFFECTED SYSTEM DID NOT INDICATE SHARP EDGES. HOWEVER, IT IS POSSIBLE THAT THE METAL PLATE COULD HAVE CAUSED THIS TYPE OF INJURY WITH APPLIED PRESSURE. THERE IS NO INDICATION OF ANY INJURY PRACTICES WERE FOLLOWED (BLOOD SAMPLES AND VACCINATIONS) AS A PRECAUTION. IN 2006, THE MANUFACTURER IMPLEMENTED A SYSTEM WITH RADIUS CORNERS AND EDGES. THE OLA2500 SYSTEM MENTIONED IN THIS EVENT WAS PRODUCED PRIOR TO THAT MODIFICATION. THE FOREIGN MANUFACTURER WILL INITIATE A FIELD MODIFICATION TO BE IMPLEMENTED IN THE USA BY OLYMPUS AMERICA INC. FIELD SERVICE ENGINEERS. THE MODIFICATION WILL PROVIDE A PROTECTIVE COATING TO BE APPLIED TO POTENTIALLY SHARP EDGES. NO INCIDENT OF THIS TYPE HAS OCCURRED IN THE USA, THAT OLYMPUS AMERICA INC. IS AWARE OF. | 0
1011217 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< CUSTOMER REPORTED RECEIVING ERRATIC READINGS ON THEIR FREESTYLE FREEDOM BLOOD GLUCOSE MONITOR. CUSTOMER REPORTED RECEIVING READINGS OF 389 MG/DL AND 50 MG/DL WITHIN 10 MINUTES. ALL TESTS WERE PERFORMED ON THE FINGER. THE RESULTS WHEN PLOTTED ON A PARKES ERROR GRID FELL INTO THE 'C' ZONE SHOWING HE DIFFERENCE IN VALUES TO BE CLINICALLY SIGNIFICANT. THE CUSTOMER REPORTS SQUEEZING THE SITE EXCESSIVELY TO OBTAIN SAMPLE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1062213 | THERE IS NO INDICATION THAT THE DEVICE MALFUNCTIONED OR FAILED TO PERFORM AS DESIGNED AND INTENDED. THEREFORE, THE MFR AND UF HAVE AGREED THAT THERE IS NO NEED TO EVALUATE THE DEVICE. WE ARE CONSIDERING THAT A MISUNDERSTANDING OF THE ALARM SETTINGS ON THE PART OF THE HOSP NURSING STAFF RESULTED IN A DELAY IN PROVIDING THERAPY FOR THIS PT AND THIS DELAY CANNOT BE RULED OUT AS A FACTOR IN THE PT'S DEATH. ADDITIONALLY, THE AVAILABLE INFO SUGGESTS THAT THE HOSP NURSING STAFF MAY HAVE FAILED TO CHECK AND VERIFY ALARM LIMITS EACH TIME A PARAMETER WAS USED TO ENSURE THEY WERE APPROPRIATE FOR PT BEING MONITORED. NOTE ALSO THAT THE USERS HAD THE INCORRECT IFU FOR THE MONITOR (SEE BELOW). DURING THE INVESTIGATION, THE USER STATED THAT THEY FELT THE TRAINING REC'D WHEN THE SYS WAS INITIALLY INSTALLED WAS AN 'OVERVIEW ONLY' WITH NO SPECIFIC INSTRUCTIONS ON HOW TO CONFIGURE THE POWER UP DEFAULT SETTINGS. HOWEVER, REVIEW OF THE POWER UP DEFAULT SETTINGS OF THE UF'S BEDSIDE MONITORS SHOWED THAT ALL BEDSIDE MONITORS WERE NOT SET TO FACTORY DEFAULT SETTINGS AS ORIGINALLY INDICATED BY THE UF FOLLOWING THE PT EVENT. THIS WOULD INDICATE THAT THE DEFAULT SETTINGS WERE BEING CHANGED BY HOSP'S STAFF. DURING THE INVESTIGATION, IT WAS FOUND THAT THE USER POSSESSED THE INCORRECT OPERATIONS MANUAL FOR THEIR MODEL NUMBER OF BESIDE MONITOR. THE BILL OF MATERIALS FOR THE SHIPMENT TO THE CUSTOMER WAS ORDERED AND THE CORRECT MANUAL WAS RECORDED AS SHIPPED. HOWEVER, THE CUSTOMER CLAIMS THE INCORRECT MANUAL WAS RECORDED AS SHIPPED. HOWEVER, THE CUSTOMER CLAIMS THE INCORRECT MANUAL WAS RECEIVED BY THEM. THE MFR REVIEWED ALL RECORDS AND FOUND THAT THERE WERE NO REPORTS MADE THAT THE UF REC'D AN INCORRECT OPERATIONS MANUAL. ADDITIONALLY, THERE HAVE BEEN NO REQUESTS BY THIS UF FOR ADD'L TRAINING ON THE MONITORING SYS IN THE TWO YRS OF MONITORING SINCE THE INSTALLATION. ONLY AFTER THE OCCURRENCE OF THIS PT EVENT WERE THESE PROBLEMS REPORTED TO THE MFR. UPON INSTALLATION OF BESIDE MONITORS AND CENTRAL STATION SYS, USERS ARE ROUTINELY PROVIDED WITH OPERATIONS MANUALS, QUICK REFERENCE GUIDES AND IN-SVC TRAINING CONDUCTED BY A CLINICAL APPLICATIONS SPECIALIST. DURING IN-SVC TRAINING, CHECKLISTS AND IN-SVC GUIDES ARE UTILIZED BY THE CLINICAL APPLICATIONS SPECIALIST TO ENSURE THAT ALL IMPORTANT POINTS ARE CONSISTENTLY COVERED WITH EACH USER BEING TRAINED. A REVIEW OF THE CHECKLISTS AND IN-SVC GUIDES UTILIZED IN INITIAL USER TRAINING SHOWED THAT THE TRAINING ADEQUATELY COVERS THE SETTING OF ALARMS AND POTENTIAL ALARM ISSUES, INCLUDING: ALARM SUSPENDED FEATURE, TAILORING PT ALARMS TO FIT THE CLINICAL SCENARIO, THE EFFECTS OF TURNING OFF ALARMS AT THE CENTRAL STATION VS. BEDSIDE, AND THE SYMBOLS DISPLAYED BY THE MONITORING SYS WHEN ALARMS ARE ON (BELL) VS. OFF (BELL WITH 'X' THROUGH IT). FINALLY FOLLOWING INITIAL INSTALLATION AND IN-SVC TRAINING, USERS HAVE ACCESS TO THE MFR'S TECHNICAL SUPPORT STAFF, WHICH IS AVAILABLE 24 HRS A DAY TO GET ADD'L INFO OR REPORT ANY PROBLEMS WITH THE MONITORING SYS. AT THE REQUEST OF THIS UF, THE MFR HAS PROVIDED FIVE ADD'L IN-SVC TRAINING CLASSES FOR THE UF STAFF. THE CLASSES INCLUDED STAFF FROM BOTH DAY AND NIGHT SHIFT AND EACH SESSION INCLUDED TRAINING WITH REGARDS TO OPERATION OF THE MONITOR: TRACE SELECTION, A DESCRIPTION OF EACH ALARM, STEP BY STEP INSTRUCTIONS FOR USE FOR EACH PARAMETER AND CONFIGURATION FOR POWER UP DEFAULT SETTINGS. FOLLOWING THIS TRAINING, THE UF STATE THAT THEY WERE SATISFIED WITH THE LEVEL OF TRAINING THAT WAS PROVIDED DURING THIS SITE VISIT AND STATED THAT THIS TRAINING MET THEIR NEEDS. >< IT WAS REPORTED TO THE MFR BY THE USER FACILITY (UF) THAT THE UF WAS HAVING PROBLEMS WITH ALARMS AND NEEDED IN-SVC TRAINING ON THE SYS. AFTER FURTHER EVAL, IT WAS DISCLOSED BY THE UF THAT A PT HAD CODED AND SUBSEQUENTLY EXPIRED. | 1
1046468 | EVAL IN PROGRESS, BUT NOT CONCLUDED. >< DURING PREPARATION FOR USE FOR CARDIOPULMONARY BYPASS, THE PRESSURE SENSOR REPORTEDLY COULD NOT BE CALIBRATED. THERE WAS NO ADVERSE CONSEQUENCE TO A PT AS A RESULT OF THIS EVENT. | 0
1022469 | DATASCOPE REPLACED THE BRACKET. THE BRACKET WAS RETURNED TO THE VENDOR. THEY CONFIRMED A MATERIAL DEFECT IN THE LATCHING MECHANISM. >< THE CUSTOMER ADVISED DATASCOPE SALES THAT THE MONITOR FELL OFF THE QUICK DISCONNECT MOUNT, STRIKING A NURSE ON THE HEAD. SHE REPORTEDLY 'SAW STARS' AND WAS DIZZY. A CT SCAN CAME UP NEGATIVE. THE BIOMED STATED THAT HE THINKS THE NURSES ARE ABUSIVE WITH THE EQUIPMENT. | 0
976494 | THE CENTRAL TELEMETRY MONITOR SUDDENLY LOST A SIGNAL FROM 4 PATIENT ROOMS. THE PATIENTS WERE IMMEDIATELY MOVED TO OTHER ROOMS OR PLACED ON PORTABLE MACHINES FOR CARDIAC MONITORING. THE MANUFACTURER WAS IMMEDIATELY NOTIFIED, AND THEY WERE ABLE TO QUICKLY CORRECT THE PROBLEM. THE TELEMETRY PATIENTS DID NOT SUFFER ANY HARM. MANUFACTURER RESPONSE FOR TELEMETRY MONITOR SYSTEM, APEX PRO: THE MANUFACTURER CAME TO THE HOSPITAL AND REPAIRED THE DEVICE. | 1
1245343 | I RECEIVED A PHONE CALL FROM THE NURSE REQUESTING MY REVIEW OF THE TRACING IN THE PATIENT'S ROOM. TRACING CLEARLY SHOWING A DOUBLED FETAL HEART RATE (FHR) WITH NEARLY EVERY CONTRACTION, BEGINNING APPROXIMATELY 20 MINUTES PRIOR TO MY REVIEW. FHR BASELINE AT 125BPM. TRACING RATE DURING CONTRACTIONS APPEAR LIKE INVERTED DECELS STARTING AT 210 AND INCREASING TO 240, OR SIMPLY RECORDING AT 210-240. FHR WAS AUDIBLE IN THE 110'S. THE DOUBLING OCCURRED DURING AND IMMEDIATELY AFTER THE CONTRACTION, RAISING THE CONCERN OF WHETHER THE DECELERATIONS WERE LATES. I WITNESSED SEVERAL OF THESE EVENTS. THE FETAL MONITOR DISPLAYED AT TIMES A FHR OF 115 THEN JUMPED TO 234 WHILE THE AUDIBILE SOUND WAS SLOWING (CONSISTENT WITH A DECELERATION). THIS SITUATION OCCURRED IN THE ABSENCE OF MATERNAL MOVEMENT, STAFF INTERFERENCE OR MANIPULATION OF THE ULTRASOUND DEVICE. THE DISPLAYED FHR IN OBTV WAS NEVER CONSISTENT WITH THE READING ON THE MONITOR DEVICE. IT WOULD READ 129 WHEN THE MONITOR READ 234 FOR INSTANCE. THE OBTV DELAY MADE IT APPARENT THAT THE NUMBER ON THE FETAL MONITOR WAS THE ONE TO RESPOND TO. THE PHYSICIAN CHOSE AROM AND PLACEMENT OF AN INTERNAL SCALP ELECTRODE RATHER THAN TO SWITCH TO THE OLD MONITOR. THE FSE READING IMMEDIATELY WAS CONTINUOUS AND READABLE AND AS OF THIS REPORT CONTINUES TO SHOW NO DOUBLING. | 1
1053036 | BRAINLAB PROGRAM NOT WORKING, SOFTWARE NOT WORKING. THE REP FOR BRAINLAB WAS CONTACTED AND ABLE TO PULL UP THE SOFTWARE ON HIS LAPTOP COMPUTER AND ASSIST WITH THE CASE. THE SURGERY WAS DELAYED ABOUT 45 MINUTES. THERE WAS NO PATIENT HARM. AFTER THE CASE, THE TECH INSTALLED AN UPDATED LICENSE ON THE FACILITY'S PLANNING STATIONS TO INCLUDE THE REQUIRED ARC LOCALIZER FOR IPLAN STEREOTAXY. THE SYSTEM WAS CHECKED BY BIOMED AND RETURNED TO SERVICE. | 1
1055507 | IN 2005, A DENTAL PRACTICE USED SARGENTI RC2B ROOT CANAL PASTE CONTAINING PARAFORMALDEHYDE, LEAD AND MERCURY DURING A ROOT CANAL ON TOOTH #18. AN OVERFILL OCCURRED INTO MY MANDIBULAR NERVE CHANNEL. MY INFERIOR ALVEOLAR NERVE WAS PERMANENTLY DAMAGED AND SINCE THE DENTISTS FAILED TO DISCLOSE THE OVERFILL, OR THE USE OF SARGENTI PASTE, MY ABILITY TO HAVE THE OVERFILL REMOVED BEFORE PERMANENT, DEBILITATING DAMAGE WAS DONE WAS LOST. SARGENTI PASTE, AKA N2, RC2B, RC2W, ONE STEP ENDODONTIC WHITE, N2 UNIVERSAL, ETC.. HAS NOT BEEN APPROVED BY THE FDA AND ON FEB. 12, 1993 WAS DENIED NDA APPROVAL BECAUSE THE SAFETY AND EFFICACY OF THE MATERIAL HAD NOT BEEN PROVEN -HTTP://WWW.DENTALWATCH.ORG/QUESTIONABLE/ARGENTI/HEARING-1993-.RTF-. I NOW SUFFER FROM PERMANENT NUMBNESS AND DEBILITATING BURNING PAIN IN MY LOWER LEFT LIP, CHIN, AND JAW WHICH REQUIRES EXTENSIVE MEDICATIONS TO REDUCE PAIN. NOTHING ELIMINATES IT; THEREFORE, I AM LEFT WITH 24/7 PAIN AND SUFFERING. ANTI DEPRESSANTS ARE USED TO TREAT THIS CONDITION. THE FDA RECENTLY CAME OUT WITH A WARNING THAT ANTIDEPRESSANTS CAN CAUSE DANGEROUS SIDE EFFECTS ON PTS, SUCH AS INCREASED SUICIDAL ATTEMPTS. WHY WOULD YOU ALLOW AN ILLEGAL DRUG/DEVICE TO BE USED THAT WHEN DAMAGE OCCURS REQUIRES YET ANOTHER DANGEROUS DRUG TO COUNTERACT? THE PROPONENTS OF SARGENTI PASTE, N2 PRODUCTS, CLAIM THAT RC2B, RC2W ARE GENERICS FOR N2 UNIVERSAL AND THEREFORE, THE GENERICS ARE BEING USED BEFORE THE APPROVAL OF THE NON-GENERIC N2. HOW IS THIS ALLOWED TO HAPPEN? THE AMERICAN ENDODONTIC SOCIETY, WHICH SUPPORTS THE N2 PRODUCTS NDA ACTIVITIES, CLAIMS IN ONE OF THEIR NEWSLETTERS THAT THE FDA HAS APPROVED THE SAFETY AND EFFICACY OF N2. THIS IS NOT TRUE. WHY CAN THE FDA NOT DO SOMETHING TO STOP THE FALSE ADVERTISING OF THE STATUS OF SARGENTI PASTES? HOW CAN A GENERIC VERSION OF AN UNAPPROVED DRUG BE ALLOWED ON THE MARKET? SMALL COMPOUNDING PHARMACIES ARE MAKING THIS MATERIAL, SELLING IT IN BULK QUANTITIES INTERSTATE AGAINST THE FDA POSITION THAT STATES SHIPMENTS OF 5 GRAMS OR LESS ARE GRANTED A PHARMACY EXEMPTION. THE PHARMACY, WAS SELLING INTERSTATE TO MY DENTISTS IN QUANTITIES OF 6 TIMES THE FDA ALLOWED LIMIT PER A 1991 COMMISSIONER LETTER THAT THE FDA HAS AFFIRMED REMAINS ITS POSITION TODAY. THERE ARE OTHER PHARMACIES THAT HAVE BEEN FOUND SELLING ROOT CANAL PASTES WITH PARAFORMALDEHYDE CONTENTS OF 46% AND 98% IN INTERSTATE COMMERCE. THE FDA NEEDS TO ENFORCE THEIR OWN LAWS AND STOP THE USE OF ANY PARAFORMALDEHYDE CONTAINING ROOT CANAL MATERIALS UNLESS AND UNTIL A NDA IS GRANTED. PEOPLE CONTINUE TO BE INJURED BY THIS MATERIAL. DENTISTS CONTINUE TO USE IT IN SECRECY -EVEN AFTER THE PT IS INJURED- BECAUSE THEY KNOW IT IS BELOW THE STANDARD OF CARE. THE FDA HAS THE JURISDICTION TO STOP THE MFG OF THIS MATERIAL AND SHOULD BE DOING SO. IN 1996, THE. IN 2007, THE FDA IS DOING NOTHING. WHAT HAS CHANGED BETWEEN 1996 AND TODAY THAT HAS MADE THIS DRUG NON-TOXIC AND ACCEPTABLE TO USE? IN 1991, THE FDA ISSUED A WARNING LETTER TO A COMPOUNDING PHARMACY FOR ILLEGALLY ACTING AS A MFG SELLING BULK QUANTITIES ACROSS STATE LINES. HTTP://WWW.CASEWATCH.ORG/FDAWARNING/COMP/SARGENTS.SHTML-. WHY IS IT THAT TODAY NO WARNING LETTERS ARE GENERATED YET THE FDA HAS AFFIRMED THAT THEIR POSITION HAS NOT CHANGED ON THIS MATERIAL SINCE 1991? IN SOME CASES, THE FDA CAN REGULATE DENTISTS FOR USE OF UNAPPROVED DRUGS/DEVICES/ETC. I REQUEST THAT THE FDA SUBMIT LETTERS TO EACH STATE DENTAL AND PHARMACY BOARD TO DEMAND THAT THE USE OF SARGENTI AND OTHER PARAFORMALDEHYDE CONTAINING DENTAL MATERIALS BE REMOVED FROM THE MARKET AND THE USE BE STOPPED. IT IS CLAIMED BY SARGENTI SUPPORTERS THAT US PHARMACIES ARE SELLING SARGENTI PASTE INTERNATIONALLY TO COUNTRIES. THEY RECENTLY TOUTED THAT THEY HAVE SUPPLIED 400 ROOT CANAL'S WORTH OF THIS MATERIAL FOR USE IN OTHER COUNTRY. IF INTERSTATE BULK SALES ARE NOT ALLOW, HOW ARE INTERNATIONAL SALES ALLOWED? DOSE OR AMOUNT: >3 CANALS, FREQUENCY: ONCE, ROUTE: DENTAL. DIAGNOSIS OR REASON FOR USE: TOOTH ACHE- #18-. | 0
977849 | THE EVENT IS UNDER INVESTIGATION. THE PART ASSOCIATED WITH THIS EVENT WAS RECEIVED BY OLYMPUS AMERICA INC. ON JAN. 8, 2008 AND FORWARDED TO MANUFACTURER ON JAN. 8, 2008. THE GAS SPRING ON COVER WAS REPLACED IN 2007 AND IS PERFORMING AS EXPECTED. IF REQUIRED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED AFTER THE INVESTIGATION IS COMPLETED. >< AN ACCOUNT REPORTED THAT WHILE PERFORMING MAINTENANCE ON THE ANALYZER, THE LID OF THE ANALYZER FELL ON HAND, CAUSING INJURY. EMERGENCY ROOM TREATMENT (X-RAY) DETERMINED HAND WAS NOT BROKEN BUT NEEDED SPLINT. | 1
989281 | THE CUSTOMER'S PRODUCT HAS BEEN REQUESTED BACK FOR AN INVESTIGATION. A FOLLOW-UP REPORT WILL BE FILED ONCE INVESTIGATION RESULTS ARE AVAILABLE. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE LETTER. >< A CUSTOMER REPORTED RECEIVING AN ERROR 658 ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SVC THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MGMT SYSTEM. THIS IS A KNOWN MALFUNCTION WITH THE SOFTWARE THAT CAUSES INCORRECT TRENDING OF GLUCOSE RESULTS. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
971416 | IT WAS REPORTED TO TOSOH THAT A PT WAS TESTED FOR CNTL AT 10 AM ON 8/30 THE RESULT WAS 3.98. THIS RESULT WAS REPORTED. LATER THAT SAME DAY AT 4PM, 6 HRS LATER, THE PT WAS RE-DRAWN AND THE RESULT WAS NOW <0.06. THE LAB THEN WENT AND GOT THE ORIGINAL 10 AM PT SAMPLE AND RE-RAN IT. THE RESULT WAS NOW <0.06 ON THE REPEAT. THIS PT ALSO HAD A MB RUN AT 10 AM ON THE SAME SAMPLE AND THE RESULT WAS 5.2. THE 4 PM SAMPLE HAD A MB RUN AND THE RESULT WAS 2.3. NO PT INJURY REPORTED. | 0
1220677 | HELLO, I HAVE BOUGHT A DIGITAL THERMOMETER FROM STORE: IN 2008 AT 6:37 PM. THE PRODUCT MODEL# MT19R1. THIS PRODUCT TO BE USED FOR MY 2-MONTH OLD SON. THE INTRODUCTION ON THE COVER OF THE PRODUCT SAYS: FOR USE ORALLY, RECTALLY OR UNDER THE ARM. THEREFORE, I BOUGHT THIS PRODUCT TO USE IT RECTALLY FOR MY SON UNDER THE RECOMMENDATION OF HIS PHYSICIAN. HOWEVER, WHEN I OPENED THE PRODUCT TO MEASURE MY SON'S TEMPERATURE, I READ FIRST THE ENCLOSED INTRODUCTION THAT SAYS: NOTE: THIS THERMOMETER IS NOT RECOMMENDED FOR TAKING RECTAL TEMPERATURES.' I BELIEVE THIS INACCURATE DESCRIPTION IN ORDER TO MARKET THE PRODUCT IS NOT ACCEPTABLE FROM ONE OF THE LARGEST PHARMACY CHAINS IN THE COUNTRY. IN MY OPINION, THE INTRODUCTIONS SHOULD MATCH TO GET THE CONSUMER'S SATISFACTION. I WOULD APPRECIATE IT IF THE MEDWATCH CONTACT THE STORE AND ASK FOR MORE EXPLANATION ABOUT THIS ISSUE. DIAGNOSIS OR REASON FOR USE: TEMPERATURE MEASURING. | 0
1038066 | THE CONMED MODEL 7550 ARGON BEAM COAGULATOR (ABC) WAS USED DURING A COLONOSCOPY PROCEDURE TO REPAIR A BLEED IN THE BOWEL. SEVERAL ATTEMPTS WERE MADE TO COAGULATE THE TISSUE WITHOUT SUCCESS. IN ADDITION, THE VIDEO IMAGE ON THE SONY LMD1950MD DISPLAY WOULD BECOME DISTORTED THEN BLACKOUT WITH AN ERROR MESSAGE. THE ERROR READS NO RGB CONNECTION. THE VIDEO IMAGE RETURNS TO NORMAL WHEN THE ABC ENERGY OUTPUT IS TURNED OFF AT THE FOOT SWITCH. THE COLONOSCOPY PROCEDURE WAS COMPLETED USING THIS SAME DEVICE IN THE BIPOLAR MODE. THE VIDEO SYSTEM WORKED FINE DURING THE USE OF BIPOLAR. | 1
1006299 | ONE BANK OF CENTRAL MONITORS CONNECTED IN 13 ROOMS WENT BLANK. MONITORS WERE FUNCTIONING IN THE PATIENT ROOMS, BUT WERE NOT VISIBLE IN THE MONITOR ROOM. CLINICAL ENGINEERING AND GE TECHNICIAN NOTIFIED. THE HARD DRIVE WAS REPAIRED AND CENTRAL MONITOR FUNCTION RETURNED FOUR HOURS AFTER MALFUNCTION. | 1
1047522 | THE CUSTOMER'S PRODUCT HAS BEEN REQUESTED BACK FOR AN INVESTIGATION. A FOLLOW-UP REPORT WILL BE FILED ONCE INVESTIGATION RESULTS ARE AVAILABLE. >< A CUSTOMER REPORTED RECEIVING A GLUCOSE RESULT OF 98 MG/DL ON THEIR FREESTYLE FLASH BLOOD GLUCOSE MONITOR, WHICH WAS REPORTED AS BEING LOWER THAN HOW SHE WAS FEELING. THE CUSTOMER THEN REPORTED FEELING SLEEPY AND TWITCHING, AND THEN REPORTED EXPERIENCING A SEIZURE. PARAMEDICS WERE CALLED, AND THE CUSTOMER WAS TRANSPORTED TO A LOCAL HOSPITAL, WHERE AN UNSPECIFIED IV TREATMENT WAS ADMINISTERED IN ORDER TO STABILIZE THE CUSTOMER. THERE WAS NO REPORT OF DEATH OR PERMANENT IMPAIRMENT ASSOCIATED WITH THIS EVENT. | 0
857371 | CONCLUSIONS - DEVICE REPLACED FOR THE CUSTOMER. >< INFO WAS RECEIVED THAT THE DEVICE REPORTEDLY WOULD NOT ALARM. THE DEVICE WAS NOT REPORTED TO BE IN PATIENT USE AND THERE WAS NO REPORTED PATIENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE MFR FOR EVALUATION. DURING THE REPAIR EVALUATION, IT WAS FOUND THAT THE DEVICE WOULD NOT ALARM DURING CERTAIN ALARM TEST POINTS OF THE CHECK OUT PROCEDURE. THE PROBLEM WAS INITIALLY DIAGNOSED AS A FAULTY MAIN PCA BOARD. THE DEVICE WAS FORWARDED TO ENGINEERING FOR FURTHER INVESTIGATION. THE DEVICE WAS TESTED BY ENGINEERING USING THE CHECK OUT PROCEDURE AND FOUND TO ALARM TO SPECIFICATION. THE CHECK OUT PROCEDURE WAS REPEATED SEVERAL TIMES WITH THE DEVICE ALARMING TO SPECIFICATION EACH TIME. ENGINEERING COULD NOT DUPLICATE THE PROBLEM. THE DEVICE WAS SCRAPPED DUE TO THE POTENTIAL FOR AN INTERMITTENT MAIN PCA BOARD FAILURE. | 1
1007960 | IN 2008, A PT INFORMED NORTH COUNTRY HOSPITAL THAT SHE HAD BEEN INJURED WHILE UNDERGOING A MAMMOGRAM AT THEIR FACILITY. SHE STATED THAT SHE HAS A HERNIATED DISC IN HER NECK, CAUSED BY AN INCIDENT THAT OCCURRED WHILE AT THEIR FACILITY. SHE REPORTED THAT DURING A MAMMOGRAM IN '2008', THE SELENIA MAMMOGRAPHY SYSTEM INADVERTENTLY MOVED UPWARD WHILE SHE WAS UNDER COMPRESSION. SHE REPORTED THAT THE SYSTEM LIFTED HER UP ON HER TOES. THE PT DID NOT REPORT THIS EVENT TO THE TECHNOLOGIST DURING THE EXAM OR IMMEDIATELY FOLLOWING THE EXAM. THE TECH REPORTED THAT SHE DID NOT OBSERVE ANY INADVERTENT MOTION OF THE MAMMOGRAPHY SYSTEM WHILE WITH THIS PT. >< INVESTIGATION RESULTS: ON FEB 4, 2008, THE SELENIA MAMMOGRAPHY SYSTEM WAS EVALUATED BY THE MAMMOGRAPHY TECHNOLOGIST AND WAS FOUND TO BE FUNCTIONING NORMALLY. AN HOLOGIC TECHNICAL SUPPORT ENGINEER REVIEWED THE DYNAMIC LOG FILES FROM THE SYSTEM. THIS LOG RECORDS SYSTEM OPERATION EVERY FEW SECONDS. THE LOG FOR ALL VIEW STUDIES OF THIS PARTICULAR PT RECORDED IN LATE 2007, SHOW NO ANOMALIES IN SYSTEM OPERATION. HOLOGIC COULD NOT CONFIRM THAT INADVERTENT MOVEMENT OF THE MAMMOGRAPHY SYSTEM OCCURRED DURING THE PT'S EXAMINATION ON THE SAME DAY. THE TECHNOLOGIST REPORTED THAT SHE DID NOT OBSERVE ANY INADVERTENT MOTION OF THE MAMMOGRAPHY SYSTEM WHILE WITH THIS PT. THE TECHNOLOGIST ALSO REPORTED THAT THE PT NEVER INDICATED TO HER THAT INADVERTENT MOVEMENT OF THE MAMMOGRAPHY SYSTEM OCCURRED OR THAT SHE (PT) HAD BEEN INJURED. THIS SITE HAS CONTINUED TO USE THE MAMMOGRAPHY SYSTEM WITHOUT INCIDENT. HOLOGIC FIELD ENGINEER EVALUATED THE MAMMOGRAPHY SYSTEM ON MARCH 3, 2008, AND FOUND IT TO BE FUNCTIONING PROPERLY. | 1
964225 | USER EXPERIENCING FALSE NEGATIVE URINE LEUKOCYTES TESTS RESULTS ON APPROXIMATELY 3-4 PATIENTS PER DAY SINCE 2007. THE FOLLOWING EXAMPLE WAS PROVIDED: INITIAL RESULT NEGATIVE FOR URINE LEUKOCYTES. A MICROSCOPIC EXAM WAS PERFORMED WHICH SHOWED 20-25 WBC'S PER HPF. NO ERRONEOUS RESULTS WERE REPORTED. IF ADDITIONAL INFORMATION IS RECEIVED, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
1018238 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) AND REPORTED THAT INITIAL AND REDRAW SAMPLE GENERATED ELEVATED WHITE BLOOD CELLS (WBCS) RESULTS THAT WERE FLAGGED WITH THE INSTRUMENT GENERATED 'X' APERTURE ALERTS. THE RESULTS WERE REPORTED OUT OF THE LAB AND A PHYSICIAN REQUESTED THE REDRAWN SAMPLE TO BE SENT TO A REFERENCE LAB FOR REPEAT TESTING. WBC RESULTS FROM THE REFERENCE LAB RECOVERED WITHIN NORMAL RANGES. UPON FURTHER REVIEW OF THE DATA, IT WAS NOTED THAT HEMOGLOBIN (HGB) AND PLATELET (PLT) RESULTS OBTAINED FROM THE ACT DIFF INSTRUMENT DID NOT CORRELATE WITH THE REFERENCE LAB RESULTS. WBC PARAMETER IS NOT A SOLE DETERMINANT; HOWEVER, HGB AND PLT RESULTS ARE CONSIDERED SOLE DETERMINANTS. BASED ON AVAILABLE INFO, THERE WAS NO AFFECT TO PT OR USER. >< THE SPECIMENS WERE COLLECTED IN VACUTAINER. THE ERRONEOUS RESULTS OCCURRED IN A WHOLE BLOOD MODE OF OPERATION FOR A SPECIFIC SAMPLE. QC IS RUN ONCE A DAY AND WAS RUN PRIOR TO THE EVENT. THE INSTRUMENT IS CURRENTLY WORKING WITHIN QC SPECIFICATIONS. PREVIOUSLY TESTED, PT SAMPLES WERE NOT RERUN TO CONFIRM ACCURACY BACK TO THE LAST ACCEPTABLE QC RUN. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S LAB: THE FSE VERIFIED INSTRUMENT OPERATION. THE FSE RAN REPRODUCIBILITY, CARRY-OVER TESTING, CONTROL RECOVERY, AND A START-UP. ALL PARAMETERS WERE WITHIN SPECIFICATIONS. PER PHONE CONVERSATION WITH AN OPERATOR IN 2008, THE INSTRUMENT IS FULLY FUNCTIONAL WITH NO OTHER PROBLEMS. A CLEAR ROOT CAUSE HAS NOT BEEN DETERMINED TO DATE. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. | 0
1220708 | TELEMETRY UNIT REPORTED THAT THERE WERE ISSUES ASSOCIATED WITH A PHILIPS MONITOR DISPLAYING A DISCHARGED PATIENT'S RHYTHM AND OTHER PROTECTED HEALTH INFORMATION (PHI) SHOWING UP ON A NEWLY ADMITTED PATIENT'S RHYTHM AND PHI. RISK MANAGEMENT SPOKE WITH A PHILIPS CUSTOMER SERVICE REP WHO HAD ALREADY BEEN CONTACTED BY OUR FACILITY BIOMEDICAL DEPT ABOUT IT, AND THE MONITOR WAS SENT TO THEIR CORPORATE OFFICE FOR ANALYSIS AND FOR A REPORT TO BE PROVIDED. A PHILIPS FIELD SERVICE ENGINEER (FSE) CAME ON-SITE TO THE FACILITY TO DOWNLOAD THE MONITOR EQUIPMENT'S HISTORY FOR EVALUATION AND COMPLETE A PERFORMANCE TESTING ON THE DEVICE. TWO MONTHS AFTER THE EVENT, A PHILIPS HEALTHCARE INFORMATICS/PT MONITORING QUALITY & REGULATORY SPECIALIST PROVIDED A REPORT SPECIFYING THAT ON THE DATE OF THE REPORT, A NEWLY ADMITTED PATIENT HAD AN ALARM OCCUR AND STAFF NOTED AN ALARM STRIP WITH A SIMILAR RHYTHM WAS IN THIS NEW PATIENT'S ALARM HISTORY, BUT THE STRIP HAD AN EARLIER DATE ON IT WHICH WAS PRIOR TO THE ADMISSION DATE OF THE SECOND PATIENT. THE REPORT SAID THEY HAD INVESTIGATED THE ISSUE AND DETERMINED THERE WAS INSUFFICIENT INFORMATION PROVIDED TO CONFIRM THAT ANY DEVICE MALFUNCTION OCCURRED. THE FSE VERIFIED THE DEVICE WAS PERFORMING ACCORDING TO SPECIFICATIONS AND WAS ALARMING AND STORING/RECORDING ALARM STRIPS APPROPRIATELY. THE FSE COLLECTED AND ANALYZED LOG FILES, BUT A COMPLETE ANALYSIS COULD NOT BE PERFORMED DUE TO THE PRIOR PATIENT'S ADMISSION AND DISCHARGE HAD NOT BEEN PROVIDED. THE LOGS DID SHOW THE LAST DISCHARGE/REMOVE DATA ACTION PERFORMED FOR A PATIENT PRIOR TO THE INCIDENT, TWO DAYS PRIOR TO THE INCIDENT @ 16:22. IN SOME CASES, A PATIENT MAY BE DISCHARGED FROM THE CENTRAL, BUT MAY REMAIN ON THEIR MONITORING DEVICE FOR SHORT TIME AND MAY CONTINUE TO GENERATE ALARMS WHICH WILL THEN BE STORED FOR THAT BED. IF ANOTHER DISCHARGE IS NOT PERFORMED BEFORE A NEW PATIENT IS ADMITTED, THIS PRIOR ALARM DATA WILL BE APPENDED TO ANY NEW PATIENT ADMITTED TO THE BED. PHILIPS INDICATED INFORMATION ABOUT ADMITTING/DISCHARGING COULD BE FOUND IN THE INTELLIVUE INFORMATION CENTER INSTRUCTIONS FOR USE, PART # M3150-9001J, IN THE SECTION ON PATIENT MANAGEMENT. PHILIPS HC INDICATED THEY WOULD CONTINUE TO REVIEW AND MONITOR INTERNAL QUALITY RECORDS ASSOCIATED WITH THE INVOLVED PRODUCT AND WOULD TAKE APPROPRIATE ACTIONS AS NECESSARY. | 1
1032073 | PHILIPS IS IN THE PROCESS OF OBTAINING MORE INFORMATION CONCERNING THIS EVENT AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE COMPLETED ONCE THE INVESTIGATION IS COMPLETE. THE ISSUE REPORTED BY THE CUSTOMER IS CONSISTENT WITH THE SOFTWARE PROBLEM RESOLVED. >< YESTERDAY, THE WORST CASE SCENARIO OCCURRED. THE PRESCRIPTION MODULE DID NOT WORK RELIABLE [SIC]. DUE TO INCOMPLETE INSTRUCTIONS (GENERATE THROUGH ICIP) A MEDICAL TREATMENT (OP) WAS REQUIRED ON A PATIENT! AT 23.03, THE PATIENT WAS TO RECEIVE PANTOPRAZOL. THIS MEDICATION DID NOT ADMINISTER TO THAT PATIENT UNTIL 26.03, BECAUSE THE PRESCRIPTION WAS NOT OBSERVABLE IN THE WORK FOLDER AT ANY TIME!! THE PATIENT GOT A DISTINCT ULCER AND THAT REQUIRED AN EMERGENCY GASTRECTOMY. AT 26.03, THE OLD PRESCRIPTION WAS DIMMED AND AN NEW ONE STARTED UP. NOW, IN THE WORK FOLDER, YOU CAN SEE THE OLD' AND NEW PRESCRIPTION. THE MEDICAL PRESCRIPTION SHOWS, THAT WITH THE PRESCRIPTION FROM 23.03, THE FREQUENCY WAS SIGN IN [SCHEDULED] BUT IT WERE NOT BUILT UP [POPULATED] IN THE SCHEDULE PLAN. MAYBE, THIS IS THE REASON WHY THE PRESCRIPTION DID NOT GENERATE. | 1
1274224 | IMAGES FROM THE SIEMENS AXIOM ARTIS SYSTEM ARE SUPPOSED TO BE ARCHIVED AUTOMATICALLY BY THE MCKESSON HORIZON CARDIOLOGY INFORMATION SYSTEM. SOMETIMES THIS DOES NOT HAPPEN. DEPARTMENTAL POLICY IS THAT AT THE END OF EACH CASE THE MONITORING TECH IS SUPPOSED TO VERIFY THAT THE IMAGES HAVE TRANSFERED OVER. ONCE THE TRANSFER AND ARCHIVING IS VERIFIED, THE IMAGES CAN BE DELETED FROM THE SIEMENS SYSTEM.IN THIS CASE THE TEAM MEMBER NOTED THAT THE IMAGES WERE MISSING, BUT DIDN'T ATTEMPT TO ARCHIVE THEM. THE TEAM MEMBER THEN ACCIDENTALLY DELETED THE IMAGES FROM SIEMENS. ONCE DELETED, THESE IMAGES CAN'T BE RECOVERED. THE CATH PROCEDURE WAS RE-DONE. | 1
874577 | THE CUSTOMER IS UNAWARE OF THE CAUSE OF THE HEMOLYSIS. PROVUE MALFUNCTION MOST LIKELY NOT THE ROOT CAUSE OF THE ISSUE. HOWEVER, THE ANALYZER HAS NOT BEEN EVALUATED BY OCD PERSONNEL TO DATE, AND PROVUE MALFUNCTION HAS NOT BEEN RULED OUT. THE CUSTOMER INDICATED THAT FOLLOWING THIS INCIDENT, THE PROVUE FUNCTIONED AS EXPECTED. REPAIRS HAVE NOT BEEN PERFORMED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR INCIDENTS. >< THE CUSTOMER INDICATED THAT THEY HAD PERFORMED TESTING WITH OCD 0.8% AFFIRMAGEN REAGENT RED CELLS AND REMOVED THE RED CELLS FROM THE PROVUE AFTER A FULL DAY OF TESTING. THE CUSTOMER INDICATED THAT THEY NOTICED HEMOLYSIS IN THE REAGENT VIALS PRIOR TO LOADING THEM ON THE PROVUE THE FOLLOWING DAY. THE CUSTOMER IS UNSURE OF WHEN THE CONTAMINATION OCCURRED. THE USE OF HEMOLYZED REAGENT RED CELLS MAY AFFECT THE INTERPRETATION IN DETECTING ANTIBODY-INDUCED HEMOLYSIS. HEMOLYSIS IS A POSITIVE RESULT IN TESTS FOR ANTIBODIES TO RED CELL ANTIGENS. IF THIS INCIDENT WERE TO RECUR UNDETECTED, ERRONEOUS RESULTS MAY OCCUR, LEADING TO POSSIBLE TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1084725 | A FLUOROSCOPIC PERFORMANCE EVALUATION WAS CONDUCTED BY A PHYSICIST PER THE ADMINISTRATIVE CODE (TAC) REGULATIONS AS WELL AS ADDITIONAL IMAGE QUALITY STANDARDS (E.G., RADIATION SAFETY, IMAGE QUALITY, ETC.) THE SYSTEM WAS EVALUATED IN 2008, AND AGAIN THREE MONTHS LATER, JUST AFTER THIS INCIDENT WAS REPORTED. THERE WERE NO PROBLEMS NOTED AND THE SYSTEM PASSED ALL REQUIREMENTS. ERROR LOG DATA AND STUDY INFORMATION FROM THE DATA BASE HAVE BEEN OBTAINED FROM SITE FOR INVESTIGATION INTO THE CAUSE OF THIS INCIDENT. A REVIEW OF THIS INFORMATION DID NOT REVEAL A SYSTEM PROBLEM, BUT DID REVEAL THAT THE STUDY EXPOSURE TIME WAS APPROXIMATELY 160 MINUTES AS REPORTED. TOSHIBA DOES NOT BELIEVE FURTHER ACTIONS ARE REQUIRED AT THIS TIME. >< IN 2008, IT WAS REPORTED THAT A PATIENT ALLEGEDLY EXPERIENCED A RADIATION BURN. IT WAS FOUND THAT THE PATIENT WAS EXPOSED FOR APPROXIMATELY 160 MINUTES. | 1
1109352 | CUSTOMER REPORTED THAT HE WAS ATTEMPTING TO USE EXPIRED TEST STRIPS AND WAS UNABLE TO TEST AND LOST CONSCIOUSNESS. SHE STATED SHE WAS SEEN AT A LOCAL HOSPITAL AND TREATED WITH INSULIN. HOWEVER, NO DIAGNOSIS WAS REPORTED. THERE WAS NO REPORT OF DEATH OR PERMANENT IMPAIRMENT ASSOCIATED WITH THIS CASE. >< NO PRODUCT IS EXPECTED BACK FOR INVESTIGATION, THIS IS A FINAL REPORT. | 0
879493 | THE CUSTOMER INDICATED THAT THE ORTHO PROVUE ANALYZER PRODUCED FALSE NEGATIVE RESULTS FOR A PATIENT'S SAMPLES CONTAINING ANTI-K. RETESTING OF THE SAMPLES IN MANUAL METHOD AND THE PROVUE PRODUCED POSITIVE RESULTS. THE SAMPLES WERE FROM AN ONCOLOGY PATIENT WHO RECEIVED BLOOD TRANSFUSION ON WEEKLY BASIS. THE PATIENT WAS TRANSFUSED WITH ONE K POS UNIT. TRANSFUSION REACTION WAS NOT REPORTED AS A RESULT OF THIS INCIDENT. IF A SIGNIFICANT ANTIBODY/ANTIGEN INTERACTION IS NOT DETECTED DURING ANTIBODY SCREEN TESTING, OR IS NOT IDENTIFIED UPON FURTHER TESTING, THE PATIENT MAY BE TRANSFUSED WITH ANTIGEN-POSITIVE BLOOD AND SUFFER A HEMOLYTIC TRANSFUSION REACTION. >< THE CUSTOMER DID NOT SUBMIT TEST MATERIALS TO MTS FOR INVESTIGATION. THEREFORE, THE COMPLAINT COULD NOT BE CONFIRMED. NO DEFINITIVE ROOT CAUSE COULD BE DETERMINED. THE CUSTOMER INDICATED THAT THE INSTRUMENT WAS OPERATING AS EXPECTED AND THAT SERVICE WAS NOT REQUIRED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR INCIDENTS. | 0
1015844 | CUSTOMER REPORTED THAT THE PROBE IS LEAKING. OBSERVATION OF LIQUID ON THE UPPER PORTION OF THE CARD, AND CELLS APPEAR TO HAVE EVIDENCE OF CONTAMINATION. TESTING WAS ABORTED; NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND REPLACED THE NECESSARY COMPONENTS AND PERFORMED ADJUSTMENTS TO RETURN THE ANALYZER TO EXPECTED OPERATION. THE CUSTOMER HAS NOT LOGGED ANY COMPLAINTS REGARDING THIS ISSUE SINCE THIS INCIDENT. | 0
1015845 | THE CUSTOMER REPORTED THE DISCOVERY OF A BENT LEAKING PROBE. THE CUSTOMER INDICATES THAT THE DISCOVERY WAS MADE WHILE ADDRESSING NUMEROUS LLD ERRORS. WHILE RETRYING THE LLD, THE CUSTOMER OBSERVED THE PROBE LEAK INTO REAGENT VIALS. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE REPLACED THE NECESSARY COMPONENTS AND PERFORMED ADJUSTMENTS TO RETURN THE ANALYZER TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS REGARDING THIS ISSUE SINCE THIS INCIDENT. | 0
1060816 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND INDICATED THAT THE CUSTOMER HAD REPORTED THE PROBE WAS BENT. REPLACEMENT OF THE PROBE AND THE APPROPRIATE ADJUSTMENTS HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED FLUID. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1009148 | THE CUSTOMER REPORTED THE SYRINGE LEAKED FLUID ONTO THE REAGENT CAROUSEL OF THE ORTHO PROVUE ANALYZER DURING QC TESTING. THE CUSTOMER OBSERVED SALT RESIDUE ON THE OUTSIDE OF THE REAGENT CAROUSEL AND THE SYRINGE. ALTHOUGH CONTAMINATION OF REAGENTS DID NOT OCCUR, THE CUSTOMER INDICATED THAT THE REAGENTS WERE DISCARDED AS A PRECAUTIONARY MEASURE; NO ERRONEOUS RESULTS WERE REPORTED. FLUID LEAK CAN LEAD TO DILUTION OF REAGENT, SAMPLE CARRY-OVER, CROSS CONTAMINATION, AND/OR ERRONEOUS TEST RESULTS. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED ON SITE AND DETERMINED THAT THE CAVRO VALVE WAS RUSTED AND HAD A BUILD UP OF SALT RESIDUE. THE FE CLEANED THE VALVE AND REPLACED SYRINGE AND WASH STATION TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. | 0
1009149 | NO DEFINITIVE CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER PERFORMED CAMERA ADJUSTMENTS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< AN OCD FIELD ENGINEER (FE) REPORTED THAT WHILE AT THE CUSTOMER SITE PERFORMING SERVICE, THE ORTHO PROVUE ANALYZER INTERPRETED A WEAKLY POSITIVE REACTION AS NEGATIVE. VISUAL INSPECTION OF THE PROCESSED GEL CARDS SHOWED THE REACTIONS WERE WEAKLY POSITIVE. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. THE USER DETECTED THE ISSUE PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. | 0
1007842 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THAT THE PROBE WAS BENT. A BENT PROBE COULD LEAD TO A LOSS OF VACUUM/PRESSURE IN THE FLUIDICS SYSTEM AND PROBE DRIP. THE FE REPLACED THE PROBE TO RETURN THE INSTRUMENT TO EXPECTED OPERATIONS. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED FLUID WHILE PERFORMING PATIENT TESTING. THE CUSTOMER ABORTED TESTING AND DISCONTINUED THE USE OF THE ANALYZER. CONTAMINATION OF REAGENTS OR SAMPLES WAS NOT OBSERVED. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. THE USER DETECTED THE ISSUE, PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. | 0
1007843 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THAT THE PRESSURE WAS OUT OF SPECIFICATION. INCORRECT PRESSURE AND/OR VACUUM IN THE FLUIDICS SYSTEM CAN LEAD THE PROBE TO DRIP. THE FE REPLACED THE PRESSURE REGULATOR AND PERFORMED THE APPROPRIATE ADJUSTMENTS AND REPAIRS TO RETURN THE INSTRUMENT TO EXPECTED OPERATIONS. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED THE PROBE ON THE ORTHO PROVUE ANALYZER LEAKED FLUID, WHICH RESULTED IN CONTAMINATION OF THE REAGENTS ON BOARD. THE CUSTOMER ALSO INDICATED THAT THE DILUTION CUP OVERFLOWED. THE CUSTOMER ABORTED TESTING AND DISCARDED ALL THE CONTAMINATED REAGENTS; NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. THE USER DETECTED THE ISSUE, PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. | 0
969966 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER LEAKED INTO THE REAGENT RED CELLS VIALS CAUSING CONTAMINATION. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND /OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. ERRONEOUS TEST RESULTS WERE NOT REPORTED AS A RESULT OF THIS INCIDENT. >< AN OCD FIELD ENGINEER (FE) VISITED THE SITE. A POSSIBLE ROOT CAUSE WAS DETERMINED. THE PROBE WAS CLOGGED AND THE WASH STATION WAS CRACKED. AN OBSTRUCTION IN THE FLUIDICS SYSTEM COULD LEAD TO LOSS OF VACUUM PRESSURE, WHICH COULD RESULT IN A PROBE LEAK. INFO WAS NOT PROVIDED REGARDING POSSIBLE RESULT CODES INTENDED TO PREVENT ERRONEOUS RESULTS FROM BEING REPORTED. THE OPERATOR DETECTED THE ISSUE, PREVENTING THE POSSIBILITY OF ERRONEOUS RESULTS FROM BEING REPORTED. INSTRUMENT MALFUNCTION COULD NOT BE RULED OUT. A COMPLAINT REVIEW INDICATES INCIDENT HAS NOT RECURRED AT THE SITE POST SERVICE. | 0
1014950 | WE ARE AWAITING RECEIPT OF THE COMPLAINT DEVICE TOWARDS CONDUCTING A ROOT CAUSE FAILURE ANALYSIS. WHEN THE INVESTIGATION IS COMPLETED, THE RESULTS WILL BE THE SUBJECT OF THE FOLLOW UP REPORT. >< IT WAS REPORTED BY THE SALES REP. THAT THE PRESSURE DISPLAYED ON THE PUMP IS NOT ACCURATE AND IS READING LOWER THAN ACTUAL PRESSURE. THE SALES REP. STATED THAT IN THE LAST SEVERAL PROCEDURES, THERE WAS EXCESSIVE SWELLING OF THE KNEE JOINT. THERE WAS NO EXTENDED RECOVERY TIME DUE TO THE ISSUE. HOWEVER, IN THE 1ST OF THE TWO KNEE PROCEDURES IN 2008, THE SURGEON RESORTED TO MAKING A PORTAL AT THE POSTERIOR OF THE KNEE TO FREE THE FLUID AND RELIEVE THE COMPARTMENT PRESSURE TO AVOID PATIENT INJURY. IN THE 2ND PROCEDURE ON THREE DAYS LATER, A SIMILAR EVENT OCCURRED. IN THIS INSTANCE, THE OUTFLOW CANNULA WAS REMOVED FROM ITS PORTAL ALLOWING FOR A GREATER OUTFLOW OF FLUID TO REDUCE AND RELIEVE PRESSURE IN THE JOINT SPACE THUS PREVENTING DAMAGE. IN BOTH CASES THE INTERVENTION WAS SUCCESSFUL IN PREVENTING INJURY. BOTH PROCEDURES WERE COMPLETED SUCCESSFULLY WITHOUT FURTHER INCIDENT OR HARM TO THE PATIENT. THIS IS ALL OF THE INFORMATION THAT COULD BE ATTAINED. ALSO SEE ASSOCIATED MDR 1221934-2008-00141. | 0
1083961 | AT THIS POINT IN TIME MITEK IS AWAITING RECEIPT OF THE COMPLAINT DEVICE TOWARDS CONDUCTING A FAILURE ANALYSIS. THE RESULTS OF THE ANALYSIS WILL BE THE SUBJECT OF A FOLLOW UP REPORT. >< THE USER FACILITY IS REPORTING THAT DURING AN ARTHROSCOPIC PROCEDURE, BOTH THE SURGEON AND THE PATIENT RECEIVED AN ELECTRICAL SHOCK WITH THE USE OF AN FMS PUMP, FMS INTERFACE CABLE AND A COMPETITORS SHAVER SYSTEM, A SMITH & NEPHEW DYONICS EP1. THEY ARE REPORTING THAT THERE WERE NO INJURIES, AND THAT THE PROCEDURE WAS CONCLUDED SUCCESSFULLY WITHOUT FURTHER INCIDENT OR HARM TO THE PATIENT. SEE ALSO ASSOCIATED MDR 1221934-2008-00358. | 0
1031160 | SAMPLES FROM THE SAME LOT WERE EXAMINED. MICROBIAL TESTING CONFIRMED PRESENCE OF ASPERGILLUS IN ALL SAMPLES. PRODUCTION DOCUMENTATION WAS REVIEWED FROM RAW MATERIALS THROUGH FINISHED PRODUCT. THE REVIEW INDICATED THAT THE PRODUCT COMPLIED WITH SPECIFICATIONS AND QA SYSTEM REQUIREMENTS. PRODUCT IS NOT INTENDED TO BE STERILE AND IS NOT LABELED AS STERILE. ADDITIONAL TESTING IS ONGOING, AND A FOLLOW-UP REPORT WILL BE FILED. >< ACCORDING TO THE INFORMATION RECEIVED, A 6-MONTH OLD STOMY PATIENT WAS INFECTED BY MUCOR. PATIENT TREATED WITH AMBISOME, AND TREATMENT OF FUNGAL INFECTION WAS SUCCESSFUL. PATIENT CONDITION IS STILL CRITICAL, BUT STABLE. HOSPITAL LABORATORY TESTING OF TWO PRODUCT SAMPLES IDENTIFIED MUCOR ON BOTH. COLOPLAST TESTING OF PRODUCT FROM REFERENCED LOT IDENTIFIED ASPERGILLUS. | 0
1063151 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS ON A PATIENT SAMPLE RUN ON GALILEO USING CAPTURE R READY ID. ANTI K WAS IDENTIFIED BUT NOT ANTI FYA. >< REACTIVITY OF THE FYA ANTIGEN WAS CONFIRMED ON RETENTION CAPTURE-R READY-SCREEN, LOT X242, AND RETENTION CAPTURE-R READY-ID, LOT ID100. THESE ARE THE PRODUCTS USED BY THE CUSTOMER ON GALILEO AT THE TIME OF THE EVENT. THE CUSTMER DID NOT RETURN PRODUCT OR SAMPEL FOR INVESTIGATION TESTING. IT IS NOT POSSIBLE TO RULE OUT THE SAMPLE AS A CAUSE OF THE EVENT. | 0
968183 | IT WAS REPORTED THAT A HOSPITAL'S VNS THERAPY PROGRAMMING HANDHELD WAS FREEZING UP. TROUBLESHOOTING VIA PERFORMING A RESET OF THE HANDHELD RESOLVED THE ISSUE. HANDHELD IS NOT BEING RETURNED TO MANUFACTURER AT THIS TIME, AS USER FACILITY IS STILL GOING TO USE IT. >< DEVICE MALFUNCTION SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1067009 | REPORTER INDICATED THAT A VNS PATIENT WAS FOUND TO BE SET TO 0MA DURING INTERROGATION. THE VNS WAS REPROGRAMMED SUCCESSFULLY AND THE PATIENT WAS ASYMPTOMATIC. ATTEMPTS FOR PROGRAMMING HISTORY AND FURTHER INFORMATION ARE IN PROGRESS. | 0
968180 | DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. >< REPORTER INDICATED, THAT HIS VNS THERAPY PROGRAMMING HANDHELD 'WAS NOT RESPONDING, WHEN HE WOULD TAP IT WITH THE STYLUS.' TROUBLESHOOTING DID NOT RESOLVE THE ISSUE. GOOD FAITH ATTEMPTS ARE CURRENTLY BEING MADE TO OBTAIN THE HANDHELD FOR PRODUCT ANALYSIS. | 0
1067007 | REPORTER INDICATED A VNS PATIENT WAS INTERROGATED AND FOUND TO BE SET TO 0MA. THE PATIENT WAS ASYMPTOMATIC AND WAS SUCCESSFULLY REPROGRAMMED TO THE APPROPRIATE SETTINGS. THE REPORTER STATED HE IS NOT SURE IF A FINAL VNS INTERROGATION WAS DONE AT THE PATIENT'S LAST VISIT SIX MONTHS PRIOR. ATTEMPTS FOR PROGRAMMING HISTORY TO ESTABLISH IF A PROGRAMMING ANOMALY OCCURRED ARE IN PROGRESS. | 0
1000773 | DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. >< REPORTER INDICATED THAT A DELL HANDHELD COMPUTER WOULD NOT CHARGE EVEN AFTER SEVERAL HOURS ON THE CHARGER. THE REPORTER COULD NOT TURN ON THE HANDHELD COMPUTER. PRODUCT HAS BEEN RETURNED TO THE MANUFACTURER FOR ANALYSIS. | 0
997544 | DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. >< IT WAS REPORTED THAT A HP JORNADA HANDHELD COMPUTER WOULD NOT POWER ON. IT HAD BEEN A YEAR SINCE THE LAST THE LAST TIME THE HANDHELD HAD BEEN USED. PRODUCT IS EXPECTED TO BE RETURNED TO THE MFR. | 0
1048322 | REPORTER INDICATED THAT A DELL X5 VNS COMPUTER WAS GIVING REPEATED ERROR MESSAGES AND THAT IT WOULD TAKE SEVERAL ATTEMPTS TO INTERROGATE VNS PATIENTS. THE EXACT ERROR MESSAGES ARE UNKNOWN. THE DELL X5 COMPUTER AND FLAHSCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. >< DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1048320 | REPORTER INDICATED A VNS COMPUTER WOULD NOT POWER ON. THE COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. >< DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
947048 | BASED ON INFORMATION THAT PMS RECEIVED ON 10/11/2007, PMS HAS DETERMINED THAT THIS REPORT IS SIMILAR TO OTHER MDR REPORTABLE EVENTS INVOLVING THE TEAL. HERE IS A SUMMARY OF THE ISSUE. THE USER REPORTED AN: ELECTRIC SMELL NEAR THE CONSOLE. GANTRY IS SPARKING. THE INITIAL REPORT DID NOT DOCUMENT THAT THE TEAL PDU WAS CAUSING THE PROBLEM. >< THE TEAL HAS UL ELECTRICAL SAFETY APPROVAL, AND THE RISK OF INJURY TO A PATIENT IS REMOTE. THIS ISSUE WAS FOUND DURING THE REVIEW OF INTERNAL SERVICE RECORDS. VOLUNTARY MEDICAL DEVICE RECALL REPORT WAS FILED. IT IS A CLASS 2 RECALL. PMS ISSUED A FIELD CHANGE ORDER TO CORRECT THIS ISSUE IN MAY, 2007. THIS FCO HAS BEEN CHANGED TO MANDATORY ACTION TO ENSURE AGREEMENT WITH THE C&R REPORT. THE ROOT CAUSE OF THE ISSUE IS A COMPONENT FAILURE ON THE VENDED DEVICE. | 1
1065055 | DUE TO A TECHNICAL PROBLEM ON 05/25/08, THE MULTIVIEW WORKSTATION (CENTRAL STATION) TURNED OUT OF OPERATION. THE CENTRAL STATION WAS INSPECTED ON 05/26/08 AND SPARE PARTS WERE ORDERED. IN THE MEANTIME, THE CUSTOMER INSTALLED A PATIENT MONITOR INSTEAD OF THE CENTRAL STATION TO BE ABLE TO SEE THE SERIOUS AND LIFE THREATENING ALARMS OF THE BEDSIDE MONITORS ON THE STATUS BAR (GERM: HINWEISZEILE). THE ALARM VOLUME OF ALL MONITORS WAS SET TO THE MAXIMUM. DUE TO NOT SUCCESSFUL REPAIR, I ORGANIZED A RENTAL CENTRAL. THE RENTAL CENTRAL COULD NOT BE PUT INTO IMMEDIATE OPERATION BECAUSE THE SOFTWARE HAD TO BE INSTALLED. AT NOON, A STAFF OF THE MEDICAL TECHNOLOGY DEPT TOLD ME THAT A PT DIED IN THE MEANTIME, BECAUSE THE STAFF DID NOT HEAR THE ALARM. IN THE AFTERNOON, I INSTALLED THE RENTAL CENTRAL AND CHECKED IT, EVERYTHING WAS OK. >< OUR EVALUATION OF THE INFO AVAILABLE AT THIS TIME AND WITH COMMUNICATION OF THE DRAEGER TECHNICAL SERVICE REP ON SITE WITH BIOMED ENGINEERING, WE HAVE REACHED THE CONCLUSION THAT OUR DEVICE DID NOT CAUSE OR CONTRIBUTE TO THIS ADVERSE EVENT. ACCORDING TO THE BIOMED ENGINEERING, THE STAFF ON SITE DID HEAR THE ALARM OF THE CONCERNED MONITOR/BED, AND THE ALARM FROM ANOTHER BED AT THE SAME TIME. BY THE TIME THE STAFF REACHED THE PT IN QUESTION, THE PT HAD ALREADY PASSED AWAY. DRAEGER DEVICE WORKED AS INTENDED. | 1
1038410 | CLOGGED/OBSTRUCTED; HGB FLOWCELL, WBC TRANSDUCER, SILICON TUBING, SYRINGE AND PROBE DRIVES. THE CUSTOMER STATED THEY ALSO NOTED THAT THE RESULTS FROM AN API PROFICIENCY TEST DID NOT MATCH THE RESULTS GENERATED ON A CELL-DYN 1800 ANALYZER. THE CUSTOMER'S CELL-DYN 1700 HAD NOT HAD PREVENTIVE MAINTENANCE FOR SEVERAL YEARS AND SERVICE WAS REQUESTED TO OPTIMIZE THE INSTRUMENT. THE FIELD SERVICE REPRESENTATIVE (FSR) FOUND THE HGB FLOWCELL READING OUT OF SPECIFICATION. A WORN 1 ML SAMPLE SYRINGE AND WORN SILICON TUBING WERE REPLACED AND THE SYRINGE AND PROBE DRIVES WERE CLEANED AND LUBRICATED. THE HGB FLOWCELL AND WBC TRANSDUCERS WERE CLOGGED/OBSTRUCTED AND THEREFORE WERE CLEANED AND ADJUSTED. THE VACUUM/PRESSURE, METERING, SETTINGS, AND GAINS WERE VERIFIED. PRECISION AND QC PASSED. AS THE INSTRUMENT MET SPECIFICATIONS, NO FURTHER INVESTIGATION WAS REQUIRED. THE CELL-DYN 1700 SYSTEM OPERATOR'S MANUAL (03H58-01, REV D) PROVIDES GENERAL TROUBLESHOOTING INFORMATION IN SECTION 10, PAGES 10-9, 10-10. THE TROUBLESHOOTING GUIDE FOR 'ABNORMAL OR ERRATIC HGB, MCH AND/OR MCHC RESULTS' (10-13) INDICATES A DIRTY HGB FLOWCELL MAY BE A PROBABLE CAUSE IF THE HGB REFERENCE VALUE IS <1700. CLEANING THE FLOWCELL AND CLEANING/CHANGING THE SAMPLE SYRINGE ARE NON- SCHEDULED MAINTENANCE ACTIVITIES (TABLE 9.1: NON-SCHEDULED MAINTENANCE FREQUENCY PAGE 9-19). THE FLOWCELL CLEANING PROCEDURE IS GIVEN ON PAGES 9-39 TO 9-41. THE SAMPLE SYRINGE REPLACEMENT/CLEANING PROCEDURE IS GIVEN ON PAGES 9-31 TO 9-33. A REVIEW OF COMPLAINTS FOR THE PERIOD OF 2007 THROUGH 2008, DID NOT INDICATE AN ADVERSE TREND FOR THE CELL-DYN 1700, LIST 03H53-01 FOR HGB ISSUES. THERE IS NO SYSTEMIC ISSUE. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATED A PATIENT GENERATED A HEMOGLOBIN VALUE OF 8.6 G/DL ON THE CELL-DYN 1700 ANALYZER. THE PATIENT WAS SENT TO THE HOSPITAL TO BE TRANSFUSED WHERE A NEW SPECIMEN WAS DRAWN AND YIELDED A HEMOGLOBIN VALUE OF 9.8 G/DL. NO TRANSFUSION WAS GIVEN. NO INJURY WAS REPORTED. | 1
1082667 | WHITE UNKNOWN MATERIAL WAS OBSERVED ON THE JEL INSIDE OF THE TRANSDUCER. >< EVALUATION CODES: DEVICE HAS NOT BEEN RETURNED AT THIS TIME FOR EVALUATION. | 0
1082668 | EVALUATION: DEVICE HAS NOT YET BEEN RETURNED FOR EVALUATION. >< CRACK WAS OBSERVED ON THE FEMALE LUER ON THE TRANSDUCER. | 0
1082666 | IT WAS REPORTED THE CONNECTOR PARTS BETWEEN TRANSDUCER AND IV LINE ARE BROKEN. >< EVALUATION: DEVICE HAS NOT BEEN RETURNED AT THIS TIME FOR EVALUATION. | 0
1082662 | PRESSURE MEASUREMENT DROPPED DOWN SLOWLY DURING USE. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION AT THIS TIME. | 0
1082670 | LEAKAGE WAS OBSERVED FROM THE UV BONDING CONNECTION, BETWEEN THE TRUWAVE AND THE THREE WAY STOP COCK UNDER 300MMHG PRESSURE. >< DEVICE HAS NOT RETURNED FOR EVALUATION AT THIS TIME. | 0
1082679 | IT WAS REPORTED THERE WAS LEAKAGE. >< DEVICE HAS NOT RETURNED FOR EVALUATION. | 0
1034507 | DEVICE NOT RETURNED FOR EVALUATION. >< IT WAS REPORTED THERE WAS LEAKAGE FOUND. | 0
984547 | IT WAS REPORTED THAT THE ARTERIAL TUBING BECAME DISCONNECTED FROM THE SAMPLING PORT. FOLLOW UP WITH THE HOSPITAL IN 2007, DID NOT REVEAL IF THE REPORTED EVENT OCCURRED BEFORE OR DURING USE. INFORMATION CONCERNING THE STATUS OF THE PT WAS NOT GIVEN, AS NO FURTHER DETAILS WERE PROVIDED. THE HOSPITAL CONTACT INDICATED THAT THEY WOULD FOLLOW UP CONCERNING FURTHER DETAILS. NO RESPONSE HAS BEEN RECEIVED AS OF THE DATE OF THIS MEDWATCH REPORT. >< THE DEVICE WAS NOT RETURNED FOR EVALUATION. | 0
998074 | THE ACCOUNT STATED THE RECAPPER TAPE SENSOR ON THE INPECO SYS IS RUNNING OUT OF TAPE WITHOUT GIVING AN ERROR AND CONTINUES TO RECAP TUBES WITHOUT TAPE. THE UNSEALED TUBES ARE PLACED IN STORAGE MODULE AND THEN SENT TO WASTE AFTER A WEEK. THE UNSEALED TUBES CREATE A BIOHAZARD WHEN SPILLED. SVC OBSERVED THE INPECO SYS SENSORS WERE NOT OPERATIONAL ON THE RECAPPER TAPE FEEDER AND FORMER HEAD. SVC CLEANED SENSORS, RESEATED THE FIBER OPTIC SENSORS ON THE CONTROL MODULE FOR RESOLUTION. NO INJURY WAS REPORTED. >< THE ABBOTT FIELD SVC REP CLEANED THE SENSORS AND RESEALED THE FIBER OPTICS AT THE SENSORS ON THE CONTROL PANEL. AFTER THE FIELD SVC VISIT, NORMAL OPERATION WAS OBSERVED FOR THE INPECO SYSTEM. REFERENCE WHERE IN LABELING THE EVENT IS ADDRESSED: THE ACCELERATOR APS OPERATIONS MANUAL INDICATES THE FOLLOWING: SECTION 8 HAZARDS - BIOHAZARD SAFETY. PRECAUTIONS. SPILL CLEAN-UP. WASTE HANDLING AND DISPOSAL. SECTION 9 MAINTENANCE. - EMPTY STORAGE AND RETRIEVAL MODULE WASTE. THIS IS A FINAL REPORT. | 0
1009102 | PT WAS IN ICU ON COMBINED TREATMENT WITH A THROMBOLYTIC AND ANTICOAGULANT WITH AN ARTERIAL AND VENOUS SHEATH IN PLACE. PT WAS LAST SEEN AT 0300 AND LATER FOUND AT 0400 OUT OF BED ON FLOOR, BLEEDING, AND UNRESPONSIVE WITH SHEATHS AND MONITOR LEADS DISLODGED. IT APPEARS THAT THE SHEATHS AND LEADS MAY HAVE BECOME DISLODGED WHEN THE PATIENT GOT OUT OF BED. REVIEW OF THE MONITOR HISTORY REFLECTS THE PATIENT WAS OFF THE CARDIAC MONITOR AT 0324, THE OXYGEN SATURATION MONITOR SHOWED A DECREASED READING FOLLOWED BY AN INABILITY TO READ AT 0343. IN THIS INCIDENT IT IS UNKNOWN WHETHER THE MONITOR FAILED TO ALARM OR IF THEY DID ALARM AND WERE NOT HEARD. HOWEVER, A FEW DAYS LATER, STAFF WITNESSED A FAILURE TO ALARM WITH THE SAME OVERHEAD MONITOR OR WHEN THE LEADS WERE DISCONNECTED. THE SAME SCENARIO WAS ALSO OBSERVED WITH A DIFFERENT MONITOR ON THE SAME UNIT. | 1
1056394 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006LETTER. >< A CONSUMER REPORTED THEIR PRECISION XTRA METER HAD SPONTANEOUSLY CHANGED THE DATE AND TIME AT LEAST THREE TIMES AN SHE HAD THE METER DOWNLOADED BY A LOCAL HOSPITAL. ABBOTT DIABETES CARE IS AWARE THAT THE HOSPITAL IS A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 1
992867 | PHILLIPS HAS RECEIVED NO ADDITIONAL INFORMATION INCLUDING NO ADDITIONAL TESTING RESULTS. PHILLIPS IS IN THE PROCESS OF OBTAINING MORE INFORMATION CONCERNING THIS EVENT AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SENT ONCE THE INVESTIGATION IS COMPLETED. >< THE CUSTOMER REPORTED THAT THE INVOLVED DEVICE ALARMED AT THE BEDSIDE. THE INVOLVED DEVICE DID NOT ALARM IN THE OTHER ROOMS, IT WAS NETWORKED TO ALARM IN. | 0
996978 | THE CUSTOMER STATES THAT THE IMX ANALYZER HAD BEEN POWERED OFF IN ORDER TO REPLACE DISPENSE PARTS. WHEN THE IMX ANALYZER WAS POWERED BACK ON, THE CUSTOMER NOTICED A LOUD NOISE, AND A SMALL AMOUNT OF SMOKE COMING FROM THE FAN AREA ON THE BOTTOM OF THE ANALYZER. THE CUSTOMER IMMEDIATELY POWERED OFF THE ANALYZER AND A SERVICE CALL WAS INITIATED. THERE WERE NO REPORTS OF INJURY OR DAMAGE TO THE SURROUNDING ENVIRONMENT. >< AN ABBOTT FIELD SERVICE REPRESENTATIVE VISITED THE CUSTOMER'S SITE AND REPLACED THE IMX ANALYZER'S AIR HEATER/THERMISTOR AND INSULATOR ASSEMBLY AS WELL AS THE EXHAUST FAN. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THIS IS A FINAL REPORT. | 0
1106619 | THE CATHETER USED IN THIS PROCEDURE WAS A NAVISTAR 4MM, 7F, D-CURVE CATHETER. HOWEVER, NO OTHER DETAILED INFO ON THE CATHETER WAS GIVEN. >< IT WAS REPORTED THAT WHEN ABLATING IN THE RIGHT ATRIUM SEPTUM IN THE SLOW PATHWAY WITH A CATHETER, A STEAM POP OCCURRED WHICH RESULTED IN A THIRD DEGREE / COMPLETE HEART BLOCK IN THE PT. THERE WAS NO NOTICE OF JUMP IN IMPEDANCE. THE CATHETER HAS BEEN DISCARDED. THE PT WAS TRANSFERRED IN GOOD CONDITION. | 0
1001138 | THE PROFESSIONAL EDUCATION SPECIALIST (PROF ED) REVIEWED THE M1-M4 CHANNELS ON THE FRONT OF THE PT INTERFACE UNIT (PIU) WITH THE CUSTOMER. THE PROF ED EXPLAINED THAT THE FRONT M1-M4 WILL BYPASS THE PACING ROUTING SYSTEM IN CARTO AND WILL ALLOW THEM TO PACE DIRECTLY THROUGH THE CATHETER. THIS INFO IS ALSO CONTAINED IN THE CARTO XP USER MANUAL. CUSTOMER STATED THAT THEY WOULD LIKE TO TEST THIS OPTION BEFORE THEIR NEXT PROCEDURE. THE PROF ED ALSO RECOMMENDED ADDITIONAL INSERVICING OF THE EP LAB STAFF ON THIS METHOD. >< CUSTOMER CALLED BIOSENSE WEBSTER (BWI) TO REVIEW PACING PROTOCOL WHEN MAPPING CATHETER IS OUT OF RANGE. CUSTOMER REPORTED THE FOLLOWING: DURING AN ABLATION PROCEDURE USING THE CARTO XP SYSTEM, THE PROCEDURE BECAME A CODE SITUATION. WHEN THE BED WAS LOWERED TO PERFORM CPR, THE LOCATION MAGNET MOVED OUT OF THE MAPPING RANGE. CUSTOMER ALSO REPORTED THAT WHEN THE PHYSICIAN ATTEMPTED TO PACE THROUGH THE MAPPING CATHETER, THE CATHETER WOULD NOT PACE DUE TO AN 'OUT OF RANGE' ERROR MESSAGE. HOWEVER, AFTER NOTICING A SPIKE ON THE PRUKA SYSTEM, THE PHYSICIAN ASSUMED THE CATHETER WAS NOT CAPTURING, AND STARTED TO APPLY MORE FORCE ON THE CATHETER. THE PHYSICIAN WAS CONCERNED ABOUT PERFORATION AND THE FACT THAT THERE WAS NO ERROR MESSAGE ON THE CARTO STATING THAT THE MAPPING CATHETER WOULD NOT PACE WHILE OUT OF RANGE. PER CUSTOMER, THE PT REMAINED IN THE HOSP, AND THE PT'S CONDITION WAS DESCRIBED AS 'NOT GREAT'. SUBSEQUENT FOLLOW UP WITH THE CUSTOMER INDICATED THAT THE PT DIED OF CARDIAC ARREST. PER CUSTOMER THE CARTO XP SYSTEM OR THE ABLATION DID NOT DIRECTLY CAUSE THIS EVENT AS THE PT WAS ALREADY SUFFERING FROM SEVERE CORONARY ARTERY DISEASE (CAD) AND WAS CHRONICALLY ILL. | 0
1252604 | THE PT'S IMAGES WERE SENT TO THIS PICTURE ARCHIVING COMPUTER SYSTEM (PACS). HOWEVER, ANOTHER PT'S X-RAY IMAGE WAS SENT TO THIS PACS AND WAS EXCHANGED FOR THE ORIGINAL PATIENT IMAGE. NO PT INJURY OR MISDIAGNOSIS HAPPENED, DUE TO THIS EVENT. | 1
1031610 | PT ON THE ARCTIC SUN HYPOTHERMIA MANAGEMENT SYSTEM, AN OUTLINE OF THE VEST WRAP ON THE PT'S TRUNK, AND THE THIGHS FROM THE ARCTIC GEL PADS CAN BE NOTED. IT APPEARS LIKE THE SKIN IS BURNED. | 0
1084407 | POSSIBLE DIAGNOSTIC CELLS OUTSIDE FIELD OF VIEW. ACCORDING TO CUSTOMER, 22 FIELDS OF VIEW WERE REVIEWED BY A CYTOTECHNOLOGIST AND NO ABNORMAL CELLS WERE PRESENT. CASE WAS SIGNED OUT AS NEGATIVE. PATHOLOGIST FOUND AREAS OF LSIL ON THE SLIDE. CYTOLOGY APPLICATIONS SPECIALIST (CAS) REVIEWED THE SLIDE AT HOSP AND FOUND TRIGGERS (ASCUS AND 2 AREAS OF LSIL) WITHIN THE 22 FIELDS OF VIEW THAT SHOULD HAVE PROMPTED FULL AUTOSCAN OF THE SLIDE. CAS REVIEWED THE SLIDE WITH THE CUSTOMER WHO AGREED WITH CAS FINDINGS. CASE IS CONSIDERED A USER INTERPRETIVE ERROR AND NOT AN IMAGE MISS. CUSTOMER AGREES THAT THE IMAGER HAD PERFORMED AS EXPECTED. CAS WAS NOT ABLE TO ACCESS THE CUSTOMER SITE AND SLIDE AND REPORT BACK FINDINGS UNTIL 2008. >< CASE IS CONSIDERED A USER INTERPRETIVE ERROR AND NOT AN IMAGE MISS. CUSTOMER AGREED THAT THE IMAGER HAD PERFORMED AS EXPECTED. | 0
1050067 | PARENTS OF INFANT USED CVS BRAND DIGITAL THERMOMETER AT HOME TO CHECK CHILD'S TEMPERATURE. TEMPERATURE READ AS 102.5. FAMILY PRESENTED TO ER WITH INFANT. IN ER SETTING TEMPERATURE OBTAINED WAS 98.9. SEPSIS WORK-UP PERFORMED, INCLUDING MULTIPLE LUMBAR PUNCTURE ATTEMPTS. IV ANTIBIOTICS ADMINISTERED. PARENTS BROUGHT THERMOMETER IN FROM HOME. WHEN TESTED, THE INFANT TEMPERATURE WAS 103 ON THE CVS THERMOMETER AND 99.2 ON THE HOPSITAL THERMOMETER. | 0
996628 | . >< REPORTED BY THE COMPLAINANT AS FOLLOWS: NURSE CALLED TO REPORT RECTAL BLEEDING WITH THE FMS IN ONE PT AT TWO DIFFERENT TIMES. THE FIRST EPISODE WAS IN 2007 AND THE SECOND WAS 2008. SHE ADVISED THAT THE PT IS A MALE WITH AN ADMISSION DATE OF TWENTY SIX DAYS EARLIER. SHE IS UNABLE TO PROVIDE DATE OF FLEXI-SEAL INSERTION. THE PT IS POSITIVE FOR C-DIFFICILE AND THAT IS THE REASON FOR THE DIARRHEA; WHICH HAD A DURATION OF APPROX ONE WEEK. SHE DESCRIBES THE PT AS HAVING BRIGHT RED RECTAL BLEEDING WHILE THE FLEXI-SEAL WAS IN PLACE WITH HEMODYNAMIC INSTABILITY AND A DECREASED HEMOGLOBIN REQUIRING BLOOD TRANSFUSIONS WHICH TOTALED AT LEAST 12 UNITS OF PACKED RBCS. COLONOSCOPY WAS PERFORMED AND IDENTIFIED ULCERATION PRESENT WITHIN THE ANTERIOR RECTAL WALL. NURSE STATES THAT A SURGICAL REPAIR WAS PERFORMED IN 2008 TO ATTEMPT TO CONTROL THE RECTAL BLEEDING. | 0
1077210 | THE PATIENT DEVELOPED LOOSE STOOLS. A FLEXI-SEAL WAS APPLIED TO CONTAIN THE STOOL AND PREVENT SKIN BREAKDOWN. THE PATIENT BECAME FEBRILE AND FOUL ODOR WAS NOTED, DESCRIBED AS ROTTING FLESH BY STAFF. THE FLEXI-SEAL BALLOON WAS DEFLATED AND NOTED TO HAVE 85 ML OF FLUID (THE RECOMMENDED AMOUNT IS 45ML). SOME NECROTIC TISSUE WAS NOTED ON THE TIP, AND WAS EXPELLED FROM THE PATIENT. THE PATIENT HAD A RECTAL SCOPE WHICH DID NOT SHOW ADDITIONAL RECTAL EROSION OR NECROSIS. THE FLEXI-SEAL HAD BEEN IN PLACE FOR THREE DAYS. BALLOON OVER INFLATION COULD POTENTIALLY LEAD TO NECROSIS. | 0
1008369 | TWO DIFFERENT PATIENT SAMPLES GOT THE SAME NAME (INCORRECT ASSIGNMENT). RESULT OF ROOT CAUSE ANALYSIS: TANGO HAS IN THE PATIENT AREA 12 TRACKS FOR 12 PATIENT RACKS. THE PROCESS OF LOADING RACKS IS CONTINUOUSLY POSSIBLE. WORKFLOW WHICH LED TO THE INCORRECT ASSIGNMENT. REQUESTS AND PT DATA ARE IN THE TANGO OPTIMO DATABASE. ON TRACK ONE THE DATA RELATED PT SAMPLES WERE LOADED. THE CUSTOMER LOADED AND REMOVED THE RACK 2 SEVERAL TIMES (OFTEN CAUSED BY BAD QUALITY OF THE BARCODES OF THE SAMPLE TUBES). THE IMPORT DIALOG OF THE PT DATA IS OPEN AND POSITION 2 ON TRACK ONE IS SELECTED. REMOVAL OF RACK 2 ON TRACK 2, FLAP WAS CLOSED AND THEREFORE, THE IMPORT DIALOG IS CLOSED TOO. THE PT DATA ARE TRANSPOSED FROM RACK1 POSITION2 TO RACK 2 POSITION2. ROOT CAUSE FOR THE PROBLEM IS A MULTIPLE LOADING OF RACKS DURING OPEN PATIENT IMPORT DIALOG CONCERNING THE BAD QUALITY OF THE BARCODES (BARCODES PRINTED BY CUSTOMER. CORRECTIVE ACTION: CUSTOMER TRAINING REG. THE INSTRUCTION OF THE CORRECT PROCESS AND ) SOFTWARE PATCH WILL BE INSTALLED TO STABILIZE THE OPERATIONAL HANDLING AT THE TIME OF 2 MONTHS. | 0
1130483 | WHILE SURGEON USING NEPTUNE WASTE MANAGEMENT SYSTEM DURING AN ARTHROSCOPY, HE NOTED THAT THE UNIT WAS ALLOWING BACKFLOW OF SUCTION WASTE FLUID FROM THE NEPTUNE WASTE CANNISTER AND WAS ENTERING THE WOUND FROM THE SHAVER SUCTION. THE SHAVER WAS IN THE PT'S KNEE JOINT WHEN LOSS OF SUCTION WAS RECOGNIZED. EVERYTHING WAS IMMEDIATELY REMOVED. UPON PULLING THE SUCTION TUBING FREE FROM PORT, FLUID CONTINUED TO POUR FROM THE NEPTUNE FILTER PORT. WOUND WAS IRRIGATED WITH 9000CC OF ANTIBIOTIC IRRIGANT AND INTRAOP IV ANTIBIOTICS GIVEN. THIS MACHINE HAD 2 PREVIOUS PTS CONTAMINATED FLUIDS IN THE CANNISTER. PT WAS TREATED WITH HOME ANTIBIOTICS FOR 2 WEEKS, HIV/HEP PROPHYLAXIS AND LAB TESTING X 1 YEAR. STRYKER TESTING MACHINE, FOUND UNUSUAL PIECE OF PLASTIC IN CANNISTER. WE HAD ALSO NOT CHANGED MANIFOLD PROPERLY. | 0
985927 | EVALUATION IN PROGRESS BUT NOT CONCLUDED. >< DURING PREPARATION FOR USE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE ARTERIAL MONITORING MODULE DID NOT POWER ON. THE DEVICE WAS REPLACED AND THE PROCEDURE CONCLUDED WITHOUT INCIDENT. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT AS A RESULT OF THIS EVENT. | 0
1253323 | THE MACHINE WAS TESTED BY THE LOCAL BIOMED WHO REPORTED THAT IT PASSED ALL TESTS, AND WAS RETURNED TO USE. THE ANESTHESIOLOGIST AND DIRECTOR OF THE SICU (SURGICAL ICU) STATED THAT HE DID NOT BELIEVE AT THIS TIME THAT ANY MACHINE MALFUNCTION WAS PRESENTED OR CONTRIBUTED TO THE NEGATIVE PATIENT OUTCOME. THE INVESTIGATION OF THE MANUFACTURER IS ONGOING. THE RESULTS WILL BE REPORTED IN A FOLLOW UP REPORT. >< IT WAS REPORTED THAT DURING A CASE, THE DOCTOR WAS HAVING DIFFICULTY OXYGENATING THE PATIENT AND NOTICED THAT THE VENTILATOR WAS NOT RUNNING. IT WAS REPORTED THAT THE RED ALARM LEDS WERE NOT LIT AND THE MACHINE WAS NOT ALARMING. IT WAS REPORTED THAT THE DOCTOR WAS ABLE TO ACTIVATE MAN/SPONT MODE BY PRESSING THE MAN/SPONT MODE KEY AND WAS ALSO ABLE TO ACTIVATE THE VOLUME MODE BY PRESSING THE VOLUME MODE KEY. THE CASE WAS COMPLETED USING THIS MACHINE. IT WAS REPORTED THAT THE PATIENT DID NOT REGAIN CONSCIOUSNESS AND EXPIRED THE NEXT DAY. | 0
993111 | THE CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH THE GALILEO (EXTEND DN ASSAY) ON A PATIENT SAMPLE. THE CUSTOMER LATER STATED THAT THIS PROBLEM OCCURRED WITH 2 DIFFERENT SAMPLES. >< THE CUSTOMER'S RETURNED SAMPLE WAS TESTED WITH RETENTION CAPTURE-R READY-ID EXTEND II, LOT DN023, ON AN IN-HOUSE GALILEO. WEAK (1+) REACTIVITY WAS OBSERVED WITH CELL 14 (D+). THE SOURCE OF THE ANTI-D IS DUE TO ADMINISTRATION OF RHIG. | 0
900346 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER, WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMER AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED, AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
952507 | THIS IS A PICTURE ARCHIVING COMMUNICATIONS SYSTEM. THE CUSTOMER WAS POST PROCESSING PT B'S STUDY AT THE VIEWFORUM AND FOUND THE IMAGE FROM PT A. THE LAST IMAGE OF PT A'S STUDY WAS FOUND TO BE ALSO THE FIRST IMAGE OF PT'S B STUDY. NO PT WAS INJURED OR MISDIAGNOSED. >< (EVAL METHOD): CONCLUSIONS: THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED, A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. | 1
1034932 | EXAMINATIONS WERE DISCOVERED AT THIS SITE THAT DID NOT PRESENT THE PROPER STUDY DATE IN THE GE PICTURE ARCHIVING AND COMMUNICATION SYSTEM (PACS). THE EXAMINATIONS SHOW THE DATE AT WHICH THEY WERE ACQUIRED INTO THE PACS SYSTEM INSTEAD OF THE DATE, THE STUDY WAS ACTUALLY PERFORMED. THIS SITUATION CAN GIVE THE APPEARANCE THAT THE PT'S IMAGES WERE TAKEN MORE RECENTLY THAN WHAT HAD ACTUALLY OCCURRED. THE SITUATION WAS IDENTIFIED AT THE SITE WHILE MIGRATING IMAGES INTO THE PACS SYSTEM AND THERE WAS NO REPORT OF THIS SITUATION RESULTING IN ANY ADVERSE PT OUTCOME. THE SITE EXPRESSED CONCERN THAT IF THE STUDY DATE DISCREPANCY WAS NOT NOTICED, IT COULD POTENTIALLY ALLOW A CLINICAL DECISION TO BE BASED OFF INCORRECT INFO. | 1
1047865 | MCKESSON MEDICAL IMAGING COMPANY ENGINEERING EXAMINATION REVEALED A SITUATION ON HORIZON MEDICAL IMAGING 11.0.X SOFTWARE ON WHICH THE IP LIBRARY HAS A REFRESH ERROR CAUSING ANNOTATION(S) MOVING IN THE VIEWPORT. THIS ERROR OCCURS ON AN IMAGE THAT HAS NON SQUARE PIXELS AND ANNOTATION(S), AND THE FIRST OPERATION APPLIED TO THAT IMAGE AFTER APPLYING A CUSTOM DISPLAY PROTOCOL IS ADJUST WINDOW/LEVEL. THE MCKESSON AFFECTED PRODUCTS ARE HORIZON RAD STATION ADVANCED OR HORIZON RAD STATION DISTRIBUTED VERSIONS 11.0.3 THROUGH 11.0.8 THAT IS PART OF HORIZON MEDICAL IMAGING 11.0.3 THROUGH 11.0.8 WORKSTATIONS. SOFTWARE UPDATES FOR HORIZON MEDICAL IMAGING 11.0.X WORKSTATIONS THAT ELIMINATE THE REPORTED PROBLEM WERE RELEASED IN 2008. SINCE THIS RELEASE DATE, ALL THE CLINICAL SITES WITH THE ABOVE LISTED PROGRAMS HAVE BEEN IN PROCESS OF BEING UPDATED WITH THE ABOVE LISTED PROGRAM UPDATES. >< THE ANNOTATIONS FOUND ON AN IMAGE WITH NON SQUARE PIXELS ARE SLIGHTLY MOVED IF THE FIRST OPERATION APPLIED TO THAT IMAGE AFTER APPLYING A CUSTOMER DISPLAY PROTOCOL IS ADJUST WINDOW/LEVEL. | 1
1048544 | PATIENT WAS IN VTACH/VFIB. THERE WAS NO AUDIBLE ALARM HEARD. PATIENT'S ICD FIRED REPEATEDLY (X3), AS IT SHOULD HAVE, BUT PATIENT'S RHYTHM DID NOT CONVERT ON 3RD FIRING. WE WERE NOT AWARE OF THESE OCCURRENCES UNTIL PATIENT EXPIRED. WE HAVE HAS SYSTEM TESTED AND FOUND THAT IT MEETS MANUFACTURER'S SPECIFICATIONS. ALL ROOMS AND MONITORS TESTED IN THIS PROCESS, AND DETAIL IS AVAILABLE IF REQUESTED. | 1
1020512 | THE METER HAS BEEN CONFIRMED EXHIBIT THE MEMORY OVERWRITE MALFUNCTION CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
1024634 | ONE PATIENT SAMPLE WITH DISCREPANT URINE GLUCOSE RESULTS. INITIAL RESULT WAS NORMAL; VISUAL READING SAMPLE WAS 250 MG/DL. SAMPLE RETESTED GIVING NORMAL RESULT (USING ANALYZER); VISUAL READING GAVE 250 MG/DL. A FINAL VISUAL READING OF SAMPLE WITHOUT USING ANALYZER GAVE 250 MG/DL. THE ERRONEOUS RESULT WAS REPORTED AND PT WAS TREATED WITH THE ANTIBIOTIC BACTRIM, BUT WAS NOT ADVERSELY AFFECTED. IF ADDITIONAL INFO IS REC'D, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
979249 | ONE PT URINE SAMPLE WITH DISCREPANT RESULTS. INITIAL URINE GLUCOSE RESULT NORMAL. SAME SAMPLE PREVIOUSLY TESTED USING A DIFFERENT ANALYZER, DIFFERENT METHOD, GAVE RESULT OF>=1000 MG/DL. VISUAL DIPSTICK OF SAME SAMPLE SHOWED NORMAL FOR GLUCOSE. SAME SAMPLE WAS REPEATED AGAIN USING THE METHOD WHICH INITIALLY GAVE A NORMAL RESULT. THE RESULT WAS 1000 MG/DL ANOTHER VISUAL DIPSTICK TEST WAS PERFORMED, SHOWING 3+. ERRONEOUS RESULT WAS REPORTED. PT NOT ADVERSELY AFFECTED. IF ADD'L INFO IS RECEIVED, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
979265 | DURING PREPARATION FOR USE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE CENTRAL CONTROL MONITOR OF THE HEART LUNG CONSOLE CONTINUOUSLY REPEATED ITS POWER-ON INITIALIZATION SEQUENCE. THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT AS A RESULT OF THIS EVENT. >< EVALUATION IN PROGRESS, BUT NOT CONCLUDED. | 1
872098 | BECAUSE OF THE REMOTE POSSIBILITY OF INJURY RESULTING FROM A SLIP AND FALL EVENT CAUSED BY WASTE FLUID OVERFLOW, THE FIRM IS IN THE PROCESS OF IMPLEMENTING AN IMPROVED WASTE SENSOR DESIGN AS A REPLACEMENT TO THE CURRENT SENSOR. THIS NEW WASTE SENSOR DESIGN SHOULD ELIMINATE THE OBSERVED FAILURE MODE. THE SITE DECLINED TO HAVE THE NEW UPGRADED SENSOR INSTALLED AND HAS ASSUMED THE RISK OF FUTURE SPILLS. COMPLETE FIELD IMPLEMENTATION OF THE NEW SENSOR DESIGN IS EXPECTED TO BE COMPLETED BY 2007. THE NEW WASTE SENSOR DESIGN WAS INTRODUCED TO THE MANUFACTURING FLOOR ON AUGUST 3RD, 2006. A FINAL REPORT WILL BE SENT PENDING THE COMPLETION OF THE FIELD IMPLEMENTATION OF THE NEW WASTE SENSOR. >< A CUSTOMER SITE REPORTED, WASTE FLUID HAD OVERFLOWED FROM THE WASTE CARBOY ON A BENCHMARK XT INSTRUMENT ONTO THE FLOOR WITHOUT AN ALARM. NO ONE WAS INJURED AND PT CARE WAS NOT AFFECTED. THE SITE DECLINED TO HAVE AN UPGRADED SENSOR INSTALLED. THE FIELD SERVICE ENGINEER (FSE) CHECKED THE WASTE SENSOR AND IT FUNCTIONED AS INTENDED. THE FSE BELIEVES THE LEAK TO BE CAUSED BY USER ERROR. THERE WAS NO EVIDENCE OF A MALFUNCTION AS THE LEAK COULD NOT BE DUPLICATED. WHILE THERE IS ALWAYS A POSSIBILITY THAT A SLIP AND FALL EVENT, DUE TO A FLUID SPILL COULD RESULT IN DEATH OR SERIOUS INJURY, ALL EVIDENCE INDICATES THAT THE POSSIBILITY OF DEATH OR SERIOUS INJURY IS REMOTE. | 0
1055465 | CUSTOMER REPORTED THAT DONOR SEGMENTS TYPE AS B POSITIVE ON THE GALILEO, BUT ARE ACTUALLY AB POSITIVE. >< FWD_ABORH TESTING PERFORMED ON IN-HOUSE GALILEO WITH IN-HOUSE DONOR SAMPLES OF VARIOUS ABO AND RH TYPES USING RETENTION ANTI-A, LOT 101679, ANTI-B SERIES 3, LOT 203234, ANTI-D SERIES 4, LOT 504697, AND MONOCLONAL CONTROL, LOT 492028. ALL IN-HOUSE DONOR SAMPLES TYPED AS EXPECTED WITH NO DISCREPANCIES. TESTING PERFORMED WITH RETENTION ANTI-A, LOT 101673, 101674 WITH IN HOUSE DONOR SAMPLES REVEALED TYPE A RED CELL STICKING IN THE ANTI-A WELLS. TO PREVENT THE STICKING, IN HOUSE STUDIES FOUND THAT PIPETTING THE REAGENTS INTO THE WELLS BEFORE THE SAMPLE RED CELLS GREATLY REDUCED THE STICKING. IN HOUSE COMPARISON TESTING OF THE NEW REAGENT-FIRST PIPETTING ORDER TO THE EXISTING SAMPLE-FIRST PIPETTING ORDER SHOWED NO UNEXPECTED NEGATIVE ANTI-A REACTIONS WITH THE NEW PIPETTING ORDER. SOME UNEXPECTED NEGATIVE REACTIONS WERE OBSERVED DURING TESTING WITH THE CURRENT ASSAY. CUSTOMERS WERE NOTIFIED OF THE MODIFICATION TO THE ASSAY ON 9/20/07.THIS FACILITY DID NOT HAVE THE MODIFIED ABO ASSAY SOFTWARE INSTALLED. ACCORDING TO THE GALILEO OPERATOR MANUAL,FORWARD ONLY ABO-RH TESTING HAS A HIGHER RISK OF MISTYPE DUE TO THE ABSENCE OF THE REVERSE TYPE RESULTS. HAZARDOUS MISTYPES MAY OCCUR, SUCH AS A SAMPLE BEING INTERPRETED AS A GROUP AB, OR AN RH (D) NEGATIVE SAMPLE BEING INTERPRETED AS RH(D) POSITIVE. FOR THIS REASON, ABO-RH RESULTS SHOULD ALWAYS BE COMPARED TO THE PATIENT OR DONOR'S HISTORY. | 0
1232005 | IN 2008, THE STAFF AT THE HOSPITAL DISCOVERED A TELEMETRY PT SLUMPED IN TOILET AREA IN CARDIAC ARREST. THE STAFF RESUSCITATED THE PT AND TRANSFERRED TO INTENSIVE CARE. THE STAFF ALLEGED THAT NO PRIORITY ALARM WAS SOUNDED DESPITE VENTRICULAR FIBRILLATION BEING OBSERVED ON CARDIAC MONITOR. THIS INCIDENT WAS REPORTED TO MHRA AND FORWARDED TO SPACELABS IN 2008. >< THIS INCIDENT APPEARS TO HAVE BEEN REPORTED TO MHRA 10 DAYS AFTER THE INCIDENT OCCURRED. SPACELABS FIRST LEARNED ABOUT THIS INCIDENT THROUGH THE INFO FORWARDED BY MHRA. THE HOSPITAL REPORTED THAT THEY TESTED THE DEVICE AND TELEMETRY CHANNEL AFTER THE INCIDENT AND BEFORE THE DEVICE WAS PLACED BACK INTO SVC. THE HOSPITAL'S MEDICAL ENGINEER WAS ABLE TO CONFIRM THAT THE DEVICE SUCCESSFULLY TRIGGERED ALL ALARM CONDITIONS. SPACELABS RESPONDED BY VISITING THE HOSP WITHIN 24 HOURS OF BEING NOTIFIED OF THE INCIDENT. SPACELABS CONDUCTED SIMULATOR TESTING OF THE SUSPECT DEVICE, AND ALL OTHER SIMILAR TELEMETRY DEVICES. SPACELABS TESTING DEMONSTRATED THAT ALL EQUIPMENT WAS OPERATING TO SPECS, ALL ALARMS WERE FULLY FUNCTIONAL. SPACELABS WAS NOT ABLE TO REVIEW ANY RELEVANT ALARM OR HEART RATE RECORDS AS THE HOSP HAD THE PRINTER OFF DURING THE INCIDENT. SPACELABS DID ATTEMPT TO DETERMINE IF THIS INCIDENT WAS RELATED TO THE CURRENT TELEMETRY FIELD CORRECTIVE ACTION. OUR INVESTIGATION WAS INCONCLUSIVE. THE CONDITIONS BEING ADDRESSED IN THE RECALL ARE A FAILURE OF THE TELEMETRY RECEIVER'S ALARM FUNCTION THAT IS ONLY CORRECTABLE BY REMOVING POWER FROM THE RECEIVER MODULE FOR MORE THAN 10 MINUTES. DUE TO THE 10-DAY DELAY IN BEING NOTIFIED OF THIS MATTER, WE ARE UNABLE TO ABSOLUTELY CONFIRM THAT THE RECEIVER HAD NOT BEEN RESET. ALL TESTING DID CONFIRM THAT ALARMS WERE FUNCTIONING PROPERLY SET. ALL TESTING DID CONFIRM THAT ALARMS WERE FUNCTIONING. NO FURTHER ACTION IS ANTICIPATED. | 1
982398 | THIS INCIDENT APPEARS TO BE AN ISOLATED MALFUNCTION DUE TO USER ERROR. IT STATES IN THE USER GUIDE THAT THE COVER IS TO BE CLOSED WHEN OPERATING THE SYSTEM. POST INFECTIOUS MATERIAL EXPOSURE PRACTICES WERE FOLLOWED (BLOOD SAMPLES DRAWN FROM THE OPERATOR WITH NEGATIVE RESULTS) AS A PRECAUTION. ADDITIONAL BLOOD SAMPLES WILL BE DRAWN FROM THE OPERATOR IN A 6 MONTHS. OLYMPUS IS CONTINUING TO EVALUATE THIS EVENT. IF ACTION IS INDICATED OR OTHER SIGNIFICANT INFO IS REC'D, A FOLLOW-UP TO THIS EVENT WILL BE FILED. >< INFORMATION WAS RECEIVED THAT A LID COVER FELL ON AN OPERATOR WHILE LOADING SAMPLES. THE SAMPLES REPORTEDLY SPLASHED ON THE OPERATOR'S FACE. | 0
897958 | AN OCD FIELD ENGINEER (FE) VISITED THE SITE; REMOVED AND RESEATED ALL ASSOCIATED TUBING, PERFORMED ALIGNMENT TO THE CAMERA AND RUN DIAGNOSTICS TESTS WITH NO ISSUES. THE CUSTOMER PERFORMED AND ACCEPTED QC RESULTS. NO DEFINITIVE ROOT CAUSE WAS DETERMINED. >< THE CUSTOMER OBSERVED DROPLETS OF RED CELLS ON THE FOIL OF THE MTS GEL CARDS ON THE ORTHO PROVUE ANALYZER. THE PROVUE ANALYZER POSTED AN ERROR MESSAGE INDICATING THE INABILITY TO IDENTIFY LIQUID LEVELS. FLUID ON TOP OF THE GEL CARD FOIL CAN LEAD TO CARRY OVER AND / OR CROSS CONTAMINATION, WHICH CAN LEAD TO ERRONEOUS TEST RESULTS AND TRANSFUSION OF INCOMPATIBLE BLOOD. NO ERRONEOUS RESULTS WERE REPORTED. | 0
861371 | INVESTIGATION SUMMARY: QC WAS PERFORMED BEFORE THE EVENT. THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS. THE SAMPLES WERE COLLECTED IN 5CC, VACUTAINER TUBES AND WERE STORED AT AMBIENT TEMPERATURE FOR 2 HOURS PRIOR TESTING. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S LAB: THE FSE INSPECTED THE INSTRUMENT AND DISCOVERED A BROKEN SOLENOID 18 AND MIS-THREADED TUBING AT VL4. THE FSE REPLACED SOLENOID 18 AND PINCH VALVE FOR VL4. THE FSE RAN STARTUP, REPRODUCIBILITY, CONTROLS AND PATIENT SAMPLES WITH ACCEPTABLE RESULTS. AS OF 05/24/07, NO OTHER PROBLEMS WERE REPORTED FROM THIS ACCOUNT. HARDWARE ISSUES ADDRESSED BY THE FSE MAY HAVE CONTRIBUTED TO THIS EVENT. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. >< A CUSTOMER CONTACTED BECKMAN COULTER AND STATED THAT THE COULTER GEN-S INSTRUMENT INTERMITTENTLY RECOVERED APPROX 1 GRAM HIGHER HEMOGLOBIN (HGB) RESULTS ON 4 PATIENTS WHEN COMPARED WITH HGB RESULTS OBTAINED FROM DIFFERENT INSTRUMENTS. THE CUSTOMER INDICATED THAT THE BIAS IN THE HGB RESULTS WERE DISCOVERED WHEN DELTA CHECKS FAILED AND THE REPEATED RESULTS FROM THE BACKUP INSTRUMENTS DID NOT MATCH THE COULTER GEN -S RESULTS. PATIENT A: THE HGB RESULT GENERATED BY THE COULTER GEN-S INSTRUMENT WAS 16.5G/DL, AND 15.5G/DL AND 15.4G/DL FROM THE OTHER 2 INSTRUMENTS. PATIENT B: THE HGB RESULT GENERATED BY THE COULTER GEN-S INSTRUMENT WAS 12.9G/DL, AND 11.9G/DL AND 11.8G/DL FROM THE OTHER 2 INSTRUMENTS. PATIENT C: THE HGB RESULT GENERATED BY THE COULTER GEN-S INSTRUMENT WAS 10.7G/DL, AND 10.1G/DL AND 9.9G/DL FROM THE OTHER 2 INSTRUMENTS. PATIENT D: THE HGB RESULT GENERATED BY THE COULTER GEN-S INSTRUMENT WAS 12.8G/DL, AND 12.0G/DL AND 11.9G/DL FROM THE OTHER 2 INSTRUMENTS. THE ERRONEOUS HGB RESULTS WERE NOT REPORTED OUT OF THE LAB AND THERE WAS NO CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS EVENT. | 0
961645 | SEVERAL DISCUSSIONS WITH PT REVEALED WHAT APPEARS TO BE A SKIN REACTION TO THIS OR OTHER PRODUCTS IN A GERIATRIC PT. RARE BUT KNOWN REACTION TO THIS PRODUCT. >< PT HAD EEG AND DEVELOPED SWOLLEN EYES AND SKIN IRRITATION. VERY SLOW TO SUBSIDE. PUFFINESS AROUND EYES RETURNED. ALSO HAD SCALING AND DRY SKIN AT FORMER ELECTRODE SITES. LOSS OF HAIR. | 0
993151 | OLYMPUS HAS CONDUCTED AN INVESTIGATION OF THIS COMPLAINT. THE INVESTIGATION CONFIRMED THE USER'S CLAIM OF AN OVERWRITTEN PT PROCEDURE NOTE. OLYMPUS' INVESTIGATION REVEALED THAT UNDER CERTAIN UNIQUE AND RARE CONDITIONS, THE PROCEDURE NOTES DICTATED FOR A PT WITH THE CHARTNOTE MODULE COULD BE OVER-WRITTEN BY THE PROCEDURE NOTE OF ANOTHER PT WHEN THE PROCEDURE NOTE IS BEING SAVED TO A DATABASE. THERE IS NO CORRUPTION OR OVER-WRITE OF PT IMAGES BY THIS MALFUNCTION. ARCHIVED PT RECORDS ARE PROPERLY IDENTIFIED AND MAINTAINED IN THE DATABASE FILES AND PT IMAGES ARE CORRECTLY ARCHIVED AND MAINTAINED IN THE DATABASE FILES. THIS MALFUNCTION IMPACTS ANY CHARTNOTE PROCEDURE NOTE DICTATED BY THE PHYSICIAN AND SAVED IN A DATABASE. THE IMAGE MANAGER SOFTWARE, CLEARED, IS A DIGITAL IMAGE STORAGE DEVICE UTILIZED WITH ENDOSCOPY IMAGING SYSTEMS TO CAPTURE, LABEL, STORE, SEARCH, REPORT AND PRINT ENDOSCOPIC IMAGES. CHARTNOTE IS AN OPTIONAL MODULE TO THE IMAGE MANAGER SOFTWARE THAT ENABLES THE PHYSICIAN TO DOCUMENT AND MAINTAIN PROCEDURE NOTES FOR A PT ENDOSCOPY EXAM. OLYMPUS IS UPGRADING 666 CUSTOMERS, INCLUDING THE SUBJECTS FACILITY, DUE TO THE POTENTIAL FOR A DICTATED PROCEDURE NOTE TO BE OVERWRITTEN BY THE PROCEDURE NOTE OF ANOTHER PT. NOT ALL DATABASES WILL HAVE OVERWRITTEN PROCEDURE NOTES. IN DATABASES THAT OAI HAS EXAMINED THAT HAVE THIS OVERWRITE MALFUNCTION, THE OCCURRENCE OF THIS MALFUNCTION RANGES FROM 0.05% TO 0.002% OF THE DATABASE TOTAL. THESE 666 DEVICES WERE DISTRIBUTED FROM 1997 TO PRESENT. ATTACHED PLEASE FIND A COPY OF THE CUSTOMER COMMUNICATION AS WELL AS THE LIST OF THE CONSIGNEES. NO INJURIES OR ILLNESS HAVE OCCURRED WITH THE USE OF THIS DEVICE. >< THE HOSPITAL REPORTED THAT THE ARCHIVED CHARTNOTE PROCEDURE NOTE OF A PT WAS OVERWRITTEN BY THE PROCEDURE NOT OF ANOTHER PREVIOUS PT. CHARTNOTE IS AN OPTIONAL MODULE OF THE IMAGE MANAGER SOFTWARE THAT ENABLES THE PHYSICIAN TO DOCUMENT AND MAINTAIN PROCEDURE NOTES FOR A PT ENDOSCOPY EXAM. | 1
964974 | DRAEGER MEDICAL SYSTEMS, INC. HAS RECEIVED INFORMATION FROM A CUSTOMER THAT OUR INFINITY CENTRAL STATION DISPLAY, INCLUDING WAVEFORMS ON THE MAIN SCREEN BECAME FROZEN AND REMAINED FROZEN FOR AN UNK LENGTH OF TIME. THE HOSPITAL BIOMED TRIED TO TELNET INTO THE DEVICE, BUT WAS UNSUCCESSFUL, A MANUAL RESET WAS PERFORMED ON THE INFINITY CENTRAL STATION USING THE POWER SWITCH ON THE FRONT OF THE UNIT. AFTER REBOOTING, THE INFINITY CENTRAL STATION FUNCTIONED NORMALLY. NO PATIENT INJURY WAS REPORTED. >< THE COMPLAINT IS CURRENTLY UNDER INVESTIGATION. | 1
1122170 | BREAKING OF THE BLUE TAB DURING USE. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. | 0
1080765 | UPON THE INITIAL CONTACT TO BAXA, THE CUSTOMER PROVIDED THE ASSOCIATED DOCUMENTATION THAT IS PRODUCED WHEN A BAG IS COMPOUNDED ON THE MICROMACRO 12 (MM12). A REVIEW BY BAXA TECHNICAL SUPPORT OF THE ABACUS FORMULA AND SOLUTION LABELS, MIX CHECK REPORT AND THE MM12 BLACK BOX DATA (THE MM12 BLACKBOX DATA IS AN ONGOING DIALOGUE OF SYSTEM RECORDED BARCODE READS, SYSTEM LOGINS, AND OTHER SIGNIFICANT OPERATIONAL EVENTS STORED IN A LOG FILE FORMAT ON THE MM12 SYSTEM) FOR THE SUBJECT BAG SHOWED THAT THE MM12 COMPOUNDER PERFORMED AS DESIGNED AND DELIVERED THE VOLUME OF CALCIUM AS ORDERED IN THE ABACUS SOFTWARE. INVESTIGATION HAS CONFIRMED THAT ABACUS SOFTWARE AND MICROMACRO 12 COMPOUNDER PERFORMED AS DESIGNED AND DELIVERED THE VOLUME OF CALCIUM AS ORDERED IN THE ABACUS SOFTWARE. >< IN 2008, BAXA WAS NOTIFIED OF AN INCIDENT WHERE A PT EXPERIENCED AN 'ADVERSE EFFECT' DURING ADMINISTRATION OF A TPN BAG. THE PATIENT'S LAB RESULTS INDICATED ELEVATED LEVEL OF CALCIUM. THE CUSTOMER REPORTED THAT THEY WERE AIMING FOR 3.91% CONCENTRATION OF CALCIUM IN THE TPN ORDER AND GOT A READING OF 4.2% CONCENTRATION OF CALCIUM. BAXA WAS NOT INFORMED OF HOW THE CALCIUM CONCENTRATION WAS TESTED WITHIN THE TPN BAG. UPON FOLLOW-UP WITH THE CUSTOMER, THE CUSTOMER COULD NOT PROVIDE ANY ADDITIONAL INFO, EXCEPT TO SAY THAT THEY HAD RULED THE COMPOUNDER OUT AS A SOURCE OF THE EVENT. UPON REVIEWING THE PATIENT'S LAB RESULTS, ALL OF THE PATIENT'S ELECTROLYTES WERE OUT OF RANGE AND THAT THE ADVERSE EFFECT' PERTAINED TO THE PATIENT'S CLINICAL CONDITION. | 1
1013232 | DEVICE HAS NOT RETURNED FOR EVALUATION. >< IT WAS REPORTED THAT THE ARTERIAL LINE BACKED UP BLOOD INTO THE RESERVOIR AND PUSHED THE RESERVOIR PLUNGER BACK TO ABOUT THE 1 ML MARK. PROBLEM PERSISTED DESPITE THE ASSURANCE THAT THE PLUNGER WAS FULLY SEATED AND LINES FLUSHED. | 0
1120543 | PHILIPS IS IN THE PROCESS OF OBTAINING MORE INFO CONCERNING THIS EVENT AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SENT ONCE THE INVESTIGATION IS COMPLETED. >< THE CUSTOMER REPORTED THAT A PT WITH AN IMPLANTED PACER, EXPIRED. HOWEVER, THE MONITOR STILL SHOWED A HEART RATE OF 60 AND DID NOT ALARM. | 1
1252366 | CONVATEC FLEXI SEAL FECAL MANAGEMENT SYSTEM, FOR FECAL DIVERSION LED TO RECTAL BLEEDING REQUIRING SURGICAL INTERVENTION AND OVERSEW OF BLEEDING RECTAL VESSELS AND RIGID PROCTOSCOPY. DATES OF USE: 2008. DIAGNOSIS OR REASON FOR USE: INCONTINENCE OF DIARRHEA. EVENT ABATED: YES. | 0
977325 | THE OPERATOR HAD THE ISOCENTER CENTROID DISTANCE SET A LARGE DISTANCE AWAY FROM THE MARKERS. THE SOFTWARE WAS SET TO 'ROTATION AND TRANSLATION' MODE. IF THE TARGET POINT IS NOT IN THE SAME ANATOMY (ORGAN) AS THE MARKERS, IT MAY MOVE DIFFERENTLY THAN THE MARKER DOES AND RECOMMEND SHIFTS WHICH DO NOT MATCH WHAT THE CLINICIAN EXPECTS. REVIEW OF THE PATIENT MARKER LOCATIONS WITH THE IDENTIFIED ISOCENTER PLACEMENT INDICATED THE SOFTWARE SHOULD RECOMMEND A SHIFT OF 1.2 CM. SWITCHING THE SOFTWARE TO 'TRANSLATION ONLY' MODALITY ALLOWED THE SOFTWARE TO CORRECTLY CALCULATE A RECOMMENDED SHIFT OF 1.17 CM. A NOTIFICATION LETTER WAS GENERATED AND MAILED OUT TO ALL USERS ON 4/14/06 TO HELP USERS UNDERSTAND WHICH MODALITY TO UTILIZE FOR OPTIMAL TREATMENT CALCULATIONS. [SEE SCANNED PAGES.] >< CUSTOMER REPORTED THAT THE SOFTWARE WAS RECOMMENDING SHIFTS TO THE PATIENT FOR TREATMENT. HOWEVER, THE CLINICIAN DID NOT FEEL THAT SHIFTS WERE NECESSARY BASED ON THE SEED PLACEMENT. THIS WAS A NEW PATIENT SET-UP, AND NO TREATMENT HAD BEEN CONDUCTED ON THE INDIVIDUAL. | 1
977324 | CUSTOMER REPORTED THAT THEY USE TWO LATERAL OBLIQUE IMAGES DURING TREATMENT. BY ERROR, THEY SWAPPED THE LEFT AND RIGHT IMAGES BY MISTAKE. THIS WAS IDENTIFIED PRIOR TO ANY TREATMENT, BUT THERE IS A POSSIBILITY OF A MISTREATMENT TO OCCUR. THE SOFTWARE DOES NOT HAVE A WAY TO DETECT THIS. IF THE USER DOES NOT VERIFY THE IMAGES PRIOR TO TREATMENT, THE SOFTWARE CAN POSITION THE PATIENT IN AN INCORRECT LOCATION FOR TREATMENT. >< INSTRUCTIONS FOR USE ON THIS DEVICE INFORM THE USER HOW TO SELECT THE CORRECT ORIENTATION OF THE IMPORTED PICTURE OR TO CHANGE AS APPROPRIATE. THIS IS INTENDED AS A VERIFICATION OF THE PATIENT'S ORIENTATION. | 1
870707 | DRAEGER MEDICAL SYSTEMS, INC. HAS RECEIVED INFORMATION OF A PRODUCT PROBLEM INVOLVING OUR DELTA PATIENT MONITOR AND SP02 SENSOR. THE CUSTOMER REPORTED THAT THE DRAEGER REUSABLE SPO2 CONNECTED TO OUR DELTA MONITOR WAS FOUND LYING ON FLOOR AND THE SPO2 CURVE WAS DISPLAYED UNCHANGED ON THE MONITOR. THE SPO2 VALUE WAS 95, EVEN THOUGH THE SENSOR WAS NOT ON THE PATIENT'S FINGER AND NO ALARM WAS INDICATED ON THE MONITOR. THE CUSTOMER ALSO REPORTED THAT THE CONDITION WAS NOT REPRODUCIBLE WHEN A SIMULATION OF A SENSOR DETACHMENT WAS PERFORMED. NO PATIENT INJURY WAS REPORTED. | 1
996370 | PT HAD INTRATHECAL OPIATE SYSTEM IN PLACE IN MID DEC. HE WAS DOING WELL FOR 1 1/2 WEEKS AND THEN EXPERIENCED SYMPTOMS OF ACUTE OPIATE WITHDRAWAL. THE PT EXPERIENCED REPEAT EPISODES OF WITHDRAWAL SYMPTOMS. THE PT WAS OFFERED A REVISION OF THE CATHETER, IN ORDER TO CORRECT THE INTERMITTENT NON-DELIVERY OF THE OPIATE, AND HE CONSENTED. THE PUMP WAS REMOVED AND ASPIRATION FROM THE CATHETER YIELDED NO CSP, AND IT WAS FELT THAT THE CATHETER MAY BE KINKED. FURTHER INVESTIGATION REVIEWED THE CATHETER WAS FRACTURED AT APPROXIMATELY THE 15CM MARK. THE SURGEON CONTACTED A MEDTRONIC NEUROSURGEON FOR ADVICE ON HOW TO PROCEED. THE NEUROSURGEON AT MEDTRONIC AND OTHER ONE STATED TO DATE NO COMPLICATIONS HAVE BEEN REPORTED WITH A PIECE OF THE CATHETER REMAINING INTRATHECALLY AND TO REMOVE IT WOULD INVOLVE AN EXTENSIVE SURGICAL PROCEDURE. A NEW CATHETER WAS INSERTED AND CHECKED FOR FUNCTION. | 0
1047719 | A SEDECAL X PLUS LP STAND, BEING PART OF IMAGING DYNAMICS CO LTD. XPLORER 1600 SYSTEM, WAS BEING POSITIONED FOR A CHEST X-RAY. THE TECHNOLOGIST WAS LOWERING THE STAND WITH THE ARM IN THE HORIZONTAL (CHEST) POSITION. THE STAND APPEARED TO ENCOUNTER SOME SORT OF OBSTRUCTION AND ACTED AS IF THE ANTI-CRUSH MECHANISM HAD BEEN ACTIVATED. IT STOPPED DOWNWARD MOTION, BRIEFLY REVERSED AND CAME TO A STOP. AGAIN THE DOWN BUTTON WAS PRESSED AND THE STAND RESUMED DOWNWARD MOTION. WHEN THE BUTTON WAS RELEASED THE STAND DID NOT STOP BUT CONTINUED DOWNWARD UNTIL IT HIT THE FLOOR. THE STAND MOTOR CONTINUED TO RUN AFTER THE DETECTOR HIT THE FLOOR AND THEREBY UNSPOOLED THE SUPPORT CABLE. NO INJURIES OCCURRED. >< EVAL SUMMARY: THE ROOT CAUSE OF THIS INCIDENT HAS BEEN IDENTIFIED AS A CIRCUIT BOARD WHICH CONTROLS THE STAND CONTROLLER. AN UPDATE WILL BE PROVIDED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1016198 | AS REPORTED, THE SCREWS FROM THE GCX MOUNTING BRACKET CAME LOOSE CAUSING DRAEGER MONITOR TO FALL FORWARD. THE GCX MOUNTING BRACKET IS A NON DRAEGER PRODUCT. THE DRAEGER DEVICE DID NOT CAUSE THIS EVENT. WE HAVE NOTIFIED THE MOUNTING BRACKET MFR (GCX CORP) OF THIS EVENT FOR FURTHER INVESTIGATION AND POSSIBLE CORRECTIVE ACTION BY GCX CORP. DRAEGER. >< IT WAS REPORTED THAT DURING TREATMENT OF A PT A NURSE TOUCHED THE DRAEGER MONITOR WHICH FELL FORWARD AND HIT HER IN THE FACE. SHE SUFFERED MINOR CUTS AND BRUISING AND HAD TO BE RELIEVED OF HER DUTIES. A FAULT WAS DIAGNOSED WITH GCX BRACKET DUE TO THE TENSION BOLTS WORKING LOOSE CAUSING THE 5 INCH PLUNGE PLATE TENSIONING MECHANISM TO BECOME LOOSE AND ALLOWING THE DRAEGER MONITOR TO FALL. THE DRAEGER MONITOR WAS MOUNTED ON THE GCX BRACKET. THE NURSE RETURNED TO DUTY THE FOLLOWING DAY. THE BRACKET WAS TENSIONED ON INSTILLATION, BUT SUBSEQUENTLY WORKED LOOSE. | 1
983337 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REGARDING LOW INITIAL CBC RESULTS COMPARED TO RERUN RESULTS GENERATED BY THE COULTER ACT DIFF 2 ANALYZER FOR ONE PT SAMPLE. CUSTOMER RAN A PT SAMPLE AFTER THE INSTRUMENT WAS IDLE AND THE SYSTEM GENERATED LOW CBC RESULTS. UPON RERUN, THE SAME SPECIMEN, ALL CBC PARAMETERS RECOVERED HIGHER. THE INITIAL HEMOGLOBIN (HGB) RESULT GENERATED BY THE ACT DIFF 2 ANALYZER WAS 9.2 G/DL AND THE RERUN RESULTS WERE 11.6 G/DL, 11.5 G/DL AND 11.6G/DL RESPECTIVELY. THE INITIAL RESULT FOR PLATELET (PLT) COUNT WAS 235X10 TO THE 3RD POWER CELLS/UL AND THE RERUN RESULTS WERE IN THE RANGE OF 295-304 X 10 TO THE 3RD POWER CELLS/UL. ERRONEOUS RESULTS WERE NOT REPORTED OUTSIDE OF THE LABORATORY. THERE WAS NO DEATH, INJURY OR CHANGE TO PT TREATMENT AS A RESULT OF THIS EVENT. >< THE SAMPLE WAS DRAWN VIA VENIPUNCTURE AND COLLECTED IN A 4 ML BD CONTAINER. QC IS RUN ONCE A DAY AND WAS RUN BEFORE, BUT NOT AFTER THE EVENT, WITH ACCEPTABLE RESULTS. THE UNIT IS CURRENTLY WORKING WITHIN QC SPECIFICATIONS. DATA PROVIDED BY CUSTOMER SHOWS THAT THE INSTRUMENT WAS IDLE FOR APPROXIMATELY SEVEN HOURS PRIOR TO RUNNING THE INITIAL SAMPLE. PER TECHNICAL SUPPORT ENGINEER, THIS INSTRUMENT DOES NOT REQUIRE A 'SAMPLE PRIME' CYCLE AFTER IDLE PERIODS UNLESS IT IS OVER TWO HOURS AND THE INSTRUMENT ENTERS THE 'SLEEP STATE'. THE INSTRUMENT AUTOMATICALLY PERFORMS A PRIME CYCLE AFTER A TWO HOUR PERIOD WHEN TAKEN OUT OF THE SLEEP STATE BY THE OPERATOR. AT PRESENT, THERE HAVE BEEN NO FURTHER ISSUES OF ERRONEOUS SAMPLE RESULTS FROM THIS CUSTOMER. A FIELD SERVICE ENGINEER (FSE) HAS BEEN DISPATCHED TO THE CUSTOMER'S LAB FOR THIS ISSUES; HOWEVER SERVICE INFO WAS NOT PROVIDED. A CLEAR ROOT CAUSE HAS NOT BEEN DETERMINED TO DATE. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. | 0
947722 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REGARDING LOW HEMOGLOBIN (HGB) AND PLATELET (PLT) RESULTS THAT WERE GENERATED BY THE COULTER AC T DIFF 2 INSTRUMENT FOR A SINGLE PATIENT SAMPLE. A PATIENT SAMPLE WAS TESTED FOR HGB AND PLT AND RESULTS OF 4.2G/DL AND 240 X 10 TO THE THIRD POWER CELLS/UL WERE OBTAINED RESPECTIVELY. THE RESULTS WERE REPORTED OUT OF THE LAB AND THE PATIENT WAS SENT TO A HOSPITAL AND REDRAWN. THE HOSPITAL'S RESULTS WERE HIGHER FOR BOTH ANALYTES; HGB RESULT WAS 9.1G/DL. PLT RESULT WAS 416 X 10 TO THE THIRD POWER CELLS/UL. THE HOSPITAL'S RESULTS WERE CONSIDERED CORRECT. ON THE NEXT DAY, THE CUSTOMER RE-TESTED THE ORIGINAL SAMPLE FOR HGB AND PLT AND RESULTS WERE: HGB RESULT WAS 9.7G/DL. PLT RESULT WAS 385 X 10 TO THE THIRD POWER CELLS/UL. BASED ON INFORMATION PROVIDED, THERE WAS NO AFFECT TO PATIENT OR USER AS A RESULT OF THIS EVENT. >< QC IS RUN ONCE A DAY. QC WAS PERFORMED BEFORE AND AFTER THE EVENT AND RESULTS WERE IN RANGE. THE SPECIMEN WAS COLLECTED IN A BD, 4ML TUBE AND RAN LESS THAN 15 MINUTES AFTER COLLECTION. THE SAMPLE WAS TESTED IN CLOSED VIAL WHOLE BLOOD MODE. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S LAB: A) PER FSE, QC WAS PERFORMING PROPERLY. B) THE FSE INSPECTED FLUIDICS AND VERIFIED MIXING BUBBLES. C) THE FSE PERFORMED REPRODUCIBILITY USING NORMAL CONTROL MATERIAL AND OBTAINED PASSING RESULTS. D) THE FSE VERIFIED THE INSTRUMENT'S PERFORMANCE. A CLEAR ROOT CAUSE HAS NOT BEEN DETERMINED IN THIS EVENT. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. | 0
1002543 | EVAL IN PROGRESS BUT NOT CONCLUDED. >< DURING PREVENTIVE MAINTENANCE, THE DEVICE SPONTANEOUSLY RESET ITSELF WHILE OPERATING, SHUTTING DOWN AND RESTARTING. THE BEHAVIOR COULD NOT BE REPRODUCED. NO PT INJURY WAS REPORTED AS A RESULT OF THIS EVENT. | 1
1045860 | THE CUSTOMER REPORTED, THAT IN 2008 WHILE A PT WAS CONNECTED TO AN MP30 INTELLIVUE MONITOR, THAT THE PT'S SPO2 LEVEL DROPPED TO 21% AND THE MONITOR DID NOT AUDIBLY ALARM. THE PT'S MOTHER CAME OUT AND ASKED A NURSE AND STUDENT NURSE TO COME IN THE ROOM. THE PT WAS VERY PALE AND BLUE AROUND THE MOUTH. THE CUSTOMER ALSO STATED, THAT THIS WAS THE SECOND EPISODE THAT DAY, WHEN THE PT WAS APNEIC. THE FIRST EPISODE WAS APPROX 20:20 AND WAS NOT WITNESSED BY ANY HOSPITAL STAFF. AFTER THE SECOND EPISODE, THE PT WAS TRANSFERRED TO PICU. PHILIPS IS CONSIDERING THAT EMERGENT CARE IS EXPECTED FOR THIS PT CONDITION AND THE CLINICIANS HAD NOT RESPONDED EFFECTIVELY TO THE CHANGE IN THE PT CONDITION, SO WE WILL CONSIDER THIS AS A SERIOUS INJURY, WHERE MONITORING WAS A FACTOR. A PHILIPS FIELD SVC ENGINEER (FSE), WENT ONSITE POST-INCIDENT AND TESTED THE INVOLVED DEVICE AND COLLECT DATA LOGS. THE CUSTOMER'S BIOMED TOLD THE FSE THAT HE TESTED THE MONITOR AND FOUND NO PROBLEM, THE MONITOR TESTED TO SPECIFICATIONS. THE LOG FILES WERE REVIEWED BY A PHILIPS CLINICAL SPECIALIST, WHO DETERMINED THAT THE MP30 MONITOR ALARMED FOR BOTH SPO2 DESATURATION INSTANCES. THE ALARM LOG SHOWS THAT IN 2008, AT 20:14:53 THERE WAS A RED DESAT ALARM REPORTED FOR THIS PT WITH A DESAT LEVEL OF 41, WHICH WAS SILENCED AT THE CENTRAL STATION AT 20:15:08, 55 SECONDS AFTER THE ALARM SOUNDED. THE SECOND REPORTED INCIDENT SHOWS ON THE SAME DAY, AT 23:22:32 THERE WAS A RED DESAT ALARM REPORTED FOR THIS PT WITH A DESAT LEVEL OF 21, FULLY CONSISTENT WITH IT BEING SILENCED AT THE BEDSIDE. FURTHER INVESTIGATION INTO THE CUSTOMER'S CONFIGURATION FILE WAS DONE BY PHILIPS LEARNING PRODUCTS CLINICIAN, WHO DETERMINED THAT THE BEDSIDE ALARM IN THE CUSTOMER'S CONFIGURATION IS SET TO VOLUME 5 (OUT OF A POSSIBLE RANGE OF 01-10) WITH THE LOW VOLUME LIMIT AT 4. NO DEVICE MALFUNCTION OCCURRED. >< THE CUSTOMER REPORTED, THAT THE PT QUIT BREATHING AND THERE WAS NO ALARM. THE NEONATE HAD A SPO2 DESATURATION OF 21 PERCENT. THE CUSTOMER CLAIMED THERE WAS NO AUDIO ALARM GENERATED. | 1
1243610 | NURSE APPROACHED OMNICELL ON NORTH SIDE OF UNIT, JUST AFTER 12 MIDNIGHT - SHE ENTERED HER USER ID WHICH OMNICELL MESSAGE STATED WAS 'INVALID'. THEN, AS SHE REACHED OR POSSIBLY TOUCHED THE 'OK' BUTTON ON SCREEN WITH HER RIGHT INDEX FINGER, SHE FELT A 'HUGE SHOCK' GO THROUGH HER RIGHT THUMB, UP HER ARM AND TO HER BACK. HER THUMB WAS REDDENED, AS IF BURNED. THE OMNICELL WENT DARK, AND THE NURSE WENT TO THE EMERGENCY ROOM FOR EVALUATION. SHE ALSO WENT TO EMPLOYEE HEALTH SERVICE IN THE MORNING. SHE WENT HOME AND WAS TAKEN OFF WORK SECONDARY TO 'BACK SPASMS'. SHE WAS OFF ONE NIGHT, AND CALLED IN THE NEXT NIGHT DUE TO PAIN/BACK SPASMS. SHE IS STILL OFF, AND SEEING EMPLOYEE HEALTH SERVICE TODAY FOR FOLLOW-UP.BIOMED WAS CALLED. THEY INSTRUCTED CLINICIANS TO CALL MAINTENANCE-ELECTRICIAN. ELECTRICIAN CAME UP TO ASSESS PLUGS, ETC. IN THE MORNING, AND THOUGHT 'STATIC ELECTRICITY' CAUSED THE SHOCK. BIOMED CALLED AGAIN, AND CLINICAL ENGINEER CAME UP IN THE AFTERNOON TO EVALUATE.COMPANY (OMNICALL) CAME IN LATER IN THE MORNING AND REPLACED THE 'OMNI BOX' . THE OMNICELL WAS PLACED BACK ON LINE. BIOMED TO FOLLOW UP TO EVALUATE THIS WITH THE COMPANY. | 1
1005952 | THE CUSTOMER REPORTED THAT PATIENTS (NUMBER NOT SPECIFIED) FELT A SHOCK SENSATION DURING THE SCAN WITH THE BIOMETRY PROBE. THEY ALSO REPORTED FINDING OPACIFICATION ON THE CENTRAL CORNEA AFTER BIOMETRY IN A FEW PATIENTS. NO INTERVENTION HAS BEEN REPORTED TO DATE. THEY FOUND EXPOSED METAL WIRES ON THE PROBE CONTACT SURFACE WHEN EXAMINED UNDER A MICROSCOPE, AND SUSPECT THAT THIS IS CAUSING THE PROBLEM. ADDITIONAL INFORMATION HAS BEEN REQUESTED. >< THE CUSTOMER WAS ADVISED TO USE A DIFFERENT MACHINE. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. | 1
1009897 | PER USER FILED MEDWATCH REPORT, ' DATEX-OHMEDA CARDIOCAP/5 MONITORS INTERMITTENTLY DISPLAY INCORRECT SPO2 SATURATION. LEVELS, 5-12% TOO LOW, LEADING TO POTENTIAL FOR PT TREATMENT RELATED TO LOW SATURATION LEVELS.' GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
897506 | DRAEGER MEDICAL SYSTEMS INC. HAS RECEIVED INFORMATION FROM A CUSTOMER THAT REPORTED AN EVENT OCCURRED WHILE MONITORING A PATIENT'S SPO2 VALUES USING OUR SC9000 PATIENT MONITOR. THE CUSTOMER DID INDICATE THAT THE PATIENT WAS ALREADY IN BREATHING DISTRESS AND MULTIVISCERAL FAILURE BEFORE THIS EVENT: AND NO ADVERSE CONSEQUENCES ARE ATTRIBUTED TO THE EVENT. THE PATIENT WAS THEN MONITORED USING OUR DEVICE FOR SPO2 AND THE CUSTOMER QUESTIONED THE HIGHER THAN EXPECTED SPO2 VALUE, WHICH WAS REPORTED TO BE 95%. AS A RESULT, THE CUSTOMER PERFORMED A BLOOD GAS ANALYSIS AS A COMPARATIVE MEANS AND DETERMINED THAT HYPOXY (LOW SAT) WAS NOT DETECTED BY THE SC9000 MONITOR, BUT SHOWN ONLY WITH THE SECONDARY BLOOD ANALYSIS TEST. | 1
924170 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CLINICAL CHEMISTRY ICT CALIBRATOR LOT NUMBER 0505017 HAS SHOWN SEVERAL CALIBRATOR SETS WITHIN THE LOT TO PRODUCE DEPRESSED, BUT ACCEPTABLE POTASSIUM CALIBRATION CURVES THAT FAIL BY QC BEING OUT OF RANGE LOW, WHEN USED IN CONJUNCTION WITH THE ICT MODULE ON THE AEROSET, ARCHITECT C8000 AND ARCHITECT C16000. THIS ISSUE IS NOT SEEN IN ALL SETS OF THE CALIBRATOR VIALS AND IS LIMITED TO THE LOW CALIBRATOR LEVEL AND POTASSIUM VALUES. | 0
951820 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER IS QUESTIONING THE LOWER CALIBRATION SLOPE AND THE OUT OF RANGE LOW QUALITY CONTROL RESULTS OF POTASSIUM GENERATED USING THE CLINICAL CHEMISTRY SERUM ICT CALIBRATOR, LOT# 0505017. THERE WAS NO IMPACT TO PT MANAGEMENT REPORTED. | 0
1058939 | THE CUSTOMER REPORTED THAT WHILE THE CENTRAL STATION WAS IN USE MONITORING PTS ALONG WITH PASSPORT 2'S AND TELEPACKS, THE CENTRAL STATION DISPLAYED THE ERROR MESSAGE 'WINDOWS HAS DETECTED AN ERROR' AND STOPPED MONITORING ALL PTS. NO PT INJURY WAS REPORTED. >< THE COMPANY REP REPLACED THE CENTRAL STATION. AFTER THE NEW CENTRAL WAS INSTALLED, THE SYSTEM WAS TESTED TO FACTORY SPECIFICATIONS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. THE ORIGINAL CENTRAL WAS RETURNED TO THE FACTORY FOR EVAL. THE ERROR LOGS SUGGESTED A PROBLEM WITH THE VIDEO CARD. THE VIDEO CARD IS BEING REPLACED. | 1
1001861 | OUR REP IS REPORTING THAT DURING A PROCEDURE, A PT SUFFERED A MINOR 1ST DEGREE BURN AT THE WORKING PORTAL. THIS BURN WAS PRODUCED FROM A LINVATEC CAUTERY DEVICE, NOT A MITEK PRODUCT. HOWEVER, THERE SEEMS TO BE SOME CONCERN THAT THERE MAY HAVE BEEN SOME SORT OF 'ELECTRICAL' INTERFERENCE FROM AN FMS PUMP THAT CAUSED THE CAUTERY DEVICE TO ACTIVATE INCONVENIENTLY. THE PROCEDURE WAS CONCLUDED SUCCESSFULLY WITHOUT FURTHER INCIDENT OR HARM TO THE PT. THERE WERE 2 FMS PUMPS IN THE ROOM AND NEITHER THE SURGEON NOR THE STAFF CAN IDENTIFY WHICH ONE OF THE 2 DEVICES IS THE COMPLAINT DEVICE. ALSO SEE ASSOCIATED MDR 1221934-2008-00098. >< THE COMPLAINT DEVICE HAS BEEN RECEIVED AT OUR EVALUATION AND REPAIR FACILITY. THE DEVICE WILL GO THROUGH COMPLETE VISUAL AND FUNCTIONAL TESTING. THE RESULTS OF THE TESTING WILL BE EVALUATED. OUR FINDINGS WILL BE THE SUBJECT OF A FOLLOW-UP REPORT. | 0
1069176 | IT WAS REPORTED THAT AT APPROX 7:00AM IN 2008, A TELEMETRY PT EXPIRED WHILE THE TELEMETRY TRANSMITTER (IDENTIFIED IN LOGS AS T8) WAS IN STANDBY MODE. THE CUSTOMER HAS REQUESTED ANALYSIS OF THE MVWS LOG FILES. THE CUSTOMER HAS REQUESTED CLARIFICATION OF THE EFFECT OF REPEATEDLY CHANGING THE TELEMETRY TRANSMITTER STATUS INTO AND OUT OF STANDBY ON EVENT DISCLOSURE DATA STORAGE. >< OUR EVAL OF THE INFO AVAILABLE AT THIS TIME INDICATES THAT OUR DEVICE DID NOT CAUSE OR CONTRIBUTE TO THIS ADVERSE EVENT. HOWEVER, WE WILL CONTINUE FURTHER INVESTIGATION AND IF FOUND THE RESULTS TO BE DIFFERENT, WE WILL REPORT TO FDA IMMEDIATELY. DRAEGER DEVICE WORKED AS INTENDED. | 1
1150083 | AS IT IS DESIGNED, THERE SHOULD BE A 5 SECOND DELAY BETWEEN THE ALARM CONDITION AND WHEN THE NOTIFICATION PAGE IS SENT. IF THERE IS ONLY 1 PAGER ASSIGNED TO THE PULSE OXIMETER, THE SYSTEM BEHAVES OK. IF THERE ARE MULTIPLE PAGERS (I.E. ASSIGNED BACKUP NURSE) THERE IS A 7 SECOND DELAY BETWEEN EACH PAGE REQUEST). DEPENDING UPON WHICH PAGER THE PRIMARY NURSE WAS HOLDING, THERE COULD BE A SEVERAL SECOND DELAY BEFORE THEY GET THE INITIAL PAGE. ANOTHER AREA OF THE HOSPITAL IS RUNNING THE SAME SOFTWARE VERSION WITH THE SAME ISSUE. | 1
1034130 | IT WAS REPORTED THAT THE TUBING DETACHED AT THE IV SIDE NEAR THE FLUSH TAB. THIS WAS NOTED WHEN NURSE HAD WALKED IN TO CHECK ON PT. >< NO DEVICE HAS BEEN RETURNED FOR EVALUATION AT THIS TIME. | 0
1019981 | THE GE MACLAB PHYSIOLOGIC MONITORING FAILED DURING ACUTE INTERVENTION OF ST ELEVATION MYOCARDIAL INFARCTION PREVENTING CONTINUATION OF THE PROCEDURE IN THIS ROOM. THE FOLLOWING MORNING A GE SERVICE ENGINEER STATES THAT THE UPS FAILED. THE UPS SYSTEM IS A FAILSAFE INCORPORATED IN THIS SYSTEM BY GE TO SPECIFICALLY PREVENT EQUIPMENT FAILURE DURING PROCEDURES. IN ADDITION, ALL THE UPS HAD ALSO FAILED. FOR THIS PT DOOR TO BALLOON TIMES WAS DELAYED. FORTUNATELY, A SECOND WAS AVAILABLE TO COMPLETE THE PROCEDURE. | 1
996685 | THE ISSUE INVOLVES FUNCTIONALITY IN CERNER MILLENNIUM POWERCHART POWERORDERS AND AFFECTS USERS THAT UTILIZE THE MEDICATION ADMINISTRATION RECORD. IF A PRO RE NATA (PRN) MEDICATION ORDER IS MODIFIED TO A SCHEDULED ORDER AND BACK TO PRN, THE LAST GIVEN DATE AND TIME ARE DISPLAYED INCORRECTLY FOR THE TASK. THIS ISSUE OCCURS IF AN ADMINISTRATION OR EVENT WAS DOCUMENTED WHILE THE MEDICATION WAS A NON-PRN ORDER. THE SYSTEM DOES NOT RECOGNIZE THIS ADMINISTRATION AS THE PREVIOUS ADMINISTRATION, AND INCORRECT INFO IS DISPLAYED IN THE LAST GIVEN DETAIL. PT CARE COULD BE AFFECTED IF A CLINICIAN BASES THE LAST ADMINISTRATION OFF OF THE INCORRECT LAST GIVEN DATE AND TIME, AS A MEDICATION DOSE COULD BE GIVEN WITHOUT THE PROPER TIME INTERVAL BETWEEN DOSES. CERNER HAS NOT RECEIVED COMMUNICATION OF ANY ADVERSE PT EVENTS AS A RESULT OF THIS ISSUE. >< CERNER HAS DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION 2008 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE AND AN INTERIM WORKFLOW ADJUSTMENT TO PREVENT THE MALFUNCTION. A SOFTWARE MODIFICATION IS BEING DEVELOPED TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORPORATION WILL PROVIDE A FOLLOW-UP REPORT WHEN THE SOFTWARE CORRECTION IS AVAILABLE. | 1
1157270 | SALES REP REPORTED THAT WHILE USING AN ARTHROSCOPIC PUMP WITH THE NEPTUNE ROVER, THE ROVER WAS ALLOWED TO OVERFILL WHICH CAUSED THE SUCTION TO DEACTIVATE. THE ARTHROSCOPIC PUMP CONTINUED TO FORCE FLUID INTO THE ROVER WHICH CAUSED BACKFLOW TOWARD THE SURGICAL SITE. THE SURGEON NOTICED THIS RIGHT AWAY, AND THE PT WAS IRRIGATED WITH 9 LITERS OF NORMAL SALINE, AND ANTIBIOTICS. THE PT ALSO RECEIVED ANCEF, AND VANCOMYCIN AT THE HOSPITAL. THE PT WAS SENT HOME WITH ONE MONTH OF HIV PROPHYLAXIS, AND WILL HAVE LABS DRAWN AFTER ONE YEAR. THE PT WAS ALSO SENT HOME WITH LEVAQUIN FOR TWO WEEKS. >< THE DEVICE WAS BEING OPERATED WITH A 'FLUID VOLUME ERROR.' THIS MEANS THAT THE FLUID VOLUME SENSOR WAS DISABLED WHICH MEANS THAT THE SYSTEM COULD NOT SENSE THE VOLUME LEVEL. THE FIELD TECHNICIAN REPORTED TO HAVE FOUND A SMALL PIECE OF PLASTIC (NOT PART OF THE NEPTUNE SYSTEM) IN THE CANISTER. THIS PIECE OF PLASTIC HAD WEDGED ITSELF BETWEEN THE BOTTOM OF THE CANISTER, AND LEVEL SENSE FLOAT. THE LEVEL SENSE FLOAT IS WHAT DETERMINES THE VOLUME OF THE CANISTER. BASED ON THE SERVICE INFO, FLUID WAS PUMPED BACK TOWARDS THE SURGICAL SITE BECAUSE AN EXTERNAL PUMPING DEVICE WAS ALLOWED TO CONTINUE TO PUMP FLUID WHEN THE ROVER CANISTER WAS OVERFILLED BY THE USER. THE INSTRUCTIONS FOR USE INCLUDE A WARNING STATING 'VERIFY THAT THE ROVER'S FLUID SUCTION IS ACTIVATED WHEN USING AN EXTERNAL PUMPING DEVICE. RUNNING AN UNSUPERVISED EXTERNAL PUMPING DEVICE WHILE THE ROVER'S FLUID SUCTION IS DEACTIVATED MAY RESULT IN PT INJURY. THE ROVER'S FLUID SUCTION DEACTIVATES WHEN THE FLUID COLLECTION CONTAINER IS FILLED AT 20 LITERS.' | 0
1199947 | ZOLL CIRCULATION HAS REQUESTED CONTACT INFO FOR OBTAINING THE AUTOPSY REPORT AND MAKING DIRECT CONTACT WITH THE MEDICAL PROFESSIONALS AND EMT CREW INVOLVED. THE AUTOPSY AND DIRECT CONTACT WILL BE NECESSARY TO COMPLETE THE INVESTIGATION. THE LIFEBAND WAS EVALUATED AND FOUND TO HAVE EVIDENCE OF ROUGH HANDLING AND POTENTIAL IMPROPER USE. THE SIDE CLIPS ON THE CARTRIDGE WERE BROKEN, THE BELT HAS EVIDENCE OF FOLDS, AND THE BELT COVER WAS ABRADED. ZOLL CIRCULATION IS FILING THESE INITIAL REPORTS BECAUSE OUR INVESTIGATION IS ONGOING AND MORE INFO IS NEEDED BEFORE A CONCLUSION CAN BE MADE. WHILE INTERNAL INJURIES SUCH AS THOSE DESCRIBED CAN OCCUR WITH GOOD, VIGOROUS CPR EITHER MANUAL OR AUTOMATIC, THE EXTENT OF THE INJURIES AND THE CONTRIBUTION OF THE DEVICE ARE NOT KNOWN PENDING THE AUTOPSY REPORT AND FURTHER INTERVIEWS OF THOSE INVOLVED. IT IS UNUSUAL THAT INTERNAL INJURIES OCCURRED WITH NO RIB OR STERNAL FRACTURES. ADDITIONALLY, THERE IS SOME EVIDENCE THAT THE DEVICE MAY HAVE BEEN IMPROPERLY DEPLOYED. >< THIS IS AN INITIAL REPORT. IN 2008, ZOLL CIRCULATION RECEIVED A REPORT THAT AN AUTOPULSE RESUSCITATION SYSTEM WAS DEPLOYED ON A PT WHO HAD BEEN FOUND ON THE COUCH COLD, CYANOTIC, NOT BREATHING AND WITH NO PULSE. MANUAL CPR WAS PERFORMED BEFORE THE AUTOPULSE WAS APPLIED. PT DID NOT RESUSCITATE. NO INFO IS AVAILABLE ABOUT THE CONDITION OF THE PT PRIOR TO THE EVENT. ACCORDING TO THE REPORTER, AN AUTOPSY FOUND LIVER, SPLEEN, AND LUNG DAMAGE. NO RIB OR STERNAL FRACTURES WERE REPORTED. ACCORDING TO THE REPORTER, THE AUTOPSY CONCLUDED THAT THE AUTOPULSE HAD CAUSED THE INJURIES, BUT WAS NOT IMPLICATED AS THE CAUSE OF DEATH. THE AUTOPULSE AND THE LIFEBAND DISPOSABLE USED DURING THE EVENT WERE RETURNED TO ZOLL CIRCULATION FOR EVAL. REFERENCE MEDWATCH REPORT # 3003793491-2008-00007 FOR THE AUTOPULSE INFO. | 1
1245769 | I WAS ATTENDING ON A DIRECT CURRENT CARDIOVERSION FOR A MALE PT WITH ATRIAL FLUTTER WHO WAS ON OUR RECENTLY INSTALLED PHILLIPS MONITOR SYSTEM, IN OUR MONITORED STEP-DOWN UNIT. HE WAS ON A MONITOR THAT I WATCHED DURING THE PROCEDURE. AFTER I ADMINISTERED THE SYNCHRONIZED DIRECT-CURRENT SHOCK, THE FLUTTER CONTINUED WITHOUT CHANGE. I WAS SURPRISED BY THIS AND STARTED TO TURN UP THE ENERGY FOR ANOTHER ATTEMPT WHEN THE MONITOR THEN SHOWED THE SHOCK AND CONVERSION -WITH A DELAY OF ABOUT 5 SECONDS. THERE WAS NO HARM IN THIS CASE, BUT THE DELAY WAS SURPRISING AND CLEARLY COULD CAUSE HARM IN OTHER SITUATIONS. FOR EXAMPLE, DURING PACEMAKER THRESHOLD TESTING, WE TURN DOWN THE PACER OUTPUT UNTIL LACK OF CAPTURE IS SEEN IN ORDER TO FIND THE PACING THRESHOLD AND THEN QUICKLY TURN BACK UP THE OUTPUT SO THAT THERE IS NOT A LONG PAUSE WITHOUT A HEARTBEAT; WITH A 5-6 SECOND DELAY ADDED TO THE BRIEF PAUSE AFTER LOSS OF CAPTURE, A PT WOULD LIKELY LOSE CONSCIOUSNESS AND MIGHT EVEN EXPERIENCE VENTRICULAR FIBRILLATION RELATED TO THE SUDDEN ASYSTOLE. I HAD HEARD OF THIS PROBLEM AT ANOTHER HOSPITAL WITH THE SAME BRAND MONITOR. IT IS CLEAR THAT THE HARD-WIRED MONITORS DO NOT HAVE SUCH A LONG DELAY, BUT THE DOCTOR OR NURSE TAKING CARE OF THE PT MAY NOT BE AWARE WHETHER THE PT IS ON TELEMETRY OR NOT. ONE EXPECTS THAT THE MONITOR OVER THE PT'S HEAD IS SHOWING REAL-TIME CARDIAC ACTIVITY. THE COMPANY REP TOLD US THAT THE MONITOR IS WITHIN ACCEPTED STANDARDS AND THAT NO OPTION IS AVAILABLE TO CORRECT THE PROBLEM. OF NOTE, THIS DELAY WAS NEVER MENTIONED DURING NEGOTIATION FOR PURCHASE OR ORIENTATION AFTER THE SALE. IT IS NOTED IN THE EXTENSIVE PRODUCT INFO PROVIDED. OBVIOUSLY, PLACING A HARD-WIRED MONITOR IS ONE SOLUTION, BUT THIS FORCES THE PT TO STAY IN BED AND ADDS RISK ASSOCIATED WITH INACTIVITY. WE REVIEWED THIS THROUGH OUR QI PROCESS AT UNIVERSITY OF WASHINGTON MEDICAL CENTER AND WE HAVE PLACED SIGNS AT EACH MONITOR REMINDING CAREGIVERS THAT A SUBSTANTIAL DELAY MAY BE PRESENT WHEN THE PT IS ON TELEMETRY. THIS IS ONLY A STOP-GAP MEASURE, AND I WISH WE HAD A MONITOR THAT DISPLAYED THE RHYTHM OVER THE PT ON TELEMETRY WITH REAL-TIME OR NEAR REAL-TIME DISPLAY. FURTHERMORE, IN MY OPINION, MONITORS SHOULD BE REQUIRED TO PROVIDE THIS CLINICAL SERVICE. | 1
1051475 | IT WAS REPORTED THERE WAS A LOSS OF AUDIBLE ALARMS AT MULTIPLE CIC'S SUPPORTING TELEMETRY PTS. NO PT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A F/U REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
953792 | UNDER ANOTHER COUNTRY LAW, PT INFO IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSP. >< CUSTOMER REPORTED THAT THE PT MONITOR WENT BLANK DURING A CASE. NO PT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ON GOING. A F/U REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1041224 | THE CUSTOMER REPORTED THAT HIS PRECISION XTRA METER HAD SPONTANEOUSLY CHANGED THE DATE AND TIME. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN HIS METER WERE NOT PROPERLY SET, AND HE REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED. | 0
1001638 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED, RECEIVING AN ERROR 658 DISPLAYING ON THEIR PRECISION XTRA METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MGMT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
1056335 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1001625 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTICED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN ERROR 658 DISPLAYING ON THEIR PRECISION XTRA METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
1001624 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< CUSTOMER REPORTED THAT THE DATE AND TIME SET IN THEIR PRECISION XTRA BLOOD GLUCOSE METER HAD CHANGED SPONTANEOUSLY. THE CUSTOMER REPORTS USING THE P-LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
1001627 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE ADC FA21DEC2006 LETTER. >< CUSTOMER REPORTED THAT THE DATE AND TIME SET IN THEIR PRECISION XTRA BLOOD GLUCOSE HAD CHANGED SPONTANEOUSLY. THIS IS A KNOWN MALFUNCTION WITH THE SOFTWARE THAT CAUSES INCORRECT TRENDING OF GLUCOSE RESULTS. THE CUSTOMER REPORTS USING THE PRECISION LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
1001626 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE INFORMED THROUGH ADC FA21DEC2006 LETTER. >< THE CUSTOMER REPORTED THAT HIS PRECISION XTRA METER HAD SPONTANEOUSLY CHANGED THE DATE AND TIME. HE ALSO REPORTS USING PRECISION LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
993129 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED. >< CUSTOMER REPORTED RECEIVING AN ERRC 6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER UPON TEST STRIP INSERTION, AND NOTICED THE DATE AND TIME HAD CHANGED SPONTANEOUSLY. THE CUSTOMER REPORTS USING THE P-LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
993120 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE LETTER. >< CUSTOMER REPORTED THAT THE DATE AND TIME SET IN THEIR PRECISION XTRA BLOOD GLUCOSE METER HAD CHANGED SPONTANEOUSLY. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
993119 | CUSTOMER REPORTED THAT THE DATE AND TIME HAD CHANGED SPONTANEOUSLY IN THEIR PRECISION XTRA METER. THE CUSTOMER REPORTS USING THE P-LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. >< THE CUSTOMER'S PRODUCT HAS BEEN REQUESTED BACK FOR AN INVESTIGATION. A FOLLOW-UP REPORT WILL BE FILED ONCE INVESTIGATION RESULTS ARE AVAILABLE. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE LETTER. | 0
993117 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE LETTER. >< CUSTOMER REPORTED RECEIVING ERRC 663 MESSAGE WITH THEIR PRECISION XTRA BLOOD GLUCOSE METER WHEN UPON TEST STRIP INSERTION, AND THE DATE AND TIME HAD CHANGED SPONTANEOUSLY. THE CUSTOMER REPORTS USING THE P-LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. | 0
993115 | CUSTOMER REPORTED RECEIVING ERRC 658 MESSAGE WITH THEIR PRECISION XTRA BLOOD GLUCOSE METER WHEN UPON TEST STRIP INSERTION AND THE DATE AND TIME HAD CHANGED SPONTANEOUSLY. THE CUSTOMER REPORTS USING THE P-LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE. INCORRECT TRENDING OF RESULTS CAN OCCUR WHEN RESULTS OBTAINED ON A METER SET WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE LETTER. | 0
1054593 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SVC THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MGMT SYS. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1001610 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< THE CUSTOMER REPORTED THAT THE PRECISION XTRA METER HAD SPONTANEOUSLY CHANGED THE DATE AND TIME. THE CUSTOMER ALSO REPORTED USING THE PRECISION LINK SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1057618 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED. | 0
1018490 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH ADC FA21DEC2006 LETTER. >< THE CUSTOMER REPORTED THAT HER FREESTYLE LITE BLOOD GLUCOSE MONITOR HAD SPONTANEOUSLY CHANGED THE DATE AND TIME. SHE ALSO REPORTS USING PRECISION LINK DATA SOFTWARE. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS DEATH. | 0
1057617 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED. | 0
1054581 | A CUSTOMER REPORTED RECEIVING AN ERROR 6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT, THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED. | 0
1054584 | A CUSTOMER REPORTED THAT THE DATE AND TIME ON IN HIS PRECISION XTRA BLOOD GLUCOSE METER WAS INCORRECT. THE CUSTOMER WAS IDENTIFIED BY ADC CUSTOMER SERVICE TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
972108 | . >< THIS CASE WAS REPORTED BY A CONSUMER AND DESCRIBED THE OCCURRENCE OF NEUROPATHY IN A FEMALE PT WHO RECEIVED POLIGRIP (SUPER POLIGRIP EXTRA CARE WITH POLISEAL) PASTE INTERMITTENTLY OVER A PERIOD OF 20 TO 25 YEARS FOR LOOSE DENTURES. A PHYSICIAN OR OTHER HEALTH CARE PROFESSIONAL HAS NOT VERIFIED THIS REPORT. THE PT'S PAST MEDICAL HISTORY INCLUDED APPENDECTOMY AND HYSTERECTOMY. CONCURRENT MEDICAL CONDITIONS INCLUDED HIGH CHOLESTEROL, HIGH TRIGLYCERIDES, INDIGESTION AND TYPE 2 DIABETES. CO-SUSPECT MEDICATION INCLUDED FIXODENT. CONCURRENT MEDICATIONS INCLUDED SODIUM RABEPRAZOLE (ACIPHEX), GEMFIBROZIL, SIMVASTATIN, GLIBENCLAMIDE (GLYBURIDE), METFORMIN HYDROCHLORIDE (METFORMIN), HYDROCODONE + PARACETAMOL (HYDROCODONE APAP) AND PREGABALIN (LYRICA). ON AN UNKNOWN DATE, THE PATIENT STARTED POLIGRIP (DENTAL). THE PATIENT REPORTED THAT FOR THE PAST 20 TO 25 YEARS, SHE SWITCHED BETWEEN USE OF POLIGRIP AND FIXODENT DENTURE ADHESIVE. APPROXIMATELY TWO TO THREE MONTHS PRIOR TO THE REPORT, THE PATIENT EXPERIENCED NEUROPATHY. THE PATIENT STATED THAT HER NEUROPATHY MAY BE RELATED TO HER HISTORY OF DIABETES, BUT WAS CONCERNED THAT ZINC IN BOTH POLIGRIP AND FIXODENT PRODUCTS COULD BE THE CAUSE. AT THE TIME OF REPORTING ZINC LEVELS HAD NOT BEEN TESTED. THIS CASE WAS ASSESSED AS MEDICALLY SERIOUS BY GSK. TREATMENT WITH POLIGRIP WAS DISCONTINUED. AT THE TIME OF REPORTING, THE EVENT WAS UNRESOLVED. MANUFACTURER'S COMMENT: THE MANUFACTURER'S REPORT NUMBER FOR THIS CASE IS 9681138-2007-00014. SUPER POLIGRIP EXTRA CARE WITH POLISEAL IS MANUFACTURED IN ANOTHER COUNTRY. THE PRODUCT IS NOT AVAILABLE. NO CORRECTIVE ACTIONS HAVE BEEN REQUIRED BASED ON THIS REPORT. | 0
868342 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< MICROBIAL CONTAMINATION WITHIN THE ARCHITECT SYSTEM (ARCHITECT I2000 / I2000SR PROCESSING MODULES AND ARCHITECT WASH BUFFER) THROUGH GENERATION OF FOLATE-LIKE BY-PRODUCT (MICROBIAL FOLATE) MAY CAUSE ARCHITECT FOLATE CALIBRATION FAILURES, AND SHIFTS IN ARCHITECT FOLATE RESULTS (QC AND PT RESULTS). THERE WAS NO IMPACT TO PT MGMT REPORTED. | 0
1046517 | THE ACCOUNT STATED THAT THE ARCHITECT I2000 ANALYZER IS GENERATING DISCREPANT AFP RESULTS. IN 2008, THE AFP =3 NG/DL, THREE MONTHS LATER, THE AFP RESULT = 141 NG/DL, ONE WEEK LATER, AFP RESULT =3 NG/DL AND THE NEXT DAY, THE AFP RESULT = 2 NG/DL. THE ACCOUNT REPORTED OUT THE RESULT OF 141 NG/DL FROM ONE WEEK PRIOR, AND THE DOCTOR QUESTIONED THE RESULT AND STATED IT DID NOT MATCH THE PATIENT'S CLINICAL PICTURE. THE SPECIMEN FROM THE SAME DAY, HAS BEEN SENT TO A REFERENCE LAB. THE ACCOUNT RETESTED THE SPECIMEN AFTER IT WAS RETURNED FROM THE REFERENCE LAB AND THE RESULTS WERE 2 AND 3 NG/DL. THERE WAS NO IMPACT TO THE PATIENT AS THE SAMPLE WAS RETESTED WHEN QUESTIONED BY THE DOCTOR. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
981126 | A REVIEW OF THE CUSTOMER'S MAINTENANCE PROCEDURES INDICATED THAT THEY WERE NOT REPLACING THE SAMPLE PROBE EVERY 3 MONTHS AS INSTRUCTED IN A PREVIOUS PRODUCT INFO LETTER. ABBOTT HAS DETERMINED THAT REGULARLY SCHEDULED REPLACEMENT OF THE PROBE AT THE SAMPLE PIPETTOR POSITION CAN HELP MAINTAIN OPTIMUM SYSTEM PERFORMANCE AND REQUIRES THAT THE PROBE BE REPLACED EVERY 3 MONTHS. THE ARCHITECT TOTAL B-HCG PACKAGE INSERT STATES THAT INDIVIDUAL SAMPLES MAY EXHIBIT EVALUATED CONCENTRATION DUE TO BUILD UP OF PROTEINS ON THE SAMPLE PIPETTOR PROBE. THE CUSTOMER WAS TRAINED ON REQUIRED MAINTENANCE PROCEDURES AND NO FURTHER ASSISTANCE WAS REQUESTED. A REVIEW OF THE COMPLAINT HISTORY FOR ARCHITECT ISYSTEM WAS PERFORMED FOR THE TIME PERIOD OF 2007 THROUGH EARLY 2008, AND NO TRENDS RELATED TO THIS ISSUE WERE IDENTIFIED. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATED THAT A FEMALE PT GENERATED A RESULT OF 472 MIU/ML ON THE ARCHITECT TOTAL B-HCG ASSAY. THE SAMPLE WAS TESTED WITH A RESULT OF <1.2 MIU/ML. ANOTHER SAMPLE WAS OBTAINED FROM THIS PT AND RECOVERED A RESULT OF <1.2 MIU/ML. NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
982048 | THE ACCOUNT STATED THEY SAW SMOKE COMING OUT OF THE TOP OF THE PRISM ANALYZER. THE ACCOUNT TURNED OFF THE PRISM ANALYZER. ADDITIONALLY, THE PRISM ANALYZER GAVE ERROR 40-909, XY TABLE CONTROLLER BAD RESPONSE (254) DURING SAMPLING IN THE MORNING PRIOR TO SEEING THE SMOKE. SERVICE WAS REQUESTED FOR THE PRISM ANALYZER. NO INJURY TO THE OPERATOR WAS REPORTED. >< ABBOTT SERVICE DISCOVERED SOME LOOSE WIRE STRANDS FROM Z2 AXIS PIPETTOR DRIVE. SOME WIRE STRANDS WERE GOING INSIDE THE XY AXIS CARD CAGE FROM THE BOTTOM. SERVICE REPLACED THE Z1 AND 2 AXIS DRIVE CABLES AND THE XY CONTROLLER CARD CAGE FOR RESOLUTION. THE ABBOTT PRISM OPERATIONS MANUAL SHOULD BE REFERENCED: HAZARDS, NOTES 'CAUTION' ASSOCIATED WITH ELECTRICAL, MECHANICAL AND PHYSICAL HAZARDS. OPERATING INSTRUCTIONS, INCLUDES INSTRUCTIONS FOR AN EMERGENCY POWER OFF THE SYSTEM. THIS IS A FINAL REPORT. | 0
1079046 | DEVICE EVAL IS IN PROCESS, NO CONCLUSION CAN BE DRAWN AT THIS TIME. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT GENERATED ERRATIC AXSYM VALPROIC ACID AND CARBAMAZEPINE RESULTS ON A SPECIMEN. THE SPECIMEN INITIALLY TESTED AXSYM VALPROIC ACID = 0 AND AXSYM CARBAMAZEPINE = 0. THE SPECIMEN WAS REPEATED WITH RESULTS OF AXSYM VALPROIC ACID = 12.9 AND AXSYM CARBAMAZEPINE = 5.3, 5.4, 8.2. NO UNITS OF MEASUREMENT WERE PROVIDED FOR AXSYM VALPROIC ACID OR AXSYM CARBAMAZEPINE. NO LIQUID SURFACE DETECTION ERRORS WERE FOUND ON THE AXSYM ANALYZER. NO PATIENT INFORMATION IS AVAILABLE. THE AXSYM VALPROIC ACID AND AXSYM CARBAMAZEPINE ERRATIC RESULTS WERE NOT REPORTED OUTSIDE OF THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
1048323 | INVESTIGATION SUMMARY: THE CUSTOMER REPORTED DISCREPANT RESULTS GENERATED FOR 3 PT SAMPLES WHEN TESTED USING THE ARCHITECT BHCG ASSAY. THE COMPLAINT TEXT INDICATES THAT THE ANALYZER PERFORMED WELL DURING THE DAY WITHOUT ERROR AND CONTROLS RUN DURING THE DAY WERE ACCEPTABLE. WHEN THE SAMPLES GAVE THE INITIALLY LOW RESULTS, THEY WERE RERUN AND RECOVERED SIMILARLY LOW VALUES, WHICH WERE REPORTED. THE RESULT FOR ONE OF THE SAMPLES WAS QUESTIONED SO THE SAMPLES WERE REPEATED GETTING DIFFERENT VALUES. THE TEXT INDICATES THAT ALTHOUGH THE RESULTS WERE REPORTED THERE WAS NO IMPACT TO PT MANAGEMENT AS THE ERROR WAS CAUGHT QUICKLY AND REPORTED TO THE PHYSICIANS QUICKLY. PER THE TEXT ENGINEERING CHANGED THE BUFFER CONTAINER, CHECKED THE PROBES AND DRAIN SENSORS AND THE INSTRUMENT WORKED. THE TECHNICAL MARKETING REP (TMR) ALSO VISITED THE SITE. THE TMR INSPECTED THE INSTRUMENT LEARNING THAT DURING THE NIGHT SHIFT THE TECHNOLOGIST BENT THE R2 PROBE BUT REPLACED IT WITH THE STAT PROBE. THE TMR INSTALLED A NEW REAGENT PROBE, RECALIBRATED BOTH PROBES AND RAN THE CONTROLS. THE TMR ALSO PERFORMED AN INTERNAL DECONTAMINATION. THE TMR ALSO RECOMMENDED THE TECHNOLOGISTS RUN CONTROLS EACH SHIFT SAMPLES ARE PROCESSED. A REVIEW OF THE COMPLAINT HISTORY FOR ARCHITECT ISYSTEM WAS PERFORMED FOR THE TIME PERIOD 10/01/2007 THROUGH 05/15/2008. THE INSTRUMENT WAS INSTALLED AT THE SITE IN JANUARY 2008 AND SINCE THAT TIME NO OTHER COMPLAINTS OF THIS NATURE WERE FOUND. A REVIEW OF COMPLAINTS LOGGED AGAINST EQUIPMENT SUBTYPE 116 AND LN 7K78 WAS PERFORMED FOR THE TIME PERIOD 10/01/2007 THROUGH 05/15/2008. THE RESULTS OF THE REVIEW IDENTIFIED OTHER COMPLAINTS, WHICH DESCRIBE THE ARCHITECT ISYSTEM GENERATED DISCREPANT OR DISCORDANT RESULTS FOR THE ARCHITECT BHCG ASSAY. THE REVIEW DID NOT FIND THAT THE COMPLAINTS IDENTIFIED A DEFICIENCY OF THE PRODUCT OR THAT THE PRODUCT WAS PERFORMING OUTSIDE ITS INTENDED USE. THE EVENT IS ADDRESSED IN LABELING OF THE ARCHITECT TOTAL B-HCG PACKAGE INSERT LN 7K78 INDICATES THAT ADEQUATE LABELING IS PROVIDED RELATED TO THE CUSTOMER'S OBSERVATION. REFER TO THE FOLLOWING SECTIONS: QC PROCEDURES INTERPRETATION OF RESULTS LIMITATIONS OF THE PROCEDURE. FOR DIAGNOSTIC PURPOSES, HCG RESULTS SHOULD ALWAYS BE USED IN CONJUNCTION WITH OTHER DATA; E.G., PT'S MEDICAL HISTORY, SYMPTOMS, RESULTS OF OTHER TESTS, CLINICAL IMPRESSIONS, ETC. ECTOPIC PREGNANCY CANNOT BE DISTINGUISHED FROM NORMAL PREGNANCY BY HCG MEASUREMENTS ALONE. AT 10, 11 THE RESULTS FROM THIS OR ANY OTHER DIAGNOSTIC KIT SHOULD BE USED AND INTERPRETED ONLY IN THE CONTEXT OF THE OVERALL CLINICAL PICTURE. IF THE HCG LEVEL IS INCONSISTENT WITH, OR UNSUPPORTED BY, CLINICAL EVIDENCE, RESULTS SHOULD BE CONFIRMED BY AN ALTERNATE HCG METHOD. THIS MAY INCLUDE THE QUALITATIVE TESTING OF URINE. AT 25 THE ABSENCE OF URINARY HCG MAY SUGGEST A FALSELY ELEVATED SERUM RESULT. RESULTS MAY ALSO BE CONFIRMED BY PERFORMING SERIAL DILUTIONS OF THE SAMPLE. USUALLY, BUT NOT ALWAYS, SAMPLES THAT CONTAIN INTERFERING SUBSTANCES EXHIBIT NONLINEAR RESULTS WHEN DILUTED. THE ABBOTT ARCHITECT SYSTEM OPERATIONS MANUAL (JUNE, 2007) INDICATES THE FOLLOWING UNDER SECTION 7, OPERATIONAL PRECAUTIONS AND LIMITATIONS, 'LIMITATIONS OF RESULT INTERPRETATION' THAT ASSAY RESULTS MUST BE USED WITH OTHER CLINICAL DATA, FOR EXAMPLE, SYMPTOMS, OTHER TEST RESULTS, PT HISTORY, CLINICAL IMPRESSIONS, INFO AVAILABLE FROM CLINICAL EVAL, AND OTHER DIAGNOSTIC PROCEDURES. ALL DATA MUST BE CONSIDERED FOR PT CARE MANAGEMENT. IF ASSAY RESULTS ARE INCONSISTENT WITH CLINICAL EVIDENCE, ADD'L TESTING IS SUGGESTED TO CONFIRM THE RESULT. THE ARCHITECT SYSTEM HAS BEEN VALIDATED FOR ITS INTENDED USE. HOWEVER, ERRORS CAN OCCUR DUE TO POTENTIAL OPERATOR ERRORS AND ARCHITECT SYSTEM TECHNOLOGY LIMITATIONS. THE OPERATIONS MANUAL ALSO INCLUDES A PROBABLE CAUSE FOR THE CUSTOMER'S ISSUE THAT THE PROBE TIP IS BENT OR DAMAGED UNDER SECTION 10 TROUBLESHOOTING AND DIAGNOSTICS, OBSERVED PROBLEMS 'ERRATIC ASSAY RESULTS (I SYSTEM)'. CONCLUSION: THE DEVICE INVOLVED IN THE EVENT WAS EVALUATED. A REVIEW OF MECHANISM AND PERFORMANCE TESTS WAS PERFORMED. LABELING WAS REVIEWED. INTERMITTENT FAILURE OF THE REAGENT PROBE WAS DISCOVERED AND REPLACED TO RESOLVE THE ISSUE. USE ERROR ALSO MAY HAVE CONTRIBUTED TO THE ISSUE, AS THE CUSTOMER APPEARS TO HAVE REPLACED A REAGENT PROBE WITH A STAT PROBE THAT MAY HAVE BEEN ON THE INSTRUMENT. NO IMPACT TO PT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. SEE SCANNED PAGE. >< THE CUSTOMER STATES THAT DISCREPANT RESULTS WERE GENERATED FOR PT SAMPLES WHEN USING THE ARCHITECT BHCG ASSAY. (MIU/ML) INITIAL RESULT: <1.20. REPEAT RESULT: 99760. THE INITIAL RESULTS WERE OBSERVED TO BE LOW AND WERE REPEATED, YIELDING THE SAME RESULTS AND WERE REPORTED FROM THE LAB. BUFFER AND A BENT PROBE WERE REPLACED ON THE ANALYZER AND PT SAMPLES REPEATED GENERATING POSITIVE RESULTS. PHYSICIANS WERE IMMEDIATELY NOTIFIED WITH NO ADVERSE IMPACT TO PT MANAGEMENT REPORTED DUE TO THIS ISSUE. | 0
1048324 | INVESTIGATION SUMMARY: THE CUSTOMER REPORTED DISCREPANT RESULTS GENERATED FOR 3 PT SAMPLES WHEN TESTED USING THE ARCHITECT BHCG ASSAY. THE COMPLAINT TEXT INDICATES THAT THE ANALYZER PERFORMED WELL DURING THE DAY WITHOUT ERROR AND CONTROLS RUN DURING THE DAY WERE ACCEPTABLE. WHEN THE SAMPLE GAVE THE INITIALLY LOW RESULTS, THEY WERE RERUN AND RECOVERED SIMILARLY LOW VALUES, WHICH WERE REPORTED. THE RESULT FOR ONE OF THE SAMPLES WAS QUESTIONED SO THE SAMPLES WERE REPEATED GETTING DIFFERENT VALUES. THE TEXT INDICATES THAT ALTHOUGH THE RESULTS WERE REPORTED THERE WAS NO IMPACT TO PT MANAGEMENT AS THE ERROR WAS CAUGHT QUICKLY AND REPORTED TO THE PHYSICIANS QUICKLY. PER THE TEXT ENGINEERING CHANGED THE BUFFER CONTAINER, CHECKED THE PROBES AND DRAIN SENSORS AND THE INSTRUMENT WORKED. THE TECHNICAL MARKETING REP (TMR) ALSO VISITED THE SITE. THE MFR INSPECTED THE INSTRUMENT LEARNING THAT DURING THE NIGHT SHIFT THE TECHNOLOGIST BENT THE R2 PROBE BUT REPLACED IT WITH THE STAT PROBE. THE TMR INSTALLED A NEW REAGENT PROBE, RECALIBRATED BOTH PROBES AND RAN THE CONTROLS. THE MFR ALSO PERFORMED AN INTERNAL DECONTAINATION. THE TMR ALSO RECOMMENDED THE TECHNOLOGISTS RUN CONTROLS EACH SHIFT SAMPLES ARE PROCESSED. A REVIEW OF THE COMPLAINT HISTORY FOR ARCHITECT ISYSTEM WAS PERFORMED FOR THE TIME PERIOD 2007 THROUGH 2008. THE INSTRUMENT WAS INSTALLED AT THE SITE IN 2008 AND SINCE THAT TIME NO OTHER COMPLAINTS OF THIS NATURE WERE FOUND. A REVIEW OF COMPLAINTS LOGGED AGAINST EQUIPMENT SUBTYPE 116 AND LN 7K78 WAS PERFORMED FOR THE TIME PERIOD 2007 THROUGH 2008. THE RESULTS OF THE REVIEW IDENTIFIED OTHER COMPLAINTS, WHICH DESCRIBE THE ARCHITECT ISYSTEM GENERATED DISCREPANT OR DISCORDANT RESULTS FOR THE ARCHITECT BHCG ASSAY. THE REVIEW DID NOT FIND THAT THE COMPLAINTS IDENTIFIED A DEFICIENCY OF THE PRODUCT OR THAT THE PRODUCT WAS PERFORMING OUTSIDE ITS INTENDED USE. THE EVENT IS ADDRESSED IN LABELING OF THE ARCHITECT TOTAL B-HCG PACKAGE INSERT LN 7K78 (77-4440/R2) INDICATES THAT ADEQUATE LABELING IS PROVIDED RELATED TO THE CUSTOMER'S OBSERVATION. REFER TO THE FOLLOWING SECTIONS: QC PROCEDURES INTERPRETATION OF RESULTS LIMITATIONS OF THE PROCEDURE. FOR DIAGNOSTIC PURPOSES, HCG RESULS SHOULD ALWAYS BE USED IN CONJUNCTION WITH OTHER DATA; E.G., PT'S MEDICAL HISTORY, SYMPTOMS, RESULTS OF OTHER TESTS, CLINICAL IMPRESSIONS, ETC. ECTOPIC PREGNANCY CANNOT BE DISTINGUISHED FROM NORMAL PREGNANCY BY HCG MEASUREMENTS ALONE. TEN, ELEVEN THE RESULTS FROM THIS OR ANY OTHER DIAGNOSTIC KIT SHOULD BE USED AND INTERPRETED ONLY IN THE CONTEXT OF THE OVERALL CLINICAL PICTURE. IF THE HCG LEVEL IS INCONSISTENT WITH, OR UNSUPPORTED BY, CLINICAL EVIDENCE, RESULTS SHOULD BE CONFIRMED BY AN ALTERNATE HCG METHOD. THIS MAY INCLUDE THE QUALITATIVE TESTING OF URINE. TWENTY FIVE THE ABSENCE OF URINARY HCG MAY SUGGEST A FALSELY ELEVATED SERUM RESULT. RESULTS MAY ALSO BE CONFIRMED BY PERFORMING SERIAL DILUTIONS OF THE SAMPLE. USUALLY, BUT NOT ALWAYS, SAMPLES THAT CONTAIN INTERFERING SUBSTANCES EXHIBIT NONLINEAR RESULTS WHEN DILUTED. THE ABBOTT ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-104) 2007 INDICATES THE FOLLOWING OPERATIONAL PERCAUTIONS AND LIMITATIONS, 'LIMITATIONS OF RESULT INTERPRETATION' THAT ASSAY RESULTS MUST BE USED WITH OTHER CLINICAL DATA, FOR EXAMPLE, SYMPTOMS, OTHER TEST RESULTS, PT HISTORY, CLINICAL IMPRESSIONS, INFO AVAILABLE FROM CLINICAL EVAL, AND OTHER DIAGNOSTIC PROCEDURES. ALL DATA MUST BE CONSIDERED FOR PT CARE MANAGEMENT. IF ASSAY RESULTS ARE INCONSISTENT WITH CLINICAL EVIDENCE, ADD'L TESTING IS SUGGESTED TO CONFIRM THE RESULT. THE ARCHITECT SYSTEM HAS BEEN VALIDATED FOR ITS INTENDED USE. HOWEVER, ERRORS CAN OCCUR DUE TO POTENTIAL OPERATOR ERRORS AND ARCHITECT SYSTEM TECHNOLOGY LIMITATIONS. THE OPERATIONS MANUAL ALSO INCLUDES A PROBABLE CAUSE FOR THE CUSTOMER'S ISSUE THAT THE PROBE TIP IS BENT OR DAMAGED UNDER SECTION 10 TROUBLESHOOTING AND DIAGNOSTICS, OBSERVED PROBLEMS 'ERRATIC ASSAY RESULTS (I SYSTEM)'. CONCLUSION: THE DEVICE INVOLVED IN THE EVENT WAS EVALUATED. A REVIEW OF MECHANICAL AND PERFORMANCE TESTS WAS PERFORMED. LABELING WAS REVIEWED. INTERMITTENT FAILURE OF THE REAGENT PROBE WAS DISCOVERED AND REPLACED TO RESOLVE THE ISSUE. USE ERROR ALSO MAY HAVE CONTRIBUTED TO THE ISSUE, AS THE CUSTOMER APPEARS TO HAVE REPLACED A REAGENT PROBE WITH A STAT PROBE THAT MAY HAVE BEEN ON THE INSTRUMENT. NO IMPACT TO PT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATES THAT DISCREPANT RESULTS WERE GENERATED FOR PT SAMPLES WHEN USING THE ARCHITECT BHCG ASSAY. (MIU/ML) INITIAL RESULT: <1.20. REPEAT RESULT: 331.3. THE INITIAL RESULTS WERE OBSERVED TO BE LOW AND WERE REPEATED, YIELDING THE SAME RESULTS AND WERE REPORTED FROM THE LAB. BUFFER AND A BENT PROBE WERE REPLACED ON THE ANALYZER AND PT SAMPLES REPEATED GENERATING POSITIVE RESULTS. PHYSICIANS WERE IMMEDIATELY NOTIFIED WITH NO ADVERSE IMPACT TO PT MANAGEMENT REPORTED DUE TO THIS ISSUE. | 0
960282 | PER THE COMPLAINT TEXT, THE CUSTOMER HAS HAD MULTIPLE INSTRUMENT ISSUES SINCE JUNE 2007 THAT REQUIRED FIELD SERVICE REPRESENTATIVE (FSR) INTERVENTION. THE RECENT ISSUES WERE FOR ERROR CODE 5015 MOTOR STEP LOSS, PROBE UP/DOWN, PROCESSING CENTER. THE FSR RECALIBRATED THE SAMPLE PROBE, REPLACED THE SMALL VOLUME HEATER DUE TO A LEAK, CHANGED BOTH PROBE LINK TUBINGS AND VERIFIED MUP AND WASH DISPENSE VOLUMES. THE FSR PERFORMED A MEIA STATION VERIFICATION BUT IT FAILED. THE PROCEDURE WAS REPEATED SUCCESSFULLY AFTER REPLACING THE MEIA LAMP AND PERFORMING MEIA VERIFICATION. THE FSR ALSO PERFORMED A REAGENT PACK CALIBRATION AND SAMPLE BARCODE CALIBRATION SUCCESSFULLY. THE FSR RAN CONTROLS AND THE QUESTIONABLE SAMPLE (3X) GETTING ACCEPTABLE RESULTS. A REVIEW OF THE COMPLAINT HISTORY FOR AXSYM SYSTEM OVER THE TIME PERIOD 2006 THROUGH 2007 DID NOT FIND OTHER COMPLAINTS FOR THE CUSTOMER'S CURRENT OBSERVATION. THE ABBOTT AXSYM SYSTEM OPERATIONS MANUAL VOLUME-1 (66-6749/R3-FEBRUARY 1996) DESCRIBES THE FOLLOWING UNDER THE SECTION SHOWN BELOW: SECTION 7 OPERATIONAL PRECAUTIONS AND LIMITATIONS. RESULT INTERPRETATION/REPORTING. AXSYM ASSAY RESULTS/ANALYTE DETERMINATIONS MUST BE USED IN CONJUNCTION WITH OTHER CLINICAL DATA, E.G., SYMPTOMS, RESULTS OF OTHER TESTS, PATIENT HISTORY, CLINICAL IMPRESSIONS, INFORMATION AVAILABLE FROM CLINICAL EVALUATION, AND OTHER DIAGNOSTIC PROCEDURES. ALL OF THIS DATA MUST BE CONSIDERED FOR DIAGNOSIS AND GOOD PT CARE MANAGEMENT. SECTION 10 TROUBLESHOOTING AND DIAGNOSTICS, OBSERVED PROBLEMS PROVIDES A PROBABLE CAUSE FOR THE CUSTOMER'S ISSUE A DISPENSE ISSUE WITH THE BULK SOLUTION 1 AND 3 WHICH IS DESCRIBED UNDER MEIA TEST RESULTS TOO LOW AND RESULTS ERRATIC, DISCREPANT, AND/OR EXPERIENCING IMPRECISION. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATED THAT THE AXSYM ANALYZER LIST # 7A83 GENERATED A DISCREPANT HCV RESULT FOR A PT SAMPLE. THE PATIENT SAMPLE INITIALLY RECOVERED A VALUE OF 0.7 S/CO (SAMPLE/CUTOFF), WHICH IS NEGATIVE BUT WAS QUESTIONED, AS IT DID NOT MATCH THE PT HISTORY. AN OLD SAMPLE FROM THE PT WAS ANALYZED WITHOUT CENTRIFUGING FIRST AND RECOVERED A VALUE OF 1.25 S/CO. THE 'PARENT' SAMPLE WAS CENTRIFUGED THEN RUN RECOVERING A VALUE OF 1.23 S/CO. THE FIRST SAMPLE WAS RERUN WITHOUT CENTRIFUGING RECOVERING A VALUE OF 1.25 S/CO, WHICH WAS REPORTED. THERE WAS NO IMPACT TO PT MANAGEMENT REPORTED. | 0
1036329 | THE CUSTOMER STATES THAT THE AXSYM ANALYZER IS GENERATING ERROR CODE 3550: SYS PRESSURE LIMIT EXCEEDED, SOLN 4. DURING TROUBLESHOOTING FOR THIS ISSUE, SMOKE BEGAN COMING FROM THE ANALYZER'S SYRINGE MOTOR. THE ABBOTT CUSTOMER TECHNICAL ADVOCATE ADVISED THE CUSTOMER TO SHUTDOWN/POWER OFF THE ANALYZER AND INITIATED A SERVICE CALL. THERE WERE NO REPORTS OF INJURY OR DAMAGE TO THE SURROUNDING LABORATORY ENVIRONMENT. THERE IS NO IMPACT TO PT MANAGEMENT REPORTED. >< AN ABBOTT FIELD SERVICE REP (FSR) REPLACED THE PRESSURE MONITOR PCB, THE SAMPLE SYRINGE AND THE PRESSURE MONITOR SENSOR. THE FSR THEN PERFORMED THE FOLLOWING VERIFICATION PROCEDURES: VP03 STARTUP, VP47 FLUSH, VP49 FLUIDICS, VP06 TEAH CLEANING, VP22 P-PROBE CAL, VP25 S-PROBE CAL, VP55 TOUCH SCREEN CAL AND A PRESSURE MONITOR CHECK. THE AXSYM ANALYZER MET ALL SPECIFICATIONS. NO INJURIES WERE REPORTED AND NO PT RESULTS WERE IMPACTED. NO CAUSE WAS ATTRIBUTED TO THE REPORT OF SMOKE. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE ABBOTT AXSYM SYSTEM SOFTWARE VERSION 5.20 ADDENDUM: MARCH 2005 SOFTWARE VERSION 5.20 ADDENDUM: PRESSURE MONITORING 3550 SYSTEM PRESSURE LIMIT EXCEEDED, SOLUTION 4. PROBABLE CAUSE: MALFUNCTION OF PRESSURE MONITORING SYSTEM. THIS IS A FINAL REPORT. | 0
1203923 | THE COMPLAINT TEXT FOR ARCHITECT ISYSTEM (LN 03M74-01) INDICATES THAT THE UPS (UNINTERRUPTIBLE POWER SUPPLY) THAT THE CUSTOMER IS USING (NON-ABBOT MANUFACTURED PRODUCT), GENERATED SMOKE AND CAUGHT FIRE. THE CUSTOMER EXTINGUISHED THE FIRE WITH A FOAM FIRE EXTINGUISHER. NO INJURIES WERE REPORTED. THE ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED AND REPORTED THAT THE UPS CABLES HAD BEEN TORN WHILE THE CUSTOMER WAS MOVING THE UPS. EXPOSED WIRES FROM THE TORN CABLES SUBSEQUENTLY TOUCHED ONE ANOTHER AND SMOKE AND FIRE WERE PRODUCED. THE FSR REPAIRED THE WORN CABLES AND VERIFIED THAT THE INSTRUMENT WAS OPERATIONAL. THE ABBOTT ARCHITECT SYSTEM OPERATIONS MANUAL ((PN 201837-105) 2008) ADDRESSES THE CUSTOMER'S ISSUE IN SECTION 7 OPERATIONAL PRECAUTIONS AND LIMITATIONS, SECTION 5 OPERATING INSTRUCTIONS - INSTRUCTIONS FOR THE EMERGENCY SHUTDOWN PROCEDURE AND SECTION 8 HAZARDS, SUBSECTION ELECTRICAL HAZARDS. A MALFUNCTION OF THE SYSTEM WAS NOT IDENTIFIED. THE COMPLAINT TEXT INDICATES THAT THE UPS CABLES HAD BEEN TORN WHILE THE CUSTOMER WAS MOVING THE UPS. BASED ON THE REVIEW OF THE INSTRUMENTS SERVICE HISTORY, NO REPEATS OF THIS ISSUE HAVE BEEN REPORTED SINCE THE FSR REPAIRED THE WORN CABLES. THE INVESTIGATION DEMONSTRATED THAT THE ARCHITECT I2000SR ANALYZER IS PERFORMING WITHIN ITS INTENDED USE, LABEL CLAIMS AND SPECIFICATIONS. NO DEFICIENCY RELATED TO THE PERFORMANCE OF THE DEVICE WAS IDENTIFIED. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT THE ARCHITECT I2000SR NETWORK HUB HAS BURNED (THE CUSTOMER REPORTS SEEING FLAMES) CAUSING SOME CABLE DAMAGE. THE NETWORK HUB IS AN EXTERNAL DEVICE THAT JOINS COMMUNICATION LINES AND ENABLES THE ELECTRONIC TRANSFER OF INFORMATION BETWEEN THE SYSTEM CONTROL CENTER (SCC) AND PROCESSING MODULE(S). NO INJURIES WERE REPORTED NOR DAMAGE TO THE SURROUNDING ENVIRONMENT. A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1015495 | THE CUSTOMER STATES THAT THE IMX ANALYZER GENERATED AN ALARM ALONG WITH ERROR CODE 74 FOR A BAD HIGH VOLTAGE POWER SUPPLY. WHEN THE CUSTOMER WENT TO TURN OFF THE ALARM VIA THE ANALYZER'S KEYPAD, SHE RECEIVED A 'PRETTY BAD SHOCK.' THE ABBOTT CUSTOMER TECHNICAL ADVOCATE (CTA) VERIFIED THAT THE CUSTOMER WAS NOT HARMED AND DID NOT REQUIRE ANY MEDICAL ATTENTION. THE CTA INSTRUCTED THE CUSTOMER TO POWER OFF THE ANALYZER AND A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PT MANAGEMENT REPORTED. >< AN ABBOTT FIELD SERVICE REP (FSR) VISITED THE CUSTOMER SITE AND REPLACED THE ANALOG BOARD AND PERFORMED VOLTAGE VERIFICATIONS SUCCESSFULLY. CONTROLS WERE WITHIN SPECS. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE IMX SERVICE AND SUPPORT MANUAL, MANUAL REVISION NUMBER 45987-104, FRONT MATTER DOCUMENT CONTROL, IN SECTION 2 - TROUBLESHOOTING, ERROR CODE 74, WHICH INSTRUCTS THE FIELD SERVICE REP TO CHECK THE VOLTAGES AT THE ANALOG BOARD TEST CONNECTOR J12-2 TO DETERMINE IF THE VOLTAGE FALLS BETWEEN 9V TO 10.1V. IF THE VOLTAGE IS NOT WITHIN THIS RANGE, THE FSR IS INSTRUCTED TO REPLACE THE ANALOG BOARD (IP-34 ERROR CODE 74ERROR CODE 74 BAD HIGH VOLTAGE POWER SUPPLY). THIS IS A FINAL REPORT. | 0
1048371 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT ONE PATIENT'S AXSYM VANCOMYCIN ASSAY RESULT OF 0.04 MG/L WAS REPORTED OUT OF THE LAB. THE PATIENT'S PHYSICIAN QUESTIONED THE RESULT AND ASKED THAT THE SAMPLE BE RETESTED. IN THE MEANTIME, A DOSE OF THE DRUG WAS ADMINISTERED BASED ON THE LOW RESULT. THE SAMPLE RETESTED AT 10.97 MG/L. THE SAMPLE IS DESCRIBED AS A BIT TURBID. CONTROLS HAVE BEEN WITHIN SPECIFICATIONS. THE PATIENT IS DOING FINE AND NO OTHER ISSUES HAVE OCCURRED. THERE IS NO FURTHER IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
981182 | DRAEGER MEDICAL SYSTEMS, INC HAS RECEIVED INFO FROM A CUSTOMER THAT ONE OF OUR MONITORS WAS FOUND WITH AN OVER HEATED INTERNAL BATTERY, BUT DID NOT EFFECT THE FUNCTIONALITY OF THE MONITOR. NO PT INJURY WAS REPORTED. >< WE HAVE REQUESTED THE RETURN OF THE CITED MONITOR AND BATTERY FOR EVALUATION. RESULTS WILL BE INCLUDED IN A FOLLOW UP REPORT. | 1
879295 | . >< A PHILLIPS TECHNICIAN MADE SOME CHANGES TO AN ECHO CART AT HOSPITAL IN 2007; THE PURPOSE OF THE CHANGES WERE TO ENABLE 'SCREEN SCRAPING', I.E. CAPTURE OF SCREEN DATA TO BE STORED IN PATIENT FILES ON THE AGFA HEARTLAB CARDIOVASCULAR SYSTEM. A HOSPITAL TECHNICIAN RAN A TEST CASE AFTER THE CHANGES WERE MADE AND THE MEASUREMENTS FROM THE TEST WERE ERRONEOUSLY MERGED INTO A LIVE PATIENT FOLDER. THIS DATA CORRUPTION, INCORRECT DATA ATTRIBUTED TO A PATIENT, COLD HAVE RESULTED IN THE AFFECTED PATIENT BEING SCHEDULED INTO SURGERY. SINCE THE HOSPITAL TECHNICIAN ATTRIBUTED THAT THE TEST DATA APPEARED IN THE PATIENTS STUDIES WHICH WERE RUN POST TEST, THE DATA WAS CORRECTED AND THE PATIENT WAS NOT SCHEDULED FOR SURGERY. THEREFORE, THERE WERE NO CONSEQUENCES TO THE INCORRECT MERGING. NO INJURY RESULTED. | 1
1024677 | NO DEFINITIVE CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED ON SITE AND PERFORMED THE APPROPRIATE CAMERA ADJUSTMENTS AND REPAIRS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER REPORTED FALSE NEGATIVE RESULTS. THE VISUAL INSPECTION SHOWED THE REACTIONS WERE POSITIVE. NO ERRONEOUS TEST RESULTS WERE REPORTED. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1027005 | DURING THE PROCESS OF TROUBLESHOOTING A FAILURE OF THE CARE UNIT'S TELEMETRY SYSTEM, A PATIENT WHO WAS NOT BEING MONITORED BY ECG TELEMETRY CODED AND REQUIRED MEDICAL INTERVENTION. THE PATIENT WAS TRANSFERRED TO ICU AND DIED TWO DAYS LATER. THE PATIENT WAS THE LAST OF TWO THAT COULD NOT BE PROVIDED WITH A SPARE TRANSMITTER DURING THE TIME THE TELEMETRY SYSTEM WAS BEING REPAIRED. THE CARE UNIT WAS ADULT MEDICAL DIABETES. THE TELEMETRY SYSTEM WAS BROUGHT BACK ON LINE EARLY THE FOLLOWING MORNING. | 1
1000791 | THE CUSTOMER REMOVED THE VENT HEAD ASSEMBLY AND SAW NO EVIDENCE OF DAMAGE TO THE VENT HEAD ASSEMBLY OR THE ASPIRATION PROBE. THE CUSTOMER CYCLED POWER BUT THIS DID STOP FLUID FROM SHOOTING OUT OF THE VENT HEAD ASSEMBLY. THE CUSTOMER ALSO AUTOCLEANED 3 TIMES BUT THE INSTRUMENT STILL SPORADICALLY PRODUCED RESULTS OF ZERO AND SHORT SAMPLE ERRORS. A FIELD SERVICE REP (FSR) WAS DISPATCHED AND REPLACED THE MANIFOLD VALVE ASSEMBLY, CONTAINING 8 SOLENOID VALVES, BECAUSE THE RETIC REAGENT WAS NOT FILLING. THE FSR ALSO REPLACED, AS A PRECAUTION, A CIRCUIT BOARD, WHICH CONTROLS THE SOLENOID VALVES, AND A PRESSURE SENSOR IN THE ASPIRATION TUBING. THE VENT HEAD WAS VERIFIED TO BE ALIGNED PROPERLY. LABELING: THE CELL-DYN SAPPHIRE SYSTEM OPERATOR'S MANUAL (08H10-01, REV D) IN SECTION 8, OUTLINES HAZARDS AND RESPONSES TO HAZARDS. PAGE 8-2 OUTLINES THE OPERATOR RESPONSIBILITY. PAGES 8-5 THROUGH 8-10 DISCUSSES BIOLOGICAL, CHEMICAL AND PHYSICAL HAZARDS. THE OPERATOR SHOULD WEAR GLOVES, LAB COATS AND PROTECTIVE EYEWEAR WHEN HANDLING HUMAN-SOURCED MATERIALS, REAGENTS OR WORKING AROUND THE INSTRUMENT WHEN IT IS OPERATING. THE TEXT DOES NOT DESCRIBE HOW THE OPERATOR LOOKED INSIDE THE INSTRUMENT TO OBSERVE THE ACTION OF THE VENT HEAD ASSEMBLY. WHEN THE VENT HEAD ASSEMBLY IS ASPIRATING, IT IS NOT VISIBLE BECAUSE OF THE PROCESSOR COVER. THE PROCESSOR COVER IS LOCKED AND THE LOCKED STATUS INDICATOR IS ON WHEN THE INSTRUMENT IS IN ACTIVE USE [OPEN OR CLOSED MODE]. THE PROCESSOR COVER IS NOT INTENDED TO BE REMOVED BY THE OPERATOR DURING ROUTINE OPERATION. WHEN THE COVER IS REMOVED DURING MAINTENANCE PROCEDURES, THE SYSTEM IS UNABLE TO PROCESS SAMPLES. (P 1-14) THE SENSOR OF THE PROCESSOR COVER HAS TO BE DELIBERATELY INACTIVATED FOR THE INSTRUMENT TO OPERATE WITHOUT THE PROCESSOR COVER. THE VENT HEAD ASSEMBLY CAN BE VIEWED, WHEN THE VENT HEAD ASSEMBLY HAS MOVED BACKWARDS AND IS DISPENSING BY OPENING THE RIGHT HAND DOOR. THE INSTRUMENT WILL CONTINUE TO OPERATE WHILE THE DOOR IS OPEN BUT THE VIEWER CANNOT GET CLOSE TO THE VENT HEAD BECAUSE OF THE PROCESSOR COVER. PAGES 8-10, 8-11 CAUTIONS AGAINST OVERRIDING SAFETY DEVICES OR ALLOWING THE BODY TO ENTER THE REGION OF MOVEMENT WHILE THE INSTRUMENT IS OPERATING. A REVIEW OF COMPLAINTS FROM APRIL 2007 THROUGH JANUARY 2008 DID NOT IDENTIFY ANY ADVERSE TRENDS FOR THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATED THAT THE VENT HEAD ASSEMBLY OF A CELL-DYN SAPPHIRE WAS SPORADICALLY SHOOTING FLUID INTO THE AIR WHILE RUNNING PT SAMPLES. THE CUSTOMER LOOKED INSIDE THE INSTRUMENT TO OBSERVE THE ACTION OF THE VENT HEAD ASSEMBLY. FLUID DID NOT COME INTO CONTACT WITH THE CUSTOMER AND NO INJURY WAS REPORTED. THE CUSTOMER ALSO OBSERVED THE ASPIRATION PROBE WAS DISPENSING INTO THE WASHCONE INSTEAD OF THE DILUTION CUP CAUSING RESULTS TO BE ZERO. | 0
1056768 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT STATED THAT THE CELL-DYN 3200 ANALYZER IS GENERATING DISCREPANT RBC RESULTS ON ONE PATIENT SAMPLE. THE ACCOUNT ALSO STATED THAT THE DIFFERENTIAL IS NOT MATCHING THE RESULTS. THE INITIAL RBC RESULT WAS 4.91 M/UL AND THE REPEAT RESULT WAS 7.34 M/UL. FIELD SERVICE WAS REQUESTED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1056766 | THE CUSTOMER REPORTED THAT LOW RESULTS WERE RECOVERED FROM A PATIENT SAMPLE RUN IN CLOSED MODE. WHEN REPEATED IN OPEN MODE THE RESULTS WERE ACCEPTABLE. THE INSTRUMENT IN USE WAS THE CELL-DYN SAPPHIRE. ALL QC RESULTS WERE IN RANGE. THE SAMPLE TUBE CONTAINED ABOUT 20 ML WHOLE BLOOD. THERE WAS NO SHORT SAMPLE MESSAGE GENERATED. HEMOGLOBIN (HGB) WAS UNDERLINED PER INTERNAL LIMITS AND THEREFORE THE SAMPLE WAS REPEATED IN OPEN MODE WITH ACCEPTABLE RESULTS. THE REPEAT RESULTS (OPEN MODE) WERE REPORTED OUT TO THE PHYSICIAN, SO THERE WAS NO IMPACT TO PATIENT MANAGEMENT. THE CUSTOMER PROVIDED FOLLOWING DATA: SEE SCANNED TABLE. >< THE CUSTOMER STATES THAT THE CELL-DYN SAPPHIRE ANALYZER GENERATED LOW RESULTS FOR ONE PATIENT SAMPLE IN CLOSED MODE OF OPERATION. IT APPEARED THAT THE PROBE WAS NOT DESCENDING FAR ENOUGH INTO THE TUBE TO ASPIRATE A LOW SAMPLE VOLUME ADEQUATELY. THE CUSTOMER TECHNICAL ADVOCATE (CTA) ASKED THE CUSTOMER TO CHANGE THE VENT ASSEMBLY, DO A ZERO PARK ALIGNMENT AND REPEAT THE PATIENT SAMPLE. THE VENT HEAD ASSEMBLY WAS CHANGED AND THE ZERO PARK ALIGNMENT WAS PERFORMED. THE INSTRUMENT WAS FOUND TO BE PERFORMING PER LABELING IN BOTH AN OPEN AND A CLOSED MODE. THE ASSIGNABLE CAUSE WAS GIVEN AS THE VENT HEAD ASSEMBLY CAUSING PARTIAL ASPIRATION. REVIEW OF THE COMPLAINTS FOR THE MONTHS OF AUGUST 2007 THROUGH APRIL 2008, DID NOT FIND ANY ADVERSE TREND FOR THE COMPLAINT ISSUE. THERE IS NO SYSTEMIC ISSUE. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE CELL-DYN SAPPHIRE SYSTEM OPERATOR'S MANUAL (08H10-01, REV. D), SECTION 3: PRINCIPLES OF OPERATION; SYSTEM INITIATED MESSAGES AND DATA FLAGS; DISPERSIONAL DATA ALERTS; 3-55 AND SECTION 10: TROUBLESHOOTING AND DIAGNOSTIC; TROUBLESHOOTING DATA RELATED PROBLEMS; METHODOLOGY FOR DATA-RELATED TROUBLESHOOTING; 1-121, 1-122. THIS IS A FINAL REPORT. | 0
1034502 | THE ACCOUNT STATED THAT THE WBC ON THE LOW CONTROL WAS OUT OF RANGE HIGH ON THE CELL-DYN 1700 ANALYZER. IN ADDITION, A PT GENERATED A DISCREPANT WBC RESULT. THE INITIAL WBC VALUE WAS 17.3 K/UL (NORMAL RANGE 4.1 - 10.9 K/UL) BUT UPON REPEAT, THE VALUE WAS 6.0 K/UL. THE HIGH WBC RESULT WAS NOT REPORTED AND THERE WAS NO IMPACT TO PT MANAGEMENT. >< THE CUSTOMER TECHNICAL ADVOCATE VERIFIED THAT ALL REAGENT LINES WERE FREE OF CRIMPS; SYRINGES FREE OF BUBBLES; LYSE LOT IN DATE, RECENTLY OPENED; LYSE CAP INTACT; MAINTENANCE UP TO DATE; CONTROLS STORED AND HANDLED APPROPRIATELY; GOOD BUBBLE MIX IN THE PRE-MIX CUP AND TRANSDUCERS; APERTURE PLATE AND SUPPLEMENTAL APERTURE PLATE CLEANING JUST COMPLETED. THE CUSTOMER RAN CONTROLS FROM A NEW BOX, AGAIN RECOVERING HIGH WBC RESULTS (LOW, NORMAL AND HIGH CONTROLS) A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPLACED AND FOUND THAT THE BUBBLE MIX TUBING WAS CLOGGED. THE FSR CLEANED THE TUBING AND VERIFIED THAT CONTROLS WERE WITHIN SPECIFICATION. NO PARTS WERE REPLACED. LABELING CELL-DYN 1700 SYSTEM OPERATOR'S MANUAL (03H58-01, REV D) STATES; A RESULT THAT FALLS OUTSIDE A LABORATORY ACTION LIMIT CAN INDICATE THE NEED FOR THE OPERATOR TO FOLLOW A LABORATORY PROTOCOL, SUCH AS REPEATING THE SAMPLE, PERFORMING A SMEAR REVIEW, OR NOTIFYING A PHYSICIAN. (3-23). SECTION 10 STATES IF QC SPECIMENS ARE OUTSIDE ACCEPTABLE LIMITS THERE MAY BE INSUFFICIENT OR NO DILUTION MIXING (OBSERVE BUBBLE MIX). (PAGE 10-24) A REVIEW OF COMPLAINT ANALYSIS TRENDING SYSTEM (CATS) AND TRENDING ANALYSIS SUPPORT SYSTEM (TASS) REPORTS, FOR THE PERIOD AUGUST 2007 THROUGH MARCH 2008, DID NOT INDICATE AN ADVERSE TREND FOR THE CELL-DYN 1700, LIST 03H53-01 (WHICH INCLUDES 03H57-01) FOR WBC ISSUES. THERE IS NO SYSTEMIC ISSUE. THIS IS THE FINAL REPORT. | 0
1022520 | THE ACCOUNT NOTICED WBC IMPRECISION WITH THE CELL-DYN 4000 ANALYZER. ONE PT GENERATED A WBC = 3000/UL BUT WHEN THE SPECIMEN WAS TESTED ON ANOTHER ANALYZER THE WBC = 9000 (NO UNITS OF MEASUREMENT WERE GIVEN). THE WBC = 3000/UL WAS NOT REPORTED OUTSIDE OF THE LAB. THROUGH TROUBLESHOOTING, IT WAS DISCOVERED THAT WBC IMPRECISION HAD BEEN OCCURRING FOR THE PREVIOUS 2 HOURS AND ABOUT 50 PT SPECIMENS WERE REPORTED WITH INACCURATE RESULTS. THE 50 SPECIMENTS WERE RERUN AND THE REPORTS WERE CORRECTED. NO SPECIFIC PT INFO OR DATA IS AVAILABLE. NO IMPACT TO PT MGMT WAS REPORTED. >< (INVESTIGATION PENDING). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROGRESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
1000758 | THE CUSTOMER CONTACTED ABBOTT REGARDING THE HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERROR MESSAGES GENERATED FROM THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER HAD REMOVED THE SYRINGE AND CLEANED IT WITH NO RESOLUTION. THE CUSTOMER REPORTED THE FAULTY SYRINGE WAS DISCARDED, INSTALLED A NEW SYRINGE AND THE 'FAIL TO HOME' ISSUE WAS RESOLVED. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO PT MANAGEMENT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1036331 | INVESTIGATION SUMMARY: THE CUSTOMER STATES THAT THEY ARE HAVING AN ISSUE WITH HEMOGLOBIN (HGB) RECOVERY USING A CELL-DYN 3000CS ANALYZER. APPROXIMATELY TEN PT SAMPLES WERE AFFECTED, HOWEVER, THE CUSTOMER PROVIDED ONLY ONE PT DATA EXAMPLE. THE CUSTOMER HAD CLEANED THE HGB FLOWCELL, EXERCISED VALVE 14 AND CHANGED THE PINCH TUBING FOR VALVES 14 AND 15 WITHOUT RESOLVING THE HGB ISSUE. QC (R&D CONTROL) WAS IN RANGE. A CUSTOMER TECHNICAL ADVOCATE (CTA) WALKED CUSTOMER THROUGH EXERCISES VALVES 13, 15, 41, 42, 44. CUSTOMER ALSO REPLACED PINCH TUBING IN VALVES 13, 41, 42, 44--ALL VALVES APPEARED TO OPERATE PROPERLY. THE HGB LYSE AND DILUENT SYRINGE WERE PROPERLY SEATED, NO BUBBLES WERE SEEN AND THE SYRINGE MOVEMENT WAS SMOOTH. HOWEVER, THE CUSTOMER OBSERVED THAT THERE WAS TOO MUCH REAGENT IN THE HGB MIXING CHAMBER AFTER CYCLING, WHICH DILUTED THE NEXT SAMPLE, CAUSING LOW HGB READINGS. FIELD SERVICE WAS DISPATCHED. FSR VISITED ON 4/17/2008 AND REPLACED, AS HARD FAILURE, FIVE WORN SOLENOID VALVES (V34, V35, V36, V37, V38). VERIFICATION PROCEDURE 3 (VP3) FOR SOLENOID OPERATION PASSED. VP8 (HGB FLOWCELL CLEANING AND ADJUSTMENT) WAS PERFORMED AND PASSED. THE CELL-DYN SYSTEM 3000 OPERATOR'S MANUAL, 92430-01, REV H: THE TROUBLESHOOTING GUIDE (SECTION 10 PAGES 10-25 TO 10-26) STATES THAT IN THE PROBLEM IDENTIFICATION STEP THE CUSTOMER SHOULD BE OBSERVING WHAT IS RIGHT AND WHAT IS WRONG TO FURTHER ISOLATE THE CATEGORY OF PROBLEM (HARDWARE, SOFTWARE OR MEASUREMENT). SECTION 3, SAMPLE ANALYSIS CYCLE OVERVIEW OUTLINES THE PROCESS OF ANALYSIS (PAGES 3-2 TO 3-4). PRECISE VOLUMES OF BLOOD ARE DILUTED WITH SPECIFIC VOLUMES OF REAGENT ACCORDING TO THE TYPE OF ANALYSIS BEING DONE. THE HGB DETERMINATION IS MADE FROM 1.2 UL OF BLOOD ADDED TO 1.5ML OF DILUENT AND 1.5ML HEMOGLOBIN LYSE REAGENT, WHICH IS MIXED BY BUBBLING ACTION IN THE HGB MIXING CHAMBER. THE MIXING CHAMBER AND FLOW CELLS ARE RINSED AFTER EACH CYCLE. THE VOLUME OF LIQUID IN THE MIXING CHAMBER SHOULD BE CONSISTENT CYCLE TO CYCLE. THE CUSTOMER OBSERVED WHILE CHECKING THE MOVEMENT OF SYRINGES AND VALVES, THAT THERE WAS TOO MUCH REAGENT IN THE MIXING CHAMBER, WHICH INDICATED A MEASUREMENT PROBLEM FOR HGB ANALYSIS AT MIXING CHAMBER AND A POSSIBLE HARDWARE ISSUE RELATED TO THE FILL/DRAIN OF THE MIXING CHAMBER. TRENDING ANALYSIS: A REVIEW OF THE COMPLAINT REPORTS, FOR THE PERIOD AUGUST 2007 THROUGH MARCH 2008, DID NOT INDICATE ANY ADVERSE TREND WITH THE CELL-DYN 3000, LIST BASE 91250-01, (WHICH INCLUDED 91323-03) FOR ISSUES WITH ERRATIC/LOW HGB RESULTS. THERE WAS NO SYSTEMIC ISSUE. LABELING: THE EVENT IS ADDRESSED IN THE CELL-DYN 3000 SYSTEM OPERATOR'S MANUAL, 92420-01, REV H SECTION 3: PRINCIPLES OF OPERATION: SAMPLE ANALYSIS CYCLE OVERVIEW; 3-2 TO 3-4 SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS: TROUBLESHOOTING GUIDE; 10-25 TO 10-26. CONCLUSION: THE DEVICE INVOLVED IN THE EVENT WAS EVALUATED. SERVICE REPLACED SOLENOID VALVES THAT WERE INTERMITTENTLY FAILING TO RESOLVE THE ISSUE. NO IMPACT TO PT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATES THAT HEMOGLOBIN RESULTS ON PT SAMPLES HAVE BEEN ERRATIC AND LOW WHEN GENERATED USING A CELL-DYN 3000 CS ANALYZER. AN EXAMPLE FOR ONE PT WAS PROVIDED; HGB=5 G/DL COMPARED TO A PREVIOUSLY OBTAINED RESULT OF 15 G/DL. INCORRECT RESULTS WERE NOT REPORTED OUT OF THE LAB. THERE WAS NO IMPACT TO PT MANAGEMENT REPORTED DUE TO THIS ISSUE. | 0
970099 | CUSTOMER REPORTED OUT FLAGGED WBC RESULT WITHOUT CONFIRMING WITH SLIDE OR ALTERNATE METHOD. THE CUSTOMER STATES THAT THEY REPORTED OUT A HIGH WBC COUNT GENERATED USING A CD 1800 ANALYZER. THE PATIENT WAS A BABY. TWO SEPARATELY DRAWN SPECIMENS WERE TESTED USING THE CD 1800 ANALYZER GENERATING WBC VALUES OF 33.4 AND 32.4 K/UL. THE CD 1800 GENERATED R0 FLAGS FOR BOTH RESULTS. THE CTA INFORMED THE CUSTOMER THAT WBC RESULTS WITH R0 FLAGGING SHOULD BE CONFIRMED WITH A SLIDE REVIEW OR AN ALTERNATE METHOD PRIOR TO REPORTING OUT RESULTS. TESTING WAS PERFORMED AT ANOTHER HOSPITAL OBTAINING A WBC RESULT OF 16.5 K/UL. THE CTA QUERIED THE CUSTOMER REGARDING IF THE HOSPITAL CORRECTED THE WBC COUNT FOR NRBCS WITH NO ADDITIONAL INFORMATION PROVIDED AT THIS TIME. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< INVESTIGATION SUMMARY: CELL-DYN 1800 GENERATED HIGH WBC RESULT WITH FLAGGING. ALL QC WAS IN RANGE. RESULTS REPORTED OUT WITHOUT VERIFICATION OF RESULTS PER INSTRUCTIONS FOR R0 FLAG. THE WBC RESULT WAS REPORTED OUT AND ANOTHER FACILITY (HOSPITAL) REPEATED THAT PATIENT WITH A LOWER WBC RESULT. THE CUSTOMER TECHNICAL ADVOCATE (CTA) VERIFIED ONLY ONE PATIENT SAMPLE WAS AFFECTED. THE BLOOD WAS NOT CLOTTED AND SAMPLE WAS DRAWN IN A SMALL VACUTAINER. THE WBC RECOVERED WAS: RUN 1 WBC=33.4; RUN 2 WBC=32.4 K/UL. THE CUSTOMER STATED BOTH RESULTS HAD A R0 FLAG, BUT A SMEAR WAS NOT REVIEWED BEFORE THE HIGH WBC WAS REPORTED. THE DOCTOR SENT THE PATIENT TO ANOTHER HOSPITAL WHERE THE WBC WAS FOUND TO BE 16.5 K/UL. IT WAS UNK WHAT INSTRUMENT WAS USED AT THE OTHER FACILITY. THE CUSTOMER RAN ALL QC AFTER THIS PATIENT AND ALL WITHIN SPECIFICATION. ALL MAINTENANCE WAS UP TO DATE. THE CTA RECOMMENDED THAT WHEN ANY FLAG IS DISPLAYED, THE RESULTS SHOULD BE CONFIRMED BY REVIEWING A SLIDE OR USING AN ALTERNATIVE METHOD. CTA REQUESTED THEY FIND OUT IF THE HOSPITAL LAB CORRECTED THE WBC COUNT FOR NRBC, BUT NO INFORMATION WAS PROVIDED. THE CTA FOLLOWED UP IN LATE 2007. THE CUSTOMER STATED THAT THEY SENT FIVE PATIENT SAMPLES TO THE HOSPITAL AND ALL MATCHED CLOSELY FOR THE WBC RESULTS ON THE SAME SAMPLE. QC WAS WITHIN RANGE. THE ISSUE WAS RESOLVED BY EDUCATION OF THE CUSTOMER. THE INSTRUMENT AND ASSAY WERE PERFORMING PER LABELING. THE CELL-DYN SYSTEM 1800 OPERATOR'S MANUAL, 07H80-01, REV D, PROVIDES INFORMATION ON R0 FLAGGING ON PAGES 3-26, 3-31, 10-42. R0 FLAGS ARE SUSPECT PARAMETER FLAGS, WHICH INDICATE THE INSTRUMENT IS HAVING DIFFICULTY MEASURING A PARAMETER. IT IS RECOMMENDED THAT THE OPERATOR FOLLOW THE LABORATORY REVIEW CRITERIA OR REVIEW A STAINED SMEAR TO CONFIRM THE DIFFERENTIAL RESULTS AND VERIFY THE WBC. TRENDING: A REVIEW OF COMPLAINT REPORTS, FOR THE PERIOD MARCH 2007 THROUGH OCTOBER 2007, DID NOT INDICATE ANY ADVERSE TREND WITH THE CELL-DYN 1800, LIST BASE 07H77-01, FOR ISSUES WITH DISCREPANT WBC RESULTS. NOVEMBER 2007 REPORTS WERE STILL PROCESS. LABELING: THE EVENT ADDRESSED IN THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL, 07H80-01, REV M SECTION 3: PRINCIPLES OF OPERATION: SYSTEM INITIATED MESSAGES AND DATA FLAGS: SUSPECT PARAMETER FLAGS; P. 3-26, CELL-DYN 1800 REGION ALERTS: FIGURE 3.1: REGIONAL ALERTS AND POSSIBLE CAUSES; P. 3-31 SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS; INDEX OF ERROR MESSAGES AND CONDITIONS; SUSPECT PARAMETER FLAG, LYM R0 OR RM; P, 10-42. CONCLUSION: THE CELL-DYN 1800 ANALYZER WAS EVALUATED FOR PERFORMANCE AND LABELING. A DEVICE MALFUNCTION WAS NOT IDENTIFIED PER THE INVESTIGATION. THE DEVICE WAS FOUND TO BE PERFORMING WITHIN SPECIFICATIONS. USER ERROR CONTRIBUTED TO THE EVENT AS LABELING INSTRUCTIONS WITH REGARD TO FLAGGED RESULTS WERE NOT FOLLOWED. CUSTOMER EDUCATION RESOLVED THE ISSUE. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. THIS IS THE FINAL REPORT. END OF REPORT. | 0
997547 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROGRESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER CONTACTED ABBOTT REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERROR MESSAGES GENERATED FROM THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER STATED IT WAS NECESSARY TO REPLACE THE HEMOGLOBIN SYRINGE TWICE. THE FIRST SYRINGE GENERATED THE ERROR IMMEDIATELY AND THE CUSTOMER REPLACED IT. THE SECOND SYRINGE WORKED ONLY A FEW DAYS, HOWEVER, THE SYRINGE LEAKED AND IT WAS REPLACED AGAIN. THE REPLACEMENT SYRINGE GENERATED THE ERROR MESSAGE AND THE SYRINGE WAS REPLACED AGAIN. NO FURTHER 'FAIL TO HOME' ERROR MESSAGES WERE GENERATED AFTER THE LAST REPLACEMENT. ADDITIONALLY, THE CUSTOMER STATED ALL QUALITY CONTROLS WERE WITHIN RANGE AND ISSUE WAS RESOLVED. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO PT MGMT. | 0
997549 | THE CUSTOMER CONTACTED ABBOTT REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERROR MESSAGES GENERATED FROM THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER REPLACED THREE SYRINGES IN FOUR DAYS, ORIGINALLY DUE TO A LOOSE SYRINGE PLUNGER SEAL, HOWEVER, SUBSEQUENT SYRINGE REPLACEMENTS GENERATED THE ERRORS UNTIL THE LAST REPLACEMENT. THE CUSTOMER STATED THERE WERE NO FURTHER ERRORS AND THE CUSTOMER WAS INSTRUCTED TO MONITOR THE SYRINGE FOR SEVERAL DAYS. UPON FOLLOW UP, THE CUSTOMER STATED THE ANALYZER WAS PERFORMING WITHIN SPECS, ALL QUALITY CONTROLS WERE WITHIN RANGE AND THE ISSUE WAS RESOLVED. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO PATIENT MANAGEMENT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
997548 | THE CUSTOMER CONTACTED ABBOTT REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERRORS GENERATED ON THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER REPLACED THE FAULTY SYRINGE WITH A NEW SYRINGE, AND THE ISSUE WAS RESOLVED. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO PT MANAGEMENT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROGRESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
989520 | THE CUSTOMER CONTACTED ABBOTT REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERRORS GENERATED FROM THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE ERROR MESSAGES OCCURRED SHORTLY AFTER INSTALLATION AND THE CUSTOMER CLEANED AND REPLACED THE SYRINGE TWICE WITH NO RESOLUTION. THE ABBOTT CUSTOMER TECH ADVOCATE DISPATCHED ABBOTT FIELD SVC REP (FSR) TO THE CUSTOMER FACILITY. THE FSR PERFORMED MAINTENANCE AND VERIFIED THE PERFORMANCE ANALYZER CONTROLS AND RESOLVED THE ISSUE. SHORTLY AFTER, THE CUSTOMER REPORTED THE ISSUE WAS UNRESOLVED AND DISCOVERED ANOTHER ISSUE WITH THE SYRINGE. THE FSR RETURNED TO THE CUSTOMER FACILITY, REPLACED THE SYRINGE, VERIFIED CONTROLS WERE WITHIN RANGE AND RESOLVED THE HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERROR MESSAGES. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO PT MGMT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
989519 | THE ABBOTT FIELD SERVICE ENGINEER (FSE) CONTACTED ABBOTT CUSTOMER SERVICE CENTER ON BEHALF OF THE CUSTOMER REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERRORS GENERATED FROM THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER REPORTED INTERMITTENT ERROR MESSAGES. THE FSE PERFORMED MAINTENANCE, REPLACED PARTS, VERIFIED THE PERFORMANCE OF THE ANALYZER CONTROLS AND THE ISSUE WAS RESOLVED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
997554 | THE CUSTOMER CONTACTED ABBOTT REGARDING FREQUENT HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERROR MESSAGES GENERATED ON THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER STATED THE SYRINGE HAS BEEN REPLACED WITHOUT RESOLUTION AND REQUESTED FIELD SERVICE. THE ABBOTT FIELD SERVICE REP (FSR) REPLACED PARTS, REMOVED AND CLEANED THE SYRINGE AND DETERMINED THE ANALYZER IS PERFORMING WITHIN SPECIFICATIONS. ADDITIONALLY, THE FSR STATED THE QUALITY CONTROLS WERE WITHIN RANGE. THE DAY FOLLOWING THE VISIT BY FIELD SERVICE, THE CUSTOMER ABBOTT TO REPORT THE ISSUE WAS NOT RESOLVED. UPON RETURN TO THE CUSTOMER FACILITY, THE FSR PROACTIVELY REPLACED PARTS, VERIFIED THE QUALITY CONTROLS PASSED, THE ANALYZER WAS PERFORMING WITHIN SPECIFICATIONS. THE FSR DETERMINED THE 'FAIL TO HOME' SYRINGE ERROR MESSAGE ISSUE WAS RESOLVED. THE CUSTOMER REPORTED THERE WAS NO IMPACT TO THE PT MANAGEMENT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
988652 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER CONTACTED ABBOTT REGARDING HEMOGLOBIN SYRINGE 'FAIL TO HOME' ERRORS GENERATED ON THE CELL-DYN SAPPHIRE HEMATOLOGY ANALYZER. THE CUSTOMER STATED THE ISSUE HAS BEEN ON-GOING AS THE ERROR MESSAGE OCCURS INTERMITTENTLY. THE ABBOTT CUSTOMER TECHNICAL ADVOCATE (CTA) INSTRUCTED THE CUSTOMER TO REPLACE THE SYRINGE. THE CUSTOMER REPORTED THEY WERE UNABLE TO PERFORM TROUBLESHOOTING PROCEDURES AND REQUESTED ABBOTT FIELD SERVICE. THE CTA SENT THE CUSTOMER REPLACEMENT SYRINGES. AN ABBOTT FIELD SERVICE REPRESENTATIVE PERFORMED MAINTENANCE, VERIFIED ASSAY CONTROLS WERE WITHIN SPECIFICATION AND THE ISSUE WAS RESOLVED. THE CUSTOMER REPORTED, THERE WAS NO IMPACT TO PT MANAGEMENT. | 0
1174884 | A CERNER REPRESENTATIVE CALLED TO REPORT AN ISSUE WITH AN INCORRECT SAMPLE ID BEING SENT TO THE LAB INFORMATION SYSTEM (LIS) AND THE RESULTS WERE REPORTED TO THE WRONG PATIENT. THE ACCOUNT STATED THAT THE SAMPLE IS ALWAYS AN 8 DIGIT NUMBER SO IF A SAMPLE ID TRANSMITS ACROSS WITH FEWER DIGITS, THE ACCOUNT FILLS IT IN WITH NUMBERS THAT CORRESPOND WITH THE JULIEN DATE FOR THAT DAY. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). OTHER: INSUFFICIENT INFORMATION TO MAKE ANY HYPOTHESIS. NO HARDWARE OR SOFTWARE ISSUE IDENTIFIED WITH AVAILABLE DATA. ON (B)(6) 2008, OUR (B)(4) OF (B)(6) HOSPITAL IN (B)(6) REPORTED THAT THERE WERE THREE (3) INSTANCES WHEN SOME SAMPLES IDS, RUN CORRECTLY ON A NON-CELL-DYN SAPPHIRE INSTRUMENT AT (B)(6) HOSPITAL, SHOWED UP ON THE PENDING LIST OF THE CELL-DYN SAPPHIRE A FEW MONTHS LATER. NO DATA PERTAINING TO THIS ISSUE WAS SUBMITTED FOR INVESTIGATION. AS PART OF THE INVESTIGATION, THE ISSUE WAS CONSULTED WITH A SENIOR SOFTWARE ENGINEER, SUBJECT MATTER EXPERT. THE INVESTIGATION WAS BASED ON THE INFORMATION DOCUMENTED IN THE REPORTED EVENT, AS NO DATA WAS SUBMITTED FOR THIS ISSUE. THE INVESTIGATION INTO THE ISSUE CONCLUDED THAT THERE WERE 3 INSTANCES WHERE A SAMPLE WAS RUN CORRECTLY AT A (B)(6) HOSPITAL (NON-CELL-DYN SAPPHIRE INSTRUMENT) AND THEN MONTHS LATER THE SAME SAMPLE ID SHOWED UP IN A PENDING LIST AT OUR (B)(6) HOSPITAL. THE ONLY RESULTS PENDING WHERE THE EXTRA TESTS THAT THE CELL-DYN SAPPHIRE PERFORMED. IT APPEARED THAT WHEN A BARCODE WAS READ, THE SAMPLE ID WAS BEING EDITED AND THE 10-DIGIT NUMBER WAS NOT ENTERED COMPLETELY. WHEN CERNER, THE COMPUTER SYSTEM PROCESSING DATA, SAW 6 DIGITS INSTEAD OF 10, IT FILLED IN THE FRONT NUMBERS USING THE (B)(6) DATE, WHICH BROUGHT IT BACK TO A SPECIMEN ID ALREADY RUN AT THE OTHER SITE. THE SPECIMEN THAT SHOWED UP ON THE PENDING LIST AS (B)(4) HAD RUN AT THE OTHER SITE AS (B)(4). IT APPEARED TO BE AN ISSUE WITH THE CUSTOMER'S LABORATORY INTERFACE SPECIFICATIONS (LIS) FILLING IN MISSING DIGITS WITH THE (B)(6) DATE, WHICH MAY CAUSE IT TO INADVERTENTLY CREATE A DUPLICATE SPECIMEN ID. THE SUBJECT MATTER EXPERT SUGGESTED THAT THIS SHOULD HAPPEN AGAIN; THE LOG FILES AND CONFIGURATION FILES FROM THE CELL-DYN SAPPHIRE SHOULD BE COLLECTED FROM THE CUSTOMER SOON AFTER THE EVENT HAPPENS. A PRINTOUT OF THE DATALOG AND THE SPECIMEN REPORT WOULD ALSO BE HELPFUL IN ORDER TO FULLY INVESTIGATE THE ISSUE. SECTION 2 OF THE CELL-DYN SAPPHIRE, LIST NUMBER (B)(4), PROVIDES INFORMATION REGARDING PROPER LIS SET UP. SECTION 5 RECOMMENDS SETTING UP A LABORATORY PROCEDURE TO REQUIRE ANY UNPROCESSED WORK LIST ENTRIES BE VIEWED AND CLEARED AT THE END OF EACH SHIFT OR DAY. USE OF THIS PROCEDURE WILL MAINTAIN AN UP-TO-DATE WORK LIST AND REDUCE THE OPPORTUNITY FOR ANY UNPROCESSED SPECIMEN IDS, LEFT IN THE WORK LIST FOR AN EXTENDED TIME, TO BE MATCHED WITH A DIFFERENT PATIENT WITH THE SAME SPECIMEN ID. APPENDIX A PROVIDES LIST NUMBER (B)(4) FOR THE LABORATORY INTERFACE SPECIFICATIONS FOR THE CELL-DYN SAPPHIRE AVAILABLE FOR CUSTOMERS TO ORDER. A REVIEW OF DOMESTIC AND INTERNATIONAL COMPLAINT ANALYSIS TRENDING SYSTEM (CATS) AND TRENDING ANALYSIS SUPPORT SYSTEM (TASS) TREND ANALYSIS REPORTS ON THE (B)(4) DATABASE FOR THE MONTHS OF (B)(6) 2007 THROUGH (B)(6) 2008 HAVE BEEN EVALUATED FOR THE COMPLAINT ISSUE. NO ADVERSE TRENDS HAVE BEEN IDENTIFIED FOR THE CELL-DYN SAPPHIRE INSTRUMENT, LIST NUMBER (B)(4), REGARDING THIS ISSUE. THE (B)(6) 2008 REPORT IS NOT YET AVAILABLE. THIS IS THE FINAL REPORT. | 0
1015496 | INVESTIGATION SUMMARY; THE CUSTOMER OBSERVED DISCREPANT RESULTS FROM OPEN TO CLOSED MODE. THE INSTRUMENT IN USE WAS A CELL-DYN 3700SL. THE ISSUE WAS DISCOVERED THE DAY BEFORE WHEN PATENT CONTROLS WERE RUN (USED TO MONITOR CLOSED MODE). THERE WERE NO ERROR MESSAGES SEEN. THE CUSTOMER PERFORMED TROUBLESHOOTING WITHOUT RESOLUTION. THE CUSTOMER TECHNICAL ADVOCATE (CTA) VERIFIED THERE WERE NO OBSTRUCTIONS IN THE CLOSED MODE ASPIRATION PATHWAY. CALIBRATION WAS DONE IN OCTOBER 2007. CTA REQUESTED THAT THE CUSTOMER RE-CLEAN THE NEEDLE AND Y FITTING IN CLOSED MODE, CLEAN THE SHEAR VALVE AND THEN RUN PRECISION IN OPEN AND CLOSED MODES PRIOR TO CALIBRATING. THE CUSTOMER REPORTED THE RECOMMENDED TROUBLESHOOTING HAD BEEN DONE BUT WHILE RUNNING THE PRECISION TESTING, FOUND THE RESULTS OUT OF SPEC. THE CUSTOMER OBSERVED THE BLOOD WAS NOT BEING CLEARED FROM THE TUBING PRIOR TO GOING TO THE WASTE CHAMBER. THE FIELD SERVICE REP (FSR) WAS DISPATCHED TO THE CUSTOMER SITE. THE FSR FOUND A STICKY SOLENOID VALVE (VALVE 3 ON THE VPM- VALVE PRESSURE MODULE). VALVE 3 CONTROLS THE P2 READING. THE FSR CLEANED THE TUBING FOR THE WASTE LINES THEN VERIFIED THE INSTRUMENT BY RUNNING CONTROLS. THE SYSTEM WAS PERFORMING WITHIN SPEC. NO FURTHER INVESTIGATION WAS REQUIRED. THE CD3700 SYSTEM OPERATOR'S MANUAL (REV M): TROUBLESHOOTING CHARTS FOR ISSUES WITH IMPRECISE OR INACCURATE DATA GIVE INADEQUATE MIXING OR DRAINING OF THE MIXING CHAMBER, AS A RESULT OF A PROBLEM IN THE PRESSURE SYSTEM (E.G. LEAK OR OBSTRUCTION) AS A CAUSE OF IMPRECISION. (10-56 TO 10-60). TRENDING: A REVIEW OF COMPLAINT REPORTS, FOR THE PERIOD APRIL 2007 THROUGH DECEMBER 2007, DID NOT INDICATE ANY ADVERSE TREND FOR CELL- DYN 3700, LIST BASE 02H31-01 FOR THE COMPLAINT ISSUE. LABELING: THE EVENT IS ADDRESSED IN THE CELL-DYN 3700 SYSTEM OPERATOR'S MANUAL, 06H89-01, REV F TROUBLESHOOTING: WOC DATA ARE IMPRECISE OR INACCURATE; 10-58. WIC DATA ARE IMPRECISE OR INACCURATE; 10-59. HGB DATA ARE IMPRECISE OR INACCURATE; 10-60. CONCLUSION: THE DEVICE INVOLVED IN THE ISSUE WAS EVALUATED BY VISUAL EXAMINATION AND PERFORMANCE TESTING. A COMPONENT/SUBASSEMBLY MALFUNCTION CONTRIBUTED TO THE EVENT. THIS ISSUE WAS RESOLVED BY A SERVICE REP AND VERIFYING THE ANALYZER WAS PERFORMING TO SPECS. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATES THAT THE CELL-DYN 3700 SL ANALYZER HAS GENERATED DISCREPANT RESULTS WHEN COMPARING RESULTS OBTAINED IN OPEN VS CLOSED MODE OF OPERATION. THE CUSTOMER PROVIDED THE FOLLOWING DATA. OPEN MODE; WCB=4.1 K/UL, HGB=14.7 G/DL. CLOSED MODE; WBC=1.5 K/UL, HGB=5.6 G/DL. RESULTS WERE NOT RELEASED FROM THE LAB WITH NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
998076 | A TECHNICAL SERVICE SPECIALIST (TSS) WAS SENT TOT HE ACCOUNT. THE TSS RAN 7 DEGREE LATEX PARTICLES, WHICH RECOVERED WITHIN SPECIFICATION. WHEN SAMPLES AND CONTROLS WERE RUN, THE PLT RESULTS WERE ERRATIC. THE DILUENT/SHEATH FILTER WAS CHECKED AND FOUND BLOCKED. A NEW INLINE SHEATH FILTER WAS INSTALLED WHICH RESOLVED THE ISSUE. THE CD3200 SYSTEM OPERATOR'S MANUAL (REV M), ON PAGES P 3-31 STATES: 'A RESULT THAT FALLS OUTSIDE A LABORATORY ACTION LIMIT CAN ALSO INDICATE THE NEED FOR THE OPERATOR TO FOLLOW A LABORATORY PROTOCOL, SUCH AS REPEATING THE SAMPLE, NOTIFYING THE PHYSICIAN OR PERFORMING A SMEAR REVIEW ...' THE PLT VALUES RECOVERED WERE OUT OF RANGE LOW AND WERE REPEATED. REPEATING THE SAMPLES IDENTIFIED THAT PLT RESULTS WERE ERRATIC. THE INLINE SHEATH FILTER IS MONTHLY MAINTENANCE ITEM: 'THE DILUENT/SHEATH FILTER IS LOCATED ON THE FRONT PANEL TO THE LEFT OF WASTE CHAMBER #3. REPLACE THE FILTER ONCE A MONTH, OR WHENEVER CONTAMINATION IS SUSPECTED. A SIGN OF CONTAMINATION IS USUALLY MANIFESTED BY HIGH PLATELET BACKGROUND, POORLY DEFINED PLT-RBC (0 DEGREES / 10 DEGREES) SCATTERPLOT, EXCESSIVE WBC FLAGGING, OR ERRONEOUS 5-PART DIFFERENTIAL RESULTS.' (PP. 9-25/26). IN ADDITION, THE COMPLAINT ANALYSIS TRENDING SYSTEM (CTAS) AND TRENDING ANALYSIS SUPPORT SYSTEM (TASS) REPORTS WERE REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THIS IS THE FINAL REPORT. END OF REPORT. >< THE ACCOUNT STATED THAT THE CELL DYN 3200 SL ANALYZER GENERATED INCORRECT PLATELET (PLT) RESULTS. AN INITIAL PLT RESULT OF 96,000/UL WAS GENERATED WHICH REPEATED AT 12,000/UL AND WAS 7,000/UL ON THE CELLDYN 4000 ANALYZER. | 0
1033689 | AROUND 09:00 AM, A NURSE WAS IN ROOM << SOINS5 >> AND AN ASYSTOLE OCCURRED IN ROOM <<SOINS1 >>. THE AUDIBLE ALARM COULD BE HEARD AT THE BEDSIDE 'SOINS1' AND AT THE INTELLIVUE INFORMATION CENTER (IIC). ALTHOUGH THE USER REPORTED THAT THE CAREGROUP AND AUTOMATIC WINDOW POPUP WAS ON, NO WINDOW POPUP WAS SEEN ON THE MONITOR IN ROOM 'SOINS5'. THERE WAS NO DOUBLE BIP (WHICH MEANS THE APPARITION OF THE POPUP WINDOW) EITHER ON THE MONITOR. BUT THE NURSE HEARD THE ALARM OF THE IIC A TIME LATER. THE PATIENT HAD A CARDIAC ARREST AND DIED DUE TO THE LATE INFORMATION. >< TESTING SUPPORTS THAT THE INVOLVED DEVICE DID NOT MALFUNCTION. IT WAS REPORTED TO PHILIPS THAT THE USERS WERE AWARE OF THE ALARMING. THE USERS DID NOT RESPOND TO THE ALARMS. PHILIPS IS NOT ABLE TO DETERMINE IF THE HIGHER PRIORITY MONITORING INTERACTIONS PREVENTED CAREGROUP ALARMS OR IF A USER HAD DISABLED THE CAREGROUP ALARMS FEATURE. THE AVAILABLE INFORMATION DOES NOT SUPPORT THAT THERE WAS A MALFUNCTION. | 0
1074286 | THE CUSTOMER DID NOT PROVIDE THE SERIAL NUMBER. THE HOSPITAL BIOMEDICAL ENGINEER TESTED THE DEVICE USING A SIMULATOR AND COULD NOT DUPLICATE THE ISSUE. A PHILIPS FIELD SERVICE ENGINEER VISITED THE SITE AND FOUND THAT ALL ECG ALARMS WERE DISABLED AT THE BEDSIDE AND NOT AT THE CENTRAL STATION. A NURSE STATED THAT SHE DISABLED THE YELLOW ALARMS AT THE CENTRAL STATION. THE USER THOUGHT THAT THEY WERE ONLY DISABLING THE YELLOW ALARMS, HOWEVER, WHEN DISABLING ECG ALARMS, ALL ALARMS ARE DISABLED (NOT JUST YELLOW ALARMS). ALARM LOG FILES AND RECORDING LOGS WERE RETRIEVED AND REVIEWED BY A PHILIPS CLINICIAN AND AN ENGINEER. THE ALARM LOGS SHOW THAT THE PT'S BED HAD NO DATA OR ACTIVITY FOR FOUR MONTHS IN 2008, WHICH SUPPORTS THAT THE BESIDE MONITOR WAS DISCONNECTED FROM THE NETWORK AND RECONNECTED FOUR DAYS LATER. ANALYSIS OF THE RECORDER LOGS ALSO SUPPORTS THE NETWORK DISCONNECT WHERE AUTOMATIC RECORDING WAS NOT AVAILABLE FOR FOUR DAYS, HOWEVER, DURING THIS TIME, MANUAL RECORDING WAS AVAILABLE AND 3 MANUAL RECORDINGS WERE PERFORMED. BASED ON THE LOG FILES THE AUTOMATIC RECORDING CAPABILITY WAS REESTABLISHED. THE DEVICE DID NOT MALFUNCTION. THE DEVICE REMAINS IN USE AT THE CUSTOMER SITE. >< THE CUSTOMER REPORTED THAT A PT WENT INTO ASYSTOLE AND VTACH AND EXPIRED. NO RED ALARMS WERE PRESENT. | 0
989183 | A DATASCOPE CLINICAL EDUCATION SPECIALIST REVIEWED THE DETAILS OF THE INCIDENT WITH THE CUSTOMER OVER THE PHONE SHORTLY AFTER THE EVENT. SHE ASSISTED THE CUSTOMER IN CHECKING THE CENTRAL ALARM SETTINGS AND CONFIRMED THAT THEY WERE SET APPROPRIATELY. DATASCOPE SERVICE THEN CONTACTED THE CUSTOMER AND DETERMINED THAT THE CENTRAL HAD PRODUCED A VISUAL BUT NOT AN AUDIBLE ALARM AT THE TIME OF THE EVENT. HE DETERMINED THAT THE CUSTOMER HAD MOVED THE LONGVIEWS (REMOTE KEYBOARD, VIDEO, MOUSE) TO THE NURSES STATION APPROXIMATELY ONE WEEK PRIOR TO THE INCIDENT. THEIR ORIGINAL LOCATION AS INSTALLED BY DATASCOPE HAD BEEN ON A WALL, ATTACHED BY BRACKETS. THE CUSTOMER HAD NEGLECTED TO CONNECT TWO OF THE LONGVIEW'S AFTER MOVING THEM. THE CUSTOMER RECONFIGURED THE LONGVIEWS. >< THE CUSTOMER ADVISED DATASCOPE SERVICE THAT WHILE A PATIENT WAS BEING MONITORED WITH A PANORAMA CENTRAL STATION AND A TELEPACK AMBULATORY TRANSMITTER, THE PATIENT WENT INTO CARDIAC ARREST AND NO ALARM SOUNDED AT THE CENTRAL. THE PATIENT'S FAMILY ADVISED THE CLINICIANS AS TO THE PATIENT'S CONDITION. THE PATIENT EXPIRED. | 1
937826 | A CUSTOMER SITE REPORTED WASTE FLUID HAD OVERFLOWED FROM THE WASTE CARBOY ON A BENCHMARK XT INSTRUMENT ONTO THE FLOOR WITHOUT AN ALARM. NO ONE WAS INJURED AND PATIENT CARE WAS NOT AFFECTED. THE FIELD SERVICE ENGINEER EVALUATED THE UPGRADED WASTE SENSOR AND FOUND THE TUBE TO BE BROKEN FROM MISHANDLING PREVENTING THE SENSOR FROM FUNCTIONING PROPERLY. THE FSE (FIELD SERVICE ENGINEER) INSTALLED A NEWER DESIGNED WASTE SENSOR WHICH HAS THE TUBE COUNTER SUNK PREVENTING LATERAL MOVEMENT WHICH PREVENTS THE CUSTOMER FROM BREAKING THE TUBE IF MISHANDLED. WHILE THERE IS ALWAYS A POSSIBILITY THAT A SLIP AND FALL EVENT DUE TO A FLUID SPILL COULD RESULT IN DEATH OR SERIOUS INJURY, ALL EVIDENCE INDICATES THAT THE POSSIBILITY OF DEATH OR SERIOUS INJURY IS REMOTE. >< THE MOST PROBABLE CAUSE FOR THE BROKEN TUBE IS THE CUSTOMER MISHANDLING THE WASTE SENSOR. THIS IS A KNOWN ISSUE AND A NEWER DESIGNED WASTE SENSOR HAS BEEN INSTALLED ON THE INSTRUMENT TO PREVENT THIS FROM HAPPENING IN THE FUTURE. BECAUSE OF THE REMOTE POSSIBILITY OF INJURY RESULTING FROM A SLIP AND FALL EVENT CAUSED BY WASTE FLUID OVERFLOW, THE FIRM IS IN THE PROCESS OF IMPLEMENTING AN IMPROVED WASTE SENSOR DESIGN AS A REPLACEMENT TO THE CURRENT SENSOR. THIS NEW WASTE SENSOR DESIGN SHOULD ELIMINATE THE OBSERVED FAILURE MODE. THE WASTE SYSTEM ON THIS INSTRUMENT WAS UPGRADED WITH THE NEW SENSOR DESIGN ON 09/04/07. COMPLETE FIELD IMPLEMENTATION OF THE NEW SENSOR DESIGN TO BE COMPLETED BY 2007/10/31. THE NEWER WASTE SENSOR DESIGN WAS INTRODUCED TO THE MANUFACTURING FLOOR ON 2007/02/08. A FINAL REPORT WILL BE SENT PENDING THE COMPLETION OF THE FIELD IMPLEMENTATION OF THE NEW WASTE SENSOR. | 0
988476 | OVERHEATED ELECTRICAL COMPONENT INSIDE UNIT MELTED HOLE IN CASE. POTENTIAL FOR PT BURN OF FIRE. NO INJURY OCCURRED. ACCORDING TO WELCH ALLYN REPRESENTATIVE, POSSIBLE L2 CONDUCTED ISSUE. | 1
1217044 | THE OPTIA MACHINE BEGAN MAKING NOISE SECONDS INTO THE TREATMENT FROM THE CENTRIFUGE COMPARTMENT. A LOUD BANGING SOUND WAS HEARD WHILE MACHINE WAS BEING PAUSED. ON INVESTIGATION THE STEM OF THE CENTRIFUGE TUBE WAS TORN OFF AND LEAKING BLOOD.====================== MANUFACTURER RESPONSE FOR HEMAPHERESIS MACHINE, OPTIA SPECTRA======================THE SERVICE REPRESENTATIVE FROM THE MANUFACTURER OF THE DEVICE STATED THAT IT MALFUNCTIONED FOR ONE OF THREE REASONS: THE SETUP MAY NOT HAVE BEEN LOADED CORRECTLY, AFTER SET UP WAS LOADED A CAP MAY HAVE FALLEN INTO THE CENTRIFUGE CAUSING THE REALIGNMENT OF THE TUBING, OR THE TUBING COULD HAVE BEEN DEFECTIVE. | 0
998175 | OUR EVALUATION CONFIRMED THAT THE TOP COVER OF THE DEVICES ACTUALLY SEPARATED FROM THE BASE AS REPORTED. WE ALSO FOUND THAT ONE OF THE UNITS WAS MODIFIED (TOP COVER GLUED TO BASE) BY THE CUSTOMER. THIS IS A KNOWN SUPPLIER ISSUE AND WE HAVE IMPLEMENTED SUPPLIER CORRECTIVE ACTION AND INITIATED A FIELD CORRECTION. THE DEVICES IDENTIFIED IN THIS INCIDENT FALL IN THE RANGE OF THE POTENTIALLY DEFECTIVE DEVICES OF THE FIELD CORRECTION. >< DRAEGER MEDICAL SYSTEMS, INC. HAS RECEIVED INFORMATION FROM A CUSTOMER WHO HAS REPORTED THAT TWO OF OUR INFINITY DOCKING STATIONS (IDS) UNITS WERE FOUND TO HAVE SEPARATED (TOP COVER FROM THE BASE COMPONENT). NO PATIENT INJURY WAS REPORTED. | 1
998058 | THE CUSTOMER REPORTED THAT SOMETIMES PRESCRIBED ORDERS ARE NOT SHOWING IN THE WORKLIST. >< PHILIPS HAS DETERMINED THAT SOME PHYSICIAN ENTERED ORDERS ARE NOT SHOWING UP ON WORKLISTS. NEW INFORMATION RECEIVED 07 FEBRUARY 2008 IDENTIFIED PATIENT RISK. PHILIPS IS IN THE PROCESS OF OBTAINING MORE INFORMATION CONCERNING THIS EVENT, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 1
953966 | A CUSTOMER CONTACTED BECKMAN COULTER REGARDING LOW HEMOGLOBIN (HGB), AND PLATELET (PLT) RESULTS THAT WERE GENERATED BY THE COULTER ACT DIFF INSTRUMENT. A PATIENT SAMPLE WAS TESTED FOR HGB AND PLT: HGB RESULT WAS 9.3 G/DL, AND PLT RESULT WAS 137 X 10 TO THE THIRD POWER CELLS/UL. BOTH RESULTS WERE PRINTED WITH AN OPERATOR DEFINED MESSAGES 'L' (L-LOW RESULT. FOR PATIENT SAMPLES, RESULT IS LOWER THAN THE LOW PATIENT SAMPLE LIMIT'). HGB AND PLT RESULTS WERE REPORTED OUT OF THE LAB AND THE PATIENT WAS SENT TO A HOSPITAL. UPON REDRAW AT THE HOSPITAL, HGB AND PLT RESULTS WERE 'NORMAL' AND THE PATIENT WAS RELEASED FROM THE HOSPITAL WITHOUT TREATMENT. HOSPITAL PRINTOUT WAS NOT PROVIDED. THE CUSTOMER STATED THAT THE DAY OF THE EVENT (2007), THE BATHS WERE NOT DRAINING PROPERLY. AFTER CORRECTING THE DRAIN ISSUE, THE CUSTOMER RERAN THE SAMPLE FOR HGB AND PLT AND THE REPEATED RESULTS WERE: 14.1 G/DL AND 169 X 10 TO THE THIRD POWER CELLS/UL RESPECTIVELY. THE REPEATED RESULTS CORRELATED WITH THE HOSPITAL RESULTS WHICH WERE CONSIDERED CORRECT. BASED ON AVAILABLE INFORMATION, THERE WAS NO IMPACT TO PATIENT OR USER. >< QC WAS RUN BEFORE AND AFTER CORRECTING THE BATH OVERFLOW ISSUE AND ALL LEVELS RECOVERED WITHIN EXPECTED RANGES. THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS. PREVIOUSLY TESTED PATIENT SAMPLES WERE RERUN TO CONFIRM ACCURATE BACK TO THE LAST ACCEPTABLE QC RUN. THE SAMPLE WAS DRAWN IN A 4.0 ML VACUTAINER AND WAS STORED AT AMBIENT TEMPERATURE FOR 15 MINUTES PRIOR TESTING. THE ERRONEOUS RESULTS OCCURRED IN AN OPEN VIAL WHOLE BLOOD MODE OF OPERATION AND CUSTOMER STATED THAT THIS WAS THE LAST SAMPLE TESTED IN 2007. BASED ON AVAILABLE INFORMATION, CUSTOMER REPLACED THE PERISTALTIC PUMP TUBING ON THE WASTE PUMP AND CHECK VALVE IN THE DRAIN SYSTEM. THERE WAS NO PROBLEM WITH THE DRAINING, BUT THE SYSTEM GENERATED A VACUUM ERROR AND CUSTOMER REPLACED THE FLUID BARRIER AND ADJUSTED THE VACUUM. AFTER SERVICING THE INSTRUMENT, CUSTOMER PERFORMED A STARTUP AND QC AND OBTAINED ACCEPTABLE RESULTS. AS PER PRODUCT LABELING, WHEN ALL COUNTED PARAMETERS ARE CONSISTENTLY LOWER THAN NORMAL AND MCV IS NORMAL, THIS CAN BE ATTRIBUTED TO THE FOLLOWING: SHORT SAMPLE, POOR BATH DRAIN, DILUTED SAMPLE. THE LOW RESULTS CAN BE ATTRIBUTED TO A DILUTED SAMPLE CAUSED BY THE BATH NOT DRAINING PROPERLY. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. | 0
1230760 | BAXTER INVESTIGATION REVEALED THE FOLLOWING DETAILS: THE HOMECHOICE INSTRUMENT WAS RECEIVED IN BAXER UK, BUT NOT EVALUATED YET. REGARDING THE PRIMARY AND SECONDARY CAUSE OF DEATH, NO INFORMATION WAS PROVIDED. THE PATIENT'S PHYSICIAN STATED THAT, THE CAUSE OF DEATH DOES NOT NEED TO BE KNOW BY BAXTER BECAUSE THE DEATH WAS UNRELATED TO THE HOMECHOICE INSTRUMENT. BAXTER REQUESTED THE INSTRUMENT FOR EVALUATION. THE RESULTS WILL BE PROVIDED UPON COMPLETION. >< BAXTER COMPLAINT COORDINATOR IN ANOTHER COUNTRY REPORTED THAT PATIENT PASSED AWAY WHILST ON THE MACHINE. BAXTER FIELD SERVICE ENGINEER WAS CALLED BY THE PATIENT'S DAUGHTER FOR HELP TO END THE THERAPY, REMOVE THE CASSETTE, AND TURN OFF THE MACHINE. NO FURTHER INFORMATION WAS AVAILABLE AT THIS TIME. FURTHER INFORMATION RECEIVED FROM INT'L AFFILIATE. THE PATIENT'S PHYSICIAN, HAS CONFIRMED THAT THE PATIENT WAS NOT CONNECTED TO THE HOMECHOICE MACHINE AT THE TIME OF DEATH. THE PATIENT HAD SET UP HIS MACHINE, AND THEN WENT TO THE BATHROOM. HE SUDDENTLY PASSED AWAY IN THE BATHROOM, PRIOR THE PD THERAPY. THE FAMILY CALLED THE TECHNICAL SERVIES HELPLINE TO ASK HOW TO DISCONNEST THE MACHINE, NOT HOW TO DISCONNECT THE PATIENT. NO FURTHER INFORMATION AVAILABLE AT THIS TIME. | 0
1024339 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1024344 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1024342 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
986027 | ADD'L EVALUATION INCLUDED ENGINEERING REVIEW OF THE CODE THAT HANDLES MPR IMAGE ORIENTATION. AFTER INVESTIGATING THIS SITUATION, IT IS PHILIPS' OPINION THAT THERE IS MINIMAL RISK OF PT INJURY IF THE MPR (RECONSTRUCTED) IMAGE IS INCORRECTLY DISPLAYED IN ISITE PACS BECAUSE: THE ORIGINAL (CORRECT) IMAGE IS ALWAYS AVAILABLE IN THE PT TIMELINE. THE ORIGINAL (CORRECT) IMAGE IS ALWAYS DISPLAYED ON THE MONITOR NEXT TO THE POSSIBLE INCORRECT IMAGE (NOT ALL MPRS ARE INCORRECT, ONLY THOSE IMAGES THAT ARE ACQUIRED AS DESCRIBED. ANATOMICAL MARKERS ARE OFTEN PRESENT (MAKING IDENTIFICATION OF THE INCORRECT DISPLAY LIKELY). ANY FINDINGS WOULD GENERATE ADD'L IMAGING PRIOR TO SURGERY OR ACTIVE INTERVENTION CONFIRMING ANATOMICAL LOCATION. THIS IS THE SAME ISSUE PRESENTED IN REPORTED MDR 295704-2008-00001 AND SUBJECT TO THE SAME CORRECTIVE RECALL (2954704-01042008-001). >< IN EARLY 2008, A CUSTOMER COMPLAINT WAS REPORTED THAT WHEN USING ISITE PACS TO PERFORM MPR THE LEFT AND RIGHT ORIENTATION MARKERS WERE BEING DISPLAYED INCORRECTLY. FURTHER INVESTIGATION REVEALED THE REVERSE ORIENTATION MARKER DISPLAY AFFECTS ONLY RECONSTRUCTED CORONAL AND SAGITTAL IMAGES. ADDITIONALLY, THE ORDERING OF CORONAL AND SAGITTAL IMAGE SLICES CAN ALSO BE REVERSE. FOR AN EXAMPLE, A SCOUT LINE ON THE RIGHT SIDE OF AN AXIAL IMAGE COULD INTENTIONALLY DISPLAY THE CORRESPONDING SAGITTAL SLICE POSITION ON THE LEFT SIDE. SUBSEQUENT IN-HOUSE TESTING REVEALED THAT THE REPORTED ISSUE IS REPRODUCIBLE. THE ORIGINAL AXIAL SLICES ARE ALWAYS DISPLAYED AND ANNOTATED CORRECTLY IN ALL CASES. THE ORIGINAL DICOM IMAGES ARE ON ISITE PACS AND PRESENTED IN CORRECT ORIENTATION NEXT TO THE MPR IMAGES WHEN THEY ARE DISPLAYED ON THE MONITOR. THERE ARE NO REPORTS OF PT INJURY RELATED TO THIS ISSUE. | 1
1036866 | . >< A LAWYER REPORTED A FEMALE CONSUMER USED CUSHION GRIP AND ORAFIX. HE REPORTED AS A RESULTS SHE SUSTAINED SEVERE, PERMANENT AND DISABLING INJURIES INCLUDING ACUTE MYELOGENOUS LEUKEMIA FROM WHICH SHE DIED IN 2005. NO ADDITIONAL INFO WAS PROVIDED. | 0
983089 | IT WAS REPORTED THAT THE PT WAS ON A MONITOR, WHEN SHE EXPERIENCED A WITNESSED ARREST AND THE MONITOR'S ALARM REPORTEDLY DID NOT SOUND. THE HOSP FURTHER STATED, THAT THE ALARM HAD BEEN PREVIOUSLY TURNED OFF AT THE CIC BY HOSP STAFF AS THE PT WAS IN THE PROCESS OF RECEIVING AM CARE. THE ALARM REPORTEDLY HAS NOT BEEN TURNED BACK ON. THERE IS NO ALLEGATION OF EQUIPMENT MALFUNCTION. A PT DEATH WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1000159 | A CUSTOMER REPORTED THAT A PATIENT BEING TREATED FOR LUNG CANCER HAD BEEN TREATED FOR 5 FRACTIONS WITH AN INCORRECT ISOCENTER (5 CM ERROR). THE ISOCENTER ERROR WAS INTRODUCED BY AN IMPORT APPLICATION PROBLEM IN FOCAL. HOWEVER, THE ERROR IS HIGHLY DETECTABLE AND IN THIS CASE, THE XIO VISUAL DISPLAY CLEARLY ALERTED DOSIMETRY TO THE PROBLEM. BASED ON CUSTOMER INPUT, THE INFORMATION LEAVING DOSIMETRY TO THE TREATMENT ROOM WAS CORRECT, BUT ADDITIONAL USER ERRORS SEEMED TO CAUSE THE ACTUAL MISTREATMENT. CLINICIANS ADJUSTED THE REMAINING FRACTIONS TO COMPENSATE. THE IMPORT APPLICATION PROBLEM WAS INTRODUCED IN FOCAL 4.3.0 AND CORRECTED IN FOCAL 4.34. >< A CUSTOMER ADVISORY IS BEING DRAFTED FOR TRANSMITTAL TO ALL CUSTOMERS OF OUR FOCAL WORKSTATION TO ADVISE THEM OF THIS SITUATION. | 1
1174977 | THE X-RAY MACHINE MALFUNCTIONED WHEN TECH WAS ATTEMPTING TO TAKE AN X-RAY ON A PATIENT. THE MACHINE DID NOT BEEP WHEN THE X-RAY WAS BEING TAKEN. WHILE TECHS WERE ASSESSING THE PROBLEM, THE PATIENT'S MOTHER STATED THAT THERE WAS SMOKE. IT WAS COMING FROM THE DIGITAL DETECTOR THAT WAS NEAR THE PATIENT. WHEN THE TECHNICIAN PICKED UP THE DETECTOR BY THE HANDLE, SHE TOUCHED THE CABLE WHICH WAS VERY HOT. DEVICE WAS REMOVED FROM SERVICE.====================== MANUFACTURER RESPONSE FOR X-RAY MACHINE, PORTABLE, DIGITAL======================GE REPLACED BOTH THE CABLE THAT GOT HOT AS WELL AS THE POWER SUPPLY. GE CONTENDS THAT A CIRCUIT BREAKER WITHIN THE POWER SUPPLY SHOULD HAVE PREVENTED THIS TYPE OF EVENT. POWER SUPPLY WAS SENT TO GE HEADQUARTERS FOR FURTHER EVALUATION. | 1
1052453 | ONE PT WITH FALSE NEGATIVE LEUKOCYTE RESULT THAT CONFIRMED 20-30 WBC/HPF. INITIAL RESULT NOT REPORTED, PT NOT ADVERSELY AFFECTED. IF ADDITIONAL INFO IS RECEIVED, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
1063021 | CUSTOMER SERVICE DOCUMENTED CUSTOMER RECEIVING ERRATIC READINGS ON THEIR FREESTYLE FLASH BLOOD GLUCOSE MONITOR. CUSTOMER REPORTED TWO READINGS OF 75 MG/DL, 37 MG/DL AND 'HI' (GREATER THAN 500 MG/DL) WITHIN 10 MINUTES. ALL TESTS WERE DONE ON THE ARM. THE RESULTS WHEN PLOTTED ON A PARKES ERROR GRID FELL INTO THE 'C' ZONES SHOWING THE DIFFERENCE IN VALUES TO BE CLINICALLY SIGNIFICANT. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THE COMPLAINT WAS NOT CONFIRMED AND NO NEW ISSUES WERE OBSERVED, ALL RESULTS WERE WITHIN THE RANGE SPECIFICATION AND NO ERRORS WERE OBSERVED DURING CONTROL SOLUTION TESTING. ACCORDING TO THE MEMORY LOG OF THE RETURNED METER THE VALUES OF THE CUSTOMER RECEIVED ARE 75 MG/DL, 37 MG/DL, 75 MG/DL, AND A 'HI' WITHIN 10 MINUTES WERE FOUND IN 2008. | 0
1003917 | PATIENT HAD AAA REPAIR. AAA REPAIRED WITH STENT GRAFT AND CARDIOMEM SENSOR PLACED BY SURGEON IN ANEURYSM SAC. IT WAS NOTED AT THE END OF THE SURGICAL PROCEDURE THAT A COUNTABLE ITEM WAS MISSING. AN ANTERIOR AND POSTERIOR X-RAY TAKEN OF THE ABDOMEN. THERE WERE NOT ANY COUNTABLE ITEMS IN THE ABDOMINAL CAVITY ON X-RAY. PATIENT HAD A POST-OP FOLLOW-UP VISIT WITH THE SURGEON APPROXIMATELY A MONTH LATER. ANOTHER ABDOMINAL FILM WAS TAKEN. THIS FILM REVEALED A LONG OBJECT (~12 CM) RESTING NEXT TO THE ANEURYSM AREA. SURGEON AND RADIOLOGIST REVIEWED THE FILM, AND THEY WERE UNABLE TO DETERMINE THE IDENTITY OF THE OBJECT. THE SURGEON DECIDED AT THIS TIME TO LEAVE THE RETAINED OBJECT BECAUSE IT DID NOT SEEM TO BE CREATING ANY PROBLEMS FOR THE PATIENT. PATIENT RETURNED TO ED 26 DAYS LATER WITH A RUPTURED AAA. FAMILY MADE DECISION IN THE ED TO NOT TREAT SURGICALLY. PATIENT WAS ADMITTED TO PALLIATIVE CARE. PATIENT EXPIRED TWO DAYS LATER. AUTOPSY PERFORMED AND REVEALED RETAINED DEVICE TO BE THE TIP OF THE CARDIOMEM DELIVERY CATHETER. DELIVERY CATHETER TIP WAS SENT FOR EXAMINATION. PRELIMINARY AUTOPSY RESULTS REVEALED THAT A SHARP BONE SPUR ON THE PATIENT'S SPINE RUPTURED THE POSTERIOR WALL OF THE ANEURYSM, AND THE RETAINED DEVICE WAS NOT THE CAUSATIVE AGENT CREATING THE AAA RUPTURE. AFTER AUTOPSY RESULTS WERE RETURNED TO SURGEON, HE REMEMBERED THAT THE FIRST CARDIOMEM DELIVERY CATHETER WAS SLIGHTLY BENT AND HAD TO BE EXCHANGED DURING THE PROCEDURE. IT IS BELIEVED THAT DURING THE EXCHANGE, THE CATHETER TIP FROM THE FIRST CARDIOMEM WAS UNKNOWINGLY RETAINED. THE SECOND CARDIOMEM CATHETER WAS INSERTED WITHOUT DIFFICULTY. SPECIFIC DEVICE INFORMATION UNKNOWN. | 0
1061444 | THE EKG CENTRAL MONITORS ON THE 4TH FLOOR NURSING UNIT WENT BLANK LATE ONE EVENING AND WERE NOT RESTORED TO OPERATION UNTIL THE EARLIER HOURS OF THE NEXT MORNING. | 1
897405 | DRAEGER MEDICAL SYSTEMS, INC. HAS RECEIVED INFORMATION FROM OUR LOCAL OFFICE INDICATING THAT AN INCIDENT OCCURRED INVOLVING AN ACCESSORY PRODUCT, (PATIENT MONITOR WALL MOUNTING BRACKET) THAT BROKE AND FELL TO THE FLOOR WITH A PATIENT MONITOR ATTACHED. NO PATIENT INJURY WAS REPORTED. WE HAVE REQUESTED THE RETURN OF THE CITED PATIENT MONITOR AND WALL MOUNTING BRACKET ACCESSORY FOR EVALUATION. | 0
981914 | USED A SAFETY SYRINGE 'VANISHPOINT' AUTOMATED RETRACTION TECHNOLOGY TO ADMINISTER HALOPERIDOL-DECANOATE INJECTION. DURING ADMINISTRATION OF THE MEDICATION, THE SYRINGE SEPARATED BETWEEN NEEDLE AND HUB, CAUSING THE MEDICATION TO SPILL OUT OVER CLIENT'S BODY. DATES OF USE: USED ONE TIME. DIAGNOSIS OR REASON FOR USE: SCHIZOPHRENIA. | 0
1053198 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED. SERVICE WAS NOT DISPATCHED SINCE THE PROVUE HAND HELD SCANNER WAS UNABLE TO READ THE SAMPLE BARCODES AND THE ANALYZER MISREAD THE BARCODE UPON SECOND ATTEMPT. THE ISSUE MAY HAVE BEEN RELATED TO THE SAMPLE BARCODES. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE READ A SAMPLE BARCODE ID INCORRECTLY. SAMPLE MISIDENTIFICATION MAY LEAD TO THE REPORTING OF ERRONEOUS TEST RESULTS. THE CUSTOMER IDENTIFIED THE ISSUE, PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. | 0
931350 | MICROBIAL CONTAMINATION OF THE ARCHITECT I2000/I2000SR SYSTEMS WASH BUFFER FLUIDICS PATHWAY CAN OCCUR DURING THE MFG OF THE INSTRUMENTS. THE ROOT CAUSE WAS IDENTIFIED AS THE USE OF CONTAMINATED PARTS FROM VENDORS AND/OR STORAGE AND SHIPPING OF INSTRUMENTS UNDER FAVORABLE MICROBIAL CONDITIONS, OR MICORBIALLY CONTAMINATED ARCHITECT SYS AUTOMATED RECONSTRUCTION MODULE (ARM) TO PREPARE WASH BUFFER. THE CURRENT ANTIMICROBIAL AGENT IN THE WAS BUFFER IS UNABLE TO PREVENT GROWTH OF MICROORGANISMS TO HIGH LEVELS. ADDITIONALLY, THE CURRENT DECONTAMINATION PROCEDURE FOR THE WASH BUFFER FLUIDICS PATHWAY REDUCES THE MICROBIAL CONTAMINATION BUT DOES NOT COMPLETELY ELIMINATE IT. CERTAIN MICROORGANISMS THAT CONTAMINATE THE WASH BUFFER (LINE DILUTENT) FLUIDICS PATHWAY ON THE ARCHITECT I2000/I2000SR SYS CAN PRODUCE FOLIC ACID, OR STRUCTURALLY SIMILAR METABOLITE. THIS MICROBIAL FOLIC ACID IS THEN ADDED TO THE REACTION MIXTURE, AS THE WASH BUFFER IS USED TO DILUTE BOTH SAMPLE AND CONJUGATE AND TO WASH THE MICROPARTICLES IN THE FOLATE ASSAY. IF CONTAMINATION INCREASES TO A HIGH ENOUGH LEVEL (APPROX >25,000 COLONY FORMING UNITS [CFU'S]), THE ARCHITECT FOLATE ASSAY PERFORMANCE CAN BE AFFECTED BY THE MICROBIALLY PRODUCED FOLIC ACID IN THE WASH BUFFER. AS A RESULT OF THE INVESTIGATION, THE FOLLOWING CORRECTIVE ACTIONS WILL BE IMPLEMENTED: ARCHITECT I2000/I2000SR SYSTEM ARE NOW FLUSHED WITH DEIONIZED WATER FOLLOWED BY AIR WITHIN 2 HRS OF COMPLETING FUNCTIONAL TESTING. ADDITIONALLY, THE INSTRUMENTS ARE STORED AND SHIPPED WITH AIR IN THE LINES. PREVIOUSLY MFG ARCHITECT I2000 INSTRUMENT MFG IN DALLAS WERE STORED AND SHIPPED WITH A WASH BUFFER FILTER INSTALLED. THESE INSTRUMENTS ARE NOW STORED AND SHIPPED WITHOUT THE WASH BUFFER FILTER INSTALLED. UPON INSTRUMENT INSTALLATION, ABBOTT FIELD SVC NOW DECONTAMINATES THE ARCHITECT I2000 SYSTEM PER THE INTERNAL DECONTAMINATION PROCEDURES AND INSTALLS A NEW WASH BUFFER FILTER. THE WASH BUFFER TRANSFER PUMP, WHICH IS REC'D DIRECTLY FROM THE VENDOR, WAS FOUND TO BE BACTERIALLY CONTAMINATED. THE VENDOR HAS CHANGED THE MFG PROCEDURE TO INCLUDE FLUSHING THE PUMP WITH ISOPROPYL ALCOHOL PRIOR TO SHIPMENT TO ABBOTT. THE INVESTIGATION REVEALED THE ARCHITECT ARM MODULE RECEIVED DIRECTLY FROM THE VENDOR, WAS BACTERIALLY CONTAMINATED. ABBOTT IS WORKING WITH THE VENDOR TO ADDRESS THE MICROBIAL CONTAMINATION ISSUE. IMPROVEMENTS TO THE INSTRUMENT DECONTAMINATION PROCEDURE ARE IN PROCESS. THE IMPROVEMENTS WILL ENSURE THE DECONTAMINATION SOLUTION CONTACTS ALL PORTIONS OF THE BUFFER FLUIDICS PATHWAY. INTERNAL DECONTAMINATION PROCEDURES WERE REVISED TO CORRECT THE FLUID LOADING ISSUES THAT COULD SOMETIMES PREVENT THE PROCEDURE FROM EXECUTING PROPERLY. SODIUM AZIDE AND/OR OTHER PRESERVATIVES WILL BE ADDED TO THE WASH BUFFER TO IMPROVE CONTROL OF MICROBIAL GROWTH. THIS IS THE FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATED THAT AFTER DECONTAMINATING THE ANALYZER, THE ARCHITECT FOLATE CONTROLS ARE FIRST ACCEPTABLE AND THEN AFTER A WEEK GO BACK TO BEING OUT OF RANGE HIGH. THE CUSTOMER IS HAVING THE ISSUE ON BOTH ARCHITECT I2000 ANALYZERS. THE CUSTOMER WAS TRAINED TO PERFORM A COMPLETE DECONTAMINATION OF THE ANALYZERS. THERE WAS NO IMPACT TO PT MGMT REPORTED. | 0
970734 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED FOR THE ELECTRIC SHOCK. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THAT THE VOLUME VERIFICATION BOARD WAS DEFECTIVE. HOWEVER, IT IS UNKNOWN IF THE DEFECTIVE BOARD RESULTED IN THE SHOCK EXPERIENCED BY THE CUSTOMER. REPLACEMENT OF THE BOARD HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< AN OCD BLOOD BANK TECHNICAL SPECIALIST (BBTS) REPORTED ON BEHALF OF THE CUSTOMER THAT IN 2007, A TECH TOUCHED THE REAGENT CAROUSEL AND EXPERIENCED AN ELECTRIC SHOCK WHILE THE ORTHO PROVUE ANALYZER WAS NOT OPERATING. NO HARM WAS REPORTED AS A RESULT OF THIS INCIDENT. | 0
1082695 | DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. >< DURING USE OF THE FIRST SET THE CONNECTION RIGHT AFTER THE VAMP TO THE STOP COCK CAME LOOSE. ANOTHER SET WAS OPEN AND WHEN THEY TESTED IT BEFORE USE, IT WAS THE SAME PROBLEM. THE WELDING SEEMS TO BE UNSTABLE. | 0
989066 | PATIENT WAS IN NURSING HOME AND MONITORED USING POSEY 8361 SITTER SELECT ALARM AND POSEY 8290A OVER THE MATTRESS BED SENSOR. IN 2007, PATIENT WAS FOUND LYING ON THE FLOOR. HOSPITAL CLAIMS ALARM LIGHTS WERE BLINKING, BUT NEVER SOUNDED. PATIENT SUSTAINED CUTS AND FRACTURED FACIAL BONES. PATIENT WAS HOSPITALIZED AND RETURNED TO THE NURSING HOME. PREVIOUSLY, PATIENT HAD HISTORY OF DEMENTIA AND FALLS. IT WAS FOUND THAT THE ALARM USER RECORDED A CUSTOM VOICE MESSAGE IN PLACE OF THE DEFAULT VOICE MESSAGE, BUT PROBABLY INADVERTENTLY RECORDED A VERY SHORT LENGTH MESSAGE WITHOUT ANY CONTENT. AS A RESULT, WHEN THE ALARM WAS SET OFF, NO CUSTOM VOICE MESSAGE WAS HEARD, AND THE DEFAULT FACTORY MESSAGE DID NOT PLAY. USER CLAIMS TO HAVE TESTED ALARM BEFORE LEAVING PATIENT UNATTENDED, AS INSTRUCTED IN THE PRODUCT INSTRUCTIONS AND LABELS ON BOTH THE ALARM AND SENSOR; HOWEVER, THIS TESTING SHOULD HAVE BROUGHT THE EMPTY MESSAGE TO THE USER'S ATTENTION. ALARM AND SENSOR WERE RETURNED TO POSEY AND INSPECTION SHOWED BOTH PRODUCTS WORK ACCORDING TO SPECIFICATIONS. ALARM WAS MANUFACTURED EARLY MARCH 2005. IN FEBRUARY 2006, POSEY UPDATED THE ALARM FIRMWARE SO THAT IN CASES OF AN EMPTY CUSTOM VOICE MESSAGE, THE DEFAULT ALARM MESSAGE AND TONE WOULD AUTOMATICALLY PLAY. | 1
1076516 | THE ISSUE WAS DETERMINED TO BE CAUSED BY A CODE ERROR. THE RANGE FLAG FOR NUMERIC RESULTS WERE INCORRECTLY OVERWRITTEN. >< IT WAS REPORTED THAT LOW REFERENCE RANGE FLAGS WERE NOT REPORTED AS EXPECTED IN HL7 MESSAGES. THIS ISSUE WAS DISCOVERED IN THE CUSTOMER'S TEST ENVIRONMENT. | 1
1064729 | SOMETIMES IN 2008 AT 3:45 AM THE MICU CENTRAL MONITORING SYSTEM WENT DOWN. 'NO BED SIGNAL' APPEARED ON THE MONITOR SCREEN. NURSING STAFF HAD ATTEMPTED TO REBOOT THE SYSTEM WITHOUT SUCCESS. PATIENTS WERE MONITORED USING THE BEDSIDE SYSTEM WITH ALARMS ON AND AUDIBLE BY STAFF. THE AOD WAS REQUESTED TO CALL BIOMED ENGINEERING TO REPAIR THE MONITORS. AT 4:30 AM BIOMED ON-CALL STAFF SPOKE WITH ICU STAFF TO DETERMINE IF A REBOOT WAS ATTEMPTED. BIOMED THEN PLACED A CALL TO THE CONTRACT SERVICE REPRESENTATIVE FOR ASSISTANCE. AT 5:30 AM THE SERVICE REPRESENTATIVE REQUESTED THE ICU STAFF TO REBOOT THE SYSTEM AGAIN WHICH WAS UNSUCCESSFUL. THE REPRESENTATIVE REMINDED THE STAFF TO CHECK THE BEDSIDE ALARMS AND ESTIMATED TIME OF ARRIVAL WAS 1 1/2 TO 2 HOURS. THE REPRESENTATIVE ARRIVED AT 6:45 AM TO BEGIN WORK ON THE CENTRAL MONITOR SYSTEM. REPAIRS WERE COMPLETED AT 11:45 AM. AT 7:10 AM A NURSE WENT INTO A PATIENT'S ROOM AND CALLED OUT, 'GET THE CRASH CART, CALL A CODE.' A PT WITH SIGNIFICANT CHF, CAD, PVOD, DM AND RENAL INSUFFICIENCY WAS FOUND IN VENTRICULAR FIBRILLATION. NO ATTEMPTS AT RESUSCITATION WAS MADE AS THE PATIENT WAS A DNR/DNI. THE PATIENT WAS PRONOUNCED DEAD AT 7:30 AM. THE BEDSIDE ALARM DID NOT TRIGGER A RED ALARM WHEN THE PATIENT WENT INTO VENTRICULAR FIBRILLATION. THE MONITOR WAS CHECKED AND THE ALARM VOLUME WAS SET AS ZERO. ALL BEDSIDE ALARMS WERE CHECKED AND ALL THE ALARM VOLUMES WERE READING ZERO. ALL ALARM VOLUMES WERE TURNED UP WITHOUT DIFFICULTY. UPDATE: ABOUT 3 DAYS LATER, ONE OF THE QUESTIONS WE ARE TRYING TO DETERMINE IS WHETHER THE BEDSIDE MONITOR VOLUMES RESET TO ZERO OR HAD NEVER BEEN RAISED INITIALLY WHEN THE CENTRAL STATION WENT DOWN. AT THIS TIME, THE INFORMATION WE HAVE FROM PATIENT CARE SERVICES IS THAT THE VOLUME CONTROL AT THE BEDSIDE MONITORS APPEARED TO HAVE SET THEMSELVES TO ZERO AT THE TIME OF THE REBOOT. WE HAVE SPOKEN WITH BIOMEDICAL ENGINEERING AND WE HAVE VERIFIED THAT THE CENTRAL STATION SHOULD NOT/DOES NOT CONTROL THE VOLUMES AT THE BEDSIDE, SO WE DO NOT HAVE A CLEAR REASON FOR THE VOLUMES TO HAVE RESET TO ZERO. A CALL HAS BEEN PLACED TO PHILLIPS AND WE ARE AWAITING A RESPONSE FROM THEM ON THIS ISSUE. IT IS NOT AN EXPECTATION THAT THE BEDSIDE MONITOR VOLUMES BE SET AT A CERTAIN NUMBER AS WE HAVE THE CENTRAL STATION LOCKED, SO THAT THE VOLUME THERE CAN NOT BE CHANGED. WE ARE WRITING A PROCEDURE FOR SAFETY CHECKS OF THE BEDSIDE MONITORS FOR EVERY DOWNTIME OF THE CENTRAL STATION AND/OR REBOOT OF THE CENTRAL STATION SO THAT WE CAN ASSURE THAT THE VOLUMES ARE SET AS APPROPRIATE WHEN NEEDED. WE EXPECT THE USE OF THE CHECKLIST TO PREVENT THIS TYPE OF EVENT IN THE FUTURE. THIS SHOULD BE COMPLETED BY JUNE 20, 2008. IN THE MEANTIME A DEBRIEFING HAS OCCURRED ON THE UNIT AND STAFF ARE AWARE OF THE NECESSITY TO CHECK BEDSIDE ALARM VOLUMES. REGARDING MAINTENANCE OF EQUIPMENT: ON AN EARLIER REPAIR VISIT, JUNE 7, THE PHILLIPS TECHNICIAN FOUND THE SCC - NETWORK INTERFACE CABINET - TO BE VERY DUSTY AND NOTED IT WAS LIKELY THE CAUSE OF THE MALFUNCTION - NOT DEFINITIVELY - ON JUNE 7, 2007. THE SCC DOES NOT HAVE A CIRCULATION FAN OR FILTERS AND THUS DUSTING AND CLEANING IS NOT ON THE PM RECOMMENDATIONS. PHILLIPS IS OBVIOUSLY NOW AWARE OF THIS AND IS STARTING A PM CYCLE THROUGH THE ICU'S AND TELEMETRY UNIT THIS WEEK. WE WILL FOLLOW-UP TO ASSURE THIS IS KEPT UP. THE ISSUE OF BIOMEDICAL ENGINEER ON CALL NOT COMING TO THE FACILITY HAS BEEN ADDRESSED WITH THE APPROPRIATE STAFF MEMBERS. IN THE FUTURE, THE ON-CALL STAFF WILL COME TO THE FACILITY, ASSESS THE PROBLEM AND THEN NOTIFY PHILLIPS WHEN NEEDED. | 1
1023404 | A FOLLOW-UP REPORT WILL BE SUBMITTED IF THE PRODUCT IS RETURNED AND INVESTIGATION RESULTS ARE AVAILABLE. >< CUSTOMER REPORTS USING THE INCORRECT CALIBRATION CODE ON HER FREESTYLE FLASH BLOOD GLUCOSE METER SINCE 2007 AND 'NO ONE CAUGHT THAT THE CODE WAS WRONG' UNTIL THE DATE OF THE EVENT. SHE REPORTS HER DR HAS CHANGED HER MEDICATION BASED ON HER 'INCORRECT READINGS', HOWEVER, NO READINGS ARE REPORTED. SHE REPORTS EXPERIENCING SYMPTOMS OF NAUSEA AND VOMITING DUE TO THE MEDICATION CHANGE. PARAMEDICS WERE CALLED AND TREATED HER WITH 'IV MEDICATION' (EXACT SOLUTION NOT REPORTED). SHE WAS TRANSPORTED TO A LOCAL HEALTH CARE FACILITY, HOWEVER, SHE 'DOES NOT KNOW' IF SHE WAS DIAGNOSED WITH HYPO/HYPERGLYCEMIA AND 'DOES NOT RECALL' THE EXACT MEDICATION GIVEN STATING SHE WAS TREATED WITH 'MANY DIFFERENT DRINKS TO HER STOMACH PAINS'. THERE WAS NO REPORT OF DEATH OR PERMANENT IMPAIRMENT ASSOCIATED WITH THIS CASE. | 0
1157864 | I PURCHASED THIS BP MONITOR TO SELF-MONITOR BLOOD PRESSURE. I RETURNED IT. THEY REFUSED TO REPAIR OR REPLACE. MONITOR DOES NOT GIVE ACCURATE READINGS. IT IS THEREFORE USELESS IN DETERMINING MY NEED FOR BLOOD PRESSURE MEDICATION. | 0
965559 | EVAL ANTICIPATED BUT NOT BEGUN. >< DURING PREPARATION FOR USE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE CENTRAL CONTROL MONITOR OF THE HEART LUNG CONSOLE DID NOT FUNCTION. CONTROL OF PUMPS AND SAFETY SENSORS REMAINED AVAILABLE THROUGH REDUNDANT LOCAL CONTROLS. THE PUMP CONSOLE WAS REPLACED. THERE WAS NO ADVERSE CONSEQUENCE TO THE PT AS A RESULT OF THIS EVENT. | 1
1082613 | DURING AN EXHIBITION OF THE DEVICE, THE TOUCH PANEL OF THE CENTRAL CONTROL MONITOR OF THE HEART LUNG CONSOLE DID NOT FUNCTION AS EXPECTED. CONTROL OF PUMPS AND SAFETY SENSORS REMAINED AVAILABLE THROUGH REDUNDANT LOCAL CONTROLS. THERE WAS NO REPORTED ADVERSE CONSEQUENCE TO A PATIENT AS A RESULT OF THIS EVENT. >< EVALUATION IN PROGRESS, BUT NOT CONCLUDED. | 1
1027007 | QUESTIONED PT. RHYTHM WHEN TELEMETRY DID NOT MATCH PT. PALPATED RATE. VERIFIED TELEMETRY BOX NUMBER #57 ON BOX, AND #57 ON MONITOR SCREEN. EKG DONE TO VERIFY RHYTHM, AND EKG DID NOT MATCH TELEMETRY RHYTHM. IT WAS DISCOVERED THAT THE TELEMETRY BOX # IS MARKED ON THE REMOVABLE BATTERY COVER OF THE TELEMETRY BOX, AND APPARENTLY DURING CLEANING IN THE TELEMETRY ROOM, THE TELEMETRY #57 BATTERY COVER WAS PLACED ON THE TELEMETRY #67 BOX. THIS HAD THE POTENTIAL TO BE A SERIOUS EVENT WITH NEGATIVE PT. OUTCOME. | 1
1221083 | CARDIAC TELEMETRY BOX WAS SOUNDING OFF ALARM WITH WRONG READING. | 1
1084318 | THE ICU CENTRAL MONITOR STOPPED WORKING. THE SCREENS WERE BLANK. | 1
1027618 | THE CUSTOMER STATES THAT A PT SAMPLE TESTED USING A CELL-DYN 1800 ANALYZER GENERATED A LOW HEMOGLOBIN RESULT OF 7.5G/DL. THE SAMPLE WAS SENT TO A REFERENCE LAB OBTAINING A HIGHER RESULT OF 12.6 G/DL (TEST METHOD NOT PROVIDED). THERE WAS NO ADVERSE IMPACT TO PT MANAGEMENT DUE TO THIS ISSUE. SERVICE WAS DISPATCHED TO TROUBLESHOOT THE ISSUE. >< INVESTIGATION SUMMARY: ISSUE: DISCREPANT RESULTS FOR HEMOGLOBIN FOR ONE PT SAMPLE. THE INSTRUMENT IN USE WAS A CELL-DYN 1800 ANALYZER. DAILY QC WAS IN RANGE. THE CUSTOMER WAS UNABLE TO SAY IF THE SAMPLE WAS CLOTTED. THE CUSTOMER CONTACTED A FIELD SERVICE REP (FSR) WHO REQUESTED A FIELD REPAIR ACTIVITY BE OPENED. NO ADD'L DATA WAS AVAIL. THE FSR VISITED 2008, AND VERIFIED THE ISSUE (READING OF HGB FLOW CELL WAS OUT OF RANGE). THE HGB FLOWCELL WAS CLEANED AND ASSOCIATED VERIFICATION PROCEDURES WERE DONE. THE INSTRUMENT WAS PERFORMING TO SPECIFICATION FOR PRECISION AND QC RUNS. THE CELL-DYN SYSTEM 1800 OPERATOR'S MANUAL, 07H80-01, REV D: THE VALUE OF 7.5 G/DL SHOULD HAVE GENERATED A DISPERSIONAL DATA ALERT AS IT IS OUT OF THE NORMAL ADULT RANGE (APPENDIX B, PAGE B-3). PT LIMITS SETS AND PANIC LIMITS CAN BE SET BY THE OPERATOR TO SUIT THE NORMAL RANGES FOUND IN THE POPULATION SERVED BY THE LAB. THE OPERATOR MANUAL RECOMMENDS: THAT ONE PT LIMIT SET OR THE PANIC LIMITS BE USED TO ENTER INSTRUMENT-SPECIFIC LAB ACTION LIMITS. A RESULT THAT FALLS OUTSIDE A LAB ACTION LIMIT CAN ALSO INDICATE THE NEED FOR THE OPERATOR TO FOLLOW A LAB PROTOCOL, SUCH AS REPEATING THE SAMPLE, PERFORMING A SMEAR REVIEW OR NOTIFYING THE PHYSICIAN (P 3-25), INACCURATE HGB READINGS MAY BE CAUSED BY CONTAMINATION OF THE HGB FLOW CELL WITH ORGANIC RESIDUE. THE CORRECTIVE ACTION IS TO CLEAN THE FLOW CELL (PAGE. 10-36). CLEANING THE FLOW CELL IS AN 'AS-REQUIRED' MAINTENANCE ACTIVITY (PAGES 9-35 TO 9-38). TRENDING: A REVIEW OF COMPLAINT REPORTS, FOR THE PERIOD 2007 THROUGH 2008, DID NOT INDICATE ANY ADVERSE TREND WITH THE CELL-DYN 1800, LIST BASE 07H77-01, FOR ISSUES WITH DISCREPANT RESULTS ON MULTIPLE PARAMETERS. LABELING: THE EVENT IS ADDRESSED IN THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL 07H80-01, REV D SECTION 3: PRINCIPLES OF OPERATION: PARAMETER DATA FLAGS: DISPERSIONAL DATA ALERTS: P 3-25 SECTION 5: OPERATING INSTRUCTIONS: SETUP INSTRUCTIONS: PT LIMIT SETS: P 5-18 TO 5-21, SPECIMEN ANALYSIS: P 5-51, SPECIMEN STABILITY & SPECIMEN COLLECTION: 5-52 SECTION 9: SERVICE AND MAINTENANCE: AS - REQUIRED MAINTENANCE: CLEANING THE HGB (HEMOGLOBIN) FLOW CELL: 9-35 TO 9-38, SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS; INDEX OF ERROR MESSAGES AND CONDITIONS: DATA PROBLEMS; INACCURABLE HGB READINGS: 10-36 APPENDIX B: REFERENCE TABLES; TABLE B-2: REFERENCE INTERVALS (NORMAL VALUES) FOR AUTOMATED BLOOD COUNTERS; B-3. CONCLUSION: THE DEVICE INVOLVED IN THE ISSUE WAS EVALUATED. SERVICE DISCOVERED THAT THE FLOW CELL READING WAS OUT OF RANGE AND THIS PART WAS REPLACED. THE ANALYZER WAS CLEANED AND ASSOCIATED VERIFICATION PROCEDURES WERE PERFORMED WITH ACCEPTABLE RESULTS. THERE WAS NO IMPACT TO PT MANAGEMENT REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. | 0
1068287 | A SERVICE CALL WAS MADE. THE CENTRIFUGE WAS REPLACED. AFTERWARDS, VERIFICATION TESTS WERE PERFORMED; THE UNIT OPERATED WITHIN SPECIFICATION. >< CUSTOMER REPORTED AN ABO DISCREPANCY WITH A DONOR UNIT. TESTED ON GALILEO THE FWD_ABORH TYPE WAS AB-POSITIVE; THE DONOR UNIT WAS TYPED A-POSITIVE. | 0
1020044 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED. | 0
1001416 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT STATED THE COMMANDER FPC DILUENT SYRINGE WAS BROKEN WHEN A NEW BOX WAS OPENED. ANOTHER FPC DILUENT SYRINGE WAS PLACED ON THE FPC AND BROKE UPON FIRST USE. BOTH SYRINGES ARE BROKEN AT THE SAME PLACE WHERE THE SILVER METAL PART ON TOP OF THE SYRINGE MEETS THE GLASS. NO INJURY DUE TO THE BROKEN SYRINGES WAS REPORTED. | 0
1011565 | CUSTOMER RETURNED PRECISION XTRA BLOOD GLUCOSE METER. METER'S DATE AND TIME WERE SET, METER WAS PROPERLY CALIBRATED, AND METER WAS THEN SET ASIDE FOR APPROX 12 HOURS. SETTINGS WERE THEN RE-CHECKED, AND THE METER PROPERLY RETAINED ALL SETTINGS. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED. >< A CUSTOMER REPORTED THAT THEIR PRECISION XTRA BLOOD GLUCOSE METER KEPT LOSING DATE AND TIME. THIS ISSUE WAS REPORTED TO HAVE AFFECTED THEIR PRECISION LINK DATE MGMT SYSTEM, WHICH WAS ALSO SHOWING DATE AND TIME ISSUES. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
949751 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< INTERNAL TESTING BY ABBOTT IDENTIFIED TWO SOFTWARE DEFECTS AFFECTING THE USE OF SAMPLE IDENTIFICATION (SID) OR PATIENT IDENTIFICATION (PID) WITH THE ACCELERATOR DECISION MANAGER (ADM). ABBOTT DETERMINED THAT SID OR PID NUMBERS THAT CONTAIN MORE THAN 12 CHARACTERS ARE TRUNCATED (SHORTENED ) TO THE FINAL 12 CHARACTERS BY ADM. A SURVEY OF EXISTING ADM USERS INDICATED THAT NO LABORATORIES ARE USING MORE THAN 12 CHARACTERS FOR SID/PID. THE TRUNCATION OF SID OR PID MAY RESULT IN TEST ORDERS AND RESULTS MATCHED TO THE INCORRECT SAMPLE OR WHEN A MATCHING SID TO THE TRUNCATED SID DOES NOT EXIST, A NEGATIVE QUERY MAY OCCUR. ABBOTT HAS NOT RECEIVED ANY REPORTS OF ADVERSE EVENTS RELATED TO THIS ISSUE. | 1
1015800 | CUSTOMER REPORTED THAT A DASH MONITOR ALLEGEDLY FAILED TO ALARM FOR ASYSTOLE. NO PT DEATH OR INJURY WAS REPORTED. THE DOCTOR WAS AT THE BEDSIDE TO WITNESS THE EVENT. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1021228 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1021236 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1021202 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1021211 | IT WAS REPORTED THAT 32 EAGLE MONITORS SHUT DOWN FROM 12:00 TO 16:30. NO PT DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
900547 | DRAEGER MEDICAL SYSTEMS HAS RECEIVED INFORMATION FROM A CUSTOMER THAT OUR SC7000 PATIENT MONITOR MALFUNCTIONED WHEN PERFORMING A PATIENT THERMO MONITORING PROCESS USING OUR TEMP 'Y' CABLE. THIS FUNCTION PROVIDES THE CAPABILITY OF MEASURING A PATIENT'S CORE AND SURFACE BODY TEMPERATURE AND DISPLAYING THE VALUES ON THE PATIENT MONITOR. IT WAS REPORTED THAT TEMPERATURE (TEMP B) READOUT WAS NOT VISIBLE ON THE MONITOR, BUT THE PROBLEM WAS RESOLVED WHEN THE MONITOR WAS RESTARTED. THE SECOND ISSUE INVOLVED A 'FROZEN' TEMPERATURE DISPLAY, (TEMP B) VALUE WHICH COULD BE CORRECTED BY RESTARTING THE PATIENT MONITOR, BUT IT WAS INDICATED THAT THE TEMPERATURE DISPLAY WAS THEN SPORADIC. NO PATIENT INJURY WAS REPORTED. | 1
981941 | CUSTOMER CLAIMED THAT TELEMON DID NOT CAPTURE HEART RHYTHM WHILE BEING MONITORED BY A TELEMETRY PAK. >< THE PT WAS ON A DIALYSIS UNIT BEING MONITORED ON THE TELEMON. WHEN A LEADS OFF CONDITION OCCURRED, THE MONITOR TECHNICIAN CALLED THE DIALYSIS UNIT TO INFORM THEM OF THE SITUATION. THE PTS LEADS WERE NOT RE-APPLIED FOR SOME TIME AND WHEN THEY WERE, THE EXTREME BRADYCARDIA WAS NOTED. THE PT WAS CODED AND ULTIMATELY EXPIRED. | 1
975379 | CONTRARY TO THE UFR, IT WAS REPORTED TO THE MFR, THAT THE CRNA (HOSPITAL STAFF) WAS UNABLE TO CALIBRATE THE FLOW SENSOR DURING THE PRE-USE CHECK OF THE DEVICE. THE BIOMED DEPT WAS INFORMED ABOUT THE DEVICE PROBLEM, BUT THE OPERATOR/CRNA DID NOT WAIT FOR FIXING THE FLOW SENSOR PROBLEM AND CONTINUED THE PROCEDURE, INTUBATED THE PT AND STARTED THE CASE. IT WAS FURTHER REPORTED, THAT THE DEVICE SEEMED TO BE FUNCTIONING, BUT IT WAS NOTED THAT THE TIDAL VOLUME READINGS WAS 200 ML/MIN EVEN THOUGH THE SET VALUE WAS 700 ML/MIN. ALL OTHER INFO IS IN CORRESPONDENCE TO THE EVENT DESCRIPTION IN THE UFR. >< THE REPORTED EVENT AND CONDITIONS COULD NOT BE REPRODUCED. THE CHEMICAL/PHYSICAL ANALYSIS OF THE AFFECTED/BURNED PARTS AND ANALYTICAL INVESTIGATION OF THE EVENT DID NOT REVEAL A CLEAR ROOT CAUSE OF THE REPORTED DEFLAGRATION IN THE BREATHING SYSTEM. WITHIN THE OVERALL PRODUCT SURVEILLANCE OF THE FABIUS DEVICE FAMILY, THIS IS THE FIRST REPORTED CASE OF THIS NATURE. THE MFR COMES TO THE CONCLUSION THAT THERE IS NO INDICATION WHILE USING THE FABIUS GS ACCORDING TO THE INSTRUCTIONS FOR USE (FABIUS GS) WITH APPROVED BREATHING GASES AND VOLATILE ANESTHETIC AGENTS A FLAMMABLE GAS MIXTURE COULD OCCUR. SAFETY PRECAUTIONS AND CLEANING RECOMMENDATIONS ARE REFERRED IN THE INSTRUCTIONS FOR USE (OPERATOR'S MANUAL). THEREFORE, IT MUST BE ASSUMED THAT ALIEN SUBSTANCES LIKE CLEANING OR DISINFECTANTS ON ALCOHOL BASE OR MEDICINAL NEBULIZER OR AN AEROSOL WERE PRESENT IN THE BREATHING SYSTEM AND LED TO THE REPORTED DEFLAGRATION. THESE ASPECTS WERE QUESTIONED AND DISCUSSED WITH THE CUSTOMER/USER, BUT WERE EXCLUDED BY HOSPITAL STAFF. THE EVENT HAS BEEN FINALLY ASSESSED TO BE AN ISOLATED CASE WITH UNKNOWN ROOT CAUSE. | 1
1192585 | CUSTOMER REPORTED A SAMPLE BARCODE WAS MISREAD ON GALILEO. DURING VALIDATION TESTING. >< THE DIAGNOSTIC ARCHIVE WAS NOT ABLE TO BE RETRIEVED FROM THE INSTRUMENT, SINCE IT WAS DELETED AFTER 30 DAYS. A SERVICE CALL WAS MADE. THE RAM TEST WAS PERFORMED 3 TIMES ON EACH OF THE CUSTOMER'S GALILEOS, AND PASSED EACH TIME. THE SYSTEM OPERATED WITHIN SPECIFICATIONS WITH NO PROBLEMS OBSERVED. | 0
978482 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED A SAMPLE FROM A POSTPARTUM PT WAS RECEIVED FROM A HEALTHCARE CENTER AND GENERATED REPEAT REACTIVE RESULTS ON THE AXSYM AG/AB COMBO ASSAY. THE SAMPLE WAS SENT TO ANOTHER LAB FOR CONFIRMATION. THE SAMPLE WAS REACTIVE FOR THE FOLLOWING COMBO ASSAYS: UNIFORM. MUREX, AXSYM AND MUREX AG MAB (AG P24). THE PT WAS INSTRUCTED TO SUSPEND LACTATION OF HER BABY. A SECOND SAMPLE OBTAINED FORM THIS PT GENERATED NONREACTIVE RESULTS ON AXSYM AG/AB COMBO. THE FIRST SAMPLE WAS THOUGHT TO BE CONTAMINATED. NO INJURY WAS REPORTED. | 0
1059523 | IT WAS REPORTED THAT THERE WAS A LINE DETACHMENT. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. | 0
1068293 | POSSIBLE OVERCORRECTION DURING VISION CORRECTION PROCEDURE. UPDATE OF SOFTWARE ON VISX EXCIMER LASER MACHINE WAS DONE IN MID-MAY. PTS WHO UNDERWENT CORRECTION OF LOW-CORRECTION VISION --1 OR LESS- APPEAR TO HAVE BEEN OVER-CORRECTED -TO +2- DUE TO SOFTWARE UPDATE. HIGHER-CORRECTION VISION NOT IMPACTED BY THIS PROBLEM. SEVEN PATIENTS IDENTIFIED WHO MAY HAVE EXPERIENCED THIS OVER-CORRECTION, AND MIGHT REQUIRE RECORRECTION. MACHINE HAS BEEN TAKEN OUT OF SERVICE UNTIL FURTHER NOTICE AND COMPANY -AMO/ADVANCED MEDICAL OPTICS- TECHNICIANS WORKING ON PROBLEM. DATES OF USE: 2008. DIAGNOSIS OR REASON FOR USE: CORRECTIVE VISION PROCEDURE. | 1
1007846 | THE ACCOUNT STATED THE COMMANDER PPC LID GAS HINGE WAS NOT WORKING PROPERLY. THE ACCOUNT REQUESTED SERVICE FOR THE COMMANDER PPC ANALYZER. THERE WAS NO INJURY REPORTED DUE TO THE COMMANDER PPC LID GAS HINGE NOT WORKING PROPERLY. >< RESULTS- OTHER: GAS SPRING WHICH HOLDS THE TOP LEFT COVER UP. THE COMPLAINT INDICATED THE FIELD SERVICE REP (FSR) ORDERED THE PART FOR INSTALLATION. THE COMPLAINT INDICATES NO ONE HAD BEEN INJURED DUE TO THE COVER NOT REMAINING OPEN. THE FSR REPLACED THE PART AND VERIFIED. IT WAS FUNCTIONING AS REQUIRED. A REVIEW OF THE COMPLAINT HISTORY FOR PPC WAS PERFORMED FOR THE TIME PERIOD EARLY 2006 THROUGH 2008. NO OTHER COMPLAINTS WERE FOUND RELATED TO THIS OBSERVATION. REFERENCE WHERE IN LABELING THE EVENT IS ADDRESSED: ALTHOUGH LABELING IS NOT AVAILABLE RELATED TO THE CUSTOMER'S SPECIFIC OBSERVATION, THE FOLLOWING INFO IS PROVIDED: THE COMMANDER PARALLEL PROCESSING CENTER (PPC) OPERATIONS MANUAL (69-5745/R3 DECEMBER 2001) STATES THE FOLLOWING RELATED TO THE CUSTOMER OBSERVATION: SECTION 8 HAZARDS, PHYSICAL AND MECHANICAL SAFETY. BASIC RULES OF MECHANICAL OPERATION ARE ESSENTIAL TO SAFE OPERATION OF THE PPC. OPERATORS SHOULD PRACTICE SOUND MECHANICAL SAFETY HABITS THAT INCLUDE BUT ARE NOT LIMITED TO THE FOLLOWING: KEEP ALL PROTECTIVE COVERS IN PLACE WHEN PROCESSING SPECIMENS. VERIFY THAT ALL COVERS ARE SECURELY OPEN TO PREVENT INADVERTENT CLOSURE. USE CARE WHEN CLOSING COVERS SO AS NOT TO PINCH HANDS OR FINGERS. PER THE INSTRUMENT SERVICE ADVISORY (ISA) 50-076, THE FSR IS REQUIRED TO CHECK THE COVER ON ALL PREVENTATIVE MAINTENANCE (PM). THE FOLLOWING PROCEDURE IS RECOMMENDED FOR CHECKING THE INTEGRITY OF THE PPC GAS SPRING, FOR THE TOP LEFT COVER. SLOWLY LIFT THE TOP LEFT COVER UNTIL RESISTANCE IS ENCOUNTERED APPROX 30.5 CENTIMETERS +/- 2.5 CENTIMETERS / 12 INCHES +/- 1 INCH. CONTINUE TO LIFT THE COVER UNDER RESISTANCE ANOTHER 2.5 - 5 CENTIMETERS / 1 TO 2 INCHES. RELEASE AND OBSERVE. REPLACE THE GAS SPRING IF COVER BEGINS TO CLOSE. THIS IS A FINAL REPORT. END OF REPORT. | 0
1001272 | RESULT (GAS SPRING). THE ISSUE WAS RESOLVED BY REPLACING THE GAS SPRING. ABBOTT FIELD SVC ENGINEERS ROUTINELY CHECK THE GAS SPRINGS ON THE TOP LEFT COVER DURING PREVENTATIVE MAINTENANCE SVC CALLS. A REVIEW OF THE COMPLAINT HISTORY FOR PPC FOR THE TIME PERIOD OF 2006 THROUGH 2008 DID NOT IDENTIFY ANY OTHER COMPLAINTS RELATED TO THIS ISSUE. THIS IS A FINAL REPORT. >< THE CUSTOMER STATED THAT THE GAS SPRING FOR THE TOP LEFT COVER OF THE COMMANDER PARALLEL PROCESSING CENTER (PPC) IS NOT ENGAGING AS EXPECTED AND THE LID DOES NOT STAY OPEN. THERE WAS NO REPORT OF INJURY. | 0
1001274 | 'OTHER' (GAS SPRING). THE ISSUE WAS RESOLVED BY REPLACING THE GAS SPRING. ABBOTT FIELD SVC ENGINEERS ROUTINELY CHECK THE GAS SPRINGS ON THE TOP LEFT COVER DURING PREVENTATIVE MAINTENANCE SVC CALLS. A REVIEW OF THE COMPLAINT HISTORY FOR PPC FOR THE TIME PERIOD OF AUGUST 01, 2008 THROUGH FEBRUARY 18, 2008 DID NOT IDENTIFY ANY OTHER COMPLAINTS RELATED TO THIS ISSUE. THIS IS A FINAL REPORT. >< THE CUSTOMER STATED THAT THE GAS SPRING FOR THE TOP LEFT COVER OF THE COMMANDER PARALLEL PROCESSING CENTER (PPC) IS WORN OUT AND WILL NOT HOLD THE COVER OPEN. NO INJURIES WERE REPORTED. | 0
1233365 | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR X-RAY MFR TO SUBMIT THIS INCIDENT THAT OCCURRED IN ANOTHER COUNTRY. THE PROBLEM: THE PT WAS CONNECT TO THIS PHYSIOLOGICAL MONITOR. THEIR INVASIVE BLOOD PRESSURE (ON CHANNEL 1 AND CHANNEL 3) WAS DROPPING IN VALUE DURING THIS EXAMINATION. THEN THE PT DIED WHILE STILL CONNECTED TO THIS DEVICE. | 1
1174283 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). THE INITIAL REPORT WAS SUBMITTED UNDER BRAND NAME APS INPUT OUTPUT MODULE, (B)(4) MANUFACTURING SITE. IT WAS DETERMINED THAT THE INVESTIGATION SHOULD BE PERFORMED ON BRAND NAME ACCELERATOR DECISION MANAGER, (B)(4) MANUFACTURING SITE. THE DOCUMENTATION FOR THIS ISSUE FOR THE ACCELERATOR DECISION MANAGER (ADM) INDICATES THAT THE CUSTOMER IS RUNNING THE ACCELERATOR AUTOMATED PROCESSING SYSTEMS (APS) WORKCELL WITH DECISION MANGER (ADM) MIDDLEWARE. AN ARCHITECT C16000 SYSTEM CONNECTED TO THE APS WORKCELL REPORTED A HIGH MICRO ALBUMIN (>500 MG/L) FOR SAMPLE IDENTIFICATION (B)(4). THE CREATININE ASSAY FOR THIS SAMPLE HAD INITIAL RESULT AS <400 UMOL/ L. A SAMPLE PRESENTATION ERROR WAS GENERATED FOR THIS SAMPLE AND A SAMPLE QUEUE ERROR WAS GENERATED FOR THE NEXT SAMPLE ((B)(4)). THE SAMPLE PRESENTATION AND SAMPLE QUEUE ERRORS WERE BOTH DOCUMENTED AT THE APS, HOWEVER THE RESULTS RECEIVED BY THE ADM DID NOT INCLUDE ERROR FLAGS FOR THESE ERRORS AS EXPECTED. THE REVIEW OF THE ADM LOGS DETERMINED THAT THE CONFIGURATION OF ADM WAS BASED ON THE PREVIOUS VERSION OF INSTRUMENT SERVICE ADVISORY (B)(4), WHICH WAS BASED ON APS SOFTWARE V1.2. WITH THE UPGRADE TO APS SOFTWARE V1.3.1, THE SOFTWARE WAS CHANGED AND ADM SHOULD HAVE BEEN CONFIGURED ACCORDING TO (B)(4). AS A RESULT OF THE CONFIGURATION NOT BEING UPDATED IN ADM, THE MIDDLEWARE WAS UNABLE TO CAPTURE THE NOTIFICATION MESSAGES AND THEREFORE DID NOT FLAG THE TEST RESULTS FOR THE AFFECTED SAMPLES. THE ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS INSTRUCTED TO COMPLETE (B)(4) AT THIS ACCOUNT. AFTER THIS ISA WAS COMPLETED, ADM WAS STILL NOT PROPERLY CAPTURING THE NOTIFICATION MESSAGES FROM THE APS FOR SAMPLE PRESENTATION/ SAMPLE QUEUE ERRORS. IT WAS DETERMINED THAT THE CUSTOMER USES THE SAMPLE TYPE AS PART OF THE BARCODE IDENTIFICATION (ID). FOR EXAMPLE, IF THE REQUEST ID IS (B)(6), THEN SAMPLE ID (B)(6) IS FOR A SERUM SAMPLE AND (B)(6) IS FOR A URINE SAMPLE. ADM IS CONFIGURED TO ASSOCIATE SAMPLE PRESENTATION/SAMPLE QUEUE ERRORS BASED ON A SID TO THE INTERNAL ADM REQUEST ID. THE DEFAULT CONFIGURATION FOR ADM HAS THE SID EQUIVALENT TO THE REQUEST ID, BUT THIS IS NOT THE CASE WITH THIS CUSTOMER. BASED ON THE AVAILABLE INFORMATION, IT WAS DETERMINED THAT FOR (B)(4) (THE SAMPLE TUBE WITH SAMPLE PRESENTATION ERROR), THE URINE CREATININE WAS REPEATED AND THIS RESULT OF (B)(6) WAS RELEASED TO THE LABORATORY INFORMATION SYSTEM (LIS). THIS URINE CREATININE FOR THIS SAMPLE WAS THEN REPEATED THE FOLLOWING DAY AND A RESULT OF (B)(6) WAS REPORTED OUT TO THE LIS. FOR THIS SAME SAMPLE, THE MICRO ALBUMIN RESULT WAS NOT REPORTED UNTIL THE FOLLOWING DAY, WITH A VALUE 0.9. FOR THE NEXT SAMPLE ((B)(6)) WHICH WAS FLAGGED WITH SAMPLE QUEUE ERROR, THE POTASSIUM AND CREATININE RESULTS WERE REPORTED TO THE LIS. PER THE APS OPERATIONS MANUAL IN APPENDIX B UNDER WARNING MESSAGE W026, THESE SAMPLE RESULTS SHOULD NOT HAVE BEEN REPORTED OUT. THIS SAMPLE SHOULD HAVE BEEN RECOLLECTED AND THEN HAVE ALL ASSAYS REPEATED. A SEARCH OF COMPLAINT DATABASES DID NOT FIND OTHER COMPLAINTS OF THIS NATURE. ON (B)(6) 2007, A PRODUCT CORRECTION LETTER (FA02AMAR2007) WAS ISSUED TO PROVIDE CUSTOMERS WITH INFORMATION CONCERNING SAMPLE PRESENTATION AND SAMPLE QUEUE ERRORS ON THE APS WORKCELL AND CONFIGURATION OF THEIR MIDDLEWARE, INCLUDING ACCELERATOR DECISION MANAGER. A REVIEW OF ACCELERATOR DECISION MANAGER INTEGRATED PRODUCT CONFIGURATION GUIDE ((B)(6) SEPTEMBER 2007) DETERMINED THAT AS PART OF THIS ISA, THE ABBOTT FIELD SERVICE REPRESENTATIVE IS INFORMED TO UPDATE THE HANDLING OF AUTOMATION WARNING MESSAGES FROM THE APS. THE PURPOSE IS THE APS MAY GENERATE ERRORS THAT REQUIRE ADM TO PREVENT THE RESULTS FROM BEING REPORTED. WHEN A SAMPLE PRESENTATION OR A SAMPLE QUEUE ERROR IS DISPLAYED AT THE APS EXCEPTION SCREEN, AN AUTOMATION WARNING MESSAGE IS GENERATED WITH A WARNING CODE OF X01 OR X02. THE ABBOTT ARCHITECT SYSTEM OPERATIONS MANUAL (PART NUMBER 201837-105, MAY 2008) INDICATES UNDER SECTION 7, OPERATIONAL PRECAUTIONS AND LIMITATIONS, LIMITATIONS OF RESULT INTERPRETATION: ASSAY RESULTS MUST BE USED WITH OTHER CLINICAL DATA, FOR EXAMPLE, SYMPTOMS, OTHER TEST RESULTS, PATIENT HISTORY, CLINICAL IMPRESSIONS, INFORMATION AVAILABLE FROM CLINICAL EVALUATION, AND OTHER DIAGNOSTIC PROCEDURES. ALL DATA MUST BE CONSIDERED FOR PATIENT CARE MANAGEMENT. IF ASSAY RESULTS ARE INCONSISTENT WITH CLINICAL EVIDENCE, ADDITIONAL TESTING IS SUGGESTED TO CONFIRM THE RESULT. THE ARCHITECT SYSTEM HAS BEEN VALIDATED FOR ITS INTENDED USE. HOWEVER, ERRORS CAN OCCUR DUE TO POTENTIAL OPERATOR ERRORS AND ARCHITECT SYSTEM TECHNOLOGY LIMITATIONS. A REVIEW OF THE ACCELERATOR APS OPERATIONS MANUAL ((B)(6) - SEPTEMBER 2008) DETERMINED TROUBLESHOOTING ASSISTANCE IS PROVIDED IN APPENDIX B FOR SAMPLE PRESENTATION AND SAMPLE QUEUE ERRORS. THE ACCELERATOR DECISION MANAGER (ADM) IS DESIGNED FOR WORKFLOW AND DATA MANAGEMENT FOR ALL OF THE IMMUNOCHEMISTRY AND HEMATOLOGY SYSTEMS TO WHICH IT IS CONNECTED, IN CLINICAL LABORATORIES HAVING MULTIPLE INSTRUMENTS. A SINGLE BI-DIRECTIONAL CONNECTION WITH THE LIS OR HOSPITAL INFORMATION SYSTEM (HIS) ALLOWS THE ACCELERATOR DECISION MANAGER TO RECEIVE WORK ORDERS AND OUTPUT ANALYTICAL RECORDS. A MALFUNCTION OF THE SYSTEM WAS IDENTIFIED. THE COMPLAINT TEXT INDICATES IT WAS DETERMINED THE ADM MIDDLEWARE IS NOT PERFORMING AS INTENDED AND PER DESIGN WHEN USING THE SAMPLE TYPE AS PART OF THE BARCODE ID. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT MICROALBUMIN RESULTS FOR ONE PATIENT WERE GENERATED AT GREATER THAN 500 MG/L ON THE ARCHITECT C16000 ANALYZER/APS SYSTEM. THE SAMPLE WAS DILUTED AND GENERATED A FINAL RESULT OF 42 MG/L. THE SAMPLE WAS THEN RETESTED NEAT AND GENERATED A RESULT OF 6.9 MG/L. THE ABBOTT CUSTOMER TECHNICAL ADVOCATE (CTA) INSPECTED THE REACTION GRAPHS AND SUSPECTS THE INITIAL RUN WAS PERFORMED ON SERUM AND NOT URINE. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. IT IS SUSPECTED THAT THE ANALYZER ASPIRATED FROM AN INCORRECT SAMPLE. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 1
862701 | AFTER A RECENT INTERNAL COMPANY COMPLAINT, IT HAS BEEN IDENTIFIED THAT SPRING CLIPS ON DISPOSABLE TIP (DITI) RACKS MAY HAVE SHARP PROTRUDING EDGES. IN THE ASSOCIATED COMPLAINT, AFTER REMOVING THE DITI RACK FROM A DEVELOPMENTAL VERSION OF THE M2000SP DURING EVALUATION, TESTING AND REMOVING THE TIPS FROM THE RACK, A GLOVE WAS TORN AND AN ABBOTT MOLECULAR EMPLOYEE'S FINGER WAS SCRATCHED. THIS OBSERVATION OF SHARP SPRING CLIP EDGES HAS NO EFFECT ON M2000SP OPERATION OR FUNCTIONALITY, AND NO IMPACT ON PATIENT RESULTS. THE SHARP EDGE OF THE DITI RACK'S SPRING CLIP MAY CAUSE PERSONAL INJURY OF A SCRATCH OR CUT. THIS INJURY CARRIES RISK OF INFECTION. THE EXPOSED SPRING CLIPS DO NOT NORMALLY CONTACT BIOHAZARDOUS MATERIAL, BUT ARE INSIDE THE INSTRUMENT IN CLOSE PROXIMITY TO POTENTIALLY BIOHAZARDOUS SUBSTANCES. WEEKLY DECONTAMINATION PROCEDURES ARE RECOMMENDED WITHIN THE OPERATIONS MANUAL MAINTENANCE PROCEDURES AND OPERATORS ARE TO BE TRAINED IN COMPLIANCE WITH GOOD LABORATORY PRACTICES AND GENERAL BIOHAZARD AND SAFETY. >< IT HAS RECENTLY BEEN IDENTIFIED THAT SPRING CLIPS ON DITI RACKS MAY HAVE SHARP PROTRUDING CORNERS. AFTER REMOVING THE DITI CAN FROM A DEVELOPMENTAL M2000SP, GLOVES WERE TORN AND THE FINGER SCRATCHED OF AN ABBOTT MOLECULAR EMPLOYEE. AN ADDITIONAL ABBOTT MOLECULAR EMPLOYEE CUT A FINGER WHILE WORKING ON A DITI CAN IN A MACHINE SHOP (NON-STANDARD USER ENVIRONMENT) ENOUGH TO DRAW BLOOD. THREE ADDITIONAL OBSERVATIONS OF SHARP EDGES WERE DOCUMENTED, ALTHOUGH THESE OBSERVATIONS DID NOT RESULT IN INJURY. | 0
996792 | ACCOUNT INITIALLY GOT A NEGATIVE RESULT FOR LEUKOCYTES ON A PT URINE SAMPLE. WHEN EXAMINED MICROSCOPICALLY, THE RESULT WAS 30 - 49 WBC'S/HPF. VISUAL DIPSTICK OF THE SAME SAMPLE SHOWED NEGATIVE FOR WBC'S. CUSTOMER PERFORMED MAINTENANCE AND TROUBLESHOOTING ACTIVITIES WHICH RESOLVED THE ISSUE. | 0
1167344 | DYONICS ARTHROSCOPIC FLUID MANAGEMENT SYSTEM WAS BEING USED DURING ORTHO PROCEDURE, AND WAS MALFUNCTIONING BY PUTTING TOO MUCH FLUID INTO THE KNEE JOINT, EVEN AT THE LOWEST SETTING, CAUSING THE KNEE TO BECOME OVEREXTENDED WITH FLUID. THE PT DID NOT SUFFER PERMANENT DAMAGE OR INJURY. | 0
961064 | THE ABBOTT FIELD SERVICE REP (FSR) CLEANED THE WASH CUPS, CLEANED AND ALIGNED THE CUVETTE WASHER TO ADDRESS OVERFLOW AND REPLACED 1 ML SYRINGES. THE FSR ADVISED CUSTOMER ON THE NEED TO COMPLETE QUARTERLY MAINTENANCE. THE CUSTOMER CALLED BACK TO REPORT THE ISSUE WAS NOT RESOLVED. THE FSR RETURNED TO REPLACED THE ICT MODULE AND TUBING. THE SYSTEM WAS CALIBRATED SUCCESSFULLY. THE FSR RAN CONTROLS AND PERFORMED A PRECISION RUN WITH ACCEPTABLE RESULTS TO RESOLVE THE ISSUE. THE CUSTOMER ISSUE IS ADDRESSED IN THE AEROSET CUSTOMER TRAINING SUPPLEMENT-ICT PREVENTIVE MAINTENANCE (95132-101-FEBRUARY 2003) WITH INFO RELATED TO ADDRESSING ICT CONCERNS AND THE AEROSET SYSTEM OPERATIONS MANUAL (200155-101-NOVEMBER 2004) IN THE SECTIONS ENTITLED: SECTION 7- OPERATIONAL PRECAUTIONS AND LIMITATIONS, SECTION 9- SERVICE AND MAINTENANCE, AND SECTION 10- TROUBLESHOOTING AND DIAGNOSTICS. A MALFUNCTION OF THE SYSTEM OCCURRED DUE TO MULTIPLE ISSUES (DIRTY CUVETTE DRYER TIPS AND MIXER WASH CUPS AND LEAKING WAS PUMP SYRINGES). ALTHOUGH THE ISSUE RETURNED AFTER THE INITIAL FIELD SERVICE INTERVENTION, RESOLUTION CAME WITH THE REPLACEMENT OF THE SYRINGES, TUBING AND ICT MODULE. IT SHOULD BE NOTED THAT ON THIS OCCASION THE FSR ADVISED THE CUSTOMER TO COMPLETE QUARTERLY MAINTENANCE WHICH WAS NOT DOCUMENTED TO HAVE BEEN COMPLETED FOR THIS INSTANCE OF ANY OTHER BASED ON COMPLAINT HISTORY REVIEW. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATES THAT ERRATIC PT RESULTS ARE BEING GENERATED BY THE AEROSET SODIUM ASSAY. ONE PT RESULTS OF 126 MMOL/L WAS REPORTED OUT OF THE LAB. WHEN THIS LOW RESULT WAS NOTICED, THE SAME SAMPLE WAS TESTED ON ANOTHER AEROSET ANALYZER IN THE LAB AND GENERATED A RESULT OF 141 MMOL/L AND A CORRECTED REPORT WAS ISSUED. THE CUSTOMER THEN RAN A 10 POINT PRECISION RUN WITH THE FOLLOWING RESULTS: 123, 135, 131, 135, 119, 134, 134, 130, 131, AND 135 MMOL/L. UPON EXAMINING THE AEROSET ANALYZER IN QUESTION, THE CUSTOMER NOTICED THAT THE CUVETTE WIPERS AND MIXER CUPS ARE DIRTY AND THAT THE WASH PUMP SYRINGES ARE LEAKING. THE CUSTOMER DECLINED TO TROUBLESHOOT WITH THE CUSTOMER TECHNICAL ADVOCATE AND IS REQUESTING A SERVICE CALL. THERE IS NO IMPACT TO PT MGMT REPORTED. | 0
975127 | ADD'L MODEL- 5.4, 5.5, 5.6. THE SOFTWARE CORRECTION MODIFIES THE PROGRAMS WHICH CHECKS PT BLOOD TYPE FOR MERGING PT RECORDS. EACH OF THE LAB DATABASES WILL NOW PASS ONE OF THE FOLLOWING VALUES TO THE MEDICAL RECORDS MERGE: AN ERROR CODE IF THERE IS A BLOOD TYPE MISMATCH BETWEEN THE SELECTED MEDICAL RECORDS WITHIN ANY DATABASE. IF THERE IS ONLY ONE LOCAL HISTORY, THAT HISTORY'S BLOOD TYPE, OR NIL IF IT HAS NONE. IF THERE IS AT LEAST ONE BLOOD TYPE, THE BLOOD TYPE. IF THERE ARE TWO RECORDS AND NEITHER HAS A BLOOD TYPE, THE TEMPORARY BLOOD TYPE. IF THERE ARE NO BLOOD TYPES AND NO TEMP BLOOD TYPES, NIL. THE MEDICAL RECORDS MERGE NOW USES THE VALUES SENT FROM EACH OF THE LAB DATABASES TO COMPARE THE BLOOD TYPES ACROSS MULTIPLE LAB DATABASES CORRECTLY. IF THERE IS A DISCREPANCY, THE MERGE IS PREVENTED. IN ADDITION, TO CORRECT DATA IN HOSP SYSTEMS, CUSTOMERS WILL RECEIVE A LISTING OF SUSPECT PT RECORDS TO REVIEW FOR DISCREPANCIES AND RESOLVE. THE METHOD FOR IDENTIFYING RECORDS WILL CHECK ANY BBK HISTORY WITH A MERGE, CHECK ALL OTHER LAB DATABASES FOR OTHER BBK HISTORIES FOR THE SAME MEDICAL RECORD, COMPARE BLOOD TYPES AMONG ALL SUCH HISTORIES, AND RETURN A LISTING OF ALL MEDICAL RECORDS WITH BLOOD TYPE DISCREPANCIES. DEVICE EVALUATED BY MFR. THE MAGIC BLOOD BANK MODULE AT THE CUSTOMER LOCATION IS AVAILABLE TO MEDITECH STAFF THROUGH A DIAL-UP CONNECTION. THE CUSTOMER SITE HAS TWO SEPARATE DATABASES WITH THE SAME VERSION OF THE SOFTWARE. A 'TEST' DATABASES FOR RESEARCH, TESTING, AND VALIDATION; AND A 'LIVE' DATABASE USED FOR RECORDING DAILY OPERATIONS OF THE HOSP'S BLOOD BANK. INVESTIGATION FOUND THE SOFTWARE ERROR IN THE HOSP'S USER-SITE TEST DATABASE AND IN MEDITECH'S OWN IN-HOUSE CONTROLLED TESTING ENVIRONMENT. THE ERROR OCCURRED ONLY IN THE FOLLOWING SITUATION: THE HEALTHCARE ORGANIZATION MUST BE SET UP WITH MULTIPLE FACILITIES OF DATABASE CONFIGURATION MUST INCLUDE ONE MEDICAL RECORDS DATABASE LINKING TO TWO LAB DATABASES. THE MEDICAL RECORDS BEING MERGED MUST HAVE PT VISITS IN BOTH LAB DATABASES. THE MEDICAL RECORDS BEING MERGED MUST HAVE BOTH BLOOD BANK HISTORICAL DATA. THE MEDICAL RECORDS BEING MERGED MUST HAVE DIFFERENT BLOOD TYPES. MEDITECH'S BLOOD BANK INTERNAL REVIEW BOARD EVALUATED THE REPORT OF THE SOFTWARE ERROR. THE SEVERITY, LIKELIHOOD, AND DETECTABILITY WERE ANALYZED. THE SEVERITY OF THE ERROR WAS RATED AS HIGH, BASED ON THE POSSIBILITY THAT UNSUITABLE BLOOD PRODUCT COULD BE RELEASED IF THE HEALTHCARE PROFESSIONAL RELIED ONLY ON THE HISTORICAL RECORD OF BLOOD TYPE. STEPS FOR CORRECTIVE AND PREVENTIVE ACTIONS WERE INITIATED PER MEDITECH'S ESTABLISHED PROCEDURES. >< MEDITECH HAS IDENTIFIED A SAFETY CONCERN WITH VERSIONS 4.9, 5.4 AND 5.5 OF THE MAGIC BLOOD BANK SOFTWARE. THESE VERSIONS COULD MALFUNCTION, RESULTING IN INCORRECT PT HISTORICAL DATA. THE CUSTOMER RECENTLY RECEIVED A SOFTWARE UPDATE AND WAS VALIDATING THE BLOOD BANK SOFTWARE IN A TEST DATABASE. IN EARLY 2000, THE INFO SYSTEMS TESTER PERFORMED A MERGE OF TWO TEST PT ACCOUNTS WITH BLOOD BANK DATA, INCLUDING HISTORICAL BLOOD TYPE. ACCORDING TO THE VALIDATION SCRIPT, A WARNING MESSAGE SHOULD HAVE DISPLAYED AND THE TESTER PREVENTED FROM COMPLETING THE MERGE. OUR INVESTIGATION FOUND THAT THE MALFUNCTION COULD OCCUR ONLY WITH A SPECIFIC MULI-FACILITY DATABASE CONFIGURATION WHERE ONE MEDICAL RECORDS DATABASE REFERENCES TWO LAB DATABASES. THE MALFUNCTION HAS NOT BEEN REPORTED AS AFFECTING ANY ACTUAL HOSPITAL PTS. THE POTENTIAL FOR MISDIAGNOSIS BASED ON AN INCORRECT HISTORICAL BLOOD TYPE WAS DEEMED SERIOUS. COMPARISON OF HISTORICAL RECORDS WITH CURRENT RECORDS ARE REQUIRED PRIOR TO A TRANSFUSION. INCORRECT HISTORICAL RECORDS COULD LEAD TO MISDIAGNOSIS OR DELAY WHILE DISCREPANCIES ARE INVESTIGATED. | 1
1048325 | (ISSUE RESOLVED THROUGH TROUBLESHOOTING. NO DEFINITE CAUSE FOR THE LOW PHOSPHORUS RESULTS). THE CUSTOMER OBSERVED LOW PHOSPHORUS RESULTS FOR PT SAMPLES. IN ORDER TO RESOLVE THE ISSUE, THE CUSTOMER REPLACED THE PHOSPHORUS REAGENT CARTRIDGE ON THE INSTRUMENT; HOWEVER, THE NEW REAGENT CARTRIDGE DID NOT RESOLVE THE ISSUE. THE CUSTOMER TECHNICAL ADVOCATE (CTA) INSTRUCTED THE CUSTOMER TO PERFORM A SERIES OF TROUBLESHOOTING STEPS. THE CUSTOMER PERFORMED THE RECOMMENDED TROUBLESHOOTING. NO INSTRUMENT PARTS WERE REPLACED; HOWEVER, THE CUSTOMER DID TIGHTEN THE CONNECTIONS AROUND THE TUBING AND STYLETTED THE CUVETTE WASHER NOZZLES. THE CUSTOMER STATED AFTER THE COMPLETION OF TROUBLESHOOTING, THERE HAVE BEEN NO OTHER OCCURRENCES OF PHOSPHORUS RESULTS LESS THAN 0.5 MG/DL. THE CONTROLS HAVE BEEN IN RANGE AND THE INSTRUMENT IS PERFORMING WITHIN SPECS. A REVIEW OF THE COMPLAINT HISTORY FOR THE ARCHITECT CSYSTEM OVER THE TIME PERIOD OF 2007 THROUGH 2008 WAS PERFORMED. THE REVIEW DID NOT FIND OTHER COMPLAINTS RELATED TO THIS ISSUE. THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-104 2007) INDICATES THE FOLLOWING RELATED TO THE CUSTOMER ISSUE: SECTION 9: SERVICE MAINTENANCE MONTHLY MAINTENANCE DESCRIPTION (C SYSTEM PREOCESSING MODULE). MONTHLY MAINTENANCE FOR THE C SYSTEM IS REQUIRED ON THE PROCESSING MODULE ONLY. PERFORM THESE PROCEDURES MONTHLY: 6016 CHECK DISPENSE COMPONENTS, PAGE 9-27, 6018 CLEAN CUVETTE WASHER NOZZLES, PAGE 9-27, 6026 CHECK SYRINGES AND VALVES, PAGE 9-28, 6300 CLEAN ICT DRAIN TIP, PAGE 9-28. SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS, UNDER THE OBSERVED PROBLEMS - ERRATIC RESULTS, POOR PRECISION - PHOTOMETRIC RESULTS (C SYSTEM) AND CONTROLS OUT OF RANGE (C-SYSTEM) SECTIONS, TWO TABLES OF PROBABLE CAUSES AND CORRECTIVE ACTIONS HAVE BEEN PROVIDED FOR THE CUSTOMER TO TROUBLESHOOT THEIR SYSTEM. CONCLUSION: ALTHOUGH THERE IS NOT DEFINITE CAUSE FOR THE LOW PHOSPHORUS RESULTS, THE ISSUE APPEARS TO BE RESOLVED BY TIGHTENING TUBING CONNECTIONS AND CLEANING THE CUVETTE WASHER NOZZLES, WHICH IS PART OF MONTHLY MAINTENANCE. THIS IS A FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATES THAT THE ARCHITECT C8000 ANALYZER GENERATED PHOSPHORUS RESULTS OF LESS THAN 0.5 MG/DL FOR PTS. WHEN REPEATED, THE RESULTS WERE BETWEEN 5.0 MG/DL AND 6.0 MG/DL. THE CUSTOMER OBSERVED THE SAME PHENOMENON ON AT LEAST ONE CONTROL VALUE. THE LOW PHOSPHORUS RESULTS WERE NOT REPORTED OUT OF THE LABORATORY. THERE WAS NO REPORT OF IMPACT TO THE PT MANAGEMENT. | 0
1036330 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< THE CUSTOMER SUSPECTED TWO CONSECUTIVE SAMPLES REPORTED SIMILAR RESULTS FOR ALL TESTS. THE CUSTOMER REVIEWED ALL RESULTS AND FOUND THAT SAMPLE #136 WAS ASSIGNED SIMILAR RESULTS TO THAT OF SAMPLE #137. AN ASPIRATION ERROR HAD OCCURRED ON SAMPLE #135 JUST PRIOR TO SAMPLES #136 AND #137. NO INCORRECT RESULTS WERE REPORTED OUT OF THE LABORATORY. THE CUSTOMER MESSAGE HISTORY LOGS WERE REVIEWED AND CONFIRMED THAT A MISS-ASSOCIATION DID OCCUR AND THAT SAMPLE #136 WAS NOT ASPIRATED. SAMPLE #137 WAS INAPPROPRIATELY ASPIRATED AND USED TO GENERATE RESULTS FOR SAMPLE #136. SAMPLE #137 WAS AGAIN ASPIRATED AND GENERATED RESULTS FOR SAMPLE #137. NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
1211964 | (B)(4). THE ERRATIC RESULT ISSUE WAS RESOLVED BY THE FIELD SERVICE REPRESENTATIVE (FSR) REPLACING AND CALIBRATING THE REAGENT PROBES. POTASSIUM CONTROL RESULTS WERE VERIFIED TO BE AT TARGET VALUES AFTER THE REAGENT PROBES WERE REPLACED. THE COMPLAINT TEXT ALSO NOTES THE CUSTOMER SITE HAD IMPURITIES IN THE WATER SUPPLY PIPE. LABELING REVIEW: THE AEROSET OPERATIONS MANUAL LIST NO. 09D06-05 WAS REVIEWED AND WAS FOUND TO CONTAIN ADEQUATE INFORMATION ON PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR ERRATIC RESULTS. SECTION 10, TROUBLESHOOTING AND DIAGNOSTICS, OBSERVED PROBLEMS LISTS MULTIPLE PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR ERRATIC RESULTS OR PRECISION IS POOR. THESE PROBABLE CAUSES INCLUDE LEAKING, DRIPPING, DAMAGED OR OBSTRUCTED REAGENT PROBES, AND DIRTY WATER BATH. CORRECTIVE ACTIONS INCLUDE REPLACING THE REAGENT PROBES AND CHANGING THE WATER IN THE BATH. SERVICE HISTORY REVIEW: A REVIEW OF COMPLAINTS FOR THE TIME PERIOD OF (B)(6) 2008 TO (B)(6) 2008 FOUND NO ADDITIONAL INCIDENTS OF AEROSET SERIAL NUMBER (B)(4) GENERATING ERRATIC RESULTS SINCE THE REAGENT PROBES WERE REPLACED BY THE FSR ON (B)(6) 2008. A REVIEW OF US AND INTERNATIONAL COMPLAINTS FOR ERRATIC, INCONSISTENT AND ABERRANT RESULTS FROM (B)(6) 2007 THROUGH (B)(6) 2008 DETERMINED THE RATES ARE CURRENTLY WITHIN ESTABLISHED LIMITS. NO PRODUCT DEFICIENCY WAS IDENTIFIED. THIS IS THE FINAL REPORT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED A PATIENT GENERATED A POTASSIUM RESULT OF 9.7 MMOL/L ON THE AEROSET ANALYZER. THE SAMPLE WAS RETESTED YIELDING A POTASSIUM RESULT OF 4.1 MMOL/L. THE VALUE OF 4.1 MMOL/L ALIGNED WITH THE PATIENT'S PREVIOUS VALUE AND WAS CONFIRMED CORRECT ON ANOTHER AEROSET ANALYZER. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
1183384 | (B)(4). ABBOTT FIELD SERVICE WAS DISPATCHED TO THE ACCOUNT AND FOUND ICT REFERENCE SOLUTION LEAKING FROM THE ASPIRATION LINE AND A STICKY ICT ASPIRATION SYRINGE. FIELD SERVICE RESOLVED THE CUSTOMER ISSUE ON SITE BY CLEANING AND REPLACING THE ICT ASPIRATION SYRINGE FOLLOWED BY A PASSING CALIBRATION AND ACCEPTABLE QC. LABELLING IN THE ARCHITECT OPERATIONS MANUAL LN 06E28-07 ((B)(4) 2007) PROVIDES SUFFICIENT INFORMATION WITH REGARD TO SYSTEM MAINTENANCE AND TROUBLESHOOTING THE CUSTOMER'S ISSUE. TROUBLESHOOTING PERFORMED BY FIELD SERVICE TO RESOLVE THE ISSUE WAS CONSISTENT WITH TROUBLESHOOTING INFORMATION IN THE ARCHITECT OPERATIONS MANUAL. NO DEFICIENCY WAS IDENTIFIED. THIS IS THE FINAL REPORT. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED THE FOLLOWING RESULTS ON A PATIENT: SODIUM = 107 MMOL/L, POTASSIUM = 2.8 MMOL/L, AND CHLORIDE = 141 MMOL/L. THE SAMPLE WAS REPEATED ON THE SAME ANALYZER WITH THE FOLLOWING RESULTS: SODIUM = 144 MMOL/L, POTASSIUM = 3.8 MMOL/L, AND CHLORIDE = 104 MMOL/L. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. | 0
1098663 | THE CUSTOMER STATES THAT THE ARCHITECT C8000 ANALYZER GENERATED ERRATIC PHOSPHORUS ASSAY RESULTS ON A PATIENT SAMPLE. A PATIENT SAMPLE GENERATED AN INITIAL PHOSPHORUS RESULT OF 7.9 MG/DL, AND THE PHOSPHORUS RESULTS OF 4.4 MG/DL AND 4.3 MG/DL WHEN THE SAMPLE WAS REPEATED. THE LOWER RESULT WAS REPORTED OUTSIDE OF THE LABORATORY. THE CONTROLS WERE IN RANGE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT DUE TO THIS ERRATIC RESULT. >< AN INVESTIGATION WAS PERFORMED IN RESPONSE TO THE CUSTOMER'S COMPLAINT OF A DISCREPANT PHOSPHORUS PATIENT RESULT GENERATED BY THE ARCHITECT C8000 ANALYZER, IN 2008. THE INVESTIGATION OF THE ISSUE CONSISTED OF A REVIEW OF CUSTOMER COMPLAINTS, CURRENT C8000 LABELING, AND THE INFORMATION PROVIDED INCLUDING THE COMPLAINT TEXT. THE REPLACEMENT OF THE DAMAGED MIXER 1 APPEARS TO HAVE RESOLVED YOUR C8000 DISCREPANT RESULT ISSUE. NO ADDITIONAL DISCREPANT RESULT OCCURRENCES BY C8000 HAVE BEEN REPORTED SINCE THE BENT MIXER 1 WAS REPLACED. A REVIEW OF THE COMPLAINT HISTORY FOR THE ARCHITECT CSYSTEM WAS PERFORMED OVER THE TIME PERIOD 2008. THIS SEARCH DID NOT FIND OTHER COMPLAINTS CONCERNING ERRATIC RESULTS FOR THE PHOSPHORUS ASSAY. THE C8000 OPERATIONS MANUAL PROVIDES MULTIPLE PROBABLE CAUSES AND CORRECTIVE ACTIONS TO ASSIST THE CUSTOMER WITH RESOLVING ERRATIC/DISCREPANT RESULTS: OPERATIONAL PRECAUTIONS AND LIMITATIONS UNDER, LIMITATIONS OF RESULT INTERPRETATION STATING THAT ASSAY RESULTS MUST BE USED WITH OTHER CLINICAL DATA, FOR EXAMPLE, SYMPTOMS, OTHER TEST RESULTS, PATIENT HISTORY, CLINICAL IMPRESSIONS, INFORMATION AVAILABLE FROM CLINICAL EVALUATION, AND OTHER DIAGNOSTIC PROCEDURES. ALL DATA MUST BE CONSIDERED FOR PATIENT CARE MANAGEMENT. IF ASSAY RESULTS ARE INCONSISTENT WITH CLINICAL EVIDENCE, ADDITIONAL TESTING IS SUGGESTED TO CONFIRM THE RESULT. THE ARCHITECT SYSTEM HAS BEEN VALIDATED FOR ITS INTENDED USE. TROUBLESHOOTING AND DIAGNOSTICS DESCRIBES THE PROBABLE CAUSE AND CORRECTIVE ACTION(S) FOR MIXER MALFUNCTION UNDER ERRATIC RESULTS, POOR PRECISION - PHOTOMETRIC RESULTS (C SYSTEM). A MALFUNCTION OF THE SYSTEM AND REAGENT APPEARS TO HAVE OCCURRED. THE COMPLAINT TEST STATES THE CUSTOMER REPLACED MIXER 1 AND NO ADDITIONAL ISSUES WERE OBSERVED. THIS IS A FINAL REPORT. END OF REPORT. | 0
1192608 | THE PATHOLOGIST STATED THAT THE PATIENT'S LABORATORY RESULTS INDICATED THAT THE CAUSE OF DEATH WAS A MYOCARDIAL INFARCTION. THERE IS NO INDICATION THAT THE ARCHITECT C8000 CREATININE RESULTS CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. >< THE CUSTOMER STATED THAT AN ARCHITECT C8000 CREATININE RESULT OF 76 UMOL/L WAS REPORTED ON A FEMALE PATIENT IN 2008. A DIFFERENT SAMPLE FROM THE SAME PATIENT WAS TESTED THE PREVIOUS NIGHT AT A DIFFERENT LABORATORY USING THE BECKMAN CREATININE METHOD AND THE RESULT WAS 400 UMOL/L. THE PHYSICIAN REQUESTED THAT THE ARCHITECT C8000 RESULT BE VERIFIED. THE SAMPLE RETESTED AT 370 UMOL/L. THE PATIENT HAD BEEN GIVEN RADIOGRAPHIC CONTRAST MEDIA. THE CUSTOMER STATED THAT THE PATIENT EXPIRED AND THAT THE PATHOLOGIST COULD NOT CONFIRM THE CAUSE OF DEATH, BUT STATED THAT THE PATIENT'S LABORATORY RESULTS INDICATED THAT THE PATIENT HAD A MYOCARDIAL INFARCTION. THERE IS NO INDICATION THAT THE ARCHITECT C8000 CREATININE RESULTS CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. >< (B)(4). BASED ON DISCUSSION BETWEEN DR. (B)(6) (ABBOTT DIAGNOSTICS DIVISION MEDICAL DIRECTOR) AND (B)(6) (FOOD & DRUG ADMINISTRATION, SCIENTIFIC REVIEWER) ON (B)(6), 2008, THIS FOLLOW-UP MEDWATCH 3500A REPORT IS BEING SUBMITTED TO INDICATE ADVERSE EVENT, DEATH PER (B)(6) RECOMMENDATION. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). AN INVESTIGATION WAS PERFORMED IN RESPONSE TO THE CUSTOMER'S COMPLAINT OF A DISCREPANT CREATININE PATIENT RESULT GENERATED BY ARCHITECT C8000 (B)(4) ON (B)(4) 2008. THE INVESTIGATION OF THE ISSUE CONSISTED OF A REVIEW OF CUSTOMER COMPLAINTS, CURRENT ARCHITECT C8000 LABELING, AND THE INFORMATION PROVIDED INCLUDING THE COMPLAINT TEXT. THE REPLACEMENT OF POPPET VALVES ON THE C8000 APPEARS TO HAVE RESOLVED THE CUSTOMER'S DISCREPANT RESULT ISSUE. NO ADDITIONAL DISCREPANT RESULT OCCURRENCES BY ARCHITECT C8000 (B)(4) HAVE BEEN REPORTED SINCE THE POPPET VALVES WERE REPLACED. RESULTS OF THE INVESTIGATION DETERMINED THAT THERE WAS NO PRODUCT DEFICIENCY. THIS IS THE FINAL REPORT. | 0
953134 | AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO INSPECT THE CUSTOMER'S INSTRUMENT AND FOUND ALL B-LINE NOZZLES WERE SEVERELY CLOGGED. THE B-LINE NOZZLES WERE CLEARED AND START-UP AND CONTROLS WERE PERFORMED TO VERIFY SYSTEM PERFORMANCE. NO FURTHER FOLLOW-UP OR PARTS WERE NEEDED. SUBSEQUENT INSTRUMENT OPERATIONS WERE ACCEPTABLE. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE AEROSET PHOSPHORUS REAGENT PACKAGE INSERT (SPECIMEN COLLECTION AND HANDLING) AND THE AEROSET SYSTEM OPERATIONS MANUAL (P/N 200154-101: NOVEMBER 2004) IN SECTION 7 - OPERATIONAL PRECAUTIONS AND LIMITATIONS, IN SECTION 9 - SERVICE AND MAINTENANCE UNDER 'OVERVIEW', AND IN SECTION 10 - TROUBLESHOOTING AND DIAGNOSTICS, UNDER OBSERVED PROBLEMS AND ERROR LOG MESSAGES. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT ERRATIC RESULTS ARE BEING GENERATED ON THE AEROSET ANALYZER. THE CUSTOMER GAVE THE GENERAL EXAMPLE OF PATIENT PHOSPHORUS ASSAY RESULTS INITIALLY READING 1.8 MG/DL AND RETESTING AT 3.8 MG/DL. CONTROLS WERE WITHIN SPECIFICATIONS EARLIER IN THE DAY. THE CUSTOMER REQUESTED A SERVICE CALL. THERE IS NO IMPACT TO PT MANAGEMENT REPORTED. | 0
998072 | THE ACCOUNT STATED 3 PATIENT SPECIMENS IN A ROW WERE FLAGGED AS LESS THAN THE LINEARITY LIMIT AND GENERATED ERRATIC RESULTS ON THE ARCHITECT C8000 ANALYZER. THE ACCOUNT REPEATED THE SPECIMENS ON ANOTHER C8000 ANALYZER WITH ACCEPTABLE RESULTS. THE ACCOUNT SENT CORRECTED REPORTS OUT OF THE LABORATORY. IT IS UNKNOWN IF THERE WAS IMPACT TO PATIENT MANAGEMENT DUE TO THE LOW OR ERRATIC RESULTS. THROUGH TROUBLESHOOTING, IT WAS DISCOVERED THAT THE ACCOUNT BIOMEDICAL ENGINEER PERFORMED QUARTERLY MAINTENANCE ON THE ARCHITECT C8000 EARLIER IN THE DAY. AFTER THE LOW AND ERRATIC RESULTS WERE GENERATED, THE ACCOUNT NOTICED LEAKING FROM THE R2 SYRINGE AND THE WASH PUMP PROBE. SERVICE WAS REQUESTED FOR THE ARCHITECT C8000 ANALYZER. SPECIMEN DATA : INITIAL SODIUM = 178 (UNITS OF MEASUREMENT NOT GIVEN). REPEAT SODIUM = 142 (UNITS OF MEASUREMENT NOT GIVEN). INITIAL TOTAL BILIRUBIN = 1.7 (UNITS OF MEASUREMENT NOT GIVEN). REPEAT TOTAL BILIRUBIN = 0.9 (UNITS OF MEASUREMENT NOT GIVEN). INITIAL POTASSIUM = 4.9 (UNITS OF MEASUREMENT NOT GIVEN). >< THE ABBOTT FIELD SERVICE ENGINEER REPLACED THE PUMP ON THE ARCHITECT C8000 ANALYZER. VERIFICATION PROCEDURES WERE COMPLETED TO CONFIRM THAT THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. REFERENCE WHERE IN LABELING THE EVENT IS ADDRESSED: ABBOTT ARCHITECT SYSTEM. SYMPTOMS OF IMPRECISE RESULTS, AND/OR ERRATIC RESULTS OBSERVED RESULTS. ERRATIC RESULTS, POOR PRECISION - ICT RESULTS (C SYSTEM) REAGENT TROUBLESHOOTING VARIABLES (C SYSTEM). THIS IS A FINAL REPORT. | 0
946585 | THE CUSTOMER REPORTED THE ORTHO PROVUE ANALYZER IS RANDOMLY NOT DISPENSING RED CELLS INTO THE MTS A/B/D MONOCLONAL AND REVERSE GROUPING CARDS. INCORRECT OR NO DISPENSE CAN LEAD TO ERRONEOUS TEST RESULTS AND TRANSFUSION OF INCOMPATIBLE BLOOD. THE OPERATOR DETECTED THE ISSUE PREVENTING THE POSSIBILITY OF ERRONEOUS RESULTS FROM BEING REPORTED. >< A DEFINITIVE ROOT CAUSE WAS DETERMINED FOR THE REPORTED EVENT. A LEAKING SYRINGE MAY LEAD TO A LOSS OF VACUUM/PRESSURE IN THE FLUIDICS SYSTEM AND INCORRECT/NO-DISPENSE. INFORMATION WAS NOT PROVIDED REGARDING ERROR CODES POSTED AS A RESULT OF THIS ISSUE. REPAIRS HAVE RETURNED THE ANALYZER TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. | 0
1030591 | INVESTIGATION SUMMARY: THE CUSTOMER STATES THAT DISCREPANT AFP RESULTS WERE GENERATED USING THE ARCHITECT AFP ASSAY. THE INSTRUMENT WAS EVALUATED BY A FIELD SVC REP AND THE TRIGGER/PRE-TRIGGER VALVE FUNCTION WAS OBSERVED. THE TRIGGER DRAIN VALVE WAS NOT FUNCTIONING CORRECTLY. THE MANIFOLD VALVE WAS REPLACED AND THE VALVE TEST PASSED. TRENDING: A REVIEW OF COMPLAINTS WAS PERFORMED FOR THE TIME PERIOD OF 1/1/2008 THROUGH 4/1/2008. NO TRENDS RELATED TO AFP RESULTS WERE FOUND. LABELING: THE EVENT IS ADDRESSED IN THE ABBOTT ARCHITECT SYS OPERATIONS MANUAL (WITH C16000) SECTION 10-555 (PN 201837-102) OCTOBER, 2006 TROUBLESHOOTING AND DIAGNOSTICS SECTION 10 SAMPLE RESULTS OBSERVED PROBLEMS (I SYS) ERRATIC ASSAY RESULTS (I SYS). THIS PROBLEM MAY BE OBSERVED ON AN ARCHITECT I SYS. PROBABLE CAUSE; SYRINGE ASSEMBLIES OR VALVES ARE LEAKING. CORRECTIVE ACTION; CHECK FOR SALT CRYSTALS AND/OR LIQUID AROUND ALL SYRINGES OR VALVES. CONTACT YOUR AREA CUSTOMER SUPPORT IF CRYSTALS OR LIQUID ARE OBSERVED. CONCLUSION: ACTUAL DEVICE INVOLVED IN INCIDENT WAS EVALUATED. ANALYSIS OF LABELING WAS PERFORMED. A COMPONENT/SUBASSEMBLY FAILURE (VALVE) ON THE DEVICE WAS FOUND TO BE MALFUNCTIONING CONTRIBUTING TO THE EVENT. THE MANIFOLD VALVE WAS REPLACED AND THE VALVE TEST PASSED RESOLVING THE ISSUE. THERE WAS NO IMPACT TO PT MGMT REPORTED DUE TO THIS EVENT. THIS IS THE FINAL REPORT. >< THE CUSTOMER STATES THAT UNEXPECTED HIGH RESULT WAS GENERATED FOR A PT SAMPLE TESTED USING THE ARCHITECT AFP ASSAY THAT WHEN REPEATED, A LOWER RESULT WAS OBTAINED. NO IMPACT TO PT MGMT WAS REPORTED. PT RESULTS: INITIAL RESULT 192.29 NG/ML, REPEAT RESULT 3.39 NG/ML. | 0
1097963 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH CAPTURE-R READY ID (CRRID) ON A PATIENT SAMPLE TESTED BY THE ECHO. THE UNEXPECTED REACTIONS WERE AS FOLLOWS: NEGATIVE WITH CELL 14 ( E-K+K+FYA+FYB+), NEGATIVE WITH CELL 8 ( E-K+K+FYA-) >< REACTIVITY OF THE FYA, E AND K ANTIGENS WERE CONFIRMED ON RETENTION CRRID, LOT ID101, THE LOT USED BY THE CUSTOMER. THE CUSTOMER DID NOT RETURN PRODUCT OR SAMPLE FOR INVESTIGATION TESTING. IT IS NOT POSSIBLE TO RULE OUT THE SAMPLE AS A CAUSE OF THE EVENT. | 0
885665 | THE CUSTOMER STATED TEST RESULTS ON THE ARCHITECT C16000 SYSTEM WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. FOR EXAMPLE, SAMPLE 1 SHOWED THE RESULTS FROM SAMPLE 2, SAMPLE 2 SHOWED THE RESULTS FROM SAMPLE 3, ETC. THE ROBOTIC SAMPLE HANDLER (RSH) GENERATED ERROR CODES 0274, TRAY IN BAY (X) WAS REMOVED BEFORE ACCESS WAS GRANTED, AND 0526, EMPTY TRAY DETECTED IN BAY (X), INDICATING THAT THE OPERATOR HAD INAPPROPRIATELY REMOVED A TRAY FROM THE RSH. CODE 0304, TRANSPORTER ERROR ALSO OCCURRED. THE RSH WENT TO STOPPED STATUS WHILE THE PROCESSING MODULE REMAINED IN RUNNING. THE OPERATOR ORDERED RERUNS FROM THE EXCEPTION SCREEN AND BEGAN TO RUN AGAIN. THE OPERATOR SUSPECTED AN ABERRANT RESULT ON A SAMPLE AND STOPPED THE RUN. REVIEW SHOWED THAT THE RESULTS WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. SAMPLE #4 HAD THE RESULTS FROM SAMPLE #5. THE OPERATOR FOUND 28 CONSECUTIVE SAMPLES THAT WERE AFFECTED. THIRTEEN OF THE RESULTS HAD ALREADY BEEN REPORTED TO THE DOCTOR, BUT NO IMPACT TO PT MANAGEMENT WAS REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. FOUR CONDITIONS WERE IDENTIFIED THAT CAN LEAD TO A STOP OF THE ROBOTIC SAMPLE HANDLER (RSH) WHEN THE SAMPLE IS PRESENTED TO THE CHEMISTRY MODULE FOR ASPIRATION: REMOVAL OF A TRAY FROM A ROUTINE BAY WHILE THE ACCESS INDICATOR IS AMBER; - OPENING AN RSH COVER WHILE THE RSH IS IN THE RUNNING STATE; - USER REQUESTING STOP EITHER ON THE RSH KEYPAD OR FROM THE SNAPSHOT SCREEN ON THE SYSTEM CONTROL CENTER; - A NON-RECOVERABLE HARDWARE FAILURE OF AN RSH MECHANISM. INVESTIGATION OF THE ISSUE DISCOVERED A TIMING-DEPENDENT DEFECT WHERE THE TRAY-REMOVAL DOES NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CLINICAL CHEMISTRY MODULE. THIS CONDITION CAUSES THE MODULE TO RECEIVE AN ERRONEOUS SAMPLE ARRIVAL FOR TESTS/SAMPLES THAT WERE PREVIOUSLY DELETED WHEN THE RSH NOTIFIED THE INTERFACE SOFTWARE OF THE TRAY REMOVAL ERROR CONDITION. AFTER A SUBSTANTIAL DELAY THE MODULE SENDS AN EQUALLY ERRONEOUS SAMPLE FREE MESSAGE THAT CAUSES THE RSH TO INDEX THE CARRIER TO THE NEXT POSITION. AT THAT POINT THE SEQUENCE OF SAMPLING IS OFF BY ONE CARRIER POSITION. THE ISSUE WAS IDENTIFIED IN SOFTWARE VERSION 3.00 FOR THE ARCHITECT C16000 SYSTEM, BUT ALL SOFTWARE VERSIONS THAT SUPPORTED AN INTERFACE TO THE C8000 AND/OR THE C16000 AND RSH ARE SUSCEPTIBLE TO THIS ISSUE (V2.00, V2.01, V2.02, V2.03, V2.10, V2.11, V2.20, V2.20 DB UPDATE, V2.60, V2.30, V3.00 AND V3.10). THE ISSUE CAN OCCUR ON THE C8000 AND C16000 SYSTEMS SINCE THE RSH AND THE INTERFACE SOFTWARE ARE COMMON TO BOTH CONFIGURATIONS. THE ISSUE WAS FOUND TO BE DUE TO A TIMING- DEPENDENT DEFECT WHERE THE TRAY REMOVAL WOULD NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CHEMISTRY MODULE. LABELING: THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS DESCRIPTIONS OF THE USE AND FUNCTION OF THE RSH IN SECTION 1. STATUS INDICATORS INDICATES THE STATUS OF SAMPLE PROCESSING AND WHEN THE OPERATOR CAN ACCESS SAMPLES. A STEADY AMBER LIGHT INDICATES THAT SAMPLES ARE PROCESSING AND CANNOT BE ACCESSED. SECTION 5: OPERATING INSTRUCTIONS; SUBSECTION: SAMPLE MANAGEMENT PROVIDES PROCEDURES FOR LOADING SAMPLES ONTO THE RSH FOR PROCESSING. THE PROCEDURE 'ACCESS A SAMPLE WITH TESTS IN PROCESS' ADVISES THE OPERATOR TO PERFORM THIS PROCEDURE TO SUSPEND A BAY ON THE RSH WHEN IMMEDIATE ACCESS IS NEEDED TO A SAMPLE THAT IS BEING PROCESSED. IMPROPER REMOVAL OF A TRAY FROM THE RSH WILL CAUSE THE SYSTEM TO GENERATE ERROR CODE 0273 (REMOVAL FROM THE PRIORITY BAY) OR 0274 (REMOVAL FROM THE ROUTINE BAY). THE STOPPED STATUS FOR THE RSH IS DESCRIBED IN TABLE 1.3. CONDITIONS FOR THE RSH TO BE STOPPED INCLUDE:- THE OPERATOR HAS SELECTED F6- STOP FROM THE SNAPSHOT SCREEN OR THE STOP KEY ON THE SAMPLE HANDLER KEYPAD;- ONE OF THE RSH COVERS WAS OPENED WHILE THE SAMPLE HANDLER WAS RUNNING;- SAMPLE HANDLER DETECTED A FATAL ERROR WHILE PROCESSING. THE OPERATOR IS ADVISED IN SECTION 7: OPERATIONAL PRECAUTIONS AND LIMITATIONS TO 'KEEP ALL PROCESSING MODULE AND SAMPLE HANDLER DOORS CLOSED AND COVERS IN PLACE UNLESS INSTRUCTED OTHERWISE IN A MAINTENANCE OR TROUBLESHOOTING PROCEDURE' WHILE OPERATING THE SYSTEM. OPENING A COVER DURING OPERATION WILL RESULT IN THE GENERATION OF ERROR CODE 0709, RSH STOPPED, COVER OPENED. TROUBLESHOOTING ASSISTANCE FOR ALL ERROR CODES IS PROVIDED IN SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS. | 0
885662 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. FOUR CONDITIONS WERE IDENTIFIED THAT CAN LEAD TO A STOP OF THE ROBOTIC SAMPLE HANDLER (RSH), WHEN THE SAMPLE IS PRESENTED TO THE CHEMISTRY MODULE FOR ASPIRATION: REMOVAL OF A TRAY FROM A ROUTINE BAY WHILE THE ACCESS INDICATOR IS AMBER; - OPENING AN RSH COVER WHILE THE RSH IS IN THE RUNNING STATE; - USER REQUESTING STOP EITHER ON THE RSH KEYPAD OR FROM THE SNAPSHOT SCREEN ON THE SYSTEM CONTROL CENTER; - A NON-RECOVERABLE HARDWARE FAILURE OF AN RSH MECHANISM. INVESTIGATION OF THE ISSUE DISCOVERED A TIMING-DEPENDENT DEFECT WHERE THE TRAY-REMOVAL DOES NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CLINICAL CHEMISTRY MODULE. THIS CONDITION CAUSES THE MODULE TO RECEIVE AN ERRONEOUS SAMPLE ARRIVAL FOR TESTS/SAMPLES THAT WERE PREVIOUSLY DELETED, WHEN THE RSH NOTIFIED THE C8000II.DLL OF THE TRAY REMOVAL ERROR CONDITION. AFTER A SUBSTANTIAL DELAY, THE MODULE SENDS AN EQUALLY ERRONEOUS SAMPLE FREE MESSAGE THAT CAUSES THE RSH TO INDEX THE CARRIER TO THE NEXT POSITION. AT THAT POINT, THE SEQUENCE OF SAMPLING IS OFF BY ONE CARRIER POSITION. THE ISSUE WAS IDENTIFIED IN SOFTWARE VERSION 3.00 FOR THE ARCHITECT C16000 SYSTEM, BUT ALL SOFTWARE VERSIONS THAT SUPPORTED AN INTERFACE TO THE C8000 AND/OR THE C16000 AND RSH ARE SUSCEPTIBLE TO THIS ISSUE (V2.00, V2.01, V2.02, V2.03, V2.10, V2.11, V2.20, V2.20 DB UPDATE, V2.60, V2.30, V3.00 AND V3.10). THE ISSUE CAN OCCUR ON THE C8000 AND C16000 SYSTEMS SINCE THE RSH AND THE INTERFACE SOFTWARE (C8000II.DLL) ARE COMMON TO BOTH CONFIGURATIONS. THE ISSUE WAS FOUND TO BE DUE TO A TIMING- DEPENDENT DEFECT, WHERE THE TRAY REMOVAL WOULD NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CHEMISTRY MODULE. LABELING: THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS DESCRIPTIONS OF THE USE AND FUNCTION OF THE RSH IN SECTION 1. STATUS INDICATORS INDICATES THE STATUS OF SAMPLE PROCESSING AND WHEN THE OPERATOR CAN ACCESS SAMPLES. A STEADY AMBER LIGHT INDICATES THAT SAMPLES ARE PROCESSING AND CANNOT BE ACCESSED. SECTION 5: OPERATING INSTRUCTIONS; SUBSECTION: SAMPLE MANAGEMENT PROVIDES PROCEDURES FOR LOADING SAMPLES ONTO THE RSH FOR PROCESSING. THE PROCEDURE 'ACCESS A SAMPLE WITH TESTS IN PROCESS' ADVISES THE OPERATOR TO PERFORM THIS PROCEDURE TO SUSPEND A BAY ON THE RSH, WHEN IMMEDIATE ACCESS IS NEEDED TO A SAMPLE THAT IS BEING PROCESSED. IMPROPER REMOVAL OF A TRAY FROM THE RSH WILL CAUSE THE SYSTEM TO GENERATE ERROR CODE 0273 (REMOVAL FROM THE PRIORITY BAY) OR 0274 (REMOVAL FROM THE ROUTINE BAY). THE STOPPED STATUS FOR THE RSH IS DESCRIBED IN TABLE 1.3. CONDITIONS FOR THE RSH TO BE STOPPED INCLUDE:- THE OPERATOR HAS SELECTED F6- STOP FROM THE SNAPSHOT SCREEN OR THE STOP KEY ON THE SAMPLE HANDLER KEYPAD;- ONE OF THE RSH COVERS WAS OPENED WHILE THE SAMPLE HANDLER WAS RUNNING;- SAMPLE HANDLER DETECTED A FATAL ERROR WHILE PROCESSING. THE OPERATOR IS ADVISED IN SECTION 7: OPERATIONAL PRECAUTIONS AND LIMITATIONS TO 'KEEP ALL PROCESSING MODULE AND SAMPLE HANDLER DOORS CLOSED AND COVERS IN PLACE UNLESS INSTRUCTED OTHERWISE IN A MAINTENANCE OR TROUBLESHOOTING PROCEDURE' WHILE OPERATING THE SYSTEM. OPENING A COVER DURING OPERATION WILL RESULT IN THE GENERATION OF ERROR CODE 0709, RSH STOPPED, COVER OPENED. TROUBLESHOOTING ASSISTANCE FOR ALL ERROR CODES IS PROVIDED IN SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS. >< THE CUSTOMER STATED, TEST RESULTS ON THE ARCHITECT C16000 SYSTEM WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. FOR EXAMPLE, SAMPLE 1 SHOWED THE RESULTS FROM SAMPLE 2, SAMPLE 2 SHOWED THE RESULTS FROM SAMPLE 3, ETC. THE ROBOTIC SAMPLE HANDLER (RSH) GENERATED ERROR CODES 0274, TRAY IN BAY (X) WAS REMOVED BEFORE ACCESS WAS GRANTED, AND 0526, EMPTY TRAY DETECTED IN BAY (X), INDICATING THAT THE OPERATOR HAD INAPPROPRIATELY REMOVED A TRAY FROM THE RSH. CODE 0304, TRANSPORTER ERROR ALSO OCCURRED. THE RSH WENT TO STOPPED STATUS WHILE THE PROCESSING MODULE REMAINED IN RUNNING. THE OPERATOR ORDERED RERUNS FROM THE EXCEPTION SCREEN AND BEGAN TO RUN AGAIN. THE OPERATOR SUSPECTED AN ABERRANT RESULT ON A SAMPLE AND STOPPED THE RUN. REVIEW SHOWED THAT THE RESULTS WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. SAMPLE #7 HAD THE RESULTS FROM SAMPLE #8. THE OPERATOR FOUND 28 CONSECUTIVE SAMPLES THAT WERE AFFECTED. THIRTEEN OF THE RESULTS HAD ALREADY BEEN REPORTED TO THE DOCTOR, BUT NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
885664 | THE CUSTOMER STATED TEST RESULTS ON THE ARCHITECT C16000 SYSTEM WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. FOR EXAMPLE, SAMPLE 1 SHOWED THE RESULTS FROM SAMPLE 2, SAMPLE 2 SHOWED THE RESULTS FROM SAMPLE 3, ETC. THE ROBOTIC SAMPLE HANDLER (RSH) GENERATED ERROR CODES 0274, TRAY IN BAY (X) WAS REMOVED BEFORE ACCESS WAS GRANTED, AND 0526, EMPTY TRAY DETECTED IN BAY (X), INDICATING THAT THE OPERATOR HAD INAPPROPRIATELY REMOVED A TRAY FROM THE RSH. CODE 0304, TRANSPORTER ERROR ALSO OCCURRED. THE RSH WENT TO STOPPED STATUS WHILE THE PROCESSING MODULE REMAINED IN RUNNING. THE OPERATOR ORDERED RERUNS FROM THE EXCEPTION SCREEN AND BEGAN TO RUN AGAIN. THE OPERATOR SUSPECTED AN ABERRANT RESULT ON A SAMPLE AND STOPPED THE RUN. REVIEW SHOWED THAT THE RESULTS WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. SAMPLE #5 HAD THE RESULTS FROM SAMPLE #6. THE OPERATOR FOUND 28 CONSECUTIVE SAMPLES THAT WERE AFFECTED. THIRTEEN OF THE RESULTS HAD ALREADY BEEN REPORTED TO THE DOCTOR, BUT NO IMPACT TO PT MANAGEMENT WAS REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. FOUR CONDITIONS WERE IDENTIFIED THAT CAN LEAD TO A STOP OF THE ROBOTIC SAMPLE HANDLER (RSH) WHEN THE SAMPLE IS PRESENTED TO THE CHEMISTRY MODULE FOR ASPIRATION: REMOVAL OF A TRAY FROM A ROUTINE BAY WHILE THE ACCESS INDICATOR IS AMBER; OPENING AN RSH COVER WHILE THE RSH IS IN THE RUNNING STATE; USER REQUESTING STOP EITHER ON THE RSH KEYPAD OR FROM THE SNAPSHOT SCREEN ON THE SYSTEM CONTROL CENTER; A NON-RECOVERABLE HARDWARE FAILURE OF AN RSH MECHANISM. INVESTIGATION OF THE ISSUE DISCOVERED A TIMING-DEPENDENT DEFECT WHERE THE TRAY-REMOVAL DOES NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CLINICAL CHEMISTRY MODULE. THIS CONDITION CAUSES THE MODULE TO RECEIVE AN ERRONEOUS SAMPLE ARRIVAL FOR TESTS/SAMPLES THAT WERE PREVIOUSLY DELETED WHEN THE RSH NOTIFIED THE C8000II.DLL OF THE TRAY REMOVAL ERROR CONDITION. AFTER A SUBSTANTIAL DELAY THE MODULE SENDS AN EQUALLY ERRONEOUS SAMPLE FREE MESSAGE THAT CAUSES THE RSH TO INDEX THE CARRIER TO THE NEXT POSITION. AT THAT POINT THE SEQUENCE OF SAMPLING IS OFF BY ONE CARRIER POSITION. THE ISSUE WAS IDENTIFIED IN SOFTWARE VERSION 3.00 FOR THE ARCHITECT C16000 SYSTEM, BUT ALL SOFTWARE VERSIONS THAT SUPPORTED AN INTERFACE TO THE C8000 AND/OR THE C16000 AND RSH ARE SUSCEPTIBLE TO THIS ISSUE (V2.00, V2.01, V2.02, V2.03, V2.10, V2.11, V2.20, V2.20 DB UPDATE, V2.60, V2.30, V3.00 AND V3.10). THE ISSUE CAN OCCUR ON THE C8000 AND C16000 SYSTEMS SINCE THE RSH AND THE INTERFACE SOFTWARE (C8000II.DLL) ARE COMMON TO BOTH CONFIGURATIONS. THE ISSUE WAS FOUND TO BE DUE TO A TIMING-DEPENDENT DEFECT WHERE THE TRAY REMOVAL WOULD NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CHEMISTRY MODULE. LABELING: THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS DESCRIPTIONS OF THE USE AND FUNCTION OF THE RSH. STATUS INDICATORS INDICATES THE STATUS OF SAMPLE PROCESSING AND WHEN THE OPERATOR CAN ACCESS SAMPLES. A STEADY AMBER LIGHT INDICATES THAT SAMPLES ARE PROCESSING AND CANNOT BE ACCESSED. OPERATING INSTRUCTIONS; SUBSECTION: SAMPLE MANAGEMENT PROVIDES PROCEDURES FOR LOADING SAMPLES ONTO THE RSH FOR PROCESSING. THE PROCEDURE 'ACCESS A SAMPLE WITH TESTS IN PROCESS' ADVISES THE OPERATOR TO PERFORM THIS PROCEDURE TO SUSPEND A BAY ON THE RSH WHEN IMMEDIATE ACCESS IS NEEDED TO A SAMPLE THAT IS BEING PROCESSED. IMPROPER REMOVAL OF A TRAY FROM THE RSH WILL CAUSE THE SYSTEM TO GENERATE ERROR CODE 0273 (REMOVAL FROM THE PRIORITY BAY) OR 0274 (REMOVAL FROM THE ROUTINE BAY). CONDITIONS FOR THE RSH TO BE STOPPED INCLUDE: THE OPERATOR HAS SELECTED F6- STOP FROM THE SNAPSHOT SCREEN OR THE STOP KEY ON THE SAMPLE HANDLER KEYPAD; ONE OF THE RSH COVERS WAS OPENED WHILE THE SAMPLE HANDLER WAS RUNNING; SAMPLE HANDLER DETECTED A FATAL ERROR WHILE PROCESSING. OPERATIONAL PRECAUTIONS AND LIMITATIONS TO 'KEEP ALL PROCESSING MODULE AND SAMPLE HANDLER DOORS CLOSED AND COVERS IN PLACE UNLESS INSTRUCTED OTHERWISE IN A MAINTENANCE OR TROUBLESHOOTING PROCEDURE' WHILE OPERATING THE SYSTEM. OPENING A COVER DURING OPERATION WILL RESULT IN THE GENERATION OF ERROR CODE 0709, RSH STOPPED, COVER OPENED. | 0
885663 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. FOUR CONDITIONS WERE IDENTIFIED THAT CAN LEAD TO A STOP OF THE ROBOTIC SAMPLE HANDLER (RSH) WHEN THE SAMPLE IS PRESENTED TO THE CHEMISTRY MODULE FOR ASPIRATION: REMOVAL OF A TRAY FROM A ROUTINE BAY WHILE THE ACCESS INDICATOR IS AMBER; - OPENING AN RSH COVER WHILE THE RSH IS IN THE RUNNING STATE; - USER REQUESTING STOP EITHER ON THE RSH KEYPAD OR FROM THE SNAPSHOT SCREEN ON THE SYSTEM CONTROL CENTER; - A NON-RECOVERABLE HARDWARE FAILURE OF AN RSH MECHANISM. INVESTIGATION OF THE ISSUE DISCOVERED A TIMING-DEPENDENT DEFECT WHERE THE TRAY-REMOVAL DOES NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CLINICAL CHEMISTRY MODULE. THIS CONDITION CAUSES THE MODULE TO RECEIVE AN ERRONEOUS SAMPLE ARRIVAL FOR TESTS/SAMPLES THAT WERE PREVIOUSLY DELETED WHEN THE RSH NOTIFIED THE INTERFACE SOFTWARE OF THE TRAY REMOVAL ERROR CONDITION. AFTER A SUBSTANTIAL DELAY THE MODULE SENDS AN EQUALLY ERRONEOUS SAMPLE FREE MESSAGE THAT CAUSES THE RSH TO INDEX THE CARRIER TO THE NEXT POSITION. AT THAT POINT THE SEQUENCE OF SAMPLING IS OFF BY ONE CARRIER POSITION. THE ISSUE WAS IDENTIFIED IN SOFTWARE VERSION 3.00 FOR THE ARCHITECT C16000 SYSTEM, BUT ALL SOFTWARE VERSIONS THAT SUPPORTED AN INTERFACE TO THE C8000 AND/OR THE C16000 AND RSH ARE SUSCEPTIBLE TO THIS ISSUE (V2.00, V2.01, V2.02, V2.03, V2.10, V2.11, V2.20, V2.20 DB UPDATE, V2.60, V2.30, V3.00 AND V3.10). THE ISSUE CAN OCCUR ON THE C8000 AND C16000 SYSTEMS SINCE THE RSH AND THE INTERFACE SOFTWARE ARE COMMON TO BOTH CONFIGURATIONS. THE ISSUE WAS FOUND TO BE DUE TO A TIMING- DEPENDENT DEFECT WHERE THE TRAY REMOVAL WOULD NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CHEMISTRY MODULE. LABELING: THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS DESCRIPTIONS OF THE USE AND FUNCTION OF THE RSH IN SECTION 1. STATUS INDICATORS INDICATES THE STATUS OF SAMPLE PROCESSING AND WHEN THE OPERATOR CAN ACCESS SAMPLES. A STEADY AMBER LIGHT INDICATES THAT SAMPLES ARE PROCESSING AND CANNOT BE ACCESSED. SECTION 5: OPERATING INSTRUCTIONS; SUBSECTION: SAMPLE MANAGEMENT PROVIDES PROCEDURES FOR LOADING SAMPLES ONTO THE RSH FOR PROCESSING. THE PROCEDURE 'ACCESS A SAMPLE WITH TESTS IN PROCESS' ADVISES THE OPERATOR TO PERFORM THIS PROCEDURE TO SUSPEND A BAY ON THE RSH WHEN IMMEDIATE ACCESS IS NEEDED TO A SAMPLE THAT IS BEING PROCESSED. IMPROPER REMOVAL OF A TRAY FROM THE RSH WILL CAUSE THE SYSTEM TO GENERATE ERROR CODE 0273 (REMOVAL FROM THE PRIORITY BAY) OR 0274 (REMOVAL FROM THE ROUTINE BAY). THE STOPPED STATUS FOR THE RSH IS DESCRIBED IN TABLE 1.3. CONDITIONS FOR THE RSH TO BE STOPPED INCLUDE:- THE OPERATOR HAS SELECTED F6- STOP FROM THE SNAPSHOT SCREEN OR THE STOP KEY ON THE SAMPLE HANDLER KEYPAD;- ONE OF THE RSH COVERS WAS OPENED WHILE THE SAMPLE HANDLER WAS RUNNING;- SAMPLE HANDLER DETECTED A FATAL ERROR WHILE PROCESSING. THE OPERATOR IS ADVISED IN SECTION 7: OPERATIONAL PRECAUTIONS AND LIMITATIONS TO 'KEEP ALL PROCESSING MODULE AND SAMPLE HANDLER DOORS CLOSED AND COVERS IN PLACE UNLESS INSTRUCTED OTHERWISE IN A MAINTENANCE OR TROUBLESHOOTING PROCEDURE' WHILE OPERATING THE SYSTEM. OPENING A COVER DURING OPERATION WILL RESULT IN THE GENERATION OF ERROR CODE 0709, RSH STOPPED, COVER OPENED. TROUBLESHOOTING ASSISTANCE FOR ALL ERROR CODES IS PROVIDED IN SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS. >< THE CUSTOMER STATED TEST RESULTS ON THE ARCHITECT C16000 SYSTEM WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. FOR EXAMPLE, SAMPLE 1 SHOWED THE RESULTS FROM SAMPLE 2, SAMPLE 2 SHOWED THE RESULTS FROM SAMPLE 3, ETC. THE ROBOTIC SAMPLE HANDLER (RSH) GENERATED ERROR CODES 0274, TRAY IN BAY (X) WAS REMOVED BEFORE ACCESS WAS GRANTED, AND 0526, EMPTY TRAY DETECTED IN BAY (X), INDICATING THAT THE OPERATOR HAD INAPPROPRIATELY REMOVED A TRAY FROM THE RSH. CODE 0304, TRANSPORTER ERROR ALSO OCCURRED. THE RSH WENT TO STOPPED STATUS WHILE THE PROCESSING MODULE REMAINED IN RUNNING. THE OPERATOR ORDERED RERUNS FROM THE EXCEPTION SCREEN AND BEGAN TO RUN AGAIN. THE OPERATOR SUSPECTED AN ABERRANT RESULT ON A SAMPLE AND STOPPED THE RUN. REVIEW SHOWED THAT THE RESULTS WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. SAMPLE #6 HAD THE RESULTS FROM SAMPLE #7. THE OPERATOR FOUND 28 CONSECUTIVE SAMPLES THAT WERE AFFECTED. THIRTEEN OF THE RESULTS HAD ALREADY BEEN REPORTED TO THE DOCTOR, BUT NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
885660 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. FOUR CONDITIONS WERE IDENTIFIED THAT CAN LEAD TO A STOP OF THE ROBOTIC SAMPLE HANDLER (RSH) WHEN THE SAMPLE IS PRESENTED TO THE CHEMISTRY MODULE FOR ASPIRATION: REMOVAL OF A TRAY FROM A ROUTINE BAY WHILE THE ACCESS INDICATOR IS AMBER; - OPENING AN RSH COVER WHILE THE RSH IS IN THE RUNNING STATE; - USER REQUESTING STOP EITHER ON THE RSH KEYPAD OR FROM THE SNAPSHOT SCREEN ON THE SYSTEM CONTROL CENTER; - A NON-RECOVERABLE HARDWARE FAILURE OF AN RSH MECHANISM. INVESTIGATION OF THE ISSUE DISCOVERED A TIMING-DEPENDENT DEFECT WHERE THE TRAY-REMOVAL DOES NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CLINICAL CHEMISTRY MODULE. THIS CONDITION CAUSES THE MODULE TO RECEIVE AN ERRONEOUS SAMPLE ARRIVAL FOR TESTS/SAMPLES THAT WERE PREVIOUSLY DELETED WHEN THE RSH NOTIFIED THE INTERFACE SOFTWARE OF THE TRAY REMOVAL ERROR CONDITION. AFTER A SUBSTANTIAL DELAY THE MODULE SENDS AN EQUALLY ERRONEOUS SAMPLE FREE MESSAGE THAT CAUSES THE RSH TO INDEX THE CARRIER TO THE NEXT POSITION. AT THAT POINT THE SEQUENCE OF SAMPLING IS OFF BY ONE CARRIER POSITION. THE ISSUE WAS IDENTIFIED IN SOFTWARE VERSION 3.00 FOR THE ARCHITECT C16000 SYSTEM, BUT ALL SOFTWARE VERSIONS THAT SUPPORTED AN INTERFACE TO THE C8000 AND/OR THE C16000 AND RSH ARE SUSCEPTIBLE TO THIS ISSUE (V2.00, V2.01, V2.02, V2.03, V2.10, V2.11, V2.20, V2.20 DB UPDATE, V2.60, V2.30, V3.00 AND V3.10). THE ISSUE CAN OCCUR ON THE C8000 AND C16000 SYSTEMS SINCE THE RSH AND THE INTERFACE SOFTWARE ARE COMMON TO BOTH CONFIGURATIONS. THE ISSUE WAS FOUND TO BE DUE TO A TIMING- DEPENDENT DEFECT WHERE THE TRAY REMOVAL WOULD NOT CANCEL THE SAMPLE-ARRIVAL MESSAGE THAT WAS TO BE SENT TO THE CHEMISTRY MODULE. LABELING: THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS DESCRIPTIONS OF THE USE AND FUNCTION OF THE RSH IN SECTION 1. STATUS INDICATORS INDICATES THE STATUS OF SAMPLE PROCESSING AND WHEN THE OPERATOR CAN ACCESS SAMPLES. A STEADY AMBER LIGHT INDICATES THAT SAMPLES ARE PROCESSING AND CANNOT BE ACCESSED. SECTION 5: OPERATING INSTRUCTIONS; SUBSECTION: SAMPLE MANAGEMENT PROVIDES PROCEDURES FOR LOADING SAMPLES ONTO THE RSH FOR PROCESSING. THE PROCEDURE 'ACCESS A SAMPLE WITH TESTS IN PROCESS' ADVISES THE OPERATOR TO PERFORM THIS PROCEDURE TO SUSPEND A BAY ON THE RSH WHEN IMMEDIATE ACCESS IS NEEDED TO A SAMPLE THAT IS BEING PROCESSED. IMPROPER REMOVAL OF A TRAY FROM THE RSH WILL CAUSE THE SYSTEM TO GENERATE ERROR CODE 0273 (REMOVAL FROM THE PRIORITY BAY) OR 0274 (REMOVAL FROM THE ROUTINE BAY). THE STOPPED STATUS FOR THE RSH IS DESCRIBED IN TABLE 1.3. CONDITIONS FOR THE RSH TO BE STOPPED INCLUDE:- THE OPERATOR HAS SELECTED F6- STOP FROM THE SNAPSHOT SCREEN OR THE STOP KEY ON THE SAMPLE HANDLER KEYPAD;- ONE OF THE RSH COVERS WAS OPENED WHILE THE SAMPLE HANDLER WAS RUNNING;- SAMPLE HANDLER DETECTED A FATAL ERROR WHILE PROCESSING. THE OPERATOR IS ADVISED IN SECTION 7: OPERATIONAL PRECAUTIONS AND LIMITATIONS TO 'KEEP ALL PROCESSING MODULE AND SAMPLE HANDLER DOORS CLOSED AND COVERS IN PLACE UNLESS INSTRUCTED OTHERWISE IN A MAINTENANCE OR TROUBLESHOOTING PROCEDURE' WHILE OPERATING THE SYSTEM. OPENING A COVER DURING OPERATION WILL RESULT IN THE GENERATION OF ERROR CODE 0709, RSH STOPPED, COVER OPENED. TROUBLESHOOTING ASSISTANCE FOR ALL ERROR CODES IS PROVIDED IN SECTION 10: TROUBLESHOOTING AND DIAGNOSTICS. >< THE CUSTOMER STATED TEST RESULTS ON THE ARCHITECT C16000 SYSTEM WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. FOR EXAMPLE, SAMPLE 1 SHOWED THE RESULTS FROM SAMPLE 2, SAMPLE 2 SHOWED THE RESULTS FROM SAMPLE 3, ETC. THE ROBOTIC SAMPLE HANDLER (RSH) GENERATED ERROR CODES 0274, TRAY IN BAY (X) WAS REMOVED BEFORE ACCESS WAS GRANTED, AND 0526, EMPTY TRAY DETECTED IN BAY (X), INDICATING THAT THE OPERATOR HAD INAPPROPRIATELY REMOVED A TRAY FROM THE RSH. CODE 0304, TRANSPORTER ERROR ALSO OCCURRED. THE RSH WENT TO STOPPED STATUS WHILE THE PROCESSING MODULE REMAINED IN RUNNING. THE OPERATOR ORDERED RERUNS FROM THE EXCEPTION SCREEN AND BEGAN TO RUN AGAIN. THE OPERATOR SUSPECTED AN ABERRANT RESULT ON A SAMPLE AND STOPPED THE RUN. REVIEW SHOWED THAT THE RESULTS WERE INCORRECTLY ASSOCIATED WITH THE SUBSEQUENT SAMPLE. SAMPLE #9 HAD THE RESULTS FROM SAMPLE #10. THE OPERATOR FOUND 28 CONSECUTIVE SAMPLES THAT WERE AFFECTED. THIRTEEN OF THE RESULTS HAD ALREADY BEEN REPORTED TO THE DOCTOR, BUT NO IMPACT TO PT MANAGEMENT WAS REPORTED. | 0
1178643 | IT WAS REPORTED THAT THE ULTRASOUND MEASUREMENTS WERE RECEIVED FROM THE CORO 120 MONITOR EVEN THOUGH THE FETUS HAD DIED 30 MINUTES PRIOR. >< GE HEALTHCARE HAS EVALUATED THE UNIT AND IT WAS FOUND TO PERFORM IN ACCORDANCE WITH ITS SPECS. NO MALFUNCTION WAS DETECTED. THE TRACING RECEIVED FROM THE HOSP INDICATED THAT THE CHILD WAS BORN ALIVE AND SUBSEQUENTLY PASSED AWAY AFTER 5 HOURS. | 1
1079830 | LINE DETACHMENT AT SAMPLING BOARD. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION AT THIS TIME. | 0
972696 | INVESTIGATION INTO THIS EVENT FOUND THAT THE TUBE IN QUESTION HAD MULTIPLE BAR CODES ATTACHED. THE DEVICE LABELING CONCERNING PROPER BARCODE USE WAS NOT FOLLOWED. INSTRUMENT MALFUNCTION DID NOT OCCUR. REPAIRS WERE NOT REQUIRED TO RETURN THE ANALYZER TO EXPECTED OPERATION. INCIDENT OCCURRED AS A RESULT OF USER ERROR. >< THE CUSTOMER REPORTED THAT ORTHO PROVUE READ A SAMPLE BARCODE ID INCORRECTLY. A MISS-ASSOCIATION OF RESULT TO THE WRONG SAMPLE COULD RESULT IN INAPPROPRIATE PHYSICIAN ACTION. THE USER IDENTIFIED THE ISSUE DURING SAMPLE IDENTIFICATION AND ABORTED TESTING, PREVENTING ERRONEOUS RESULTS FROM BEING RELEASED. | 0
1046994 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED. THE CUSTOMER INDICATED THAT THE SAMPLE REACTED WEAKLY POSITIVE IN MANUAL GEL METHOD AND NEGATIVE ON THE PROVUE. SAMPLES WITH WEAK ANTIGEN EXPRESSIONS OR LOW ANTIBODY TITERS MAY CONTRIBUTE TO FALSE NEGATIVE REACTIONS ON THE PROVUE. THE CUSTOMER INDICATED THAT DAILY QC WAS ACCEPTABLE AND MAINTENANCE WAS UP TO DATE. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED THAT A PREVIOUSLY KNOWN SAMPLE CONTAINING ANTI-K AND ANTI- ANTI- FYA FAILED TO REACT ON THE ORTHO PROVUE ANALYZER. VISUAL INSPECTION OF THE PROCESSED GEL CARD CONFIRMED THE REACTIONS WERE NEGATIVE. REPEAT TESTING IN MANUAL GEL TEST USING THE SAME LOT OF REAGENTS SHOWED WEAK POSITIVE REACTIVITY WITH THE SAMPLE. NO ERRONEOUS RESULTS WERE REPORTED. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1046988 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED ON SITE AND DETERMINED THAT THE CONNECTIONS FROM THE FLUIDICS SYS WASH/WASTE TUBING WERE SWITCHED WITH THE CONNECTOR FROM THE WASH SOLUTION A CONTAINER, RESULTING IN THE PROBE DIP. THE FE CONNECTED THE TUBINGS APPROPRIATELY TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THE CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS SINCE THIS INCIDENT. INCIDENT IS ISOLATED. >< THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER DRIPPED FLUID INTO THE REAGENT VIALS. THE CUSTOMER INDICATED THAT TESTING COULD NOT BE INITIATED AS A RESULT OF THIS INCIDENT. THE CUSTOMER TOOK THE INSTRUMENT OF USE. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND /OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATABLE BLOOD. | 0
970077 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED. THE OEM ANALYZED THE LOG FILES SUBMITTED FOR INVESTIGATION AND INDICATED THAT THE TIF/PHOTOGRAPHIC IMAGE OF THE AFFECTED GEL CARD APPEARED BRIGHTER THAN EXPECTED. THE EXCESS BRIGHTNESS MAY RESULT IN WEAKER REACTIONS. THE FIELD ENGINEER RELOCATED THE ANALYZER AWAY FROM THE WINDOWS AND READJUSTED THE CAMERA ALIGNMENT AND OPTICS TO RETURN THE ANALYZER TO EXPECTED OPERATION. THE USER DETECTED THE ISSUE, PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. PROVUE MALFUNCTION COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR TO THIS INCIDENT. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED, THE ORTHO PROVUE ANALYZER INTERPRETED POSITIVE REACTIONS AS NEGATIVE DURING ANTIBODY SCREEN TESTING. VISUAL INSPECTION OF CARDS PROCESSED BY PROVUE SHOWED POSITIVE REACTIVITY. THE USER DETECTED THE ISSUE PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. THEREFORE, NO ERRONEOUS TEST RESULTS WERE REPORTED. | 0
970079 | THE CUSTOMER REPORTED, THAT THEY USED REAGENT RED CELLS FOR TESTING ON THE ORTHO PROVUE ANALYZER AFTER OBSERVING THE REAGENTS WERE BROWN IN COLOR. THE CUSTOMER INDICATED THAT QC TESTING WAS PERFORMED AND PASSED, SO THEY CONTINUED TESTING WITH THE REAGENTS. INCIDENT OCCURRED IN 2007. THE CUSTOMER REPORTED, THAT THE SCREENING CELLS WERE INITIALLY OPENED THREE DAYS EARLIER AND APPEARANCE WAS ACCEPTABLE. HOWEVER, THE REAGENTS WERE KEPT ON THE INSTRUMENT FOR APPROX 12-15 HOURS DAILY AND THEN PLACED IN REFRIGERATOR AT 2-8OC. PROBE DRIP WAS NOT OBSERVED ON THE INSTRUMENT. CONTAMINATION OF REAGENTS COULD LEAD TO ERRONEOUS RESULTS. ERRONEOUS RESULTS WERE NOT REPORTED. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. PROBE DRIP WAS NOT OBSERVED ON THE INSTRUMENT. SERVICE WAS NOT REQUIRED SINCE INSTRUMENT MALFUNCTION WAS NOT IDENTIFIED. BASED ON THE AVAILABLE INFORMATION PROVIDED, USER ERROR COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR. SINCE THE SOURCE OF CONTAMINATION IS UNKNOWN, INSTRUMENT MALFUNCTION COULD NOT BE RULED OUT. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THE INCIDENT. | 0
1009150 | THE CUSTOMER INDICATED OBSERVING FALSE POSITIVE REACTIONS WHILE ANALYZING 7 SAMPLES ON THE ORTHO PROVUE ANALYZER. THE VISUAL INSPECTION OF THE PROCESSED GEL CARDS CONFIRMED THE REACTIONS WERE POSITIVE. THE CUSTOMER REPEATED THE TESTING IN THE MANUAL GEL; ALL REACTIONS WERE NEGATIVE. THE USER DETECTED THE ISSUE, PREVENTING ERRONEOUS RESULTS FROM BEING REPORTED. FALSE POSITIVE RESULTS MAY RESULT IN ABO MISCLASSIFICATION AND THE TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE SITE AND REPLACED THE NECESSARY COMPONENTS AND PERFORMED ADJUSTMENTS TO THE READER CAMERA TO RETURN THE ANALYZER TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS REGARDING FALSE POSITIVE RESULTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. | 0
965128 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER LEAKED. THE USER NOTICED THE ISSUE AND ABORTED TESTING. INFO WAS NOT PROVIDED REGARDING RESULTS OF PT TESTING OR POSSIBLE RESULT CODES INTENDED TO PREVENT ERRONEOUS RESULTS FROM BEING REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED; HOWEVER, REPAIRS HAVE RETURNED THE INSTRUMENT TO ITS EXPECTED OPERATION. INFO WAS NOT PROVIDED REGARDING POSSIBLE ERROR MESSAGES/RESULT CODES INTENDED TO SUPPRESS ERRONEOUS TEST RESULTS. PROVUE MALFUNCTION COULD NOT BE RULED OUT. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THE REPAIR. | 0
965126 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED. THE USER NOTICED THE ISSUE AND ABORTED TESTING. INFO WAS NOT PROVIDED REGARDING RESULTS OF PT TESTING OR POSSIBLE RESULT CODES INTENDED TO PREVENT ERRONEOUS RESULTS FROM BEING REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED; HOWEVER, REPAIRS HAVE RETURNED THE INSTRUMENT TO ITS EXPECTED OPERATION. INFO WAS NOT PROVIDED REGARDING POSSIBLE ERROR MESSAGES/RESULT CODES INTENDED TO SUPPRESS ERRONEOUS TEST RESULTS. PROVUE MALFUNCTION COULD NOT BE RULED OUT. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THE REPAIR. | 0
1040407 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED FLUID, RESULTING IN REAGENT CONTAMINATION. ERRONEOUS RESULTS WERE NOT REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND DETERMINED THAT THE PRESSURE WAS OUT OF SPECS. THE FE CLEANED AND ADJUSTED THE PRESSURE REGULATOR TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. | 0
988687 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED FOR THE ELECTRIC SHOCK. AN OCD FIELD ENGINEER ARRIVED AT THE CUSTOMER SITE AND PERFORMED VOLTAGE, CONTINUITY CHECKS AND ELECTRICAL INSPECTION ON THE INSTRUMENT. ALL WERE IN GOOD CONDITION WITH NO NOTICEABLE BREAKS OR EVIDENCE OF SHORTS. THE FE INDICATED THAT THEY OBSERVED TWO WHITE WIRES FOR THE GRIPPER SOLENOID THAT WERE DRESSED INCORRECTLY AND PROTRUDING OUT. THE FE REDRESSED THE WIRES PROPERLY AND FOUND NO EVIDENCE OF CRACKED INSULATION OR SHORTS. NO CURRENT OR VOLTAGE ERRORS WERE FOUND IN THE SYSTEM LOG. THE CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< AN OCD FIELD ENGINEER INDICATED THAT THE CUSTOMER REPORTED EXPERIENCING AN ELECTRIC SHOCK ON THE ORTHO PROVUE ANALYZER WHILE ATTEMPTING TO MOVE THE HANDLER FROM A PARKED POSITION UNDER THE CENTRIFUGE AFTER A PT'S RUN. THE CUSTOMER REPORTED A 'SLIGHT TINGLY FEELING' WITHOUT ANY BURNS OR INJURY IN THE HAND. NO HARM WAS REPORTED AS A RESULT OF THIS INCIDENT. | 0
990013 | THE CUSTOMER STATED THAT FALSE NEGATIVE AXSYM CMV IGG RESULTS WERE GENERATED FOR TWO DIFFERENT PATIENTS. THE FOLLOWING RESULTS ARE FOR PT 1 OF 2. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. AXSYM: 12.4 AU/ML, AXSYM: 114.9 AU/ML, AXSYM: 113 AU/ML . >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROGRESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1036604 | PER INFO RECEIVED BY PHILIPS ON 04/14/08, PHILIPS IS CONSIDERING THAT USE OF MONITORING A FACTOR IN THE DEATH BECAUSE THE CLINICIANS WERE UNAWARE OF THE STATUS OF THE INFANT EVEN THOUGH THEY WERE MONITORING. PHILIPS IS IN THE PROCESS OF OBTAINING MORE INFO CONCERNING THIS EVENT AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SENT ONCE THE INVESTIGATION IS COMPLETED. >< THE MOTHER HAD GESTATIONAL DIABETES AND WAS ADMITTED IN SPONTANEOUS LABOR. TWO HOURLY BLOOD SUGARS WERE TAKEN. FURTHER, CTG WAS IN PROGRESS THROUGHOUT LABOR. AT 7CM 0830, 7CM 1030, 7CM 1310. AT 1530, INVOLUNTARY PUSHING. AT 1551, BABY WAS DELIVERED WITH NO HEARTBEAT. PEDIATRICIANS SUMMONED FOR ASSISTANCE. STILL BORN INFANT. APGAR 0/1, 0/3. CORD PH 6.983(A), PH 7.269(V). THE CORD HAD SLIPPED OVER THE NECK. | 1
1082680 | LINE DETACHMENT AT THE VAMP RESERVOIR. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. | 0
1059519 | IT WAS REPORTED THE TUBING IS COMING AWAY FROM THE CONCERTINA SECTION ON THE VAMP. NO ADVERSE PT EVENT. >< DEVICE HAS NOT YET BEEN RETURNED FOR EVALUATION. | 0
1008617 | WE HAVE NOTICED A PROBLEM IN IMRT FIELDS TRANSFERRED FROM BRAINSCAN OR IPLAN TREATMENT PLANNING TO THE NOVALIS ACCELERATOR AFTER UPGRADE FROM VARIS TO THE ARIA 8.1 SYSTEM. IT IS REQUIRED IN ARIA 8.1 THAT ALL IMRT FIELDS HAVE TO BE VERIFIED THROUGH THE MLC VERIFICATION FUNCTION OF ARIA BEFORE TREATMENT CAN BE ACCEPTED AND PERFORMED. HOWEVER, WHEN CERTAIN IMRT FIELDS ARE 'VERIFIED' WITH THIS FUNCTION, CHANGES MAY HAPPEN IN THE DMLC SEQUENCE. THE MLC VERIFICATION CREATES NEW OPEN SEGMENTS (CA. 10 MM) IN THE BEGINNING OF THE DMLC SEQUENCE. ACCORDING TO CUSTOMER DOSIMETRIC TESTS, THE DOSE ERROR MAY BE NEAR 10% CLOSE TO THE FIELD EDGE. THE MLC VERIFICATION IN ARIA IS REQUIRED FOR BOTH EXISTING PATIENTS (FROM THE PERVIOUS DATABASE) AND NEW PATIENTS. IN BOTH CASES SIMILAR CHANGES MAY OCCUR. >< SOFTWARE TESTS PERFORMED IN OUR DEVELOPMENT FACILITY IN FINLAND ON THE SAME VERSION OF SOFTWARE AS INSTALLED AT THE CUSTOMER SITE RESULTS: THE ROOT CAUSE WAS A SOFTWARE DESIGN ERROR, COMPOUNDED BY THE USER IGNORING A MESSAGE INDICATING AN UNUSUAL EVENT (THE CHANGING OF THE COLLIMATOR LEAF POSITIONS) WAS OCCURRING. CONCLUSIONS: SOFTWARE CONTRIBUTED TO THE EVENT, THE USER ALSO CONTRIBUTED BY IGNORING THE MESSAGE GIVEN (THE MESSAGE IS INFORMATION, NOT A WARNING, HOWEVER), THE EVENT RESULTED IN AN ALTERATION IN THE DOSE DELIVERED TO THE ONE PT BEFORE IT WAS DISCOVERED BY THE CUSTOMER. | 1
1013302 | DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. >< A 104 PLUS 100 BLACK MARKS AND IN BEDDED CONTAMINANTS. | 0
1028953 | THE SCOPE (AN MP20 COMBINED WITH ROOM) ALARMED ALL THE NIGHT ON NON-SUSTAIN VT (YELLOW ALARM) BUT NURSES DIDN'T HEAR AND SEE THE YELLOW ALARMS. ACCORDING TO THE CUSTOMER, THE SYSTEM ALARMED IN RED. THEY NOTICED IT IN THE MORNING BUT IT WAS TOO LATE. THE PT DIED AROUND 13H50 WHILE HE WAS IN REANIMATION ROOM. THE PT WAS TRANSFERRED TO THE REANIMATION ROOM IN THE MORNING. >< THE CUSTOMER REPORTED THAT A PT HAD NON SUSTAIN VT ALARMS OCCURRING THROUGHOUT THE NIGHT IN 2008, BUT STAFF WERE NOT AWARE AS SOME USERS ARE NOT FULLY RESPONSIVE TO YELLOW LEVEL ALARMS. AT ONE POINT, A RED ALARM DID OCCUR THE FOLLOWING MORNING AND STAFF DID RESPOND, TRANSFER THE PT TO A RESUSCITATION ROOM AND ATTEMPTED TO RESUSCITATE THE PT BUT THE PT EXPIRED. A PHILIPS FIELD SERVICE ENGINEER DID GO ONSITE AND CONFIRMED THAT THE DEVICE WAS PERFORMING TO SPECIFICATION. NO PT STRIPS, OR PT DEMOGRAPHIC INFO WAS PROVIDED. NO DEVICE MALFUNCTION OCCURRED. THIS FAILURE TO RECOGNIZE AND RESPOND TO DETERIORATION OF THE PT'S CONDITION IS A HUMAN USE ISSUE, AND THE HOSPITAL IS WORKING WITH ITS STAFF TO ADDRESS THE ISSUE. | 1
1062519 | PATIENT WAS ON A TELEMETRY FLOOR AND WAS BEING MONITORED BY AN AMBULATORY TRANSCEIVER. PATIENT HEARD A SUDDEN LOUD NOISE AND FELT A SLIGHT VIBRATION OF THE AMBULATORY TRANSCEIVER. THE CENTRAL MONITORING STATION NOTED A LOSS OF SIGNAL FROM THE TRANSCEIVER, INDICATING A BATTERY FAILURE, SO A NURSE WENT TO THE PATIENT TO CHANGE THE BATTERY. WHILE THE NURSE WAS CHANGING THE BATTERY, (9V DURACELL PROCELL PRODUCT NUMBER 4F21 G305 17) THE BATTERY 'POPPED' AND THE CARDBOARD BOTTOM CAME OFF THE BATTERY THE INTERNAL CYLINDERS STARTED TO PUSH OUT THE BOTTOM OF THE BATTERY. THE PLASTIC CAP ON THE END OF SOME OF THE CYLINDERS ALSO POPPED OFF OF THE CYLINDER. THE TRANSCEIVER WAS TAKEN OUT OF SERVICE AND THE EVENT REPORTED TO BIOMEDICAL ENGINEERING. ENGINEERING WILL BE LOOKING AT THE TRANSCEIVER TO SEE IF THERE WAS A MALFUNCTION OF THE TRANSCEIVER THAT CAUSED THE BATTERY FAILURE. ENGINEERING BELIEVES THAT THIS IS A BATTERY PROBLEM, BUT FELT A REPORT SHOULD BE COMPLETED IN CASE OTHER FACILITIES ARE REPORTING THIS TYPE OF PROBLEM. THE NURSING STAFF FOUND OTHER BATTERIES THAT POPPED. THESE WERE DISCARDED/USED BATTERIES IN THE DISPOSAL BIN AND WOULD POP IN THE BIN. DURACELL GLOBAL MARKETING WAS NOTIFIED OF THE BATTERY FAILURE AND THEY REMOVED ALL THE BATTERIES WITH LOT #'S 4F15 AND 4F21. THE FAILED BATTERIES WILL BE RETURNED TO THEM SO THAT THEY CAN LOOK INTO THE BATTERY FAILURE. | 1
1019599 | THE CUSTOMER STATED A PATIENT GENERATED A REACTIVE RESULT ON THE AXSYM HBSAG ASSAY BUT CONFIRMED NEGATIVE ON THE AXSYM HBSAG CONFIRMATORY ASSAY. THE SAMPLE WAS TESTED BY PCR AND WAS POSITIVE. THE SAMPLE WAS TESTED ON ANOTHER AXSYM ANALYZER AND WAS REACTIVE ON THE HBSAG ASSAY AND CONFIRMED POSITIVE. THE PATIENT IS A KNOWN HEPATITIS B CARRIER AND HAS THE FOLLOWING SEROLOGICAL PROFILE: NONREACTIVE FOR ANTI-HBS, ANTI-HBC IGM AND HBEAG. REACTIVE FOR ANTI-HBC AND ANTI-HBE. HBV PCR POSITIVE (479 U/ML). NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< EVALUATION: THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROGRESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
982070 | IT WAS REPORTED THAT THE UNIT ALLEGEDLY DID NOT ALARM FOR A PACED RHYTHM. NURSE FOUND PT AND CALLED CODE. PT EXPIRED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1054271 | INVESTIGATION SUMMARY: THE CUSTOMER STATES THE SYSTEM GENERATED DISCREPANT RESULTS WITH THE ARCHITECT BHCG (LN 7K78) FOR ONE PT SAMPLE. THE COMPLAINT INDICATES THAT A BHCG PT SAMPLE WAS COLLECTED IN A GEL TUBE. THE FIELD SVC SPECIALIST CHECKED THE SPEED OF THE CENTRIFUGE, WHICH WAS ACCEPTABLE. THE TEXT ALSO INDICATES THEY HAVE EXPERIENCED ASPIRATION ERRORS IN THE PAST BUT NO OCCURRENCES WERE NOTED ON THE DAY OF THIS INCIDENT. THE CUSTOMER SUPPORT SPECIALIST ADVISED THE CUSTOMER TO REPLACE ALL PROBES ON THE INSTRUMENT THEN PERFORM A PRECISION STUDY. THE COMPLAINT TEXT INDICATES THE CUSTOMER WAS TO STOP USING THIS SYSTEM AS THEY WERE INSTALLING AN ARCHITECT I4000 SYSTEM, SO TROUBLESHOOTING WAS NOT COMPLETED. A REVIEW OF COMPLAINTS LOGGED AGAINST EQUIPMENT SUBTYPE AND LN 7K59 AND LN 7K78 WAS PERFORMED OVER THE TIME PERIOD 10/01/2007 THROUGH 05/19/2008. THE REVIEW DID NOT FIND THAT THE COMPLAINTS IDENTIFIED A DEFICIENCY OF THE PRODUCT OR THAT THE PRODUCT WAS PERFORMING OUTSIDE ITS INTENDED USE. REVIEW OF THE ARCHITECT TOTAL BHCG PACKAGE INSERT INDICATES THAT ADEQUATE LABELING IS PROVIDED RELATED TO THE CUSTOMER'S OBSERVATION. REFER TO THE FOLLOWING SECTIONS: LIMITATIONS OF THE PROCEDURE FOR DIAGNOSTIC PURPOSES, HCG RESULTS SHOULD ALWAYS BE USED IN CONJUNCTION WITH OTHER DATA; E.G., PT'S MEDICAL HISTORY, SYMPTOMS, RESULTS OF OTHER TESTS, CLINICAL IMPRESSIONS. THE RESULTS FROM THIS OR ANY OTHER DIAGNOSTIC KIT SHOULD BE USED AND INTERPRETED ONLY IN THE CONTEXT OF THE OVERALL CLINICAL PICTURE. IF THE HCG LEVEL IS INCONSISTENT WITH, OR UNSUPPORTED BY, CLINICAL EVIDENCE, RESULTS SHOULD BE CONFIRMED BY AN ALTERNATE HCG METHOD. THIS MAY INCLUDE THE QUALITATIVE TESTING OF URINE. THE ABSENCE OF URINARY HCG MAY SUGGEST A FALSELY ELEVATED SERUM RESULT. RESULTS MAY ALSO BE CONFIRMED BY PERFORMING SERIAL DILUTIONS OF THE SAMPLE. USUALLY, BUT NOT ALWAYS, SAMPLES THAT CONTAIN INTERFERING SUBSTANCES EXHIBIT NONLINEAR RESULTS WHEN DILUTED. THE ARCHITECT SYSTEM HAS BEEN VALIDATED FOR ITS INTENDED USE. HOWEVER, ERRORS CAN OCCUR DUE TO POTENTIAL OPERATOR ERRORS AND ARCHITECT SYSTEM TECHNOLOGY LIMITATIONS. SECTION 10 TROUBLESHOOTING AND DIAGNOSTICS, OBSERVED PROBLEMS ERRATIC ASSAY RESULTS (I SYSTEM) PROVIDES SEVERAL POSSIBLE PROBABLE CAUSES AS WELL AS RESOLUTION FOR THE CUSTOMER'S OBSERVATION, WHICH INCLUDE THE PROBE TIP IS BENT OR DAMAGED AND THE SAMPLE CONTAINS FIBRIN CLOTS OR PARTICULATE MATTER. CONCLUSION: THE DEVICE INVOLVED IN THE INCIDENT WAS EVALUATED AND ANALYSIS OF LABELING WAS PERFORMED. THE DEVICE WAS FOUND TO BE PERFORMING OUT OF SPECIFICATION DUE TO COMPONENT/SUBASSEMBLY ISSUES. THE CUSTOMER WAS ADVISED TO REPLACE ALL PROBES ON THE ANALYZER AND PERFORM PRECISION TESTING. THE CUSTOMER DECLINED TO FURTHER TROUBLESHOOT THE ISSUE STATING THAT THEY HAVE STOPPED USING THE I2000 ANALYZER AND HAVE INSTALLED AN I4000 ANALYZER RESOLVING THE ISSUE. THE EXACT CAUSE OF THE MALFUNCTION WAS NOT IDENTIFIED. NO IMPACT TO PT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. THIS IS THE FINAL REPORT. END OF REPORT. >< THE CUSTOMER STATES THAT A PT SAMPLE GENERATED DISCREPANT RESULTS FOR BHCG CONCENTRATION. AN INITIAL TEST WAS PERFORMED USING THE ARCHITECT BHCG ASSAY OBTAINING A RESULT OF 35 MIU/ML. THE SAMPLE WAS REPEATED USING AN ALTERNATE METHOD (RAPID TEST) OBTAINING A NEGATIVE RESULT. THE SAMPLE WAS THEN REPEATED USING THE ARCHITECT BHCG ASSAY OBTAINING A RESULT OF 1.0 MIU/ML RESULT. THE RAPID TEST WAS ALSO REPEATED, YIELDING A NEGATIVE RESULT. THERE WAS NO IMPACT TO PT MANAGEMENT REPORTED DUE TO THIS ISSUE. | 0
955226 | THE HEMO MONITOR CRASHED WHEN SEEKING IN A 'DATA LOG' NOTE IN FULL DISCLOSURE PLAYBACK. THE CIRCUMSTANCES TO CAUSE THIS WERE UNUSUAL IN THAT THERE WAS ONLY DISK SPACE AVAILABLE FOR ONE FULL DISCLOSURE DATA FILE. THE CRASH SCENARIO IS AS FOLLOWS: THEY STARTED, PERFORMED AND CLOSED A STUDY ON PT A. THEY DID SAME FOR PT B AFTER PT A WAS DONE. WHEN THEY STARTED THE STUDY WITH PT MONITORING ON PT B, IT DELETED PT A'S FULL DISCLOSURE FILE. SOME TIME AFTER THEY WERE DONE WITH PT B, THEY RE-OPENED PT A WITHOUT PT MONITORING, I.E. NO FULL DISCLOSURE RECORDING. THEY STARTED FULL DISCLOSURE PLAYBACK AND WHEN THEY SELECTED A 'DATA LOG' NOTE TO SEEK TO THE HEMO MONITOR CRASHED. THE CRASH OCCURRED AS A RESULT OF THE USER ATTEMPTING TO PLAYBACK A FULL DISCLOSURE THAT HAD BEEN PREMATURELY DELETED. THE PREMATURE DELETION WAS THE RESULT OF THE USER FACILITY IGNORING THE MFR'S WARNING THAT THE DISK SPACE WAS GETTING LOW, AND ACTION NEEDED TO BE TAKEN BY THE CUSTOMER TO FREE UP DISK SPACE. ACTION BEING TAKEN BY EMAGEON: AS THIS INCIDENT WAS THE RESULT OF USER ERROR, AND THE APPROPRIATE LABELING EXISTS TO WARN USERS OF THIS POTENTIAL ISSUE, NO ACTION IS BEING TAKEN BY EMAGEON. >< A SYSTEM CRASH OCCURRED ON THE HEARTSUITE HEMODYNAMICS MONITOR WHEN THE USER TRIED TO RETRIEVE A FULL DISCLOSURE THAT HAD PREVIOUSLY BEEN DETECTED. THERE WAS CONCERN THAT THE DELETION OCCURRED PREMATURELY. AFTER FURTHER INVESTIGATION (SEE ADDITIONAL MFG. NARRATIVE BELOW), IT WAS DISCOVERED THAT THE PREMATURE DELETIONS WERE THE RESULT OF THE USER IGNORING THE MANUFACTURER'S WARNINGS THAT DISK SPACE WAS GETTING LOW. | 1
1170067 | WHEN ATTEMPTING TO ACCESS SPACELAB HEALTHCARE INTESYS CLINICAL SUITE -ICS- 'FULL DISCLOSURE' MONITORING INFO SYSTEM. TEN -10- HOURS OF CARDIAC MONITORING INFO WAS NOT AVAILABLE ON ELEVEN -11- PTS IN THE CARDIAC INTENSIVE CARE AND TELEMETRY UNITS. INFO SYSTEMS NOTIFIED SPACELABS OF ISSUE. DATA WAS RESTORED 2 HOURS LATER. | 1
972139 | AN ABBOTT TECHNICAL SERVICE SPECIALIST (TSS) HAD BEEN IN THE LAB THAT DAY TO ADDRESS AN ISSUE RELATED TO CALIBRATIONS NOT HOLDING. THE TSS REPLACED THE BOOM, SYRINGE DRIVE ASSEMBLY, OPTICS AND ANALOG OPTICS BOARD. THE SAMPLES WERE RUN AFTER SERVICE INTERVENTION. THE LAB IS USING REAGENT LOT NUMBER 52180Q100 WITH CALIBRATOR LOT NUMBER 53621Q100. DOCUMENTATION INDICATES THE INSTRUMENT WILL BE SWAPPED OUT FOR ANOTHER AS RESOLUTION FOR ONGOING ISSUES. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE THEOPHYLLINE MONOCLONAL II REAGENT PACKAGE INSERT LN 8A53 (34-4570/R6) AND THE TDX SYSTEM OPERATIONS MANUAL (LN 9520-22/31789-122). SECTIONS ON ASSAY AND INSTRUMENT CALIBRATIONS AND QUALITY CONTROL ENSURE PROPER OPERATION OF THE ASSAY AND ANALYZER. INSTRUMENT DIAGNOSTIC CHECKS AND CALIBRATIONS ENSURE OPTIMAL ANALYZER PERFORMANCE. THIS IS A FINAL REPORT. >< AN ABBOTT FIELD SERVICE REP (FSR) WAS CALLED TO THE CUSTOMER SITE TO INVESTIGATE A COMPLAINT BY THE CUSTOMER OF CALIBRATIONS NOT HOLDING ON THE TDX FPIA ANALYZER. THE FSR REPLACED THE BOOM ASSEMBLY, SYRINGE DRIVE ASSEMBLY, OPTICS ASSEMBLY AND AN ANALOG OPTICS BOARD. THE CUSTOMER THEN RAN PT SAMPLES WITH PATIENT #2 GENERATING AN INITIAL TDX THEOPHYLLINE II ASSAY RESULT OF 176 UMOL/L. THE RESULT WAS QUESTIONED BY THE MEDICAL STAFF AS BEING HIGHER THAN EXPECTED. THE SAMPLE WAS RETESTED AT A VALUE OF 79 UMOL/L. THERE IS NO IMPACT TO PT MANAGEMENT REPORTED. AN INSTRUMENT SWAP OUT WAS DISCUSSED WITH THE CUSTOMER. | 0
997538 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER VISITED THE SITE AND REPLACED THE NECESSARY FLUIDICS COMPONENTS AND VERIFIED OPERATIONS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER REPORTED THAT THE PROBE LEAKED FLUID FROM THE HANDLER AND THE DILUTION CUP OVERFLOWED ON THE ORTHO PROVUE ANALYZER. ALL TESTING WAS ABORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. THE USER DETECTED THE ISSUE; PREVENTING ERRONEOUS RESULTS FORM BEING REPORTED. | 0
1001469 | THE MOST PROBABLE CAUSE FOR THE BROKEN TUBE IS THE CUSTOMER MISHANDLING THE WASTE SENSOR. THIS IS A KNOWN ISSUE AND A NEWER DESIGNED WASTE SENSOR WAS INSTALLED ON THIS INSTRUMENT IN EARLY 2008 TO PREVENT THIS FROM HAPPENING IN THE FUTURE. BECAUSE OF THE REMOTE POSSIBILITY OF INJURY RESULTING FROM A SLIP AND FALL EVENT CAUSED BY WASTE FLUID OVERFLOW, THE FIRM IS IN THE PROCESS OF IMPLEMENTING AN IMPROVED WASTE SENSOR DESIGN AS A REPLACEMENT TO THE CURRENT SENSOR. THIS NEW WASTE SENSOR DESIGN SHOULD ELIMINATE THE OBSERVED FAILURE MODE. COMPLETE FIELD IMPLEMENTATION OF THE NEW SENSOR DESIGN TO BE COMPLETED BY 05/21/2008. THE NEWER WASTE SENSOR DESIGN WAS INTRODUCED TO THE MFG FLOOR ON 02/08/2007. A FINAL REPORT WILL BE SENT PENDING THE COMPLETION OF THE FIELD IMPLEMENTATION OF THE NEW WASTE SENSOR. >< A CUSTOMER SITE REPORTED WASTE FLUID HAD OVERFLOWED FROM THE WASTE CARBOY ON A BENCHMARK XT INSTRUMENT ONTO THE FLOOR WITHOUT AN ALARM. NO ONE WAS INJURED AND PT CARE WAS NOT AFFECTED. THE FIELD SERVICE ENGINEER EVALUATED THE UPGRADED WASTE SENSOR AND FOUND THE TUBE TO BE BROKEN FROM MISHANDLING PREVENTING THE SENSOR FROM FUNCTIONING PROPERLY. THE FSE ( FIELD SERVICE ENGINEER) INSTALLED A NEWER DESIGNED WASTE SENSOR WHICH HAS THE TUBE COUNTER SUNK PREVENTING LATERAL MOVEMENT WHICH PREVENTS THE CUSTOMER FROM BREAKING THE TUBE IF MISHANDLED. WHILE THERE IS ALWAYS A POSSIBILITY THAT A SLIP AND FALL EVENT DUE TO A FLUID SPILL COULD RESULT IN DEATH OR SERIOUS INJURY, ALL EVIDENCE INDICATES THAT THE POSSIBILITY OF DEATH OR SERIOUS INJURY IS REMOTE. | 0
1078989 | REPORTEDLY, SOMETIMES SVO2 DOESN'T WORK EITHER PARAMETERS ARE NOT CONSISTENT, SOMETIMES THEY DON'T GET SVO2. TRIED GOOD CABLES BUT NO CHANGE. TRIED DIFFERENT VIG WITH SAME CABLES, WORKS FINE. >< EVAL SUMMARY: CO NOT WORKING - ERROR MESSAGE 'ALERT: ELECTRICAL INTERFERENCE' AND 'FAULT SVO2: GAIN' ERROR. | 1
1252367 | PATIENT HAD A FLEXI SEAL FECAL MANAGEMENT SYSTEM THAT LED TO RECTAL BLEEDING REQUIRING VASOPRESSORS, BLOOD TRANSFUSIONS, FFP AND DDAVP. | 0
1034139 | WENT TO PRIME ARTERIAL LINE TUBING AND IT BEGAN TO LEAK AT THE VAMP. OBSERVED TO BE COMING OFF. >< DEVICE HAS NOT BEEN RETURNED FOR EVAL AT THIS TIME. | 0
898720 | THE MOST PROBABLE CAUSE FOR THE BROKEN TUBE IS THE CUSTOMER MISHANDLING THE WASTE SENSOR. THIS IS A KNOWN ISSUE AND A NEWER DESIGNED WASTE SENSOR HAS BEEN INSTALLED ON THE INSTRUMENT TO PREVENT THIS FROM HAPPENING IN THE FUTURE. BECAUSE OF THE REMOTE POSSIBILITY OF INJURY RESULTING FROM A SLIP AND FALL EVENT CAUSED BY WASTE FLUID OVERFLOW, THE FIRM IS IN THE PROCESS OF IMPLEMENTING AN IMPROVED WASTE SENSOR DESIGN AS A REPLACEMENT TO THE CURRENT SENSOR. THIS NEW WASTE SENSOR DESIGN SHOULD ELIMINATE THE OBSERVED FAILURE MODE. THE WASTE SYSTEM ON THIS INSTRUMENT WAS UPGRADED WITH THE NEW SENSOR DESIGN IN 2007. COMPLETE FIELD IMPLEMENTATION OF THE NEW SENSOR DESIGN TO BE COMPLETED BY THE NEXT FIVE MONTHS. THE NEWER WASTE SENSOR DESIGN WAS INTRODUCED TO THE MFG FLOOR EIGHT MONTHS EARLIER, A FINAL REPORT WILL BE SENT PENDING THE COMPLETION OF THE FIELD IMPLEMENTATION OF THE NEW WASTE SENSOR. >< A CUSTOMER SITE REPORTED WASTE FLUID HAD OVERFLOWED FROM THE WASTE CARBOY ON A BENCHMARK LT INSTRUMENT ONTO THE FLOOR WITHOUT AN ALARM. NO ONE WAS INJURED AND PT CARE WAS NOT AFFECTED. THE FIELD SVC ENGINEER EVALUATED THE UPGRADED WASTE SENSOR AND FOUND THE TUBE TO BE BROKEN FROM MISHANDLING, PREVENTING THE SENSOR FROM FUNCTIONING PROPERLY. THE FSE (FIELD SERVICE ENGINEER) INSTALLED A NEWER DESIGNED WASTE SENSOR WHICH HAS THE TUBE COUNTER SUNK, PREVENTING LATERAL MOVEMENT WHICH PREVENTS THE CUSTOMER FROM BREAKING THE TUBE IF MISHANDLED. WHILE THERE IS ALWAYS A POSSIBILITY THAT A SLIP AND FALL EVENT DUE TO FLUID SPILL COULD RESULT IN DEATH OR SERIOUS INJURY, ALL EVIDENCE INDICATES THAT THE POSSIBILITY OF DEATH OR SERIOUS INJURY IS REMOTE. | 0
1022528 | GE WAS REQUESTED TO VERIFY THAT ALARMS WERE WORKING AT THE CCU CIC CENTRAL STATION. AFTER BRIEF INVESTIGATION, IT WAS DISCOVERED THAT THE SPEAKERS HAD BEEN UNPLUGGED FROM THE CIC, AND THEREFORE NO AUDIBLE ALARMS WERE ABLE TO BE HEARD. A PT DEATH WAS REPORTED. THE SPEAKERS WERE RECONNECTED TO THE SYSTEM AND THE NURSE MGR WAS MADE AWARE OF THE FINDINGS. >< THE LOG FILES FROM CIC AT THE TIME OF EVENT OCCURENCE WERE NOT AVAILABLE, SO THE ALARM VOLUME SETTING CANNOT BE DETERMINED AT THIS TIME. IT WAS DETERMINED THAT THE FACILITY WAS USING PREMIERE MULTIMEDIA SPEAKER SYSTEM MODEL PDM/5. IF THE EXTERNAL SPEAKERS ARE DISCONNECTED FROM THE CIC, AN INTERNAL SPEAKER IS ACTIVATED AS A BACKUP AND A WARNING MESSAGE IS DISPLAYED INDICATING THAT THE EXTERNAL SPEAKERS ARE UNPLUGGED. THE FACILITY DID REPORT THAT WHEN THE EXTERNAL SPEAKERS WERE UNPLUGGED, THE ERROR MESSAGE INDICATING SPEAKERS ARE UNPLUGGED APPEARED, WHICH SUGGESTS THAT INTERNAL SPEAKERS WERE ACTIVE. GE INVESTIGATION TEAM SUSPECTS THAT THE VOLUME LEVEL OF INTERNAL SPEAKERS WAS SET TOO LOW FOR THE USER TO HEAR. THE CIC SERVICE MANUAL CONTAINS THE FOLLOWING NOTED: 'EXTERNAL SPEAKERS ARE CONNECTED DURING NORMAL OPERATION. AN ALERT MESSAGE DISPLAYS WHEN NO EXTERNAL SPEAKER CONNECTION IS PRESENT. THIS IS TRUE EVENT IF THE CIC PRO CENTER IS USED AS A MIRRORED CIC PRO CENTER AND ITS AUDIBLE ALARM VOLUME IS SET TO OFF'. | 1
997520 | RTI WAS NOTIFIED OF A POSITIVE HEPATITIS C SCREENING TEST IN 2007. THE RECIPIENT HAD TWO ALLOGRAFT IMPLANTS IN 2004 THAT ARE ASSOCIATED WITH A RECALL OF BTS TISSUE. THE SURGEON WAS CONTACTED TO OBTAIN CONFIRMATORY RESULTS. THE SURGEON WAS UNAWARE THAT THE PT TESTED POSITIVE FOR HEPATITIS C. UNABLE TO LOCATE INFO ON THE PRIMARY PHYSICIAN. >< THE GRAFTS REMAIN IMPLANTED. REVIEW OF INTERNAL RECORDS, DONOR RECORDS, SEROLOGICAL RESULTS, AND MANUFACTURING RECORDS INCLUDING QA/QC REVIEWS. THE GRAFT PASSED ALL RELEASE REQUIREMENTS AT THE TIME OF DISTRIBUTION. NO DEVIATIONS WERE NOTED. THE GRAFT UNDERWENT DEMINERALIZATION AND IRRADIATION WHICH ARE BOTH VALIDATED METHODS OF STERILIZATION TO ELIMINATE THE POTENTIAL OF DISEASE TRANSMISSION. IT IS LIKELY THAT THE RECIPIENT COULD HAVE CONTRACTED HEPATITIS C, IF TRULY POSITIVE FOR THE VIRUS, FROM A SOURCE OTHER THAN THE IMPLANT. THIS REPORT IS ASSOCIATED WITH REPORT #302719998-2007-011 AND PREVIOUSLY REPORTED ON 11/22/07. | 0
1054847 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND DETERMINED THAT THE CONNECTOR TO THE WASTE BOTTLE WAS NOT CONNECTED PROPERLY, RESULTING IN FLUID LEAKING. THE FLUID ENTERED INTO THE PRESSURE TRANSDUCER, RESULTING IN THE COMPONENT DAMAGE. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. >< THE CUSTOMER INDICATED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER WAS NOT DISPENSING THE CORRECT AMOUNT OF FLUID INTO THE MTS GEL CARDS. NO ERRONEOUS RESULTS WERE REPORTED. INCORRECT OR NO DISPENSE CAN LEAD TO ERRONEOUS TEST RESULTS AND TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1082663 | SILICONE OIL WAS OBSERVED INSIDE OF THE TUBING. >< DEVICE HAS NOT BEEN RETURNED FOR EVALUATION AT THIS TIME. | 0
1037680 | THE ARCTIC SUN 2000 TEMPERATURE MANAGEMENT SYSTEM WAS USED AS A MONITOR DURING THERAPEUTIC HYPOTHERMIA IN A PATIENT FOLLOWING AN OUT-OF-HOSPITAL CARDIAC ARREST. THE FOLEY TEMPERATURE PROBE, WHICH USED FOR MONITORING THE PATIENT'S TEMPERATURE, WAS BROKEN. THE END CONNECTOR PIECE WAS FOUND TO BE MISSING. THERE DID NOT SEEM TO BE ANY REPLACEMENT CABLES AVAILABLE SO THE CABLE WAS USED WITHOUT THE CONNECTOR. THE WIRES WERE INSERTED INTO THE ARTIC SUN MACHINE, AND AS A RESULT, THE PATIENT'S TEMPERATURE WAS NOT ACCURATE. THE PATIENT WAS OVERCOOLED TO A TEMP OF 29.9 DEGREES. HE BECAME HYPOTENSIVE BUT DID NOT HAVE BRADYCARDIA OR OSBORNE WAVES. THE THERAPEUTIC HYPOTHERMIA WAS DISCONTINUED AND THE PATIENT WAS RE-WARMED. | 1
1018448 | PATIENT IN CATH LAB. BP CUFF ON RIGHT ARM, SET TO REPEAT EVERY 3 MINUTES. RN NOTICED BP CUFF WAS NOT WORKING AFTER INITIAL BP. BP CUFF DISCONNECTED FROM TUBING AND NEW CUFF PLACED ON PATIENT'S LEG. AT END OF THE CASE NURSE NOTED REDNESS AROUND RIGHT ARM. ONE OF THE STAFF NURSES SAID RIGHT ARM CUFF STILL RETAINED PRESSURE AT END OF CASE. THE PATIENT HAD REDNESS FROM HAND TO JUST PROXIMAL TO THE ELBOW THAT CORRELATES WITH THE DISTAL END OF THE BP CUFF. PULSES INTACT AT END OF CASE BUT PATIENT COMPLAINED OF NUMBNESS AND TINGLING OF ARM WHICH WAS DECREASING BUT STILL PRESENT AT DISCHARGE (JUST OVER 24 HOURS AFTER THE PROCEDURE) FROM THE HOSPITAL. PATIENT ALSO HAD DECREASED MUSCLE STRENGTH. | 1
1082665 | IT WAS REPORTED THAT THE DEVICE HAD SEPARATED FROM Z SITE. >< EVALUATION: CUSTOMER REPORT WAS CONFIRMED. VAMP TUBING WAS FOUND TO BE COMPLETELY DETACHED FROM SOLVENT BOND JOINT WITH SAMPLE SITE (Z-SITE). AT SOME PART OF TUBING BOND AREA INDICATION OF BONDING SOLVENT WAS PRESENT BUT APPEARED TO BE THIN. AND IT DID NOT APPEAR THAT THE ADHESIVE CHEMICALLY ETCHED TO THE SURFACE OF THE Z-SITE TUBE HOUSING. TUBING O.D. WAS .139' WHICH WAS ON THE LOWER SIDE OF THE SPEC. (141'+/-.002' PER DRAWING 111467-001 REV. AD). | 0
1013225 | DEVICE HAS NOT BEEN RECEIVED FOR AN EVALUATION AT THIS TIME. >< LEAKAGE WAS OBSERVED FROM THE CONNECTION BETWEEN THE ONE WAY STOP COCK OF THE VAMP PLUS AND THE TUBING. | 0
1013226 | DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AT THIS TIME. >< IT WAS REPORTED THAT THE INTERNAL RUBBER OF THE SYRINGE UNFASTENED. | 0
967281 | DRAEGER MEDICAL SYSTEMS, INC. HAS RECEIVED INFORMATION FROM A CUSTOMER THAT WHEN EDITING ECG WITH OUR INFINITY MEGACARE SOFTWARE APPLICATION, THE CORRECT TEXT IS NOT ALWAYS SHOWN ON THE PRINT OUT. THERE IS A DIFFERENCE FROM WHAT IS DISPLAYED ON THE SCREEN TO WHAT IS PRINTED OUT. THERE ARE DIFFERENT FAULTS THAT OCCUR, HOWEVER, THE NEW SIGNATURE/HEADLINE PRINTED TOGETHER WITH THE OLD TEXT IS A SERIOUS VARIANT. THIS INTERMITTENT FAULT HAS OCCURRED DURING THE LAST 18 MONTHS AND THE LAST OCCURRENCE WAS IN 2007. THE CUSTOMER IS CONCERNED THAT AN INCORRECT INTERPRETATION MAY OCCUR, BUT HAS REPORTED THAT THERE IS NO KNOWN CASE OF THIS OCCURRING. ALL WRONG PRINTOUTS HAVE BEEN DISCOVERED, AND THE STAFF CAREFULLY CONTROLS THE PRINTOUTS BY EXAMINING EVERY PRINTOUT. NO PATIENT INJURY WAS REPORTED. >< THE COMPLAINT IS CURRENTLY UNDER INVESTIGATION. | 1
879342 | THE INCIDENT ORIGINATED OUTSIDE THE U.S. AND THEREFORE REQUIRED REVIEW BY BOTH FOREIGN DIVISIONS OF SYSMEX CORPORATION PRIOR TO U.S. NOTIFICATION. THE TECHNOLOGIST'S INJURY DID NOT RESULT FROM A FAILURE OF THE SYSMEX CA-1500 TO OPERATE CORRECTLY, A FAILURE OF THE DEVICE TO MEET SPECIFICATIONS NOR WAS IT THE RESULT OF A DESIGN DEFECT OR FAILURE OF A SAFETY FEATURE OR SAFETY COMPONENT OF THE DEVICE. FROM THE AVAILABLE INFORMATION, THE TECHNOLOGIST'S INJURY APPEARS TO HAVE OCCURRED ACCIDENTALLY WHILE PERFORMING ROUTINE ANALYZER MAINTENANCE AS DESCRIBED IN THE CA-1500 OPERATOR'S MANUAL. BASED ON OUR ASSESSMENT, THE ANALYZER OPERATOR'S INJURIES DO NOT QUALIFY AS 'SERIOUS'., AS THEY DID NOT APPEAR TO BE LIFE-THREATENING, DID NOT RESULT IN PERMANENT IMPAIRMENT OR DAMAGE TO BODY FUNCTION OR STRUCTURE, AND DID NOT NECESSITATE MEDICAL OR SURGICAL INTERVENTION TO PRECLUDE PERMANENT IMPAIRMENT OF BODY FUNCTION OR PERMANENT DAMAGE TO A BODY STRUCTURE. SYSMEX REPORTS THIS INJURY FOLLOWING FDA'S NOVEMBER 12, 2002 GUIDANCE DOCUMENT: NEEDLESTICKS: MEDICAL DEVICE REPORTING GUIDANCE FOR USER FACILITIES, MANUFACTURERS AND IMPORTERS, WHERE ADMINISTRATION OF ANTI-RETROVIRAL DRUGS AND PRESCRIPTION MEDICATION WOULD BE CLASSIFIED AS 'MEDICAL INTERVENTION' AND REQUIRE AN MDR REPORT. >< WHILE CLEANING THE SAMPLE PROBE (PIERCER NEEDLE) OF A CA-1500 AUTOMATED COAGULATION ANALYZER AS PART OF ROUTINE WEEKLY INSTRUMENT MAINTENANCE, A TECHNOLOGIST EXPERIENCED A NEEDLE STICK. THE INSTRUMENT POWER WAS OFF DURING THE MAINTENANCE, BUT THE TECHNOLOGIST COULD NOT VERIFY WHETHER THE AUTOMATED MAINTENANCE SEQUENCE PRIOR TO POWERING OFF THE ANALYZER HAD BEEN COMPLETED. THE TECHNOLOGIST IMMEDIATELY CONTACTED A PHYSICIAN, AND BEGAN TREATMENT WITH ANTI-RETROVIRAL DRUGS. TWO HRS LATER, THE TECHNOLOGIST CONTACTED AN OUTPATIENT CENTER, BECAUSE SHE WAS FEELING NAUSEOUS. AS A RESULT, SHE WAS PUT ON A ONE MONTH REGIMEN OF THE ANTI-NAUSEA/ ANTI VOMITING DRUG, PREIMPERAN, FROM SANOFI AVENTIS. | 0
1149859 | DEVICE NOT RETURNED. >< REPORTEDLY, NURSE SPOTTED SPARKS AND SMOKE COMING OUT OF IT. NO PATIENT INJURY REPORTED. | 1
1051961 | REPORTEDLY, SMOKED. NO PATIENT INJURY REPORTED. >< DEVICE NOT RETURNED. | 1
1051960 | REPORTEDLY, INCORRECT, DRIFTS-OXYGEN PROFILE DOES NOT WORK. NO PATIENT INJURY REPORTED. >< DEVICE NOT RETURNED. | 1
1167333 | PT ADMITTED FOR CARDIAC RHYTHM MONITORING. PHYSICIAN UNABLE TO RETRIEVE ARCHIVE DATA FROM SPACELAB HEALTHCARE INTESYS CLINICAL SUITE - ICS- 'FULL DISCLOSURE' MONITORING INFO SYSTEM. INFO SYSTEMS - IS- WAS ABLE TO RESTORE DATA, BUT UNABLE TO CORRECT MALFUNCTION. IS DETERMINED THE 'FULL DISCLOSURE' COMPONENT OF THE MONITORING SYSTEM IS ON A RUNNING CLOCK. ONCE THE DATA IS RETRIEVED, IT MUST BE REVIEWED IMMEDIATELY, OR THE DATA WILL BE LOST OVER TIME. THE DATA AT THE BEGINNING OF THE MONITORING PERIOD WAS LOST. | 1
1006882 | THE CUSTOMER STATED THAT AN AXSYM DIGOXIN III RESULT OF 3.61 NG/ML WAS GENERATED FOR A PATIENT SAMPLE THAT TESTED AT 1.57 NG/ML AT A REFERENCE LABORATORY (OLYMPUS DIGOXIN METHOD). THE CUSTOMER REPORTED THE RESULT FROM THE REFERENCE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1001563 | THE CUSTOMER INDICATED THAT THE PATIENT DID NOT SURVIVE THE CARDIAC ARREST. ADDITIONAL CLARIFICATION / INVESTIGATION PERFORMED ON BEHALF OF LIFESYNC CORPORATION WITH NURSES THAT WERE ON SITE AT THE TIME OF THE INCIDENT INDICATED THAT THE LEADWEAR BROKE DURING CHEST COMPRESSIONS, THE PATIENT WAS GRAVELY ILL AT THE TIME OF THE INCIDENT, AND THE OUTCOME OF RESUSCITATION WAS NOT IMPACTED IN ANY WAY BY THE LEADWEAR BREAK. IN THIS CASE, A DEVICE FAILURE OCCURRED, BUT WAS NOT RELATED TO THE OUTCOME OF TREATMENT TO THE PATIENT. DEVICE FAILURE WAS NOT RELATED TO THE DEATH OF THE PATIENT. >< A LIFESYNC CORPORATION SALES REPRESENTATIVE WAS RECORDING FEEDBACK OBTAINED DURING A MEETING WITH A CUSTOMER. THE CUSTOMER INDICATED TO THE LIFESYNC REPRESENTATIVE THAT THEY WERE ATTEMPTING TO RESUSCITATE A PATIENT THAT WAS IN CARDIAC ARREST, WHILE PERFORMING CPR, THE LEADWEAR BROKE. THE CUSTOMER INDICATED THAT THE BREAK IN THE LEADWEAR MADE IT DIFFICULT TO MONITOR THE PATIENT DURING THE CARDIAC ARREST. MONITOR OF THE PATIENT RESUMED WHEN DEFIBRILLATOR PADS WERE PLACED. BROKEN LEADWEAR DURING RESUSCITATION HAS THE POTENTIAL TO LEAD TO A DELAY IN TREATMENT DURING AN EMERGENCY PROCEDURE. | 1
1060863 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH CELL #12, HETEROZYGOUS S CELL, OF CAPTURE-R READY ID, LOT ID100 WHEN TESTING A PATIENT SAMPLE WITH A HISTORY OF ANTI-E AND ANTI-S ON THE ECHO. NO ADVERSE REACTIONS OCCURRED AS A RESULT OF THE UNEXPECTED NEGATIVE REACTIONS. CUSTOMER WAS PERFORMING VALIDATIONS WITH A CURRENT PATIENT SAMPLE. >< REACTIVITY OF THE S ANTIGEN WAS CONFIRMED ON RETENTION CAPTURE-R READY-ID, LOT ID100. THE CUSTOMER DID NOT RETURN PRODUCT OR SAMPLE FOR INVESTIGATION. IT IS NOT POSSIBLE TO RULE OUT THE SAMPLE AS A CAUSE OF THE EVENT. | 0
1099063 | THE TWO RETURNED CUSTOMER SAMPLES AND IN-HOUSE DONOR SAMPLES OF VARIOUS RH PHENOTYPES INCLUDING WEAK D CELLS, WERE TESTED ON AN IN-HOUSE ECHO FOR WEAK D ASSAY USING RETENTION CAPTURE-R SELECT PLATES, LOT SC079 AND SC083, WITH RETENTION ANTI-D SERIES 4, ANTI-D SERIES 5. THE SAMPLES EXHIBITED NEGATIVE WEAK D AND ALL IN-HOUSE DONOR SAMPLES EXHIBITED THE EXPECTED REACTIVITY. THE RETURNED SAMPLES WERE TESTED ON IN-HOUSE ECHO FOR GROUP ASSAY. NTD WAS OBSERVED. WEAK REACTIVITY WAS OBSERVED WITH ANTI-D SERIES 4,AND NO REACTIVITY WAS OBSERVED WITH ANTI-D SERIES 5. HA TESTS WERE PERFORMED WITH THE SAMPLES WITH RETENTION ANTI-D SERIES 4, AND ANTI-D SERIES 5, AND WITH OTHER ANTI-D MANUFACTURERS' BLOOD GROUPING REAGENTS. NO REACTIVITY WAS OBSERVED IN ALL PHASES OF TESTING WITH ALL ANTI-D REAGENTS TESTED. THE SAMPLES TYPED AS D-. THE ECHO PROBE WAS EXAMINED AND THERE IS NO EVIDENCE OF THE LEAK TO CAUSE CARRYOVER. TESTED B-NEGATIVE DONOR SEGMENTS ON AN IN-HOUSE ECHO WITH RETENTION ANTI-A, ANTI-B, ANTI-D SERIES 4, ANTI-D SERIES 5, AND MONOCLONAL CONTROL. THE EXPECTED REACTIVITY WAS OBSERVED. THE CUSTOMER'S COMPLAINT APPEARS TO BE SAMPLE-RELATED. >< CUSTOMER REPORTED AN UNEXPECTED POSITIVE REACTION ON THE ECHO FOR A PATIENT SAMPLE WITH AN HISTORY OF RH NEGATIVE. 1+ RESULTS WERE OBSERVED WITH ANTI-D SERIES 4. THIS SAMPLE WAS NEGATIVE WITH ANTI-D SERIES 5. | 0
1138304 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIVITY ON A PATIENT SAMPLE TESTED ON THE ECHO. THE PATIENT HAD A PREVIOUS HISTORY OF ANTI-E AND ANTI-JKA THAT WAS CONFIRMED IN GEL AND BY A REFERENCE LAB TESTING. THE SAMPLE WAS NEGATIVE, BUT SHOULD HAVE BEEN POSITIVE WITH CELL 7 (E-JKA+JKB+). THE OTHER CELLS REACTED AS EXPECTED. >< REACTIVITY OF THE JKA ANTIGEN WAS CONFIRMED ON RETURNED AND RETENTION CRRID, LOT ID102; THIS REAGENT WAS USED BY THE CUSTOMER. THE RETURNED SAMPLE WAS TESTED ON AN IN-HOUSE ECHO. POSITIVE REACTIVITY WAS OBSERVED. A SERVICE VISIT WAS MADE. NO OUT OF TOLERANCE RESULTS WERE OBSERVED AND NO ERRORS OCCURRED. THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. | 0
1055470 | REACTIVITY OF THE C, E, K AND V ANTIGENS WERE CONFIRMED ON RETENTION CRRID, LOT ID102. THIS LOT WAS USED ON THE ECHO BY THE CUSTOMER. THE RETURNED PATIENT'S SAMPLE EXHIBITED NO REACTIVITY WITH CELLS # 1, 2, 5, 7, 9, 10, 11, 12 ON RETENTION CRRID, LOT ID102. SLIGHT REACTIVITY WAS OBSERVED WITH CELLS # 3, 4, 6, 8, AND 13. THE RETURNED SAMPLE WAS TESTED IN HEMAGGLUTINATION TESTS WITH SELECTED CELLS FROM PANOCELL-20, LOT 13600, USING IMMUADD AS THE POTENTIATOR. BECAUSE OF PREVIOUS HISTORY OF COLD ANTIBODIES, A ROOM TEMPERATURE TESTING WAS NOT INCLUDED. ANTI-C, -E AND -V WERE NOT DETECTED AT THIS TIME. THE PATIENT'S ANTI-K WAS DETECTED. THE RETURNED SAMPLE WAS TESTED WITH PANOSCREEN I, II AND III, LOT 14622 AT ROOM TEMPERATURE INCUBATION; ONE SET WAS INCUBATED AT 15' 37 DEGREES C AND ANTI-IGG WAS ADDED. THE SECOND SET, INSTEAD OF ANTI-IGG, SALINE WAS ADDED. CELL I (K+K+) EXHIBITED 3+ AT ROOM TEMPERATURE AND 1+ AT ANTI-IGG. CELLS II AND CELL III WERE NONREACTIVE IN ALL PHASES OF TESTING. ALL THREE REAGENT RED CELLS WERE NONREACTIVE WHEN SALINE WAS ADDED. THE EVENT APPEARS TO BE RELATED TO THE NATURE OF THE SAMPLE. >< AN UNEXPECTED NEGATIVE REACTION WAS REPORTED FOR THE ECHO. A PATIENT SAMPLE WAS HISTORICALLY POSITIVE FOR ANTI-C, -K, -V, AND -E. A SELECTED CELL GEL PANEL WAS PERFORMED ON THE PATIENT AND THE SAMPLE WAS 1+ WITH A C+ CELL, 3+ WITH A K+ CELL, AND NEGATIVE WITH AN E+ CELL. NO V+ CELL WAS AVAILABLE. THE PATIENT WAS TRANSFUSED. A NEW SAMPLE WAS DRAWN AND TESTED ON THE ECHO USING CAPTURE-R READY-ID (CRRID). THE SAMPLE WAS 3+ WITH CELL 3 (C+, E+, K+), 2+ WITH CELL 4 (C+, V+), 3+ WITH CELL 8 (C+, K+), 1+ WITH CELL 9 (C+), 1+ WITH CELL 10 (C+), 2+ WITH CELL 11 (C+), 1+ WITH CELL 12 (C+), AND 3+ WITH CELL 14 (K+). THE SAMPLE WAS NEGATIVE WITH CELLS 5, 6, 7, AND 13, ALL OF WHICH ARE C+. | 0
993113 | REVIEW OF THE RESULTS FILES FROM THE ECHO SHOWED THAT THE REACTIONS APPEAR AS 1+ AND NEGATIVE WITH WELL SCORES OF 17 AND 0. ADVISED CUSTOMER THAT ACCORDING TO THE OPERATOR MANUAL, REACTIONS TESTING 1+ OR LESS IN TUBE MAY BE NEGATIVE ON THE ECHO. >< CUSTOMER REPORTED THAT A PATIENT WITH A PREVIOUS HISTORY AS D POSITIVE HAD TESTED NEGATIVE WITH ONE OF THE ANTI-D REAGENTS AND 1+ POSITIVE WITH THE OTHER ANTI-D REAGENT ON THE ECHO (SOFTWARE V 1.1.1.7). | 0
993112 | CUSTOMER IS REPORTING AN ABO DISCREPANCY ON THE GALILEO WITH THE FWD_ABO ASSAY. THE FIRST RUN ON THE GALILEO RESULTED IN O NEGATIVE AND THE SECOND RUN RESULTED IN A NEGATIVE. >< THE INSTRUMENT IMAGES WERE REVIEWED AND APPEARED AS EXPECTED. TESTING PERFORMED WITH RETENTION ANTI-A, LOT 101673, 101674 WITH IN HOUSE DONOR SAMPLES REVEALED TYPE A RED CELL STICKING IN THE ANTI-A WELLS. TO PREVENT THE STICKING, IN HOUSE STUDIES FOUND THAT PIPETTING THE REAGENTS INTO THE WELLS BEFORE THE SAMPLE RED CELLS GREATLY REDUCED THE STICKING. IN HOUSE COMPARISON TESTING OF THE NEW REAGENT-FIRST PIPETTING ORDER TO THE EXISTING SAMPLE-FIRST PIPETTING ORDER SHOWED NO UNEXPECTED NEGATIVE ANTI-A REACTIONS WITH THE NEW PIPETTING ORDER. SOME UNEXPECTED NEGATIVE REACTIONS WERE OBSERVED DURING TESTING WITH THE CURRENT ASSAY. CUSTOMERS WERE NOTIFIED OF THE MODIFICATION TO THE ASSAY ON 9/20/07.THIS FACILITY DID HAVE THE MODIFIED ABO ASSAY SOFTWARE INSTALLED. ACCORDING TO THE GALILEO OPERATOR MANUAL, FORWARD ONLY ABO-RH TESTING HAS A HIGHER RISK OF MISTYPE DUE TO THE ABSENCE OF THE REVERSE TYPE RESULTS. HAZARDOUS MISTYPES MAY OCCUR, SUCH AS AN A SAMPLE BEING INTERPRETED AS A GROUP AB, OR AN RH (D) NEGATIVE SAMPLE BEING INTERPRETED AS RH(D) POSITIVE. FOR THIS REASON, ABO-RH RESULTS SHOULD ALWAYS BE COMPARED TO THE PATIENT OR DONOR'S HISTORY. | 0
1055469 | FWD_ABORH TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH IN-HOUSE DONOR SAMPLES OF VARIOUS ABO AND RH TYPES USING RETENTION ANTI-A, LOT 101679, ANTI-B SERIES 3, LOT 203234, ANTI-D SERIES 4, LOT 504697, AND MONOCLONAL CONTROL, LOT 492028. ALL IN-HOUSE DONOR SAMPLES TYPED AS EXPECTED WITH NO DISCREPANCIES. TESTING PERFORMED WITH RETENTION ANTI-A, LOT 101673, 101674 WITH IN HOUSE DONOR SAMPLES REVEALED TYPE A RED CELL STICKING IN THE ANTI-A WELLS. TO PREVENT THE STICKING, IN HOUSE STUDIES FOUND THAT PIPETTING THE REAGENTS INTO THE WELLS BEFORE THE SAMPLE RED CELLS GREATLY REDUCED THE STICKING. IN HOUSE COMPARISON TESTING OF THE NEW REAGENT-FIRST PIPETTING ORDER TO THE EXISTING SAMPLE-FIRST PIPETTING ORDER SHOWED NO UNEXPECTED NEGATIVE ANTI-A REACTIONS WITH THE NEW PIPETTING ORDER. SOME UNEXPECTED NEGATIVE REACTIONS WERE OBSERVED DURING TESTING WITH THE CURRENT ASSAY. CUSTOMERS WERE NOTIFIED OF THE MODIFICATION TO THE ASSAY ON 9/20/07.THIS FACILITY DID HAVE THE MODIFIED ABO ASSAY SOFTWARE INSTALLED. ACCORDING TO THE GALILEO OPERATOR MANUAL,FORWARD ONLY ABO-RH TESTING HAS A HIGHER RISK OF MISTYPE DUE TO THE ABSENCE OF THE REVERSE TYPE RESULTS. HAZARDOUS MISTYPES MAY OCCUR, SUCH AS A SAMPLE BEING INTERPRETED AS A GROUP AB, OR AN RH (D) NEGATIVE SAMPLE BEING INTERPRETED AS RH(D) POSITIVE. FOR THIS REASON, ABO-RH RESULTS SHOULD ALWAYS BE COMPARED TO THE PATIENT OR DONOR'S HISTORY. >< CUSTOMER REPORTED THAT FIVE DONOR SPECIMENS WERE TESTED WITH THE GALILEO (FWD_ABORH ASSAY) AND FOUND TO BE GROUP O; HISTORICAL TYPING OF THESE DONORS REVEALED A TYPE OF GROUP A. | 0
1055466 | BLUD_DIRECT WAS UTILIZED TO OBTAIN INSTRUMENT IMAGES AND TO DETERMINE THE TIMES THE INDICATOR CELLS WERE PLACED ON THE GALILEO. THE WELL FILL IMAGES FOR THE SCREEN PLATE APPEARED ACCEPTABLE. IT APPEARED THE INDICATOR CELLS WERE CHANGED OUT APPROPRIATELY AND WERE USED WITHIN EXPIRATION. THE UNEXPECTED NEGATIVE REACTIONS WERE NOT REPRODUCED WHEN REPEATED WITH THE SAME REAGENTS. >< CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIVITY WITH A PATIENT SAMPLE USING THE 2 CELL SCREEN ON GALILEO. | 0
1055463 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WITH THE POOL_CELL SCREEN ASSAY ON THE GALILEO. A FIRST TIME THE DONOR SAMPLE TESTED REACTIVE 2+ . IT WAS REPEATED ON ANOTHER POOL_CELL PLATE USING THE SAME LOTS OF REAGENTS AND IT TESTED NEGATIVE. >< UPON REVIEW OF THE INSTRUMENT IMAGES, THE IMAGES APPEARED NORMAL AND THE ROIS DID NOT APPEAR TO BE MISALIGNED. THE WELL IMAGE ON THE ORIGINAL RUN EXHIBITED MODERATE ADHERENCE, BUT ON THE REPEAT RUN THE WELL IMAGE EXHIBITED SLIGHT ADHERENCE WITH A DEFINED BUTTON. ACCORDING TO THE CUSTOMER, FLOW CHECKS HAVE BEEN PASSING. STIR BALLS WERE ADDED TO THE CAPTURE-R INDICATOR CELLS AND NO HEMOLYSIS WAS NOTED, NO WAS LEAKING OBSERVED IN PIPETTOR OR SYRINGE AREAS, NO BUBBLES/FOAM WERE IN SYSTEM LIQUID LINES OR REAGENT VIALS, AND READER MAINTENANCE HAS BEEN PASSING. POOL_CELL TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH THE CUSTOMER'S RETUREND DONOR SAMPLE, USING RETENTION CAPTURE-R READY-SCREEN (POOLED), LOT N151, AND RETENTION CAPTURE-R INDICATOR RED CELLS, LOT 221168. THE SAMPLE EXHIBITED 2+ REACTIVITY. POOL_CELL TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH CUSTOMER'S SAMPLE USING RETURNED CAPTURE-R READY-SCREEN (POOLED), LOT N151, AND RETENTION CAPTURE-R INDICATOR RED CELLS, LOT 221168. THE SAMPLE EXHIBITED 1+ REACTIVITY. POOL_CELL TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH THE CUSTOMER'S SAMPLE USING RETURNED CAPTURE-R READY-SCREEN (POOLED), LOT N151, AND RETURNED CAPTURE-R INDICATOR RED CELLS, LOT 221168. THE SAMPLE EXHIBITED 1+ REACTIVITY. | 0
1033203 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS ON THE ABS2000. PATIENT SAMPLES WITH A HISTORY ANTI-KELL AND ANTI-FYA, RESPECTIVELY, WERE REPORTED AS ANTIBODY NEGATIVE ON THE ABS2000. >< THE REACTIVITY OF THE K ANTIGEN WAS CONFIRMED FOR RETENTION CAPTURE-R REASY-SCREEN (4), LOT G158, USED AT THE TIME OF THE COMPLAINT. REACTIVITY OF THE FYA ANITGEN WAS CONFIRMED THROUGH DHR REVIEW, SINCE PRODUCT EXPIRED DURING INVESTIGATION TESTING. A SERVICE VISIT WAS MADE. A 100UL BENT SYRINGE WAS REPLACED AND THE READER WINDOW WAS CLEANED. THE LEVEL SENSE FREQUENCY WAS OUT OF RANGE AND THE CENTRIFUGE SHAKE TEST OUT OF RANGE; BOTH SPECIFICATIONS WERE ADJUSTED. THE REPAIR HAS RETURNED THE INSTRUMENT TO OPERATING WITHIN SPECIFICATIONS. | 0
1099083 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH CAPTURE-R READY SCREEN(3), LOT R023 WHEN TESTING A PATIENT SAMPLE ON THE ECHO. GROUP/SCREEN TESTING ON THE ECHO RESULTED AS NEGATIVE. >< ANTIBODY SCREENING TEST WAS PERFORMED WITH THE CUSTOMER'S RETURNED SAMPLE USING RETURNED CAPTURE-R READY-SCREEN (3), LOT R023, ON AN IN-HOUSE ECHO. THE CUSTOMER'S SAMPLE EXHIBITED STRONG POSITIVE (4+) REACTIVITY WITH CELLS I AND II AND WAS NONREACTIVE WITH CELL III. THE TESTING WAS REPEATED AND CONSISTENT RESULTS WERE OBTAINED. MANUAL TESTING WAS PERFORMED WITH THE CUSTOMER'S PATIENT SAMPLE USING RETENTION CAPTURE-R READY-ID AND CAPTURE-R INDICATOR RED CELLS. THE SAMPLE EXHIBITED POSITIVE REACTIVITY WITH ALL D+ CELLS AND WAS NONREACTIVE WITH ALL D- CELLS. | 0
1240023 | CUSTOMER REPORTED A NEGATIVE ANTIBODY SCREEN ON A PATIENT SAMPLE CONTAINING ANTI-E WHEN TESTED ON GALILEO. THE CUSTOMER REPORTED VISUALLY THAT THE REACTIONS LOOKED REACTIVE. >< AN AUTOMATIC CAMERA ADJUSTMENT WAS DONE. ASKED CUSTOMER TO REPEAT THE SAMPLE; CUSTOMER RERAN THE SAMPLE AND STILL GOT A NEGATIVE RESULT. A SERVICE CALL WAS MADE. THE CAMERA MODULE WAS REPLACED. RAN QC_ASSAY, 2_CELL, SCREEN AND ABORH WITH ONE PATIENT SAMPLE WITH ACCEPTABLE RESULTS. THE INSTRUMENT WAS RETURNED TO CUSTOMER USE. | 0
1102363 | INSTRUMENT IMAGES WERE REVIEWED. THE SAMPLE APPEARS NEGATIVE IN ALL WELLS. THE REACTIVITY OF THE C AND E ANTIGENS WERE CONFIRMED ON RETURNED CAPTURE-R READY SCREEN, LOTS K184 AND E007.THE CUSTOMER'S RETURNED SAMPLE WAS TESTED WITH RETURNED CAPTURE-R READY SCREEN 4, LOT K184 ON AN IN-HOUSE GALILEO. THE RETURNED SAMPLE WAS NONREACTIVE WITH ALL CELLS TESTED. THE RETURNED SAMPLE WAS ALSO TESTED WITH RETURNED CAPTURE-R READY SCREEN 3, LOT E007ON AN IN-HOUSE ECHO. THE SAMPLE WAS NONREACTIVE WITH ALL CELLS TESTED. THE CUSTOMER'S SAMPLE WAS ALSO TESTED WITH RETENTION PANOSCREEN I, II AND III, LOT 20724 WITH CAPTURE SELECT LOT SC089 ON AN IN-HOUSE GALILEO. THE SAMPLE WAS NONREACTIVE WITH ALL CELLS TESTED. THE EVENT APPEARS TO BE RELATED TO THE NATURE OF THE SAMPLE. >< CUSTOMER REPORTED MISSED ANTIBODIES ( ANTI E AND ANTI C) ON THE ECHO, IN AN ANTENATAL SAMPLE TAKEN FROM A MOTHER. THE BABY WAS TRANSFUSED 1 UNIT OF O NEGATIVE BLOOD. | 0
1226172 | CUSTOMER REPORTED AN ABO DISCREPANCY ON THE GALILEO. A PATIENT TYPED AS 'B POSITIVE' ON THE GALILEO, HOWEVER, UPON REVIEW OF PATIENT HISTORY, SHE HAD PREVIOUSLY TYPED ON THE GALILEO, AS AB POSITIVE. >< THE CUSTOMER DID NOT SEND SPECIMEN FOR INVESTIGATION TESTING. | 0
1062614 | THE RETURNED PATIENT SAMPLE WAS TESTED IN HEMAGGLUTINATION USING SELECTED K+K+, K+K- AND K-K+ FROM PANOCELL-16, LOT 14611 USING IMMUADD AS THE POTENTIATOR. A ROOM TEMPERATURE INCUBATION WAS INCLUDED. THE SAMPLE WAS POSITIVE AT THE IGG PHASE ONLY.REACTIVITY OF THE K ANTIGEN WAS CONFIRMED ON RETENTION CAPTURE-R READY SCREEN (CRRS) (4), LOT G159 AND G160. THIS LOT OF CAPTURE WAS USED BY THE CUSTOMER. THE PATIENT SAMPLE WAS TESTED IN MANUAL SOLID PHASE WITH RETENTION CRRS(4), LOT G159 AND G160. THE SAMPLE EXHIBITED WEAK REACTIVITY (SCORE OF 4 ON K+K+ CELLS ON G159 AND SCORE OF 6 WITH G160 ON SCALE OF 0-8). THE SAMPLE WAS ALSO TESTED ON AN IN-HOUSE ABS2000 USING RETENTION CRRS(4), LOT G160; IT EXHIBITED POSITIVE REACTIVITY WITH K+K+ REAGENT RED CELLS. THE PATIENT'S SAMPLE WAS TESTED ON THE IN-HOUSE ABS2000 USING RETENTION CRRS(4), LOT G159; IT EXHIBITED NO REACTIVITY REACTIVITY WITH K+K+ REAGENT RED CELLS. POSITIVE REACTIVITY WITH K+K+ REAGENT RED CELLS WAS OBSERVED WITH THE PATIENT'S SAMPLE WHEN TESTED ON AN IN-HOUSE ECHO USING RETENTION CRRS(3). >< CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION ON THE ABS2000. A PATIENT SAMPLE WITH A KNOWN ANTI-KELL TESTED NEGATIVE ON THE ABS2000 AND 1+ POSITIVE IN GEL. THE ANTI-KELL WAS IDENTIFIED IN GEL AND PEG. A NEW PATIENT SAMPLE WAS OBTAINED APPROXIMATELY 5 DAYS LATER, AND GEL TESTING WAS PERFORMED ON BOTH SAMPLES. 2+ REACTIVITY WAS OBSERVED WITH BOTH SAMPLES. | 0
1062623 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WITH ANTI-D SERIES 5 ON A PATIENT SAMPLE RUN ON THE ECHO. CUSTOMER HAD RUN THE SAMPLE IN TUBE USING ANTI-D MONOCLONAL BLEND AND HAD POSITIVE (3+) REACTIONS. >< A LIMITATION OF THE ECHO, ACCORDING THE OPERATOR MANUAL, IS THAT HEMAGGLUTINATION REACTIONS GRADED 1+ OR LESS IN TUBE, MAY NOT BE DETECTED ON THE ECHO. | 0
1102343 | ABORH WAS TESTING PERFORMED ON AN IN-HOUSE GALILEO WITH IN-HOUSE DONOR SAMPLES OF KNOWN ABO/RH TYPES AND THE CUSTOMER'S RETURNED SAMPLE, USING RETENTION ANTI-A, LOT 101682, AND ANTI-B SERIES 3, LOT 203236. A1 REFERENCELLS, LOT 111655, AND B REFERENCELLS, LOT 113655, WERE USED FOR REVERSE GROUP TESTING. ALL IN-HOUSE DONOR SAMPLES TYPED AS EXPECTED. THE RETURNED SAMPLE WAS INTERPRETED AS A POSITIVE, AS EXPECTED. >< CUSTOMER REPORTED AN ABO DISCREPANCY ON GALILEO. A PATIENT SAMPLE RESULTED AS O POSITIVE. THE PATIENT IS A HISTORICAL A POSITIVE AND PREVIOUSLY RECEIVED 2 UNITS OF A POSITIVE RED BLOOD CELLS. | 0
1163324 | CUSTOMER INDICATED THAT AN ANTIBODY PANEL COMPLETED ON GALILEO SHOWED WEAK TO NO REACTIVITY ON A PATIENT WHICH HAD AN ANTI-C. THE CUSTOMER STATED THAT BASED ON THE REACTIVITY OF THE ANTIBODY ID PANEL, ANTI-C WAS RULED OUT. THE CUSTOMER INDICATED THAT THE UNEXPECTED REACTIONS DID NOT RESULT IN AN ADVERSE PATIENT EVENT. >< REACTIVITY OF THE E, C AND JKA ANTIGENS WERE CONFIRMED ON RETENTION CRRS(I AND II), LOT X253, CRRID, LOTS ID104 AND DP029. CUSTOMER RETURNED SAMPLE FOR INVESTIGATION TESTING. TESTING IS ONGOING. THE CUSTOMER'S RETURNED SAMPLES WERE TESTED WITH RETENTION CRRS(I AND II), LOT X253 ON AN IN-HOUSE GALILEO INSTRUMENT. MODERATE TO STRONG REACTIVITY WAS OBSERVED WITH CELL II {(C-E+E-C+;JK(A+B-)}. THE SAMPLE WAS TESTED WITH RETENTION CRRID, LOT ID104 ON AN IN-HOUSE GALILEO. IDENTICAL REACTIVITY BETWEEN CUSTOMER'S TESTING AND OUR TESTING WAS OBSERVED WITH CELL 1 TO CELL 8, 10 AND 14. OUR IN-HOUSE GALILEO INSTRUMENT GAVE 1+ TO 2+ REACTIVITY IN CELL # 11, 12 AND 13 (DID NOT GIVE EQUIVOCAL (?) AS CUSTOMER OBSERVED). IN CELL 9, OUR GALILEO GAVE NEGATIVE REACTIVITY (CUSTOMER GOT (?) EQUIVOCAL REACTIVITY). WE CANNOT RULE OUT OTHER CLINICALLY SIGNIFICANT ALLOANTIBODIES. THE CUSTOMER'S SAMPLE WAS TESTED IN HEMAGGLUTINATION TESTS WITH SELECTED CELLS FROM PANOCELL-20, LOT 28841, USING IMMUADD AS THE POTENTIATOR. ONLY VERY WEAK MICROSCOPIC REACTIVITY WAS OBSERVED. | 0
1234620 | IN-HOUSE DONOR SAMPLES OF KNOWN D TYPES WERE TESTED WITH RETENTION ANTI-D SERIES 4, LOTS 504698 AND 504700 AND ANTI-D SERIES 5, LOT 505548 ON AN IN-HOUSE ECHO. THE EXPECTED REACTIVITY PATTERN WAS OBSERVED WITH ALL SAMPLES TESTED. THE RETURNED SAMPLES WERE RIMMED WITH APPLICATOR STICKS AND CHECKED FOR CLOTS - NONE WERE OBSERVED IN THE SUBMITTED SAMPLES. THE RETURNED SAMPLES WERE TESTED WITH RETENTION ANTI-D SERIES 4, LOTS 504698 AND 504700 AND ANTI-D SERIES 5, LOT 505548 ON AN IN-HOUSE ECHO. THE SAMPLE HISTORICAL TYPE (A NEGATIVE) WAS CONFIRMED. THE CUSTOMER STATES THAT THEY DO NOT ROUTINELY CHECK THEIR SAMPLES FOR SMALL CLOTS. THE POSSIBILITY OF SAMPLE INTEGRITY COULD NOT BE ELIMINATED, AS A POSSIBLE CAUSE FOR THE INITIAL POSITIVE REACTIONS NOTED BY THE CUSTOMER. >< CUSTOMER REPORTED A PATIENT SAMPLE TYPED A POSITIVE ON THE ECHO WITH BOTH ANTI-D REAGENTS ( 1+). THE SAMPLE WAS RUN IN TUBE USING ANTI-D4, BUT A DIFFERENT LOT THAN ON THE ECHO, AND IT TESTED NEGATIVE. | 0
1102384 | CUSTOMER REPORTED AN ABO DISCREPANCY ON THE ABS 2000. AN A POSITIVE PATIENT WAS TYPED AB POSITIVE. >< A SERVICE VISIT WAS PERFORMED. TROUBLESHOOTING THE EQUIPMENT DID NOT REVEAL ANY ISSUE WITH THE SYSTEM. CLEANED THE READER WINDOW, FILTERS, FILTER LENS AND LAMP DEFECTOR. THE CUSTOMER RAN 4 PATIENT SAMPLES; EXPECTED RESULTS WERE OBTAINED. | 0
1035460 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WHEN TESTING A GROUP A2 PATIENT SAMPLE ON THE ECHO. INSTRUMENT RESULTS WERE GROUP O. NO ADVERSE REACTIONS OR TRANSFUSIONS OCCURRED AS A RESULT OF THE UNEXPECTED NEGATIVE REACTIONS. >< IN-HOUSE DONOR SAMPLES OF KNOWN ABO TYPES INCLUDING A2 SUBGROUPS WERE TESTED ON AN IN-HOUSE ECHO. THE EXPECTED REACTIVITY WAS OBSERVED IN ALL TESTING. THE RETURNED PATIENT'S SAMPLE EXHIBITED NTD RESULTS WHEN TESTED WITH THE SAME RETENTION LOTS ON AN IN-HOUSE ECHO INSTRUMENT. THE SAMPLE WAS TESTED BY MANUAL TUBE METHOD USING THE SAME REAGENTS USED ON FOR ECHO TESTING. THE FORWARD TYPE WAS GROUP O AND THE REVERSE TYPE WAS GROUP B. THE GALILEO ECHO OPERATOR MANUAL 'CERTAIN BLOOD SAMPLES FACTORS HAVE A HIGH LIKELIHOOD OF CAUSING NO- TYPE-DETERMINED(NTD) ABO OR RH (D) INTERPRETATIONS. THESE INCLUDE SEROLOGICAL FACTORS DUE TO THE INHERITANCE OF WEAKLY EXPRESSED GENE PRODUCTS, BY CERTAIN DISEASES (SUCH AS LEUKEMIA) THROUGH THE TRANSFUSION OR TRANSPLANTATION OF ABO AND RH (D) ALLOGENEIC RED BLOOD CELL PRODUCTS, OR DUE TO THE PATIENT'S AGE'. THE NATURE OF THE SAMPLE CANNOT BE RULED OUT AS THE CAUSE OF THE EVENT. | 0
1035462 | FWD_ABORH TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH IN-HOUSE DONOR SAMPLES OF VARIOUS ABO/RH TYPES AND CUSTOMER'S RETURNED PATIENT SAMPLE. ALL IN-HOUSE DONOR SAMPLES TYPED AS EXPECTED. THE CUSTOMER'S SAMPLE WAS INTERPRETED AS GROUP B, RH POSITIVE.HEMAGGLUTINATION TUBE TESTING WAS PERFORMED WITH THE PATIENT SAMPLE USING ANTI-A, LOT 101680, AND ANTI-B, LOT 203234. THE SAMPLE EXHIBITED VERY WEAK (+W) REACTIVITY AT IS WITH ANTI-A REAGENT AND STRONG POSITIVE (4+) REACTIVITY WITH ANTI-B. AFTER 15'RT INCUBATION, SAMPLE EXHIBITED 3+ MIXED-FIELD REACTIVITY. THE GALILEO OPERATOR MANUAL STATES ?THE GALILEO CANNOT RELIABLY DETECT HEMAGGLUTINATION REACTIONS THAT ARE GRADED AS 1+ OR LESS IN TEST TUBE METHODOLOGY.' >< CUSTOMER REPORTED AN ABO DISCREPANCY ON THE GALILEO FOR A NEONATAL PATIENT SAMPLE, USING THE FWD_ABO ASSAY. THE SAMPLE WAS B POSTIIVE ON THE INSTRUMENT, BUT AB POSITIVE IN MANUAL TESTING. | 0
1035461 | THIS CUSTOMER ARCHIVES DAILY AND DISCARDS THE DISK THAT THE DATA IS ARCHIVED TO; IT WAS NOT POSSIBLE TO REVIEW INSTRUMENT IMAGES FOR THE REFLEX ABO TESTING PLATE. RETREIVE THE IMAGE OF THE WEAK D PLATE AND THE OAL FILES ASSOCIATED WITH THE REFLEX ABO AND WEAK D ASSAYS. SAMPLE WAS IN POSITION E4 AND PRODUCED A NEGATIVE RESULT ON THE PRINTOUT. THE VISUAL APPEARANCE OF THE WELL SHOWS A NEGATIVE REACTION WITH A SOLID CELL BUTTON THAT IS CENTRALLY LOCATED WITHIN THE WELL. THERE IS A MONOLAYER PRESENT IN THE WELL.THE CUSTOMER DID NOT RETURN SAMPLE FOR INVESTIGATION TESTING, IT IS NOT POSSIBLE TO RULE OUT THE SAMPLE AS A CAUSE OF THE EVENT. >< CUSTOMER REPORTED THAT A DONOR SAMPLE WAS TESTED AND RESULTED AS B NEGATIVE ON GALILEO (B, WEAK D PENDING WITH REFLEX ABO, AND WEAK D ASSAY RESULTED AS NEGATIVE). THE FACILITY WAS NOTIFIED THAT THIS SAMPLE WAS HISTORICALLY TYPE B, WEAK D POSITIVE. | 0
1035464 | THE SAMPLE WAS TESTED AND EXPECTED RESULTS WERE OBTAINED. THE INSTRUMENT IS OPERATING AS EXPECTED. >< CUSTOMER REPORTED AN ABO DISCREPANCY ON A DONOR SAMPLE TESTED ON THE GALILEO USING THE FWD_ABO ASSAY. | 0
1026971 | RETENTION CRRS (I AND II), LOT X231 WAS POSITIVE FOR THE K ANTIGEN. THE CUSTOMER RETURNED A SAMPLE. 2_CELL TESTING WAS PERFORMED ON AN IN-HOUSE GALILEO WITH CUSTOMER'S SAMPLE USING RETENTION CAPTURE-R READY-SCREEN (I AND II), LOT X231. THE SAMPLE WAS NONREACTIVE WITH BOTH CELLS. HEMAGGLUTINATION TUBE TESTING WAS PERFORMED WITH THIS SAMPLE USING CELL I (K+K+) AND CELL II (K-) FROM RETENTION PANOSCREEN I, II, AND III, LOT 09551. IMMUADD, WAS USED AS POTENTIATOR. THE SAMPLE WAS NONREACTIVE WITH BOTH CELLS IN ALL PHASES OF TESTING. >< CUSTOMER REPORTED A MISSED KELL ANTIBODY ON A PATIENT SAMPLES USING THE 2_CELL SCREEN ASSAY ON GALILEO . | 0
972936 | DBSS ALLOWS THE ENTRY OF A PHLEBOTOMY TO AN INCORRECT REGISTRATION WHICH PREDATES AN ACTIVE DEFERRAL WITHOUT WARNING. | 1
868415 | THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
951762 | THE CUSTOMER REPORTED A SPEAKER MALFUNCTION ON THE MONITOR. >< THE HOSPITAL BIOMEDICAL ENGINEER, CALLED THE REMOTE TECHNICAL ASSISTANCE CENTER TO REPORT A FAILED SPEAKER. A PHILLIPS FIELD SERVICE ENGINEER (FSE), WENT ONSITE AND TESTED THE INVOLVED DEVICE. THE FAILED SPEAKER WAS REPLACED. THE REPAIRED DEVICE REMAINS IN USE AT THE CUSTOMER SITE. | 1
1018586 | ACCOUNT REPORTS EXPERIENCING AN INTERMITTENT ISSUE WITH THEIR MODULAR PRE-ANALYTICS LABORATORY INTEGRATION SYSTEM GENERATING BUBBLES IN SAMPLE CUPS PRIOR TO BEING LOADED ONTO THE ANALYZER FOR TESTING. ACCOUNT STATES THIS ISSUE HAS CAUSED DISCREPANT PT RESULTS FOR APPROXIMATELY 10 PT EXAMPLES STARTING IN 2008. ONLY THE FOLLOWING PT SAMPLE WAS PROVIDED: INITIAL RESULT GAVE 5.9 MG/DL; REPEAT GAVE 9.6 MG/DL. ERRONEOUS RESULT WAS NOT REPORTED. THE FIELD SERVICE REPRESENTATIVE DETERMINED THE CAUSE TO BE PROBLEMS WITH THE ALIQUOT Z DEPTH, ALIQUOT TIP POSITIONING AND TSM SETTINGS. PERFORMED ADJUSTMENTS TO THE ALQ TIP POSITIONING, TIP RACK POSITION, AND SET TSM SETTINGS. PERFORMANCE TESTS WERE PERFORMED AND WITHIN SPECIFICATION. | 0
1068374 | THE GREATER CHESAPEAKE AND POTOMAC REGION, MANUFACTURING MANAGER SELF- DISCOVERED THAT RED BLOOD CELL (RBC), WHICH WAS INVOLVED IN AN NTL TESTING ISSUE WHEREIN THE CMV RESULT CHANGED FROM NEGATIVE TO POSITIVE, WAS RELEASED OUT OF QUARANTINE PRIOR TO THE CMVN (CMV NEGATIVE) PROPERTY BEING REMOVED. THIS RESULTED IN THE RBC BEING LABELED AS CMVN. THE REGION GAINED CONTROL OF THE RBC AND IT WAS IMMEDIATELY QUARANTINED PENDING REMOVAL OF THE PROPERTY. | 0
1035456 | DURING PREOPERATIVE ANESTHESIA MACHINE CHECKS, THE DOCTOR NOTED A NIBP 'CUFF OCCLUSION' ERROR MESSAGE ON MACHINE S5 MONITOR DISPLAY. THE M-PRETN MULTIPARAMETER VITAL SIGNS MODULE WOULD NOT ALLOW NONINVASIVE BLOOD PRESSURE MEASUREMENT. THE DOCTOR CHECKED ALL CABLES AND CUFF TUBING FOR KINKING, AND THERE WAS NONE. THIS ERROR HAS BEEN RECURRING ON SEVERAL OF THESE MODULES AND HAS REQUIRED REPAIR BY THE MANUFACTURER. IN EVERY CASE THE MANUFACTURER HAS REPLACED THE NIBP INTERFACE PCB TO CORRECT ERROR. THE PROBLEM WITH THE REMAINING MODULES KEEPS OCCURRING, WHICH CAUSES DELAY AND OR CANCELLATION OF SURGICAL PROCEDURES. THE MANUFACTURER HAS BEEN NOTIFIED OF THESE FAILURES, BUT HAS NOT RESOLVED THIS ONGOING ISSUE. THIS CURRENT MEDWATCH REPORT IS FOR TWO ADDITIONAL MODULE FAILURES. | 1
1084865 | RADIANT WARMER USED FOR THERMOREGULATION OF NEWBORN PREMATURE INFANT. TEMPERATURE PROBE READING INCORRECTLY. TEMPERATURE PROBE REPLACED 3 TIMES FOR READING LOW SKIN TEMPERATURE WHEN BABY FELT WARM AND AUXILIARY TEMP WITHIN NORMAL LIMITS. AFTER 3RD PROBE PLACED ON INFANT, WARMER TEMPERATURE SET TO 36 C SKIN. FORTY FIVE MINUTES LATER, INFANT'S SKIN NOTED TO BE BRIGHT RED ON BACK. AUXILIARY TEMPERATURE 38.3C AND RED ERYTHEMA EVIDENT ON EXPOSED SKIN. CHILD PLACED ON NEW WARMER. WARMER SEQUESTERED FOR BIOMED EVALUATION. NO FURTHER ISSUES WITH TEMPERATURE OR SKIN. ERYTHEMA RESOLVED WITHIN HOURS. BIOMED ENGINEER INVESTIGATED THE RADIANT WARMER ISSUE THAT OCCURRED. GE HEALTHCARE REP CAME TO HOSPITAL TO EVALUATE THE UNIT. THE REP FULLY TESTED THE UNIT AND FOUND THE UNIT OPERATING AS PRESCRIBED BY THE MFR. HE STATED THAT THE UNITS ARE SELF CONTAINED AND REQUIRE NO CALIBRATION. HE FOUND NO REASON WHY THIS INCIDENT HAPPENED. HE SPOKE TO NICU CHARGE NURSE ON DUTY AND INFORMED HER OF HIS FINDINGS. HE ASKED THAT THE FACILITY CONTINUE TO MONITOR THIS AND CONTACT GE CLINICAL APPLICATIONS IF THIS ISSUE REAPPEARS. DIAGNOSIS OR REASON FOR USE: NEWBORN PREMATURE INFANT REQUIRING THERMOREGULATION. | 0
977219 | ACUDOSE CABINET ERRORS AT 0600, 1230, AND 2300 HOURS. THE DRAWERS OF THE ACUDOSE WOULD LOCK UP AND DENY STAFF ACCESS TO THE MEDICATIONS. STAFF WOULD HAVE TO GO ELSEWHERE FOR THE MEDICATIONS. IF THERE WAS AN EMERGENCY IT WOULD TAKE SEVERAL MINUTES TO GET THE MEDICATIONS. THESE EVENTS WERE NOT EMERGENCIES BUT PREVIOUS EVENTS HAVE BEEN. | 1
1069878 | PHARMACY TECH WAS ABLE TO OVER-RIDE COMPUTER SYSTEM PASSCODE PROTECTION ENABLING HIM TO REMOVE CONTROLLED MEDICATIONS BY CHANGING THEIR STATUS. MCKESSON UNAWARE OF THIS POTENTIAL, ADVISED OF SAME IN 2008. EMPLOYEE TERMINATED FOR DRUG DIVERSION. | 1
1183929 | THIS MALE HAS CALCIFIC AORTIC STENOSIS AND SEVERE LEFT VENTRICULAR SYSTOLIC DYSFUNCTION. PREVIOUS ECHOCARDIOGRAPHY AND CARDIAC CATHETERIZATION YIELDED CONFLICTING ESTIMATES OF THE SEVERITY OF AORTIC STENOSIS. HE WAS REFERRED FOR REPEAT CARDIAC CATHETERIZATION, TO INCLUDE DOBUTAMINE STIMULATION TO DETERMINE SURVIVAL PROBABILITY IF AORTIC VALVE REPLACEMENT IS ATTEMPTED. THE MCKESSON HORIZON CARDIOLOGY HEMODYNAMIC MONITORING AND RECORDING SYSTEM FAILED TO ALLOW THERMODILUTION MEASUREMENT OF CARDIAC OUTPUT, DATA CRUCIAL FOR ASSESSMENT OF AORTIC STENOSIS AND RESPONSE TO DOBUTAMINE. THE THERMODILUTION CATHETER WAS CHANGED, THE PROBLEM PERSISTED. THE CONNECTING CABLE WAS CHANGED, THE PROBLEM PERSISTED. THE PT WAS TRANSFERRED TO ANOTHER ROOM, THE PROBLEM PERSISTED. A THIRD THERMODILUTION CATHETER WAS EMPLOYED, WITHOUT BENEFIT. FAILURE OF THE CARDIAC OUTPUT MEASUREMENT FUNCTION IN THE MCKESSON SYSTEM HAS BEEN INTERMITTENT BUT RECURRENT SINCE THIS SYSTEM WAS INSTALLED. COMPANY TECHNICIANS WITNESSED ONE SUCH EVENT, AND COMPANY EXECUTIVES HAVE BEEN INFORMED OF THE PROBLEM, YET IT HAS NOT BEEN RESOLVED. THIS STUDY WAS INCOMPLETE, PROLONGED AND EXPOSED THE PT TO HAZARDS AS A RESULT OF EQUIPMENT MALFUNCTION. THE CLINICAL QUESTION REMAINS UNANSWERED, DOBUTAMINE STRESS ECHOCARDIOGRAPHY WILL BE PERFORMED FOR LEFT VENTRICULAR SYSTOLIC CONTRACTION RESPONSIVENESS, BUT CARDIAC OUTPUT DATA WILL NOT BE AVAILABLE. | 1
953843 | UNDER ANOTHER COUNTRY LAW, PATIENT INFORMATION IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASE BY THE HOSPITAL. >< CUSTOMER REPORTED MECHANICAL VENTILATION STOPPED AND MONITOR SCREEN BLANKED OUT. THERE WAS NO REPORTED PATIENT INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
989767 | IT WAS REPORTED THAT THE M-PRENT WAS USED WITH PT DURING SURGERY. THE DEVICE ALLEGEDLY DISPLAYED AN ERROR 6 CODE WHICH INDICATES THAT NIBP VALUES ARE NOT AVAILABLE. IT IS NOT KNOWN AT THIS TIME WHETHER THE PT DEATH IS ATTRIBUTABLE TO THE ALLEGED EVENT, THEREFORE, THE EVENT IS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
881023 | OASIS MEDICAL RECEIVED AN EMAIL FROM A CUSTOMER STATING THAT THEY 'WITHIN THE LAST 2 WEEKS WE HAD AT LEAST 7 SURGEONS COMPLAINING ABOUT THE QUALITY.' THERE WERE NO FORMAL COMPLAINTS FILED WITH THE CUSTOMER FOR ANY OF THESE COMPLAINTS MENTIONED. OASIS MEDICAL FOLLOWED UP FOR FURTHER INFO REGARDING THE COMPLAINTS. WE WERE INFORMED THAT NONE OF THE DOCTORS USED THE PRODUCT AND NOTICED THAT UPON OPENING THE PRODUCT THAT THE KNIFE HAD PUNCTURED THROUGH THE FOAM AND PIERCED THE PACKAGING. THESE COMPLAINTS WERE ALL REGARDING THE OASIS MEDICAL PE SLIT KNIFE. >< THE PRODUCT WAS NOT RETURNED TO OASIS MEDICAL FOR ANALYSIS. THE CUSTOMER WAS NOT GOING TO RETURN THE PRODUCT. THE CUSTOMER DID NOT KNOW THE LOT NUMBERS OR SPECIFIC PRODUCT CODES/NUMBERS. THE CUSTOMER ONLY REFERRED TO THE PRODUCT LINE AS PE SLIT KNIVES. THIS PROBLEM IS THE SAME AS RECALL AND WE HAVE SINCE EXPANDED THE RECALL TO INCLUDE THESE PRODUCTS. THE RECALL WAS EXPANDED ON 07/12/2007. | 0
976895 | ACUDOSE CABINET ERROR AT 0110. CABINET LOCKED UP. | 1
1006519 | OASIS MEDICAL WAS UNABLE TO OBTAIN PRODUCT REFERENCE NUMBER AND LOT NUMBERS FOR EITHER PT OR PRODUCT AFTER MULTIPLE ATTEMPTS WITH THE REPORTING DOCTOR. A REVIEW OF PRODUCT BATCH PRODUCTION RECORDS SHOWS THAT THERE WERE NO PROCESS DEVIATIONS, ANOMALIES OR REMARKABLE EVENTS DURING RECENT PRODUCTION OF SOFT PLUG SILICONE PUNCTUM PLUGS. A REVIEW OF THE GAMMA RADIATION STERILIZATION CYCLE SHOWS THAT THE STERILIZATION CYCLE WAS WITHIN VALIDATED PARAMETERS, NO ANOMALIES WERE NOTED. A PRODUCT BIO-BURDEN AUDIT IS PERFORMED RANDOMLY THROUGHOUT THE YEAR AND HISTORICALLY THE BIO-BURDEN LEVELS FOR THE SILICONE PLUGS ARE WITHIN ACCEPTABLE RANGE. A REVIEW OF THE OASIS MEDICAL, INC CUSTOMER COMPLAINT HISTORY FILES SHOWS THAT THERE HAVE BEEN NO OTHER REPORTED ADVERSE EVENTS OF THIS NATURE FOR SOFT PLUG SILICONE PUNCTUM PLUGS. >< TWO PTS WITH THE IDENTICAL PROBLEM WITH PLUGS. PLUGS HAVE BEEN EXTRUDED, BUT TETHERED TO THE LID MARGIN BY WHAT APPEARS TO BE PYOGENIC GRANULOMATOUS TISSUE. NO PROBLEMS REMOVING THE PLUGS AND THE PTS HAVE NO EFFECTS. IN ONE PT, DOCTOR WAS UNABLE TO PLACE ANOTHER PLUG PRESUMABLY DUE TO OCCLUSIN OF THE INTERIOR ASPECT OF THE PUNCTUM BY GRANULATION TISSUE. BOTH PTS REMAINED SYMPTOM-FREE. | 0
981883 | PATIENT IN OPERATING ROOM ON VENTILATOR. VENTILATOR WAS TURNED OFF TO ASSESS PATIENT. WHEN VENTILATOR SWITCHED BACK ON, VENTILATOR WOULD NOT TURN ON AND GAVE AN ERROR ALERT OF 'CANNOT DRIVE BELLOWS.' ERROR ALERT WAS NOT CONSISTENT WITH THE VENTILATOR MALFUNCTION SO STAFF DID NOT KNOW THE REAL PROBLEM. THE REAL PROBLEM WAS MOISTURE ON THE INSPIRATORY/EXPIRATORY SENSORS. CONTACTED VENDOR WHO SUGGESTED USING THE AESPIRA FLOW SENSORS FOR BETTER MOISTURE TOLERANCE IN FUTURE. NO REAL SOLUTION GIVEN TO THE MISLEADING ALERT. NO PATIENT HARM OCCURRED BECAUSE PT WAS BEING MANUALLY VENTILATED AT THE TIME OF VENTILATOR FAILURE. | 0
1085550 | THE SITE RADIOLOGIST REPORTED THAT THE DEVICE DISPLAYED INFORMATION OF A DIFFERENT PATIENT WHEN A PATIENT JACKET (A DROP DOWN LIST CONTAINING A LIST OF THE PATIENT'S EXAMS) WAS OPENED. THE ISSUE RECURRED AFTER THE PATIENT JACKET WAS CLOSED AND REOPENED. NO INJURY OR MISDIAGNOSIS WAS REPORTED. >< THE REPORTED ISSUE WAS DUPLICATED THROUGH INTERNAL INVESTIGATION. IT WAS FOUND THAT THE PACS SYSTEM DOES PROVIDE THE USER WITH A NOTICE IN THE JACKET HEADER THAT THE JACKET DOES NOT MATCH THE CURRENT EXAM WHEN AN INCORRECT PATIENT JACKET APPEARS. THE PATIENT JACKET, ALTHOUGH NOT MATCH WITH THE CURRENT EXAM, DOES PROVIDE THE CORRECT INFORMATION TO THE USER ABOUT ITS CONTENTS, INCLUDING THE PATIENT NAME AND IDENTIFIER. THE CONTENTS, IF SELECTED, MATCH THE PATIENT JACKET AND DISPLAY THE SAME NAME AND IDENTIFIER. THIS MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. THIS ACTIVITY IS BEING COMPLETED AS PART OF A CORRECTIVE ACTION TAKEN PURSUANT TO AN FDA INSPECTION CONDUCTED AT THE FACILITY IN 2008. | 1
1085557 | THE SITE REPORTED THAT THE APPLICATION WOULD INTERMITTENTLY OPEN THE REPORT PALLET WITH A REPORT THAT WAS HIGHLIGHTED IN THE BROWSE WORKLIST WITHOUT REQUEST. THE RESULTING REPORT DID NOT MATCH THE EXAM THAT WAS BEING VIEWED. NO INJURY OR MISDIAGNOSIS WAS REPORTED. >< THE REPORTED ISSUE WAS DUPLICATED THROUGH INTERNAL INVESTIGATION. IT WAS FOUND THAT THE PACS SYSTEM DOES PROVIDE THE USER WITH A NOTICE IN THE JACKET HEADER, THAT THE JACKET DOES NOT MATCH THE CURRENT EXAM WHEN AN INCORRECT PATIENT JACKET APPEARS. THE PATIENT JACKET, ALTHOUGH NOT MATCHED WITH THE CURRENT EXAM, DOES PROVIDE THE CORRECT INFORMATION TO THE USER ABOUT ITS CONTENTS, INCLUDING THE PATIENT NAME AND IDENTIFIER. THE CONTENTS, IF SELECTED, MATCH THE PATIENT JACKET AND DISPLAY THE SAME NAME AND IDENTIFIER. THE MDR IS BEING SUBMITTED LATE, AS THIS ISSUE WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF COMPLAINT FILES. THIS ACTIVITY IS BEING COMPLETED AS PART OF A CORRECTIVE ACTION TAKEN PURSUANT TO AN FDA INSPECTION. | 1
1025003 | PRIOR TO EACH USE OF THE ANESTHESIA UNIT ANESTHESIOLOGISTS PERFORM THE MANUFACTURER RECOMMENDED COMPLETE SYSTEM CHECKOUT PROCEDURE TO ENSURE THE INTEGRITY OF THE UNIT AND BREATHING ACCESSORIES. IN RECENT CASES, THE SYSTEM CHECKOUT HAD 'FAILED' DUE TO A LEAK LOCATED AT THE EXPIRATORY PORT ON THE COMPACT BLOCK ASSEMBLY. THIS LEAK IS ATTRIBUTED TO A LOOSENING OF THE LOCKING RING ON THE EXPIRATORY PORT AND RESOLVED WITH THE RETIGHTENING OF THE RING ITSELF. THIS SYSTEM CHECK FAILURE DID IMPACT PT CARE, THE PT WAS RELOCATED INTO ANOTHER SURGERY SUITE AND ALL STERILE INSTRUMENTS WERE OPENED AND UNABLE TO BE USED. THE POSSIBILITY OF COMPROMISING THE CRITICAL VENTILATION SUPPORT TO A PT WHILE THE UNIT IS IN USE HAS BECOME APPARENT AND NOW HAS AN ADVERSE EFFECT ON OUR PT CARE. | 0
1250243 | DURING A MEDICATION PASS, THE BAR CODE SCANNER FAILED AND THE NURSE WAS UNABLE TO SCAN THE PATIENT'S ID/MEDICATION. THE DEVICE FAILED AND THERE WAS NOT ENOUGH TIME TO TROUBLESHOOT OR RESET THE DEVICE. THE MEDICATION HAD TO BE GIVEN STAT AND THE ERROR OCCURRED WHEN THE SCANNER FAILED AND THE RN WASN'T ABLE TO VERIFY THAT THE WRONG DOSE WAS BEING GIVEN TO THE PATIENT. | 1
1013906 | RUI JOYSTICK EMERGENCY STOP BUTTON MISSING. REMOTE USER INTERFACE (RUI) PROVIDES REMOTE C-ARM MECHANICAL MOVEMENTS. THE EMERGENCY STOP BUTTON WAS DAMAGED OR DESTROYED AND WOULD NOT BE AVAILABLE FOR USE IF NEEDED. >< A GE SERVICE REPRESENTATIVE INVESTIGATED THE ISSUE AND FOUND A DAMAGED REMOTE USER INTERFACE THAT WAS REPLACED. WHILE ON SITE, THE SYSTEM DISPLAYED A LOW MA ERROR AND THE CALIBRATION FILES WERE RELOADED TO RESTORE FUNCTION. THE SYSTEM WAS TESTED, FOUND TO BE OPERATING AS INTENDED, AND RELEASED TO THE CUSTOMER FOR USE. NO PATIENT INJURY WAS REPORTED AS RESULT OF THIS COMPLAINT. THE FACILITY WAS UNABLE TO, OR REFUSED TO PROVIDE ANY FURTHER PATIENT INFORMATION WITH RESPECT TO THIS ISSUE. | 1
956723 | IT WAS REPORTED TO GE THAT THE TIP OF THE INVASIVE ULTRASOUND TRANSDUCER MELTED DUE TO APPARENT OVERHEATING OF THE TRANSDUCER ARRAY. NO INJURY WAS REPORTED. CUSTOMER HAS BEEN ADVISED TO STOP USING THE SYSTEM AND THE PROBE IS RETURNED TO GET FOR FURTHER INVESTIGATION. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1013707 | CUSTOMER ALLEGES THAT WHILE MONITORING AN INFANT WHO REPORTEDLY HAD A TRACHEOSTOMY AND HYPOXIA, THE TRUSAT OXIMETER DID NOT ALARM AT A SETTING OF 90 OR BELOW. THE INFANT REPORTEDLY DIED. PER DISTR, THEY DO NOT HAVE THE UNIT NOR HAVE THEY BEEN ABLE TO REACH THE FAMILY TO OBTAIN THE UNIT. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE UNIT FROM THE PT'S FAMILY. THE OXIMETER HAS NOT BEEN RETURNED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
1212025 | RAPID RESPONSE CODE CALLED FOR THIS PT. THE 12 LEAD DIAGNOSTIC EKG MACHINE REQUESTED BY PHYSICIAN WAS HOOKED UP TO THE PT, BUT THE TECHNICAL PARTNER WAS UNABLE TO OBTAIN AND EKG BECAUSE THE EKG MACHINE HAD 'MEMORY WAS FULL' MESSAGE DISPLAYED. THE EKG MACHINE WOULD NOT DISPLAY THE WAVEFORM EVEN AFTER THE CARDIOLOGIST ATTEMPTED TO RESOLVE THE ISSUE BY POWER CYCLING THE EKG MACHINE. THIS CAUSED A DELAY IN THE RAPID RESPONSE CODE. CARDIOLOGIST WAS FINALLY ABLE TO GET A WAVEFORM STRIP FROM CENTRAL WIRELESS MONITORING SYSTEM. | 1
1266724 | REPORT REC'D OF A USER ERROR RESULTING IN A SUCTION SWAB DISENGAGEMENT. RPTR STATED HER BROTHER (PT) WAS ADMITTED IN 2008, TO HOSP FOR A CRANIOPLASTY PROCEDURE. THE PT WAS TRANSFERRED TO REHAB FACILITY THE FOLLOWING DAY, AT 5:30 PM. AT APPROX 1:45 AM THE NEXT DAY, THE PT'S VITAL SIGNS BECAME UNSTABLE. AT APPROX 2:00 PM, A CODE BLUE WAS CALLED. SHORTLY AFTER THE CODE BLUE, THE PT WAS TRANSFERRED BACK TO HOSP. A DR PERFORMED A BRONCHOSOCPY PROCEDURE AND RETRIEVED A SWAB HEAD FROM THE PT'S LUNG. RPTR STATED THE SWAB HEAD WAS FROZEN AND SENT TO PATHOLOGY FOR EVAL. RPTR STATED THE PT SUBSEQUENTLY DEVELOPED PNEUMONIA. RPTR STATED THE INVOLVED SUCTION SWAB WAS A SAGE #6550, TOOTHETTE ORAL CARE SUCTION SWAB. LOT INFO WAS NOT AVAILABLE FROM RPTR. THOUGH REQUESTED, NO ADD'L INFO WAS AVAILABLE. FACILITY CONTACT STATED ON ORIGINAL DATE, THE PT WAS ADMITTED TO HOSP FOR A CRANIOPLASTY PROCEDURE. THE PT REC'D ORAL CARE WITH A #6550, TOOTHETTE SUCTION SWAB AND REPORTEDLY BIT DOWN THROUGH THE STICK OF THE SWAB. REPORTEDLY, X-RAYS WERE PERFORMED; HOWEVER, VISUALIZATION OF THE SWAB HEAD WAS NOT POSSIBLE. REPORTEDLY IT WAS ASSUMED THAT THE PT SWALLOWED THE SWAB HEAD. THE PT WAS TRANSFERRED TO REHAB FACILITY THE FOLLOWING DAY. CONTACT STATED THAT SOMETIME THE NEXT DAY, THE PT DEVELOPED OXYGEN DESATURATION ISSUES AND AMBU BAGGING VIA THE PT'S TRACHEOSTOMY TUBE WAS INITIATED. THE CONTACT SPECULATED THAT THE SWAB HEAD WAS ABOVE THE TRACHEOSTOMY TUBE CUFF AND WHEN THE TRACHEOSTOMY CUFF WAS MANIPULATED AT FACILITY, THE SWAB HEAD MIGRATED INTO THE PT'S LUNG. IT WAS REPORTED THAT THE PT WAS TRASNFERRED BACK TO HOSP THAT DAY AND A BRONCHOSCOPY WAS PERFORMED. CONTACT STATE THE PIECE OF FOAM WAS REMOVED FROM THE PT'S LUNG AND SENT FOR PATHOLOGY TESTING. THE PT RETURNED TO CENTER THE FOLLOWING DAY. >< FACILITY CONTACT: RN, HOSPITAL -RISK MGMT. INSTRUCTIONS FOR USE STATE, 'A BITE BLOCK SHOULD BE USED WHEN PERFORMING ORAL CARE ON PTS WITH ALTERED LEVELS OF CONSCIOUSNESS OR THOSE WHO CANNOT COMPREHEND COMMANDS.' CONTACT STATED A BITE BLOCK WAS NOT USED AT THE TIME OF THE INCIDENT. CONTACT STATED THE INVOLVED SUCTION SWAB WAS SAVED; HOWEVER, DECLINED TO RETURN TO SAGE PRODUCTS, INC FOR EVAL. LOT INFO WAS NOT AVAILABLE FROM CONTACT. THOUGH REQUESTED, NO ADD'L INFO WAS AVAILABLE. | 0
976896 | THERE WAS AN ACUDOSE CABINET ERROR. THE CABINET LOCKED UP AND BECAME INOPERABLE. THERE HAVE BEEN MULTIPLE PROBLEMS OF THIS NATURE WITH THIS CABINET IN THE PAST. OUR ADMINISTRATION AND PHARMACY DIRECTOR IS DEALING WITH MCKESSON, BUT THEY ARE SLOW TO RESPOND. MCKESSON HAS HAD TECH SUPPORT HERE BUT THEY CAN NOT FIGURE OUT THE PROBLEM. ADMINISTRATION FEELS THAT MCKESSON HAS NOT BEEN RECEPTIVE TO OUR ISSUES WITH THE ACUDOSE CABINET, NOR ARE THEY CONFIDENT ABOUT SOLVING THE PROBLEMS. | 1
1201963 | REPORTER STATES THAT IN 2008, HER BROTHER'S VITAL SIGNS DIMINISHED WHILE IN THE HOSPITAL, AND A 'CODE BLUE' HAD TO BE CALLED. WHEN THE DOCTOR TRIED TO VENTILATE THE PT, AN OBSTRUCTION WAS FELT AND A BRONCHOSCOPE WAS DONE. A PIECE OF GREEN SPONGE WAS FOUND IN THE PATIENT'S LUNG, REPORTER WAS TOLD THAT THIS WAS THE CAUSE OF HER BROTHER'S LUNGS TO COLLAPSE. REPORTER STATES THAT SHE HAS RESEARCHED AND FOUND SIMILAR REPORTS ON WEB. | 0

## 2009
1016382 | this computed radiography system with no radiology information system (ris) sent images to facility, picture archiving and communication system (pacs). however duplicate unique identifiers (uid) caused patient images in the pacs to be merged together into one folder. the files are kept intact so, the values are not overwritten or changed. there have been no patient injuries or misdiagnosis. it is not known at this time if the phillips or the candelis system is causing this problem. >< the investigation is still ongoing on this event. when the investigation is completed, a follow-up report will be sent to the fda. | 1
1480250 | the user reported that, after experiencing problems with their primary monitoring device (see report 1051786-2009-00009), the user switched to their secondary monitoring device. the user reported that they experienced difficulty obtaining a good ecg waveform with the secondary monitoring device and this was the same problem they experienced with the primary monitoring device. while the user was troubleshooting the monitoring devices, it was reported that the pt coded. the pt was successfully revived. the user claimed that the pt was connected to both the primary and secondary monitoring devices at the same time when event occurred. >< the user facility's biomedical engineer reported to the device manufacturer that he evaluated the device, and could not find any problems, but was requesting the device manufacturer's assistance in evaluating the device further. the device manufacturer's field service representative traveled to the user facility, and evaluated the device involved in the reported event. the device manufacturer's field service representative could not find any problems with the device while on site. the device manufacturer's field service representative performed operational and functional testing on the device according to internal procedures, using the same accessories that were in use during the reported event. all testing conducted by the device manufacturer's field service representative passed, and no malfunction of the monitoring system was detected. the device manufacturer's field service representative consulted with the user facility's biomedical engineer regarding maintenance inspections that, per the instruction for use, should be done routinely on the device as well as on an annual and semi-annual basis. also, in-service training covering basic operation of the device (per the instructions for use) was requested by the user facility and this training was conducted during this site visit, as well in depth application training was also requested by the user facility. this training is currently scheduled to occur in early october. the user facility reported to the device manufacturer that the device is back in service and will continue to be utilized as a secondary/back-up monitoring device. the device manufacturer cannot rule out the use of the device as a contributing factor in the event. while the device manufacturer's field service representative was conducting his site visit, the user stated that they were not familiar with this monitoring equipment, and that there was a lot of confusion in the mri suite at the time of the event. the users stated that they were fumbling with the monitoring device while trying to establish monitoring for the pt. while the user was troubleshooting the monitoring devices, the pt coded but was successfully revived. | 1
1480251 | the user reported experiencing difficulty obtaining a good ecg waveform with the device, so the user switched to another monitoring device. the user reported that they experienced the same problem with the secondary monitoring device (see report 1051786-2009-00010). while the user was troubleshooting the monitoring devices, it was reported that the pt coded. the pt was successfully revived. the user claimed that the pt was connected to both the primary and secondary monitoring devices at the same time when the event occurred. >< the user facility's biomedical engineer reported to the device manufacturer that he elevated the device and could not find any problems, but was requesting the device manufacturer's assistance in evaluating the device further. the device manufacturer's field service representative traveled to the user facility and evaluated the device involved in the reported event. it was found that the ecg cable used during reported event was physically damaged at the ra and la leads. additionally, the device's wireless spo2 module was cracked, the top of the device's water trap was missing and the wireless module batteries had broken side locking tabs. both the missing top of the water trap and the damaged ecg lead set would be obvious in either any visual inspection or in specified maintenance. the cracked wireless spo2 module and batteries with broken side locking tabs are cosmetic problems only and would not have an impact on product performance. the device manufacturer's field service representative performed operational and functional testing on the device according to internal procedures. all testing was conducted with a new ecg lead wire set and a new water trap. all testing conducted by the device manufacturer's field service representative passed and no malfunction of the monitoring system was detected. the device manufacturer's field service representative consulted with the user facility's biomedical engineer regarding maintenance inspections that, per the instructions for use, should be done routinely on the device as well as on an annual and semi-annual basis. also, in-service training covering basic operation of the device (per the instruction for use) was requested by the user facility and this training was conducted during this site visit as well. in depth application training was also requested by the user facility. this training is currently scheduled to occur in early october. the device manufacturer concludes that the condition of the ecg lead wires was the cause of the device's reported failure to function as expected. use of the user facility's damaged ecg lead wires while attempting to monitor a pt produced an ecg signal/waveform of non-clinical quality. this quality of ecg signal/waveform would have been recognizable, and obvious to the user as incorrect. the device manufacturer believes that the device provided an appropriate response to the deteriorated condition of the ecg lead wires. the device manufacturer does not consider the condition of the ecg lead wires or the poor ecg signal to be a malfunction of the device since the instructions for use provides precautions against the use of accessories that are not in proper working condition. per the instructions for use: cleaning: 'inspect the accessory for any cracks, holes, tears, cuts, etc., that could affect operation, and replace as necessary. 'before connecting to a pt': 'step 3. ensure that all pt accessories being utilized (nibp hoses, cuffs, ecg lead wires, etc.) are listed in the accessories section of this manual. visually inspect all accessories for damage. replace any accessories that show signs of damage or abuse.' the user facility reported to the device manufacturer that the device is back in use. additionally, during their site visit, the device manufacturer's field service representative observed several mri cases with the device in use and no issues were noted. the device manufacturer cannot rule out the use of the device as a contributing factor in the event. while the device manufacturer's field service representative was conducting his site visit, the user stated that they noticed the pt's etco2 reading deteriorating. despite the deterioration in the pt's etco2 reading, the user claimed that they believed that the device's reading was erroneous due to the ecg problems they had been experiencing with the device (as described above). as a result, the pt was not checked; the pt subsequently coded, but was successfully revived. | 1
1083564 | baxa was notified on 06/25/2008 that a patient had received a tpn bag containing a higher-than-desired quantity of trace elements. the bag was infused for 18 hours and then removed once the issue was discovered. the patient's blood work indicated higher-than-expected levels of zinc and chromium. however, no additional medical intervention was needed. the patient is in stable condition and did not experience any adverse symptoms as a result of the infusion. >< the customer's subsequent investigation revealed that the pharmacist, who entered the subject order within the abacus calculating software, mistakenly entered a value of 3ml/kg of pediatric trace elements in addition to 3ml of adult trace elements. the facilities procedure states that patients should receive adult trace elements and no pediatric trace elements. pharmacist verification of the ordered ingredients failed to identify the additional 84ml of pediatric trace elements. with the assistance of baxa technical support, the importance of abacus software warning limits and ingredients dosing rules were addressed with the customer and a new catastrophic warning limit was created for pediatric trace elements for all pediatric and neonatal patients. this warning limit will notify the user when the value for pediatric trace elements has exceeded the maximum value of 5ml. investigation has confirmed that the abacus calculating software performed as designed and delivered the volume of pediatric trace elements as ordered. | 1
1357884 | baxa was notified of an issue in 2009 related to a pt involvement incident. it was reported to baxa that a tpn bag created for this pt was infused at an incorrect rate (501 ml/hr instead of 50.1 ml/hr). after the infusion, the pt showed signs of volume overload and hyperglycemia. the pt was given furosemide and additional insulin. the pt was subsequently released from the hospital one week later with no apparent long term injury or illness resulting from this issue. >< pharmacy reported to baxa that the flow rate on the tpn bag label which read 50.1 ml/hr was misread as 501 ml/hr and infused into a pt. order-entry software was used to produce the bag label for the tpn solution. the customer was able to provide their database and copies of labels produced by software for the subject bag created in 2009. only the software and labels were evaluated; as the pump used for infusion was not manufactured or supplied by baxa. an eval of the software bag label for the subject bag indicates the solution to be administered via central line only at a flow rate of 51.76 ml/hr for 24 hours. the initial reporter was contacted about the discrepancy of what was originally reported for the flow rate value. the customer indicated that they are not sure why 50.1 ml/hr was reported to him. he was told that the flow rate was misread as 501 ml/hr. the label was assessed to determine if the text was legible and correct. in conclusion, the text was found to be adequate. in order to investigate the issue, the additional following items were evaluated: database review summary: the database shows the software settings in place at the time of database retrieval. results of eval: an eval of the database for the subject order revealed a rate of 51.76 ml/hr for 24 hours. all templates and formulary items were examined for additional issues; none were found. all values for ingredients are confirmed correct and accurate. based on an eval of the documentation associated with this event, it was concluded that the software performed as designed and produced a correct and legible bag label for the tpn solution. to prevent recurrence of this incident, the facility worked with baxa technical support to modify the label so that it displays only the whole number portion of the rate. | 1
1300683 | a hospital was able to provide their abacus database, copies of the prescriber order form, copies of labels produced by abacus and the compounder mixcheck report for the subject bag created in late 2008. the prescriber order has instructions to include dextrose 15% in the subject bag. an evaluation of the compounder mixcheck report and the abacus solution label indicates that the compounder delivered dextrose as ordered in abacus (595.7 ml). it was also determined that the mixcheck report indicates that the total compounded volume of the subject bag was within +/-3% of the ordered volume (actual % difference -0.09%). in order to investigate the issue, the additional following item was evaluated: abacus database review summary: the abacus database shows the software settings in place at the time of database retrieval. results of evaluation: the abacus database revealed that this facility created a 12 hour cyclic infusion for the subject order. dextrose was ordered using a glucose infusion rate unit (i.e. mg/kg/min) which is based off of a 24 hour period. the pharmacist, who entered the subject order, mistakenly entered a value of 16.9 mg/kg/min for the dextrose ingredient; thus, doubling the dextrose concentration. abacus displayed the glucose infusion rate calculation on the order screen and showed the discrepancy between the 'as ordered' amount and the calculated amount. the abacus database showed that the software displayed a warning that the dextrose concentration order value of 30.89% (twice the ordered concentration) exceeded the limit value of 25.00%. rational was given as 'under curve' and the order was completed. the customer's subsequent investigation revealed that the pharmacists calculated glucose infusion rate was incorrect and the abacus software was correct based on their data entry. if they wanted a final dextrose concentration of 15% according to the prescriber order, the glucose infusion rate should have been 8.21 mg/kg/min. the facility looked at another pediatric patient with a cyclic infusion, and did not see any problems. to prevent recurrence of this incident, the facility changed their ordering methods. dextrose will now be ordered as gram/day and not as mg/kg/min. based on an evaluation of the documentation associated with this event, it was concluded that the higher-than-desired quantity of dextrose was the result of the user entering a glucose infusion rate of 16.9 mg/kg/min rather than 8.21 mg/kg/min. investigation has confirmed that the abacus software performed as designed and correctly calculated the concentration of dextrose. actual device involved in incident was evaluated. device was evaluated at customer site by baxa personnel. results: device performed according to specification. conclusions: device operated according to specifications. user error caused event. >< in late 2008, baxa was notified by the customer of a patient involved incident. the customer uses abacus v2.0 tpn calculating software for tpn production. during the same month, tpn feeding was administered to a home care patient who subsequently became hyperglycemic. it was determined that the tpn bags contained a higher-than-desired quantity of dextrose. no additional medical intervention was needed and no apparent long-term injury or illness resulted from this issue. | 1
1541737 | pt is receiving dorpenem 1g in 120 ml ns mixed in an accuflo device. accuflo devices are leaking. medication had to be replaced. dates of use: 2009. diagnosis or reason for use: 494.1. | 0
1410786 | during surgery on the anesthesia machine, while being used, it alarmed displaying 'machine will shut down in 8 seconds'. the anesthesiologist turned on alternate oxygen control then manually bagged pt. he then turned the machine off and on. the machine re-booted continuing to work finishing the surgery without any further events. no alteration in surgery or pt surgical care. there was not any negative outcome. diagnosis or reason for use: exploratory laparatomy. | 0
1413879 | pt a presented for dosing of methadone. identification verified-dosing screen appeared-dose given was later discovered to be pt bs dose. ams-software- did not record pt as dose although pt as signature was there. pt a received 5mg more of methadone syrup than ordered. serious software malfunction-company notified-no harm to pt. | 1
1275443 | ge / datex / ohmeda call. we experienced two caster failures in four months on a s/5 avance anesthesia machine. the first failure was in 2008, the second failure was four months later. the caster failed when going over an elevator gap of about one inch. the previous failure was related to a simple floor threshold that was not much more than 1/4' high. these failures have created dangerous conditions of the machine becoming unstable, tipping and almost falling over on its side. i have the caster that failed on the same day, and will return it to ge healthcare/datex/ohmeda when provided with shipping information. a review of our data on the other aestiva / avance anesthesia machines we have here show one other failure, which was given to our service rep at the time, a failure of s/5 avance machine happened in 2007. in the same month, this wheel was given to rep, for additional ge /datex/ ohmeda failure analysis. we never received a response, although it was not a formal complaint. for our population of 11 avance systems, our overall caster mtbf is 10 years, but we have two units with caster mtbf ranges of 1.19 years to 2.46 years, which may indicate a lot/batch problem with a series of casters, or a problem with the machine bases. | 0
990982 | it was reported that the monitor was sounding a distorted alarm followed by no alarm sound. heart rate beep was available. no indicator received that the reported situation caused or contributed directly or indirectly to a death or serious injury. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 1
992275 | it was reported that the customer noticed the patient was turning light so she was about to administer a medication, but didn't get a chance to, the patient started to de-sat. the primus gave pinsp no attained. she checked for leaks the situation got worse. she switched to man spont once she did that, she attempted to push the o2 flush but, nothing came out. she switched to pressure mode and tried the o2 flush also nothing, the patient was turning blue at this point. she then switched man spont and bagged the patient. the o2 supply for the ambu bag was the flow meter on the side of the primus, and it supplied the o2 required. no permanent injury reported. >< event is under investigation. results will be reported in the follow up report. | 0
1518011 | the apnea monitor was returned for evaluation and tested in accordance with product labeling (ref. smartmonitor 2, parent's guide, 02/25/2003 and smart monitor 2 checkout manual, part number 1020818, 09/07/2006). the device passed both the self-test procedure identified in the parent's guide, and the service checkout procedure identified in the device's 'checkout manual' (this procedure is used to verify the safety and efficacy of devices by the manufacturer's service centers after devices are serviced and before they are returned to customers). the monitor's audible alarm output met product specifications and is compliant with the fda guidance, class ii special controls guidance document: apnea monitors; guidance for industry and fda, july 17, 2002). data downloaded from the monitor's memory log revealed that on the day of the reported event (2009) twelve bradycardia events occurred. the memory log also indicates the monitor was turned off. based on this level of alarm activity, we believe the monitor operated as designed prior to and during the reported event. although our attempts to acquire clarification from the customer regarding their specific complaint were unsuccessful, we reviewed the smart monitor 2 product labeling to determine if it contained errors or inadequacies related to the intended use of the monitor. the smartmonitor 2 international parent's guide (warnings and cautions) provides the following warnings to a caregiver of a pt being monitored by the apnea monitor. smartmonitor 2 is a monitoring device only. it does not prevent the loss of breathing or heart activity, nor will it restore breathing or heart activity. it will not prevent death. anyone using the smartmonitor 2 to monitor an infant should be trained in current infant cardiopulmonary resuscitation (cpr), which is a proper way to restore breathing and heart activity. we believe this labeling is both appropriate and adequate for describing the inherent risks and requirements associated with apnea monitoring. based on all information available, we believe the smartmonitor 2 did not cause or contribute to the reported event and that no further action is appropriate. >< a complainant reported pt expired as a result of an apnea monitor not alarming as designed. the reported event occurred while the infant was being transported to the hospital (by his family). the reason the patient was being transported to the hospital and the nature of any pre-existing conditions could not be obtained. when contacted, the durable medical equipment supplier (dme) reported, they had received conflicting information from the patient's family; the grandmother reported the apnea monitor did not annunciate an alarm at the time of the adverse event, and the father stated the monitor did alarm, but did not 'give them enough time to do anything'. | 0
1494523 | our rep. reports that during an arthroscopic knee repair, a portion of the distal end of a fms scope sheath broke off into the patient's joint space. the fragment was easily retrieved and the procedure was concluded successfully without further issue or harm to the patient. >< mitek is at this point in time in the information gathering mode. when all that can be had, is had and thoroughly investigated and evaluated, those results will be the subject matter in a follow-up report. | 0
1403491 | a service call was made. the customer's report of misinterpreted sample barcodes was not reproduced. replaced 14-lane sample bay to prevent further barcode misreads, and verified all alignments for sample/reagent probe at all bay and plate positions. the instrument operated within specification. >< the customer stated that the galileo misread a barcode. 4 samples were in a sample rack and were being loaded for type and screen. the first time the rack was loaded, one of the samples did not read. the rack was removed and reloaded on the instrument. all of the samples read, but one of the numbers on a sample was wrong. | 1
1343269 | customer reported an anti-d antibody was missed on galileo while performing the pool cell assay. >< reactivity of the d antigen was confirmed on retention capture-r ready screen (pooled), lot n189 used by the customer at the time of the event. customer's returned sample exhibited 3+ hemolysis and was qns for galileo testing. hemagglutination tube testing was performed using retention panoscreen i, ii, and iii, lot 02212. testing was performed at all temps and immuadd was used as potentiator. sample was nonreactive in all testing. | 0
1220004 | customer reported unexpected positive reactions(1+) with anti-b on echo when running a confirmation test for a known group o donor unit. the donor unit resulted as group b instead of group o. >< the customer confirmed the test result as o positive upon repeat testing. review of the result files from the device shows a well score of 20 in well 3, anti-b well, resulting in a 1+ reaction. review of well image shows a line of red colored artifact (hair?) present. | 0
1220006 | the e antigen was confirmed on capture-r ready id, lot id106 and retention capture-r ready-sreen (3), lot r028.the customer's returned samples (same patient, 2 draw dates) were tested with retention capture-r ready-screen (3), lot r028 on an in-house echo. both samples were interpreted as negative by the echo. the returned samples were tested by tube hemagglutination tests with retention panoscreen i, ii and iii, lot 35626, using immuadd as the potentiator. very weak reactivity was observed with cell ii only at the indirect antiglobulin test. >< customer reported unexpected negative reactions using capture-r ready screen 3 (crrs) and capture-r ready id (crrid) on the echo. the patient had a known anti-e antibody. | 0
1222624 | customer reported unexpected negative reactions on the echo with capture-r ready id. a patient with a history of an anti-d resulted with a negative screen and id on the echo. the anti-d was not rhig induced. >< reactivity of the d antigen was confirmed with retention crrs(3), lot r028 and crrid, lot id105. these reagents were used by the customer at the time of the event. the customer did not return product or sample for investigation testing. | 0
1222625 | the camera was changed on the echo; however, the problem persisted. the instrument was replaced, and the new instrument functions as expected. >< customer reported a patient sample had a positive antibody screen on galileo, and negative antibody screen on echo. | 0
1234603 | customer reported unexpected negative reactions on the echo. a patient specimen with a history of anti-c tested for crossmatch produced various results on the echo. of the eight donors for crossmatch, 3 were 2+ incompatible, 1 was no interpretation, and 4 were negative (compatible). three of the four negative (compatible) donors were 4+ incompatible for c antigen and one donor was negative with tube testing @ immediate spin . >< compatibility testing was performed with customer's returned patient sample (reported to contain anti-c), using two c-negative (r1r1) and two c-positive (rr and r2r2) red cell donor segments on an in-house echo, with retention capture-r select. sample was compatible with c-negative segments. the crossmatch with the r2r2 segment was interpreted as equivocal (visually positive), and the rr segment was interpreted as incompatible. | 0
1234613 | reactivity of the e antigen on retention capture-r ready-screen (pooled), lot n174, was confirmed. the customer did not return sample or product for investigation testing. it is not possible to rule out the sampel as a cause of the event. the limitations section in the package insert for capture-r ready-screen (pooled cells) states: 'antibody screening tests employing pooled reagent red blood cells will not be as sensitive as those incorportating the red blood cells of unpooled single donors.' >< customer reported that a donor sample was tested on the galileo for antibody screen using the pool cell assay and produced a negative result. the sample was run on gel, where a 3+ result was obtained on screening cell ii, (e+e-); the antibody was identified as anti-e. | 0
1234625 | aborh testing was performed with in-house donor samples of known abo/rh types using retention anti-a, lot 101676, anti-b series 3, lot 203228, and anti-d series 4, lot 504696, on an in-house galileo. these lots were used by the customer at the time of the event. all in-house donor samples typed as expected. fwd_aborh testing was performed with customer's returned donor samples using retention anti-a, lot 101676, anti-b series 3, lot 203228, and anti-d series 4, lot 504696, on an in-house galileo. results obtained matched the customer's manual typing results. >< customer indicated abo discrepancies occurred on donor samples tested on galileo. the customer stated no adverse events occurred as a result of the abo discrepancies. | 0
1241304 | customer reported a donor sample reported as ab positive on the galileo using the reflex fwd assay. the sample was a first time donor and had been tested on the olympus pk prior to being tested on the galileo. this interpretation didn't match the olympus pk, which reported the sample as a positive. >< pcanywhere was used to retrieve the instrument images. review of the images showed a fibrous clot in the well, giving the well a positive appearance. followed up with customer, she stated that she had pulled the tube and checked for clots. she stated that there wasn't a clot in the tube but the sample was 'sticky'. sample contained fibrin. customer is aware that samples that are clotted or contain fibrin shouldn't be used for testing on the galileo. the mistype was caused by the clot in the sample. the galileo operator manual indicates the galileo can not detect very large clots; clotted samples should not be tested on galileo. | 0
1256804 | customer reported unexpected negative reactions on the echo. a patient sample had a negative panel on the echo, but an anti-fya was identified in gel. the patient was not transfused. >< the camera was adjusted using blud-direct. the alignment adjustment was made to prevent any further camera alignment issues in the future. the instrument was operating as expected. reactivity of the fya antigen was confirmed on returned capture-r ready id (crrid), lot id105. this lot was used by the customer at the time of the event. the customer's returned sample was tested on an in-house echo with returned crrid, lot id105. the sample exhibited 2+ to 3+ reactivity with all fya(+) reagent red cells. | 0
1272068 | reactivity of the and e and k antigens was confirmed on retention crrs(3), lot r031and crrid, lot id107. the customer used those regents at the time of the event. the customer did not return patient sample for investigation testing. it is not possible to rule out the sample as a cause of the event. >< customer reported unexpected negative reactions on the echo when testing a patient sample with a history of anti-e and -k. the sample resulted as negative with k+ cells on capture-r ready screen (crrs)3 and capture-r ready id (crrid). | 0
1272070 | customer reported an unexpected negative reaction on a patient sample tested on the echo. the sample tested antibody screen negative on the echo. the sample was tested by manual tube method using immuadd as potentiator and anti-e was detected. >< reactivity of the e antigen was confirmed on retention capture-r ready screen (3), lot r033, used by the customer at the time of the event. the customer did not return sample for investigation testing. | 0
1272071 | customer indicated that while performing validation testing using the pooled cell screens, known antibody positive samples were reported out as negative. the antibodies included anti-m, anti-d, anti-e and an anti-k. >< reactivity of the m, d, e, and k antigens was confirmed on retention capture-r ready-screen (crrs) pooled, lots n176 and n177. pool_cell testing was performed with customer's returned samples using retention crrs (pooled), lots n176 and n177, on an in-house galileo. one sample exhibited equivocal reactivity with crrs (pooled), lot n176, and 1+ reactivity with crrs (pooled), lot n177. all other samples were nonreactive in all testing. hemagglutination tube testing was performed with customer's samples using retention panoscreen i, ii, and iii, lot 40995. testing was performed at all temperatures; immuadd, lot 7c5511-1, was used as potentiator. two samples were nonreactive in all testing. the other samples demonstrated variable positive reactivity. the event appears to be related to the nature of the samples. | 0
1288864 | the customer did not provide additional details about the event. sample was not returned for investigation testing. >< customer reported a missed anti-e on the galileo on a patient sample known to contain anti-e. the galileo reported out negative reactions with the antibody screen and the antibody id assays. | 0
1306323 | customer indicated unexpected negative reactivity for 2 cell screen and antibody id assays on a patient sample tested by galileo. previous manual tube testing had identifed an anti-kell. the customer stated no adverse event occurred as a result of the unexpected negative reactivity. >< reactivity of the k antigen was confirmed on retention crrid, lot id107 and crrs (i and ii), lot x269. the reagents were used by the customer at the time of the event. the 2_cell testing was performed with the customer's returned samples (from same patient) using retention crrs (i and ii), lot x269 on an in-house galileo. samples were nonreactive in all testing. hemagglutination tube testing was performed with the customer's sample using retention panoscreen (i ii, and iii), lot 48123. testing was performed at all phases and immuadd, lot 357005-1, was used as potentiator. sample exhibited very weak (+w) reactivity at iat with the k+ cell and was nonreactive in all other testing. | 0
1306324 | customer reported an unexpected negative reaction on the echo. a patient with no history of an antibody resulted with a negative screen on the echo, but a 2+ incompatible autologous crossmatch was observed at ahg phase. customer performed extend i on the echo and an anti-k was identified. >< reactivity of the k antigen was confirmed on retention crrid extend i, lot dp032 and crrs(3), lot r036. the lots were used by the customer at the time of the event. the customer's returned samples were tested with retention crrs(3) lot r036 and crrid lot dp032 on an in-house echo. positive reactivity was observed with k+ reagent red cells of retention crrs(3) lot r036 and retention crrid lot dp032. we were not able to reproduce the customer's negative results. | 0
1354006 | customer reported unexpected negative reactions with capture-r ready screen pooled test wells (crrs) on a sample known to contain anti-jka. >< the customer returned unopened product, the sample and biorad control for investigation testing. returned and retention capture-r pooled, lot cw191, were tested. positive reactions with the returned and retention capture-r pooled, lot cw191 were seen with the sample, biorad control and an additional specimen known to contain anti-jka. | 0
1354007 | customer reported unexpected negative reactions with cell ii (e positive cell) of capture-r ready screen(3) when testing a patient sample containing anti-e on the echo. customer did not repeat testing on the instrument.no transfusions or adverse reactions occurred as a result of the unexpected negative reactions. >< customer's returned sample was tested with retention capture- r ready screen (crrs)(3), lot r042, used by the customer at the time of the event, on an in-house echo. the sample exhibited 1+ reactivity with the e+ reagent red cell. reactivity of the e antigen was confirmed on retention crrs(3), lot r042. | 0
1367304 | reactivity of the k antigen was confirmed on returned and retention capture-ready screen (crrs) (3), lot r041 and on retention crrid, lot id112. these products were used by the customer at the time of the event. the customer's returned sample was tested with returned and retention crrs(3), lot r041 on an in-house echo. a negative antibody screen interpretation resulted from the echo. the sample was tested with retention panoscreen i and ii, lot 07273 by tube hemagglutination tests, using immuadd as the potentiator. a room temperature (rt) incubation was included. the sample was nonreactive with all phases of testing. >< customer reported a patient sample that had a positive screen on manual capture, but was negative on the echo. | 0
1386906 | customer reported unexpected negative reactions with a patient sample containing anti-fya, tested on the echo. >< reactivity of the fya antigen was confirmed on retention crrs(3), lot r040, used by the customer at the time of the event.the customer's returned sample was tested on an in-house echo with returned crrs(3), lot r040. the sample was nonreactive. the returned sample was tested on an in-house echo with retention crrs(3), lot r040, and exhibited equivocal reactivity with cell i, the fy(a+b+) reagent red cell. the returned sample was tested by a tube hemagglutination test method with retention panoscreen i, ii and iii, using immuadd as the potentiator. a room temperature incubation was included. the sample exhibited +weak and 1+ reactivity with the fy(a+b+) and fy(a+b-) cells, respectively, at the indirect antiglobulin test phase. | 0
1468861 | a customer reported an unexpected negative reaction on 2_cell screen and capture-r ready id (crrid) assay on the galileo. customer did not retest the sample on the instrument. >< review of customer's instrument images confirmed the unexpected negative reactions. the reactivity of the c antigen was confirmed on retention capture-r ready-screen (crrs) i and ii, lot x292, and crrid, lot id117 with anti-c. these were the same lots used by the customer. the reactivity of the c antigen was confirmed on returned crrs (i and ii) and crrid with anti-c. manual testing was performed with customer's patient sample (4+ hemolysis) with returned and retention crrs (i and ii) and retention crrid. sample was nonreactive with all c+ cells. hemagglutination tube testing performed with customer's patient sample using c+c+, c+c-, and c- cells from retention panocell-10. testing was performed at all temperatures. sample exhibited very weak (w+) reactivity with c+ cells at iat only and was nonreactive with c- cell in all testing. the event appears to be related to the nature of the sample. | 0
1517910 | customer reported unexpected negative results when testing patient sample with known anti-e on the galileo. the sample resulted as negative for with both cells of capture-r ready-screen i and ii (crrs 2). >< review of instrument images: 2 cell plate. well e1- negative reaction, 28 reaction strength- visually the well appears weakly positive. well f1- negative reaction, 29 reaction strength- visually the well appears weakly positive. per the capture-r package insert: some igg antibodies have been shown to react poorly in solid phase red blood cell adherence assays. weak examples of clinically relevant antibodies may fail to react by cature-r ready-screen, even though the antibodies are detected by an alternative technique. the customer did not return product or sample for further serological testing. the sample could not be ruled as the cause of the event. | 0
1526418 | customer reported an unexpected negative results when testing patient sample using capture-r ready screen 3 (crrs 3) on the echo. the reactions look positive, but the instrument interpreted it as negative. the patient has a newly identified anti-e antibody. no adverse reactions occurred as a result of the unexpected negative reactions. >< review of the result files shows well score of 2 for cell ii, e positive cell, of crrs(3), lot #r059 resulting in a negative reaction. well image appears negative capture-r ready id (crrid) assay with same sample shows well scores of 8 for cells 1 & 3 and well score of 2 for cell 6 of crrid, lot id119 resultiing in equivocal reactions for cells 1 & 3 and negative for cell 6. well images for cell 1 appears negative and for cell 3 appears positive. customer decontaminated the instrument and repeated antibody screen and antibody id testing with the same lot of crrs(3) and crrid. crrs 3 shows well score of 2 for cell ii, e positive cell resulting in a negative reaction. well image appears negative. crrid shows wells scores of 70 for cells 1 & 3 and 1 for cell 6 of crrid resulting in a positive(3+) reaction for cells 1 & 3 and a negative reaction for cell 6. well image for cell 6 appears negative. confirmed the reactivity of the e antigen on retention crrs (3), lot r059, with anti-e. confirmed the reactivity of the e antigen on retention crrid, lot id119. these were the same lots used by the customer. customer did not return product or sample for further serological testing. | 0
1526420 | the review of the result files indicated: crrs3 strip lot r064 exp. 12/22/09. 0 strip 1 well 1=1 ( donor heterozygous for kell ). 0 strip 1 well 2=0 ( donor kell negative). 0 strip 1 well 3=0 ( donor kell negative). 4+ strip 1 well 4=100 image appears positive. new sample was drawn from same patient. crrs3 strip lot r064 exp. 12/22/09. 0 strip 2 well 1=1 ( donor heterozygous for kell )image appears negative with a halo in the background. 0 strip 2 well 2=0 ( donor kell negative) image appears negatve. 0 strip 2 well 3=0 ( donor kell negative) image appears negative. 4+ strip 2 well 4=100 image appears positive. confirmed the reactivity of the k antigen on retention crrs (3), lot r064, with anti-k. screen testing was performed with customer's patient sample (reportedly anti-k), using retention crrs (3), lot r064 on in-house echo. sample was nonreactive with all cells tested. hemagglutination tube testing was performed with customer's patient sample using selected k+ and k- cells from retention panocell-20, lot 36670. testing was performed at all temperatures. sample exhibited +w and 2+ reactivity with k+ cells and was nonreactive with k- cell. >< customer reported an unexpected negative result during validating testing on the echo. a known a positive patient with anti-k was tested using capture-r ready screen 3 (crrs 3) on the echo and the anti-k was not detected. | 0
1526421 | review of result files shows wells scores of 0 for cells 1, 3, 5, 7, 8, 11 &14, k and s positive cells of crrid, lot id120 resulting in negative reactions. well images appear negative. confirmed the reactivity of the k and s antigens on retention crrid, lot id120. nature of the sample cannot be ruled out as causing the event. >< customer reported missed antibodies when testing patient sample with known anti-k and anti-s on the echo using capture-r ready id (crrid). no transfusion or adverse reactions occurred as a result of unexpected negative reactions. | 0
1526444 | customer reported a visually positive reaction was interpreted as negative on the echo. the sample has a known anti-e. >< image files were reviewed. results file gives reaction strength of well 2 as 2. results are visually positive. the investigation is on-going. as an interim measure, customers will be notified via a technical communication with a recommendation to visually inspect all negative capture test well results. | 1
1566020 | sample was tested with capture-r ready id (crrid) on the echo. the sample was negative for all jka positive cells. cells 1, 2, 4, 6, 7, 9, 11, 12, 13 were all jka positive and given negative interpretations on the echo. visually, all these reactions appear negative. the positive and negative control wells appear as expected. sample was tested with capture-r ready screen 3 (crrs 3) on the echo. the sample was negative for all 3 cells. cells 1 and 3 were jka positive. the positive control well appears as expected. confirmed the reactivity of the jka antigen on retention crrid, lot id 121. the nature of the sample cannot be ruled out as causing the event. >< customer reported an unexpected negative reaction on the echo, when testing a patient sample with a history of an anti-jka. no transfusion reactions were reported as a result of the negative reaction. | 0
1566021 | the sample was negative for both cell 1 and cell 14 of capture-r ready id (crrid). both cells are heterozygous for e. visually both cells appear negative. the positive and negative control wells appear as expected. sample was tested with capture-r ready screen 3 (crrs 3) on the echo. the sample was 4+ positive for cell 2 (homozygous for e). the postiive control well appears as expected. confirmed the reactivity of the e antigen on retention crrid, lot id121. >< customer reported an unexpected negative reaction on the echo when testing a patient sample with a history of an anti-e. no transfusion reactions were reported as a result of the negative reaction. | 0
1566022 | customer reported an unexpected negative reaction on the echo when testing a patient sample with a history of an anti-jka. no transfusion reactions were reported as a result of the negative reaction. >< the sample tested on the echo and resulted negative with all cells of capture-r ready id (crrid). cells 3, 5, 6, 7, 8, 9, 10, 11, and 14 are all jka positive. visually all these reactions appear negative. the positive and negative control wells appear as expected. the sample was tested with capture-r ready-screen 3 (crrs 3) on the echo. the sample was 1+ positive for cell 1 and 2+ positive for cell 3. both cells are jka positive. the positive control well appears as expected. confirmed the reactivity of the jka antigen on retention crrid extend i, lot dp038, and crrs (3), lot r066, using anti-jka. these were the same lots used by the customer. | 0
1566023 | customer reported an a positive segment which was testing using the fwd abo assay on the galileo and resulted as ab positive. the sample was re-tested on the galileo and this time resulted with the expected a positive result. >< review of images: sample in row 8: a8-4+ reaction/ 99 reaction strength- visually pos. b8- 3+ reaction/ 64 reaction strength- the reaction in the anti-b well does have some irregular borders which can be suggestive of fibrin. c8- 4+ reaction/ 97 reaction strength- visually pos. d8- neg reaction/ 16 reaction strength- visually neg. repeat testing, sample in row 1: a1-4+ reaction/ 90 reaction strength- visually pos. b1-neg reaction/ 14 reaction strength- visually neg. c1-4+ reaction/ 91 reaction strength- visually pos. d1-neg reaction/ 14 reaction strength- visually neg. based on the fact that the sample repeated on the galileo as expected using the same reagents, and the visual appearance of the anti-b well in the first plate, the most likely reason for the unexpected positive reaction is due to fibrin in the segment sample. the galileo operator manual states that clotted samples should not be tested on the galileo. the instrument is operating as expected. | 0
1566028 | review of images: plate contained the sample in well e4 which resulted as negative with a reaction strength of 14. a service call was made: added pulse counts to first position to 500 from 498 for washer alignment and verified clean well roi. performed qc. galileo is operating within specifications. >< customer reported an unexpected negative reaction when testing a patient sample with a known anti-e on the galileo. | 0
1566029 | review of images: neg reactions- all wells less than or equal to 12. a service call was made: performed galileo unexpected reactions checklist to verify instrument operation and verified clean well roi. performed qc. galileo is operating within specifications. >< customer reported an unexpected negative reaction when testing a patient sample with a known anti-e on the galileo. | 0
1566030 | customer reported an unexpected positive reaction when testing a donor sample on the galileo. the sample tested weak d positive on the galileo. re-testing of the sample on a different galileo resulted as negative. >< review of images: sample was pipetted in e1 with a reaction strength of 14 and in f1 with a reaction strength of 64. f1 has moderate red cell adherence with a slight pooling of cells in the center forming a partial arc, appearing as 3+. customer did not return sample for further serological investigation. | 0
1276363 | images were download from the customer's instrument, and the well containing anti-b appeared to have a greenish tint, suggesting it contained anti-a reagent. a service call was made. the customer stated that the wash was not always as strong following addition of anti-a reagent compared to washing following addition of anti-b and anti-d reagents. new wash pumps were installed. the system was tested and operated within specifications. >< customer reported abo discrepancies with the fwd_aborh assay. a known type a donor reported as an ab on the galileo. repeat testing reported as type a.a type of ab negative was reported for a donor unit labeled as a negative. the sample was not retested either on the galileo or by another method. however, a new segment from the same unit was tested manually (i.e. tube method) and the expected a negative result was obtained. | 0
1401073 | customer called to report unexpected negative results on echo, for a patient sample known to contain anti-fyb. >< reactivity of the fyb antigen was confirmed on retention capture-r ready screen (crrs) (3), lot r043 and capture-r ready id (crrid), lot id114, which were used by the customer at the time of the event.tested the returned patient sample with retention crrid, lot id114 on an in-house echo. echo results were interpreted as invalid; the fluid level of the sample was too low. the sample was tested by tube hemagglutination test with selected fy(a-b+) and fy(a+b+) cells on retention panoscreen, lot 14362 using immuadd as the potentiator. the sample exhibited 1+ and +w reactivity at the indirect antiglobulin (iat) phase with fy(a-b+) and fy(a+b+) cells, respectively. sample was qns for further investigation. | 0
1252763 | customer reported abo discrepancy on the galileo when perform fwd abo assay testing with a donor sample. unexpected positive reactions were observed with anti-b series 3. the sample was interpreted as ab positive on galileo. tube testing results with a new sample from the donor unit were a positive. >< fwd_aborh testing was performed with in-house donor samples of known abo/rh types using retention anti-b series 3, lot 203238, on an in-house galileo. no abo discrepancies were observed. fwd_aborh testing was performed with customer's returned samples retention anti-b series 3, lot 203238, on an in-house galileo. both samples were interpreted as a positive. | 0
1407707 | the customer did not return samples for investigation testing. retention product testing is ongoing. >< customer reported unexpected negative reactions on the echo during validation testing. multiple patient specimens that were negative on the echo resulted as 1- 2+ in gel. customer stated the antibodies involved were anti-e, -s, -k, -k and -fya. | 0
1278383 | a female donor was connected to the tubing set prior to the tubing set being fully loaded onto the machine. pre-connecting tests and verifications had not been performed as instructed by the machine's operators manual. the facility's operator realized that the tubing set had not been fully loaded onto the machine, but attempted to continue through possibly two 'set load/unload' cycles. because the machine had been in-process of loading the tubing set when it was stopped, there is a possibility that the machine would not have reached a complete air evacuation state. this could potentially expose the donor to air because the donor was connected to the tubing while in this state. the facility's technician observed air bubbles in the tubing set and the donor complained of not feeling well. the donor was taken to emergency for observation and was released approx. 2 hours later. there was no alleged machine malfunction which was also confirmed through the manufacturer's investigation. >< the manufacturer's investigation included run data analysis of the event, simulated testing, review of health risk and a review of product labeling. the manufacturer's investigation concluded the device performed to specification. | 0
1171501 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< corrected data: type of reportable event was changed from malfunction to serious injury. the customer received an electrical shock and a small blister to one finger. there was no report of further impact or treatment due to this injury. . investigation summary: in 2008, a customer reported getting an electrical shock while performing maintenance on the cell-dyn sms instrument. the customer had a small blister on one (1) of the fingers as a result of the electrical shock. a field service representative visit found that the shielding of the power cable of one of the pumps was damaged, exposing the core of the wire, which caused a short circuit with the chassis of the instrument. three months later, the subject matter expert (sme) arrived at the customer site and inspected the cell-dyn sms instrument with suggested measurements to ground from leica biosystems, the original equipment manufacturer (oem). the laboratory equipment was maintained by qualified electrical engineers, as indicated by certification stickers on each piece of equipment. conductivity measurements with a calibrated digital multimeter (resistance to ground) were similar to cell-dyn sms units at the abbott manufacturing site. the customer unit was unplugged. repairs to the instrument throughout the inspection were made, as needed. in addition, the laboratory maintenance department repaired the customer's plumbing lines. the fuses were inspected and replaced according to the installation checklist for 220v use. it was found that the pressure pump fuse installed was 1.6-amp; however, the silkscreen on the main fuse pcb indicates a 1.0-amp fuse is required. the pressure pump fuse was replaced to meet specification and a review of complaint history was performed in attempts to determine the last fuse replacement. review of complaint history on this instrument showed that in 2007, the pressure was not working. on the following month, a complaint from the customer indicated that the instrument was requiring fuse replacement(s); fuses may have been replaced at this time. in late 2008, during the sme inspection, on first power on, the first error message on the screen was a slide transport error. parts were ordered to replace the motor and wire assembly, as a wire and a burned motor was found on the slide transport assembly. the load sides of the sweep arms were stuck in the autoloader. the next day, the motor, and wire assembly to the slide transport assembly were replaced, which resolved the unit's error message. an autoloader error was observed, but it was due to an extended period on non-use. the sme resolved the intermittent error. the sme observed a worn slide holder, which was replaced. the unit was checked to current manufacturing quality control procedures. a review of manufacturing records, the original equipment manufacturer of the cell-dyn sms found that the instrument was manufactured with 10 clips latching in 1998. the latch at the location beside the vacuum pump and pressure pump at the top rear of the instrument was not included in the manufacturing of this instrument. there was no requirement to have a latch in this location at the time of the manufacture since this location was implemented in 2000. the purpose of the latch is to hold the cables in place. the risk management file (rmf) for the cell-dyn sms (risk analysis version 3.0) under the risk analysis item was reviewed. the hazard source or cause indicates that electrical shock may be due to exposed ac, dc circuitry within instrument, and electrical current leakage or bad grounding. there are four controls (over current protection, warnings in manual, labels shock, and covers on instrument). the residual risk level is low and acceptable. the cell-dyn sms is manufactured according to approved standards. the cell-dyn sms system operator's manual, hazard electrical safety, page 7-5 states: warning: electrical hazard. indicate the possibility of electrical shock if procedural or engineering controls are not observed. an overview of electrical safety procedures is included on this section. section 8, maintenance - replacing the printer ribbon cartridge, page 8-30, describes the replacement procedure. a review of reports for the months of january 2008 through december 2008 did not indicate an adverse trend for list number, for the reported complaint issue. the event is addressed in product labeling. cell-dyn sms system operators manual: maintenance - replacing the printer ribbon cartridge, page 8-30. based on the investigation, the wear of the exposed pressure pump wire was likely caused by interaction with the rear panel door. the routing of the cable after pump repair caused the insulation to be worn off. based on the investigation, a product deficiency was not identified. this is the final report. >< the customer states that he received an electrical shock while performing maintenance on the cell-dyn sms. he also received a blister to one of his fingers about 2mm in diameter. no medical intervention or medications were administered for this event. a service engineer visited the customer site and found that the shielding of the power cable of one of the pumps was damaged exposing the core of the wire, causing a short-circuit with the chassis of the analyzer. the wire was repaired by the engineer. | 0
1418562 | the customer reported that the probe on the ortho provue analyzer was bent. the customer did not observe any evidence of probe leak. however, it is unknown whether contamination occurred. no erroneous results were reported. probe issues may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion of incompatible blood. >< an ocd field engineer visited the customer site and verified that the probe was bent . the fe replaced the probe, cracked wash station and o rings and performed the necessary adjustments to return the instrument to expected operation. qc passed. this customer has not logged any complaints against this analyzer since this incident. incident is isolated. | 0
1418565 | an ocd field engineer arrived at the customer site and determined that the vacuum system had vacuum restriction in the tubing connectors on the waste bottle, the wash station had a crack on the surface and the probe contained a blockage at the tip. the fe replaced the bottle and tubing connectors, the wash station and probe. replacement of the appropriate components has returned the instrument to its expected operation. this customer has not logged any complaints against this analyzer since this incident. >< the customer reported that the probe of the ortho porvue analyzer was dripping. information was not provided regarding result codes. no erroneous results were reported. probe drip may lead to dilution of sample/reagent, carry over and/or cross contamination and erroneous results which could lead to transfusion of incompatible blood. | 0
1026913 | the customer reported the ortho provue analyzer posted error messages during reagent identification indicating an inability to identify the liquid levels during testing. the customer discovered the vials to be wet. during the second lld check, the customer observed the probe drip into the reagent red cell vials. the customer aborted testing. the customer indicated that the reagents red cells on board the instrument were contaminated. probe drip may lead to dilution of sample/reagent, carry over and/or cross contamination and erroneous results which could lead to transfusion of incompatible blood. the user detected the issue, preventing erroneous results from being reported. >< a possible root cause was determined. an ocd field engineer arrived at the customer site and determined the pressure in the fluidics system was out of specification. replacement of the pressure regulator has returned the instrument to expected operation. this customer has logged a similar complaint against this analyzer since the incident, reference mdr 1056600-2008-00083. | 0
1060810 | the customer reported that the probe dripped fluid and the dilution cup overflowed on the ortho provue analyzer. no erroneous results were reported. probe drip may lead to dilution of sample/reagent, carry over and/or cross contamination and erroneous results which could lead to transfusion of incompatible blood. >< a possible root cause was determined. an ocd field engineer visited the customer site and determined that the pressure was out of specification. the fe replaced the pressure regulator and adjusted the pressure to return the instrument to expected operation. this customer has not logged any complaints against this analyzer since this incident. | 0
1316536 | an ocd field engineer visited the customer site and found the probe was bent. the fe replaced the probe and performed the necessary adjustments to return the instrument to expected operation. qc passed. this customer has not logged any complaints against this analyzer since this incident. incident is isolated. >< the customer reported that the ortho provue analyzer dripped fluid into the reagents on board the instrument. no erroneous results were reported. probe drip may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion to incompatible blood. | 0
1316537 | an ocd field engineer visited the customer site and confirmed the probe was bent. the fe replaced the probe, wash station and additionally the door electronic lock switch. the fe performed the necessary adjustments to return the instrument to expected operation. qc passed. this customer has not logged any complaints against this analyzer since this incident. incident is isolated. (b) (4). >< the customer reported that the probe of ortho provue analyzer was bent and dripped fluid into the reagents and dilution cup. no erroneous results were reported. probe drip may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion of incompatible blood. | 0
1316538 | the customer reported they tested mts gel cards exhibiting bubbles on the ortho provue analyzer. visual inspection of the gel cards showed the reactions were clearly negative with no bubbles after processing. the analyzer did not post any error messages. no erroneous results were reported. >< an ocd field engineer visited the customer site and performed handler and optics adjustments and reader camera alignments. qc passed. repairs have returned this instrument to expected operation. the gel cord product labeling requires the user to perform inspection of the gel card prior to use. any gel card that shows signs of drying, discoloration, bubbles, crystals, or other artifacts should not be used in testing. user error could not be ruled out as a contributing factor. this customer has not logged any similar complaints against this instrument since this incident. incident is isolated. (b) (4). | 0
1345127 | an ocd field engineer visited the site, inspected the provue for proper mechanical/optical operation, reviewed the log files and indicated that the analyzer was functioning as expected. the fe indicated that the antibody missed by the provue was anti-m, which tends to be temperature sensitive. sample variability may have contributed to this incident. therefore, service was not required. the customer was informed regarding the nature of the antibody. this customer has not logged any complaints against this analyzer since this incident. incident was isolated. >< the customer reported that the ortho provue analyzer resulted a positive antibody sample as negative during antibody screen testing. the sample reacted positive in manual test using the same lots of reagents. no erroneous results were reported. false negative test results can lead to transfusion of incompatible blood. | 0
1365182 | the customer reported that the ortho provue analyzer dripped fluid resulting in reagent and sample contamination. no erroneous results were reported. probe drip may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion of incompatible blood. >< an ocd field engineer visited the customer site and found the probe was bent. the fe replaced the probe, performed adjustments and preventative maintenance and run wad diagnostics to return the instrument to expected operation qc passed. this customer has not logged any similar complaints against this analyzer since this incident. incident is isolated. | 0
1436830 | the customer reported that the provue resulted an antibody screen as negative. repeat testing with the same sample seven days later and a newly drawn specimen yielded a '?' by the provue. upon visual inspection, it was noted that the reaction was a weak positive - 1+. false negative test results can lead to transfusion of incompatible blood. >< the customer did not send samples for further investigation. a field engineer replaced the camera and adjustments the instrument. this customer has not logged any similar complaints against this analyzer since this incident. | 1
1436832 | an ocd field engineer (fe) found that the read camera visor and the two rings that hold the visor in place were missing. the fe believes the gripper jammed a card onto the visor and when the customer attempted to remove the card, they knocked the visor loose and it arced off the gripper. the fe replaced the fuse on the carousel board and the visor to return the instrument to expected operation. the fe ran a diagnostics test for card reading and transport. the motor movement and initialization results were acceptable. qc passed. this customer has not logged any similar complaints against this analyzer since this incident. incident is isolated. (b) (4). >< the customer reported observing a spark and experiencing a shock while attempting to dislodge a card from the waste tray area on the ortho provue analyzer. the customer was not harmed and did not seek medical intervention. | 0
1046542 | this customer reported that the ortho provue analyzer interpreted a positive antibody screen result as negative. the issue occurred in 2008. visual inspection of the gel card showed the reaction was positive. anti-d was identified in manual gel test. no erroneous results were reported. false negative test results can lead to transfusion of incompatible blood. >< no definitive root cause was determined. a field engineer arrived at the site and performed adjustments to the reader camera and created a new reference image to return the instrument to expected operation. this customer has not logged any similar complaints against this analyzer since this incident. | 1
1060818 | a possible root cause was determined. an ocd field engineer the wash station and performed the necessary adjustments and returned the instrument to expected operation. this customer has not logged any complaints against this analyzer since the repair. >< customer reported the provue probe was dripping. probe drip may lead to dilution of sample/ reagent, carry over and/or cross contamination and erroneous results which could lead to transfusion of incompatible blood. | 0
1060819 | no definitive root cause was determined. ocd second level support indicated that the first sample was anti-m (cold agglutinin) which will not react consistently on th analyzer since pipetting is performed at 37 oc. ocd second level support reviewed the log files and indicated that the reactions in the microtube looked negative as interpreted by the provue. the field engineer indicated that repairs were not required since the analyzer was performing as expected. this customer has not logged any complaints against this analyzer since this incident. >< the customer reported two instances in which the ortho provue analyzer resulted weakly positive antibody screen samples as negative. one of the samples contained anti-m. both samples reacted weakly positive in manual gel test. no erroneous results were reported. false negative test results can lead to transfusion of incompatible blood. | 0
1105836 | a possible root cause was determined. an ocd field engineer arrived on site and indicated that the customer had reported observing fluid coming from the handler. the fe determined the probe was crashed within the handler. replacement of the probe and the appropriated adjustments to the drive motor belt and lubrication to the screw and bearings has returned to instrument to expected operation. this customer has not logged any similar complaints against this analyzer since this incident. >< the customer reported observing moisture residue in the ortho provue analyzer and on top and underneath the waste container. the customer also noticed dried wash solution on the analyzer. no erroneous result or contamination occurred as a result of this incident. fluid leak or probe issues can lead to dilution of reagent/sample, carry over/cross contamination from microtube to microtube, erroneous test results and transfusion of incompatible blood. | 0
1277330 | an ocd field engineer visited the customer site and reported that the instrument was operational upon arrival. the customer found that the connector on the waste bottle as loose. the customer replaced the affected component and performed qc without any issues. this customer has not logged any complaints against this analyzer since this incident. incident is isolated. >< the customer reported that the ortho provue analyzer dripped fluid into the reagent cells testing. no erroneous results were reported. probe drip may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion of incompatible blood. | 0
1427088 | the customer reported that the ortho provue analyzer interpreted negative reactions as positive during antibody screen testing. visual inspection of the gel cards were negative. no erroneous results were reported. a false positive result may lead to the misclassification of a patient's abo group, with the potential for transfusion of abo incompatible blood. >< an ocd field engineer visited the site and found the reader camera images were intermittently displayed with an interference. the fe replaced and adjusted the reader camera, created a new reference image and performed maintenance to return the instrument to expected operation. the customer tested qc and samples with acceptable results. this customer has not logged any complaints against this analyzer since the incident. incident is isolated. | 1
988691 | a possible root cause was determined. an ocd field engineer went to the customer site and found a clot in the drain hole. the fe cleaned the dilution cup and performed final washes to return the instrument to expected operation. user error could not be ruled out as a contributing factor. this customer has not logged any similar complaints against this analyzer since the incident. >< the customer reported that the probe on the ortho provue analyzer dripped and the dilution cup overflowed. contamination of reagents occurred as a result of this incident. info was not provided regarding possible result codes intended to prevent erroneous results from being reported. probe drip may lead to dilution of sample/reagent, carry over and/or cross contamination and erroneous results which could lead to transfusion of incompatible blood. the customer detected the issue, preventing the possibility of erroneous results from being reported. | 0
1418566 | an ocd field engineer arrived at the customer site and determined that the pressure regulator need to be replaced and the pressure was too high. the fe had also found some matter in the syringe and a slight kink in the vacuum line near the pump housing. the fe installed a new pressure regulator and made adjustments as well as cleaned the syringe and repaired the kink in the vacuum line. the instrument was returned to expected operation. this customer has not logged any complaints against this analyzer since this incident. (b) (4). >< the customer reported that the probe of the ortho provue analyzer was dripping. the instrument was taken out of operation. probe drip may lead to dilution of sample / reagent, carry over and / or cross contamination and erroneous results which could lead to transfusion of incompatible blood. | 0
1060817 | the customer reported that in 2008, the ortho provue analyzer resulted a pt's sample containing a warm-auto as positive (1+). the next day, a new sample was collected and tested on the analyzer using the same lot of reagents. the sample reacted negative. repeat testing in manual get test using the same lot of reagents showed weak positive reactivity (1-2+) with the sample. the customer repeated the sample on the provue and obtained negative results. visual inspection of the processed gel card confirmed the reactions were negative. no erroneous test results were reported. false negative test results can lead to transfusion of incompatible blood. >< no definitive cause was determined. an ocd field engineer arrived on site, reviewed the provue log images and indicated that the images correlated with the negative results. the fe determined that the incubator temperature required re-calibration. re-calibration of the incubator temperature has returned the instrument to expected operation. this customer has not logged any similar complaints against this analyzer since this incident. complaint. | 0
1476815 | images from the customer's instrument were reviewed and consistent with the results reported.review of labeling: the package insert for the anti-a and anti-b blood grouping reagents states: 'certain subgroups of a and b may produce reactions that are weaker than those obtained with a or b red blood cells of most random donors. depending on the subroup involved, some may appear nonreactivein direct agglutination tube, microtitration plate or slide tests'. instrument is operating as expected. >< customer reported that a donor sample that was tested on reflexabo assay on the galileo did not type as expected. the sample has a history of asubb but the galileo typed it as a b positive. | 0
1338663 | a customer contacted beckman coulter inc (bci) regarding erroneously low platelet (plt) results generated by the coulter gens system. the platelet result of the patient sample was 89x10^3 cells/ml. the gens instrument did not flag the sample for plt clumps and low plt (as compared to manual estimated) was reported due to the clumping. the erroneous result was reported outside the lab and later corrected with the manual estimate of 144. based on the available information, there was no affect to patient treatment. >< samples were collected in 5cc vacutainers and were sampled < 2 hours after draw. controls are run each shift and controls were run before and after the incident and recovered within qc range. in 2008, field service engineer (fse) discussed flagging criteria with the customer. fse determined instrument was working fine. a clear root cause for this event has not been determined to date. a malfunction will be assumed for the purpose of this report. | 1
956204 | (b)(4). all manufacturers/date received by manufacturer: correct from (b)(4) 2007 to (b)(4) 2007; typographical error. >< the customer states that over the past few weeks discrepant patient results have been generated by the architect c8000 calcium assay. the customer gave an example of an initial result of 3.8 mg/dl that retested at 7.4 mg/dl. no further patient results or information is available. the customer is requesting a service call. there is no impact to patient management reported. >< the field service representative (fsr) was dispatched to inspect the instrument. the text states the cuvette washer was cleaned due to the acid wash nozzle three being plugged. the field service representative flushed the water lines and performed a ten replicate precision run. the text states the customer ran multiple levels of control for several assays, which were within acceptable limits. subsequent instrument operations and test results were acceptable. the customer's issue is addressed in the clinical chemistry calcium reagent package insert in the sections addressing specimen collection and handling and reagent handling and storage. also, the abbott architect system operations manual ((pn 201837-104) june, 2007) addresses the issue under section 7, operational precautions and limitations, requirements for handling consumables and limitations of result interpretation; section 9, service and maintenance; and, section 10 troubleshooting and diagnostics - observed problems, erratic results, poor precision - photometric results (c system). this is a final report. | 0
1067406 | the customer states that one patient sample generated an initial architect c8000 potassium assay result of 8.4 mmol/l. per the customer's re-run rule, the sample was retested and generated a result of 4.2 mmol/l. the sample was also tested on another analyzer and generated a potassium result of 4.1 mmol/l. no suspect results were reported from the lab. there is no impact to patient management reported. >< concomitant medical products and therapy dates: aero/c8k ict module. device returned to manufacturer: the integrated chip technology (ict) module was returned for the investigation, not the entire architect c8000 analyzer. the investigation began with testing performed with the returned suspect integrated chip technology (ict) module returned from the customer site, which generated passing calibrations with controls within specification ranges for the sodium (na), potassium (k) and chloride (cl) assays with no error codes generated. precision runs for na, k and cl were within acceptable limits. a review of the customer's results logs determined the pre- and post- iref reading variation for the suspect k result did not exceed the defined voltage drift of 10mv, so an error was not generated. however, the variation was significantly higher than the pre- and post- iref readings of the acceptable k repeat results. currently, tracking activities are in place to monitor the drift voltage limit for the sodium, potassium and chloride assays for future enhancements and improvements. a review of the customer?s sample aspiration and dispense pressure monitoring (pm) logs did not find any abnormal or suspicious pm readings. a review of complaint tracking systems did not verify an ict module product issue, and a review of the most current corrective action/preventive action metrics did not identify an adverse trend in conjunction with the complaint issue currently under investigation. the architect system operations manual (june 2007) provides the following relevant information: section 7 provides guidance and requirements for handling specimens. specimens should be checked for proper clot formation prior to centrifugation and checked for bubbles, fibrin, and red cells prior to testing. the limitations of result interpretation section states assay results must be used with other clinical data, for example, symptoms, other test results, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. all data must be considered for patient care management. if assay results are inconsistent with clinical evidence, additional testing is suggested to confirm the result. the architect system has been validated for its intended use. however, errors can occur due to potential operator errors and architect system technology limitations. section 10 includes a list of probable causes and corrective actions applicable to the customer's issue under the observed problems 'elevated and depressed concentration - ict results', and 'erratic results, poor precision - ict results'. multiple causes are listed, including but not limited to sample integrity issues, sample or reagent probe partially obstructed, scheduled maintenance is due, and ict module expired or not performing. the investigation was unable to verify a specific cause of the elevated potassium result generated by the suspect sample. a sample integrity issue, such as latent fibrin formation, affecting the initial ict aspiration of the suspect sample is a probable cause, as the repeat testing on the same sample was successful, and no other patient or control results were affected, and testing performed on the returned module generated acceptable results. based on the available information and investigation performed, a deficiency of the ict module and/or architect system was not identified. this is a final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1215498 | lipemic patient was tested for multiple chemistry analytes. abnormal results were reported and the physician questioned the results. the sample was airfuged and rerun. the second set of results were very different. customer sent out a corrected report. no patient intervention resulted from the discrepant results. there have been no reports of adverse health consequences due to the discordant results. >< field service was not sent to the site. discrepant results were only obtained on one sample that was extremely lipemic. triglyceride interference claims in the information for use (ifu) indicated that this sample would effect other analyte results. other patient samples ran ok and analyte qc's met specifications. | 0
1257303 | the account noticed drops on the architect c8000 analyzer cuvette wash comb which caused creatinine reproducibility problems in 2 patients. one patient's serum specimen tested architect creatinine = 15 mg/l but repeated at 7 mg/l. no data was provided for the second patient. the abbott technical representative stated the architect c8000 analyzer wash comb was bent but the customer straightened it out. also, there were many sampling errors, so the abbott technical representative calibrated the assay for resolution. no results were reported outside of the laboratory. no impact to patient management was reported. >< investigation in process, a conclusion cannot be reached at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4) - erratic creatinine result, see investigation results. the operator removed the cuvette wash comb, tightened the needles, and washed the cuvettes with bleach. the operator straightened the first needle in the cuvette wash comb that was bent. the operator also recalibrated the sample probe because sampling errors were generated at the time of the incident. the instrument was inspected and troubleshooting was performed. since then, no additional issues regarding erratic results have been observed. to investigate this issue, the investigation team reviewed the complaint text, the instrument history, and architect system labeling. the review showed the architect system operations manual does address erratic results including probable causes and corrective actions. additionally the review of the instrument history showed no additional complaints have been received involving erratic results. architect system operations manual (list number 201837-104) provided adequate information about the sample probe, troubleshooting for erratic results, operational precautions, and limitations. in the clinical chemistry creatinine reagent package insert (b)(4), literature is provided in describing suitable specimens, results, and limitations of the procedure. based on the available information, the cause of the customer's issue was not identified, nor was any deficiency identified for this complaint. this is a final report. | 0
1286265 | (b)(4). the investigation was conducted by an abbott field service representative (fsr) whose inspection of the aeroset analyzer found the sample probes, reagent probes, and a mixer were out of alignment. the fsr calibrated and adjusted the components, followed by running samples and controls to verify that the issue had been resolved. the issue did not recur. subsequent instrument operations and test results were acceptable. complaint and trending data did not indicate a product issue or adverse trend. the customer's issue is addressed in the aeroset operations manual under section 9 service and maintenance, under weekly maintenance procedures and under the analyzer component replacement section; section 10, troubleshooting and diagnostics, describes the maintenance utilities screen, which provides functionality to manually move robotics for maintenance and troubleshooting purposes and also provides probable causes and corrective actions for the observed problem results are erratic, precision is poor. the aeroset failed to perform as intended when it generated discrepant results because the sample and reagent probes and a mixer were out of alignment. however, system performance was restored by the customer and abbott field service in a manner consistent with information in section 10, troubleshooting and diagnostics, of the aeroset operations manual. therefore, a product deficiency is not indicated. this is a final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer reports very low anion gap results for patient samples with corresponding elevated co2 assay results generated by the aeroset analyzer. the customer states that samples retested on the architect c8000 analyzer generated lower co2 values. the customer gave one example of an aeroset co2 result of 35 mmol/l on the same sample that generated an architect c8000 co2 assay result of 21 mmol/l. the customer states that no suspect results have been reported from the lab. controls have been recovering slightly above their current mean values. the customer states that quarterly maintenance is due and that she will begin maintenance on the syringes. a service call was also initiated. there is no impact to patient management reported. | 0
1294763 | the customer states that erratic patient electrolyte results are being generated by the aeroset analyzer. the customer gave one example of an initial sodium assay result of 111 mmol/l that retested at 147, 146, 146, and 140 mmol/l. all suspect results have initially generated low results and then retest at higher values. the customer does not have any patient information available. a service call was initiated. there has been no reports of any impact to patient management due to this issue. >< the customer initially replaced syringes 12 and 11 as part of troubleshooting, but this did not resolve the issue. the customer noted that the tubing above the integrated chip technology (ict) module probe was crimped and the ict cup was sticky and dirty. an abbott field service representative (fsr) instructed the customer on replacing the crimped tubing, but the issue was unresolved. the high concentration waste poppet valves were then replaced. subsequent instrument operations and test results were acceptable. no other instrument issues were identified through review of the service ticket history. the aeroset system operations manual provides troubleshooting assistance and contains adequate information to address the customer's concerns. the investigation demonstrated that the aeroset system is performing within its intended use, label claims and specifications. no deficiency related to the performance of the device was identified. this is a final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1330407 | architect multigent lithium assay lot#: ni. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that one patient sample generated an initial architect lithium assay result of 3.2 mmol/l. the sample was retested with results of 1.45 and 1.5 mmol/l. there is no impact to patient management reported. >< (b)(4). architect multigent lithium assay ln: 8l25-01 lot#: ni. the investigation began with a review of the complaint text, which indicates the customer observed a discrepant lithium result of 3.2 mmol/l with two repeat results of 1.45 and 1.5 mmol/l. at the recommendation of abbott customer support, the operator replaced mixer 1. the mixer was not available for return. as documented in the customer's maintenance log, daily and weekly maintenance was performed on the date of the incident. the quality control, script, message history, pressure monitoring, liquid level sense, and result instrument logs did not contain meaningful data related to this issue. the verification diagnostic procedure, 3126 mixer vibration test, was not documented in the maintenance, script, or message history log; therefore, it could not be confirmed that the customer performed the troubleshooting procedure of replacing mixer 1. based on the available information, the instrument logs could not confirm the customer's alleged issue and no product deficiency was found. a review of the complaint history for architect csystem sn# (b)(4) over the period of (b)(6) 2008 through (b)(6) 2009 was performed. the review did find two other complaints regarding erratic results that were resolved by additional customer training. the architect system operations manual ((b)(4) june, 2008) and the architect multigent lithium reagent package insert (b)(4) contain adequate information to address the customer's issue. a malfunction of the system was not identified and after the component replacement of mixer 1, the instrument has been performing as intended. this is a final report. >< (b)(4). this is a final report. | 0
1356257 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the architect c8000 analyzer has generated discrepant potassium results on several patient samples. on patient #2, the initial result was 7.18 mmol/l and the retest result was 4.62 mmol/l. there was no impact to patient management reported. >< (b)(4). evaluation: erratic potassium results. erratic potassium results and o-ring in r1 syringe. review of the complaint text indicates the customer observed discrepant potassium results generated by the architect c8000 on (b)(6) 2009. the integrated chip technology (ict) module was last replaced on (b)(6) 2009 and the sample probe was replaced during instrument maintenance in (b)(6) 2008. the customer was not performing the precision study as instructed by the sales representative, because the customer was running amylase along with potassium. the customer then performed the study with only potassium. there was no carryover of the amylase reagent from the reagent pipettor. during the syringe maintenance of the r1 syringe, the o-ring sealing was damaged and there was evidence of air bubbles or foam in the syringe. once the o-ring sealing was replaced, the syringe was performing as intended. no additional erratic potassium results were reported after the component replacement of the o-ring sealing. the architect system operations manual ((b)(4) may, 2008): section 7, operational precautions and limitations, limitations of result interpretation states; assay results must be used with other clinical data, for example, symptoms, other test results, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. all data must be considered for patient care management. if assay results are inconsistent with clinical evidence, additional testing is suggested to confirm the result. the architect system has been validated for its intended use. however, errors can occur due to potential operator errors and architect system technology limitations. section 9: service and maintenance: component replacement provides thorough instructions on the procedure to remove, replace, and verify probes. section 10, troubleshooting and diagnostics, observed problems, erratic results (c system) lists the probable causes and corrective actions for erroneous results. in the clinical chemistry integrated chip technology sodium potassium chloride (ict na+, k+ cl-) reagent package insert (b)(4), literature is provided in describing suitable specimens, results, and limitations of the procedure. fibrin clots may subsequently form in these sera and the clots could cause erroneous test results. multiple myeloma samples are known to give low results on diluted ise systems due to the high level of protein present in the sample. the probable cause of the erratic potassium results was the leaking r1 syringe caused by the damaged o-ring sealing. this is the final report. | 0
1356258 | the account stated that the architect c8000 analyzer has generated discrepant potassium results on several patient samples. on patient #1, the initial result was >6.0 mmol/l and the retest result was 3.8mmol/l. there was no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation: erratic potassium results erratic potassium results and o-ring in r1 syringe. review of the complaint text indicates the customer observed discrepant potassium results generated by the architect c8000 on (b)(6) 2009. the integrated chip technology (ict) module was last replaced on (b)(6) 2009 and the sample probe was replaced during instrument maintenance in (b)(6) 2008. the customer was not performing the precision study as instructed by the sales representative, because the customer was running amylase along with potassium. the customer then performed the study with only potassium. there was no carryover of the amylase reagent from the reagent pipettor. during the syringe maintenance of the r1 syringe, the o-ring sealing was damaged and there was evidence of air bubbles or foam in the syringe. once the o-ring sealing was replaced, the syringe was performing as intended. no additional erratic potassium results were reported after the component replacement of the o-ring sealing. the architect system operations manual ((b)(4) may, 2008): section 7, operational precautions and limitations, limitations of result interpretation states; assay results must be used with other clinical data, for example, symptoms, other test results, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. all data must be considered for patient care management. if assay results are inconsistent with clinical evidence, additional testing is suggested to confirm the result. the architect system has been validated for its intended use. however, errors can occur due to potential operator errors and architect system technology limitations. section 9: service and maintenance: component replacement provides thorough instructions on the procedure to remove, replace, and verify probes. section 10, troubleshooting and diagnostics, observed problems, erratic results (c system) lists the probable causes and corrective actions for erroneous results. in the clinical chemistry integrated chip technology sodium potassium chloride (ict na+, k+ cl-) reagent package insert (b)(4), literature is provided in describing suitable specimens, results, and limitations of the procedure. fibrin clots may subsequently form in these sera and the clots could cause erroneous test results. multiple myeloma samples are known to give low results on diluted ise systems due to the high level of protein present in the sample. the probable cause of the erratic potassium results was the leaking r1 syringe caused by the damaged o-ring sealing. this is the final report. | 0
1356260 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation: erratic potassium results. erratic potassium results and o-ring in r1 syringe. review of the complaint text indicates the customer observed discrepant potassium results generated by the architect c8000 on (b)(6) 2009. the integrated chip technology (ict) module was last replaced on (b)(6) 2009 and the sample probe was replaced during instrument maintenance in (b)(6) 2008. the customer was not performing the precision study as instructed by the sales representative, because the customer was running amylase along with potassium. the customer then performed the study with only potassium. there was no carryover of the amylase reagent from the reagent pipettor. during the syringe maintenance of the r1 syringe, the o-ring sealing was damaged and there was evidence of air bubbles or foam in the syringe. once the o-ring sealing was replaced, the syringe was performing as intended. no additional erratic potassium results were reported after the component replacement of the o-ring sealing. the architect system operations manual (b)(4): section 7, operational precautions and limitations, limitations of result interpretation states; assay results must be used with other clinical data, for example, symptoms, other test results, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. all data must be considered for patient care management. if assay results are inconsistent with clinical evidence, additional testing is suggested to confirm the result. the architect system has been validated for its intended use. however, errors can occur due to potential operator errors and architect system technology limitations. section 9: service and maintenance: component replacement provides thorough instructions on the procedure to remove, replace, and verify probes. section 10, troubleshooting and diagnostics, observed problems, erratic results (c system) lists the probable causes and corrective actions for erroneous results. in the clinical chemistry integrated chip technology sodium potassium chloride (ict na+, k+ cl-) reagent package insert (b)(4), literature is provided in describing suitable specimens, results, and limitations of the procedure. fibrin clots may subsequently form in these sera and the clots could cause erroneous test results. multiple myeloma samples are known to give low results on diluted ise systems due to the high level of protein present in the sample. the probable cause of the erratic potassium results was the leaking r1 syringe caused by the damaged o-ring sealing. this is the final report. >< the account stated that the architect c8000 analyzer has generated discrepant potassium results on several patient samples. on patient #3, the initial result was 6.27 mmol/l and the retest result was 3.74 mmol/l. there was no impact to patient management reported. | 0
1379387 | (b)(4). the customer stated that patient samples have higher than expected potassium assay results generated on an architect c8000 analyzer. recollected samples yielded results that were within the customers established normal range (3.5 - 5.3 mmol/l). the customer states that the controls for the potassium assay are acceptable and that a precision study for the potassium assay was also acceptable. the customer states that the integrated chip technology (ict) module ((b)(4)) was installed on (b)(6) 2009. the architect system operations manual states that the ict module warranty is 15,000 samples or two months post-installation, whichever occurs first. the customer indicates that the centrifuge in which the samples were spun has been replaced and that the issue appears to now be resolved. the architect system operations manual (pn 201837-105) (b)(6), 2008 and the clinical chemistry ict (na , k , cl) sample diluent package insert ((b)(4)) both contain the information necessary to address the customer's concern. a review of the complaint history files found no adverse trends related to the issue, and a review of the instrument's service history did not detect any repeats of the issue. based upon the available data and the results of this evaluation, the cause of the customer's issue was not identified, nor was any deficiency or malfunction identified. probable causes include the expiration of the ict module or that the ict module has exceeded its time or sample warranty (> 2 months after installation or >15,000 samples); also, samples were not collected and/or prepared correctly and samples contain fibrin clots or particulate matter. this is a final report. >< the customer states that one female patient generated an initial architect c8000 potassium assay result of 6.2 mmol/l that retested with the same value. the patient was redrawn the following day and this second sample generated a potassium result of 4.5 mmol/l. the customer's normal reference range is 3.5 to 5.3 mmol/l. the customer noted that six to eight patients may have been affected by this issue. no further patient information was provided by the customer. there is no impact to patient management reported. >< arch conc. ict diluent lot#:70005hw00. an investigation is in process. a final report will be submitted when the investigation is complete. an investigation is in process. | 0
1396715 | (b)(4). discrepant results were probably caused by a sample integrity issue that was not identified by the user and/or did not exceed the threshold required to generate a system pressure monitoring (pm) error. after further evaluation, the suspect medical device was changed from clinical chemistry creatinine reagent, list # 7d64-20, manufacturing site abbott manufacturing, inc. south pasadena, ca, to the architect analyzer, list # 1g06-01, manufacturing site abbott labs, registration # (b)(4). qc for all assays was performing as expected. system logs were reviewed and there was no indication of an instrument issue. the liquid level sense logs had been overwritten and were not available for review. the customer's complaint history does not include a recurrence of inconsistent or erratic results. the architect system operations manual and ict sample diluent package insert provide sufficient information regarding scheduled maintenance and component replacement for the c8000 system, sample handling, interpreting results, and troubleshooting the customer's issue. the c8000 sample pippetor assembly includes fluid sense/pressure monitoring that helps identify errors in aspiration. however, the labeling does not guarantee that all sample integrity issues will be detected and errors can occur due to potential operator errors and architect system technology limitations. based on the available information a specific cause for the inconsistent results generated by the suspect sample could not be determined. probable causes include a sample integrity related issue that was not detected by the operator and did not exceed the threshold required to generate an aspiration error. a deficiency was not identified. >< the customer stated an outpatient generated a high potassium result of 6.9 mmol/l on the architect c8000 analyzer. the result was flagged high and the sample was retested yielding a value of 6.9 mmol/l and was reported. based on the high potassium result, the physician advised the patient to go to the emergency room. another sample was obtained in the er and the potassium result was approximately 3.0 mmol/l. there was no impact to patient management reported. four days later, the customer retested the original sample yielding a potassium result of 5.4 mmol/l. other discrepant values were also noted such as an initial creatinine result of 9.81 that retested at 5.6 mg/dl. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1420461 | the account stated that the architect c8000 analyzer was leaking and caused the extension block from the system control center's (scc) that was installed on the floor, to get wet and catch on fire. there were no injuries reported. the hospital electrician stated that the extension block should be installed differently and noted that another analyzer's extension block in the lab is installed the same way. the account has requested extension leads to reach the central processing unit (cpu) to keep the extension block off the floor. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated error code 5201 (reaction vessels have not been picked up for an extended period of time) was generated on the architect i2000sr analyzer. after resolving the issue, patient samples tested before and after the error code occurred were retested and higher results were obtained. a tsh result of <0.0025 uiu/ml was generated and reported. upon retest, a value of 1.5436 uiu/ml was obtained and a corrected report issued. no impact to patient management was reported. >< leaking fluid from architect c8000 came in contact with a system control center (scc) extension block that was located on the floor, and the wet scc extension block caught fire. the technical support specialist (tss) further explained that the leak occurred when a site electrician picked up the external waste pump after loosening the screws, and the external waste pump was leaned to one side. no one was injured due to the leaked fluid or the extension block fire. the tss replaced the extension block to resolve the issue. the tss also rewired the extension leads from the ups to the c8000 scc stand to ensure there were no connectors, and to move the extension block off the floor.the architect system operations manual (96211-110) was found to contain adequate information on potentially dangerous conditions on the architect systems, and on electrical hazards. the manual notes the architect system does not pose uncommon electrical hazards to operators, electrical cabling should be inspected for signs of wear and damage, liquids should be kept away from all connectors of electrical or communication components, and spilled fluids should be cleaned immediately.the architect c8000 system service and support manual were found to contain adequate information on the troubleshooting, and parts removal and replacement for the high concentration waste pump. the manual also contains potential biohazard warnings for procedures performed on the high concentrated waste pump area, and adequate hazard information about the electrical and biological hazards of the c8000.a 6 month report search for similar complaints found no additional complaints were initiated due to burned architect c8000 extension blocks.a report search of the architect c8000 service history found no additional customer complaints of burning or smoking issues for architect c8000 have been initiated since the tss replaced the extension blocks, and rewired the extension leads from the ups to move the extension block off the floor. the abbott metric reports found no adverse trends were identified for the architect c8000 analyzer or due to burning extension block issues.the safety hazards memo was found to contain information noting abbott diagnostic equipment and accessories are certified to the appropriate safety standards, and adequate protection is provided for the operator against electrical shock or burn, excessive temperature, and spread of fire from the equipment. based on the available information and this evaluation, no deficiency was identified for the architect c8000 processing module list number 01g06-11 for the issue under evaluation. this s the final report. >< further information was received on september 25, 2009 that indicated that the extension block (cord and lead) mentioned in this investigation is not an abbott product. | 0
1509120 | (b)(4). an abbott field service representative (fsr) arrived at the customer site and inspected the analyzer. the tubing from the ict (integrated chip technology) module was discolored and the ict probe was loose. the discolored tubing and ict probe were replaced as well as the sample probe. the fsr verified the analyzer's performance with acceptable precision runs and indicated the probable cause for the customer's issue was a loose ict probe. the customer was made aware that the ict probe must be tightened in position when running. no system logs were available for review for this evaluation. complaint activity was reviewed and identified no adverse trends in association with the issue currently under evaluation. the architect system operations manual ((b)(4), (b)(6) 2009) and ict sample diluent package insert ((b)(4)) provide sufficient information regarding scheduled maintenance, component replacement, sample handling, interpreting results and troubleshooting the customer's current issue. based on the available information, the discrepant results were probably caused by the loose ict probe; however, the 1.0 ml syringes replaced by the customer and the discolored ict tubing and sample probe replaced by the fsr are also possible causes. this is a final report. >< the customer states that discrepant results were reported from the lab that were generated on the architect c8000 analyzer for the chloride assay. the customer states that controls had been within specification from 9 am through 1 pm, but by 3 pm control values had dropped. the customer retested suspect samples on the architect c16000 analyzer and issued corrected reports. no patient management was affected by the reported results. this patient generated an initial result of 84 mmol/l that retested at 95 mmol/l. a service call was initiated. >< (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1325043 | reporter (risk manager) states that the autopulse lifeband load distributing band would not stay secure. reporter states that the air bladder in the band inflated and the velcro holding the two-part band together separated and could not be placed back together. there is no further info at this time. several attempts to contact the user (biomed) were made before learning that he was no longer employed by the reporting agency, and therefore, no further info could be obtained. the autopulse (ap) is indicated only as an adjunct to manual cpr. in the event of malfunction, caregivers should revert to manual cpr. the autopulse in this case was not implicated as the cause of death or serious injury. >< the belts of the returned lifeband were cut by the user as their standard practice. as a result, simulated use testing could not be performed. however, the key components associated with any potential inflation (the air bladder and foam) as well as the fastening/alignment features (velcro and the plastic tab) were inspected, and were confirmed to be free of any damage that could contribute to the reported over inflation. all key components were found to be within our dimensional specification. the air bladder was also subjected to functional tests by the oem manufacturer to verify inflation/deflation properties. the results of this functional testing verified that the air bladder from the returned lifeband performed per the design intent, which prevents over inflation due to increased pressure inside the air bladder. the air bladder is designed to equilibrate with the atmospheric pressure by allowing air to enter the bladder when the ambient atmospheric pressure is significantly higher than the bladder's internal pressure. the valve closes when equilibrium is reached. the typical pressure gradient between atmosphere and the air bladder being small, the bladder cannot overinflate unless a high pressure air source (e.g., compressed air) is applied. the valve has 4 channels to allow the air bladder to reach the state of equilibrium with the atmosphere. with just one channel open, the air bladder may deflate slower but still cannot over inflate. in our experience, if the velcro is not fastened properly, with the alignment tab inserted completely into its matching slot per the instructions for use, the velcro fastening can come abruptly loose during normal use. | 0
1200191 | customer reported that corrected reports were issued for multiple high creatinine results. the account stated four patients received renal biopsies. potentially as an outcome of reported high results. >< the event is currently under investigation. however, our preliminary results indicate that there may have been a failure in washing of the cuvettes. if such a failure occurred, it is likely that this contributed to the event. during the run, excessive nonlinear flags appear on bun results, this is an indication of cuvette problems. in a conversations with the user facility, the olympus representative was informed that facility personnel typically ignore the bun flags (system flags are a warning that something is wrong and their significance is addressed in the user's manual). had the flags been investigated, the cuvette issue would likely have been addressed and the potential for elevated creatine results avoided. based on our preliminary findings, olympus advised the customer site to clean the cuvettes, perform a photocall and cuvette check, calibrate the reagent, and run controls. all qc checks passed and the instrument was returned to use. during an on-site service call for an unrelated event (six days after the instrument was returned to use) it was noticed that the o'rings on the wash station appeared to be worn. the worn o'rings, which may have contributed to the cuvette overflow, were replaced. | 0
1058501 | the specimen was drawn via micro collect and collected in a 300ul container. the erroneous results occurred for a specific pediatric sample. qc is run once a day, and was run prior to the event. previously tested pt samples were not rerun to confirm accuracy back to the last acceptable qc run. a field service engineer (fse) was dispatched to the customer's lab: the fse performed latex and clog detection procedure. the fse verified repair per established procedures and results met published performance specifications. a clear root cause for this event has not been determined to date. a malfunction will be assumed for the purpose of this report. >< a customer contacted beckman coulter inc. (bci) regarding elevated hemoglobin (hgb) and low platelet (plt) results generated by the coulter act diff analyzer for one pt. initial hgb result of 12.1g/dl and plt result of 286x10 to the third power cells/ul were reported out of the lab. the original sample was retested several times and repeated hgb results were: 2 results of 11.8g/dl and 12.3g/dl. repeated plt results were: 364x10 to the third power cells/ul, 378x10 to the third power cells/ul, and 362x10 to the third power cells/ul. the pt was redrawn and sample was tested for hgb and plt on a reference instrument: hgb result was 10.6g/dl, and plt result was 370x10 to the third power cells/ul. the results obtained from the reference instrument were considered correct and were reported out of the lab. based on available info, there was no impact to pt treatment. | 0
1182443 | the samples were collected in 5 ml bd vacutainer tubes, stored at room temperature and were analyzed in primary mode of operation. controls were run before and after the incident with acceptable results. a field service engineer (fse) was on-site in 2008: fse replaced laser, lens block, scatter sensor, and flow cell. fse aligned ttm and checked laser. fse verified repair per established procedures and results meet published performance specifications. a clear root cause for this event has not been determined to date. a malfunction will be assumed for the purpose of this report. >< a customer contacted beckman coulter inc. (bci) regarding erroneously generated differential results by the coulter hmx autoloader for three known oncology patients. all results had instrument generated messages. two out of the three erroneous results were reported out. corrected reports were sent out upon manual differentials or sample rerun. chemotherapy treatment was withheld for one pt. there was no impact to treatment for the other two patients. | 0
1317483 | a patient sample was tested for hemoglobin (hgb) in manual mode and a low result was obtained. hgb results obtained through an automatic mode were considered correct and correlated with results from a different instrument. the actual results were not provided. the erroneous result was not reported out of the lab. there was no death, injury, or change to patient treatment. >< the specimen was collected in edta tube and was sampled fresh. qc was within ranges on automatic mode. drift in qc results was visible on manual mode. a field service engineer (fse) was dispatched: the fse replaced blood sampling valve (bsv) actuator assembly. after the bsv actuator was replaced, the instrument is performing within qc specifications. although hardware issue was addressed by the fse, a clear root cause for this event has not been determined. a malfunction will be assumed for the purpose of this report. | 0
1407405 | the coulter lh500 instrument generated erroneously low hemoglobin (hgb) results on seven (7) patient samples. the results were reported out of the lab. the samples were re-tested and corrected reports were sent out. there was no effect to the patients or user. >< the specimens were collected in plastic corvac edta tubes and bd vacutainer. qc is run 3 times daily. the initial hgb results were printed with user defined h&h check flags (h&h = hemoglobin and hematocrit check flag) a field service engineer (fse) was dispatched: a) the fse reset the rbc and hgb calibration factors and cleaned the blood sampling valve (bsv). b) the fse verified the repair per established procedures. results meet published performance specifications. hardware issue may have contributed to this event. a malfunction will be assumed for the purpose of this report. | 0
1464066 | customer contacted beckman coulter inc. (bci) regarding erroneously low platelet (plt) results generated by the coulter act diff analyzer. erroneously low platelet results of 67x10 to the 3rd power cells/ul were generated on a patient sample. the results were questioned by the physician based on the patient's previous platelet results of 140x10 to the 3rd power cells/ul. the customer sent the patient sample to the reference lab and the platelet count recovered was 140x10 to the 3rd power cells/ul. no manual scans were performed the actual patient results were not provided. there was no affect to patient or user with regards to this event. >< qc is run daily and instrument is currently performing within qc specifications. a field service engineer (fse) was on-site. fse performed reproducibility and verified instrument operation. a clear root cause for this event has not been identified to date. a malfunction will be assumed for the purpose of this report. | 0
1551323 | (b) (4) - case opened for reporting purposes. additional data/failure investigation: the event description listed in the medwatch does not indicate harm to a pt or user. the event appears to have occurred during case preparation when a pt would not typically be present. however, this info cannot be confirmed. no medications listed in the event description, as unavailable, are life saving/life sustaining medications. in the event more info becomes available, a follow-up report will be submitted. >< company review of maude database revealed medwatch, serious injury, life threatening. company complaint management system search revealed no cases that match or resemble the event description of the medwatch. unable to identify facility, reporter , user, or date of event. unable to confirm any data/info on this medwatch. as described per (b) (4). 'event description: pyxis opened to prepare medications as part of pre-brief for open heart surgery. the inventory levels as shown in the drawer for heparin, etomidate, protamine sulfate, pancuronium bromide were zero. all the medications had to be obtained for pt from other machines in other or rooms. the pyxis software was updated the night before. the list of medications needed filled did not print as planned for two of the operating rooms. company contacted and worked on software correcting the problem.' | 1
1315772 | customer complaint received regarding wrong disinfectant type calibration made to system 83 + endoscopic reprocessor. investigation of complaint revealed that customer had provided wrong disinfectant type which led to the calibration error. svc tech was dispatched immediately and calibration error was corrected to indicate the proper disinfectant and parameters. although unlikely, this condition can put a pt at risk for infection if the colonoscope is not immersed in the facility's choice of disinfectant solution for the prescribed time suggested by the mfr. any such info must be accurately supplied by facility personnel and calibrated into the system 83+ at time of initial installation and prior to use using the prescribed parameters dictated by the disinfectant mfr. it is not known how may pts may have been put at risk. >< . | 0
1419033 | the product has not been received for evaluation and a follow-up report will be submitted when the investigation is completed. >< complainant alleged that during a routine shift check by a clinician, the device failed to power on. complainant indicated that there was no patient involvement in the reported malfunction. | 0
1419036 | the device has not been received, and this complaint is still under investigation. >< complainant alleged that during biomed testing the device failed to discharge. complainant indicated that there was no patient involvement in the reported malfunction. | 0
1423595 | complainant alleged that during a routine shift check by a clinician, the device powered on without display. complainant indicated that there was no pt involvement in the reported malfunction. >< the device has been received, and a follow-up report will be submitted when the investigation is completed. | 0
1455114 | the device has been received, and a follow-up report will be submitted when the investigation is completed. >< complainant alleged that during a routine shift check by a clinician, the device would intermittently not power up. complainant indicated that there was no patient involvement in the reported malfunction. | 0
1455116 | the device has been received and a follow-up report will be submitted when the investigation is completed. >< complainant alleged that during a routine shift check by a clinician, the device's defib output was out of specification. complainant indicated that there was no pt involvement in the reported malfunction. | 0
1495684 | complainant alleged that during training by facility staff, the device inappropriately shut down. complainant indicated that there was no patient involvement in the reported malfunction. >< the device has been received and a follow-up report will be submitted when the investigation is completed. | 0
1495686 | zoll medical corporation has not received the device for evaluation, and this complaint is still under investigation. >< complainant alleged that during biomed testing, the device powered on without display. complainant indicated that there was no patient involvement in the reported malfunction. | 0
1299143 | the customer states that while removing a reagent pack from the tdx analyzer, she scratched the back of her middle finger of her left hand with the sample probe. the probe ripped through the glove she was wearing and broke the skin of the finger. the customer was placed a course of 'pep' antibiotics (preventive treatment) with blood samples taken for hepatitis and testing. the customer has experienced nausea and diarrhea as a result of the antibiotics. the customer states that a high risk patient sample was tested that day and inquired about the sample probe washing and sample volume used by the analyzer. there is no further impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the complaint text indicates the customer opened the lid of the tdx analyzer to remove a gentamicin reagent kit and the top of the left middle finger hit the probe. the probe was stationary at this time. the operator performed first aid by removing the glove, squeezing the wound to encourage bleeding and ran the finger under running tap water. following lab policy, the operator obtained medical attention. followup indicates the wound is looking fine and showing no signs of infection. the sample run prior to the incident was from a known homosexual patient with a negative result at last test. a review of the service history of tdx analyzer indicates there were no other injury complaints since the initiation of this complaint. a review was also performed of the complaint database for the time period of 2008 - 2009 and found no similar complaints for the customer described issue. the customer's issue is addressed in the tdx system operation manual which contains adequate information in regards to this issue. the investigation demonstrated that the tdx fpia analyzer is performing within its intended use, label claims and specifications. operator error is the most probable cause of this customer described issue. no deficiency related to the performance of the device was identified. this is a final report. | 0
1060035 | the customer stated a pt sample generated discrepant wbc results on the cell-dyn 4000 analyzer. a pt generated an initial wbc result of 2.32 k/ul and upon retest, the result was 7.64 k/ul. no suspect parameter flags or invalid data flags were generated. no impact to pt management was reported. >< result - wbc flow line and fitting. the wbc syringe cup was examined and no issues with the syringe were observed. the wbc syringe cup was washed. peristaltic pump tubing was changed the previous day. the problem was not resolved after maintenance was completed and field service was dispatched. it was confirmed from the data that blood clots remained in the wbc flow line. the wbc flow line was replaced (including the fitting, which had a blood clot in it). the line from the wbc cup was washed with bleach. precision (n=20) was run twice and normal reproducibility was seen. a specimen was run and recovered as expected. cell-dyn 4000 system operator's manual outlines problem sources, probable causes and corrective actions for imprecise wbc/nrbc data. a flow chart for troubleshooting imprecise wbc is given on pages 10-185, 10-186 which includes, (page 10-186) flushing of the lines into and out of the wbc dilution cup. a review of complaints for the period september 2007 through march 2008, did not indicate an adverse trend for the cell-dyn 4000, for issues with erratic wbc. no systemic issue was identified. this is the final report. end of report. | 0
938366 | (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by (b)(4) and released for distribution on (b)(4) 2009. >< the customer contacted abbott regarding cell-dyn sapphire syringe 'failed to home' error messages observed on the cell-dyn sapphire analyzer. the customer stated the error occurred seven times in 24 hours, and the syringe was cleaned after the initial error. the customer noted the syringe plunger was very stiff. the abbott customer technical advocate (cta) instructed the customer to install new syringe and the customer was to monitor the syringe over the weekend and contact the cta if further assistance was needed. it is unknown if there was impact to patient management. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 5/8/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 6/19/07. a 100% part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. | 0
938367 | (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by (b)(4) and released for distribution on (b)(4) 2009. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 05/08/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 06/19/2007. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. end of report. >< the customer contacted abbott regarding cell-dyn sapphire syringe 'failed to home' error messages observed on the cell-dyn sapphire analyzer. the customer observed the error at least five times. the customer performed troubleshooting procedures by re-seating the syringe with no resolution. the abbott customer technical advocate (cta) advised the customer to replace the syringe. the customer's issue was resolved by replacing the cell-dyn sapphire hemoglobin syringe. it is unknown if there was impact to patient management. | 0
938369 | the customer contacted abbott regarding cell-dyn sapphire syringe 'failed to home' error messages observed on the cell-dyn sapphire analyzer. the customer reported they have tried cleaning procedures with no resolution and did not have a replacement syringe. the customer was provided a replacement syringe and the replacement syringe resolved the customer's issue. it is unknown if there was impact to pt management. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe. list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 05/08/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distributing system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08/may/2007 and 25/june/2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 06/19/2007. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specs were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. end of report. >< (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by (b)(4) and released for distribution on (b)(4) 2009. | 0
938370 | the customer contacted abbott regarding the cell-dyn sapphire hemoglobin syringe 'failed to home' error messages observed on the cell-dyn sapphire hematology analyzer. the customer removed and re-seated the syringe but continued to observe the error message. the abbott customer technical advocate (cta) recommended the customer replace the syringe, however, the customer had another syringe received at the same time the suspect syringe was shipped and requested replacement syringes. the customer discarded the suspect syringe and installed a replacement syringe sent by the abbott cta and the customer's issue was resolved. it is unknown if there was impact to patient management. >< (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted (b)(4) follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by the (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a (b)(4) and released for distribution on (b)(4) 2009. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 5/8/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 6/19/07. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. | 0
938372 | (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by (b)(4) and released for distribution on (b)(4) 2009. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 5/8/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 6/19/07. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. >< the customer contacted abbott to report observing cell-dyn sapphire 'syringe failed to home' error messages on the cell-dyn sapphire analyzer. the customer cleared the message fault, cleaned the syringe which showed no signs of leaking and re-assayed samples with no suspect results. the abbott customer technical advocate advised the customer to remove the syringe to assess syringe plunger movement and replace if indicated. about a week later, the customer contacted abbott to report the sapphire hemoglobin syringe was 'frozen' and could not be moved when cleaning procedures were preformed. a new syringe was installed and all quality control levels were within range and the syringe issue was resolved. there was no impact to patient management reported by the customer. | 0
938374 | the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 5/8/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 6/19/07. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. >< the customer contacted abbott regarding the cell-dyn sapphire hemoglobin syringe used on the cell-dyn sapphire hematology analyzer. the customer reported observing 'syringe fail to home' error message on the analyzer ten days post installation. the customer replaced the syringe and immediately observed the 'syringe fail to home' error message. the abbott customer technical advocate (cta) suggested the customer replace the syringe again, and if the issue persisted, the customer was instructed to call back and abbott field service would be sent to the customer facility. a few days later, the customer was contacted by abbott to follow up. the customer reported, the syringe was replaced as instructed and the issue was resolved. the abbott cta requested the return of the failed syringe for failure analysis. there was no impact to patient management reported by the customer. >< it was initially reported that the device was explanted after an implant duration of zero days. in 2009 (through follow-up with the health-care provider), it was learned that the device was explanted due to a sizing issue. per the operative report two weeks earlier, a 36 mm edwards ring was implanted. 'tee as the patient was briefly unloaded did show about a 2+ central leak and this was judged to be not acceptable so the patient was rearrested with an antegrade dose of cardioplegia. left atrium was reopened and the ring was removed.' the ring was replaced with a smaller size ring. >< (b)(4), concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. the customer used the suspect device at the customer facility. this follow up mdr for remedial correction 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on 30 april 2007 and was identified by the vendor to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of 08 may 2007 through 25 june 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a new vendor and released for distribution on 10 february 2009. | 0
938376 | a non-critical alarm was occurring. the device displayed eri and indicated it should be replaced by 2009. the device was running at .175ml per day and prior to that was not at a very high flow rate. the previous eri dates looked normal with the date at the last refill being 50+ months. only indication of the event was the alarm, the patient had experienced no symptoms related to the event. the device was replaced with no complications. the medication in the pump was morphine sulphate. >< the customer contacted abbott regarding cell-dyn sapphire hemoglobin syringe used on the cell-dyn sapphire hematology analyzer. the customer reported observing 'syringe fail to home' error messages for the hemoglobin syringe. the customer stated weekly syringe maintenance procedures were followed and noticed the syringe was resistant to movement. the customer put the syringe back on the analyzer, but the abbott customer technical advocate (cta) was sending a replacement syringe. a few days later, the cta contacted the customer and verified with the customer the 'syringe fail to home' error was resolved with the installation of the replacement syringe. there was no impact to patient management reported by the customer. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number provided abbott laboratories with an improved version of the syringe, which was released for use on 05/08/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 06/19/2007. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. end of report. >< (b)(4), concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on 30 april 2007 and was identified by the vendor to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of 08 may 2007 through 25 june 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a new vendor and released for distribution on 10 february 2009. | 0
938377 | (b)(4). concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02. in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by (b)(4) to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by (b)(4) and released for distribution on (b)(4) 2009. >< the customer contacted abbott regarding the cell-dyn sapphire hemoglobin syringe used on the cell-dyn sapphire hematology analyzer. the customer reported the 'failed to home' error message had occurred at least 3 times. the abbott customer technical advocate (cta) instructed the customer to replace the cell-dyn sapphire hemoglobin syringe. the customer stated no further homing failures were observed since replacing the syringe. the cta advised the customer to monitor the syringe for additional errors. upon follow up with the customer about a week later, the customer reported no additional 'syringe fail to home' error messages and the customer's issue was resolved. there was no impact to patient management reported by the customer. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 05/08/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 06/19/2007. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. end of report. | 0
938379 | (b)(4), concomitant medical products: cell-dyn sapphire hemoglobin syringe, list # 8h49-02 in the previously submitted medwatch -01 follow up, the customer complaint was inadvertently not associated with remedial action correction 2919069-8/6/07-004-c in error. the customer used the suspect device at the customer facility. this follow up mdr for 2919069-8/6/07-004-c is submitted late. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on 30 april 2007 and was identified by the vendor to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of 08 may 2007 through 25 june 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a new vendor and released for distribution on 10 february 2009. >< the customer reported that abbott field service was on site at the customer facility to address elevated hemoglobin backgrounds on the cell-dyn sapphire hemoglobin analyzer. the field service technician replaced the cell-dyn sapphire hemoglobin syringe and the customer observed 'syringe fail to home' error messages. the customer stated the syringe appeared to be seated correctly. the customer was advised to replace the cell-dyn sapphire hemoglobin syringe. the customer stated that since the new syringe was installed, no further error messages were observed and quality controls were within package insert range, therefore, the issue was resolved. there was no impact to patient management reported by the customer. >< reporting status: serious injury- medical intervention. reporting rationale: occlusion requiring medical intervention. device issue: no device malfunction has been reported. it was reported via a trail that in 2009, a total of two xience v stents (2.5 x 28 mm and 2.5 x 15 mm) were implanted in the 90% stenosed, moderately calcified, mid-rca lesion. however, the following day, the patient experienced ischemia and transient vessel occlusion and a branch occlusion, which required treatment with a 2.5 x 18 mm stent. the reported patient effects resolved the same day. no additional event or patient information is available. >< the cell-dyn sapphire hemoglobin syringe is used on the cell-dyn sapphire analyzer, an automated hematology analyzer. the vendor for the cell-dyn sapphire hemoglobin syringe, list number 08h49-02 provided abbott laboratories with an improved version of the syringe, which was released for use on 5/8/2007. the vendor sent a letter to abbott stating the hemoglobin syringes with a specific manufacturing code, were assembled with insufficient or inconsistent amount of silicone lubricant applied to the plunger tip. the absence of sufficient lubricant causes premature failure of the syringe, which generated 'syringe failed to home' error messages upon installation on the cell-dyn sapphire analyzer or shortly thereafter. in response to the vendor's letter, abbott internal nonconformance was addressed via stock sweeps to segregate non-conforming product in inventory. a world wide quality hold was issued 22 june 2007 to remove suspect product from the distribution system and a customer letter was generated to users of the cell-dyn sapphire analyzers. the defective syringes were identified in the customer letter as those packaged and shipped between 08 may 2007 and 25 june 2007. a review of abbott's complaint analysis system from january 2007 to present did not indicate an adverse trend for hemoglobin syringes. as the affected syringes did not meet the vendor's internal lubricant requirements, the vendor provided the following corrective actions: vendor employee retraining for syringe lubricant application was completed on 6/19/07. one hundred percent part inspection of the syringe will be performed and inspection results will be attached to the syringe shipment. abbott syringe specifications were defined for use by the vendor. certification documents are attached on every shipment of syringes. this is the final report. | 0
989949 | the customer is questioning the out of range hemoglobin backgrounds that recovered high on the cell-dyn sapphire analyzer using a newly installed cell-dyn sapphire hemoglobin syringe, list # 8h49-02. the issue was resolved by replacing the damaged parts and replacing the hemoglobin syringe. there was no impact to pt mgmt reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is completed. >< (b)(4). customers reported that the hemoglobin background counts were out of the specification high after installing and running the cell-dyn sapphire hemoglobin reagent syringe list number 08h49-02 on the cell-dyn sapphire instrument. the specification for the hemoglobin background is less than or equal to 0.10 grams per deciliters per the cell-dyn sapphire operators manual. background counts must be within specification prior to reporting of patient samples. investigation of this issue found that the elevated hgb background count is due to excess lubricant used by the vendor during their manufacturing process of the cell-dyn sapphire hemoglobin reagent syringe that is present on the tip of the plunger, lubricant should only be applied to the side of the plunger. the cell-dyn sapphire hemoglobin reagent syringe list number 08h49-02 is distributed for use on the cell-dyn sapphire instrument only. this issue has no impact on other platforms or instruments. the following corrective actions were put in place in order to address the issue a product recall letter fa11jan2008 was issued with its associated field action and distributed to affected customers. affected cell-dyn sapphire hemoglobin reagent syringes were identified to have packaged dates of (b)(6) 2007 through (b)(6) 2007 for fa11jan2008. the vendor, (b)(4), has implemented and updated their procedures in regards to the application of the lubricant and cleaning of the cell-dyn sapphire hemoglobin reagent syringe as part of their manufacturing process to address the excess lubricant issue. abbott (b)(4) has implemented an incoming quality inspection requirement for the cell-dyn sapphire hemoglobin reagent syringe prior to testing or use on the manufacturing line. this is a final report. | 0
1000760 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer contacted abbott regarding hemoglobin syringe 'fail to home' error messages generated from the cell-dyn sapphire hematology analyzer. the customer was advised to clean the syringe and/or replace if necessary. the customer reported, the issue was resolved by replacing the syringe. additionally, the customer reported there was no impact to pt management. >< (b)(4). this follow up mdr for remedial correction 2919069-8/6/07-004-c is submitted late. concomitant medical products: cell-dyn sapphire hemoglobin syringe. list # 8h49-02. this follow up mdr for remedial correction 2919069-8/6/07-004-c is submitted late. upon further review, the customer complaint is associated with remedial correction 2919069-8/6/07-004-c, as an unknown lot of suspect device was in use at the customer facility. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by the vendor to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a new vendor and released for distribution on (b)(4) 2009. | 0
1000764 | the customer contacted abbott regarding hemoglobin syringe 'fail to home' error messages generated from the cell-dyn sapphire hematology analyzer. the customer had no syringes to replace the faulty syringe, so the customer was sent replacement syringes. the customer reported there was no impact to pt mgmt. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). this follow up mdr for remedial correction 2919069-8/6/07-004-c is submitted late. concomitant medical products: cell-dyn sapphire hemoglobin syringe. list # 8h49-02. upon further review, the customer complaint is associated with remedial correction 2919069-8/6/07-004-c, as an unknown lot of suspect device was in use at the customer facility. the cell-dyn sapphire hemoglobin syringe, list number 8h49-02, was manufactured on (b)(4) 2007 and was identified by the vendor to have been manufactured with insufficient or inconsistent amount of silicone lubricant applied to the tip of the syringe plunger. the affected syringes with a package date of (b)(4) 2007 through (b)(4) 2007, caused the cell-dyn sapphire analyzer to generate an error message upon installation of the syringe or shortly thereafter. the cell-dyn sapphire hemoglobin syringe was distributed for the cell-dyn sapphire analyzer only and did not impact other cell-dyn analyzers. the issue was resolved when a new cell-dyn sapphire hemoglobin syringe, list 8h49-04, was manufactured by a new vendor and released for distribution on (b)(4) 2009. | 0
1056764 | investigation summary: the customer reported that a cell-dyn sapphire analyzer generated low results on one patient which showed no errors and so were reported out. the physician questioned the results. a new draw was made and this sample generated results that were approximately twice that of the first sample. the patient had to be redrawn, with a delay in reporting with no adverse impact to patient management reported. a technical service specialist (tss) ) visited in 2008 and resolved the problem by replacing the probe detector. the probe detector function is explained in the cell-dyn sapphire tm system operator's manual: during autoloader processing, the aspiration probe motor moves the probe down through the vent needle. when the probe touches the tube bottom, a sensor is tripped; this stops the downward movement of the probe and raises the probe slightly for aspiration. there were no short sample errors displayed, but if the probe detector is not working, the probe will not move upward after touching the bottom of the tube. aspiration of the sample may be impeded by close contact with the bottom of the tube. dispersional data alerts were present on the first sample for rbc, rbco, hgb, hct. the operator's manual, on page 3-55, states: a result that falls outside a laboratory action limit can also indicate the need for the operator to follow a laboratory protocol, such as repeating the sample, performing a stained blood film review or notifying the physician. the first sample was tested in closed mode. if it had been repeated in open mode, the issue might have been identified for the run in closed mode. trending: a review of the complaint reports, for 2007 through 2008, did not find any adverse trend for the complaint issue of low results on multiple parameters. labeling: the event is addressed in the cell-dyn sapphire tm system operator's manual: use or function; system hardware; aspiration probe; 1-31 section 3: principles of operation; system initiated messages and data flags; dispersional data alerts; 3-55. conclusion: the device involved in the issue was evaluated. a broken probe detector was identified as the cause for the malfunction that contributed to the event. the probe detector was replaced resolving the issue. there was no further impact to patient management reported for this issue. this is a final report. end of report. >< the customer states that a sample from a pregnant patient was tested using a cd sapphire analyzer generating the following results; wbc=5.24 k/ul, rbc=1.76 m/ul, hgb=3.70 g/dl and plt=119 k/ul. no errors were generated by the analyzer. the reported results were questioned by a physician. a new sample was drawn and all the results were about twice as high wbc=12.2 k/ul, rbc=3.98 m/ul, hgb=7.49 g/dl and plt=256 k/ul. there was a delay in reporting the retest results with no reported adverse impact to patient management reported. | 0
1056770 | the account stated that the cell-dyn 1800 analyzer has generated a discrepant platelet result when compared to the platelet result tested in another laboratory. the initial platelet result = 46 k/ul, when tested in another lab, the result was 13 k/ul and was confirmed with a blood smear. the account's qc results have been in range. the doctor questioned the results and stated that the patient has a history of having a low platelet count but the results that generated are too low. there is no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1059457 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that a pt sample generated a very low wbc result when tested using a cell-dyn 4000 analyzer. the sample was repeated on the same analyzer yielding a different wbc value. the initial low results were not reported out of the lab with no adverse impact to pt management reported due to this issue. initial wbc=0.755 k/ul. repeat wbc=9.56 k/ul. | 0
1197109 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the cell-dyn 1700 analyzer is generating hgb and hct discrepant results. the account stated that on 1 patient sample, the hct=21% and the retest result was 40%. the account also stated that there are discrepant results on multiple parameters on several patient samples. the account only provided initial results for the patients. the qc was within range. there were no results reported. field service was requested. there is no impact to patient management reported. >< (b)(4), evaluation; silicon tubing replaced as precautionary measure. on (b)(6) 2008, (b)(6), reported discrepant hemoglobin (hgb) and hematocrit (hct) results with the cell-dyn 1700 instrument. the issue had been occurring for over a week. the first run showed a hct of 21% and after repeating the run it showed 40%. quality control (qc) was within assay range and the syringes showed no adverse issues. field service representative (fsr) was dispatched for issue resolution. during the fsr visit on (b)(6) 2008, the sample syringe and 10-ml syringe were replaced due to leakage to improve precision. instrument was performing within specifications. on (b)(6) 2008, the customer faxed results generated by the cell-dyn 1700 on two patient samples. the data showed results as followed (no repeat results were submitted): wbc (k/ul): 3.6, rbc (m/ul): 5.2, hgb (g/ml): 19.1, hct(%): 15.6, mcv (fl): 81.9, mchc (g/ml): 33.3, plt (k/ul): 185; wbc (k/ul): 6.3, rbc (m/ul): 2.57, hgb (g/ml): 6.7, hct(%): 20.3, mcv (fl): 78.9, mchc (g/ml): 33.0, plt (k/ul): 203. on (b)(6) 2008, the customer was still getting discrepant results for multiple parameters. the customer faxed 5 pages of patient data for review. no results had been reported out. the customer sent the same sample to the reference lab and sample was less than 8 hours old. all samples being collected in edta/mix per abbott guidelines. qc was within range, no data given. there was no possible clot in the system. customer was informed not to run or report any patient results out from this instrument. the customer technical advocate dispatched service. faxed data showed results as followed (no repeat results were provided): wbc (k/ul): 0.6, rbc (m/ul): 0.86, hgb (g/ml): 1.9, hct(%): 5.3, mcv (fl): 61.5, mchc (g/ml): 35.8, plt (k/ul): 29.0; wbc (k/ul): 2.6, rbc (m/ul): 2.50, hgb (g/ml): 7.7, hct(%): 22.4, mcv (fl): 89.6, mchc (g/ml): 34.4, plt (k/ul): 48.0. during the fsr visit on (b)(6) 2008, the silicon tubing (s1) was replaced as a precautionary measure to improve the sample purity from the sample probe. no parts failures were detected. the instrument was performing within specifications. no further investigation was required. the cell-dyn 1700 system operators manual, list number 03h58-01, revision f, under section 10, troubleshooting and diagnostics, page 10-9, provides basic troubleshooting to assist the operator in identifying, isolating and resolving instrument problems. section 9, service and maintenance, nonscheduled maintenance frequency, table 9.1, page 9-19, recommends cleaning/replacement of the sample syringe when it is suspected to be the source of imprecision. a review of the complaint analysis trending system (cats) and trending analysis support system (tass) reports did not indicate any adverse trend for the cell-dyn 1700; list 03h53-01 for the complaint issue. a review of service tickets from (b)(6) 2008 through (b)(6) 2009 on this instrument (s/n (b)(4)) showed that no others issues related to discrepant results for multiple parameters have occurred. this is the final report. | 0
1222603 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). debris in solenoid valve, hemoglobin flow cell and/or aspiration pathway. the customer technical advocate (cta) verified that the hemoglobin reference value and hemoglobin output were within specification. all of the reagents were correct. the cta instructed the customer to clean the optical and hemoglobin flow cell and to recheck the hemoglobin reference value/output after flow cell cleaning. the cta reviewed the hemoglobin reference value and output and adjusted the hemoglobin current. a subsequent precision study yielded out of range results for hemoglobin, rbc and platelets. the cta suggested for the customer to clean the open mode aspiration pathway, solenoid valve, and possibly the hgb flow cell again. the cta followed up with the customer who stated after completing the recommended troubleshooting, precision was run and all parameters passed the specifications. the customer subsequently calibrated the instrument with hemcal plus calibrator. post-calibration, qc was well within range (only open mode was used). patient results showed good hemoglobin/hematocrit matching and also compared well with the reference lab results (customer could not compare with the cell-dyn ruby, as service for this instrument was required). the customer considered the issue on the cell-dyn 3200 instrument resolved by cleaning and calibration. no further assistance was required. the likely cause of the issue was attributed to debris in the solenoid valve, hgb flow cell, and /or aspiration pathway and the need to adjust related calibrator factors and hgb current to bring recoveries near target. the cell-dyn 3200 system operator's manual (list number 06h60-01, revision n) under section 9, page 9-29, service and maintenance, nonscheduled maintenance frequency, indicates that cleaning of the hgb flow cell should be performed when it is suspected of causing elevated hgb results or hgb imprecision. section 10, troubleshooting and diagnostics, pages 10-23 through 10-27, provides troubleshooting instructions for identifying, isolating and correcting data related issues. section 6, calibration procedures, general information, page 6-4, provides general information as to when to calibrate the cell-dyn 3200. a review of complaints for the period of (b)(4) 2007 through (b)(4) 2008, did not indicate any adverse trend for the cell-dyn 3200, list number 04h60-01, related to issues with discrepant results with rbc, hgb, mcv, and hct. based on the investigation, no product issue was identified for the cell-dyn 3200, list number 04h60-01, for issues related to discrepant results with rbc, hgb, mcv, and hct. there was no systemic issue identified for the cell-dyn 3200 product line. this is the final report. >< the customer stated results are imprecise across multiple parameters on the cell-dyn 3200 analyzer. a patient sample was tested on both the cell-dyn 3200 and cell-dyn ruby with the following results: cd 3200 cd rubyrbc 3.34 m/ul 4.07 m/ulhgb 9.23 g/dl 11.1 m/ulmcv 87.4 fl 93.4 flhct 29.2% 38%no impact to patient management was reported. | 0
1224828 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< in 2008, a hospital reported platelet (plt) scattered and poor reproducibility with the cell-dyn 3700 analyzer. the customer stated that when it showed a high value for the first time, the histogram looked like red blood cell (rbc) and plt were connected, but it was not connected with the retest, it had been clearly separated. no flagging had appeared with the first specimen, but another specimen showed rbc morphology. the customer technical advocate (cta) requested to the customer to perform reproducibility 10 times, but the customer did not agree. the cta suggested a field service representative visit in order to resolve the issue; the customer refused service due to cost related issues. nine days later, the customer performed a reproducibility testing 10 times, and the cv% settled in less than 5%. the customer investigated the issue and stated that it was not at problem level for the routine testing, and decided not to do repair service at this time. no further investigation was required. there was no impact to the reported results, as results were reported out after retesting. the cell-dyn 3700 system operator's manual, revision f, under section 10, troubleshooting guide, pages 10-27 through 10-28, provides instructions for problem identification, problem isolation, and corrective action. page 10-29 states that a specific procedure should be performed when indicated in this chapter or at the request of the abbott customer support specialist. chapter 3, principles of operation, operational messages and data flagging, page 3-37, indicates that it is import to note that there are commonly occurring substances that can affect the results reported by hematology analyzers. while the cell-dyn 3700 has been designed to detect and flag many of these substances, it may not always be possible to do. a review of the complaint analysis trending system (cats) and trending analysis support system (tass) reports, for the period october 2007 through june 2008, did not indicate any adverse trend for the cell-dyn 3700, for the reported issue. june 2008 through december 2008 reports have been evaluated and no potential adverse trends have been identified. based on the investigation, no product issue was identified for plt scattered and poor reproducibility on the cell-dyn 3700 instrument. this is the final report. >< the account stated that the cell-dyn 3700 sl analyzer generated a platelet result of 33,700/ul. the result was considered elevated and therefore was repeated. the repeat results were: 20,700/ul, 25,100/ul and 18,900/ul. the original result was not reported. there was no impact to patient management. | 0
1228323 | investigation in process, unable to determine results and conclusion at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account generated discrepant cell-dyn 1800 hemoglobin results on 2 patients. patient 1 tested cell-dyn 1800 hemoglobin = 8.8 g/dl. the low hemoglobin result was questioned by the physician as the expected result was normal. the specimen was sent to a reference laboratory and tested hemoglobin = 11.9 g/dl. no cell-dyn 1800 messages were generated with the low hemoglobin result. no patient treatment was affected or altered by the low hemoglobin result. the specimen was obtained by venipuncture. the cell-dyn 1800 was recently installed at the account. after the discrepant hemoglobin results were generated the account was unable to process patient specimens due to quality control out of range and unable to calibrate the cell-dyn 1800. service was requested for the cell-dyn 1800. service replaced the hemoglobin flow cell, pinch tubing and 3 syringe drives on the cell-dyn 1800 analyzer. >< hemoglobin (hgb) flow cell and pinch tubing (s2). service made several site visits to the account and replaced several parts on the cell-dyn 1800 analyzer. a review of service tickets from 2008 through 2009 on this instrument showed that no others issues related to discrepant hgb results have occurred. a review of the device history record (dhr) did not show any issues related to precision (open and closed modes). the product performed within specifications at the time of the build. a review of the complaint analysis trending and tracking system reports, for the period october 2007 through december 2008, did not indicate any adverse trend for the cell-dyn 1800, list 07h77-01, for the complaint issue. the cell-dyn 1800 system operators manual, list number 07h80-01, revision e, under section 10, troubleshooting and diagnostics, pages 10-11 through 10-13, provides information to assist in problem identification, isolation, and corrective actions. instructions for obtaining technical assistance are included as well. section 9, service and maintenance, preventive maintenance schedule, page 9-3, states that cleaning or replacement of the syringes is an as required service performed on the cell-dyn 1800 instrument. based on the investigation, no product issue was identified for the cell-dyn 1800 for discrepant hgb results. the cause of the reported issue could not be determined. there was no systemic issue for the cell-dyn 1800 product line. this is a final report. | 0
1228343 | (b)(4). other (b)(4) - hemoglobin (hgb) flow cell and pinch tubing (s2). service made several site visits to the account and replaced several parts on the cell-dyn 1800 analyzer. a review of service tickets from (b)(4) 2008 through (b)(4) 2009 on this instrument showed that no others issues related to discrepant hgb results have occurred. a review of the device history record (dhr) did not show any issues related to precision (open and closed modes). the product performed within specifications at the time of the build. a review of the complaint analysis trending and tracking system reports, for the period (b)(4) 2007 through (b)(4) 2008, did not indicate any adverse trend for the cell-dyn 1800, list 07h77-01, for the complaint issue. the cell-dyn 1800 system operators manual, list number 07h80-01, revision e, under section 10, troubleshooting and diagnostics, pages 10-11 through 10-13, provides information to assist in problem identification, isolation, and corrective actions. instructions for obtaining technical assistance are included as well. section 9, service and maintenance, preventive maintenance schedule, page 9-3, states that cleaning or replacement of the syringes is an as required service performed on the cell-dyn 1800 instrument. based on the investigation, no product issue was identified for the cell-dyn 1800 for discrepant hgb results. the cause of the reported issue could not be determined. there was no systemic issue for the cell-dyn 1800 product line. this is a final report. >< investigation in process, unable to determine results and conclusion at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is completed. >< the account generated discrepant cell-dyn 1800 hemoglobin results on 2 patients. patient 2 tested cell-dyn 1800 hemoglobin = 8.0 g/dl. the low hemoglobin result was questioned by the physician as the expected result was normal. the specimen was sent to a reference laboratory and tested hemoglobin = 14.0 g/dl. no cell-dyn 1800 messages were generated with the low hemoglobin result. no patient treatment was affected or altered by the low hemoglobin result. the specimen was obtained by venipuncture. the cell-dyn 1800 was recently installed at the account. after the discrepant hemoglobin results were generated, the account was unable to process patient specimens due to quality control out of range and unable to calibrate the cell-dyn 1800. service was requested for the cell-dyn 1800. service replaced the hemoglobin flow cell, pinch tubing and 3 syringe drives on the cell-dyn 1800 analyzer. | 0
1242663 | evaluation code other: device not returned, insufficient information provided; inconclusive, no device evaluation performed. in response to this issue an investigation was initiated to further examine the customer issue. the investigation included a review of the complaint text, a search for similar complaints, and a review of labeling. in 2008, hospital in foreign country, reported discrepant results on two patient samples generated by the cell-dyn sapphire. abbott complaint investigation department attempts to obtain additional information were made without success. due to insufficient information, it is not possible to identify the cause of the event; however, a review of the instrument's service tickets from 19 nov 2008 through 13 feb 2009, did not show any other complaints related to discrepant results on patient samples. a review of the complaint trending system reports, for the period october 2007 through october 2008, did not indicate any adverse trend for the cell-dyn sapphire, l/n 08h00-01, discrepant results on patient samples. the cell-dyn sapphire system operator's manual, troubleshooting and diagnostics, provides information in regards to troubleshooting data-related problems. based on the investigation, no product issue was identified for the cell-dyn sapphire for discrepant results on patient samples.this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that low results were recovered from two patient samples when analyzed on the cell-dyn sapphire analyzer. the results were acceptable when the patient samples were repeated on another cell-dyn sapphire analyzer. see data below.patient 1 cell- dyn sapphire cell- dyn sapphire wbc (k/ul) 4.45 15.0rbc (m/ul) 1.68 2.74hgb (g/dl) 4.61 8.45hct (%) 15.4 25.3platelet (k/ul) 296 224 (invalid data flag)the customer states that the suspect low patient results were not reported out of the laboratory. there is no impact to patient management reported. | 0
1242683 | (b)(4). evaluation other: (b)(4) - device not returned, insufficient information provided; (b)(4) - inconclusive, no device evaluation performed. in response to this issue an investigation was initiated to further examine the customer issue. the investigation included a review of the complaint text, a search for similar complaints, and a review of labeling. on (b)(6) 2008, (b)(6), reported discrepant results on two patient samples generated by the cell-dyn sapphire, (b)(4). abbott (b)(4) complaint investigation department attempts to obtain additional information were made without success. due to insufficient information, it is not possible to identify the cause of the event; however, a review of the instrument's service tickets from (b)(6) 2008 through (b)(6) 2009, did not show any other complaints related to discrepant results on patient samples. a review of the complaint trending system reports, for the period (b)(6) 2007 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn sapphire, l/n 08h00-01, discrepant results on patient samples. the cell-dyn sapphire system operator's manual, list number 08h10-01, revision d, under section 10, troubleshooting and diagnostics, page 10-121, provides information in regards to troubleshooting data-related problems. based on the investigation, no product issue was identified for the cell-dyn sapphire for discrepant results on patient samples. this is the final report. >< the customer states that low results were recovered from two patient samples when analyzed on the cell-dyn sapphire analyzer. the results were acceptable when the patient samples were repeated on another cell-dyn sapphire analyzer. see data below.patient 2 cell- dyn sapphire cell- dyn sapphire wbc (k/ul) 3.94 7.49rbc (m/ul) 3.39 4.05hgb (g/dl) 8.33 11.0hct (%) 26.3 31.2platelet (k/ul) 276 326the customer states that the suspect low patient results were not reported out of the laboratory. there is no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1249263 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). linearity testing performed. sample-related issue (interfering substance). investigation plan/summary: on (b)(6) 2008, (b)(6), reported seeing over counting platelet (plt) with the cell-dyn ruby instrument, serial number (b)(4). scatter had particles above plt population being included in the count, result was 2-4 times with a cell-dyn 3700 and smear estimate. during a field service representative (fsr) visit to the customer on (b)(6) 2008, the 7.0 microsphere polymer (latex) was checked and the gains were adjusted. the 3.0 microsphere polymer (latex) was checked and the gains were adjusted. instrument met specifications, controls were in range, and no further investigation was required. no parts were replaced. on (b)(6) 2008, technical support specialist (tss) reported additional information related to patient under specimen id (b)(6), including diagnosis and further review of (b)(6). the tss also performed dilution linearity studies on this instrument and the cell-dyn 3700 in the lab. the tss stated that based on very abnormal plt histogram on this patient and in combination with the 'rbc morph' flag, the customer was educated that the sample appeared abnormal and recommended reviewing smear abnormal rbc or plt morphology. after the tss performed the linearity testing for plt with one sample (from 1:1 to 1:64) on both cell-dyn ruby instrument (other s/n is (b)(4)) and the cell-dyn 3700 instrument in the lab, data indicated that the cell-dyn ruby performed well. the comparison between the cell-dyn ruby instrument, s/n (b)(4), and the cell-dyn 3700 instrument is below (note: the result in parenthesis is the expected recovery value): cell-dyn ruby s/n (b)(4): 1:1 dilution plt=345.5 (345.5). cell-dyn 3700: 1:1 dilution plt=324.5 (324.5). cell-dyn ruby s/n (b)(4): 1:2 dilution plt=169 (172.75); cell-dyn 3700: 1:2 dilution plt=189 (162.25). cell-dyn ruby s/n (b)(4): 1:4 dilution plt=81.15 (86.75); cell-dyn 3700: 1:4 dilution plt=91.55 (81.13). cell-dyn ruby s/n (b)(4): 1:8 dilution plt=40.95 (43.19); cell-dyn 3700: 1:8 dilution plt=41.9 (40.56). cell-dyn ruby s/n (b)(4): 1:16 dilution plt=18.1 (21.59); cell-dyn 3700: 1:16 dilution plt=20 (20.28). cell-dyn ruby s/n (b)(4): 1:32 dilution plt=9.45 (10.8); cell-dyn 3700: 1:32 dilution plt=8.58 (10.14). cell-dyn ruby s/n (b)(4): 1:64 dilution plt=3.39 (5.40); cell-dyn 3700: 1.64 dilution plt=2.76 (5.07). the tss stated that based on the abnormal plt scatter for sample id (b)(6) and the plt linearity study results, the issue was most likely specimen related. the cell-dyn ruby system operator's manual, list number 08h56-03, revision c, section 10, troubleshooting and diagnostics, overview, pages 10-8 through 10-9, provides troubleshooting instructions for imprecise or inaccurate data. appendix b, potential causes of spurious results, table b.1, page b-2, indicates that causes of spurious decreased may be due to cryoglobulin, cryofibrinogen hemolysis (in vivo and in vitro), microcytic red cells, red cell inclusions and white cell fragments. section 7, operational precautions and limitations, page 7-8, provides information in regards to interfering substances that may affect patient results. a review of complaint reports, for the period (b)(6) 2007 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn ruby, l/n 08h67-01, for the reported issue. (b)(6) 2008 reports are in process. based on the investigation, no product issue was identified for the cell-dyn ruby for seeing over counting plt. the event is addressed in labeling: cell-dyn ruby system operator's manual, list number 08h56-03, revision c: section 10: troubleshooting and diagnostics, overview, pages 10-8 through 10-9, provides troubleshooting instructions for imprecise or inaccurate data. appendix b: potential causes of spurious results, table b.1, page b-2 section 7: operational precautions and limitations, page 7-8 conclusion: no product deficiency was identified. the cause of the issue is likely specimen related. this is the final report. >< the customer states that platelet results for one patient appeared to be higher than expected when generated using a cell-dyn ruby analyzer. the cell-dyn ruby generated a platelet result of 48.9 k/ul compared to a platelet result of 8.5 k/ul generated using an alternate method (cell-dyn 3700). the cell-dyn 3700 result matched the platelet estimate per blood smear review. no adverse impact to patient was reported due to this issue. | 0
1249283 | (b)(4): linearity testing performed. sample-related issue (interfering substance). investigation plan/summary: on (b)(6) 2008, (b)(6), reported seeing over counting platelet (plt) with the cell-dyn ruby instrument, serial number (b)(4). scatter had particles above plt population being included in the count, result was 2-4 times with a cell-dyn 3700 and smear estimate. the customer requested field service for issue resolution. during a field service representative (fsr) visit to the customer on (b)(6) 2008, the 7.0 microsphere polymer (latex) was checked and the gains were adjusted. the 3.0 microsphere polymer (latex) was checked and the gains were adjusted. instrument met specifications, controls were in range, and no further investigation was required. no parts were replaced. the tss also performed dilution linearity studies on this instrument and the cell-dyn 3700 in the lab. the tss stated that based on very abnormal plt histogram on this patient and in combination with the 'rbc morph' flag, the customer was educated that the sample appeared abnormal and recommended reviewing smear abnormal rbc or plt morphology. after the tss performed the linearity testing for plt with one sample (from 1:1 to 1:64) on both cell-dyn ruby instrument (other s/n is (b)(4)) and the cell-dyn 3700 instrument in the lab, data indicated that the cell-dyn ruby performed well. the comparison between the cell-dyn ruby instrument, s/n (b)(4), and the cell-dyn 3700 instrument is below (note: the result in parenthesis is the expected recovery value): cell-dyn ruby s/n (b)(4): 1:1 dilution plt=359.5 (359.5); cell-dyn 3700: 1:1 dilution plt=324.5 (324.5). cell-dyn ruby s/n (b)(4): 1:2 dilution plt=180 (179.75); cell-dyn 3700: 1:2 dilution plt=189 (162.25). cell-dyn ruby s/n (b)(4): 1:4 dilution plt=89.8 (89.8); cell-dyn 3700: 1:4 dilution plt=91.55 (81.13). cell-dyn ruby s/n (b)(4): 1:8 dilution plt=39.05 (44.94); cell-dyn 3700: 1:8 dilution plt=41.9 (40.56). cell-dyn ruby s/n (b)(4): 1:16 dilution plt=18.1 (22.47); cell-dyn 3700: 1:16 dilution plt=20 (20.28). cell-dyn ruby s/n (b)(4): 1:32 dilution plt=8.26 (11.23); cell-dyn 3700: 1:32 dilution plt=8.58 (10.14). cell-dyn ruby s/n (b)(4): 1:64 dilution plt=3.67 (5.62); cell-dyn 3700: 1.64 dilution plt=2.76 (5.07). the tss stated that based on the abnormal plt scatter for sample id (b)(4) and the plt linearity study results, the issue was most likely specimen related. the cell-dyn ruby system operator's manual, list number 08h56-03, revision c, section 10, troubleshooting and diagnostics, overview, pages 10-8 through 10-9, provides troubleshooting instructions for imprecise or inaccurate data. appendix b, potential causes of spurious results, table b.1, page b-2, indicates that causes of spurious decreased may be due to cryoglobulin, cryofibrinogen hemolysis (in vivo and in vitro), microcytic red cells, red cell inclusions and white cell fragments. section 7, operational precautions and limitations, page 7-8, provides information in regards to interfering substances that may affect patient results. a review of complaint reports, for the period (b)(6) 2007 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn ruby, l/n 08h67-01, for the reported issue. (b)(6) 2008 reports are in process. (b)(6) 2008 and (b)(6) 2009 reports are not available. based on the investigation, no product issue was identified for the cell-dyn ruby for seeing over counting plt. the event is addressed in labeling: cell-dyn ruby system operator's manual, list number 08h56-03, revision c: section 10: troubleshooting and diagnostics, overview, pages 10-8 through 10-9, provides troubleshooting instructions for imprecise or inaccurate data. appendix b: potential causes of spurious results, table b.1, page b-2. section 7: operational precautions and limitations, page 7-8. conclusion: no product deficiency was identified. the likely cause is specimen related. this is the final report. >< the customer states that platelet results for one patient appeared to be higher than expected when generated using a cell-dyn ruby analyzer. the cell-dyn ruby generated a platelet result of 48.0 k/ul compared to a platelet result of 8.50 k/ul generated using an alternate method (cell-dyn 3700). the cell-dyn 3700 result matched the platelet estimate per blood smear review. no adverse impact to patient was reported due to this issue. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1251184 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation, results, (b)(4) (other), labeling information inadequate/incorrect. an investigation was conducted to solve the issue of the occasional high platelets (plt) background counts reported while executing quality control (qc) runs. during subsequent validation testing, elevated red blood cells (rbc) background was also observed. the qc carryover counting flow sequence is the same for cbc test selection which is the most frequent test used by the customer. the increased trend of plt carryover failures during a review of instrument investigation reports for the months of (b)(6) and (b)(6) of 2008 was also observed. the investigation found a design change occurred relative to a predecessor instrument that impacted carryover. this design error allows bubbles that are generated from bubble mixing (agitation by bursts of air pressure) of the specimen solution to extend into the upper surface areas of the rbc/plt mixing chamber that were not intended to be rinsed by design. this occasionally causes plt % carryover performance to exceed specifications. the rbc % carryover is impacted as well for the same reason. it was determined that the carryover performance for plt is less than or equal to 1.7% and for rbc is less than or equal to 1.2%. carryover parameters were not challenged under the cbc test selection during validation and verification and clinical testing prior to product launch. a cell-dyn ruby product deficiency was established for the plt and rbc % carryover parameters that led to field action (b)(4). the communication to customers included modified specifications for carryover to the plt and rbc parameters. the plt and rbc % carryover is less or equal to 1.0% in the cell-dyn ruby system operator's manual. it was recommended to update the carryover labeling claims for plt to 1.7% and for rbc% to 1.2%. this is a final report. >< an in-house software engineer reported that the cell-dyn ruby analyzer has shown high background counts intermittently on platelet parameter and occasionally on the red blood cells. noc flow error was also displayed under the basophiles result in the run view. it flags invalid numerical results for white blood cells and differential. this issue occurred after running high level of cell-dyn 29 plus control, lot # 80989. no impact to patient management was reported. | 0
1251186 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation, results, labeling information inadequate/incorrect. an investigation was conducted to solve the issue of the occasional high platelets (plt) background counts reported while executing quality control (qc) runs. during subsequent validation testing, elevated red blood cells (rbc) background was also observed. the qc carryover counting flow sequence is the same for cbc test selection which is the most frequent test used by the customer. the increased trend of plt carryover failures during a review of instrument investigation reports for the months of (b)(6) of 2008 was also observed. the investigation found a design change occurred relative to a predecessor instrument that impacted carryover. this design error allows bubbles that are generated from bubble mixing (agitation by bursts of air pressure) of the specimen solution to extend into the upper surface areas of the rbc/plt mixing chamber that were not intended to be rinsed by design. this occasionally causes plt % carryover performance to exceed specifications. the rbc % carryover is impacted as well for the same reason. it was determined that the carryover performance for plt is less than or equal to 1.7% and for rbc is less than or equal to 1.2%. carryover parameters were not challenged under the cbc test selection during validation and verification and clinical testing prior to product launch. a cell-dyn ruby product deficiency was established for the plt and rbc % carryover parameters that led to field action fa17nov2008. the communication to customers included modified specifications for carryover to the plt and rbc parameters. the plt and rbc % carryover is less or equal to 1.0% in the cell-dyn ruby system operator's manual. it was recommended to update the carryover labeling claims for plt to 1.7% and for rbc% to 1.2%. this is a final report. >< an in-house software engineer reported that seven cell-dyn ruby instruments have shown high background counts consistently on platelets parameter and occasionally on noc flow error. the noc flow error is displayed under basophiles result in the run view. it flags invalid numerical results for white blood cells and differential. this issue occurred after running high level of cell-dyn 29 plus control, lot # 80989. there was no impact to patient management reported. | 0
1251187 | an in house software engineer reported that the cell-dyn ruby analyzer has shown high background counts intermittently on platelet parameter and occasionally on white blood cells with an error code of noc. the noc flow error is displayed under basophiles result in the run view. it flags invalid numerical results for white blood cells and differential. this issue occurred after running high level of cell-dyn 29 plus control, lot # 80989. no impact to patient management was reported. >< (b)(4), evaluation, labeling information inadequate/incorrect. an investigation was conducted to solve the issue of the occasional high platelets (plt) background counts reported while executing quality control (qc) runs. during subsequent validation testing, elevated red blood cells (rbc) background was also observed. the qc carryover counting flow sequence is the same for cbc test selection which is the most frequent test used by the customer. the increased trend of plt carryover failures during a review of instrument investigation reports for the months of (b)(6) 2008 was also observed. the investigation found a design change occurred relative to a predecessor instrument that impacted carryover. this design error allows bubbles that are generated from bubble mixing (agitation by bursts of air pressure) of the specimen solution to extend into the upper surface areas of the rbc/plt mixing chamber that were not intended to be rinsed by design. this occasionally causes plt % carryover performance to exceed specifications. the rbc % carryover is impacted as well for the same reason. it was determined that the carryover performance for plt is less than or equal to 1.7% and for rbc is less than or equal to 1.2%. carryover parameters were not challenged under the cbc test selection during validation and verification and clinical testing prior to product launch. a cell-dyn ruby product deficiency was established for the plt and rbc % carryover parameters that led to field action fa17nov2008. the communication to customers included modified specifications for carryover to the plt and rbc parameters. the plt and rbc % carryover is less or equal to 1.0% in the cell-dyn ruby system operator's manual. it was recommended to update the carryover labeling claims for plt to 1.7% and for rbc% to 1.2%. this is a final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1268844 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that hemoglobin (hgb) and hematocrit (hct) values are not matching the rule of 3 well for patient samples tested using a cell-dyn 1800 analyzer. in late 2008, one patient sample generated a hgb=7.6 g/dl and hct=20.8%. the sample was repeated the same day yielding a hgb=7.3 g/dl and hct=20.7%. the next day, the sample was repeated 10 times yielding hemoglobin values ranging from 8.3 to 8.7 g/dl and hematocrit values ranged from 24.0 to 24.7%. no adverse impact to patient management was reported due to this issue. >< (b)(4). clogged/obstructed silicon tubing (s1). investigation summary: after reviewing the data sent by the customer, the cta suggested a smear review of the patient sample for rbc morphology and rouleau/clumps. the customer stated not having time for a smear review and that the sample was going to be sent to a reference laboratory. the customer then requested field service for issue resolution. on (b)(6) 2008, the field service representative (fsr) cleaned an obstructed silicon tubing (s1), ran controls, and passed. the instrument was performing within specifications. on (b)(6) 2008, during an fsr visit, a printed circuit board assembly was replaced on the instrument, ran controls, and qc passed. instrument was performing per specifications and no further investigation was required. the issue is addressed in labeling in the cell-dyn 1800 system operator's manual. a review of complaint) reports did not indicate any adverse trend for the cell-dyn 1800, l/n 07h77-01, for the reported complaint issue. based on the investigation, no product issue was identified for the cell-dyn 1800 for the reported issue. there was no systemic issue for the cell-dyn 1800 product line. this is the final report. | 0
1271183 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). printed circuit board ac sense module assembly. the field service engineer determined that the instrument issue was due to a defective printed circuit board (pcb) ac sense module assembly and the part was returned to abbott for further investigation. the investigation into the returned pcb ac sense module assembly showed a blown u6 voltage regulator. the root cause of the pcb ac sense module assembly failure cannot be determined. a review of pcb ac sense module assembly data for the last 12 months indicates this part has a 99.97% reliability worldwide. a review of the service history on the cell-dyn 1800, (b)(4), showed that the instrument was installed at the account on (b)(6), 2004. a total of 24 tickets have been reported throughout the instrument history, out of which four (4) complaints were for power surge/outage; one (1) required a power supply replacement ((b)(6), 2007). there has been one (1) preventive maintenance ((b)(6), 2007) performed on the instrument since installation. the cell-dyn 1800 system operator's manual, list number 07h80-01, revision e, under section 10, troubleshooting and diagnostics, provides information to assist in problem identification, isolation, and corrective actions. section 10, notes that in an emergency situation, to turn off the power switches in any order, as quickly as possible. a non-statistical trend (nst) complaint review was performed for the reported issue from (b)(6), 2009. the review showed zero (0) complaints during the searched period. no nst was identified. based on the investigation of this complaint, no product deficiency was identified on the cell-dyn 1800 instrument for the reported issue. there was no systemic issue for the cell-dyn 1800 product line. this is the final report. >< the customer stated after hearing a loud 'gunshot' noise from the cell-dyn 1800 analyzer, smoke was emitted. the customer unplugged the analyzer and field service was dispatched to inspect and service the instrument. no injuries were reported. | 0
1280903 | (b)(4). evaluation (b)(4): other, dirty aperture plates and worn silicone tubing (s2). discrepant white blood cell results were generated for a patient sample with cell-dyn 1800 analyzer, (b)(4). in response to this issue an investigation was initiated which included a review of the complaint text, a search for similar complaints, and a review of labeling. a field service representative (fsr) was dispatched to the customer location to service the cell-dyn 1800 analyzer. the fsr replaced worn silicone tubing (s2) and the sample syringe. additionally, the fsr found the aperture plates were dirty and performed a cleaning procedure. the fsr ran precision and controls, and confirmed the instrument was performing within specifications. a review of the complaint trending systems for the period (b)(4) 2007 through (b)(4) 2008, did not indicate any adverse trend for the cell-dyn 1800, list number 7h77-01, for the reported issue. the cell-dyn 1800 system operator's manual, revision e, under section 10, troubleshooting and diagnostics, pages 10-13, provides information to assist in problem identification, isolation, and corrective actions. section 9, preventive maintenance schedule, page 9-3 indicates that replacement/cleaning of syringes or aperture places are as required procedures. section 9, as-required maintenance, page 9-43, provides cleaning instructions for the aperture plates. section 9, as-required maintenance, page 9-69, provides replacement instructions for syringes. based on the investigation, no product issue was identified for the cell-dyn 1800, (b)(4). this is the final report. >< a patient sample generated a wbc result of 41.0 k/ul on the cell-dyn 1800 analyzer. the sample was repeated three times, generating the same results. the wbc result was reported and questioned by the physician. another sample was obtained the following day and yielded a wbc result of 6.4 k/ul. the original sample was retested and also generated a wbc result of 6.4 k/ul. field service was dispatched to inspect and service the instrument. no impact to patient management was reported. >< (b)(4). incorrect explanation of evaluation code of other given on previous follow-up report. evaluation (b)(4): other, dirty aperture plates and worn silicone tubing (s2) on the cell-dyn 1800 analyzer. this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1305023 | (b)(4). a field service representative (fsr) visit to the customer's site found a pinch silicon tubing (s2) that was worn out from normal use. the silicon tubing (s2) was replaced by the fsr, which resolved the issue seen by the customer. the instrument was performing within specifications. no further investigation was required. the cell-dyn 1800 system operator's manual, list number 07h80-01, troubleshooting and diagnostics, provides information to assist in problem identification, isolation, and corrective actions. instructions for obtaining technical assistance are included as well. a review of complaints for the period (b)(6) 2008 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn 1800, l/n 07h77-01, for low (b)(6) results on patient samples. based on the investigation, no product issue was identified for the cell-dyn 1800 for low (b)(6) patient results. there is no systematic issue with the cell-dyn 1800 product line. this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated the cell-dyn 1800 analyzer has generated intermittent low hemoglobin results. a patient sample yielded an initial hemoglobin result of 6.9 g/dl and upon repeat, the value was 11.4 g/dl. no suspect results were reported and there was no impact to patient management. | 0
1305053 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated the cell-dyn 1800 analyzer has generated intermittent low hemoglobin results. a patient sample yielded an initial hemoglobin result of 6.0 g/dl and upon repeat, the value was 10.6 g/dl. no suspect results were reported and there was no impact to patient management. >< (b)(4). a field service representative (fsr) visit to the customer's site found a pinch silicon tubing (s2) that was worn out from normal use. the silicon tubing (s2) was replaced by the fsr, which resolved the issue seen by the customer. the instrument was performing within specifications. no further investigation was required. the cell-dyn 1800 system operator's manual, list number 07h80-01, revision e, (b)(4), troubleshooting and diagnostics, (b)(4), provides information to assist in problem identification, isolation, and corrective actions. instructions for obtaining technical assistance are included as well. a review of complaints for the period (b)(6) 2008 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn 1800, l/n 07h77-01, for low hgb results on patient samples. based on the investigation, no product issue was identified for the cell-dyn 1800 for low hgb patient results. there is no systematic issue with the cell-dyn 1800 product line. this is the final report. | 0
1305064 | (b)(4). a field service representative (fsr) visit to the customer's site found a pinch silicon tubing (s2) that was worn out from normal use. the silicon tubing (s2) was replaced by the fsr, which resolved the issue seen by the customer. the instrument was performing within specifications. no further investigation was required. the cell-dyn 1800 system operator's manual, list number 07h80-01, revision e, (b)(4), troubleshooting and diagnostics, (b)(4), provides information to assist in problem identification, isolation, and corrective actions. instructions for obtaining technical assistance are included as well. a review of complaints for the period (b)(6) 2008 through (b)(6) 2008, did not indicate any adverse trend for the cell-dyn 1800, l/n 07h77-01, for low hgb results on patient samples. based on the investigation, no product issue was identified for the cell-dyn 1800 for low hgb patient results. there is no systematic issue with the cell-dyn 1800 product line. this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated the cell-dyn 1800 analyzer has generated intermittent low hemoglobin results. a patient sample yielded an initial hemoglobin result of 7.4 g/dl and upon repeat, the value was 10.0 g/dl. no suspect results were reported, and there was no impact to patient management. | 0
1366469 | the customer states that a physician questioned a patient's hemoglobin assay result generated with the cell-dyn 1700 analyzer. the initial result of 7.0 g/dl became 13.7 g/dl upon repeat testing. all controls have read within specifications with no background issues. the customer is also reporting numerous issues with this analyzer that include flow/clog errors and instrument lock-ups. the abbott customer technical advocate advised the customer not to report patient results from this analyzer until the issues are resolved. a service call has been initiated. there is no impact to patient management reported. >< an abbott field service representative (fsr) arrived at the customer site and flushed the hemoglobin flowcell of the cell-dyn 1700 analyzer. the fsr verified the instrument's operation. subsequent instrument operations and test results were acceptable.a review of the complaint history for the cell-dyn 1700 analyzer, since the complaint was initially opened to the conclusion of this investigation found no other complaints related to the current issue. a review of the customer's data found that the low hemoglobin results had dispersional data alerts (l flags). the cell-dyn 1700 system operator's manual (list number 3h58-01, rev f) contains information to address this issue and to address basic troubleshooting procedures for this analyzer. a review of complaints from the timeframe january 2008 through march 2009 did not indicate any adverse trend for the cell-dyn 1700 system. based on the current investigation, no product issue was identified for the cell-dyn 1700 system, list number 3h53-01. there is no systemic issue with the cell-dyn 1700 product line. no malfunction was identified. this is a final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1366471 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. an investigation is in process. >< (b)(4). an abbott field service representative (fsr) arrived at the customer site and flushed the hemoglobin flowcell of the cell-dyn 1700 analyzer. the fsr verified the instrument's operation. subsequent instrument operations and test results were acceptable. a review of the complaint history for the cell-dyn 1700 analyzer, (b)(4), since the complaint was initially opened to the conclusion of this investigation found no other complaints related to the current issue. a review of the customer's data found that the low hemoglobin results had dispersional data alerts (l flags). the cell-dyn 1700 system operator's manual (list number 3h58-01, rev f) contains information to address this issue and to address basic troubleshooting procedures for this analyzer. a review of complaints from the timeframe (b)(6) 2008 through (b)(6) 2009 did not indicate any adverse trend for the cell-dyn 1700 system. based on the current investigation, no product issue was identified for the cell-dyn 1700 system, list number 3h53-01. there is no systemic issue with the cell-dyn 1700 product line. no malfunction was identified. this is a final report. >< the customer states that a physician questioned a patient's hemoglobin assay result generated with the cell-dyn 1700 analyzer. the initial result of 6.0 g/dl became 9.0 g/dl upon repeat testing. all controls have read within specifications with no background issues. the customer is also reporting numerous issues with this analyzer that include flow/clog errors and instrument lock-ups. the abbott customer technical advocate advised the customer not to report patient results from this analyzer until the issues are resolved. a service call has been initiated. there is no impact to patient management reported. | 0
1370513 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). the initial medwatch had an incorrect date of (b)(4) 2009. the corrected date is (b)(4) 2009. evaluation (b)(4): other, sensor, photo interrupt, (b)(4). in response to this issue a field service representative (fsr) was sent to inspect the instrument. the fsr replaced the low tube height sensor [sensor, photo interrupt, (b)(4)] which was the cause of the tube misidentification and short sampling erratic results. as a precaution, the fsr also replaced the shm 1 and cleaned the closed mode needle. the fsr performed a successful open/closed mode precision, confirming resolution of the issue. a review of the complaint history for the cell-dyn 3200 (b)(4) from the complaint date ((b)(4) 2009) until present found no other complaint related to the reported issue. additionally, a review of the complaint trending system was conducted and found no adverse trends. based on the investigation, it has been determined that the cell-dyn 3200, list number 04h60-01, is performing as intended. no product deficiency was identified. this is the final report. >< the account stated that the cell-dyn 3200 analyzer is generating imprecise results which are not reproduced when retested. the account provided an example of a patient who's hgb result was repeatedly 6.0 g/dl. the patient was then sent to the hospital but was released as the hgb and hct results were 'normal' (results not provided). there was no impact to patient management. | 0
1380368 | (other), incorrect fuse was installed in the analyzer. evaluation: (other) codes were selected as no method, results, or conclusions can be made at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< an incorrect fuse was found installed in the cell-dyn analyzer at the customer site, resulting in no response to the power supply in the analyzer. no impact to patient management nor to user safety was reported. >< (b)(4). the investigation of the cell-dyn 3700 analyzer smoke and fire issue identified the root cause as the installation of an incorrect fuse (by either the customer or the field service representative) in the analyzer power supply when it is operated at 220/240 vac (volts alternating current), which may result in the possibility of smoke and fire. the cell-dyn 3700 operator's manual contains the electrical requirements and specific voltage settings for each system. there are also controls in place to reduce the risk of smoke and fire (over-current protection, fusing, safety icons and hazard warning labels). the following steps were put in place to correct this issue: a product information letter was issued on 20 march 2009 to emphasize to new customers to follow the instructions provided in the cell-dyn 3700 system operator's manual for fuse replacement. a field communication (product correction) was issued on 27 mar 2009 to all affected customers to ensure that the correct fuse was installed in the cell-dyn 3700 analyzers in the field. manufacturing instructions were changed to remove the fuse prior to shipping of the product to reduce the risk of an incorrect fuse being installed. instructions were provided to the field service personnel through an installation check list to verify that the correct fuse was installed during installation. furthermore, the preventive maintenance procedure was changed to include a check that the correct fuse was installed. | 0
1380369 | (b)(4). the investigation of the cell-dyn 3700 analyzer smoke and fire issue identified the root cause as the installation of an incorrect fuse (by either the customer or the field service representative) in the analyzer power supply when it is operated at 220/240 vac (volts alternating current), which may result in the possibility of smoke and fire. the cell-dyn 3700 operator's manual contains the electrical requirements and specific voltage settings for each system. there are also controls in place to reduce the risk of smoke and fire (over-current protection, fusing, safety icons and hazard warning labels). the following steps were put in place to correct this issue: a product information letter was issued on 20 march 2009 to emphasize to new customers to follow the instructions provided in the cell-dyn 3700 system operator's manual for fuse replacement. a field communication (product correction) was issued on 27 mar 2009 to all affected customers to ensure that the correct fuse was installed in the cell-dyn 3700 analyzers in the field. manufacturing instructions were changed to remove the fuse prior to shipping of the product to reduce the risk of an incorrect fuse being installed. instructions were provided to the field service personnel through an installation check list to verify that the correct fuse was installed during installation. furthermore, the preventive maintenance procedure was changed to include a check that the correct fuse was installed. >< incorrect fuse was installed in the analyzer. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account found an incorrect fuse (8 ampere) installed in the cell-dyn analyzer. no impact to patient management nor to user safety was reported. | 0
1392719 | (b)(4). evaluation: sample aspiration peristaltic pump tubing (ppt), medium. customer replaced the sample aspiration ppt and cleaned the rbc/plt aperture plate and the issue was resolved. on (b)(6) 2009, (b)(6) hospital from (b)(6) reported poor reproducibility with platelets (plt), red blood cell (rbc) and white blood cell (wbc) for the past two or three weeks. the results failed the delta checks and customer retested the samples. customer did maintenance on (b)(6) 2009 and also changed the peristaltic pump tubing (ppt) both medium and small on (b)(6) 2009. poor reproducibility issue is seen in both open and closed modes. customer changed the sample aspiration ppt (medium) once a month and the last change was on (b)(6) 2009. customer technical advocate (cta) suggested the customer to change the sample aspiration ppt and monitor the results. customer reported reproducibility improved after changing the sample aspiration ppt but customer wants to improve more on plt. the cta suggested customer to clean the rbc/plt aperture plate. on (b)(6) 2009, customer confirmed with cta that the issue was resolved by changing the sample aspiration ppt (medium) and cleaning the aperture plate. cta recommended the customer to monitor the results and change the ppt once a week and clean the aperture plate regularly. a review of the report found dispersional data alerts on multiple parameters (results were underlined on the printout). other wbc and plt suspect parameter flags were also present: (b)(6). the cell-dyn 3700 system operators manual, list number 06h89-01, revision f, under chapter 3, principles of operation, provides explanation and suggested actions for suspect parameter flags. also page 3-40 of the operator manual states: a result that falls outside a laboratory action limit can also indicate the need for the operator to follow a laboratory protocol, such as repeating the sample, notifying the physician or performing a smear review. in cases where a cellular abnormality is present that alters cellular morphology to the point that the cells do not fit the criteria used by the instrument to generate a flag, dispersional data alerts may be the only flag's) that will alert the operator to a potentially erroneous result. chapter 9, maintenance, states the sample aspiration ppt should be replaced weekly and the rbc/plt aperture plate should be cleaned as needed. instructions were also provided in this chapter. chapter 10, troubleshooting, pages 10-55 to 10-61, provides troubleshooting steps for imprecise or inaccurate data. if assistance is needed for any technical or operational problems, the customer is instructed to contact the local country service and support center. based on the investigation, no product issue was identified for the cell-dyn 3700, list number 02h31-01, for the reported issue. there is no systematic issue with the cell-dyn 3700 product line. this is the final report. >< the account stated that the cell-dyn 3700 analyzer has poor reproducibility issues in both the open and closed mode system. on one patient sample, the initial wbc result was 1500 and the retest result was 300. the account did not state if flags were generated. there was no impact to patient management reported. >< an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1400442 | suspect medical device, additional medical devices: cell-dyn 1700 analyzer; cell-dyn 1700cs analyzer. concomitant medical product: cell-dyn 1700/1800 system accessory kit. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< use of the current available version of the 10ml diluent syringe, part number 9212368 (without a spacer), on a cell-dyn 1700, cell-dyn 1700cs, or cell-dyn 1800 with a die-cast syringe drive may result in diluent leakage during initialization of the analyzer, resulting in diluent empty faults or platelet backgrounds out of specification. the instrument may require a field service repair potentially delaying patient's results. a product correction letter has been sent to all current, active cell-dyn 1700, 1700cs and 1800 customers. the letter provides instructions for verifying the type of syringe drive installed on the analyzer and which 10 ml diluent syringe is required. a product correction has been issued and reported under 21cfr806 to the fda. no adverse impact to patient management has been reported related to this issue. >< (b)(4). suspect medical device, additional medical devices: cell-dyn 1700 analyzer, list numbers 3h53-01 and 3h53-03; cell-dyn 1700cs analyzer, list numbers 3h57-01 and 3h57-03. new die-cast drives were designed and released for usage on the cell-dyn 1800 and cell-dyn 1700 instruments in the year 2005. the 10 ml diluent syringe, list number 04h36-01, was incorrectly listed as the replacement syringe with the new die-cast drives. customers were unable to order the proper syringe with the thicker spacer to prevent the tolerance stack-up issue as a list number had not established for the syringe. effective 29 may 2009, the 10 ml diluent syringe drive with the 0.25-inch thick washer was released under list number 09h35-01 for distribution as a drop-in replacement for the syringe and a product information letter was issued to inform new customers of the correct part number to order. a field communication was issued on 08 june 2009 to all affected customer to provide information regarding the issue and the correct syringe part number to order. the investigation concluded that the issue occurred due to an open loop process, which did not include adequate checks and balances to ensure items identified on a change impact assessment form were addressed via an engineering change order. the current process has been revised to prevent this issue from recurring. this is the final report. | 0
1400454 | the customer states that a patient sample processed using a cell-dyn 1800 analyzer generated results that were questioned by a physician. the sample was repeated twice, yielding different results. it was noticed that during the second run, blood had dripped from the sampling probe. the results of the first and third runs were considered correct. no known impact to patient care/treatment was reported due to this issue. service was dispatched. initial run: hgb=10.7 g/dl, rbc=3.75 m/ul; second run: hgb=6.9 g/dl, rbc=2.1m/ul; third run: hgb=10.7 g/dl, rbc=3.64 m/ul. >< investigation in process, no method, results or conclusion code can be chosen at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4): material deterioration from normal wear. investigation summary: the customer technical advocate (cta) had the customer run background count; all were zero (0), except for plt = 2.0 k/ul (specification of </= 10.0). the customer stated that autoclean is performed each week. the cta suggested running autoclean and observed the probe for drip. if no drip was present, the customer was instructed to run background count and run quality controls (qc). the cta instructed the customer to call back if issue was unresolved and requested a fax of the patient runs. during a follow-up with the customer, the cta was informed that autoclean was performed. the instrument generated a clog error upon running background counts and the customer cleared the orifice to resolve the issue. the customer was still running qc. the customer stated that the probe was no longer leaking. on (b)(4) 2009, the cta followed-up with the customer to request patient results from the hospital. the customer stated not having the results from the lab yet. the customer also stated having vacuum limit time-out prior to issue and again on (b)(6) 2009. a field service representative (fsr) was dispatched to resolve the issue. the issue was resolved by replacement of the 3.9 solenoid valve assembly, which had worn out from normal use. the instrument was within specification and no further investigation was required. the patient was not sent for recollection and specimen was not sent to another lab for repeat run. the physician determined that the first and last result were correct and followed patient history. the second result was not correct due to blood in the probe and then leaking out during the run. the likely cause of the probe leakage was due to vacuum limit time-out issue, which was resolved by an fsr visit. no further investigation required; instrument is within specifications. the issue is addressed in product labeling and indicates to contact abbott diagnostics customer service if a vacuum limit time-out error message is encountered by the customer. a non-statistical trend (nst) review was performed and no nst was identified. based on the investigation, no product issue was identified for the cell-dyn 1800 related to reported issue. this is the final report. | 0
1413301 | (b)(4). investigation summary: the issue began on (b)(6) 2009 with a high-pitched nose; during start-up hgb background result was 0.8 (specification of less than or equal to 0.1). the customer auto cleaned, drained and filled baths, and repeated backgrounds, which returned within specifications. the customer ran quality controls (qc), which were within range on all parameters, then began running patient samples. the sample was from a venipuncture collection on a known patient, which generated hgb results outside of patient's clinical history; this is a transfusion dependent patient undergoing transfusion therapy. the result had gran r3 flagging, plt uri flagging, and mpv was not turned on. this result was reported to the physician. the patient was sent to another lab to have blood recollected and ran on a cell-dyn 1800. the result generated lym ro and gran r3 flagging. the customer repeated the initial sample ran on the cell-dyn 1700 instrument and results closely agreed with the original results obtained from this sample. the hgb result of 8.1 g/dl closely agreed with patient history. the patient was transfused, as expected per current therapy. no adverse impact to patient management was reported by the customer. the customer decided to take the cell-dyn 1700 instrument out of service after the discrepant patient results and unusual noise and requested a field service visit for issue resolution. on (b)(6) 2009 a field service representative (fsr) found hgb imprecision, bad bubble mix tubing and sticking solenoid wbc bubble mix tubing, a bent sample aspiration probe, and a bad sample syringe. the fsr replaced the s1 and s2 silicon tubing, and the sample aspiration probe. the fsr cleaned the instrument and ran precision, which passed. the instrument was performing within specifications and no further investigation was required. the issue is addressed in product labeling. a non-statistical trend (nst) review was performed and no nst was identified. based on the investigation, no product issue was identified for the cell-dyn 1700 for the reported event. this is the final report. >< the customer states that they heard a high-pitched noise and clunking sound from behind the flow panel of the cd 1700 analyzer. the issue began in 2009 during start-up of the analyzer. backgrounds were out of specification (high) initially but returned to acceptable limits after repeating background procedures. qc was within range on all parameters. a sample from a known patient was processed on the cd 1700 analyzer generating a hemoglobin result that were unexpected for this patient's clinical history. this patient had been undergoing transfusion therapy. a hgb=11.7 g/dl and hct=25.0% were generated, reported to a physician and questioned. a new sample from this patient was collected at a different lab using a cd 1800 analyzer yielding a hgb=8.1 g/dl and hct=25.7%. the hemoglobin result of 8.1 g/dl closely agreed with the patient's history. the patient was transfused as expected per current therapy. no adverse impact to patient management was reported due to this issue. >< investigation in process, no method, results or conclusion can be chosen at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1419031 | (b)(4). new die-cast drives were designed and released for usage on the cell-dyn 1800 and cell-dyn 1700 instruments in the year 2005. the 10 ml diluent syringe, list number 04h36-01, was incorrectly listed as the replacement syringe with the new die-cast drives. customers were unable to order the proper syringe with the thicker spacer to prevent the tolerance stack-up issue as a list number had not established for the syringe. effective 29 may 2009, the 10 ml diluent syringe drive with the 0.25-inch thick washer was released under list number 09h35-01 for distribution as a drop-in replacement for the syringe and a product information letter was issued to inform new customers of the correct part number to order. a field communication was issued on 08 june 2009 to all affected customer to provide information regarding the issue and the correct syringe part number to order. the investigation concluded that the issue occurred due to an open loop process, which did not include adequate checks and balances to ensure items identified on a change impact assessment form were addressed via an engineering change order. the current process has been revised to prevent this issue from recurring. this is the final report. >< the customer received the product correction letter for the diluent syringe on the cell-dyn 1800 analyzer and discovered that the incorrect syringe was installed. the customer service advocate sent the customer the correct syringe. no impact to patient management was reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1419034 | the customer was questioning which diluent syringe should be installed on the cell-dyn 1800 analyzer. the customer technical advocate (cta) explained the product correction letter for the cell-dyn diluent syringe to the customer and determined the customer had the incorrect syringe installed. the cta sent the correct syringe to the customer. a follow-up with the customer confirmed the syringe had been received and installed successfully, resolving the issue. no impact to patient management was reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). new die-cast drives were designed and released for usage on the cell-dyn 1800 and cell-dyn 1700 instruments in the year 2005. the 10 ml diluent syringe, list number 04h36-01, was incorrectly listed as the replacement syringe with the new die-cast drives. customers were unable to order the proper syringe with the thicker spacer to prevent the tolerance stack-up issue as a list number had not established for the syringe. effective 29 may 2009, the 10 ml diluent syringe drive with the 0.25-inch thick washer was released under list number 09h35-01 for distribution as a drop-in replacement for the syringe and a product information letter was issued to inform new customers of the correct part number to order. a field communication was issued on 08 june 2009 to all affected customer to provide information regarding the issue and the correct syringe part number to order. the investigation concluded that the issue occurred due to an open loop process, which did not include adequate checks and balances to ensure items identified on a change impact assessment form were addressed via an engineering change order. the current process has been revised to prevent this issue from recurring. this is the final report. | 0
1419043 | the customer indicated that after receiving the correct diluent syringe, they felt it wasn't necessary to replace the syringe due to the absence of any issues. the customer technical advocate encouraged the customer to replace the syringe as soon as possible. no impact to patient management was reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). new die-cast drives were designed and released for usage on the cell-dyn 1800 and cell-dyn 1700 instruments in the year 2005. the 10 ml diluent syringe, list number 04h36-01, was incorrectly listed as the replacement syringe with the new die-cast drives. customers were unable to order the proper syringe with the thicker spacer to prevent the tolerance stack-up issue as a list number had not established for the syringe. effective 29 may 2009, the 10 ml diluent syringe drive with the 0.25-inch thick washer was released under list number 09h35-01 for distribution as a drop-in replacement for the syringe and a product information letter was issued to inform new customers of the correct part number to order. a field communication was issued on 08 june 2009 to all affected customer to provide information regarding the issue and the correct syringe part number to order. the investigation concluded that the issue occurred due to an open loop process, which did not include adequate checks and balances to ensure items identified on a change impact assessment form were addressed via an engineering change order. the current process has been revised to prevent this issue from recurring. this is the final report. | 0
1470717 | the cd 4000 analyzer was moved recently to another location, and it appeared that the connection of the tubing for the waste fluid was loose. the tubing subsequently became disconnected and a customer that was not the operator but standing nearby was splashed with the fluid. fluid got into the person's eyes. the person washed his/her eyes immediately. the person is to follow-up with a visit to the ophthalmologist and also testing for infectious diseases. the person was not wearing protective eye wear at the time of the event. to date, no additional information is available regarding patient's status or possible injury or treatment resulting from exposure. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). investigation summary: the customer reported that when moving the cell-dyn 4000, the connection of the waste tubing became loose and splashed an inspector that was standing nearby with waste fluid in the eyes. the inspector washed the eyes immediately after the incident. the inspector was sent for diagnosis at ophthalmology; ophthalmology results showed no problem. there was a high possibility that the waste fluid contained infectious substances; therefore, the individual was tested for infectious diseases. the testing came back negative, but the individual will be retested in 2 to 3 months. customer service and support arranged for the cell-dyn 4000 instrument to be checked and/or fixed the arrangement of the waste tubing. product labeling was reviewed and found to contain instructions for basic troubleshooting and hazardous material warnings and information. it also warns that waste in the waste tubing can be infectious and to wear appropriate personal protection equipment and follow biosafety practices as specified in the osha bloodborne pathogen rule (29 cfr part 1910.1030)1 or other equivalent biosafety procedures. a non-statistical trend (nst) review was performed and no nst was identified. based on the investigation, no product issue was identified for the cell-dyn 4000 for the reported event. this is the final report. | 0
1479556 | the customer stated a cell-dyn 1800 analyzer generated falsely elevated platelet results. the customer stated the analyzer generated platelet results of one million while the patient's correct result was 200 k/ul. the customer stated the analyzer did not generate any errors or flags, and quality controls were in range. the incorrect patient results were not reported, and no treatment was given to the patient. the customer declined to provide instrument data, specific patient data, or information on how the correct platelet value was obtained. >< (b)(4). evaluation results: spm printed circuit board assembly. in response to this issue an investigation was initiated to further investigate the customer issue. the investigation included a review of the complaint text, a service visit, a search for similar complaints, and a review of labeling. the field service representative (fsr) found the signal processor module (spm) for the rbc edit ratio (rer) and plt out of specification. the fsr replaced the spm printed circuit board assembly and performed gain adjustment for rer and plt. the instrument was then performing within specifications. tracking and trending did not indicate any adverse complaint trend exists for discrepant platelet results on the cell-dyn 1800. labeling was reviewed and found to be adequate. based on the investigation, no product issue was identified for the cell-dyn 1800 related to discrepant platelet results. >< (b)(4). investigation in process, no method, results or conclusion code can be chosen at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1496536 | (b)(4). evaluation: calibration required. in response to this issue an investigation was initiated to further investigate the customer issue. the investigation included a review of the complaint text, a search for similar complaints, and a review of labeling. the customer technical advocate (cta) suggested to the customer to perform a 4 x 4 repeatability test in open mode and closed mode. the coefficients of variation were obtained and adjusted on the analyzer. the issue was resolved by performing calibration, as recommended by the cta. the instrument was performing within specifications. tracking and trending did not indicate any adverse trend for the cell-dyn 3200 for the reported complaint issue. labeling was reviewed and found to be adequate. based on the investigation, no product issue was identified for the cell-dyn 3200 for discrepant hemoglobin results. >< (b) (4) - investigation in process, no results or conclusion code can be chosen at this time.this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated a cell-dyn 3200 analyzer generated discrepant hemoglobin results for one patient sample. the analyzer initially generated a hemoglobin result of 7.76 g/dl for the patient sample. upon repeat, the hemoglobin value generated was 11.6 g/dl. there was no adverse impact to patient management reported. | 0
1549103 | (b)(4). evaluation: wbc transducer and 0.1 ml sample syringe (worn out from normal use). in response to this issue an investigation was initiated to further investigate the customer issue. the investigation included a review of the complaint text, dispatch of a field service representative (fsr), a search for similar complaints, and a review of labeling. the fsr was dispatched to inspect and service the instrument. the fsr replaced the sample wbc transducer and verified gains. the fsr also replaced the sample syringe which was worn out to get precision to pass. the fsr performed verification procedure vp-09 (signal processor module verification/adjustment) and vp-11 (stepper motor power test and verification). the fsr ran precision and quality controls; all passed. the instrument was performing within specifications. tracking and trending did not indicate any adverse trend for the cell-dyn 1800 for the reported complaint issue. labeling was reviewed and found to be adequate. based on the investigation, no product issue was identified for the cell-dyn 1800 for discrepant wbc results. >< the customer stated a cell-dyn 1800 analyzer generated a falsely elevated wbc value of 74.0 k/ul. the incorrect value was reported outside the laboratory. the laboratory does not review patient slides. there was no adverse impact to patient management reported. | 0
1023857 | other: (investigation is in process). this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account generated a nonreactive axsym hcv v3.0 result on a pt who was known to be hcv positive. in 2007, the pt tested axsym hcv v3.0 nonreactive (0.35 s/co). the nonreactive result was questioned by the physician because the pt known to be hcv positive. the specimen was repeated with an axsym hcv v3.0 reactive result (67.87 s/co). no error messages were present in the axsym analyzer message log. the account questioned why no error message was generated by the axsym analyzer and replaced the axsym analyzer with an axsym plus analyzer. in 2008, abbott was informed of a user report to affsaps from the account. the user report stated no error message occurred during pipetting on the axsym analyzer which lead to an erroneous axsym hcv 3.0 result. the account stated in addition to axsym hcv 3.0, the specimen was also tested for hepatitis b and hiv. a sample alarm occurred on the axsym analyzer for the testing for the same pt specimen. the axsym testing and other serologies generated negative results. the nonreactive axsym hcv 3.0 result was not reported outside of the lab. no impact to pt management was reported. end of report. | 0
1046588 | to troubleshoot the problem on the axsym analyzer, solution 1 and 3 volume checks were performed and were within specifications. the meia lamp was replaced and the meia optics were calibrated. the matrix cell carousel was cleaned. a precision run of axsym ca125 medium controls yielded a % cv of 4.9%. the axsym analyzer was functioning with specifications. the most likely cause was a meia lamp failure. a review of the complaint history for abbott axsym (ln07a83-85) was performed for the time period of four months in 2008. no trends were identified. the abbott axsym system operations manual volume-269-2949/r5-march 1999 on pages (r3/r5) 10-261 and (r4/r5) 10-259 under the section troubleshooting and diagnostics section 10 observed problems: meia test results to probable cause(s): meia optical assembly malfunction. corrective action: perform an meia verification. if the verification fails, lamp replacement is listed under possible corrective actions. this is a final report. >< the account noticed a reproducibility problem with the axsym analyzer and several assays, ca125, toxo g, toxo m and hbsag. for example, one pt tested axsym ca125 = 3.37 u/ml but repeated axsym ca125 = 8.97 u/ml. the pt was previously known to test ca125 = 8 (no units of measurement provided). service was requested for the axsym analyzer. no impact to pt management was reported. | 0
1267563 | the account generated elevated axsym insulin results with lot 65003ju00 for 3 days, but when the same specimens were repeated with reagent lot 66012ju01 the results were lower. the specimens were repeated again with lot 65003ju00 and again the results were lower. due to the difference in reported values and repeated values, an amended report using the repeat values of the same lot was sent for some of the patient specimens. no patient information was provided for the specimens. no impact to patient mangement was reported. service discussed possible reagent contamination. service inspected pumps, tubing, fittings and lubricated sample/processing syringes and arm assembly. the matrix and processing carousel and v-wheels were cleaned and sample/processing probe calibrations performed. service performed meia station initialization, fpia verifications and performed fluidics which passed. the account ran quality control and precision for insulin and ferritin with passing results on the axsym analyzer. this report is for patient 12 of 24 total patients. axsym insulin data (pmol/l): patient 12: 181.9 (lot 65003jn00, reported result), 103.6 (lot 65003jn00, repeat, amended report), 117.4 (lot 66012ju01) >< (b)(4). both the axsym analyzer and axsym insulin assay were investigated. the conclusion of both investigations are provided below. axsym insulin assay investigation: no customer returned sample and/or reagent were available for investigation. the assay performance study was performed with two file axsym insulin reagent lots 65003ju00 and 66012ju00 using the file axsym insulin calibrator lot 70009ju00 and the file axsym insulin control lot 70010ju00 on one axsym. after a calibration curve is accepted on one axsym, four quality control (qc) releasing test panel members were testing in replicate of three. as a result of the assay performance study, no error occurred with the assay performance study and all panel values on the study were within the specification of qc releasing testing. additionally, the value of the panel in two lots was comparable. the tracking and trending review showed the issue of the complaint record was not identified. also, for tracking and trending review no requiring further investigation was found on the product and the complaint lot. the cause of the complaint was not identified because no error occurred with the assay performance study and all panel values on the study were within the specification of the quality releasing testing. additionally, the value of the panel in two lots were comparable. as with the other investigation results, the complaint lot was not identified with any assay performance issue. based on this investigation, the complaint lot was not identified with any assay performance issue. axsym analyzer investigation: the investigation of the issue consisted of a review of customer complaints, current axsym labeling, and the information provided including the complaint text. the axsym operations manual provides multiple probable causes and corrective actions to assist the account with resolving inconsistent/ erratic/ discrepant results. the probable causes listed include the use of malfunction of the sampling and processing syringe, valve, probe, and probe link tubing assemblies, and incorrect positioning of the sampling or processing probe. the field service representative (fsr) performed service on the axsym to verify the performance of the analyzer and prevent additional occurrences of inconsistent result generation. the fsr inspected and lubricated the sample and processing syringes, and the sample and processing arm assemblies. the fsr inspected and cleaned the matrix and processing carousel and the v-wheels. the fsr performed successful sample and processing probe calibrations, meia initialization, fpia verification, and fluidics check. please also note the insulin package insert notes samples should be tested as soon as possible after drawing, because samples may show lower insulin levels because of insulin degrading enzyme existing in the red blood cell of the samples, and the recommended temperature for storing specimens is -20 degrees celsius. tracking and trending did not identify any adverse trends. the most probable cause of the inconsistent insulin patient results was the use of a malfunctioning syringe assembly. the sample results may also have been affected by the sample storage temperature of 4 - 8 degrees in use at the account site. the actual cause of the suspect insulin patient results could not be determined with the available information. based on the available information and this investigation, no deficiency was identified for the axsym analyzer related to this issue under investigation. this is a final report. >< investigation in process, no method, results or conclusion code can be chosen at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1267564 | the account generated elevated axsym insulin results with lot 65003ju00 for 3 days, but when the same specimens were repeated with reagent lot 66012ju01 the results were lower. the specimens were repeated again with lot 65003ju00 and again the results were lower. due to the difference in reported values and repeated values, an amended report using the repeat values of the same lot was sent for some of the patient specimens. no patient information was provided for the specimens. no impact to patient mangement was reported. service discussed possible reagent contamination. service inspected pumps, tubing, fittings and lubricated sample/processing syringes and arm assembly. the matrix and processing carousel and v-wheels were cleaned and sample/processing probe calibrations performed. service performed meia station initialization, fpia verifications and performed fluidics which passed. the account ran quality control and precision for insulin and ferritin with passing results on the axsym analyzer. this report is for patient 11 of 24 total patients. axsym insulin data (pmol/l):patient 11: 21.6 (lot 65003jn00, reported result), 9.2 (lot 65003jn00, repeat),14.5 (lot 66012ju01). >< investigation in process, no method, results or conclusion code can be chosen at this time. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). both the axsym analyzer and axsym insulin assay were investigated. the conclusion of both investigations are provided below. axsym insulin assay investigation: no customer returned sample and/or reagent were available for investigation. the assay performance study was performed with two file axsym insulin reagent lots 65003ju00 and 66012ju00 using the file axsym insulin calibrator lot 70009ju00 and the file axsym insulin control lot 70010ju00 on one axsym. after a calibration curve is accepted on one axsym, four quality control (qc) releasing test panel members were testing in replicate of three. as a result of the assay performance study, no error occurred with the assay performance study and all panel values on the study were within the specification of qc releasing testing. additionally, the value of the panel in two lots was comparable. the tracking and trending review showed the issue of the complaint record was not identified. also, for tracking and trending review no requiring further investigation was found on the product and the complaint lot. the cause of the complaint was not identified because no error occurred with the assay performance study and all panel values on the study were within the specification of the quality releasing testing. additionally, the value of the panel in two lots were comparable. as with the other investigation results, the complaint lot was not identified with any assay performance issue. based on this investigation, the complaint lot was not identified with any assay performance issue. axsym analyzer investigation: the investigation of the issue consisted of a review of customer complaints, current axsym labeling, and the information provided including the complaint text. the axsym operations manual provides multiple probable causes and corrective actions to assist the account with resolving inconsistent/ erratic/ discrepant results. the probable causes listed include the use of malfunction of the sampling and processing syringe, valve, probe, and probe link tubing assemblies, and incorrect positioning of the sampling or processing probe. the field service representative (fsr) performed service on the axsym to verify the performance of the analyzer and prevent additional occurrences of inconsistent result generation. the fsr inspected and lubricated the sample and processing syringes, and the sample and processing arm assemblies. the fsr inspected and cleaned the matrix and processing carousel and the v-wheels. the fsr performed successful sample and processing probe calibrations, meia initialization, fpia verification, and fluidics check. please also note the insulin package insert notes samples should be tested as soon as possible after drawing, because samples may show lower insulin levels because of insulin degrading enzyme existing in the red blood cell of the samples, and the recommended temperature for storing specimens is -20 degrees celsius. tracking and trending did not identify any adverse trends. the most probable cause of the inconsistent insulin patient results was the use of a malfunctioning syringe assembly. the sample results may also have been affected by the sample storage temperature of 4 - 8 degrees in use at the account site. the actual cause of the suspect insulin patient results could not be determined with the available information. based on the available information and this investigation, no deficiency was identified for the axsym analyzer related to this issue under investigation. this is a final report. | 0
1288883 | axsym vancomycin ii assay ln: 5b75. an abbott field service representative (fsr) arrived at the customer site and cleaned and calibrated the fpia optics and calibrated the process and sample probes and completed an instrument preventative maintenance (pm) procedure. the fsr verified the operation of the axsym plus analyzer through instrument diagnostic testing. results were as expected following field service intervention. there was no impact to patient management reported.a review of the service history of this axsym plus analyzer indicates there were no additional erratic result complaints generated following fsr cleaning of the fpia optics and calibration of the fpia optics and probes. the customer's issue is addressed in the axsym system operation manual (ln# 9a26-06) troubleshooting and diagnostics, observed problems, results erratic, discrepant, and/or experiencing imprecision. multiple probable causes and corrective actions are listed for an inconsistent results or results not as expected. the probable causes listed include fibrin, red blood cells or particulate matter present in samples and malfunction of the sampling and processing syringe, valve, meia optics malfunction, probe and probe link tubing, inadequate washing of the probe, bulk solutions not dispensing properly and platform sensor malfunctioning. corrective actions are listed in the operations manual, which also refers the operator to the assay-specific package insert for processing samples. in operational precautions and limitations, a statement is made to evaluate every result in conjunction with other clinical data, e.g., symptoms, results of other tests, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. also, the axsym vancomycin ii package insert provides adequate information on suitable specimens, results, and limitations of the procedure. a review of the complaint database for similar complaints for the time frame of 2008 to 2009 did not find any similar complaints for the current customer described issue. the root cause for the erratic results could not be determined with the available information. field service cleaned and calibrated the fpia optics and calibrated the process and sample probes and completed the instrument preventative maintenance (pm). the axsym system operations manual provides adequate information to resolve the customer described issue. a deficiency of the axsym system was not identified. this is a final report. >< the customer states that one patient sample generated an initial axsym vancomycin ii assay result of 0.00 ug/ml. the sample retested at 12.47 ug/ml, which was the expected result. controls have been within specification. no suspect results have been reported from the lab. a service call was initiated. >< axsym vancomycin ii assay list#: 5b75. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1296003 | the account stated that the axsym analyzer generated discrepant toxoplasmosis igg result on a patient sample. the initial result was positive 24 iu/ml and the retest result was negative 0.1 iu/ml. the patient is known to be negative (0.00 iu/ml). there is no impact to patient management reported. >< (b)(4). the axsym plus serial number (b)(4) generated erratic toxoplamosis (toxo) igg patient control and a patient result. there was no adverse impact to patient management due to the erratic toxo igg results. the fsr replaced the solution 1 and solution 3 pumps to resolve the erratic result issue. the fsr performed a tubing decontamination, volume check, and ran daily controls to verify the axsym performance after service was performed. the axsym system operation manual contains adequate information on the probable causes and corrective actions for inconsistent results. the probable causes include bulk solutions 1 and 3 dispensing improperly. the toxo igg package insert was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the abbott metric reports for the axsym analyzer did not identify any adverse trends due to axsym toxo igg assay inconsistent results. the axsym system operation manual (list no. 9a26-06) section 10, troubleshooting and diagnostics, observed problems, results erratic, discrepant, and/or experiencing imprecision lists multiple probable causes and corrective actions for inconsistent, discrepant, and/or erratic results. the probable causes listed include bulk solutions 1 and 3 dispensing improperly. corrective actions listed include cleaning and flushing the dispensers, and calling the axsym customer support center. the toxo igg package insert (b)(4) was reviewed and was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the package insert notes under sample collection and preparation for analysis states it is important to follow the routine maintenance procedures defined in the axsym system operations manual for optimal axsym performance. the most probable cause of the inconsistent toxo igg patient results was the malfunctioning bulk solution 1 and 3 pumps. the actual cause of the suspect toxo igg patient results could not be determined with the information available. based on the available information and this investigation, no deficiency was identified for the axsym plus list no. 07a83-82 related to the issue under investigation. this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1344206 | the account stated that they observed smoke coming from the sample syringe board. field service was requested. there were no injuries reported. >< (b)(4). complaint text indicates smoke appeared to come from the sample syringe board. the customer technical advocate (cta) instructed the customer to shut down the system and requested the field service rep (fsr) to visit the account. the fsr identified there was a loss of liquid from a connection of the sample line. the fsr removed deposits of buffer #4 from the syringe driver cable and replaced the sample and processing probes and tubings. the syringe assembly was reinstalled and the performance tested. the power supply was confirmed to be normal and all the circuit boards were working as expected. the instrument was returned to the operator for testing and the instrument is working as expected. the system was meeting all specifications following field service intervention. no injury was documented. the axsym system operations manual provides adequate information concerning maintenance procedures, running the analyzer and emergency shutdown procedures. the axsym system has been certified to the appropriate us, (b)(4) and international safety standards with the objective to ensure the design and methods of construction used provide adequate protection for the operator and surrounding area against electrical shock, burns and excessive heat. review of abbott metric reports identified no adverse trends in conjunction with the complaint issue currently under investigation. no similar complaints were found for the customer described issue. the most probable root cause for the smoke was leaking buffer onto the sample syringe board due to loose connections. field service corrected the tubing connections and returned the instrument to the operator. the operations manual instructs the operator to verify tubing connections. a deficiency of the axsym system was not identified. this is the final report. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1344289 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that there is a crack in the shroud where the wash stations screw holds the wash station to the shroud on the imx analyzer. the customer requested service for this issue. service noted that the power supply was not providing voltage to the heaters and observed smoke while testing the heaters. no fire was observed. two more power supplies had shorted out during troubleshooting noting that all wires were burnt through due to buffer leaking and shorting across the cables. there were no injuries reported due to this issue. >< (b)(4). other (b)(4) wash station thumbscrew and wash station shroud. investigation summary: the fsr replaced multiple parts to repair the (b)(4), but the power supply issue could not be resolved. the customer technical advocate (cta) sent the customer a replacement (b)(4) was returned to manufacturer. returned (b)(4) was visually examined and was found to have bent liquid heater block posts, a cracked wash station shroud, and burned wiring. the (b)(4) operation manual was found to contain adequate information on maintenance and troubleshooting for the (b)(4) wash station. the manual replacement instructions for the (b)(4) wash station include instructions to attach the thumbscrew finger tight. the manual notes wash station alignment is critical for good washing of the probe and to prevent splashing. the manual contains directions for wash station calibration, and the manual notes as required maintenance should be performed to keep the (b)(4) in good operating condition. the manual also contains information on hazards, and contains information on precautions (b)(4) operators should follow to prevent injury. under safety, the manual cautions the operator about the risk of electrical shock. a review of complaint records for the (b)(4) analyzer did not identify any adverse trends due to the issue being evaluated. a malfunction of the (b)(4) was identified. the (b)(4) sustained power supply wiring damage due to buffer leakage, because the wash station was not held stable due to the failure of worn parts and a missing thumbscrew. the actual cause of the thumbscrew and wash station shroud failure could not be determined with the available information. based on the available information and this evaluation, no deficiency was identified for the (b)(4) wash station thumbscrew (b)(4), or the (b)(4) analyzer list no. 08389-01 related to the issue under evaluation. this is the final report. | 0
1347321 | (b)(4). static discharge within the assay system may have contributed to this event. the investigation determined the causes of the increase in frequency on static discharge events were due to the architect reaction vessels (rvs) suppliers' manufacturing material and manufacturing processes. the investigation identified variables within the suppliers' molding manufacturing process that contributed to static charge variation across rv lots. other contributing factors that can generate a static charge on the rvs include low humidity, handing, shipping and creation of change within the architect instruments themselves. in addition, static charge testing for acceptance has been implemented of all incoming rv lots from the supplier. the device manufacturer and supplier have worked in collaboration to assure consistency of incoming resin material. and to improve the supplier's molding process to reduce the potential generation of static discharge. this is the final report. >< concomitant products: architect reaction vessels list no. 7c15-01 lot no. 72248p100 and 71556p100. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer states that they are questioning a falsely elevated architect troponin assay result of 0.9 ng/ml on a patient that had a previous result of 0.013 ng/ml. the sample was repeated on the same and also a different architect analyzer yielding results of 0.013 ng/ml. no adverse impact to patient was reported due to this issue. | 0
1349943 | the account stated that an initial architect troponin result of 0.12 ng/ml was generated and reported. the sample was repeated on the same analyzer with a result of <0.03 ng/ml and on a different architect analyzer with a result of <0.03 ng/ml. there was no report of impact to patient management. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). the abbott field service representative (fsr) was dispatched and found that the wash zone 1 and 2 manifolds had excessive dried buffer buildup present. the fsr cleaned the wash zone manifolds. also, the fsr replaced and calibrated the stat pipettor as a precaution. the fsr stated that the likely cause of the erratic troponin was due to excessive buffer buildup on the washzone manifold. the complaint analysis and trending system was reviewed and no adverse trends were identified the architect system operations manual was reviewed and found to adequately address the issue of erratic assay results. a probable cause listed is wash zone manifold leaking with the corresponding corrective actions instructing the user to check for salt crystals and/or liquid on or around the wash manifold valves and/or fittings and to contact the area customer support if crystals or liquid are observed. this is the final report. | 0
1350936 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that an initial architect troponin result of 0.38 ng/ml was generated and reported. the sample was repeated 3 times with results of 0.004, 0.003, 0.004 ng/ml generated. there was no adverse impact to patient management reported. >< (b)(4). instrument logs were provided by the customer and were reviewed. the logs did not encompass the date of occurrence of the issue. no static spikes were identified. there was one occurrence of error code 1007 identified. the complaint analysis and trending system was reviewed and no adverse trends were identified. the service history for the instrument was reviewed and found no additional incidents of the architect i2000 analyzer (b)(4) generating discrepant results. the architect system operations manual was reviewed and was found to provide adequate information about the troubleshooting for erratic results, operational precautions, and limitations. in the clinical chemistry stat troponin-i reagent package insert ((b)(4)), literature is provided regarding suitable specimens, results, and limitations of the procedure. based on the available information, the cause of the customer's issue could not be identified. no deficiency / malfunction was identified. this is the final report. | 0
1352705 | the account stated that the axsym analyzer is generating discrepant cea results. in 2009, the cea result for another patient was 19.0 ng/ml, the retest result was 194 ng/ml. the patient was redrawn and the result was 291 ng/ml, the retest result was 317 ng/ml. there was no impact to patient management reported. >< (b)(4). the customer notified abbott on (b)(6) 2009 of erratic (false low) total psa results were generated on axsym sn# (b)(4) from a (b)(6) male. initial results for the patient from (b)(6) 2008 and (b)(6) 2009 were 36.4 and 31.3 ng/ml respectively. on (b)(6) 2009 a result of 0 ng/ml was generated. qc on (b)(6) 2009 was verified and all qc results were within acceptable limits. the sample was repeated on (b)(6) 2009 with a result of 5.98 ng/ml generated the customer also notified abbott on (b)(6) 2009 of erratic (false low) cea results were generated on axsym sn# (b)(4) from a (b)(6) male. initial results for the patient from (b)(6) 2009 were 19.0 ng/ml. a second sample was tested on (b)(6) 2009 with a result of 291 ng/ml. review of the patient's previous results from (b)(6) 2008 to (b)(6) 2009 were 160 to 190 ng/ml. the operator verified hardware condition was satisfactory and the qc results were within acceptable limits. repeat of the (b)(6) 2009 sample generated a result of 194 ng/ml and repeat of the (b)(6) 2009 sample generated a result of 317 ng/ml. all samples tested on (b)(6) 2009 were all retested and results were as expected. field service rep (fsr) was requested to verify instrument performance. the fsr found the process probe, wash cup and dispenser #3 dirty. the fsr cleaned the process probe, wash cup, dispenser #3 and sample probe and verified instrument performance. axsym system operation manual (ln# 9a26-06) section 10, troubleshooting and diagnostics, observed problems, results erratic, discrepant, and/or experiencing imprecision. multiple probable causes and corrective actions are listed for an inconsistent results or results not as expected. the probable causes listed include fibrin, red blood cells or particulate matter present in samples and malfunction of the sampling and processing syringe, valve, meia optics malfunction, probe and probe link tubing, inadequate washing of the probe, bulk solutions not dispensing properly and platform sensor malfunctioning and sample not loaded correctly. corrective actions are listed in the operations manual, which also refers the operator to the assay-specific package insert for processing samples. in section 7, operational precautions and limitations, a statement is made to evaluate every result in conjunction with other clinical data, e.g., symptoms, results of other tests, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. section 9, service and maintenance, weekly maintenance provides instructions to clean the dispenser nozzles, probes and wash stations (cups) weekly. the axsym system operations manual provides adequate information concerning erratic/ discrepant results and weekly maintenance. the axsym total psa package insert provides literature in describing suitable specimens, results, and limitations of the procedure. failure to follow operations manual and package insert instruction may cause erroneous results. review of the abbott metric metrics identified no adverse trends in conjunction with the complaint issue currently under investigation. the exact root cause for the customer described issue could not be determined. field service cleaned multiple components including the sample and process probes, wash cup and dispenser #3 and verified instrument performance. a deficiency of the axsym system was not identified. this is the final report. >< the initial mdr was filed with both the discrepant total psa and cea results which were from two different patients. for correction, this mdr is being filed to report the discrepant cea results only, as the cea results were from another patient. this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1371871 | the customer reports smoke accompanied by the smell of burning plastic coming from the architect i2000sr analyzer. a service call was initiated. the abbott field service representative (fsr) inspected the analyzer and attributed the issue to a leaking reagent 1 (r1) syringe valve. buffer leaking from the valve accumulated in the space beneath the syringe and eventually came into contact with the radio frequency (1,2) antenna, causing the smoke and odor. there were no injuries reported or damage to the surrounding environment. there is no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. an investigation is in process. >< (b)(4). an abbott field service engineer (fse) visited the customer site and discovered that the architect i2000sr analyzer generated liquid level sense (lls) errors from the r1 and sample pipettor. the r1 syringe valve was leaking liquids onto the lls antenna. the liquid shorted the circuit board and released a burning smell that the operator detected. the burning of the circuit board became a small flame that was contained to the lls antenna. a new lls antenna, r1 syringe valve, and transfer pump were installed into the instrument. subsequent instrument operations and test results were acceptable. no injuries occurred due to this issue. the objective of the safety standards for laboratory equipment is to ensure there is no spread of fire outside the equipment in normal conditions or in single fault condition. this objective is met even though this particular failure mode caused multiple simultaneous shorts on the board. based on the available information, no product deficiency was found for this issue. the source of the liquid that contacted the antenna board was from a leaking reagent 1 syringe valve. a review of the complaint history for architect isystem (b)(4) over the period of time of (b)(6) 2008 through (b)(6) 2009 was performed. the review did not find other complaints related to this issue. the architect system operations manual ((pn 201837-104) (b)(6) , 2007) does not provide information related to smoke generated from the analyzer or components but does provide information related to potential safety and electrical hazards in section 8: hazards, electrical hazards. a malfunction of the system was identified. the investigation documents that an abbott fse replaced the rv antenna, r1 syringe valve, and transfer pump to resolve the issue. this is a final report. | 0
1372772 | the account stated that the architect i2000sr analyzer is generating discrepant total psa results on 2 patient samples. for patient #1, in 2009 the result was 47.5 u/ml, two weeks later, the result was 50.5 u/ml and on april 3, the result was 38.8 u/ml, this sample was then sent to a reference lab and the result was 18.8 u/ml. there was no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). it was stated that the total psa results were discrepant which is incorrect. the ca125 ii results were discrepant. all other information and data is correct. the complaint text for architect isystem (ln 03m74-01) s/n (b)(4) indicates that two different patients had discrepant ca 125-ii assay results using an architect i2000sr analyzer and another analyzer. the additional analyzer was at a reference laboratory and the name was not provided. however, the methodology listed is chemiluminescence. patient 1 had the following ca 125 ii assay results: date: (b)(6) 2009, ca 125 ii result: 46.8 u/ml, flag: none; (b)(6) 2009, 47.5 u/ml, exp; (b)(6) 2009, 50.5 u/ml, exp; (b)(6) 2009, 38.8 u/ml, exp. the patient sample from (b)(6) 2009 was sent to a reference lab and a result of 18.8 u/ml was obtained. patient 2 had the following result generated: date: 03/26/2009, ca 125 ii result: 59.4 u/ml flag: exp. the patient sample from (b)(6) 2009 was sent to a reference lab and a result of 34 u/ml was generated. the complaint text indicates that for patient 1, the sample with a result of 47.5 u/ml was analyzed on (b)(6) 2009. review of the customer returned results list report indicates that the sample analyzed on (b)(6) 2009 had a result of 47.5 u/ml, and the result from (b)(6) 2009 had a result of 46.8 u/ml. the customer returned 5 pages of data. page 1 is a results list report, which contains the results mentioned above. pages 2, 3 and 4 are the levey-jennings reports for the ca 125 ii assay. the only quality control (qc) results which correspond to patient sample testing dates are (b)(6) 2009. also, both qc results have a flag of exp listed. no qc results are contained for the dates of (b)(6) 2009. page 5 is the calibration curve details report for the ca 125 ii assay, and indicates the date of calibration was 01/24/2009. the customer support specialist (css) additional communication email indicates that the system logs and database are not available for return. the lot number of the ca 125 ii reagent was 66060m100, expiration date may 21, 2009. also, the css states that the exp flag next to the patient results is because the ca 125 ii reagent was expired on board, due to the reagent being on board more than 30 days. additionally, the css states that no controls were ran on (b)(6) 2009, which is the date that patient 2 sample, was analyzed. the abbott architect system operations manual ((pn 201837-105) may, 2008) provides the following information related to addressing the customer issue: section 7 discusses the operational precautions and limitations, in section 5 operating instructions, it discusses the descriptions of patient result flags, descriptions of processing codes and the reagent inventory management. in the architect ca 125 ii assay package insert (016-622 5/07): storage instructions are provided and states the architect ca 125 ii reagent kit may be stored on board the architect i system for a maximum of 30 days. after 30 days, the reagent kit must be discarded. for information on tracking onboard time, refer to the architect system operations manual, section 5. a review of the service history for architect isystem s/n (b)(4) was performed for the time period of (b)(6) 2009. no additional complaints of this nature have been documented against architect s/n (b)(4) for the time period of (b)(6) 2009. this is the final report. | 0
1372784 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the architect i2000sr analyzer is generating discrepant total psa results on 2 patient samples. for patient #2: in 2009, the result was 59.0 u/ml, this sample was then sent to a reference lab and the result was 34.0 u/ml. there was no impact to patient management reported. >< (b)(4). it was stated that the total psa results were discrepant which is incorrect. the ca125 results were discrepant. all other information and data is correct. the complaint text for architect isystem (ln 03m74-01) s/n (b)(4) indicates that two different patients had discrepant ca 125-ii assay results using an architect i2000sr analyzer and another analyzer. the additional analyzer was at a reference laboratory and the name was not provided. however, the methodology listed is chemiluminescence. patient 1 had the following ca 125 ii assay results: date: (b)(6) 2009, ca 125 ii result: 46.8 u/ml, flag: none; (b)(6) 2009, 47.5 u/ml, exp; (b)(6) 2009, 50.5 u/ml, exp; (b)(6) 2009, 38.8 u/ml, exp. the patient sample from (b)(6) 2009 was sent to a reference lab and a result of 18.8 u/ml was obtained. patient 2 had the following result generated: date: (b)(6) 2009, ca 125 ii result: 59.4 u/ml, flag: exp. the patient sample from (b)(6) 2009 was sent to a reference lab and a result of 34 u/ml was generated. the complaint text indicates that for patient 1, the sample with a result of 47.5 u/ml was analyzed on (b)(6) 2009. review of the customer returned results list report indicates that the sample analyzed on (b)(6) 2009 had a result of 47.5 u/ml, and the result from (b)(6) 2009 had a result of 46.8 u/ml. the customer returned 5 pages of data. page 1 is a results list report, which contains the results mentioned above. pages 2, 3 and 4 are the levey-jennings reports for the ca 125 ii assay. the only quality control (qc) results which correspond to patient sample testing dates are (b)(6) 2009. also, both qc results have a flag of exp listed. no qc results are contained for the dates of (b)(6) 2009. page 5 is the calibration curve details report for the ca 125 ii assay, and indicates the date of calibration was 01/24/2009. the customer support specialist (css) additional communication email indicates that the system logs and database are not available for return. the lot number of the ca 125 ii reagent was 66060m100, expiration date may 21, 2009. also, the css states that the exp flag next to the patient results is because the ca 125 ii reagent was expired on board, due to the reagent being on board more than 30 days. additionally, the css states that no controls were run on (b)(6) 2009, which is the date that patient 2 sample, was analyzed. the abbott architect system operations manual ((pn 201837-105) may, 2008) provides the following information related to addressing the customer issue: section 7 discusses the operational precautions and limitations, in section 5 operating instructions, it discusses the descriptions of patient result flags, descriptions of processing codes and the reagent inventory management. in the architect ca 125 ii assay package insert (016-622 5/07): storage instructions are provided and states the architect ca 125 ii reagent kit may be stored on board the architect i system for a maximum of 30 days. after 30 days, the reagent kit must be discarded. for information on tracking onboard time, refer to the architect system operations manual, section 5. a review of the service history for architect isystem s/n (b)(4) was performed for the time period of (b)(6) 2009. no additional complaints of this nature have been documented against architect s/n (b)(4) for the time period of (b)(6) 2009. this is the final report. | 0
1372785 | the account stated that the architect analyzer has generated erratic total psa results results on a patient sample. the initial result was 7.68, the retest result was 8.24. the qc was within range. the patients value for total psa was known to be 13.9 when tested in 2005. units of measurement were not provided. there was no impact to patient management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation: erratic free prostate specific antigen and total prostate specific antigen results. a review of the complaint text indicates on (b)(6) 2009, the ticket documents the customer observed free prostate specific antigen (psa) values greater than the total psa serum patient values. the total psa result was 7.68 ng/ml with a rerun result of 8.244 ng/ml. the patient's previous total psa result of 13.877 ng/ml was generated on (b)(6) 2005. the free psa patient value was 8.517 ng/ml with a rerun value of 8.273 ng/ml. the last free psa calibration performed on the instrument was (b)(6) 2009 with free psa calibrator lot#66115lf00, expiration date (b)(6) 2009 and free psa reagent lot#68228lf00, expiration date 9/1/2009. the sample needle was replaced on (b)(6) 2009 and the reagent 1 needle was replaced on (b)(6) 2009. the architect system operations manual ((b)(4) june, 2007): section 7, operational precautions and limitations. limitations of result interpretation: assay results must be used with other clinical data, for example, symptoms, other test results, patient history, clinical impressions, information available from clinical evaluation, and other diagnostic procedures. all data must be considered for patient care management. if assay results are inconsistent with clinical evidence, additional testing is suggested to confirm the result. the architect system has been validated for its intended use. however, errors can occur due to potential operator errors and architect system technology limitations. section 10, troubleshooting and diagnostics, observed problems, erratic results (i system) lists the probable causes and corrective actions for erroneous results. in the architect free prostate specific antigen (free psa) package insert (b)(4), literature is provided regarding suitable specimens, results, and limitations of the procedure. in the architect total prostate specific antigen (total psa) reagent package insert (b)(4), literature is provided in describing suitable specimens, results, and limitations of the procedure. based on the available information, the cause of the customers issue was not identified, nor was any deficiency identified for this complaint. this is the final report. | 0
1408348 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< when a field service representative (fsr) was performing troubleshooting on the architect i2000 analyzer for a clog in the waste line, waste fluid splashed from the tube when it was removed from its connection into the fsr's eye. the eye was washed immediately. the fsr reported eye pain and went to the ophthalmology department at the hospital urgently. the injured person revisited the hospital the next day still complaining of eye pain. there was damage to the cornea of the eye. the eye is being treated with ophthalmic antibiotic eye drops (fluorometholone 0.05% and ofloxasin 0.3%). the person was not admitted to the hospital. vision has not been impaired due to this issue and at present there is no eye pain. the fsr was not wearing eye protection (glasses) at the time of the incident. >< (b)(4). other evaluation code: wash zone 2 tubing. the field service representative ( fsr) was given eye drops for treatment of an eye injury caused by a splash of waste fluid. the fsr is currently experiencing no pain, and has recovered from the injury. architect labeling instructs the reader to consult material safety data sheets (msds) for safe use instructions and precautions and avoid contact with skin and eyes; if contact with material is anticipated, the reader is instructed to wear impervious gloves, protective eye wear, and clothing. additional warnings caution that failure to follow safe use instructions could cause injury. tracking and trending showed that other potentially related complaints have been received for liquid splashing, but no other similar complaints were returned for liquid splashing into the eyes during maintenance on the wash zone. no adverse trends were identified related to this issue, or the architect i2000sr instrument. a malfunction of the system was not identified. no product deficiency was identified. the system is meeting current safety standards. this is the final report. | 0
1408376 | buffer reagent not returned. review of msds indicated no expected irritant effect when following instructionsthe material safety data sheet (msds) for the architect concentrated wash buffer confirmed that the product contains a low concentration of sodium azide. contact of the architect wash buffer with the eyes or skin is not expected to cause any irritant effect, however, as a precautionary measure rinsing well with water is recommended. regulatory information contained in both the material safety data sheet and product label instruct the user not to breathe gas/fumes/vapour/spray from this product and to wear suitable protective clothing and eye/face protection.the customer was referred to the material safety data sheet (msds) for the architect concentration wash buffer.a review of our complaint log revealed no adverse trend for this list number or the issue under investigation. no product deficiency was identified. this is the final report. >< the account stated that the architect wash buffer reagent is causing allergic reactions when reconstituting the buffer washer. when the technicians handle the buffer reagent without gloves, they experience issues with irritation of the skin, eyes and face (redness, skin blotches),as well as difficulty breathing, due to emanations from the buffer. this issue has existed since 2006, but the account did not indicate this until now. the account resolved the issue by handling the buffer with gloves and protective eyewear. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1420462 | (b)(4). the architect i2000sr (b)(4) generated discrepant troponin results on one patient sample on (b)(6) 2009. the discrepant results were not reported outside of the laboratory, and there was no impact to patient management due to the discrepant troponin patient result. the customer ran a troponin control precision run using all 3 control levels and found the %cvs were a little high. an fsr checked the analyzer. the field service representative (fsr) replaced the sample probe and stat probe. the fsr calibrated the troponin assay and ran controls to verify the analyzer performance. the fsr reran the patient sample x5 and noted the i2000sr instrument was working to the expected specifications. the architect system operations manual (revision 96211-110 section 10) was found to contain adequate information to resolve discrepant result issues. the manual lists multiple probable causes for discrepant results including bent, damaged or misaligned probes. the stat troponin-i package insert ((b)(4)) was reviewed and was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the package insert notes under sample collection and preparation for analysis it is important to follow the routine maintenance procedures defined in the architect system operations manual for optimal analyzer performance. a history review found no additional incidents of architect i2000sr (b)(4) generating discrepant results, no other issues have been documented since the fsr serviced the i2000sr and replaced the sample and stat probes. the current rate for architect i2000sr erratic occurrences/ million tests and complaints/ million tests are within expected action limits. the probable cause of the discrepant troponin results was a malfunctioning sample and/or stat probe. the actual cause of the suspect patient results could not be determined with the information available. based on the available information and this investigation, no deficiency was identified for the architect i2000sr instrument list no. 03m74-02 related to the issue under investigation. this is the final report. >< an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the architect i2000 sr analyzer has generated a false positive troponin result on one patient sample. the initial result was negative at 0.12 ng/ml. the sample was then centrifuged and retested twice, the results were negative at 0.08 ng/ml and positive at 0.24 ng/ml. the account stated that >0.1 ng/ml is considered positive in their lab. the account then tested the sample six times on another i2000 analyzer and the results were all negative. there was no impact to patient management reported. | 0
1474288 | (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the customer stated results are not reproducible on the axsym toxo igg assay. a patient generated a positive result of 42 iu/ml and when retested in duplicate, both results were negative at 0 iu/ml. no impact to patient management was reported. >< (b)(4). the customer stated they use sst gel tubes and centrifuge the tubes at 1,700 g for 10 minutes. the complaint documentation does not state the integrity of the patient sample was checked. the customer ran a x10 control precision run on the axsym and % cv was within package insert specifications. a field service representative (fsr) checked the axsym analyzer and noted the axsym was working according to expected specifications. the axsym system operation manual contains adequate information on the probable causes and corrective actions for inconsistent results. the probable causes listed include fibrin, red blood cells or particulate matter in samples. the toxo igg package insert was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the insert notes specimens should be free of bubbles, fibrin, red blood cells, or other particulate matter. the insert also notes specimens showing particulate matter, red blood cells or turbidity must be centrifuged at >/= 10,000 x g for 10 minutes. a service history review found no additional incidents of axsym serial number (b)(4) generating inconsistent results since the fsr verified the axsym performance. a review of quality metrics showed no adverse trends for toxo igg assay inconsistent results generation. the current rate for axsym erratic occurrences/ million tests and complaints/ million tests are within expected action limits. the most probable cause of the inconsistent toxo igg patient results on one patient sample was an issue with sample integrity. the actual cause could not be determined with the information available. based on the available information and this evaluation, no deficiency was identified for the axsym analyzer list no. 07a83-01 related to the issue under investigation. | 0
1478282 | (b)(4). investigation summary: an evaluation was performed in response to troponin reproducibility issues on an architect (b)(4) analyzer. the evaluation consisted of a review of the complaint text, instrument service history, and a review of the architect system operations manual. the abbott customer support specialist (css) assisted the customer with replacing the wash zone probes. the architect operations manual provides multiple probable causes, which include wash zone probes are bent or damaged, and corrective actions to assist with resolving erratic results. a malfunction of the system was identified. based on a review of the instrument's service history record, since the customer replaced all 3 wash zone probes, no repeats of the issue have been reported. no adverse trends or product deficiency were identified related to the issue. this is the final report. >< the customer states that they are observing reproducibility issues when processing patient specimens using the architect i1000 analyzer with the troponin-i assay. the initial results appear to be falsely elevated compared to results obtained upon repeat testing. the customer provided the following patient data.results (ng/ml)initial result: 0.63 retest results: 0.00 & 0.00.no adverse outcomes were reported related to this issue. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1481818 | the customer stated an emergency room doctor questioned an elevated troponin-i result of 0.51 ng/ml (customer's normal range 0.0 - 0.03 ng/ml) generated on the axsym analyzer. controls had previously been within range earlier that morning, but when retested, control values were out of range. the customer replaced the sampling probe, replaced the bulk solutions, and used new lots of reagent to troubleshoot the control issue. controls were run and were within range. the patient sample was retested yielding a value of 0.03 ng/ml and a corrected report was issued. no impact to patient management was reported. >< the customer replaced the sampling probe, replaced the bulk solutions, and used new lots of reagent to troubleshoot the controls out of range results. the customer ran a successful fluidics check and meia verification. after troubleshooting, the customer ran controls and results were within expected ranges. the customer reran patient samples and found discrepant results had been generated.the customer requested service because they found the axsym was leaking, and there were puddles on the floor under the axsym. the field service representative (fsr) found the sampling probe was loose and in need of calibration. the fsr tightened and calibrated the sampling probe. the fsr found the waste drain was overflowing. the fsr replaced the quick disconnect kit and the waste bottle assembly to resolve the leakage problem. the fsr performed a startup and flushed fluids to verify the leakage was resolved.the axsym system operation manual contains adequate information on the probable causes and corrective actions for inconsistent results. the probable causes include dirty, damaged, misaligned or malfunctioning sampling probe. the troponin-i adv package insert was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the insert notes it is important to follow the routine maintenance procedures defined in the axsym system operations manual for optimal axsym performance. a complaint service history review found no additional discrepant result complaints have been documented for axsym serial number 8720 since the sampling probe was installed correctly and calibrated.a review of quality metrics for the axsym analyzer did not identify any adverse trends for troponin assay discrepant results generation. the current rate for axsym erratic occurrences/ million tests and complaints/ million tests are within expected action limits. based on the available information and this evaluation, no deficiency was identified for the axsym analyzer list no. 07a83-01 related to the issue under investigation. | 0
1481831 | the customer stated an emergency room doctor questioned an elevated troponin-i result of 0.62 ng/ml (customer's normal range 0.0 - 0.03 ng/ml) generated on the axsym analyzer. controls had previously been within range earlier that morning, but when retested, control values were out of range. the customer replaced the sampling probe, replaced the bulk solutions, and used new lots of reagent to troubleshoot the control issue. controls were run and were within range. the patient sample was retested yielding a value of <0.02 ng/ml and a corrected report was issued. no impact to patient management was reported. >< (b) (4) the customer replaced the sampling probe, replaced the bulk solutions, and used new lots of reagent to troubleshoot the controls out of range results. the customer ran a successful fluidics check and meia verification. after troubleshooting, the customer ran controls and results were within expected ranges. the customer reran patient samples and found discrepant results had been generated. the customer requested service because they found the axsym was leaking and there were puddles on the floor under the axsym. the field service representative (fsr) found the sampling probe was loose and in need of calibration. the fsr tightened and calibrated the sampling probe. the fsr found the waste drain was overflowing. the fsr replaced the quick disconnect kit and the waste bottle assembly to resolve the leakage problem. the fsr performed a startup and flushed fluids to verify the leakage was resolved. the axsym system operation manual contains adequate information on the probable causes and corrective actions for inconsistent results. the probable causes include dirty, damaged, misaligned or malfunctioning sampling probe. the troponin-i adv package insert was found to contain adequate information on specimen collection and preparation, cautions and limitations of the procedure. the insert notes it is important to follow the routine maintenance procedures defined in the axsym system operations manual for optimal axsym performance. a complaint service history review found no additional discrepant result complaints have been documented for axsym (b) (4) since the sampling probe was installed correctly and calibrated. a review of quality metrics for the axsym analyzer did not identify any adverse trends for troponin assay discrepant results generation. the current rate for axsym erratic occurrences/ million tests and complaints/ million tests are within expected action limits. based on the available information and this evaluation, no deficiency was identified for the axsym analyzer list no. 07a83-01 related to the issue under investigation. | 0
1487193 | (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). after further review it has been determined that this complaint is a duplicate of (b)(4) and is no longer associated with the correction and removal 1628664-7/24/09-001-c. the final investigation will be documented in 1628664-2009-00331. this is the final report. >< the architect i2000sr wash zone ground straps were replaced due to corrosion. no incident or injury were involved with this issue. | 0
1487196 | (b)(4). after further review it has been determined that this complaint is a duplicate of (b)(4). and is no longer associated with the correction and removal 1628664-7/24/09-001-c. the final investigation will be documented in 1628664-2009-00334. this is the final report. >< the customer reported the architect i2000sr wash zone ground strap was corroded and were scheduled to be replaced by service. no incident or injury were associated with this issue. >< (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1488011 | (b)(4). evaluation (b)(4) - other a single definitive cause for the erratic results could not be identified as the field service rep replaced multiple components and adjusted and tightened multiple connections. the field service rep (fsr) was requested to verify instrument performance. the fsr performed multiple part replacements and adjustments: the reagent carousel v-wheels were replaced as a precaution. the v-wheel assembly was changed at the process carousel level as a precaution and he cleaned the v-wheel assembly at the rv carousel and sample carousel. the matrix carousel v-wheels were changed as a precaution. the rv picker was changed as a precaution. the rotary cam was replaced as it was damaged. the buffer inlet tubing kit was replaced as one of the tubing was half broken at the buffer transfer pump union and it was leaking. the buffer delivery pump was found broken at the plastic union of the buffer inlet tubing. the syringe assembly tubing at the valve level was reseated as it was found loose. the connection of the tubing at the syringe assembly valve level was loose and there were some salt deposits at the union as it was leaking. the tubing was reseated and connected. the review of service history of axsym (b)(4) indicates there were no other occurrences of the customer described issue following field service visit on (b)(6) 2009. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the axsym analyzer has generated discrepant vancomycin results on 5 patient samples. for patient #3, the initial result was 1.84 mg/l, the retest result was 13.20 mg/l. the qc was within range. the account has decided to send out their samples to a reference lab. there was no adverse impact to patient management reported. | 0
1488014 | (b)(4). evaluation (b)(4) - other a single definitive cause for the erratic results could not be identified as the field service rep replaced multiple components and adjusted and tightened multiple connections. the field service rep (fsr) was requested to verify instrument performance. the fsr performed multiple part replacements and adjustments: the reagent carousel v-wheels were replaced as a precaution. the v-wheel assembly was changed at the process carousel level as a precaution and he cleaned the v-wheel assembly at the rv carousel and sample carousel. the matrix carousel v-wheels were changed as a precaution. the rv picker was changed as a precaution. the rotary cam was replaced as it was damaged. the buffer inlet tubing kit was replaced as one of the tubing was half broken at the buffer transfer pump union and it was leaking. the buffer delivery pump was found broken at the plastic union of the buffer inlet tubing. the syringe assembly tubing at the valve level was reseated as it was found loose. the connection of the tubing at the syringe assembly valve level was loose and there were some salt deposits at the union as it was leaking. the tubing was reseated and connected. the review of service history of axsym (b)(4) indicates there were no other occurrences of the customer described issue following field service visit on (b)(4) 2009. >< the account stated that the axsym analyzer has generated discrepant vancomycin results on 5 patient samples. for patient #4, the initial result was 14.9 mg/l, the retest result was 4.91 mg/l. the qc was within range. the account has decided to send out their samples to a reference lab. there was no adverse impact to patient management reported. >< (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1517913 | the account stated that the architect analyzer is generating aberrant troponin results on a patient sample. on (b) (6) 2009, a (b) (6) patient was hospitalized for feelings of dizziness. the patient's tropinin values were stable at 2 g/l on (b) (6) and (b) (6). on (b) (6), a new sample was tested and the result was 1.988 g/l. the account performed testing to eliminate heterophil antibodies for both samples and the results were both 0.979 g/l. the controls were within range. the sample from (b) (6) was then tested by the beckmann dxi method and the result was negative at < 0.010 ng/ml. there was no adverse impact to patient management reported. >< (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). evaluation, -(b)(4) other, erratic stat troponin -results. the patient was released from the hospital on (b)(6) 2009. the patient received a troponin result of 2.048 ug/l initially on (b)(6) 2009. on (b)(6) 2009, the patient received an initial and repeat troponin result of 2.019 and 0.979 ug/l, respectively. on (b)(6) 2009, a new sample was drawn from the patient. the initial troponin result of 1.988 ug/l was received. the sample was then reprocessed to eliminate heterophil antibodies and retested on the analyzer with the result of 0.792 ug/l. the sample was retested on the beckman-dxi where the result of <0.010 ng/ml was generated. the sample was also sent to a subcontracting laboratory and the result of <0.010 ng/ml was received. there was nothing in the professional, familial, or therapeutic environment that would lead to the production of heterophil antibodies in the sample. the rheumatoid factor assay was negative. the patient has been taking 5 mg of bisoprolol daily since (b)(6) 2009. the instrument logs and reaction vessel lot numbers were not available from the date of occurrence. a review of the complaint history for architect isystem (b)(4) over the period of time of (b)(6) 2008 through (b)(6) 2009 was performed. the service history review found two additional incidents of the architect (b)(4), (b)(4), generated discrepant results or imprecise controls documented before this incident that were resolved by customer support troubleshooting. based on the available information, no product deficiency was determined. since the instrument logs and reaction vessel lot numbers were not available, the cause of the erratic troponin results could not be determined. | 0
1526448 | the customer stated a patient generated an axsym total psa result of 0.00 ng/ml. a higher result was expected and was questioned by the physician. the initial sample and a redrawn sample were tested yielding expected results of 4.1 ng/ml. no impact to patient management was reported. >< (b)(4). follow-up report being submitted to correct date (b)(4) 2009 to (b)(4) 2009. >< (b)(4). this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. an investigation is in process. >< (b)(4). an evaluation was performed and consisted of a review of customer complaints, current axsym labeling, and the information provided including the complaint text. the axsym operations manual provides multiple probable causes and corrective actions to assist the customer with resolving discrepant results. the probable causes for discrepant results include bulk solution 3 dispensing improperly. complaint information notes a field service representative replaced the leaking bulk solution 3 fmi pump to resolve the discrepant results issue. a complaint review found there have been no additional incidents of discrepant result generation by axsym plus (b)(4), since the bulk solution 3 fmi pump was replaced. a review of axsym plus field performance data did not identify an increase in complaint activity for the issue the customer experienced. the most probable cause of the discrepant patient results was the improper functioning of the leaking fmi pump. the actual cause of the suspect patient results could not be determined with the information available. based on the available information and this evaluation, no deficiency was identified for the axsym plus analyzer list no. 07a83-03 or the fmi pump (b)(4) related to the issue under evaluation. | 0
1526453 | the customer states that six out of fifteen patient samples had incorrect results transmitted from the architect i2000sr analyzer to their laboratory's (b) (6) computer (lis). the results were reported from the lab. corrected reports were submitted. the customer gave one example of noticing a total protein result of 73 g/dl that retested at 7.0 g/dl. the customer also stated that some of the results had never printed from the architect even though the analyzer is configured to auto-print patient sample reports. the customer states that their lis recycles accession numbers about every two weeks and believes that the current incorrect results were transmitted from results on the same accession numbers that had been in use two weeks ago. the customer has no patient information available. there is no impact to patient management reported. >< (b) (4) this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< (b)(4). the analyzer's result log was reviewed and confirmed the repeated total protein result of 7.0 g/dl; however, the log did not capture the date of the initial sample run where the result of 73 g/dl was allegedly generated. further information was requested regarding the sample identification numbers (sid) of the six samples in question, but was not available per the customer. without the sid information, it could not be determined which samples allegedly had the incorrect results retransmitted from the instrument to the customer's laboratory information system (lis). the lis trace logs from the host lis were requested, but were not available due to privacy issues with the lis vendor. a complaint review was conducted for similar issues; however, due to the lack of information provided in the current complaint text (no sid information), any similar issues could not be determined. the abbott standard interface (b)(4) and the architect system operations manual (b)(4) provide information to address this customer's issue. no labeling deficiency associated with this issue was noted. the evaluation of the complaint text, log files, and complaint searches identified no software deficiency. no system malfunction was identified. this is a final report. | 0
1341373 | it was reported to tosoh bioscience, inc. that the calibration was not correlating. the account stated that they had two aia-600ii analyzers and one is not working properly. | 0
1063548 | this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. >< the account stated that the axsym analyzer generated a discrepant troponin result on a patient sample. the initial result was 0.409 ng/ml and the repeated result was 0.466 ng/ml. there is no impact to patient management reported. | 0
1224565 | possible false negative. diagnostic cells outside field of view (fov). customer called case negative and case was signed out. patient had a biopsy and urine specimen on the same day both of which were positive for endometrial adenocarcinoma. the ordering clinician called the lab to request that the pap be reviewed. upon review, one pathologist called the case reactive and another reviewed the case and called it positive for em adenocarcinoma. cytology application specialist (cas) reviewed the slide and did not find any trigger cells in the 22 fov. however, cas noted that the background of the slide was bloody, the slide was atrphic and hyocellular. cas found papillary clusters of atypical cells consistent with endometrial adenocarcinoma outside the 22 fov. | 0
1236122 | on 16-sep-2008, acist received notification of a user facility report, regarding the 2.5' manifold tubing in the bt2000 manifold kit developing a bubble in the tubing during a coronary angiogram. the acist bt2000 manifold kit, lot 0928w, was received by acist on 08-sep-2008. investigation and testing was conducted on 09-sep-2008. acist verified that fluid had ingressed between the two layers of the 2.5' co-extruded manifold tubing, causing a bubble to develop in the tubing. the bt2000 manifold kit was tested using a 4fr catheter and reno 60 contrast media and put through a series of injections. the bubble in the tubing increased in size with each injection. after 51 injections, the bubble ruptured. there was no evidence of air ingress into the tubing or pieces of tubing breaking off the ruptured section of tubing. per acist's complaint database, the failure rate for this failure mode is 3.8 occurrences per 1,000,000 procedures. per review by acist's medical advisory board, and based on the results of the testing, there is no evidence of a potential for air ingress with this failure mode. see scanned pages. >< user facility reported: this product was used in conjunction with a coronary angiogram. tubing in manifold system developed a bulge in the tubing. use of the product discontinued. no apparent problems, but could have been an issue if tubing completely ruptured, which would have resulted in air in the system. | 0
1124467 | operator experienced discrepant urine creatinine results when one pt sample was aliquoted by the mpa versus manual placement of the sample onto the same analyzer. initial result 0.36, repeat 60.86, units of measure not provided. erroneous result was not reported. the field service rep was unable to determine root cause. performance tests were performed. | 0
1490431 | customer states he believes the vsii is not aliquoting accurate amounts of serum into the twin tubes. he noticed a pt was run on a twin tube for a ferritin and when repeated later on the primary tube, the results were different. the initial ferritin result was 150 ug per l when run on the centaur. the result was questioned by the physician who requested a rerun. the primary tube was run on the same centaur giving 4.9 and 5.7 ug per l. customer had no history or medication info regarding the pt. pt was not adversely affected. the field service representative could not find a cause for the discrepancy. he performed test run with colored water to ensure proper pour off into correct tubes with good results. | 0
1379719 | the field service representative determined that although the root cause could not be verified, it is suspected that the aliquot cup may have been previously used and contained a urine specimen prior to this sample being aliquoted into it. he checked the aliquator operation and other aliquot result correlation with primary tube. the lab manager was informed, and will communicate to operators the importance of disposing of all cups that have been through the analyzer. >< user received erroneous results for multiple assays on one aliquoted sample that was tested on two modules. initial results were: calcium 13.3 and 12.7 mg per dl, bun 196 mg per dl, magnesium 6.0 mg per dl, bicarbonate 19 mmol per l, creatinine 17.84 mg per dl, glucose 87 mg per dl, phosphorus 10.9 mg per dl, and potassium 9.4 and 9.4 mmol per l. after the initial results were received, the primary sample tube was tested. repeat results from the primary sample tube were calcium 8.1 mg per dl, bun 14 mg per dl, magnesium 2.4 mg per dl, bicarbonate 29 mmol per l, creatinine 0.52 mg per dl, glucose 118 mg per dl, phosphorus 2.6 mg per dl, and potassium 4.0 mmol per l. the patient was not treated based on the erroneous recovery as the user noticed the issue prior to reporting the patient. | 0
1148840 | (service replaced the fan and gas springs for resolution). service replaced the commander ppc gas shock and fan and verified proper operation of the commander ppc through diagnostics. a review of tracking and trending found no other complaints for the time period of february 1, 2008 through july 31, 2008 for gas shock/fan failures. the commander ppc operations manual, section 8, hazards, physical and mechanical provides adequate instruction to resolve the issue and continue operation. this is a final report. >< the account stated the commander ppc fan is not functioning properly and the gas shock for the lid is not working as expected. service changed the gas shock and fan, verified proper operation of the commander ppc through diagnostics. there was no injury reported due to the gas shock lid or fan on the commander ppc. | 0
1057662 | a customer had cut her finger on the waste chute (diti slide) during daily maintenance while she was drying the waste chute around the cut away section near the bottom of the chute. the operator was wearing gloves at the time. the glove was torn and the customer experienced an open cut on the finger. this injury carries a risk of potential biohazardous exposure. >< at this point, a customer letter is being written to alert customers to not put their hands inside the waste chute. secondly, the waste chute fabrication will include all edges to be broken with a 0.2 mm cut at a 45 degree angle. even though the material specification does include this 45 degree angle, these waste chutes were not fabricated to the edging specification. new waste chutes will be provided to the customers once this fabrication has been actually implemented in the waste chute production. | 0
1381030 | after the spinal block was placed for the scheduled cesarean section, the patient was inadvertently administered succinylcholine instead of ephedrine. the patient became paralyzed and had awareness of difficulty breathing for about 60 seconds. the patient required a brief period of intubation and ventilatory support. the patient recovered satisfactorily from anesthesia. additional information obtained from the site: after this occurrence, pharmacy implemented a new process whereby they manually affix a yellow warning sticker that states 'warning: paralyzing agent' to the backside (white) of the label for the succinylcholine. right now, one of our pharmacists is researching switching to a different manufacturer to purchase these drugs from. the product being researched comes in pre-filled syringes and is clearly labeled from all sides/angles (without obliterating the gradiant markings). this product features better labeling. investigation: the rapid-fill automated syringe filling system (asf) automates on-line label printing by printing the label information and cutting the finished, labeled syringe from the strip. the label is affixed to one-half of the syringe barrel so that gradations on the syringe barrel are legible and it does not impair the ability to check the medication dosage. the asf is used to fill and label syringes of ephedrine (violet-colored label) and succinylcholine (fluorescent red-colored label). the labeled syringes of ephedrine and succinylcholine are stored in the same drawer of the pyxis unit. a staff member removed the syringe labeled 'succinylcholine' from the pyxis and administered it instead of the ordered dose of ephedrine. when the syringe is in the position to allow visualization of the gradation markings on the syringe barrel, only the white-colored backside of the label is visible. apparently, the staff member did not check the printed medication label prior to administration. a root cause analysis of this occurrence will be completed. an interim measure has been implemented. a yellow-colored sticker that states 'warning: paralyzing agent' is affixed to the white backsides of the succinylcholine labels before they are stored in the pyxis. there have been no other complaints with this device other than there were some issues with a batch of syringes where the label went over the markings on the syringe. | 0
1336078 | random urine sample tested negative for protein. same sample tested using a different method had a high amount of protein. a 24 hour urine sample from same patient collected the same day gave total protein result of 23.94 g/24hr when tested at another laboratory. the 24 hour urine sample was then tested on the initial instrument three times and gave result of 500 mg/dl each time. the 24 hour urine sample was also tested by immunofixation electrophoresis and gave a total protein result of 19502 mg/day. the initial result was reported. no information provided to determine if result was used to guide therapy. the field service representative determined there was a problem with the transport plate and reference pad. the instrument was repaired. | 0
1435559 | customer reported a leak from the overflow tubing onto the floor in front of the counter. the fluid is in a walkway and all over the floor. no pts were affected and no operators were harmed. >< the field service representative found the cause to be a clot in the drain line causing the drain to default to drain overflow spout. he removed the clot in the drain line, and performed probe wash. he observed analyzer was draining correctly, and ran controls to verify analyzer operation. | 0
1437735 | user experienced two pt samples which generated results for erythrocytes from the analyzer that did not compare with results from the microscopic examination. the following pt sample was determined to be discrepant. initial erythrocyte result from analyzer was negative. microscopic examination of the same urine sample showed 5 to 10 rbcs per hpf. sample was repeated on the analyzer generating a negative result for erythrocytes. both dipstick and microscopic results were reported for this pt sample. user states the pt would not have been treated for the negative result that was reported on the dipstick. the user did not have info as to whether the pt was adversely affected by not receiving treatment. >< although the field service representative was unable to verify the problem, he noted a problem with the calibration as a cause. he verified default settings for reflectance, performed a calibration and ran controls to verify analyzer performance. | 0
1418751 | the investigation revealed the cause for the result variation was unk and the original complaint was not duplicated. during repair evaluation, several loose pt strips were noted inside the analyzer. the analyzer and optics were cleaned. set up, calibration and controls were run. the overall performance passed. >< the user mentioned that a urine sample had a negative blood result but there were 4+ rbc noted on the microscopic exam. she then reran the sample on the same analyzer one hour later and the blood result was 250 rbc per ul. the protein result was originally negative and repeated 100 mg per dl. the original results were not reported. the pt was not affected. | 0
1502913 | user received false negative leukocyte results for 3 pt urine samples. microscopic examination of the same urine samples showed 34 to 50 wbc's per hpf. initial results were reported. pts were not adversely affected. the field service rep was unable to determine a cause and could not duplicate the issue. he noted upon his arrival, the customer was running the instrument with no issues. to verify analyzer performance the field service rep re-calibrated leukocytes three times, ran precision studies using samples, which were negative and positive for leukocytes. customer ran controls and pts. all results were within spec. | 0
1299325 | user experienced discrepant leukocyte results 4-5 times in the last two weeks. two examples were provided. sample 1 from female pt, read negative on leukocyte, but also had microscopic examination ordered. microscopic result for wbc was reported as 30-61/hpf. user then ran chemstrip manually and the strip read leukocyte as positive. user then reran the sample on the instrument and recovered leukocyte as 25/ul. in 2008, sample 2 from female pt, initial chemstrip reading was negative. microscopic result for wbc was reported as <123/hpf with clumps. sample was retested and chemstrip result was 25/ul. no erroneous results were reported. if additional info is received, appropriate notification will be provided. | 0
1193398 | user experienced discrepant urine leukocyte results over the last 3-4 weeks. exact number of pt samples affected is unk. two samples were affected this day. one pt example was provided. initial result was negative, repeated on another analyzer gave positive result. same sample examined microscopically gave 29 wbc/hpf. initial result was reported and a corrected report was sent. pt was not adversely affected. the field service representative determined the analyzer was contaminated, and the led reflectance was low. he decontaminated the analyzer and replaced the led and cell. performance tests were performed which were within specification. | 0
1468525 | user experiencing pt urine samples that resulted negative for leukocytes, but were positive for leukocytes upon microscopic examination. user provided the following three pt urine examples. sample 1, meter read negative for leukocytes; microscopic examination showed 20 to 50 per hpf. on (b) (6) 2009: sample 2, meter read negative for leukocytes; microscopic examination showed 20 to 50 per hpf. on (b) (6) 2009: sample 3, meter read negative for leukocytes; microscopic examination showed 20 to 50 per hpf. erroneous results were not reported. the pts were treated based on the microscopic results, not the instrument results. no pt adversely affected. if additional info is received, appropriate notification will be provided. | 0
1338343 | user experienced false negative urine results for one pt sample. initial testing generated negative results for protein, glucose, ketone and erythrocytes. upon repeat testing with a new testing strip, the protein was 500 mg per dl, glucose was 300 mg per dl, ketone was 150 mg per dl and erythrocyte was 250 per ul. these results were confirmed via microscopic examination. the erroneous results were reported. pt was not adversely affected. >< upon investigation, no cause could be found and the problem could not be verified. the measurement module was replaced and calibration test and normal and abnormal controls were run with all test results within specification. | 0
1401918 | possible delay in initiating a code response. rt weaned pt from vent 15 mins prior to event, stepped out of room and upon return to draw abg, found patient unresponsive and bedside monitor with all zeros. no audible alarms noted. main desk monitor checked and noted that the heart rate alarm was off and when turned on at the nursing station, the alarm turned itself off immediately. this occurred multiple times. fukuda was called and guided through the configuration settings. the problem was that the configuration for heart rate alarm source was set on spo2 and not on ecg on the bedside monitor. | 1
1138513 | customer reported receiving erratic readings on their blood glucose monitor within 10 minutes. results of 39 mg/dl, 177 mg/dl, 166 mg/dl and 161 mg/dl were plotted on a parkes error grid. the results fell into the 'c' zones showing the difference in values to be clinically significant. there was no report of death, serious injury or mistreatment associated with this event. >< the customer's product has been requested back for an investigation. a follow-up report will be filed once investigation results are available. | 0
898120 | dbss allows a user to edit an incoming shipment box while another user is processing the shipment box. dbss does not back track and retroactively quarantine units already in the incoming shipment box. | 0
1353951 | the customer's product has been requested back for an investigation. a follow-up report will be submitted if the meter is returned. >< a customer reported her blood glucose test did not start after the blood sample was applied to the test strip. the customer reportedly lost consciousness. it was also reported the paramedics were called and they gave to the customer a glucose medication. the customer was then reportedly transported to a hospital where she was diagnosed with severe hypoglycemia and treated with an unknown intravenous medication. there was no report of death or permanent impairment associated with this event. | 0
1380229 | this is an initial report. the product has been requested back for an investigation. the follow-up report will be sent when investigation results are available. attempts are being made to contact the customer to obtain a meter's serial number and also determine what strip lot was in use at the time of the medical event. >< a customer's family member called in in 2009, reporting that the customer was getting an error message on his meter and was unable to test, which resulted in a medical event in the second week. the caller stated that at the time of the medical episode , the customer experienced dizziness, lightheadedness and passed out. his wife transported him to a local hospital where he was diagnosed with hyperglycemia and treated with unknown intravenous solution. additionally, the customer drank diet coke to counteract his symptoms. there was no report of death or permanent impairment associated with this medical event. | 0
1178891 | customer reported receiving erratic readings on his freestyle lite blood glucose meter. customer reported receiving readings of 384 mg/dl, 349 mg/dl, 84 mg/dl, 71 mg/dl, 82 mg/dl, 155 mg/dl, 79 mg/dl, 63 mg/dl, and 66 mg/dl within 10 minutes. all tests were performed on the finger. the results when plotted on a parkes error grid fell into the 'c' zone showing the difference in values to be clinically significant. customer also reported symptoms of hypoglycemia for which he self-treated by taking his usual diabetes medication. no third-party medical intervention was required. there was no report of death, serious injury or mistreatment associated with this event. >< this is an initial report. the device has been requested back for investigation. a final report will be submitted when the investigation is complete. | 0
1369477 | the product has been returned for investigation and the complaint was not confirmed. no new issues were observed, and all results were within range specification. >< customer resprouted receiving erratic readings on their blood glucose meter. customer reported receiving readings of 158 mg/dl, 60 mg/dl and 35 mg/dl within 10 minutes. tests were performed on the finger. the results when plotted on a parkes error grid fell into the 'c' zone, showing the difference in values to be clinically significant. there was no report of death, serious injury or mistreatment associated with this event. | 0
1302592 | customer reported receiving erratic readings on their precision xtra glucose monitor. customer reported receiving readings of 429 mg/dl and 128 mg/dl within 10 minutes. the results when plotted on a parkes error grid fell into the 'c' zone showing the difference in values to be clinically significant. there was no report of death, serious injury, or mistreatment associated with this event. >< the customer's product has been requested back for an investigation. a follow-up report will be filed once investigation results are available. | 0
1289499 | this is an initial report. the product has been requested for investigation, and a final report will be submitted when the investigation is complete. >< customer reported their meter had missing segments and they were unable to test. as a result, they reported experiencing symptoms of feeling dizzy, weak and disoriented. customer was seen at a local hospital, diagnosed with hyperglycemia and treated with insulin. there was no report of death or permanent impairment associated with this event. | 0
1327944 | this is an initial report. the customer's product has been requested back for an investigation. a follow-up report will be filed once investigation results are available. >< customer reported receiving an errc 6 message from his precision xtra meter which prevents him from blood glucose testing. customer also reported experiencing symptoms of dizziness, shakiness and cold sweats. customer reported visiting his doctor who diagnosed customer with severe hyperglycemia and treated customer with insulin and oral diabetes medication. there is no report of death or mistreatment associated with this event. | 0
1387752 | i purchased an invacare model #isg4002654 blood pressure monitor after my doctor told me she wanted to monitor my blood pressure at home because she was considering putting me on high blood pressure medication. i made my first purchase in 2009. i used the monitor at home and my blood pressure according to the invacare machine was around 135/140 systolic, this was low enough that i did not need blood pressure medication. however, after having two more blood pressure checks at my doctor's that were over 150 systolic, i bought my invacare blood pressure monitor into my doctor's office and they made sure i was using it right, then we compared between the invacare model and the doctor's office bp machine. the numbers were off by more than 12 points systolic. we repeated the comparison with the same results. i called the company i purchased the machine from, united health care products, and they sent me an identical invacare bp machine. because of the above mentioned problem, i did not trust the machine so immediately compared it again with my doctor's machine. unfortunately, again the systolic and dia were always much lower than my doctor's machine. unfortunately again the systolic and dia were always much lower than my doctor's, bp machine. i called invacare for assistance and was told i must be doing something wrong. they wanted me to purchase a larger cuff, despite the fact that my bicep measurements are within their parameters for the cuff i have, and i can get two fingers in the cuff as instructed. invacare offered no further assistance with this matter. i now have two new blood pressure machines that repeatedly produce lower blood pressure states. i gave the machines to two friends that have normal blood pressure to also compare. they both reported the blood pressure machine reported consistently lower blood pressure readings. one of the friends is an army medic and he said, 'if my blood pressure was that low i would be dead.' i reported this to invacare in an e-mail and have received no response. my concern is that people like me that are border line in need of medication will rely on these machines and not get the medication they need or consult a doctor because they believe their blood pressure is ok. in my case my doctor went ahead and put me on blood pressure medication. but what of the thousands of people that purchase this machine and for whatever reason it reports very low blood pressure when they may in fact be in need of medication? i know that two machines do not make a scientific study, and i also realize that there still could be a user problem, but if that is the case, the problem is so difficult even a trained nurse and an army medic are not applying the cuff correctly, which is hard to believe. please investigate this issue, as the 'silent killer' may be taking people unnecessarily after they have relied on this machine. dates of use: 2009. diagnosis: high blood pressure monitoring. | 0
1147832 | evaluation in progress but not concluded. >< during preparation for use for a cardiopulmonary bypass procedure, the user reported the pressure sensor could not communicate with the central control monitor. an alternate device was employed. there was no adverse consequence to a patient as a result of this event. | 0
1400290 | re: wristech blood pressure monitor purchased through distributor . monitor consistently gives readings 30 to 40 points lower for systolic and 10 points lower for diastolic than my physician's professional monitor. this is no problem for me since i have regular checks for elevated blood pressure. however, for anyone unaware that he has high blood pressure and just wants to check himself occasionally, this erroneously low, misleading reading would lead to the conclusion that he has no problem. i informed distributor of the problem, and suggested that unless they knew that this one monitor was the only one malfunctioning, they should inform all other purchasers about the problem. the only response was to give me instructions about how to return the monitor for a refund. dates of use: 2009. diagnosis or reason for use: elevated blood pressure. | 0
1098743 | pt underwent a l5-s1 fusion (hemilaminectomy. partial vertebrectomy, discectomy with fusion, and transforaminal lumber interbody fusion) in 2008. approx four days postoperatively, the pt was experiencing pain, fatigue, and fluid discharge. six days later, the pt underwent an irrigation and debridement procedure. antimicrobial therapy was initiated. the recipient's blood cultures were positive for fusobacterium. >< manufacturing and donor records were re-viewed. the device underwent the validated dosage to meet the sterilization requirements and passed all release criteria at the time of distribution. no departures were noted. no other complaints have been associated with this donor. the recipient's culture report showed positive for fusobacterium in the blood and no growth of the bacteria at the surgical site. the infection and the type of microorganism identified are suggestive of contamination extrinsic to the device, most likely a nosocomial acquired infection. | 0
1511518 | patient called endoscopy to ask about resuming her prilosec. she was referred to doctor for definitive answer since she had not been told to hold it. at the time of that conversation, her bravo receiver was functioning. she called again about 30 minutes later to report that the receiver had the time, but was now flashing 'c1'. she had tried moving the position of the receiver, and turning off the tv. nurse told her that we would call the company and call her back. twenty minutes later nurse called pt back, and offered to look at the receiver now or in the am. pt states she lives close and will come now. she arrived and the receiver was flashing the time and c1. the battery was changed, and within a minute it was flashing c1 again. pt was told that the receiver or capsule was not working and will need to have another study done later. dr office notified by nurse. >< patient called endoscopy to ask about resuming her prilosec. she was referred to doctor for definitive answer since she had not been told to hold it. at the time of that conversation, her bravo receiver was functioning. she called again about 30 minutes later to report that the receiver had the time, but was now flashing 'c1'. she had tried moving the position of the receiver, and turning off the tv. nurse told her that we would call the company and call her back. twenty minutes later nurse called pt back, and offered to look at the receiver now or in the am. pt states she lives close and will come now. she arrived and the receiver was flashing the time and c1. the battery was changed, and within a minute it was flashing c1 again. pt was told that the receiver or capsule was not working and will need to have another study done later. dr office notified by nurse. | 0
1417674 | during patient telemetry ekg/arrhythmia monitoring, the transmitter battery got weak and died (stopped functioning). the central station was configured to alert the nursing staff through a secondary alarm messaging system to a rf wireless phone system when the battery became weak, and again when the battery needed to be replaced. the 'battery weak' alert was transmitted to the nurse by way of the rf wireless phone, the 'replace battery' alert was not sent to the nurse. manufacturer response (as per reporter) for central monitoring system, telemetry, intellivuea manufacturer's representative has examined the system and observed the problem. it was indicated that further tests needed to be performed at the factory to determine the cause of the problem. | 1
1325939 | it was reported that there was a loss of monitoring on two cic's lasting three hrs. the clinicians acquired alternate means of pt monitoring and connected the pts to dash bedside monitors. no pt injury or death reported. this report is the first of two mdrs. ge healthcare's investigation into the reported occurrence is still ongoing. a f/u report will be issued when the investigation has been completed. | 1
1325943 | it was reported that there was a loss of monitoring on two cic's lasting three hrs. the clinicians utilized alternate means of pt monitoring and connected the pts to dash bedside monitors. no pt injury or death reported. this report is the second of two mdrs. ge healthcare's investigation into the reported occurence is still ongoing. a f/u report will be issued when the investigation has been completed. | 1
1396684 | ge healthcare investigated the reported issue and determined that the event is related to a know condition, where the time master on the unity network broadcasts time changes and the cic reflects these changes, causing multiple time requests and excessive network traffic. this can result in the following scenarios: sluggishness or unresponsiveness of the cic pro or monitoring device user interfaces; loss of waveforms, parameters and/or alarming on the cic pro; restarting of cic pro processes or rebooting of the cic pro operating system; time changing rapidly between two or more times on unity devices; time not advancing on unity devices; inconsistent time values on devices across the unity network; rebooting of monitoring devices; excessive network traffic that delays, or otherwise interferes with, legitimate unity network communication. ge healthcare will be providing a product upgrade to prevent the potential for issues associated with network time changes on the carescape cic pro. this was reported to fda on 02/04/2009. >< it was reported that time on the cic was frozen at 03:30 am. no loss of monitoring occurred. no patient injury or death reported. | 1
1424520 | the company service representative replaced the elan switch. the icu cpu was rebooted to restore communication. the system was tested to factory specifications. it functioned normally and was returned to use. >< the customer reported that while the panorama central station was in use monitoring patients along with bedside monitors and ambulatory telepacks, the central station suddenly stopped, displaying traces for the patients in the icu, er, and telemetry. the biomed rebooted the elan switch and wireless communication was restored. traces did not return in the icu, which was hardwired. no patient injury was reported. | 1
1381754 | after adding dicom tag 0018, 1510 (primary positioner angle) to the viewport annotations, the value is displayed in vp, but the minus sign is not displayed. sample exam in pacs30 (testlab) named 'getest' tag 0018, 1510 contains value '-3,15', but only '3,15' is displayed in the viewport. this will be addressed with cw3.0.7. this mdr is being submitted late, as this is part of our ongoing quality improvement activities. we have updated the risk assessment control record and the update has resulted in the assessment that this reported complaint meets the reportability control record for ge healthcare. (b) (4). >< the customer reports that the minus sign is not shown on viewport annotation. there was no reported patient injury associated with this event. | 1
1531475 | the patient had to be re-exposed after failure of equipment to properly send exposed cassettes. the images were never transmitted to the computer from the digitizer, and could not be found on the digitizer to be pushed to the computer. the agfa representative was here and determined this to be a networking issue which was corrected. >< the patient had to be re-exposed after failure of equipment to properly send exposed cassettes. the images were never transmitted to the computer from the digitizer, and could not be found on the digitizer to be pushed to the computer. the agfa representative was here and determined this to be a networking issue which was corrected. | 1
1440023 | patient was re-exposed after agfa equipment failed to send images from digital reader to workstation. after 30+ minutes the images were transmitted. at that time the pt had already been x-rayed a second time.the device was so slow that staff believed it to be defective and ordered another x-ray exposure. the device continued to cause problems and failed to retrieve other images, causing delays in care and exposing other pts to double exposures.agfa responded and repaired the cr reader. the problem arose from certain patient medical alert info. the data held within the 'medical alerts' field of the patient record was found to cause problems when it was greater than 64 characters. this feature has been disabled and the cr readers continue to work normally. >< patient was re-exposed after agfa equipment failed to send images from digital reader to workstation. after 30+ minutes the images were transmitted. at that time the pt had already been x-rayed a second time.the device was so slow that staff believed it to be defective and ordered another x-ray exposure. the device continued to cause problems and failed to retrieve other images, causing delays in care and exposing other pt's to double exposures.agfa responded and repaired the cr reader. the problem arose from certain patient medical alert info. the data held within the 'medical alerts' field of the patient record was found to cause problems when it was greater than 64 characters. this feature has been disabled and the cr readers continue to work normally. | 1
1138975 | investigation is still under way to determine the root cause. parts are being returned to the factory for failure analysis. >< the couch system developed a 'noise' when moving vertically. upon investigating the source of the 'noise' a decision was made by the field service engineer to replace the brake. the new brake system solved the issue of the 'noise'. after the new brake system was installed the couch experienced an uncommanded motion to the lower couch limit. no patient injury was reported from this incident. | 0
1267907 | a map shift was suspected when compared to the fluoro image. no data was obtained on the amount of shift experienced. there was no pt injury reported. one of the connector pins was found to be broken. >< . | 1
998059 | philips has determined that some physician entered orders are not showing up on worklists. new information received 07 february 2008 identified patient risk. philips is in the process of obtaining more information concerning this event, and this complaint is still under investigation. a final report will be submitted when the investigation is completed. >< orders are not shown in the worklist for patient in i3box 3. | 1
1079892 | skin reaction is a rare and known risk of the use of ten20. patient had very tight ace bandage covering the electrodes and it was in place for three (3) days. the electrode application method may have contributed significantly to the severity of the injury. patient is slowly improving in the scarred areas but will likely have some visible signs of the injury remain for a significant time. >< female patient had a 3 day monitoring eeg exam, in which electrodes were placed on the forehead, behind ears, and proportionally around the scalp according to eeg practices. the electrodes were grass eeg electrodes filled in the cup portion with ten20 conductive paste according to a consulting technologist, the patient's head and forehead was wrapped with an ace style bandage and the patient was instructed to wear this for 3 days. the electrodes were hooked to an eeg monitoring device. the patient described the wrap as feeling tight. she was instructed not to remove the wrap or the electrodes for 3 days, according to her, but to 'press the electrode' if she felt any itching or discomfort. it was reported to us that the patient was motivated to go three days with no seizure activity because she would recover her driver's license if she made it through 3 days with no adverse eeg activity. therefore, even though she experienced discomfort at the electrode sites, she endured the irritation to accomplish her goal of receiving a drivers license. the two forehead electrodes, when removed, left severe imprints, redness, and yellow pustules. a consulting technician sent pictures of the injury to us through an attachment to email messages. according to the patient, there is still scarring at the sites of these electrodes. there was also some injury the electrodes that were placed adjacent to the patient's ears, but these were reported to be less severe...and of course...less noticable. | 0
1469759 | i have just switched to acuvue oasis lenses last week. i wore them for two days before one morning, i woke up and i could not open my eyes. they were blood shot and i felt like i literally had nails in them. i'm no stranger to lenses. i've been wearing them for 7 years now and i was just trying something new as they're being advertised everywhere. please, this has to be investigated. that same morning, i went to the medical clinic. the doctor at the clinic literally panicked and sent me as an emergency to an eye specialist within 2 hours. the specialist found massive swelling of my cornea to the point where my vision was completely distorted in my left eye. i am still waiting for my eyes to clear up and i have a third visit with the specialist in a week. this can cause permanent damage to the eye. i still don't have full vision back in my left eye and it's been a week! at this point, it does not look like i will be able to wear my contacts ever again. please, this has to be investigated. my eyes reacted too badly to these lenses and teared so much, i had blue bruises under both eyes from all the trauma. this is unacceptable. there is no warning at all that some people develop a reaction! dose or amount: 1. frequency: monthly. dates of use: 2 days of wear. diagnosis or reason for use: i need glasses, but i use contacts for the last 7 years. event abated after use stopped or dose reduced?: yes. event reappeared after reintroduction?: yes. | 0
1432026 | reportedly, cco value incorrect, erratic, drifts. customer changed monitors and showed values they felt were more accurate than what was reading on this monitor and they also tried a third and it correlated with the second monitor. no patient injury reported. | 0
1488022 | it was reported to boston scientific corp that a flexima biliary stent system was used during a stenting procedure in the bile duct of a male pt. during the attempt to deploy the stent, the guide catheter tore; however, the stent was able to be successfully deployed at the target site. the procedure was successfully completed with no reported pt complications. the pt was reported to be in good condition after the procedure. >< reportedly, the vigilance monitor, model vgs, 'smoked'. the customer stated that 'smoked from back of monitor'. no patient injury was reported. >< monitor display was damaged.complaint was confirmed. >< (evaluation): during monitor turn on spark was verified. no display on the screen due to pig tail of crt was broken. the monitor was not in operational condition. the pig tail of the crt can be damaged, due to miss handling of the monitor. >< on october 20th, 2009, baxter was contacted by a customer reporting that their product was broken. the intravenous tubing snapped causing the tubing to break, and it pulled the intravenous cannula out of the patient's hand. it is unknown at what time and stage the problem was noted. medical intervention was required in order to avoid patient injury; however, the clinical details of the intervention are not stated at this time. | 0
1517504 | it was reported to the manufacturer that the vns patient's generator has been moving in her chest. it was indicated that there was no patient manipulation or trauma at the device site that could have contributed to the reported event. the patient is being scheduled for surgery to reposition the generator to alleviate the discomfort from device migration. diagnostic tests performed on the device revealed proper device function. good faith attempts to obtain additional information regarding the reported event have been unsuccessful to date. >< vigilance ii monitor, model vg2, was reported by customer as having the cco value incorrect, 'cco value was too low'. no patient injury was reported. >< no defect found. could not confirm complaint. | 0
1232105 | device not returned. >< reportedly, if used with edwards in-line probe vigilance ii monitor calculates very low co values if the co is low (co ~ 1.4 - 2.0). if co is high there is a variance of bolus values of over 50%. no patient injury reported. | 0
1278883 | a cpoe device was deployed at (b)(6) hospital in 2008. this patient underwent an appendectomy two days later. the patient's care was governed by a cpoe device manufactured by cerner corp. in the care of this patient after his operation, there were 25 incidents that occurred involving flaws and defects in the device, interface defects, device user bewilderment, device caused hospital wide chaos, and device caused hospital wide near-meltdown and care disruption that resulted in neglect of this patient and his death. dates of use: 2006. diagnosis or reason for use: cpoe device was deployed by the hospital to govern all. event abated after use stopped: no. | 1
1274701 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been informed through adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 0
1046525 | a customer reported receiving an error 6 message displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been informed through adc fa21dec2006 letter. | 0
1058503 | in 2008, the dl2000 was upgraded to version 6.5 from version 6.4. a rule for low levels of vpa had been configured on the dl2000, but it was not re-validated after the software upgrade. per customer, the rule had been configured prior to the upgrade, but there had been no instances of vpa results being improperly changed to '<10ug/ml' with version 6.4. upon checking all vpa results reported from 04/01/08 to present, only samples reported since the software upgrade on original date, were incorrect. eval of the rule shows syntax errors in the rule, combined with version 6.5 software. bci did configure the rule on an in-house dl2000, and confirmed that the rule only change results with version 6.5. the errors in this rule were corrected and the issue has been forwarded to the software vendor for further investigation. a false negative result in a recurrent situation could place a patient(s) at risk with values round the toxic level threshold not properly reported. the root cause appears to be a rule with syntax errors, combined with version 6.5 software. a malfunction will be assumed for the purpose of this report. >< a customer reported to beckman coulter inc. (bci) that datalink2000 data manager (dl2000) incorrectly reported valproic acid (vpa) results to the lab info system (lis). several samples were tested for vpa and one pt questioned the results. upon data review, it was discovered that incorrect results were reported for 4 samples: actual vpa results at instrument were: 87.5ug/ml, 87.1ug/ml, 68.0ug/ml, and 23ug/ml for samples 1-4 respectively all 4 results were reported by the dl2000 as '<10ug/ml'. it is unknown if treatment was initiated or withheld as a result of this event. | 1
1075903 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customer and retailers have been informed through adc fa21dec2006 letter. | 0
1079474 | at this time, it is unknown who configured the rules and if they were verified. bci had the customer appropriately modify the set of rules to work as desired. the root cause was confirmed to be a rule issue for patient 1, but could not be identified for patient 2. no other additional information is available for this event. >< a customer reported to beckman coulter inc. (bci) that two suppressed patient creatinine (crem) results were incorrectly reported by datalink2000 data manager (dl2000) to the lab information system (lis). patient 1: a patient sample was tested for crem and a suppressed 'ea dl' result was uploaded from the instrument to the dl2000. the dl2000 uploaded the result to the lis as '<0.1'. patient 2: per customer, a suppressed 'ea' crem result was uploaded from the instrument to the dl2000. the upload from the dl2000 to the lis was not provided. based on available information, the lis reported the result as '<0.1'. the dl2000 had a series of rules to handle crem: rule 1: if results equals 'ea' - then stop (no further processing of rule per software definition). rule 2: if result contains 'dl' - then result is '<0.1'. for patient 1, the dl2000 did not stop the result for manual review. since the result did not contain the 'dl', so per the second rule, the dl2000 changed the result to '<0.1' and this was uploaded to the lis. for patient 2, the relevant data was not supplied. it is unknown what the dl2000 uploaded to the lis, but by software design when the rule states stop, no further processing of rule occurs. it is unknown if treatment was initiated or withheld. | 1
1083930 | the customer reported the time and date stored in their precision xtra meter are incorrect. the customer reported using the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified. | 0
1277340 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1277358 | a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer svc that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been informed through letter. | 0
1334800 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 0
1334801 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 0
1354496 | (b) (4). there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported incorrect time/date on the display of their precision xtra blood glucose meter. the customer additionally reported being a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1354497 | a customer reported incorrect time/date on the display of their precision xtra blood glucose meter. the customer additionally reported being a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< (b) (4). there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 0
1354498 | a customer reported incorrect time/date on the display of their precision xtra blood glucose meter. the customer additionally reported being a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified. | 0
1357126 | customers and retailers have been informed through the adc fa21dec2006 letter. >< a hospital called reporting their precision xtra meter had spontaneously changed the date and time. the customer also reports using the p link software and this is a known malfunction with the software that causes incorrect trending of glucose results. there was no report of death, serious injury or mistreatment. | 0
1359741 | a customer reported receiving an error displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 1
1359742 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 1
1359743 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified. | 1
1359749 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1371836 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1371841 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the letter. | 1
1371843 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1419441 | (b) (4). there is a known malfunction with the (b) (4) software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with (b) (4) software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the (b) (4). there was no report of death, serious injury or mistreatment associated with this event. | 1
1419446 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1448810 | (b) (4). there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by the adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1448821 | a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customer and retailers have been notified through the adc fa21dec2006 letter. | 1
1467967 | a customer reported incorrect time/date on the display of their precision xtra blood glucose meter. the customer additionally reported being a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< (b) (4). there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 1
1469766 | a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< (b) (4). there is a known malfunction with a precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. | 1
1469770 | there is a known malfunction with a precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported incorrect time/date on the display of their precision xtra blood glucose meter. the customer additionally reported being a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1480537 | a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< (b) (4). there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results, obtained on a meter with incorrect date and time, are uploaded to a computer with precision link software. customers and retailers have be notified through the adc fa21dec2006 letter. | 1
1296730 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the letter. >< a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1096777 | a customer reported receiving an error 6 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the data and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. >< there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect data and time are uploaded to a computer with precision link software. customers, and retailers have been informed through letter. | 1
1237711 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been informed through letter. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1423097 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been notified through the adc fa21dec2006 letter. >< a customer reported receiving an error 658 displaying on their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data mgmt system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1193412 | there is a known malfunction with the precision link software that can lead to incorrect trending of results. this occurs when results obtained on a meter with incorrect date and time are uploaded to a computer with precision link software. customers and retailers have been informed. >< a customer reported receiving an e6 message on the display of their precision xtra blood glucose meter. it was then additionally identified by adc customer service that the date and time settings in their meter were not properly set, and they reported to be a user of the precision link data management system. there was no report of death, serious injury or mistreatment associated with this event. | 1
1410808 | this report involves an electrical shock caused by a blood pressure monitor manufactured under the corporation named 'homedics theraip. 3000 pontiac trail, i purchased the monitor and commenced using in 2008. the monitor was plugged into the same outlet from that day until today - 2009. when i unplugged the monitor from that outlet this morning, the cover on the electrical plug detached and i received a shock that ran from my right hand to my right shoulder. a circuit board is built into the electrical plug and contact with the circuit board is what caused the shock. i called homedics today to report the incident and received an offer from the customer relations department to have the 'unit' picked up by united parcel service ups - so that homedics could 'look the unit over.' the customer relations woman stated that there had been no previous reports with the plug for the monitor. although i wasn't seriously or permanently injured i view this as a defect that should be publicized before someone is seriously injured. i attempted to report this through the consumer product safety commission who referred me to your organization. | 0
1386041 | investigation determined that when vixwin software interfaces with easy dental software, and there are two patients have the same name and no middle initial, the images will merge into both patient files. the software will be updated to correct this issue. >< cannot get two different patient files when interfacing vixwin platinum 1.1 with easy dental software, where two patients have the same name and no middle initial. | 1
1344425 | home pt found unresponsive by wife while connected to the homechoice peritoneal dialysis device. first day on home peritoneal dialysis on this equipment. paramedics called, and individual transported to local hosp ed - unable to resuscitate and pronounced dead. cause of death - cardiac arrhythmia. | 0
1425914 | in 2009, during an educational call, the patient reported he has had elevated readings up to 424 mg/dl due to not being trained on his bolus calculator. company records showed the patient had been contacted for training, but the patient did not provide his physician contact information so that parameters prescribed by the doctor could be received. a request for training was submitted. on follow up two days later, the patient's wife provided the doctor's contact information. no product will be returned for evaluation. >< no product will be returned for evaluation. | 1
1340073 | see attached complete report. quality issues exist with this product and other products included with this report. samples tested: 2. gauge endurance: both units failed the requirements under ansi. valve leakage: one of the valves exhibited leakage of 10 mmhg/min, well in excess of the 1 mmhg/10 sec limit mandated under ansi. operating instructions: not included as required under ansi. calibration info insert: incorrectly states 'if the needle rests within the zero indicator on the faceplate the unit is guaranteed to be accurate'. this is simply not true and could be dangerously misleading. any gauge could be tampered with and have the needle repositioned to appear accurate. the statement should read 'if the needle rests outside the oval, the gauge is in need of recalibration'. cuff markings: the artery mark is not positioned over the bladder midpoint as required under ansp. (exhibit a). traceability: this product does not appear to have a lot code or other identifier as per the requirements of ansi. samples tested: 2. gauge accuracy: both units failed accuracy as per the requirements of ansi, though one failed only at a single test point. gauge endurance: the one unit tested failed the requirements under ansi. operating instructions: not included as required under ansi. calibration info insert: incorrectly states 'if the needle rests within the zero indicator on the faceplate the unit is guaranteed to be accurate'. this is simply not true and could be dangerously misleading. any gauge could be tampered with and have the needle repositioned to appear accurate. the statement should read 'if the needle rests outside the oval, the gauge is in need of recalibration'. (exhibit b). cuff markings: the artery mark is not positioned over the bladder midpoint as required under ansp. cuff labeling: the cuffs do not have the proper limb range markings on them as required by ansi. quality issues exist with this product and other products included with this report. samples tested: 2. thermometer accuracy: both samples failed to meet the accuracy requirement under astm, though sample 1 was within the accuracy limits of our test machine. sample two registered one test point that was off by almost three times the allowed limit. samples tested: 2. thermometer accuracy: both samples failed to meet the accuracy requirement under astm. at one test point, both samples were off by three times the allowed limit. samples tested: 2. thermometer accuracy. both samples failed to meet the accuracy requirements under astm. at one test point, both samples were off by up to three times the allowed limit. lot number: they appear to use a date based lot numbering system (located on the rating label). with a lot dates of 02/06 on one sample it would suggest they are not within the 2 year life rating of the batteries. (exhibit d). dose or amount, frequency, route: #- na. dates of use: 2009. diagnosis or reason for use: blood pressure measurement - temperature measurement. | 0
1365176 | incident was reported with our dur-d, access sheath and 35bx guidewire. the surgeon was using this equipment and perforated a pt's ureter. the pt was suspected of having cancer in the kidney. as a result of the perforation, the surgeon had to open the pt to repair the ureter. once he opened, he discovered the pt had severe cancer and removed his kidney. >< the device has not been returned for evaluation. as a result, a determination cannot be made at this time. if further information becomes available, gyrus acmi will continue the investigation and update the agency accordingly. | 0
1265737 | the complainant reported that fluid leaked from the rotator connection during high pressure injection for a left ventriculogram procedure. there was no reported harm or injury to the pt or clinicians. >< device evaluation anticipated, but not yet begun. conclusion: the suspect device has been returned to merit for evaluation. a follow-up report will be submitted when the investigation is complete. | 0
1511240 | rotator broke during angiography using a micro catheter which occurred twice in a row. actual pressure 920 psi. (this report is one of two for reported device. the second device noted in this event is reported on mdr number 1721504-2009-00112). >< conclusion: a follow-up report will be submitted when the device evaluation has been completed. | 0
1079026 | device has not been returned for eval. >< it was reported the connector was broken. the broken area is on the px600f. | 0
1079030 | it was reported the connector was broken; broken area is located between the px600f and the 3-way connector. >< device has not been returned for evaluation. | 0
1079033 | device was has not been returned for evaluation. >< it was reported the tubing was broken: the broken area is located between the px600f and the transparent tube. | 0
1205683 | tubing was reported blocked. no pt complications were reported. >< method/results/conclusions: device has not been returned for evaluation. | 0
1208724 | device has not been returned for evaluation. >< saline leakage was observed from the dpt. location leakage was not specified in complaint. no patient complications were reported. | 0
1208745 | device has not been returned for evaluation. >< it was reported there was water leakage during use, but unsure where the leak is coming from. no patient complications were reported. | 0
1208746 | it was reported there was water leakage during use, but unsure where the leak is coming from. no patient complications were reported. >< device has not been returned for evaluation. | 0
1208747 | it was reported there was water leakage during use, but unsure where the leak is coming from. no patient complications were reported. >< device has not been returned for evaluation. | 0
1267956 | it was reported that blood leakage was observed from the connection between dpt and three way stop cock during use. no patient complications were reported. >< device has not been returned for evaluation. | 0
1267957 | device has not been returned for evaluation. >< the only information provided so far on this complaint is: they have experienced leakage on 3-4 sets during the weekend (for two days). more information will follow. incident date is yet unk. additional information received on october 22nd: summary of the main point, for additional information, please refer to complain form attached: all pts in 4 bed room of icu received problems with pressure monitoring sets: 4 cases: leakage close to the syringe near the '2 way valve', towards pt. medical staff could still aspire. when trying to aspire air was sucked into syringe. no air came out in the arterial tube to the pt. leakage close to the syringe. air was discovered in the arterial tube all the way towards the pt. air had probably been drawn into the tube by the continous flow of naci. other lot did not show any problems. no pt complications were reported. | 0
1267961 | device has not been returned for evaluation. >< abnormal wave form and pressure were reported. wave form was flat line and blood pressure measurement was 300mmhg on the monitor. no patient complications were reported. | 0
1267962 | device has not been returned for evaluation. >< it was reported that there was leakage observed from near the three way stop cock above dpt of yellow line. no pt complications were reported. | 0
1267963 | it was reported that blood leakage was observed from near the three way stop cock. no pt complications were reported. >< device has not been returned for evaluation. | 0
1392744 | device evaluation customer report was confirmed. rec'd (1) single dpt kit. tubing male connector at the end of iv line was damaged. the male luer was bent. stress mark was evident on the bent luer. the thread area of the connector was also cracked. >< a part connected with transducer was broken. | 0
1375952 | customer report was confirmed. rec'd (1) dpt - vamp adult kit. tubing was found completely broken off from uv bond joint with reservoir. tubing was bent near the site of damage. >< it was reported that the tubing pulled reservoir. no pt. complications. | 0
1247575 | it was reported that blood leaked out of the vamp flex syringe during the blood sampling three days after the artery line was indwelled. no patient complications were reported. >< device has not been returned for evaluation. | 0
1208725 | device has not been returned for evaluation. >< it was reported that the dpt could not regulate drip rate; another device from same lot number was opened and same event occurred. another device from a different lot number opened and worked fine. no patient complications were reported. | 0
1371701 | it was reported that when the nurse opened the package, it was noted that the tubing was cut/separated at approximately the vamp reservoir. no pt complications were reported. >< visual-tubing - customer report was confirmed. vamp adult-tubing was broken or torn off from bond joint with vamp reservoir. tubing was bent at the site of damage. physical appearance of bonding agent suggested that it was uv adhesive. (b) (4). | 0
1429203 | it was reported that line detachment at vamp reservoir during filling was observed.no patient complications, as it occurred before connecting the patient. | 0
1407645 | evaluation customer report was confirmed. rec'd (1) single dpt - vamp adult kit. vamp tubing was completely broken off from solvent bond joint with sample site. tubing did not appear to be bent at the site of damage. >< it was reported that line was broken before z-sampling board. incident occurred during filling but before use on patient. therefore no complications with the patient are reported. | 0
899676 | the device was not returned for evaluation. >< it was reported that the arterial line set had a faulty device that would have allowed a rapid infusion of 1000mls of saline. reportedly, the line was already set up and ready to connect to the pt when the malfunction was observed. there were no pt complications reported. | 0
1282322 | . >< the catheter was noted to be outside of the mapping shell at the completion of the procedure while the physician was maneuvering the catheter. it was not possible to place the catheter within previously created vessel tags. there was no patient injury reported. there was no indication of a map shift during a subsequent procedure. | 0
1311842 | . >< it was reported that a pericardial effusion occurred while performing a left atrial maze procedure. the pericardial effusion occurred when the physician was performing a roofline lesion on the right side of the left atrium. the patient had severe internal bleeding. drainage was attempted, however it was unsuccessful. the patient was taken to surgery and she was doing fine after the surgery. | 0
1350481 | it was reported that a map shift was experienced while using the carto rmt system. >< (b) (4). investigation of this complaint is still in progress. this report will be updated, upon completion of the investigation. | 0
1406094 | the customer was contacted by biosense webster field service engineer. the hospital ep lab staff stated that the stockert was not believed to be defective but it was hospital policy to get it checked out. investigation is still in progress. a supplemental report on device evaluation will be submitted once it is completed. >< it was reported that there was a pericardial effusion during a transceptal procedure. pericardiocentesis was performed and the physician was able to treat the vt arrythmias. physician managed to complete the procedure, however patient developed ventricular fibrillation and due to the deteriorating heart condition, it resulted in death. other medical intervention performed was cpr. the facility contact stated that none of bwi's products contributed to this incident. | 0
1320971 | it was reported that the tip of a catheter used during a ventricular tachycardia procedure was found in a folded position and was not responding to deflection at the point of attempting to remove the catheter from the patient. when consulted, the biosense webster representative suggested cutting the handle of the catheter to release the tension on the tip, however, it did not resolve the issue. aortic dissection had to be performed to remove the catheter. stent placement was used to repair the dissection. patient's condition has improved. prognosis was satisfactory. further information received mentioned that the case was not as severe as initially thought to be. upon cutting the handle of the catheter to release the tension on the tip, the catheter did not immediately become unfolded. after approximately 5 minutes, the catheter did go back to a relaxed position. meanwhile, the physician had already performed an aortic dissection and called for specialist to consult and placed the stent. the information provided mentioned that had they waited a little longer, they could have retrieved the catheter without having to perform the aortic dissection and stenting. >< the carto system was reviewed and there was no discrepancy noted. the system met all specifications upon its release. in addition, the history of customer complaints associated with this system was reviewed as well and there was no customer complaints reported that could have caused / be involved in this reported catheter issue. | 0
1474823 | investigation is pending. >< caller alleged false negative acetaminophen (apap) when testing triage tox screen with another format. caller was performing correlation with another facility. one sample was negative for apap at their facility with dln w44049; a sample was frozen and sent to the other lab and got a positive apap result with w44683. after this result, caller went back to original frozen sample and retested in her lab on dln w44063 and got positive apap. positive control ran fine on original lot. no pt reports of invasive procedures or injury. | 0
1485057 | routine antibody screen analyzed on immucor echo on this pt produced a positive result. when a panel of reagent red cells was tested for antibody identification, all cells were negative with pt plasma. antibody screen was then repeated and equipment reported negative in all 3 reagent screening cells. pt plasma was then tested using gel technology and 2+ anti-fya was identified. when reaction well pictures from cells reported as nonreactive negative by echo were reviewed, there was visible agglutination in cells that were fy-a+-. the immucor echo called positive reactions 'negative' for a pt that had developed an anti-fya. fortunately, additional testing revealed the antibody before the pt was transfused. immucor was notified or misreading, and is investigating the sample. dates of use: 2009. diagnosis or reason for use: abo, rh, antibody detection and identification. | 0
1288623 | nurses are currently manually entering vital signs data into the electronic medical record (emr). nurses have noticed that the mean pressure value that is calculated by the emr is sometimes different from the mean value displayed on the patient monitor. a few sample values were obtained: (systolic/diastolic mean) 129/65 81 emr mean 86; 146/78 95 emr 101; 134/54 81 emr 84; 150/98 116 emr 115. we were unable to verify if this occurs with a different patient monitoring vendor (such as in the er) because in that case the emr does not display the mean value. nurses reporting this problem say that they have seen much larger variations but i do not have specific numbers to report yet.====================== manufacturer response for patient monitor, (brand not provided)======================per clinical support with ge, the mean pressure is measured in the oscillometric method and is the most accurate pressure value since the systolic and diastolic are then calculated from the mean. most emr's take the systolic and diastolic numbers and then calculate the mean using a formula intended for auscultatory readings. therefore there will be variations. | 1
1200195 | it was reported that the centricity web experienced a forced log out midway through a surgical procedure where the surgeon was relying on the cweb images to be able to proceed with the procedure. the approximate time of the delay was reported to be 8 minutes. the forced log out occurred due to a bug creating the cweb force out, rather than cweb session ending at a preset time. centricity web is intended for image review and not for diagnostic use. | 1
1054751 | add'l manufacturer narrative: medical center's inpatient pharmacy contacted baxa in 2008 to report that their magnesium sulfate container ran dry before the compounder application prompted them to change the container. a total of 3ml of magnesium sulfate should have been used for the day's tpn production (container = 50ml). it was also reported that their calcium gluconate containers physical remainder was more than half filled (container = 100ml). a total of 78ml of calcium gluconate should have been used for the day's tpn production. a total of 20 tpn bags were compounded for the day's production using the compounder. the customer reported that all bags were recalled from the floor, and that three of the compounded bags had been infused into neonatal pts for 3-4 hrs. the bags were removed from infusion once the issue was discovered. these three pts are in stable condition and did not exhibit any adverse event as a result of the infusion. the facility ran tests on the three pts to determine if excess levels of magnesium were present in blood samples taken. two of the pts blood work indicated a higher than expected level of magnesium. however, this did not require any add'l medical intervention. the third pt did not have elevated levels of magnesium. to investigate the cause of the reported issue and to provide technical assistance, a site visit was conducted at medical center's inpatient pharmacy the next day by baxa associates. in order to investigate the issue, the following items were evaluated: compounder: the entire sys was isolated, including the disposables and ingredient setup to preserve the setup in use during the previous day's production. results of eval: the physical setup of the compounder was correct. all ingredients and inlets were correctly connected to their associated ports. a visual inspection of the hardware that comprises the compounder showed no physical damage. all 24 valve actuators were aligned to the home position. a visual inspection of the valve set revealed no physical damage. all 24 valve cores under the valve set were aligned to the home position. the sys check (qm-15.02 section 9 procedure) was performed with one baxa and two summerlin observers physically confirming the em2400 performance. all three test formula bags completed successfully (within +/- 3%). the functional testing revealed that the compounder performed as designed. the blackbox data: the computer running the compounder operating software and the compounder itself maintain an ongoing dialogue that is recorded in a 'blackbox' file. this provides an audit trail for reference. results of eval: the blackbox data verified that the user who set up the compounder on the day of the reported incident did bypass the barcode verification step for magnesium sulfate. the hospira brand magnesium sulfate was scanned, an audible alarm was received and the empty vial of american regent brand was scanned to bypass the alarm process. upon the magnesium sulfate container running dry, the user was presented with an audible bubble alarm error. the user acknowledged the error 19 times before aborting the bag run. the user proceeded to manually change the calcium gluconate remainder value from 21.14ml to 60.00ml. then, the magnesium sulfate container was swapped out with a new full container, primed and verified. a complete review of the blackbox data showed that the compounder performed as designed and delivered the volumes of the requested ingredients for the day's production. other observations: it was immediately noticed that vials had been retained for several ingredients that were no longer being acquired (an indication that the barcode verification process was being bypassed). hospira brand magnesium sulfate is the current ingredient in use. however, an empty vial of american regent brand magnesium sulfate was placed inside the hood. the technician staff confirmed in follow-on interviews that they scan the barcode from the american regent brand rather than the hospira brand to bypass the barcode verification process during the setup of the compounder. three add'l empty containers for other ingredients were placed on the shelf above the hood. the technician operating the equipment was observed during the site visit. the technician lacked the proper knowledge to safely use the device. the tech did not know how to power the em2400 down. during the setup process of the compounder, the tech did not use the proper techniques for connecting the inlets to the valve, spiking the source ingredients, hanging the source ingredients and labeling the inlets with their barcodes. based on an eval of the documentation associated with this event, it can be concluded that the most likely cause of this issue is that the user inadvertently connected the inlet (tubing) from the ingredient magnesium sulfate to the ingredient port (attachment at compounder site) calcium gluconate, which resulted in magnesium being pumped when calcium was intended and vice versa, thereby administering magnesium rather than calcium. this conclusion is supported by the blackbox data, the physical remainders in the containers, and the test results for the three pts. two of the pts had elevated levels of magnesium but the third pt did not have elevated levels of magnesium. the first two pts had substantial volumes of calcium in their orders. the third pt had no ordered calcium. the compounder performed as designed and delivered the requested ordered volumes of ingredients. during the site visit, the technician who was operating the compounder on the day of the incident was provided add'l training. the hospira brand barcode for magnesium sulfate was addressed to prevent problems that lead to future failure modes. the importance of the verification process during the setup of the compounder was revisited with the director. to ensure that the inlet is attached to the correct port, they were instructed to physically grab the inlet that is being primed and verified and hold on to it as the ingredient is selected for priming. this will provide a visual verification that the ingredient being primed is actually pumping through the inlet that is being held in hand. to prevent this issue from re-occurring, recommendations were made to the customer that baxa provide their facility routine scheduled training sessions and assist in establishing proper policy and procedure for safe compounding practice. device was evaluated at customer site by baxa personnel. >< in 2008, baxa was notified by the customer of two pt involved incidents using the compounder for tpn production. it was reported to baxa that this pt had received a tpn bag containing a higher-than-desired quantity of magnesium. the bag was infused into the pt for 3-4 hrs and then removed once the issue was discovered. the pts blood work indicated that a higher than expected level of magnesium was present. however, no add'l medical intervention was needed. the pt is in stable condition and did not exhibit any adverse event as a result of the infusion. >< in 2008, baxa was notified by the customer of two pt involved incidents using the em2400 compounder for tpn production. it was reported to baxa that this pt had received a tpn bag containing a higher-than-desired quantity of magnesium. the bag was infused into the pt for 3-4 hrs and then removed once the issue was discovered. the pts blood work indicated that a higher than expected level of magnesium was present. however, no add'l medical intervention was needed. the pt is in stable condition and did not exhibit any adverse event as a result of the infusion. >< add'l manufacturer narrative: medical center's inpatient pharmacy contacted baxa in 2008 to report that their magnesium sulfate container ran dry before the compounder application prompted them to change the container. a total of 3ml of magnesium sulfate should have been used for the day's tpn production (container = 50ml). it was also reported that their calcium gluconate containers physical remainder was more than half filled (container = 100ml). a total of 78ml of calcium gluconate should have been used for the day's tpn production. a total of 20 tpn bags were compounded for the day's production using the compounder. the customer reported that all bags were recalled from the floor, and that three of the compounded bags had been infused into neonatal pts for 3-4 hrs. the bags were removed from infusion once the issue was discovered. these three pts are in stable condition and did not exhibit any adverse event as a result of the infusion. the facility ran tests on the three pts to determine if excess levels of magnesium were present in blood samples taken. two of the pts blood work indicated a higher than expected level of magnesium. however, this did not require any add'l medical intervention. the third pt did not have elevated levels of magnesium. to investigate the cause of the reported issue and to provide technical assistance, a site visit was conducted at medical center's inpatient pharmacy the next day by baxa associates. in order to investigate the issue, the following items were evaluated: compounder: the entire sys was isolated, including the disposables and ingredient setup to preserve the setup in use during the previous day's production. results of eval: the physical setup of the compounder was correct. all ingredients and inlets were correctly connected to their associated ports. a visual inspection of the hardware that comprises the compounder showed no physical damage. all 24 valve actuators were aligned to the home position. a visual inspection of the em2400 valve set revealed no physical damage. all 24 valve cores under the valve set were aligned to the home position. the sys check (qm-15.02 section 9 procedure) was performed with one baxa and two summerlin observers physically confirming the device performance. all three test formula bags completed successfully (within +/- 3%). the functional testing revealed that the compounder performed as designed. the device blackbox data: the computer running the compounder operating software and the compounder itself maintain an ongoing dialogue that is recorded in a 'blackbox' file. this provides an audit trail for reference. results of eval: the blackbox data verified that the user who set up the compounder on the day of the reported incident did bypass the barcode verification step for magnesium sulfate. the hospira brand magnesium sulfate was scanned, an audible alarm was received and the empty vial of american regent brand was scanned to bypass the alarm process. upon the magnesium sulfate container running dry, the user was presented with an audible bubble alarm error. the user acknowledged the error 19 times before aborting the bag run. the user proceeded to manually change the calcium gluconate remainder value from 21.14ml to 60.00ml. then, the magnesium sulfate container was swapped out with a new full container, primed and verified. a complete review of the blackbox data showed that the compounder performed as designed and delivered the volumes of the requested ingredients for the day's production. other observations: it was immediately noticed that vials had been retained for several ingredients that were no longer being acquired (an indication that the barcode verification process was being bypassed). magnesium sulfate is the current ingredient in use. however, an empty vial of another co brand magnesium sulfate was placed inside the hood. the technician staff confirmed in follow-on interviews that they scan the barcode from the american regent brand rather than the hospira brand to bypass the barcode verification process during the setup of the compounder. three add'l empty containers for other ingredients were placed on the shelf above the hood. the technician operating the equipment was observed during the site visit. the technician lacked the proper knowledge to safely use the device. the tech did not know how to power the device down. during the setup process of the compounder, the tech did not use the proper techniques for connecting the inlets to the valve, spiking the source ingredients, hanging the source ingredients and labeling the inlets with their barcodes. based on an eval of the documentation associated with this event, it can be concluded that the most likely cause of this issue is that the user inadvertently connected the inlet (tubing) from the ingredient magnesium sulfate to the ingredient port (attachment at compounder site) calcium gluconate, which resulted in magnesium being pumped when calcium was intended and vice versa, thereby administering magnesium rather than calcium. this conclusion is supported by the blackbox data, the physical remainders in the containers, and the test results for the three pts. two of the pts had elevated levels of magnesium but the third pt did not have elevated levels of magnesium. the first two pts had substantial volumes of calcium in their orders. the third pt had no ordered calcium. the compounder performed as designed and delivered the requested ordered volumes of ingredients. during the site visit, the technician who was operating the em2400 compounder on the day of the incident was provided add'l training. the hospira brand barcode for magnesium sulfate was addressed to prevent problems that lead to future failure modes. the importance of the verification process during the setup of the compounder was revisited with the director. to ensure that the inlet is attached to the correct port, they were instructed to physically grab the inlet that is being primed and verified and hold on to it as the ingredient is selected for priming. this will provide a visual verification that the ingredient being primed is actually pumping through the inlet that is being held in hand. to prevent this issue from re-occurring, recommendations were made to the customer that baxa provide their facility routine scheduled training sessions and assist in establishing proper policy and procedure for safe compounding practice. device was evaluated at customer site by baxa personnel. | 0
1525161 | reported by the complainant as follows: acute lower gi bleed in a male post cabg pt with sepsis, extensive arteriopathy, requiring anticoagulation therapy including plavix pre and post surgery, vasopressors, renal dialysis and iv antibiotics. sigmoidoscopy showed arterial bleed in the lower rectum. >< reported to fda on nov 5, 2009. | 0
1550591 | reported by the complainant as follows... nurse calling to report, a female admitted to icu for coronary artery disease status post open heart surgery had flexiseal fms inserted approximately 1 week ago for loose stools. it was discontinued in 2009 after customer experienced rectal bleeding. lower gi series was completed the same day. gastroenterologist note indicated 'scybalous ulcer related to rectal tube.' nurse was unable to identify exact location of the ulcer. unable to obtain medications; but she is on anticoagulants. >< . | 0
1459516 | fetal monitoring performance problems have been identified by our clinical team. these have been described as:1. problems with tracing early gestation fetuses, multiples (twins, triplets). requiring separate fetal monitors to trace each fetus.2. weak fhr signal3. loss of fhr signal4. doubling/halving of fhr5. monitor volume at maximum level is still not adequate6. artifacts present as decelerations7. fhr tracing is different from actual audible fetal heart rate signalpatients are in active labor being monitored with external ultrasound fetal monitors. staff are concerned that the monitors are not projecting an accurate fetal heart rate.philips stated that they have designed a new algorithm that is currently being tested to fix the monitor incorrect signals issue. | 1
1181093 | the facility is reporting that at the beginning of an rc repair, the fms pump was not measuring pressure flow rates correctly and the outflow was not working. as a result, the patient's shoulder became exceedingly swollen. because of this, the surgeon became concerned over the probability of 'compartment syndrome' occurring, if he continued on. at this point the surgeon abandoned repairing the rotator cuff and instead performed a debridement, a subacromial decompression. the surgeon believes that this alternative means of relieving the pt's issue may work. this procedure was concluded successfully without further issue or harm to the pt. during recovery, the facility used a drain pump as a precaution. the pt was discharged under normal circumstances without issue. >< mitek is at this point in time, awaiting receipt of the complaint device towards conducting a root cause failure analysis, the results of which will be the subject of a follow-up report. | 0
1437211 | lot history record review: the reported lot number was reviewed for any abnormalities, which may have contributed to the complaint. nothing was found that would have contributed to the reported complaint. review of returned sample: the complaint sample is not available for eval. the physician agreed to send photos of the sample, but they have not yet been received. conclusion: the complaint investigation is inconclusive. the product was not returned for eval. the cause of the complaint appears to be user error. based on the info supplied by the physician, the system was not correctly set up for the case, introducing room air into the pt. a review of the mfg records indicated that all of the device spec and quality requirements were satisfied. the instructions for use state the following: the instructions for use (ic 129) states: indications: the administration of saline and contrast, through angiographic catheters. cautions: all air must be removed from the system. turn the handle to open position and allow the pt's blood pressure to back-bleed and remove the air from the fluid path. the quality dept. has been made aware of the physician's request for add'l labeling. this type of complaint will continue to be monitored for trends. no further action at this time. >< dr left a message stating that he had called to discuss the angioflush iii having serious design flaws to report to you (angiodynamics), that may have resulted in a poor outcome to a pt. during a f/u phone conversation, the physician stated the tubing was hooked up backwards, and room air was drawn in instead of the co2, resulting in a pt death. the physician feels there should be tags to mark all 4 ports. | 0
1234664 | during needle local procedures using the homer mammalok gold biopsy needle, images may become 'black' (void of image data) when the needle is placed in the breast. this anomaly required the technologist to reposition the patient and acquire another image. in 2008, in a conversation with a siemens clinical applications specialist, it was discussed that in rare cases a patient may need to have a needle removed and reinserted to acquire a clean image. on the same day in a conversation with the siemens service engineer, it was discovered that the site, on at least one occasion, had to decompress the breast, reposition the patient and recompress the breast again while a biopsy needle remained in the patient. it was also discovered that on at least one occasion, the site has to remove a place biopsy needle, reposition the patient and reinsert the needle. on the next day, in a conversation with a siemens clinical application specialist, it was discovered that another site has experienced similar results with the homer mammalok gold biopsy needle. >< at the time of this incident, it was determined not to be a reportable event as we were unaware of any harm to the patient. upon investigation, it was discovered that the site had, on at least one occasion, removed a biopsy needle and reinserted it in the patient. upon becoming aware of this information, it was determined to be a reportable event. in addition to mdr #2240869-2008-0025, siemens healthcare has also submitted an mdr for the second site we have become aware of experiencing the same results. hospital center. mdr #2240869-2008-0026. siemen healthcare has notified the manufacturer of the mammalok gold biopsy needle of this incident via certified mail as well as a copy of both mdr reports. result: needle biopsy guide. the site was provided with another manufacturer's needle. to date, we have not received any information from the site that the issue has reoccurred. | 1
1386101 | evaluation in progress, but not concluded. >< during preventative maintenance of the device, the technician observed the arterial monitor module did display temperatures as expected. the arterial monitor provides timers and temperature and pressure monitoring displays. since the event took place during preventative maintenance, there was no adverse consequence to a patient as a result of this event. | 1
1467987 | evaluation in progress but not concluded. >< during use of the device for a cardiopulmonary bypass procedure, the user reported the arterial monitor would intermittently reboot. the device was used to conclude the procedure. the user reported the surgical procedure was completed successfully, and there were no adverse consequences to the patient as a result of this event. | 1
1381655 | during preventative maintenance of the device, the central control monitor of the heart lung console had a flickering appearance. the technician installed an alternate device. since the event took place during preventative maintenance, there was no adverse consequence to a patient as a result of this event. >< evaluation in progress, but not concluded. | 1
1381656 | evaluation in progress, but not concluded. >< during preparation for use for a cardiopulmonary bypass procedure, the central control monitor of the heart lung console did not start up when the device was turned on. the user reported the surgical procedure was completed successfully. since the event took place prior to the onset of cardiopulmonary bypass, there was no adverse consequence to a patient as a result of this event. | 1
1418577 | after use of the device for a cardiopulmonary bypass procedure, the light intensity of the central control monitor of the heart lung console began to vary. the pixels were intermittent in nature on initial power up of the device. since the event took place after the cardiopulmonary bypass procedure, there was no adverse consequence to the patient as a result of this event. >< evaluation in progress, but not concluded. | 1
1420016 | the user reported, while attempting to tighten the pivot screw on the central control monitor, the mounting arm broke off. no other details regarding the nature of the event were provided. there was no reported adverse consequence to the pt as a result of this event. >< evaluation in progress, but not concluded. | 0
1496526 | during preparation for use of the device for a cardiopulmonary bypass procedure, the user reported the central control monitor would not completely boot up. the user rebooted the monitor and it began to function as expected. the device was used for the procedure. the user reported the surgical procedure was completed successfully, and there was no adverse consequence to a patient as a result of this event. >< evaluation is progress, but not yet concluded. | 1
1460056 | during installation of the device, the user reported the safety monitor was causing the connection to the cardioplegia and arterial monitors to reset. no other details regarding the event were provided. since the event occurred during installation of the device, there were no pt involvement during this event. >< evaluation in progress, but not concluded. | 1
1359066 | a baby monitored with stan s31 was born with metabolic acidosis defined by umbilical cord arterial blood sample. a first review of the case indicated that the labor was allowed to continue too long with non-reassuring fetal heart rate (fhr), which is in clear violation of the stan s31 guidelines. there are currently no indication for malfunction in the device, neither faulty labelling nor weakness in methodology. neoventa has trained the clinicians at the hosp. a f/u report will be submitted as soon as a more thorough investigation has been completed. newborn was discharged from the hosp alive. | 1
998907 | the customer stated that the calibration slope of potassium was low using the clinical chemistry ict calibrator, lot # 0505017. there was no impact to pt management reported. >< this is an initial report. an investigation is in process. a final report will be submitted when the investigation is complete. | 0
1045195 | the complainant reported that the inflation pressure monitor quit working several times throughout the day. it happened in prep, during procedure and post procedure. there was no harm or injury reported to the pts. >< the subject device is being returned to merit for eval. a f/u report will be filed when the eval is complete. | 0
1204823 | as reported by a dr, a pt developed canaliculitis after insertion of an oasis medical, inc. form fit hydrogel canalicular plug. reported observation: pt developed canaliculitis on the right upper lid after insertion of an oasis medical, inc. form fit hydrogel canalicular plug. product lot number: lh1006c. product inserted in 2007. date of complication: 2008. reported to oasis medical, inc., on 09/18/2008. >< a review of batch production and sterilization records shows that there were no remarkable events, anomalies or nonconformance's during production of form fit hydrogel canalicular plug, reference lot number lh1006c. a total of 2,016 form fit hydrogel canalicular plugs, from lot number lh1006c have been distributed. this is the only customer who has reported any adverse events. conclusion: as a result of investigation results, oasis medical cannot definitively determine that form fit hydrogel canalicular plug lot number lh1006c was directly or indirectly responsible for the reported observation of canaliculitis. issues have been resolved at this time for the pt. | 0
1473332 | camino intracranial pressure monitor has only a two hour battery life. the first indication of a low battery is a continuous screeching alarm and the machine stops functioning. there is no battery meter or alarm to indicate the battery is approaching a depleted state. device takes up to 9 hours to recharge the battery and there is no indication that the battery is fully re-charged. | 0
1359170 | boston scientific corporation (bsc) was informed by a nurse that a spark flew out of the table side controller of the intravascular ultrasound imaging system (ilab) and left a nickel-size burn mark on the side of her thigh. she could feel something warm at that time and saw an arc of electricity coming out of the table side controller. the event occurred when the nurse was getting the patient off the table and her leg was pressed up against the table. no treatment was required and the nurse is fine and it was reported that she did not consult a physician for treatment nor take any medication for the burn. the patient was not affected by the event and did not have an intravascular ultrasound (ivus) procedure with the ilab system prior to or during this event. the ilab system was on but had not been used that day. the ilab system was shut down since then, and the table side controller, imaging pc and acquisition pc of the intravascular ultrasound imaging system have been replaced. >< a. patient information: no patient involvement; the incident occurred with the nurse. evaluation code. rustles: other: the ilab system was fully functional in the operational test and met all specifications in the functionality test. the device history record review has been completed and no additional data relating to the reported problem could be found. also no additional past events in relation to the complaint of this unit were found. an investigation was conducted on the returned table side controller with cable, acquisition pc, and the image pc. there was no physical damage observed and it was determined that the products performed within specifications. various voltage and current measurements were performed, including an ilab system electrical safety test. no abnormal conditions was found. the system was fully functional in the operational test and met all specifications in the functionality test. | 1
1197623 | smoke came out of the back of the system when it was turned on. >< the system and materials have been designed and tested to iec 60601-1 to ensure that the possibility of a fire occurring is very remote. although, there was no injury that occurred due to this malfunction, there is the possibility that if it were to recur, a negative impact to other equipment and/or materials in the room (i.e. compressed gas tanks, flammable materials, ect.) could result in harm to the pt/user. this report is being sent as a notification. | 1
1354423 | actions taken: site reporting the problem has been contacted with information about the issue, and informed of the temporary solution to use to prevent the problem. other customers who may be affected have been notified of the issue and temporary solution. work on the patch of software has begun. actions to be taken: a letter describing this issue and resolution will be sent to the customer base affected by this issue. the software patch will be made available for download from our website, as per our established procedures. data innovations will assist any customer with questions or needing assistance with applying the patch. eval summary: the site reporting this issue to data innovations discovered it when manually re-running tests but not having those manually re-run tests subsequently be held by the rule that they wrote in the instrument manager software. the instrument manager software was not returned to us, but data innovations evaluated data returned to us by the site reporting the issue. specifically, communications trace, specimen event log, audit trail and specimen tracking information was gathered from the site. review of this information assisted in the investigation of the problem as reported by the site. after investigation, it was possible to reproduce the problem at data innovations and the cause of the problem was identified. the investigation also found a method of operation that will ensure that this issue does not recur for sites. this temporary work around was communicated to customers until such time as a software patch can be released. >< issue: the 'any test is rerun' data element may not trigger a rule to fire even though a test is marked as a rerun in the database. issue details: the 'any test is rerun' data element may not properly evaluate a test as being marked as a rerun in the sm database. this problem will manifest itself depending on the order in which results are received from the instrument and the order in which the tests were added to the sm database. for this reason, the problem is more likely to occur than not. | 0
1422468 | vamp on arterial line tubing disconnected with minimal tension on line. pt's alarm sounded and rn to bedside within minutes as pt bleeding from artery. no harm to patient. | 0
976130 | a customer site reported waste fluid had overflowed from the waste carboy on a benchmark xt instrument onto the floor without an alarm. no one was injured and pt care was not affected. the field service engineer eval the upgraded waste sensor and found the tube to be broken from mishandling preventing the sensor from functioning properly. the fse (field service engineer) installed a newer designed waste sensor which has the tube counter sunk preventing lateral movement which prevents the customer from breaking the tube if mishandled. while there is always a possibility that a slip and fall event due to a fluid spill could result in death or serious injury, all evidence indicates that the possibility of death or serious injury is remote. >< the most probable cause for the broken tube is the customer mishandling the waste sensor. this is a known issue and a newer designed waste sensor was installed on this instrument in late 2007 to prevent this from happening in the future. because of the remote possibility of injury resulting from a slip and fall event caused by waste fluid overflow, the firm is in the process of implementing an improved waste sensor design as a replacement to the current sensor. this new waste sensor design could eliminate the observed failure mode. complete field implementation of the new sensor design to be completed by 10/31/2007. the newer waste sensor design was introduced to the mfg floor on 02/08/2007. a final report will be sent pending the completion of the field implementation of the new waste sensor. | 0
1427928 | when multiple users simultaneously are adding patient and request, it is found that some patients are getting the wrong requests attached. there was no reported patient injury associated with this event. >< this mdr is being submitted late as this is part of our ongoing quality improvement activities. we have updated the risk assessment control record and the update had resulted in the assessment that this reported complaint meets the reportability criteria as outlined in the risk assessment control record for ge healthcare. code fix released in patch installed 28-jun-2007. | 1
1046132 | the customer experienced a problem when merging a field in field plan with the new eclipse algorithm aaa v8 1.17. when we try to merge a field in field plan with an mlc leaf travel of greater than 10cm, the mlc leaf positions for great than 10cm of travel distances are wrong in the merged plan, but still calculated for dose without any error message appearing. this was noticed on the qa plan and did not make it to treatment. >< method: software tests performed in our development facility in other country, on the same version of software as installed at the customer site. results: the root cause was a software design error. conclusions: software contributed to the event. a product notification letter will be sent to affected customers. all corrective action will be reported under the guidelines of 21 cfr part 806. no f/u reports are anticipated. | 1
1404772 | during an endodontic procedure, the irrigating syringe being used leaked sodium hypochlorite from the 'back' end of the syringe. the same type of syringe was used multiple times with a high rate of failure. it appears that the rubber tip on the plunger did not fit properly and failed to create a seal allowing liquid material to escape. | 0
1550328 | according to the user, tests were performed by the user facility's biomed dept on spo2 cable and nothing was found to be wrong with it, and it appeared to be working fine. the user returned both the spo2 cable and finger grip to the device manufacturer for additional evaluation. the device manufacturer visually examined both the spo2 cable and finger grip, and they were found to be normal. the finger grip was also turned inside out and examined under a microscope. there was no evidence of substance(s) that would have caused or contributed to a pt burn, such as metal flakes. the spo2 cable did not show evidence of overheating, and no metal particles/flakes were observed on the outside of the fiber optic ends or jacket. both the spo2 cable and finger grip were tested using the same model type monitor as that at the user facility. both the spo2 cable and finger grip functioned as expected during this testing. when tested using the returned spo2 cable and finger grip, the pt monitor displayed a normal spo2 waveform with 96% o2 and 78 bpm with a human finger inserted into the finger grip. the spo2 waveform produced by the pt monitor with the returned cable and finger grip was clean, with no noticeable anomalies. the device manufacturer asked the user if they thought the actual pt monitor malfunctioned. the user stated that they did not think the pt monitor malfunctioned and that they were only concerned about the spo2 cable and finger grip. the device manufacturer offered to send a field service rep to the user facility to evaluate the pt monitor. to date, the device manufacturer has not been notified of the user facility's desire to have the pt monitor evaluated. per the user, they continued to use the pt monitor, the spo2 cable and the spo2 finger grip that were in use during the reported incident, on pts after the reported incident. there is no evidence that the spo2 cable and finger grip caused or contributed to the pt's injury given the device manufacturer's evaluation of the returned product, along with the fact that the fiber optic cable and finger grip are nonconductive. it was reported that all of the pt's fingers (including the finger where the spo2 finger sensor was positioned) and part of the pt's palm was burned. this description of the pt's injury does not correlate with one that would have been caused by the pt monitor or an accessory of the pt monitor, and is in conflict with the initial report. the device manufacturer has requested copies of pictures that are available of the pt's injury. the user facility's risk manager has advised that the pt's family is not ready to release pictures of the pt's injury to the device manufacturer. a supplemental report will be filed if additional info is provided by the user facility or the pt's family, which impacts this report. >< it was reported to the device manufacturer that a pt was undergoing a head and cervical spine mri scan. and that spo2 was hooked up for monitoring. after the scan was complete, the user claimed that the pt immediately complained of burns on the hand where the spo2 finger grip was attached. | 0
1386048 | in late 2008, abbott point of care was contacted by a customer who reported a martel printer for the i-stat1 analyzer was hot to touch. there was no report of any injury associated with the complaint. | 0
1373108 | narrative: ge healthcare performed a checkout of the monitor and found it to function within manufacturer's specifications. ge healthcare requested a copy of the monitor tracings in order to perform a clinical evaluation of the alleged event, however, the request was denied by the hospital. no further investigation of the alleged event can take place until ge healthcare is provided a copy of the tracings. >< customer reported unit provided a maternal and fetal heart rate tracing throughout labor. pt assigned apgars of 0 at one minute and 0 at five minutes. resuscitation efforts including intubation and epinephrine administration were reportedly taken, but were unsuccessful. | 0
1477324 | since the purchase (on 10/16/2007) of multiple ge/corometrics 259cx maternal/fetal monitors, there have been innumerable intermittent problems related to the data entry keyboard (model 2116b). the buffering system built into the fetal monitor to handle communications from the keyboard (which are then sent to a paper strip printed on the 259cx) is overloaded easily, causing the printer to fail, as well as causing display failures (spo2 & nibp measurements fail to display). this intermittent failure occurs most often when multiple parameters are attempting to print at the same time (i.e. automated vs keyboard entries). ge/corometrics have only recently (a few weeks ago) admitted that they will not repair the problem (for the last 2 years, they have led us to believe that a repair was being developed for the failure). this problem only occurs with the 250 series monitors - we have used the older models without failure (116/120 series). the data entry in concern includes point-of-care annotations during bedside care (medications given, pt condition, etc). no injuries have occurred, but the failure occurs most during critical care (i.e. epidural administration), and places the [unmonitored] pt at risk: continuous vitals monitoring may be lost, and restarting the 259cx (to clear the failure) distracts the nurse from their pt care at critical times. | 1
1515362 | during pull back of pigtail across aortic valve, the waveform change on the patient monitor of invasive blood pressure had an approximately five (5) second delay or lag in display as compared to actual physical location change of the catheter. this monitor 'latency' has been noted on ecg as well. | 1
1380646 | user facility reported to the manufacturer that they were not receiving any alarm tones from the device. >< according to the initial reporter (the user facility's clinical engineer), the user facility first tried replacing the device's power supply board, but this did not rectify the no sound condition. per discussions held with the user facility's clinical engineer, the user facility decided to scrap the device due to the cost of replacing additional boards or sending it in for repair. as a result, there is no opportunity for the manufacturer to evaluate the device in question in order to confirm that an actual malfunction had occurred or to determine the cause of the reported problem. while the device display provides visual indicators of alarms occurring, the device labeling does not represent that display data is continuously monitored by staff. in the presence of life-threatening events when no sound is audible, and no staff member is continuously monitoring the display, serious injury and/or death can occur. | 1
1362990 | the pins were bent on the connector on the philips mms module. the bent pins cause the module to be inoperable. the failure of these pins presents the potential for patient safety issues, as health care providers need access to immediate monitoring of unstable patients that are brought to critical care areas. the biomedical technician replaced the necessary components, tested the device, and returned it to service. the device is available for onsite evaluation at our facility. questions and comments should be directed to the biomedical director. | 1
1362993 | the pins were bent on the connector on the philips mms module. the bent pins cause the module to be inoperable. the failure of these pins presents the potential for patient safety issues, as health care providers need access to immediate monitoring of unstable patients that are brought to critical care areas. the biomedical technician replaced the necessary components, tested the device, and returned it to service. the device is available for onsite evaluation at our facility. questions and comments should be directed to the biomedical director. | 1
1263285 | my cardiologist wanted my heart monitored with an outpatient device for 2 weeks. his office ordered a mobile outpatient cardiac device from cardionet. the device arrived in unacceptable condition: a metal base that is electrical has ripped 'rubber' layer underneath it, which is only partially glued to the metal base, allowing electrical wires to be exposed; various components are dirty, having hair and other debris on them; even the enclosed booklet was soiled. at first, i tried to clean the components with alcohol swabs, but when i got to the electrical piece with the partially exposed wires, i decided to call the co and not use this particular device. the components are packed in soft foam, and i began to wonder how clean it was if the various cords and metal pieces had debris from other pts on them. i called the co to report the specifics of the problem and stated that i would only accept a new, clean monitoring device from them, and to date, their only calls have been to find out if i plan to use the device they sent. they have expressed no care whatsoever for th dirty and defective product that they mailed me. today, they called the cardiologist's office, apparently not to express concern about the qual of their prod or my health, but to find out about my using or returning this device. i am hoping that the fda will take it from me quickly, so it does not go back into circulation and get sent to another customer. reason for use: tachycardia. | 1
1200323 | it was reported that the mri2 monitor was drawn into an mri system. the mri system was not being used on patients at the time. no injury was reported. | 1
1327591 | pt was admitted to the micu when several attempts were made to transfer the patient to the solar monitor using the pdm (tram module) and the bedside pdm docking station. each attempt was unsuccessful. additional pdms were used without success. biomed found that the eport (connector on the rear of the pdm module) on the pdms were bent and the eport connector on the pdm docking station was also bent therefore the pdm could not connect. this critical pt. had to be monitored with a defibrillator until biomed got another pdm that did not have a bent eport and the pdm docking station had to be replaced. other equipment was used to monitor the pt hemodynamics during the wait for a new pdm and pdm docking station. the eport connector on the pdm as well as the pdm docking station's eport connector is a poor design and ge has been made aware via itrack which occurs every time we encounter this problem and notify ge. this is not the first problem we have had with the pdm and the eport connectors. ge has been replacing all the pdms with bent or broken eport connectors, but something rather than just replacing parts must be done about the design of the pdm. we have also noticed this problem with the pdm and pdm docking station on the transport pro monitors.====================== manufacturer response for module, physiological monitor, pdm, tram, monitor, physiological transport pro and solar======================ge has replied to emails stating that they have forwarded our concerns to their headquarters. a ge nurse has contacted us to find out about the events involving this patient. however we have had no follow up on how ge plans to fix the design on the pdm's eport/ eport connectors and docking station. | 1
1563681 | a patient with multiple medical problems had a tracheotomy tube and had a telemetry pulse oximeter monitor which was connected to the central monitor in the nursing station. the patient was found unresponsive with the trach out. investigation revealed that the staff does not recall hearing the central monitor desaturation alarm. the alarm has settings of 1-10 and was found to be set at '1'. engineering staff can set a minimal alarm level that staff can't go below. as a result, the central monitor minimum alarm volume setting was increased for the involved unit. in addition the central monitor minimum alarm volume setting for each unit in the hospital was assessed and confirmed to be at an appropriate level based on each unit's configuration. | 1
1321464 | heart monitor found with no ability to sound alarms. biomed found that metal wire securing speaker had shifted from intended resting site to a site allowing for direct contact with the speaker connection contacts. this caused the alarm's sounds to short out. | 1
1315637 | approximately two months ago we began experiencing 'no-comm' issues with ge wireless dash 4000 monitors running on an 802.11b network. the main software = 6.4, main boot software = 6.3, dash boot software = 6.0, ecg sw = 1a, spo2 = 1599v4.0.0.1. we have been working on a daily basis with ge and our own is networking team. as many as 50% of the monitors have experienced communications issues. the symptom of 'no-comm' causes the data transfer between the bedside monitor and the nurse's central station to lose monitoring capabilities as shown at the central station. at that point only the bedside monitor can alarm and display physiological information. the following is a list of troubleshooting already performed. 1) sent log files to ge for december and january. 2) opened a ge i-trac to dedicate ge engineers and technicians to this problem. 3) replaced the central station in attempts to verify a hardware issue. 4) set up a central station complete with monitors and simulators in a test environment to troubleshoot hardware and is networking. 5) sent several more log files to ge from the grant 6 respiratory dept who have experienced (and made us aware) of a very large number of no-comm (no communications) situations that have affected patient care. 6) received reports from ge engineering who recommended several hardware and networking settings to be verify. 7) coordinated several teleconferences between ge, clinical engineering and is networking groups. 8) set up an alternative switch / hardwired network in a test environment to determine if any hardwired communications or data is being affected. no hardwired devices have experienced any issues. 9) moved 'no-comm' dash monitors to other areas to see if the problem followed them for network vs. monitor troubleshooting. 10) fully tested the 'no-comm' monitors in a test environment which confirmed the problems. 11) sent more log files of 'no-comm' monitors as verified in a test environment to ge for analysis. 12) have now determined that 'no-comm' problem is spread across all wireless dash areas. 13) have initiated a po to secure ge engineers to troubleshoot on site networking and wireless dash configurations. | 1
1511581 | . >< according to reporter, the user was hospitalized with hypoglycemia with the product in use. in was reported that the product's glucose readings were 200 mg/dl higher than the user's other meter. user was treated with oral glucose and temporary removal of pump for use. user was released from hospital care in 2009. | 0
1305987 | this is one of three systems having the same issue. the hemodynamic monitoring system has a known software problem that causes the machine to return to pre-edited measurements after a few minutes. staff need to make notes on print outs and dictate correct results. philips updated the software in september to: 1.1.1.1291 and again in january 2009 to version 1.2.1474 to resolve the latest issue. | 1
1543769 | patient's information is not available at this time. >< a customer reported that while using the m-nestpr module (which measures non-invasive blood pressure (nibp), electrocardiogram (ecg), pulse oximetry (spo2), temperature, and respiration), the nibp cuff did not deflate while on a patient's arm. the patient's arm was mildly abraded, red, and swollen, and the patient complained of numbness in the arm. the patient was reportedly discharged from the hospital with no physiological issues related to the event. the biomed tested the module following the event, and was unable to reproduce the reported issue. the module was then sent to ge healthcare's service center for testing, however, the issue was still not repeatable. the customer has received a new m-nestpr module. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 1
1244670 | spacelabs evaluated the pt monitor on-site. the parameter module in use at the time of the reported malfunction had been removed from the monitor and was unavailable. the monitor was evaluated with a replacement parameter module and found to perform to specs. spacelabs has requested that the device involved in the incident be returned to our facility for a more detailed investigation. spacelabs is waiting to receive the device. >< in 2008, spacelabs received a report from the hospital that a monitor shut down and lost all trend data. the staff commented that this has happened before. | 1
1430028 | it was reported that this unit turned off in the middle of a procedure. patient monitoring was resumed on another identical atlas monitor. there was no impact on patient care. an attempt was made to retrieve patient information, but the customer did not have any details. >< the device has been received, but additional time is required to complete the analysis. | 1
1297585 | it was reported that a pacer pt was in a room with a possibly closed door. the pt was found expired and fallen out of bed with two leads off. the healthcare staff allegedly did not hear an alarm on the system. per onsite evaluation by the hospital biomed, the icentral did not have a display with internal speakers; rather it had an external active pc speaker. the system was checked and found to have performed as designed. it is noted that the sound volume of the external speakers was turned down to a very low audible level and the alarm level was 'iso2' (beep each ~ 20 sec). the user controls the volume on the external speakers. ge does not provide external speakers for icentral systems. the display model at the customer site is not delivered from ge healthcare finland. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 1
1275307 | in 2008 an email from life systems was brought to my attention. in this email requesting routine rma's for service returns was a notation that a prodigy ii vital signs monitor had fractured the arm of the patient at golden living center in murraysville, president of omron colin medical corp. immediately contacted ms. amanda hurs, director of nursing at the facility. she was not present during the alleged incident, but described the situation as follows: 'we had a female who was hypertensive during the evening. one of my nurses was using the prodigy ii to obtain blood pressures and after multiple attempts where the machine was re-inflating, we were not able to obtain a blood pressure. the blood pressure was taken manually and was determined to be 200/98. later on that night, the patient complained that her arm and shoulder hurt. we observed a bruise on her right arm and x-rays were taken. a hairline fracture was observed on her right humerus bone along with swelling in the shoulder.' the monitor was returned to omron colin medical corp. and a thorough evaluation was performed by the engineering department. functional testing included both qc tests and a round of overpressure tests. the monitor passed all functional tests regarding overpressure, bp, temperature, and spo2 readings. all results showed that the monitor will properly perform when overpressure release according to aami standards during a blood pressure measurement. the engineering report is available for review. >< summary and conclusions: this monitor, is able to accurately measure blood pressures, temperatures, and spo2 readings. this was verified by our qc testing. the monitor also passed overpressure testing, releasing pressure appropriately and as expected. the manufacturing anomalies found have no effect on monitor performance or measurement results. these were only included for completeness. examination showed that the monitor has not received proper maintenance once outside the control of oha. because of the lack of maintenance and possibly improper maintenance, the monitor ahs fallen into disrepair. battery replacement by a non-oha entity resulted in mismatched and missing screws and exposed the inside of the unit to an uncontrolled environment. the cuff button was treated roughly and allowed to get to the point of non- functionality. this forced the user to use only interval mode to take bp measurements. a temperature probe tip was broken off inside the monitor and not retrieved immediately. the bad cuff button or broken temperature probe tip inside the unit should have been enough for a normal user to know that the monitor needed service, but the monitor continued in service. in conclusion, this monitor measures appropriately and safely according to our established qc procedures. | 1
1310498 | the pump has not been returned to animas for evaluation. animas has conducted a review of the device history record for this pump and it was operating within required specifications at the time of release. pump settings and delivery history were reviewed with the patient's certified diabetes educator and confirmed as accurate. insulin pump therapy was resumed during the hospitalization and the patient has continued to use the pump with no reported subsequent events. if the device is returned, an evaluation shall be completed and a supplemental report will be filed. no conclusions can be made at this time. >< (b)(4). device evaluation: the pump has been returned and evaluated by product analysis on (b)(4) 2011 with the following findings: a review of the total daily dose history from (b)(6) 2011 to the end of pump use on (b)(6) 2011 indicated that insulin delivery totals correctly reflected programmed values. there were no alarms or conditions that would indicate a pump malfunction noted in the pump history. evaluation revealed a cracked battery compartment. a battery compartment leak and moisture damage inside the battery compartment were observed. testing could not be adequately completed due to the moisture damage. >< the patient was hospitalized for elevated blood glucose levels and dka. | 0
1346801 | nurse reported that she administered sage oral care at 13:00. subsequent to that she noted chemical irritation to skin on right wrist and thumb. she stated she backtracked and discovered it was from oral care solution. she assessed patient and noted patient's tongue and lips were white from reaction to treatment. patient denied any pain and package was retrieved from trash. skin conditions of both patient and staff nurse resolved to normal by the next day. no harm to either individual. | 0
1058978 | the suggested use of the device is to leave the warning message turned on if a user opts to mark all studies as read when the anchor study is marked as read. however, some users do opt to turn it off because their workflow dictates that they must review all series in a study prior to marking the anchor as read. if the workflow at a site does not dictate this, the message is recommended to be left on in order to mitigate the hazard. for this particular user's site, an enhancement is being added to the software that allows the user to set the function to never mark all studies as read if the anchor is marked as read. this will be an individualized setting per user. >< the evms pacs has a feature that allows users to mark an anchor study as read, and then marks all other series in that study/correlated studies as read. this allows users to view all cases related to a study and close all as read for enhanced workflow. a warning message is provided to the user that all studies are being marked as read and the user must place a check next to each series and click 'ok' prior to closing in order to mitigate the hazard that they would be marking series as 'read' that they had not read. this feature can be turned on or off; the default is to mark all as read with the warning message turned on. this also appears in the labeling. a user called in a complaint that their series/correlated studies have been marked as read without them being aware of it. the log file showed that the warning message had been turned off using that user's login and password; however, the user was not aware that they changed it. as a result, the user missed 2 fractures on an initial read, but the er physician was aware that there were fractures so no adverse outcomes were noted. | 1
1296109 | it was reported that the customer merged the incorrect data and rejected the original images after the merge, however, the customer did not correct the dicom tags on the rejected images. the node for e-firm (non-ge remote workstation) had the permission to request rejected images. customer queried e-film for images and viewed rejected images with incorrect tags. surgery was performed on the wrong patient due to incorrect image merge. the patient outcome is unknown. >< the situation was caused by an inadequate configuration of the ge healthcare system in connection with a third party device. the setting 'request c-move from pacs (retrieve) for rejected images' has been disabled for this specific e-film workstation. a follow-up report will be filed if additional information becomes available. | 1
1424768 | a follow-up report will be filed when the investigation is complete. >< images scanned at modality are showing up flipped in pacs. there was no reported patient injury associated with this event. | 1
1181056 | the company service representative determined that the tim transceiver power supply had failed. she replaced the transceiver. the system was tested to factory specifications. it functioned normally and was returned to use. >< the customer reported that while three panorama central stations were in use monitoring patients along with spectrums, passport 2's, and ambulatory telepacks, the central stations stopped communicating. no patient injury was reported. (b) (4). | 1
1295331 | customer reported the inner tube is detaching from the outer tube, at the machine end of the circuit. there was no reported patient injury. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 0
1548003 | it was reported from a philips clinical specialist (cs) that a pt, who was a dnr, had expired during the night and the pt's history was saved on discharged and was located in the transfer list. at this customer's request, the cs went into the database server (dbs) configuration to change the unit name for a device and subsequently, staff informed philips that this saved data was no longer available for them to document in the pt's record. when the cs entered the dbs configuration and changed the unit name, existing saved history in the transfer list was lost. no device malfunction occurred, and there is no allegation that the device was a factor in the death. changes made in configuration mode to some settings can lead to a loss of saved data due to the structure of the database. the customer was informed of this occurrence. this data loss is not a health risk because it does no impact real-time monitoring. no further action or investigation is warranted. >< the customer reported that a pt death occurred and that the saved data was lost. | 1
1207868 | we are still completing our evaluation of this report. a follow up report will be submitted as soon as we complete the investigation. >< the user reported that when setting the lower alarm threshold above the current cf parameter, there was no alarm at all. | 1
1392791 | three speaker failures on three philips vm-6 bedside monitors on units with same serial number sequence within a period of 8 months. the failure results in a no audio condition. without audio, the device cannot alert medical staff to a potentially life threatening pt condition. no pt injury has occurred, however, a repetitive failure trend has been noted. philips medical was first notified march 26 2009 and again on may 11 2009. another speaker failure was repaired in 2008. a failed speaker was sent to philips from the march failure to be analyzed. | 1
1490812 | draeger is still investigating the reported incident. a follow-up report will be submitted as soon as the investigation has been completed. >< it was reported that a pt was monitored by an sc7000 monitor in the morning. a nurse discovered the pt deceased on the same day at 11:45am. the nurse says that the monitor was switched off. the user says that the monitor has been switched off for one hour, and therefore, the pt was not monitored. | 1
1062302 | our eval confirmed that the top cover of the devices actually separated from the base as reported. we also found tht one of the units was modified (top cover glued to base) by the customer. this is a known supplier issue and we have implemented supplier corrective action and initiated a field correction. the device identified in this incident falls in the range of the potentially defective devices of the field correction. >< draeger medical systems, inc. has received info from a customer who has reported that two of our infinity docking stations (ids) units were found to have separated (top cover from the base component). no pt injury was reported. note: this is the same report as 1220063-2008-00003 (copy attached) except for the second device with that customer reported in the same complaint. | 1
1196649 | the field service engineer (fse) was performing an installation and upon opening the item from its packaging, the device broke apart. >< the device in question is a mounting plate for the intellivue pt monitor. the device failure is being considered as having the risk of the monitor falling from the mount. philips is in the process of obtaining additional info concerning this event. the complaint is still under investigation. a final report will be sent once the investigation is complete. | 1
1369518 | the customer reported that while the panorama central station was in use monitoring pts along with ambulatory telepacks, the central station display went to a black screen. the pts were switched to another central and monitoring was continued. no pt injury was reported. >< the company service rep replace the display power supply. the system was tested to factory specifications. it functioned normally and was returned to use. | 1
1424517 | the customer reported that while the panorama was in use monitoring patients along with bedside monitors and ambulatory telepacks, the network went down. no patient injury was reported. >< the company service representative replaced the power supply in the tim transceiver. the system was tested to factory specifications. it functioned normally and was returned to use. | 1
1424518 | the customer reported that while the panorama central station was in use monitoring patients along with ambulatory telepacks, the central station display froze. there was no response to key requests. no patient injury was reported. >< the company service representative cleared the fault logs, reset the system, and performed a power cycle. the cause of the freeze was not determined. the system was tested to factory specifications. it functioned normally and was returned to use. they have not had any further problems with the central as of the date of this report. | 1
1369486 | the customer reported that while the panorama central station was in use monitoring patients along with ambulatory telepacks, the system experienced communication loss on all channels. no patient injury was reported. >< the company service representative rebooted the cpu to restore communication. he found that two of the two repeaters had faults, which could cause a large dropout. he reset the repeaters. the system error logs were returned to the factory for evaluation by patient monitoring engineering. there is no indication in the logs as to what may have caused the communication loss, other than that the rf connections were reporting 100% bad packets. there is insufficient information to determine whether or not the repeater faults were solely responsible for the event. | 1
1404752 | we have discovered a problem with healthcare computer; retail pharmacy management system. a prescription was entered into the synercom system and a picture of the label and the prescription were scanned into the ez-flow system. the patient instructions were consistent in all 3 places. then we used the system to fax the doctor for refills and the patient instructions came out of the system different. the correct directions were 'use one syringe 7.5mg as needed for seizures lasting longer than 5 minutes'. the directions on the fax read 'use 1 syringe //7.5mg// as needed for seizures lasting give 1 tablet by mouth once'. the last part was somehow substituted by the system. if the fax record was used to create a new prescription then instructions could be wrong creating a possible risk to the patient due to overdose, underdose or wrong route, etc.. we have seen this kind of behavior in this system before and reported it to the company. last october we had a patient prescription that had instruction that read 'give 0.4ml by mouth three times daily'. the patient decided to not pick up the prescription so it was placed on hold in the system. so 8 days later, the patient decided to fill the prescription. when we processed the prescription from hold, without modifying it, the instructions read 'every 4 to 6 hours as needed for seizures'. so the same thing happened with the same kind of risk involved. | 1
1511082 | the customer reported that a pt became apneic and experienced a lift threatening event, requiring resuscitation for 5-7 minutes. we will consider this to represent a serious injury. though the customer stated they did not get an apnea alarm, the strips and logs show that the vifib/tach alarm occurred at 03:11, which was before the monitor detected apnea, therefore, no delay in treatment occurred. the vfib/tach was responded to by staff when it occurred. the fse has indicated that on arrival, the respiration alarms were off for this bed, which would prevent respiration alarms of any kind from being annunciated. this is the default configuration at this hospital. post incident testing found the device alarming on command. the available info does not support that a device malfunction occurred. >< the customer reported that a pt became apneic and experienced a lift threatening event, requiring resuscitation for 5-7 minutes. | 1
1349388 | we have requested the return of the identified pt monitor for evaluation and will update this file when we have completed our investigation. >< the customer reported that after replacing the lithium battery in the pt monitor and then operated the device for approximately 6 hrs, the battery overheated and melted the plastic monitor case material. this condition produced enough smoke, which required an evacuation of the care unit. no pt injury was reported. | 1
1488120 | multifunction cardiorespiratory monitor with oxygen saturation capability being used on a patient in intensive care unit. there was no audible alarm for low saturation at central monitoring station. oxygen saturations dropped to 30% without audible alarm before returning to normal after staff intervention. mfr aware, but remains unable to solve problem. | 1
1063599 | it was reported that the rac 4a with tram module and mount arm disengaged from the vertical wall mount and fell while a nurse was raising the bed. the mounts were located directly over patient's bed due to bed area and available wall space logistics. no patient injury was reported. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 0
1362991 | the pins were bent on this monitor due to the pins on the philips mms module. the bent pins cause the monitor to be inoperable. this situation presents the potential for patient safety issues, as health care providers need access to immediate monitoring of unstable patients that are brought to critical care areas. the biomedical technician replaced the necessary components, tested the device, and returned it to service. the device is available for onsite evaluation at our facility. questions and comments should be directed to the biomedical director. | 1
1375123 | five previous medsun reports were filed last year. resolution from ge is still pending. over fifty pdm, carescape modules are sitting in a box in storage, which have not been deployed to clinical use since receiving them. unable to obtain the update or definite time line to when the problems will be resolved. the previous medsun reports filed were for:1) nibp hardware failure with pdm 2) nibp issues with pdm and tram 3) rebooting issues in pdm 4) spo2 issues with pdm 5) temp issues with pdm ====================== manufacturer response for physiological monitor module, pdm (carescape)======================released software upgrade for solar 8000i. upgraded monitors at the hospital. put pdm¿s (carescape module) back into clinical use into a small pod of the or¿s as a pilot. pdm¿s experienced the same failures as mentioned in the medsun reports filed. pulled devices out of use. alerted ge. solution still the hospital is pending a response from ge, with no update or timeline since. the pdm¿s have been out of clinical use for ~ 1 year at this time. | 1
1344577 | spacelab's solution was to add an additional server but this still did not work. there are two fundamental things deficient with this system;1. we have no active monitoring system that detects when the 72-hour disclosure interface is not sending info to the database. the only way staff knows the system is down is when they need to review cardiac rhythms or vital signs and the screen is blank.2. the current system configuration saves information in a flash type memory, when the interface/server is restarted all the information is purged and gone forever. the spacelab's system needs to write to a permanent archive (hard-drive) that can be accessed after the system is rebooted and restored; the flash memory does not save the tracings when it is restarted. i don't know how this patient critical system can be deficient in both these areas and still be fda approved?====================== manufacturer response for physiological monitor, module, command module======================it was our it dept's responsibility to come up with a solution to monitor the interface. | 1
1443971 | call from central monitoring to change batteries. batteries changed. there was shortly another call to change batteries due to no signal in central monitoring. unit was hot to the touch. when battery compartment was opened, batteries were starting to melt and a small spark was present. unit removed from patient. | 1
1160633 | this mdr is being submitted late, as this issue was identified during a retrospective review of complaint files. this activity is being completed as part of a corrective action. >< after the exam was completed, and transferred to picture archiving & communication sys (pacs), it was reported that the pacs scan changed the time in the study date/time, but not the date itself during this process. this change is then transmitted to pacs. no injury or misdiagnosis was reported. the incorrect study date/time can cause confusion in ordinary practice, whereby a pt's images can be perceived to be newer than they really are, and lead to improper clinical decisions. | 1
1148889 | the site reported that the workstation were showing the incorrect study date. no injury or misdiagnosis was reported. the incorrect study date/time can cause confusion in ordinary practice where by a patient's images can be perceived to be newer than they really are and lead to improper clinical decisions. >< this mdr is being submitted late, as this issue was identified during a retrospective review of complaint files. this activity is being completed as part of a corrective action taken pursuant to an fda inspection conducted at the facility in april 2008. | 1
1160642 | the reported issue was duplicated through internal investigation. it was found that the pacs system does provide the user with a notice in the jacket header that the jacket does not match the current exam when an incorrect patient jacket appears. the patient jacket, although not matched with the current exam, does provide the correct information to the user about its contents, including the patient name and identified. the contents, if selected, match the patient jacket and display the same name and identifier. this mdr is being submitted late, as this issue was identified during a retrospective review of complaint files. this activity is being completed as part of a corrective action taken pursuant to an fda inspection conducted at the facility in 2008. >< the site radiologists reported an incident where the prior patient's comparison list was displayed within the patient jacket. the radiologists noticed that the current patient packet displayed did not have any priors when the exam they had opened did have priors. no injury or misdiagnosis was reported. | 1
1160643 | the site radiologists reported an incident where the prior patient's comparison list was displayed within the patient jacket. the radiologists noticed that the current patient packet displayed did not have any priors when the exam they had opened did have priors. no injury or misdiagnosis was reported. >< the reported issue was duplicated through internal investigation. it was found that the pacs system does provide the user with a notice in the jacket header that the jacket does not match the current exam when an incorrect patient jacket appears. the patient jacket, although not matched with the current exam, does provide the correct information to the user about its contents, including the patient name and identifier. the contents, if selected, match the patient jacket and display the same name and identifier. | 1
1216358 | multiple attempts have been made by the manufacturer, regarding the outcome of the pt associated with this incident. however, the reporter has declined to provide any additional details. >< it was reported that an adult male pt in the emergency room was being treated for an open head injury. the clinicians were trying to view an exam to determine treatment of the pt. however, the exam could not be viewed via pacs. after a delay in treatment lasting approximately 40 minutes, the clinicians were able to view the exam via the modality in which it was acquired. the pt outcome was not known. | 1
1034935 | examinations were discovered at this site that did not present the proper study date in the ge picture archiving and communication system (pacs). the examinations show the date at which they were acquired into the pac system instead of the date the study was actually performed. this situation can give the appearance that the pt's images were taken more recently than what had actually occurred. there was no report of this situation resulting in any adverse pt outcome or misdiagnosis. the site expressed concern that if the study date discrepancy was not noticed it could potentially allow a clinical decision to be based off incorrect info. | 1
986026 | add'l evaluation included engineering review of the code that handles mpr image orientation. after investigating this situation, it is philips' opinion that there is minimal risk of pt injury if the mpr (reconstructed) image is incorrectly displayed in isite pacs because: the original (correct) image is always available in the pt timeline. the original (correct) image is always displayed on the monitor next to the possible incorrect image (not all mprs are incorrect, only those images that are acquired as described. anatomical markers are often present (making identification of the incorrect display likely). any findings would generate add'l imaging prior to surgery or active intervention confirming anatomical location. >< in 2007, a customer complaint was reported that when using isite pacs to perform mpr, the left and right orientation markers were being displayed incorrectly. further investigation revealed the reverse orientation marker display affects only reconstructed coronal and sagittal images when the isite mpr feature is used if the pt was scanned as supine feet first, prone head first or prone feet first. additionally, the ordering of coronal and sagittal image slices can also be reverse. for example, a scout line on the right side of an axial image could intentionally display the corresponding sagittal slice position on the left side. subsequent in-house testing revealed that the reported issue is reproducible. the original axial slices are always displayed and annotated correctly in all cases. the original dicom images are on isite pacs and presented in correct orientation next to the mpr images when they are displayed on the monitor. there are no reports of pt injury related to this issue. | 1
1034931 | examinations were discovered at this site that did not present the proper study date in the ge picture archiving and communication system (pacs). the examinations show the date at which they were acquired into the pacs system instead of the date the study was actually performed. this situation can give the appearance that the pt's images were taken more recently than what had actually occurred. the situation was identified at the site and there was no report of this situation resulting in any adverse pt outcome. the site expressed concern that if the study date discrepancy was not noticed, it could potentially allow a clinical decision to be based off incorrect info. | 1
1119287 | a defect has been identified; a behavior in isite sing the conference presentation state (cps) - where the patient data on the canvas page are not synchronized with the image data in the viewer - incorrect data will be displayed - data from different patients. a conference presentation state is saved when grouping images to present in a conference to aid a diagnosis or to use as a training tool. it has been discovered that when multiple patients are selected, and any other than the first patient on the list is pulled up, both the patient that is pulled up and the first image in the list will display on the canvas page resulting in the initial patient on canvas page verses on the dx monitors are different patients. this behavior is present in both 4.1 and 3.6. if the presentation is saved as a system presentation (which is default) this mismatch of images may not be identified and could lead to possible mis diagnosis. >< device from reserve sample evaluated. | 1
1402737 | a physician reported that while working at a workstation, they observed images from patient 'a' displayed while the report for patient 'b' was being edited. after some investigation, the reason image data from patient 'a' was being displayed while reporting on patient 'b' was that the web browser hosted within the workstation application was using previously cached web data causing a previous report to be displayed. the web browser on the pc was configured to follow the caching settings specified by the web server ('automatic') and would have not resulted in the issue, but this setting was ignored when the web browser was hosted within the workstation application. the web browser hosted within the workstation application had not been configured properly to disregard cached items. (using the web browser outside of the workstation would not have resulted in the same issue.) the work around for this issue is to change the browser setting for 'use of cached pages' to force it to get a new version of the web page ever time from the web server (and not used cached data nor follow the cashing settings specified by the web server). | 1
1003068 | there have been no reports of actual harm to patients. the product generates a report, which is then available: co systems are presently not classified as medical devices. however, any time there is data discrepancy, with clinical content elements being unavailable to the end-user, there is some potential for harm. in this particular circumstance, dealing solely with the nuclear cardiology report manager, the data elements involved are certainly not the sole determinants for clinical decision making, so the potential for harm is low. however, if the situation were to recur, with more clinically significant data elements being unavailable for clinical review, the clinical end-user of the co report might make clinical decisions that could contribute to serious injury. complicating the issue is that the co report is presented in the co to a wide variety of users, some of which are not well versed in the field of cardiology (e.g. emergency department physicians). for example, if the 'conclusions' section is not available in the his report, the final recommendations would not be available for an end-user who is not expert with the more technical elements of a cardiology study. this could prompt a repeat diagnostic study, with attendant risks (e.g. administration of contrast material, causing allergic reaction or renal failure). agfa is being conservative in reporting this product problem. >< product utilizes software, called 'report manager for nuclear reporting, outbound report, version 2.04.20. this software generates a report to be sent to the information system. this report is referred to as the 'outbound' report and is commonly retrieved from a sites his for use in patient care. in late 2007, agfa was contacted by a user that information available from the product, and present in a separate report (also generated by report manager and stored on the product) was not present in the 'outbound' report. an evaluation by agfa personnel of the request from this user for the additional information determined that based on the potential use of the 'outbound' report by hospital personnel, critical available data should have been included in the design of the 'outbound' report. risk analysis conducted by agfa personnel suggested there is a potential risk of serious injury, due to possible treatment decisions because of lack of known information in this report. information not included in the 'outbound' report is the hemodynamics table, and the conclusions from the nuclear data. | 1
1346817 | capsule endoscopy ordered. the patient swallowed the pillcam and the equipment attached to data recorder as directed. the data recorder lights were on and blinking. the patient stated that the light stopped blinking at approximately 1130. given corporation technical support notified. data recorder found not to be working properly. | 1
1073875 | the customer reported that the telemon was not providing audible alarms. >< a philips field support engineer went on site and replaced the failed speaker, and tested the device post repair. normally, both visual and audible alarms are generated at both the central station, and the telemon bedside, and a situation where there was no audible alarms at the telemon would not jeopardize pt safety, as the central station alarm would sound. however, because telemon can be used as an isolated monitoring device, and because labeling does not specify close personal observation when in use, recurrence of this failure mode could lead to serious injury or death. the repaired device remains in use at the customer site. | 1
1564745 | unplanned hospital wide cpoe and electronic care record breakdown resulted in tardiness of administering and in frank missed doses of medications. considering the size of the institution, it is possible that other patients were adversely affected by comparable delays and omissions. | 1
1550321 | the user reported that they had a pt code in the ct room, and they grabbed the device to take the pt's nibp reading. the user stated that they believed that they tried to get a nibp reading before the base monitor was in communication with the display monitor and, as a result, the device would not take nibp reading. the customer stated that the pt was revived successfully. >< the device mfr's field service rep (fse) visited the customer site to evaluate the device. the fse was able to duplicate the reported problem and replaced the device's radio as a result of his eval. although the user reported a problem with nibp parameter, the actual problem was with the device's radio. the problem with the device's radio caused the requested nibp reading to not be transmitted from the display monitor (where it is initiated) to the base monitor (which actually takes the nibp reading). following replacement of the radio, the device was functionally tested per internal test procedures and returned to service by the user. the device mfr spoke to a clinician present during the reported incident in order to determine whether or not the device was a factor in the incident and outcome of the pt. the clinician stated that they were not using the device to monitor the pt at the time of the incident. when asked if the clinician thought there was a delay in treatment of the pt as a result of the device's malfunction, the clinician replied 'absolutely not, it had nothing to do with anything'. the clinician also expressed frustration with the questions being asked by the device mfr because she felt that 'the [device mfr's] monitor had nothing to do with what happened and [the device mfr] should never have been called in the first place'. if this malfunction were top recur, it would not lead to a serious injury or death because this malfunction would be obvious to the user. if the malfunction were to recur, the user is likely to utilize back-up monitor to proceed with the case, and this is what happened in this case. if a back-up monitor was unavailable and lack of a back-up monitor was to preclude the use of the scanning device in diagnosis, alternative diagnostic methods would be implemented as warranted for the pt. | 1
1397581 | it was reported that the dpt was unable to zero, due to the dampened waveform. no patient complications were reported. >< device has not been returned for evaluation. | 1
1287143 | this patient had suffered a fractured hip and was admitted to our hospital. the patient's condition was declining, requiring transfer from our cardiac telemetry floor to our intensive care unit. the patient was intubated, placed on medication pressor agents for hypotension and sepsis, and was set up with a central line pressure monitor device to monitor his central venous pressures. the patient was receiving fluid boluses based on the readings from the monitoring equipment. after receipt of multiple liters of fluid, the readings were remaining lower than would be expected. after troubleshooting, the cable was switched out and more believable readings were obtained. this cable was taken out of service, tested, and test results showed it was not reading accurately. there was a pin broken and also the pin configuration was different than the others in the department. this patient expired. the extra fluid was not determined to be the cause of this patient's death, but it did not help the situation (it may have contributed somewhat, in regards to timeframe. this is unknown.) for this reason, we are submitting this report. | 1
1366543 | evaluation: customer report was confirmed. rec'd (1) double dpt - vamp adult kit in open original package. kit was not used. tubing was broken off from solvent bond joint with the female connector (attached to zero-stopcock). broken tubing was bent near the bond joint. >< it was reported that the piece of tubing slipped out of the luer lock before use. | 1
1080108 | facility reported to smiths medical that the luer had detached from the tubing. the sample was in use for 8 hrs. there was no pt injury or treatment reported with this event. >< add'l lot # 1284521. the subassembly in question is a560 and is mfg by smiths medical. this assembly is incorporated into the finished prod by smiths medical. the finished prod is further assembled, packaged and sterilized by smiths medical. one a560 subassembly was rec'd with the female luer lock disconnected from the tubing. the solvent ring on the tubing indicated that the tubing had been fully seated into its connector and solvent had been applied. the tubing measured within spec. there did not appear to be any mfg issues that would have contributed to a separation. the tubing assembly in question was mfg on an automatic tubing assembly machine. the prod is 100% visually inspected by the operator for proper assembly and sufficient solvent. as part of the inspection process any units removed during the 100% were reworked by the operator. as a precaution in 2007, this practice has been discontinued. any units removed during the 100% inspection are being discarded. smiths was able to confirm the issue; however, no root cause could be determined. no add'l action is necessary at this time. smiths considers this mdr closed with this report. | 0
1158006 | it was reported that leakage was observed from the crack on the female luer. >< device has not returned for evaluation. | 0
1432048 | customer has not yet returned the device to the mfr for device eval. when and if the device becomes available and is returned and evaluated the mfr will file a f/u report detailing the results of the eval. >< the reporter stated that the device tubing separated. there was no reported pt injury or treatment associated with this event. | 0
1564205 | the graph record button was pushed during a procedure. the wave forms were noted on the central monitor and the central station printer began printing. at the end of the procedure the wave forms and pressure were not printed. | 1
1156165 | reporter indicated a dell x5 vns computer was freezing during interrogation and programming events. reseating the flashcard resolved the issues. the dell handheld computer and flashcard have been returned and are currently in product analysis. | 1
1156167 | reporter indicated the light on a vns dell x5 computer was on and she was getting error message that said the back-up battery was low. the reporter reinserted the flashcard but the screen continued to be black. after performing a soft reset, two warning messages displayed: 'error message - execute cab file time out, new application not installed' and again 'back up battery low'. the reporter was able to move past these screens and was able to get to the interrogate screen. the dell computer and flashcard have been requested for return but have not been received to date. >< device failure is suspected, but did not cause or contribute to a death or serious injury. | 1
1172208 | it was reported that a dell handheld computer would not function when the physician was trying to use it during an implant in the or. troubleshooting did not resolve the issue, and it was thought that there was a possible battery failure. a week later, a company rep tried troubleshooting again on the handheld with no reaction. however, for no apparent reason, the handheld powered on and appeared to be functioning properly. the product has been returned to the manufacturer and is undergoing analysis. >< device malfunction is suspected, but did not cause or contribute to a death or serious injury. | 1
1181589 | (see scanned page). >< reporter indicated a vns dell x5 handheld computer was freezing during interrogation. reseating the flashcard resolved the screen freezing. the handheld computer and flashcard have been requested for return but have not been received to date. | 1
1181592 | reporter indicated a vns dell x50 handheld computer was not performing as intended. the computer was also not charged per the reporter. the reporter was advised to fully charge the computer before the next use. attempts to the reporter to ascertain whether the computer was working properly after charging have been unsuccessful to date. >< conclusions: device failure is suspected, but did not cause or contribute to a death or serious injury. | 1
1181595 | analysis of the returned flashcard revealed sql errors following an interrogation attempt, most likely associated with loss of power or removal of the flashcard during communication attempts. the analysis revealed that the cause for the sql error message is associated with a corrupt cyberonicsbu.cdb database. the most likely cause for the database corruption is associated with a sudden loss of power to the handheld during usage. additionally, it was identified that several archive databases were corrupt. the most likely cause for the corruptions is also associated with a sudden loss of power. no other anomalies were identified. | 1
1265758 | reporter indicated vns computer was freezing during interrogation. the computer and flashcard have been returned and are currently in product analysis. | 1
1265761 | reporter indicated a vns dell x5 computer was freezing during interrogation. the computer and flashcard have been requested for return. | 1
1374182 | we are still investigating this reported incident. a f/u report will be submitted as soon as our investigation is complete. >< it was reported by the customer that after upgrading to vf8.3, there are multiple intermittent dropouts of the telemetry pt signals on this ics central station used on their telemetry floors. no injury was reported. (b) (4). | 1
1455798 | it was reported that the metal locking tabs of six dash port 2 docking stations broke off of the monitor clamp on the pivot bracket. in this scenario the dash monitor would not be connected to the dash port 2, potentially leading to a situation where the monitor could fall on a pt or health care provider. this report is for the third of six affected devices. there was no report of pt or user injury. ge healthcare's investigation into the reported occurrence is still ongoing. a follow-up report will be issued when the investigation has been completed. | 0
1369463 | the customer reported that while the panorama central station was in use monitoring pts along with ambulatory telepacks, the central station display went blank. the pt was transferred to another central, and monitoring was continued. no pt injury was reported. >< the company service rep replaced the display power supply. the system was tested to factory specifications. it functioned normally, and was returned to use. | 1
1284703 | per user submitted mdr, 'monitors would beep then screen would go blank, a few seconds later screen would come back but data would be changed and clock time is off. this happened with several monitors so it affected several babies. ge healthcare's investigation into the reported occurrence is still ongoing. a f/u report will be issued when the investigation has been completed. | 1
1111827 | it was reported that when a radiologist was dictating a pt report using powerscribe extends software, dictation was entered for pt b when it should have been entered for pt a. it was noted that the report was stopped before reaching the physician, and no treatment was performed on the incorrect pt. there was no report of death or serious injury associated with this incident. >< the facility (b)(6) had a 3rd party powerscribe extends software (v4.7) that runs under centricity workstation. the issue has been caused by a change/fix within powerscribe software, a daylight saving time (dst) powerscribe software patch. the customer reported that prior to the installation of the dst powerscribe software patch, if a radiologist opened a study in pacs (picture archive and communication system), did not dictate anything on that study, closed it and then opened another study, the 'shell' created by powerscribe software was automatically deleted. after the dst software patch installation, if a radiologist opened a study in pacs, did not dictate anything on that study, closed it and then opened another study, the radiologist gets a message: do you want to delete report? this causes confusion as the radiologists didn't know which report they were being asked to delete and if they said 'no', the prior exam that was opened, that powerscribe report shell would then be synched with the newly opened exam. this would create an incorrect dictation on the wrong pt. | 1
1277473 | workstation intermittently opens wrong patient exam. >< a ge representative evaluated the issue with the customer's system and provided suggestions for work-around. safety letter was also provided to the customer. there was no patient injury associated with this event. this system is being monitored and remains in use at the facility. | 1
1277475 | radiologist reports that when trying to create an exam note for an open exam. he clicks on create exam note icon, the exam note that opens has a different patient name, not matching the exam he is reviewing. >< a ge representative evaluated the issue with the customers. users will be notified to watch for this issue to prevent data mix. there was no patient injury associated with this event. this system is being monitored and remains in use at the facility. | 1
1352587 | this submittal involves exactly the same issue as reported in submittal 1937649-2009-001. the difference is that it was reported by a different customer at an earlier date. >< when this issue was first reported, cms inc posted a customer bulletin discussing the issue presented in this filing. as noted earlier, at that time we did not consider this to be an mdr situation. with the reoccurrence of the issue even with the bulletin posted for all to read, we now believe it should be reported as an mdr. for a copy of the customer advisory to be sent to all users advising them of the issue presented here, please see 1937649-2009-001. the issue will be corrected in focal release which is currently scheduled to be available in 2009. | 0
1160636 | this mdr is being submitted late, as this issue was identified during a retrospective review of complaint files. this activity is being completed as part of a corrective action taken pursuant to an fda inspection in 2008. the reported issue is being investigated. a f/u report will be submitted when the investigation is complete. >< a rptr expressed concern that the standard uptake values (suv) for positron emission tomography (pet) are incorrect when the examination was performed on a scanner not mfg by ge healthcare. as suvs are useful in the determination of the malignancy of a lesion, accurate calculations are expected. there was no report of adverse pt outcome or data loss. | 1
1552426 | during left leg thrombectomy, patient on table was tilted to such an angle as to cause him to slide to the floor. | 0
1208493 | evaluation in progress, but not concluded. >< during routine testing at a service center, the arterial monitoring module did not power on after 10 seconds. there were no adverse consequences to a patient as a result of this event. | 1
1016480 | in absolute dose mode a beam with a port may be calculated with open normalization, with the result that the displayed dose reflects open normalization while the monitor units and qa plan reflect blocked normalization. this is particularly important when these beam also include compensating filters. the compensator doses will be displayed in open normalization mode, and the mu calculated in blocked, so the real, delivered dose is higher than displayed. the displayed dose is approximately the delivered dose times the compensating filter factor. an assessment can be done by looking at the weight of the beam in source index or weight page, as compared to the displayed dose for that beam at the weight point. dose differences can be of the order of the compensating filter factor. the problem is detectable by looking at the delivered dose to the weight point in the pt as compared to the source index or weight page dose. a customer advisory is being sent to each user that may see this issue. the problem will be resolved in an upcoming software release ans our final report will include an timeline for that release. >< customer had imported images/studysets/plan into focal from philips acqsim using dicom, switched the machine in focal from iec to a regular machine, and sent this to xio for dose calculation. the beam coming over to xio had port normalization set to open rather than blocked. this meant that the compensator doses were displayed in open normalization mode, and the mu was calculated in blocked resulting in the real, delivered dose being higher than displayed. the displayed dose is approximately the delivered dose times the compensating filter factor. two pts had started therapy before the problem was identified. treatment is ongoing for both and clinicians are adjusting their remaining therapy to compensate. | 0
1359124 | in connection with the use of the caresuite family of medical device data systems, it was reported that manual patient allergy data entries into our end user application were overwritten during automatic updates of patient data from the hospital information system. this configuration contributed to a patient incident, resulting in a temporary allergic reaction (shortness of breath) to a prescribed medication order of toradol. the incident was mitigated through normal clinical practice. the patient was able to proceed with their scheduled surgery, without further effects. an improper configuration of the interface database contributed to the incident, by allowing automatic third party his/cis updates to the end user application, where manual updates are in practice. customers preferring this practice should not link the interface database to their his/cis. the reporting customer site has been reconfigured to support their preferred practice. there have been no other reports of adverse patient impact associated with this issue. nonetheless, during the course of our current investigation we've received one additional site report of allergy data being overwritten. >< our investigation into the reported incident revolved around how our medical device data system interfaces work in populating caresuite end user applications with his/cis data. the caresuite family applications are delivered with 'standard' linked allergy fields. by default, these fields are configured to import data from other vendor systems, and/or to transfer data within the caresuite applications. when implementing these applications, sites must decide whether the his or caresuite is the master of the allergy and precaution data. data must not be entered through both systems in order to maintain the integrity of the data. the situation is that patient demographic data fields for allergies and precautions can be entered into care suite applications in two ways: locally by the end user or externally via an his/adt interface. our methodology is that the his is the default master. the following scenario is possible: patient is booked to go into the or. in the preop application, the pre op nurse enters an allergy. patient record is started in anesthesia and the adt interface downloads allergy information from the his. if the information is the same, then there is no problem. however, if the his allergy field is blank or different, it will overwrite the end user allergy field entered in the preop application (e.g. remove an updated allergy, leave it blank or different). it is currently possible to configure the system to prevent this, and many customers understand that the his is the master by default, but it is possible for this sequence of events to occur. through our clinical training and implementation, picis review with customers how they want to configure allergies. this review includes the fact that interface data from an his will overwrite the data entered into our applications. most customers decide not to bring allergies in through the interface database. however, if the allergy fields are not de-linked from the interface, the described scenario is possible. in developing our corrective action strategy, we have reviewed customer configuration profiles, and evaluated the level of risks associated with usage models. through more detailed understanding of customer configurations, picis has developed a strategy for recommended system interface reconfiguration, to meet our customer needs. to help ensure the most appropriate configuration and integrity of patient data, we have initiated a customer communication to our installed base, explaining the potential source of patient data conflict and to provide support in choosing the appropriate configuration option and implementation. | 1
1248163 | versacell dropped a tube over a rack of samples in the front of the system. several tubes were cross- contaminated, resulting in pt re-draws. pt treatment was not altered or prescribed. there was no report of adverse health consequences since all samples were redraw. >< a siemens healthcare field service engineer was sent to the customer site. analysis of the instrument and instrument data indicate that the gripper finger was missing o-rings which caused the dropped tubes. no further evaluation of the device is required. the instrument is performing within specifications. | 0
1546775 | the device was unavailable for return and exam. the device history record including sterilization record for lot 03r2106 was reviewed. the review concluded that the lot was mfg according to specifications and all parameters were met during the sterilization process with no deviations. an analysis of the complaint database for this lot revealed no similar occurrences have been reported, nor any failure of the disposable lot that could explain the reported occurrence. at present, there is no indication that the device caused or contributed to the event. if the kit is later received and findings differ from the info presented in this report, an addendum will be filed. >< approximately 20-25 days post collection, the customer reported a pt with a bacterial infection shortly after transfusion of an rbc product collected on a trima accel. the pt was treated with antibiotics and discharged 4 days later on oral antibiotics. a repeat blood culture 2 days after the initial positive results showed no growth. the pt is reportedly fine now. the transfused unit was part of a double rbc collection that was split using gravity techniques. thus the bags were open to each other for some amount of time in order for the product volumes in each bag to equilibrate. the second rbc unit was transfused without incident. according to the customer's medical director, 'at the time, the pt had the reaction to the transfusion, it was also determined that he had a lower lobe pneumonia. however, the presence of bacteria in the blood culture mandated iv antibiotics for at least a couple of days until the culture became negative. the source of the bacteria in his blood stream remains unk.' | 0
1468855 | reportedly, in 2009, the patient underwent a therapeutic plasma exchange (tpe) procedure with ffp as replacement fluid for ttp. the procedure was uneventful except for an air in return air detector alarm at the end of the procedure. the operator was unable to clear the alarm and performed a manual rinseback. this is part of their normal protocol, and they are familiar with performing manual rinseback. the procedure ended with no additional issues. the patient's medical condition was unchanged from prior to the procedures. two days post procedure, the patient expired. the customer does not think the alarm, the machine, or the procedure had any bearing on the death of the patient. this mdr is being filed due to the proximity of the death to the tpe procedure. >< a service call was placed to do a full checkout of the instrument with a long simulated run. there was no alarm history because the ram ic (u10) on the display board was measuring zero volts. the service technician verified that all of the other voltages were in the manufacturer's spec. the pressure sensors, valves, and the pumps were in the manufacturer's spec. calibration on the rbc detector and the collect concentration monitor function were verified to be in spec. the centrifuge speed check was also verified. a setup and simulated run were performed. the machine alarmed 'air in return chamber' several times during the run, and there was no air in the chamber. the customer biomed indicated that he had already replaced the return air sensor and the sensor cca. the service technician suggested that the air detector board, and the ram ic (u10) be replaced to bring the instrument back into manufacturer's spec. the alarms seen would not have contributed to the patient death. | 0
1448349 | a new package of cobe spectra apheresis system was opened. the tubing was kinked, and we were unable to use it because of the kink. in the past, there was another kit with the same problem, but it was not saved. unfortunately, the tubing is not visible through the packaging, so we cannot visually inspect new unopened devices from the same lot. | 0
1376656 | shower chair broke with patient on it. pt suffered laceration to scrotum. | 0
1518131 | we have discovered the natus sleepworks software (version 6.0.0, build 598) allows two recordings on different patients to be stored in one patient's medical record. the physician read this and gave recommendations to the incorrect patient. specific details: two patients were scheduled for sleep center testing on the same night. both patients had previously been at the sleep center for studies on different dates. the first patient arrives on time and is set up in his room. the second patient arrives late and is placed in a different room than originally scheduled. the sleep center tech enters the second patient's information into the computer and notes that what pulls up is the first patient's information. the tech selects 'edit' and begins to fill in the second patient's information despite a warning that there is no recording of the sleep data. the result is that the first patient has three sleep tests filed to his record (his first exam, the exam he is undergoing currently, and the second patient's exam). both recordings proceeded without difficulty. when the data was transferring to the main server, the tech discovered that all studies were linked to the first patient. the error resulted in the second patient's data stored to the first patient record. the physicians read this and gave recommendations to the incorrect patient. the sleep center staff have notified the manufacturer of this problem. >< we have discovered the natus sleepworks software (version 6.0.0, build 598) allows two recordings on different patients to be stored in one patient's medical record. the physician read this and gave recommendations to the incorrect patient. specific details: two patients were scheduled for sleep center testing on the same night. both patients had previously been at the sleep center for studies on different dates. the first patient arrives on time and is set up in his room. the second patient arrives late and is placed in a different room than originally scheduled. the sleep center tech enters the second patient's information into the computer and notes that what pulls up is the first patient's information. the tech selects 'edit' and begins to fill in the second patient's information despite a warning that there is no recording of the sleep data. the result is that the first patient has three sleep tests filed to his record (his first exam, the exam he is undergoing currently, and the second patient's exam). both recordings proceeded without difficulty. when the data was transferring to the main server, the tech discovered that all studies were linked to the first patient. the error resulted in the second patient's data stored to the first patient record. the physicians read this and gave recommendations to the incorrect patient. the sleep center staff have notified the manufacturer of this problem. | 1
1360248 | this issue involves the surescript erx functionality in powerorders and affects users that utilize powerorders electronic routing to document prescription medication orders. when an update is made to the frequency field on an existing prescription, the frequency schedule id is not simultaneously updated on new orders sent to the pharmacy via surescripts. a patient could potentially take a medication too frequently or infrequently if the pharmacy receives an electronic prescription with the wrong frequency. cerner received notification that two patients received more than the prescribed dosage of medication as a result of this issue. the first, an elderly patient, was reported to have received more than the prescribed dose of a blood thinner levoxyl for 6 weeks. no serious injury or death has been reported to cerner in association with this issue. a second patient event was reported where the patient received inappropriate dosage orders for carbamazepine and was subsequently admitted to the hospital with atypical chest pains. the client subsequently reported the patient took the correct dosage based on previous instructions and did not rely on the inappropriate dosage as indicated in the surescripts prescription message. >< cerner distributed a priority review flash notification in early 2009 to all potentially impacted client sites. the software notification includes a description of the issue and interim workflow adjustment to prevent the malfunction. a software is being developed to address the issue for all the sites that could be potentially impacted. cerner corporation will provide a follow-up report when the software modification is available. | 1
1365512 | the issue involves cerner millennium powerchart message center inox. when the results folder is filtered by event set and the event set display and event set name are different, result filtering may not function correctly. if the result filter is in message center and is an inclusive filter, results may be filtered out incorrectly. the impact associated with this issue is that patient care may be delayed or missed as follow-up may not occur. cerner has not received communication on any adverse patient events as a result of this issue. >< cerner distributed a priority review flash notification on january 28, 2009 to all potentially impacted client sites. the software notification includes a description of the issue and a software modification to address the issue for all sites that could be potentially impacted. cerner corporation will provide a follow-up report when the software modification is available. | 1
1431486 | product is a software manufacturer, called interplant, manufactured by cms software, problem being reported, due to product problem -confusion- and product user error -not behaving as expected-. used device to capture ultrasound images in planes z=0 through z=4.0cm. a plan was automatically generated that included needles containing 5 radioactive seeds with 1cm spacings. this results in seeds at slices z=0,1.2,3,4. seeds are also displayed on these slice images. software feature allows needles to be inserted/retracted. when needles were retracted, some of the seeds could no longer be visualized, because no images existed for them to be displayed upon. we expect seeds to be available only on the images. this results in a plan being generated with more radioactive seeds than we are aware of the plan needing. this also results in seeds being placed where they are not needed. the device prints out a diagram showing needle location, location of seeds within the needle, and retraction of the needle within the pt. there was no visual alert showing seeds were being placed outside the captured ultrasound images. effect on the pt should be minimal in this particular case, but in extreme cases, could result in a less than optimal treatment for prostate cancer, due to radioactive seeds being placed in location where there is no cancer. i've phoned the manufacturer and alerted them to this problem. they report the work around is to be more aware of the seed location from the print out and within the software. i've requested a change to the software, so it behaves in a more intuitive way or displays a message stating there are seeds in locations where there are no images. dose: na. diagnosis: prostate cancer. | 1
1545506 | siemens microscan walkaway 96si panel pos combo 26, giving false resistent infection - fifteen -15- pts with possible false infection secondary to software update -v 3.01-. cases are under clinical revision. dates of use: 2009. diagnosis: infection detection. event abated after use stopped or dose reduced? yes. event reappeared after reintroduction? no. | 1
1398154 | lab results were reported for the wrong patient: a patient was admitted to the ed with kidney pain. the physician's assistant assigned to the patient noted that troponin levels were reported on this patient, but were not ordered. the physician's assistant notified the lab immediately of the error. the troponins had actually been drawn for another patient admitted to the ed on the same day for a cva. there was no delay in treatment for that patient.investigation revealed that the laboratory results produced in cerner laboratory software system were resulted correctly, and the siemens his software system also resulted correctly. however, when resulted out in the picis ed pulsecheck (ibex) system, the results were for the wrong patient.lab and radiology orders are entered through ed pulsecheck, and are sent via our openlink interface engine out to our siemens his system, where an order number is assigned and communicated back to ed pulsecheck via openlink. the order is also transmitted to the siemens radiology and cerner lab along with the respective order information via openlink. once the orders are resulted, siemens radiology and cerner lab send results back via openlink directly into ed pulsecheck as well as our his system concurrently.the vendor has stated they have infrequently encountered this same issue with other customer sites. this confirms that the incorrect posting of result to the wrong patient account was due to ed pulsecheck. the incorrect posting seems to have occurred where ed pulsecheck did not recognize where the previous lab transaction terminated and took the adt information from the next radiology result that hit the interface at the same time, causing the incorrect posting.the vendor, picis, has applied a fix to insure the transactions are accurately being read by the application so this would not happen again. this has been the only incident that we have identified here, but again the vendor has experienced this issue infrequently at other customer sites. | 1
1291123 | upgraded cerner bridge software for specimen collection was installed after validation using cerner supplied protocols. nine days later the wristband of patient 'a' was barcode read and the demographic information of patient 'b' was printed on the label along with the accession number and barcode associated with patient 'a'. the error was detected in the hematology laboratory when the specimen could not be located. cerner was notified. they continue to investigate. | 1
1370286 | clinical lab chemistry analyzer was to have new software version pushed by the company. the company pushed an old version and disabled the analyzer. the equipment was inoperable for over four hours, which would delay any patient testing. | 1
1373210 | picis locks up during patient care documentation at the 'save' function. picis locks up only at the save function, allowing rns to continue to document critical patient data only to discover at the end of the case, that no data will save. in order to get out of the patient chart, the computer must be re-booted. all patient documentation is lost. nurses are forced to do their best guess as to surgical times, implants, medications, etc and re-chart the entire document. this is a safety issue for patients; rn staff direct their attention from the surgical field to the computer and spend a great deal of time trying to get the software up, and re-chart all the data. this is a healthcare safety issue because guessing what you previously documented live in a surgical environment puts rn staff at great risk for a patient care error, and an omission of critical patient surgical information, thus putting their rn license at risk.this has occurred with multiple procedures and patients. our facility has struggled with this issue for months. information specialists from our facility reported this issue to picis in february 2009. the issue has not been resolved, as of yet, and appears to be getting worse.we have considered going back to paper documentation for safety, but our electronic picis scheduling, documentation and charging system is so ingrained in our processes, it would cripple us to stop now. this ongoing software program leaves us in a difficult position - due to the risk of data loss. the current government is demanding electronic documentation as a health care industry cost initiative. we desire to be in the spirit of this order. please assist us in resolving this issue so we may move forward in the spirit of patient safety and accurate electronic documentation. | 1
1567715 | flush port of dignicare split off while flushing. now inoperable. new dignicare placed.piece that broke off was saved. the rest of the device was discarded. | 0
1083548 | a hospital facility reported on june 27, 2008, that a pt suffering from dementia arrived in the emergency department approx three months prior. the pt had a medication order placed by a physician on a custom order entry screen and received a medication to which he was allergic resulting in an allergic reaction requiring standard treatment. he fully recovered and was discharged from the hospital within 48 hrs. the custom order entry screen displayed a placeholder indicating no allergy info had been recorded on this visit. there was a historical allergy entry that did not display on the custom order entry screen, but did display on all other screens. the customer order entry screen is intended for ancillary orders (e.g., labs, x-rays, nursing orders) and is not intended for use with medication orders. the custom order entry screen requires a medical record number (obtained via registration) to display allergies from previous visits. | 1
1408143 | the pt was found with no respirations or pulse and a code was called. the clinical nurse manager responded by immediately checking the telemetry monitor at the desk and found the pt with a flat heart rate and no alarms sounding on the equipment. at the time of the code, all telemetry leads were in place. philips called to report an event and investigation. | 1
1548325 | the intellivue monitoring system has high and low heart rate alarms which can be changed by the clinician to as low as 15 bpm, and as high as 300 bpm even though the red alarm (red alarms cannot be turned off) has a limit of 40 bpm (in a grayed box that appears unable to be changed). other boxes in the red alarm screen that can be changed have a white box and up/down arrows which explicitly means that these can be changed. when the high & low rate limits are changed in the hr limits screen, the default in the red alarms also changes and can go as low as 15 or as high as 300. when this concern is raised with philips, they indicate that there must be some clinical judgement in setting alarms. while i agree, there is the potential that alarms could be changed to extremely low or extremely high, and a pt could be harmed because an alarm did not sound. | 1
1347244 | spacelabs cardiac monitor showed unrecognized rhythm with no alarm at 2:08 am. patient assessed and found with no pulse or respiration. initial strip on telemetry was 1:14 am with identified irregularities, but monitor not recognizing rhythm due to artifact. >< waveform data, trend data, error logs and monitor settings from the system were reviewed and confirmed that in the interval between 1:17 am and 2:15 am, 2009, the quality of the signal from the patient was very intermittent and noisy. during this period, the system did detect and generate alarms for tachycardia, a couplet and runs during times when the signal quality allowed. spacelabs personnel tested the system on-site on the afternoon of the same day and verified that the system was able to receive a noise-free ecg signal and that the system did generate alarms appropriately. testing at spacelabs confirmed that the receiver module was operating to specifications and that the transmitter was able to provide noise-free ecg waveform data to the receiver. the incident was discussed with the hospital patient safety coordinator, rn agreed that, based upon the hospitals internal investigation, there is no indication that spacelabs equipment may have caused or contributed to the patient's outcome. all findings are that the device is performing to specification. | 1
1243203 | hospital staff stated that a 90478 module did not give an audio or visual alarm when a patient experienced a seven beat run. the run was captured by the full disclosure feature and recorded as an alarm. other hospital staff stated that they have had similar problems on different telemetry channels. >< the module default is to alarm on a five beat run. however, the module can be configured by the user to alarm from three to ten beat runs or to turn alarms off. at this time, we do not know how the module was configured. a 'run' is classified as three or more consecutive beats of abnormal morphology. the 90478 module and full disclosure are different systems and can classify waveforms differently. as such, some waveforms can be classified by the two different systems as two different alarms. spacelabs is continuing its investigation into this issue and will supplement this mdr as appropriate. | 1
1095250 | customer was sent replacement tower and the original tower was returned to engineering for eval. engineering was unable to duplicate the reported problem. log files from the system were also analyzed. engineering found the application failed in the following scenarios three times: memory access violation to windows system built-in components. data persisting failure - this is in reference to data writing to database on the master tower. application communication failure between ui and pda component. engineering stated the first scenario is very difficult to reproduce and the log files do not have enough info for the cause of the reported incident. engineering will continue to try and reproduce the issue and determine root cause. >< in 2008, customer called to report they were monitoring patients when they received an error message which caused the q-tel rms application to shutdown and exit to the desktop. as a result of this, there was some data loss and the current patient's session date was corrupted. the end-user attempted to restore the session data through the application, but was unsuccessful. on 07/01/2008, customer had informed cardiac science sales rep of the fact that the day of the data loss occurred included a pt being monitored who was having chest pain during exercise; since the application prematurely closed and the lesion data was unrecoverable, they had no info to send to the physician. | 1
1379270 | spacelabs learned of the incident when visiting the hospital at the hospital's request to configure the pt monitoring system, such that the user could not turn off the alarms. the hospital staff advised that a nurse had inadvertently turned off the alarm watch feature at a bedside monitor, and that a pt had died. the hospital staff stated the equipment worked fine and did not take it out of service. there is no information on what level of training the nurse had. the alarm watch feature of a bedside monitor allow nurses to move away from the central station to a pt's bedside monitor and have various alarms forwarded to that particular bedside monitor. the spacelabs clinical educator met with the hospital's nurse manager and provided training on the proper use of the alarm watch feature. the nurse manager has instructed the staff to verify the alarm watch settings whenever that feature is used. this incident appears to have resulted from user error. there is no evidence of malfunction or device error. the alarm watch feature has been in use for 15 years; during that period, there has not been a significant trend of customers not knowing how to use this feature correctly. >< a pt was being monitored at a central station via a spacelabs telemetry transmitter. the responsible nurse had to leave the central station; and did so, believing the pt was being monitored with the alarm watch feature at the pt's bedside monitor. however, the nurse had mistakenly turned off the alarm watch feature at the bedside monitor. the pt collapsed and hospital staff did not hear the alarm at the central station. the pt subsequently died. | 1
1145608 | patient monitored for 2-3 days with topical eeg electrodes. when electrodes were removed, raw, irritated sites at electrodes sites were evidenced. subsequent information forwarded to us indicates two (2) patients had this type of reaction. both patients very similar in health and age profile. events occurred in 2008. >< patient1 and patient2 health severly compromised. both advanced age patients on chemotherapeutic protocols. patient1 comitose in icu prior to eeg monitoring. electrodes on patients for 2 to 3 days. | 0
1556600 | a customer reported that a fetal death occurred in 2009 following an amniocentesis procedure. when the doctor first introduced the syringe, the needle was visible at the very beginning, but then disappeared on the voluson-i screen. the doctor used a thicker needle with the second attempt, but the same issue occurred. during the third attempt, the needle again disappeared, and a vein was damaged. the customer alleges that when using this same device two days prior to death, the doctor noticed a drift of 1 centimeter between the actual needle position and the theoretical position on the screen. a ge field service engineer (fe) went onsite to examine the voluson-i device involved in the events. the same probe was used, however, the needle support, needle adapter and needle involved in the reported events were discarded by the hospital, and therefore were unavailable for investigation. the tests for the three angle positions were completed successfully. the fe collected the log files for the associated event and will send to ge healthcare engineering for further investigation. the device is currently in quarantine in the hospital's biomedical department. | 1
1495596 | the customer reported the biometry probe was giving an erroneous measurement (long-sighted). additional information received, stated the patient was under-corrected. >< the company service representative examined the system and probe, and did not find a problem. the root cause of the patient event cannot be determined in this investigation. (b) (4). (b) (4). (b) (4). | 0
1030937 | it was reported that a loss of audible alarms occurred at the cic. it is unk at this time whether the cic was supporting bedside or telemetry pts. no pt injury was reported. ge healthcare's investigation into the reported occurrence is still ongoing. a f/u report will be issued when the investigation has been completed. | 1
1378704 | device is currently being evaluated but no conclusive results are available for this report. a follow up report will be issued with a conclusion once the evaluation is complete. >< an oem customer reports smelling an odor eminating from the printer in 2009. the customer could not identify the source of the odor. sample units rejected by the oem for this reason were sent to company for investigation. the investigation is not complete but a follow up report will be issued upon completion. | 1
1051477 | it was reported there was a loss of audible alarms at multiple cic's supporting telemetry pts. no pt injury was reported. ge healthcare's investigation into the reported occurrence is still ongoing. a f/u report will be issued when the investigation has been completed. | 1
1563960 | when processing 3-d reconstructions of left atrium for electrophysiology studies and ablation on the leonardo workstation the resolution has been unacceptable to our clinicians. siemens has recently informed us that the software on this 10 month old workstation is an earlier, lower resolution, version, which captures images every 1.5 degrees. a newer version of the workstation software apparently offers the higher resolution, which captures images every 0.8 degrees, which would be required to process these images adequately on our system. the issue was discussed with staff at the siemens training center in north carolina. according to them, seven months after the install, the low resolution software is no longer available on small detectors like ours because it does not allow for adequate image quality. as this is an adjunctive imaging tool it is not essential to performing the case. atrial fibrillation ablations are usually done without this technology in many centers that have not purchased the system. the goal is to superimpose a three-dimensional shell onto the x-ray screen for better catheter navigation and to guide transeptal puncture, with the hopes of reducing fluoroscopy time and improving patient safety. atrial fib ablations can continue to be safely performed without this technology, we are just not getting the added safety benefit that was the intention of the purchase. instead, by trying to use the technology as intended we are exposing the patient to risk with pulmonary artery catheterization, contrast injection, additional anesthesia time and additional radiation exposure we would not otherwise do in an af ablation, with often unusable results. in these instances there is no benefit to offset this accumulation of small risks. | 1
1449491 | patient was in a code - radiology had responded to bedside - a chest x-ray was requested. the dr x-ray detector was under the patient, and a chest image was acquired. before the tech could remove the dr plate from under the child, the physician called for defibrillation. tech said they needed to remove the plate first but defibrillation occurred anyway. the x-ray unit is not working properly now. | 0

## 2010
1668689 | A POST OPERATIVE PT UNDERWENT THE PHYSICIAN ORDERED LEFT LOWER EXTREMITY DOPPLER TEST TO RULE OUT A DEEP VEIN THROMBOSIS. THE REPORT THAT THE FINDINGS WERE WITHIN NORMAL LIMITS WAS AVAILABLE WITHIN HOURS. ELEVEN DAYS LATER, LONG AFTER THE PT WAS DISCHARGED, A REPORT OF A CT SCAN PERFORMED ON THE SAME DAY ON THE SAME EXTREMITY ARRIVED, INDICATING SUPERFICIAL THROMBOSIS. THE SCAN WAS NOT ORDERED BY ANY PHYSICIAN. CLARITY AS TO HOW THE ORDER APPEARED SHORTLY AFTER THE DOPPLER TEST WAS UNAVAILABLE. THE SCAN WAS COMPLETED BUT DISAPPEARED AND WAS NOT READ UNTIL IT REAPPEARED 11 DAYS LATER. CLARITY IS UNAVAILABLE AS TO HOW THE CT DISAPPEARED FOR 11 DAYS AND WAS NOT INTERPRETED OVER THAT TIME FRAME. THE DIAGNOSIS WAS NOT LIFE THREATENING. THE PHYSICIANS ATTENDING THIS PT WERE NOT AWARE IT WAS DONE. CONSEQUENTLY, IT WAS NOT MISSED. | 1
1648788 | THE TUBING BROKE DURING HIGH PRESSURE INJECTION. THE TUBING WAS REMOVED AND REPLACED WITH ANOTHER SET. NO REPORT OF HARM OR INJURY. CUSTOMER REPORTED THIS HAD HAPPENED ONE TIME BUT RETURNED TWO SAMPLES, EACH HAD A BROKEN ROTATOR. THIS IS REPORT 2 OF 2 FOR THIS EVENT. REFERENCE: 1721504-2010-00074. >< DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED. THE ROTATOR WAS SEPARATED AT THE BONE BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND ONE SIMILAR COMPLAINT FOR THIS LOT NUMBER ASSOCIATED WITH THIS EVENT. THE ROOT CAUSE OF THE MALFUNCTION IS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATA BASE WAS REVIEWED. | 0
1648787 | THE TUBING IS BACKING OFF OF THE MEDRAD WHEN THEY GO TO POWER INJECT DURING A LEFT VENTRICULOGRAM PROCEDURE. NOBODY WAS SPRAYED. NO REPORT OF HARM OR INJURY. >< DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED. THE THREADS ON THE FEMALE LUER WERE DAMAGED. DAMAGE IS SIMILAR TO DAMAGE CAUSED BY OVER-TIGHTENING THE CONNECTION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATA BASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. A FOLLOW-UP REPORT WILL BE SENT IF MORE OR NEW INFO BECOMES AVAILABLE. DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATA BASE WAS REVIEWED. | 0
1648796 | DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED. THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND TWO SIMILAR COMPLAINTS FOR THIS LOT NUMBER THAT ARE ASSOCIATED WITH THIS EVENT. THE ROOT CAUSE OF THE MALFUNCTION IS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. >< DISTAL END OF THE TUBING DISCONNECTED FROM THE MAIN LINE. BLOOD SPLATTERED ON THE CEILING. NO REPORT OF HARM OR INJURY. CUSTOMER REPORTED THIS HAS HAPPENED THREE TIMES WITH THIS BATCH OF TUBING. THIS IS REPORT 1 OF 3 FOR THIS EVENT. REFERENCE: 1721504-2010-00072, 1721504-2010-00073. | 0
1648793 | DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED. THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND TWO SIMILAR COMPLAINTS FOR THIS LOT NUMBER THAT ARE ASSOCIATED WITH THIS EVENT. THE ROOT CAUSE OF THE MALFUNCTION IS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. >< DISTAL END OF THE TUBING DISCONNECTED FROM THE MAIN LINE. BLOOD SPLATTERED ON THE CEILING. NO REPORT OF HARM OR INJURY. CUSTOMER REPORTED THIS HAS HAPPENED THREE TIMES WITH THIS BATCH OF TUBING. THIS IS REPORT 2 OF 3 FOR THIS EVENT. REFERENCE: 1721504-2010-00071, 1721504-2010-00073. | 0
1473777 | THE USER RECEIVED ERRATIC READINGS FOR ONE PATIENT'S URINE SAMPLE ON THREE 'BACK TO BACK' TESTS. OF THE DATA PROVIDED,THE RESULTS FOR ERYTHROCYTES WERE DISCREPANT. THE RESULTS WERE NEGATIVE, NEGATIVE AND 50 ERYTHROCYTES PER UL. NO INFORMATION WAS PROVIDED TO DETERMINE IF ERRONEOUS RESULTS WERE REPORTED OR IF THE PATIENT WAS ADVERSELY AFFECTED. IT WAS NOTED THE USER WAS USING AN IMPROPER TECHNIQUE OF USING A DROPPER TO APPLY THE URINE TO THE STRIP INSTEAD OF DIPPING THE STRIP INTO THE URINE. | 0
1420079 | PACS RA1000 (B) (4) CT IMAGES FLIPPED LEFT TO RIGHT. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. >< SOFTWARE WILL BE UPDATED TO INCREASE THE VISIBILITY OF THE IMAGE ORIENTATION MARKER WITH SQUARE BORDER AND PROVIDE TWO ADD'L MARKERS ON NORTH AND WEST SIDES OF THE IMAGES. (B) (4). | 1
1661773 | AT 1:15 P.M., ONE OF TWO CARDIAC MONITOR SCREENS AT THE NURSING STATION ON THE UNIT WENT BLANK AND STAFF WAS UNABLE TO REBOOT THE MONITOR. THERE WERE 18 PATIENTS ON THE UNIT AT THE TIME; 16 PATIENTS WERE IMMEDIATELY RECONFIGURED TO THE 2ND CARDIAC MONITOR AT THE NURSING STATION AND 2 PATIENTS WERE PLACED ON BEDSIDE CARDIAC MONITORS. THERE WERE NO ADVERSE PATIENT OUTCOMES. BIOMED ENGINEERING RESPONDED IMMEDIATELY AND FOUND UNIT (DISPLAY) FROZEN. | 1
1789312 | NUMBNESS AND CRAMPING OF BOTH HANDS AND FEET. TROUBLE SLEEPING. TROUBLE WALKING. PINS AND NEEDLES IN ALL LIMBS. DOSE OR AMOUNT: MINIMAL. FREQUENCY: 2X. ROUTE: ORAL. DATES OF USE: 1996 - PRESENT. DIAGNOSIS OR REASON FOR USE: DENTURE WEARER. EVENT ABATED AFTER USE: YES. | 0
1861293 | THE PHILLIPS MONITORING CABLES THAT CONNECT THE PT'S TELEMETRY LEADS TO THE PHILLIPS MONITORS WERE FAILING. WHEN WE CALLED TO REPORT THE PROBLEM WE HAD IN 8 OUT OF 10 OF THE CABLES WE HAD IN STOCK, THE PHILLIPS COMPANY REP WAS AWARE OF IT AND BROUGHT 10 REPLACEMENT CABLES THE NEXT DAY. THE FIELD ENGINEER SAID THE PROBLEM WAS IN THE CABLES, NOT WITH THE MONITORS, WHICH HAD ALSO BEEN OUR EXPERIENCE. | 1
1591096 | THE USER REPORTED THAT THEY HAD A PT CODE AND THE DEVICE'S NIBP PARAMETER DID NOT WORK. >< THE DEVICE MANUFACTURER'S FIELD SERVICE REPRESENTATIVE (FSE) VISITED THE CUSTOMER SITE TO EVALUATE THE DEVICE. THE FSE WAS ABLE TO DUPLICATE THE REPORTED PROBLEM AND REPLACED THE DEVICE'S RADIO AS A RESULT OF HIS EVAL. ALTHOUGH THE USER REPORTED A PROBLEM WITH THE NIBP PARAMETER, THE ACTUAL PROBLEM WAS WITH THE DEVICE'S RADIO. THE PROBLEM WITH THE DEVICE'S RADIO CAUSED THE REQUESTED NIBP READING TO NOT BE TRANSMITTED FROM THE DISPLAY MONITOR (WHERE IT IS INITIATED) TO THE BASE MONITOR (WHICH ACTUALLY TAKES THE NIBP READING). FOLLOWING REPLACEMENT OF THE RADIO, THE DEVICE WAS FUNCTIONALLY TESTED PER INTERNAL TEST PROCEDURES AND RETURNED TO SERVICE BY THE USER. WE ARE CONSIDERING THAT THERE WAS A NEED FOR EMERGENT CARE BECAUSE THE PT WAS REPORTED AS HAVING CODED. THEREFORE, THIS IS CONSIDERED AS A SERIOUS INJURY. THE AVAILABLE INFO GIVES NO INDICATION THAT THE INABILITY OF THE USERS TO OPERATE THE NIBP FUNCTION WAS A FACTOR IN THE CODE. IF THIS MALFUNCTION WERE TO RECUR, IT WOULD NOT LEAD TO A SERIOUS INJURY OR DEATH BECAUSE THIS MALFUNCTION WOULD BE OBVIOUS TO THE USER. IF THE MALFUNCTION WERE TO RECUR, THE USER IS LIKELY TO UTILIZE A BACK-UP MONITOR TO PROCEED WITH THE CASE. IF A BACK-UP MONITOR WAS UNAVAILABLE AND LACK OF A BACK-UP MONITOR WAS TO PRECLUDE THE USE OF THE SCANNING DEVICE IN DIAGNOSIS, ALTERNATIVE DIAGNOSTIC METHODS WOULD BE IMPLEMENTED AS WARRANTED FOR THE PT. | 1
1595407 | AS THE BEDSIDE RN WAS FLUSHING THE DIGNICARE IT BURST. THE LINE WAS FLUSHABLE, WITH LITTLE RESISTANCE. PICTURE TAKEN OF BROKEN PORT. | 0
1734535 | WE CURRENTLY UTILIZE A PATHOLOGY REPORTING SOFTWARE FROM PSYCHE SYSTEMS-(B) (4), (B) (4). THIS SOFTWARE WAS INSTALLED SPRING 2009 WITH GO-LIVE IN (B) (6) 2009. WE HAVE EXPERIENCED REPEATED ERRORS AND PRINTING COMPLICATIONS. SIGNIFICANT PROBLEMS INCLUDE: PRINTING A REPORT WHILE IT WAS IN 'UNSIGNED' STATUS AND UNDERGOING DIAGNOSIS REVISION BY THE PATHOLOGIST. IN THIS INSTANCE, TWO REPORTS WITH TWO DIFFERENT DIAGNOSES PRINTED TO THE PHYSICIAN OFFICE. THERE WAS NO SAFETY NET IN PLACE THAT PREVENTED EDITING OF A 'COMPLETED' CASE. ADMISSION FROM A PSYCHE PROGRAMMER THAT HE HAS BEEN INVESTIGATING AN 'INVALID HANDLE ERROR' FOR 'SOME TIME' - STATED IN (B) (6) 2010 - ISSUE STILL EXISTS-. THE 'HANDLER ERROR' INTERRUPTS THE PRINT JOBS FOR THE PATHOLOGY REPORTS DURING THE 'COMPLETE' STATUS. AT TIMES, THE PRINTING APPEARS TO BE WORKING BUT REPORTS ARE GOING TO THE WRONG PROVIDER. THIS IS CLEARLY A POTENTIAL HIPAA ISSUE. ONGOING ATTEMPTS BY PSYCHE TO PATCH THE 'BUGS' HAS RESULTED IN NEW ERRORS AND SUBSEQUENT NEW PATCHES. AS OF 06/17/2010, OUR DEPARTMENT CEASED USING THE PRINT FUNCTION AND RESORTED TO FAXING REPORTS TO PHYSICIANS AS A TEMPORARY MEASURE. RESPONSE TO ISSUES SUBMITTED TO THE PSYCHE HELP DESK HAS NOT BEEN TIMELY RE: PRINTING ISSUES - HIPAA AND POTENTIAL DELAY IN TREATMENT RISKS-. | 1
1758785 | TELEMETRY UNIT CARDIAC MONITORING SERVER MALFUNCTIONED FOR APPROXIMATELY 45 MINUTES. TWELVE PATIENTS WERE AFFECTED AND HAD TO BE MONITORED BY NURSING STAFF WHILE BIO-MEDICAL ENGINEERING HAD TO REBOOT THE SYSTEM AND GET THE SERVER TO FUNCTION AGAIN. THERE WERE NO PATIENT ADVERSE OUTCOMES. | 1
1658817 | INVESTIGATION PENDING. >< CUSTOMER REPORTED LDX PRINTER CAUGHT ON FIRE. NO INJURY OR DAMAGE TO SURROUNDING AREA WAS REPORTED. | 1
1544799 | THE COMPLAINANT REPORTED THAT DURING A PROCEDURE, THE ROTATOR OF THE SUSPECT TUBING BLEW OFF. THE INCIDENT OCCURRED DURING INJECTION AT APPROX 600 PSI. >< DEVICE EVAL: THE SUSPECT DEVICE WAS RETURNED FOR EVAL. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED; THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. SINCE A LOT NUMBER WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORD AND COMPLAINT DATA COULD NOT BE PERFORMED. THE ROOT CAUSE OF THE MALFUNCTION WAS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. | 0
1873592 | A CUSTOMER REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED WHEN FOUR (4) CIC (CENTRAL STATIONS) REBOOTED SIMULTANEOUSLY. NO INJURY WAS REPORTED. THIS IS THE THIRD OF FOUR REPORTS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. >< THE SERIAL NUMBER OF THIS DEVICE IS UNK AT THIS TIME. THE DEVICE MANUFACTURE DATE CANNOT BE DETERMINED AT THIS TIME, WITHOUT THE SERIAL NUMBER OF THE DEVICE. | 1
1829408 | A CUSTOMER REPORTED A LOSS OF MONITORING AT THE CIC (CENTRAL STATION) DUE TO A SUSPECTED HARD DRIVE FAILURE. NO DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1317960 | IT WAS REPORTED THAT A LOSS OF MONITORING OCCURRED IN THE CARE UNIT INVOLVING A CIC SYSTEM. THE LOSS OF MONITORING REPORTEDLY LASTED MORE THAN 30 SECONDS AND INCLUDED UNEXPECTED TIME CHANGES ON BEDSIDE MONITORS, 'LOW RESOURCE' MESSAGES ON THE CIC, AND THE DISAPPEARANCE OF HARDWIRED MONITORS FROM THE CIC. NO PATIENT INJURY WAS REPORTED. >< GE HEALTHCARE INVESTIGATED THE REPORTED ISSUE AND DETERMINED THAT THE EVENT IS RELATED TO A KNOWN CONDITION WHERE THE TIME MASTER ON THE UNITY NETWORK BROADCASTS TIME CHANGES AND THE CIC REFLECTS THESE CHANGES, CAUSING MULTIPLE TIME REQUESTS AND EXCESSIVE NETWORK TRAFFIC. THIS CAN RESULT IN THE FOLLOWING SCENARIOS: SLUGGISHNESS OR UNRESPONSIVENESS OF THE CIC PRO OR MONITORING DEVICE USER INTERFACES. LOSS OF WAVEFORMS, PARAMETERS AND/OR ALARMING ON THE CIC PRO. RESTARTING OF CIC PRO PROCESSES OR REBOOTING OF THE CIC PRO OPERATING SYSTEM. TIME CHANGING RAPIDLY BETWEEN TWO OR MORE TIMES ON UNITY DEVICES. TIME NOT ADVANCING ON UNITY DEVICES. INCONSISTENT TIME VALUES ON DEVICES ACROSS THE UNITY NETWORK. REBOOTING OF MONITORING DEVICES. EXCESSIVE NETWORK TRAFFIC THAT DELAYS, OR OTHERWISE INTERFERES WITH, LEGITIMATE UNITY NETWORK COMMUNICATION. GE HEALTHCARE WILL BE PROVIDING A PRODUCT UPGRADE TO PREVENT THE POTENTIAL FOR ISSUES ASSOCIATED WITH NETWORK TIME CHANGES ON THE CARESCAPE CIC PRO, SOFTWARE VERSIONS 4.1, 5.0.3, 5.0.6, 5.0.7 AND 5.0.8. (B) (4). (B) (4). | 1
1350453 | IT WAS REPORTED THAT 14 SOLAR 8000 MONITORS FROZE SIMULTANEOUSLY. DURING THIS TIME, WAVEFORMS STOOD STILL AND ALL OF THE TRIM KNOBS ON THE SOLAR 8000 MONITORS WERE UNRESPONSIVE. ALL 14 MONITORS WERE BEING VIEWED AT A SINGLE CIC. THE ISSUE WAS RESOLVED TEMPORARILY WHEN THE CARE UNIT WAS ISOLATED FROM THE REST OF THE NETWORK. THE ISSUE RECURRED 4 DAYS LATER AND WAS RESOLVED BY REPLACING THE CIC IN QUESTION WITH A LOANED CIC. THERE HAVE BEEN NO RECURRENCES OF THE ISSUE SINCE THE CIC WAS REPLACED. BASED UPON THE INFORMATION AVAILABLE AT THIS TIME, IT IS PRESUMED THAT THE CIC MAY HAVE CONTRIBUTED TO NETWORK DISTURBANCE SEEN AT THE SOLAR MONITORS. THERE WAS NO REPORTED PATIENT INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1663656 | PUNCTSURE UNIT WAS PLUGGED IN TO CHARGE IN BETWEEN USES. NURSE WALKED BY AND SMELLED SOMETHING HOT. AFTER A QUICK SEARCH IT WAS DETERMINED THAT THE SMELL WAS ORIGINATING AT THE CHARGER TO THE PUNCTSURE UNIT. THE NURSE UNPLUGGED THE UNIT AND CALLED BIOMEDICAL ENGINEERING. NO INJURY TO NURSE OR PATIENT. BIOMED PICKED UP THE DEVICE AND DETERMINED THE SMELL WAS COMING FROM THE CHARGER. UPON OPENING THE DEVICE IT WAS FOUND THE CIRCUIT BOARD WAS CHARRED AND BURNED. A REPLACEMENT WAS ORDERED AND INSTALLED. THE PUNCTSURE WAS RETURNED TO SERVICE.====================== HEALTH PROFESSIONAL'S IMPRESSION======================POSSIBLE FIRE HAZARD. | 1
1764183 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO TOTAL BETA HUMAN CHORIONIC GONADOTROPIN 2 (TBHCG2) RESULT OF >1000 MIU/ML WAS MODIFIED BY THE DL2000 DATA MANAGEMENT TO 0.0 MIU/ML. THE RESULT WAS REPORTED OUT OF THE LABORATORY. THE MEDICAL STAFF QUESTIONED THE RESULT UPON RECEIPT. THE LAB REVIEWED THE RESULTS AND REPORTED THE DIL-HCG2 (DILUTE HCG) RESULT OF 13573 MIU/ML. PATIENT TREATMENT WAS NOT IMPACTED BY THIS EVENT. >< THE CUSTOMER HAD A RULE TO CONVERT RESULTS LOWER THAN THE LOWER LINEARITY OF THE ASSAY TO 0.00 MIU/ML. THE INFORMATION WAS FORWARDED TO THE VENDOR. THE VENDOR VERIFIED THE ISSUE TO BE SYNTAX ERROR AND PROVIDED THE CORRECT CONFIGURATION OF THE RULE. THE ROOT CAUSE FOR THIS EVENT WAS AN INCORRECTLY CONFIGURED RULE. | 1
1828303 | A CUSTOMER REPORTED A DELAY IN AUDIBLE ALARMS FOR LOW SPO2 ON THE CARESCAPE MONITOR B850. THE PT WAS REPORTEDLY ON A NON-GE VENTILATOR THAT WAS INADVERTENTLY DISCONNECTED FROM THE MAIN POWER SUPPLY AND THEN STOPPED WHEN ITS BATTERIES WERE EMPTY. THE NURSE REPORTEDLY HEARD THE VENTILATOR AUDIBLY ALARMING FOR THE LOW BATTERY CONDITION AND WENT TO THE PT'S ROOM. UPON THE NURSE'S ARRIVAL, THE PT WAS BLUE. THE NURSE REPORTEDLY BEGAN MANUALLY VENTILATING THE PT WITH A BAG, CHECKED THE SPO2 VALUE ON THE B850 SCREEN, AND NOTED THAT THE SCREEN WAS DISPLAYING A YELLOW FRAME WITH THE MESSAGE 'SPO2 LOW 53/54%'. NO AUDIBLE ALARM WAS HEARD AT THAT TIME. THE PT'S SPO2 VALUE WAS BROUGHT BACK TO 92-93%. THE PT WAS REPORTEDLY UNDER SEDATION PRIOR TO AND FOLLOWING THE EVENT, THEREFORE, IT IS UNK WHETHER THE PT SUSTAINED INJURIES FROM THE EVENT. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. >< PT INFO IS CONSIDERED CONFIDENTIAL PER COUNTRY REGULATIONS AND WILL NOT BE PROVIDED TO GE HEALTHCARE. | 1
1818467 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< (B)(4). THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. | 0
1417991 | IT WAS REPORTED THAT THERE WAS 'LINE DETACHMENT AT VAMP RESERVOIR ON THE PATIENT SIDE VERY LIKELY. THE PATIENT LOST BLOOD BUT HE IS FINE. THIS HAPPENED THREE TIMES.' | 0
1424514 | COMPANY SERVICE REPRESENTATIVES HAVE WORKED CLOSELY WITH THIS CUSTOMER TO RESOLVE THE COMMUNICATION LOSS ISSUES, WHICH CONTINUED IN FEBRUARY. IN MOST CASES, THE SYSTEM DOWNTIME WAS NOT SIGNIFICANT, DUE TO THE FACT THAT THE SYSTEM WAS FUNCTIONAL AFTER A REBOOT WAS PERFORMED. SERVICE ACTIVITY INCLUDED: REPLACING ON CENTRAL STATION. REPROGRAMMING ONE TELEPACK. RESETTING THE ELAN. REPLACING A SWITCH IN THE RACK ROOM. CORRECTING MISSPATCHED CENTRAL STATIONS AND BEDSIDE MONITORS. MONITORING NETWORK TRAFFIC AND SWITCHING LOADS TO BETTER HANDLE TRAFFIC. RESEATING A CD ROM DRIVE AND HARD DRIVE CABLES. >< THE CUSTOMER REPORTED THAT WHILE THE SYSTEM WAS IN USE, MONITORING PATIENTS, ALONG WITH BEDSIDE MONITORS AND TELEPACKS, DURING THE CUSTOMER'S 'GO-LIVE' FOLLOWING A NEW INSTALLATION, MULTIPLE CENTRAL STATIONS AND WORKSTATIONS EXPERIENCED EPISODES OF 'COMMUNICATION LOSS'. NO PATIENT INJURY WAS REPORTED. THE CENTRAL STATION SERIAL NUMBERS INVOLVED WERE: (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4). THE WORKSTATION SERIAL NUMBERS WERE (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4), (B) (4). | 1
1624743 | NO MISADMINISTRATION WAS REPORTED IN THIS CASE. THIS ISSUE IS A KNOWN MALFUNCTION, AND APPEARS TO BE A DISPLAY ONLY PROBLEM. IN PREVIOUS REPORTS, WHEN THE MLC SETTINGS HAVE BEEN PHYSICALLY EXAMINED, THEY HAVE BEEN CORRECT. NO MISADMINISTRATION HAS EVER BEEN ATTRIBUTED TO THIS MALFUNCTION. THE 'ORANGE LEAF' DISPLAY AT THE 4D TREAT CONSOLE, HOWEVER, IS INTENDED TO ALERT THE USER THAT THERE IS SOME UNCERTAINTY ABOUT THE MLC LOCATION AND THEREFORE ACTS AS A SAFETY MECHANISM TO HELP PREVENT MISADMINISTRATION. THIS SAFETY MECHANISM HAS FAILED IN THIS MALFUNCTION, WHICH INTRODUCES THE POSSIBILITY FOR MISTREATMENT SHOULD THERE ACTUALLY BE A PROBLEM WITH THE MLC, BUT THE OPERATOR CHOOSES TO IGNORE THE INDICATION DUE TO PRIOR FALSE POSITIVE REPORTS. NO FURTHER FOLLOW-UP TO THIS MDR IS EXPECTED. >< THE USER REPORTED THAT DURING PATIENT TREATMENT, THE 4D TREAT WORKSTATION INDICATED THE MULTI LEAF COLLIMATOR (MLC) LEAVES HAD NOT MOVED INTO POSITION FOR THE NEXT FIELD. THERE WAS NO INTERLOCK TO STOP THEM FROM TREATING WITH WHAT APPEARED TO BE AN INCORRECT MLC SHAPE. NO PATIENT INJURY REPORTED, AND NO PATIENT DATA PROVIDED. | 1
1752476 | STILL BIRTH OF A FULL TERM INFANT. PHILLIPS MEDICAL FETAL MONITOR INDICATED THE INFANT HAD A HEARTBEAT UNTIL THE EXACT TIME OF BIRTH. AT BIRTH THERE WAS NO HEARTBEAT OR RESPIRATORY EFFORT. THERE WAS NO RESPONSE FROM INFANT DESPITE A PROLONGED RESUSCITATION PROCESS. THERE IS NOW A QUESTION AS TO THE ACCURACY OF THE FETAL MONITORING EQUIPMENT. WAS IT RECORDING THE MATERNAL HEART RATE AFTER THE INFANT UNKNOWINGLY DIED IN UTERO? THIS WAS AN UNCOMPLICATED (B) (6) PREGNANCY. AUTOPSY WAS NORMAL EXCEPT FOR MECONIUM PRESENT IN THE LUNGS AS WELL AS HISTIOCYTES THAT HAD INGESTED MECONIUM. ALSO NOTE THAT PHILLIPS HAD RELEASED AN URGENT MEDICAL DEVICE RECALL RELATED TO THIS EQUIPMENT. (B) (4). | 1
1743648 | (B) (4). THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
1506367 | AUDIBLE ALARM TO CHECK A BLOOD GLUCOSE LEVEL DID NOT OCCUR WHEN EXPECTED. NURSES HAD TAKEN THE PATIENT OFF OF THE INSULIN DRIP PROTOCOL, BUT DID NOT DEACTIVATE THE PATIENT IN THE GLUCOCARE SYSTEM. LATE GLUCOSE CHECK WAS NOTICED BY COMPANY REPRESENTATIVE DURING A MONITORING CHECK AND CCU WAS CALLED. NURSE REPORTED THAT NO AUDIBLE ALERTS WERE OCCURRING EVEN THOUGH WEB BROWSER WAS RUNNING. THERE IS A BACKUP ALARM THAT NOTIFIES SUPPORT WHENEVER A GLUCOSE CHECK IS DUE AND THE WEB BROWSER HAS BEEN CLOSED. HOWEVER, IN THIS CASE, THE BROWSER WAS RUNNING, BUT NOT ISSUING THE ALARM, SO THE BACKUP ALARM DID NOT ACTIVATE. >< A REPRESENTATIVE OF PRONIA MEDICAL SYSTEMS DISCOVERED THIS ISSUE DURING ROUTINE MONITORING OF THE SYSTEM THE MORNING OF (B)(6) 2009. WITHIN HOURS OF RECOGNIZING THIS ISSUE, THE NURSING COORDINATORS FOR THE TWO GLUCOCARE SYSTEMS WERE NOTIFIED TO BE VIGILANT FOR ANY LATE ALARMS. IT WAS DETERMINED THAT A RECURRENCE OF THE ISSUE WOULD BE UNLIKELY TO CAUSE SERIOUS INJURY DUE TO THE NORMAL VIGILANCE OF THE NURSING STAFF REGARDING GLUCOSE CHECKS. PRONIA'S COMPLAINT REVIEW PROCESS AT THE TIME DID NOT IDENTIFY THIS AS A REPORTABLE EVENT. THE PROBABLE CAUSE OF THE ALARM FAILURE WAS QUICKLY DETERMINED TO BE EITHER A BUG IN THE INTERNET EXPLORER 6 WEB BROWSER, OR A HEADPHONE PLUGGED INTO THE SPEAKER JACK OF THE COMPUTER IN THE CCU, WHICH MUTED THE SPEAKER. TO ADDRESS THESE OR ANY OTHER ISSUES THAT COULD PREVENT THE AUDIBLE ALARM FROM SOUNDING, WE ADDED AN ADDITIONAL ALERT TO THE SYSTEM. THIS 'CRITICAL ALERT' IS A NOTIFICATION TO OUR MONITORING STAFF (EMAIL AND TEXT MESSAGE) THAT IS SENT WHENEVER ANY NURSING ACTION IS MORE THAN 15 MINUTES LATE. A SOFTWARE UPDATE CONTAINING THIS CRITICAL ALERT FEATURE (V.1.2.23) WAS DEPLOYED TO BOTH SYSTEMS ON (B)(6) 2009. IT WAS SUBSEQUENTLY DETERMINED DURING A ROUTINE AUDIT BY AN OUTSIDE CONTRACTOR, THAT THIS ISSUE MET THE DEFINITION OF A REPORTABLE EVENT. PRONIA HAS REVISED ITS COMPLAINT REVIEW PROCESS TO ENSURE FUTURE INCIDENTS ARE IDENTIFIED IN A TIMELY MANNER. | 1
1617128 | DISCUSSED INCIDENT WITH FEMALE PT AND WITH CAREGIVER. FEMALE PT HAS HAD SOME HAIR GROW BACK BUT NOT COMPLETE. SHE IS ON MEDICATIONS THAT HAVE HAD IMPLICATIONS IN HAIR LOSS, BUT IS RARE. FEMALE PT IS SEEKING FINANCIAL COMPENSATION BUT HAS NOT FILED A LAWSUIT AGAINST US OR AGAINST THE CAREGIVER. >< PATIENT HAD SLEEP STUDY DONE AT (B)(6) BY TECHNOLOGIST ON (B)(6), 2009. AFTER SLEEP STUDY AND AT HOME, PT WASHED HER HAIR AND LATER NOTICED WHERE ELECTRODES WERE PLACED, SHE HAD HAIR LOSS. THIS OCCURRED IN TWO ELECTRODES SITES AT THE BACK OF THE HEAD. FEMALE PATIENT WENT TO DERMATOLOGIST ON (B)(6) 2009 AND HAS RECEIVED STEROID SHOTS TO STIMULATE HAIR GROWTH. THIS HAS BEEN MET WITH SOME SUCCESS. | 0
1813877 | TWICE, I HAVE PURCHASED FROM (B)(6) AN OMRON BLOOD PRESSURE MONITOR, MODEL # HEM 7901T. THE (B)(6) WEB PAGE CLEARLY STATES WITHOUT DIFFERENTIATION 'ADVANCED OMRON HEALTH MANAGEMENT SOFTWARE.' I HAVE A SCREEN GRAB OF THIS WEBPAGE. NOWHERE COULD I FIND THAT THIS SOFTWARE ONLY RUNS USING WINDOWS, NOT ANY VERSION OF APPLE MAC'S OS-X. THIS IS NOT DISCLOSED UNTIL ONE READS THE VERY FIND PRINT ON THE BOX AFTER RECEIPT OF THE DEVICE OR WHEN EXAMINING THE DISK FOR THE SOFTWARE. THE (B)(6) PAGE ALSO SHOWS THE DEVICE AS 'CLEARED BY THE FDA TO HELP DETECT MORNING HYPERTENSION' AS THE VERY FIRST 'FEATURE.' I TELEPHONED (B)(6) TO COMPLAIN, AND THEIR CUSTOMER SERVICE REP PROMISED TO SEND ME A DEVICE WITH THE CORRECT SOFTWARE. A SECOND DEVICE ARRIVED YESTERDAY, AND IT ALSO WILL NOT RUN ON AN APPLE MAC USING OS-X. I BELIEVE YOUR APPROVAL OF THIS DEVICE WITH THE ACCOMPANYING SOFTWARE WAS OBTAINED BY FRAUD BECAUSE THE COMPANY OMRON DID NOT DISCLOSE TO YOU THAT THEY DO NOT OFFER THE DEVICE WITH OS-X. IT WAS ALSO FRAUDULENT FOR (B)(6) TO SELL THE DEVICE WITHOUT CLEARLY DISCLOSING IN AS LARGE TYPE AS THE PROMOTION OF THE SOFTWARE THAT IT WOULD ONLY RUN USING WINDOWS. MOREOVER, BECAUSE MILLIONS OF THE RESIDENTS OF THE USA NOW USE APPLE MAC'S WITH OS-X AND THEY NEED DEVICES TO RUN ON THOSE COMPUTERS, NO DEVICE SHOULD BE PERMITTED FOR CONSUMER SALE THAT DOES NOT RUN ON THE APPLE MAC WITH OS-X. I WAS RECENTLY DIAGNOSED WITH HIGH BLOOD PRESSURE AND STARTED TAKING MEDICATION FOR THAT PURPOSE. TODAY, MY BP WAS TESTED IN A DOCTOR'S OFFICE AT 168/84, WHICH IS HIGH, AND THUS MY INABILITY TO USE THIS DEVICE IS LIFE THREATENING TO ME. THANK YOU. | 0
1582400 | INITIAL REPORTER CALLED AND MENTIONED THAT THEY NOTICED THAT THE SERVER WAS FROZEN AND WAS NOT REACTING TO ANY COMMANDS ON KEYBOARD AND MOUSE. USER ALSO MENTIONED THAT THERE WAS 1 GREEN LIGHT ON THE SERVER. SHE THEN REBOOTED THE SERVER BY HITTING THE WHITE BUTTON ON THE SERVER AND TRIED TO RESTART IT. SHE MENTIONED THAT SHE DID THIS A FEW TIMES AND THE SERVER WAS NOT COMING UP. THERE WAS NO DATA LOSS AS A RESULT OF THIS MALFUNCTION. MT - 7976, (B) (4). >< THERE IS NO ESTABLISHED EXP DATE FOR THIS DEVICE. TECHS WENT OUT AND EVALUATED THE DEVICE IN THE FIELD ON 12/16/2009. EVAL SUMMARY - THE SAID SERVER WAS EVALUATED IN THE FIELD. AFTER FURTHER INVESTIGATION, IT WAS DISCOVERED THAT THE LIKELY CAUSE OF THE MALFUNCTION WAS DEFECTIVE HARDWARE. ALTHOUGH, THERE WAS POTENTIAL FOR PT DATA LOSS, NO DATA WAS LOST. FURTHER INVESTIGATION IS ONGOING VIA CAPA (B) (4). THIS IS NOT A SINGLE USE DEVICE. | 1
1419888 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE GAVE A 'NO SHOCK ADVISED' PROMPT FOR A RHYTHM BIOMED BELIEVED WAS SHOCKABLE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1783041 | A CHEST COMPUTERIZED TOMOGRAPHY SCAN AND SINUS RADIOGRAPHS WERE ORDERED. THE ORDERS WERE ENTERED USING MEDITECH ELECTRONIC ORDERING DEVICE. RADIOGRAPHS WERE ACQUIRED AND STORED ON A FUJI PACS DEVICE. THE RESULTS WERE REPORTED AS FOLLOWS: A CHEST RADIOGRAPH REPORT WAS PROVIDED; HERE WERE NOT ANY REPORTS OF A CHEST CT OR SINUS RADIOGRAPH. THE PT DID NOT HAVE A PLAIN CHEST RADIOGRAPH. WHAT WAS REPORTED WAS A CHEST RADIOGRAPH OF ANOTHER PT. THE TESTS THAT WERE PERFORMED ON THIS PT WERE NOT REPORTED UNTIL 24 HOURS LATER, DELAYING CARE FOR AN INFECTION. | 1
1636921 | DEVICE EVALUATION.CUSTOMER REPORT WAS CONFIRMED. REC'D (1) DOUBLE DPT - VAMP ADULT KIT. KIT APPEARED NOT USED. TUBING WAS COMPLETELY DETACHED FROM CONNECTION WITH VAMP RESERVOIR. NO VISIBLE INDICATION OF UV ADHESIVE WAS FOUND ON BONDING SURFACES OF TUBING AND RESERVOIR TUBE HOUSING. IT APPEARED THAT UV ADHESIVE HAD NOT BEEN APPLIED TO THE BOND JOINT. (B)(4) >< THE FOLLOWING WAS REPORTED: 'THE REDLINE WAS NOT CONNECTED PROPERLY, UNABLE TO USE.' | 0
1363077 | CUSTOMER INQUIRED ABOUT WHAT HAPPENS TO RESULTS FOR CPEP AND IRI IF EDTA OR CITRATED PLASMA IS USED INSTEAD OF SERUM 'HEPARINIZED PLASMA.' CUSTOMER ALSO INFORMED TOSOH BIOSCIENCE THAT ONLY RECENTLY THEY HAD BECOME AWARE THAT THEY WERE RECEIVING EDTA/CITRATED PLASMA SPECIMENS FROM SOME OF THE DOCTORS. FOR MORE THAN A YEAR THEY DID NOT NOTICE THAT EDTA/CITRATED PLASMA WAS SENT AND THEY WERE TESTING AND REPORTING RESULTS ON THE AIA-1800. >< CUSTOMER INQUIRED ABOUT WHAT HAPPENS TO RESULTS FOR CPEP AND IRI IF EDTA OR CITRATED PLASMA IS USED INSTEAD OF SERUM. CUSTOMER ALSO INFORMED TOSOH BIOSCIENCE THAT ONLY RECENTLY THEY HAD BECOME AWARE THAT THEY WERE RECEIVING EDTA/CITRATED PLASMA SPECIMENS FROM SOME OF THE DOCTORS. FOR MORE THAN A YEAR THEY DID NOT NOTICE THAT EDTA/CITRATED PLASMA WAS SENT AND THEY WERE TESTING AND REPORTING RESULTS ON THE AIA-1800. PAGE 4 OF BOTH THE ST CPEP AND ST IRI INSTRUCTIONS FOR USE ARE ATTACHED FOR REVIEW. THESE PAGES CONTAIN THE SPECIMEN COLLECTION AND HANDLING INSTRUCTIONS WHICH STATES, 'SERUM OR HEPARINIZED PLASMA IS REQUIRED FOR THIS ASSAY. EDTA AND CITRATED PLASMAS SHOULD NOT BE USED.' NO PT INJURIES OR ANY MISDIAGNOSIS HAVE BEEN REPORTED TO TOSOH BIOSCIENCE. (B)(4). | 1
1673425 | AN ERROR WAS DETECTED IN THE SOFTWARE OF THE MIDDLEWARE INSTRUMENT INTERFACE. ERROR IS A DECIMAL POINT PLACEMENT ERROR. THE ERROR CAUSED THE IMMATURE GRANULOCYTE RESULT, IN A COMPLETE BLOOD COUNT PT BLOOD SAMPLE TO BE OFF BY A FACTOR OF 10. ERROR HAS BEEN TRACED TO (B) (6) 2009 AT 10:45 AM. NO OTHER PARAMETER IN THE COMPLETE BLOOD COUNT IS AFFECTED - ONLY IMMATURE GRANULOCYTES. VENDOR HAS CORRECTED THE SOFTWARE PROBLEM - (B) (6) 20/10-. | 1
1912502 | DURING AN ISCHEMIC VT CASE, THE CARTO 3 SYSTEM FAILED TO CORRECTLY ROUTE PACING RESULTING IN 3 TOTAL CARDIOVERSIONS THAT WOULD HAVE OTHERWISE BEEN UNNECESSARY. THE FIRST INSTANCE RESULTED WHEN THE PHYSICIAN COULD NOT GET THE ABLATION CATHETER TO PACE DUE TO FAILED ROUTING. THE SECOND INSTANCE OCCURRED AS THE PHYSICIAN COULD NOT GET THE PACING ON THE ABLATION CATHETER TO DISABLE THROUGH CARTO 3 AND RF COULD NOT BE APPLIED. THE PATIENT REQUIRED CARDIOVERSION DUE TO BECOMING UNSTABLE WHILE IN THE CLINICAL VT. THE THIRD INSTANCE OCCURRED WHEN THE USER COULD NOT ABLATE DUE TO A DEFECTIVE REDEL CABLE THAT WAS READING LOW TEMPERATURES AND RF (RADIOFREQUENCY) COULD NOT BE TURNED ON. A CABLE CHANGE RESOLVED THE ISSUE. VENDOR NOTIFIED WITH A REQUEST FOR INVESTIGATION INTO THE CARTO 3 ISSUES, ESPECIALLY THE PACE ROUTING. THE REP WAS PRESENT IN THE ROOM DURING THE CASE. THE PATIENT STABILIZED DESPITE THESE EQUIPMENT MALFUNCTIONS. THE PATIENT'S OUTCOME WAS GOOD. PRODUCTS IN USE:THE CARTO 3 SYSTEMTHE STOCKERT GENERATORTHE COOL FLOW PUMPTHE NAVISTAR THERMOCOOL CATHETER.====================== HEALTH PROFESSIONAL'S IMPRESSION======================DURING ISCHEMIC VT ABLATION CASE, THE CARTO 3 SYSTEM FAILED TO CORRECTLY ROUTE PACING RESULTING IN 3 CARDIOVERSIONS THAT WOULD HAVE OTHERWISE BEEN UNNECESSARY. THE PHYSICIAN IS AN EXPERIENCED INTERVENTIONALIST AND STATED 'THE OLD SYSTEM I COULD MANAGE. WITH THE UPGRADES, IT'S IMPOSSIBLE FOR ME TO DO A CASE WITHOUT THE REP IN THE ROOM.'====================== MANUFACTURER RESPONSE FOR EP MAPPING SYSTEM, CARTO 3 SYSTEM======================INVESTIGATING | 1
1587327 | THE CUSTOMER REPORTED ON 07JAN10 THAT A MALE PT HAD A CARDIAC ARREST AND SUBSEQUENTLY EXPIRED. THE ISSUE OCCURRED UNEXPECTEDLY. WE WILL CONSIDER THAT THE DEVICE WAS A FACTOR IN THIS INCIDENT AS STAFF FOUND THE PT UNRESPONSIVE AND ALARM LOGS SHOW THE ALARMS BEGAN AT THE MOMENT THE PT WAS DISCOVERED, AS WOULD BE THE CASE IN A LEADS OFF OR BATTERY RELATED TELEMETRY INCIDENT. THE CUSTOMER COULD NOT COMMENT ON WHETHER THE LEADS WERE ON OR WHETHER THE BATTERY WAS WORKING AT THE TIME THE PT WAS FOUND. THE CUSTOMER SAID, THEY WERE CONDUCTING THEIR OWN INTERNAL INVESTIGATION AND DID NOT SAVE THE PT'S DATA TO BE ABLE TO SEE THE WAVEFORMS TO UNDERSTAND IF MONITORING WAS STILL OCCURRING OR NOT. STAFF SAID, THE PT WAS NOT HAVING LOTS OF ALARMS BEFORE THE INCIDENT. SHE DID NOT ASK IF THE LEADS WERE ON OR OFF- NO ONE MENTIONED THIS. SHE DOES NOT KNOW IF THE BATTERY WAS OK OR NOT IN THE TELE PACK. I EXPLAINED THAT THE LOGS CAPTURE RED ALARM EVENTS AND IN THIS CASE, THE ONLY ALARMS ARE THOSE THAT ARE OCCURRING AT 17:25 WHICH IS THE EXACT TIME SHE IS SAYING THE PT WAS FOUND. BASED ON THE CONDITION OF THE PT, IT WAS SUSPECTED THAT 17:25 WAS NOT THE FIRST OCCURRENCE OF A PROBLEM. PRODUCT LABELING (INSTRUCTIONS FOR USE) ADEQUATELY DESCRIBES ALARMING AND ACKNOWLEDGEMENT OF ALARMS. NO FURTHER INVESTIGATION OR ACTION WARRANTED. >< THE CUSTOMER WAS REQUESTING INFO DURING A SPECIFIC TIMEFRAME FOR A PT WHO HAD EXPIRED. | 1
1727176 | AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND PERFORMED THE READER CAMERA ALIGNMENT AND OPTICS ADJUSTMENTS. THE FE ALSO CHECKED THE GRIPPER ALIGNMENT FOR CARDS FROM THE CENTRIFUGE AND INCUBATOR. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION >< THE CUSTOMER REPORTED THAT THE GEL CAMERA MISREAD A WEAK + REACTION AS A NEGATIVE IN THE ANTI-D MICROWELL OF A TYPE AND SCREEN TEST. | 1
1766220 | APPROXIMATELY FIVE OR SIX PATIENT URINE SAMPLES GENERATED A FALSE NEGATIVE RESULT FOR LEUKOCYTES. UPON MICROSCOPIC REPEAT, EACH SAMPLE GAVE A POSITIVE READING VARYING FROM 25 TO GREATER THAN 100 LEUKOCYTES/UL. THE INITIAL RESULTS WERE NOT REPORTED. THE FIELD SERVICE REPRESENTATIVE DETERMINED THE STRIP GUIDE WAS BENT. THE INSTRUMENT WAS REPAIRED. | 1
1942567 | THE PATIENT WAS SITTING, WITH LEGS UNDER THE X-RAY DETECTOR. THE PT'S ARM WAS ON TOP OF THE DETECTOR. THE PT LEANED THEIR BODY WEIGHT ONTO THE DETECTOR. THE MACHINE SENSED A TOP COLLISION, ALERTED THE COLLISION, AND AUTOMATICALLY TRIED TO CLEAR THE COLLISION BY DRIVING THE DETECTOR DOWNWARD, PINNING THE PT'S LEG. UPON ALERTING FOR A COLLISION, THE MACHINE DISPLAYS 'PRESS AND HOLD MOVE BUTTON.' WHEN THE OPERATOR PRESSES THE UP ARROW IN ATTEMPTS TO MOVE THE DETECTOR UP AND AWAY FROM THE PT'S LEG, THE MACHINE ACTUALLY MOVES DOWNWARD EVEN MORE. THIS CAUSED THE MACHINE TO HAVE A DOWNWARD COLLISION WHICH THEN WOULD NOT ALLOW THE MACHINE TO MOVE UP OR DOWN. THE STAFF HAD TO MANUALLY LIFT UP ON THE MACHINE TO FREE THE PATIENT. NOTE: THE DOWN COLLISION DETECTOR APPEARS TO WORK IN THE SAME WAY, THAT IS, IF IT SENSES A DOWNWARD COLLISION, THE MACHINE WILL DRIVE UP A SHORT DISTANCE AUTOMATICALLY. IF THE OPERATOR PRESSES THE DOWN BUTTON, THE MACHINE WILL MOVE UP INSTEAD OF DOWN DURING A COLLISION ALERT.THE COLLISION DETECTION AUTOMATICALLY MOVES THE TABLE UP OR DOWN IN A PRE-DETERMINED DISTANCE THAT IS INHERENT TO THE DESIGN. DURING A COLLISION, THE CONTROLS THEN HAVE THE OPPOSITE OPERATION OF THE ARROW INDICATORS DURING A COLLISION ALERT. THIS CAN TRAP AND INJURE PT'S LEGS UNDER THE TABLE, CONFUSE STAFF IN AN EMERGENCY AND CAUSE MORE HARM THAN NECESSARY.THE GE FIELD SERVICE REPORT READS: 'INVESTIGATE COLLISION SENSOR OPERATION. MODEL REDACTED, SN REDACTED. DETERMINED THAT IF A COLLISION IS SENSED IN THE UP DIRECTION, THE C-ARM IS DRIVEN DOWN A PREDETERMINED AMOUNT. THE PATIENT WAS LEANING ON DETECTOR WITH UPPER BODY WHILE SEATED ON CHAIR WITH LEGS UNDER DETECTOR. WHEN COLLISION WAS SENSED IN THE UP DIRECTION, SYSTEM DROVE C-ARM DOWN A PRESET AMOUNT CRUSHING PATIENT'S KNEES. WAS NOT ABLE TO DRIVE C-ARM UP AS MAXIMUM HEIGHT WAS ALREADY ACHIEVED.' | 0
1728357 | A PHARMACY TECH WAS SHOWING PHARMACY TECH STUDENTS FOR OUR NURSING COLLEGE WHAT A PYXIS MACHINE LOOKED LIKE AND HOW TO TURN IT ON. AS SOON AS THE POWER BUTTON ON THE UPS WAS ACTIVATED, THE MACHINE POPPED AND SPARKED WITH A FLAME ON THE SIDE (OF THE UPS). THIS WAS A NON-LIVE PYXIS MACHINE THAT WAS BEING PREPPED FOR AN UPCOMING GO-LIVE. THIS WAS A BRAND NEW MACHINE. NO ONE WAS INJURED. | 1
1639023 | ACCORDING TO CUSTOMER, NO ERRONEOUS RESULTS WERE REPORTED. OCD FILED ENGINEER REPLACED THE PROBE AND PERFORMED ADJUSTMENTS.(B) (4) >< THE CUSTOMER REPORTED FLUID LEAKING FROM THE WASH STATION AND THE BOTTOM OF THE UNIT DURING TESTING. NO PERSONNEL HAD BEEN EXPOSED TO BIOHAZARDOUS FLUIDS, AND NO ERRONEOUS RESULTS HAD BEEN REPORTED. ALL TESTING HAD BEEN ABORTED. | 0
1329085 | THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. (B)(4). >< THE CUSTOMER REPORTED THAT SINUS IMAGES ARE FLIPPED. | 1
1373984 | RESULT: NEEDLE BIOPSY GUIDE. >< DURING NEEDLE LOCAL PROCEDURES USING BIOPSY NEEDLE, IMAGES MAY BECOME 'BLACK' (VOID OF IMAGE DATA) WHEN THE NEEDLE IS PLACED IN THE BREAST AND X-RAY BEAM IS DIRECTED STRAIGHT THROUGH THE HUB OF THE NEEDLE. SIMILAR INCIDENTS INVOLVING A HOMER MAMMALOK GOLD BIOPSY NEEDLE HAVE BEEN REPORTED ON NOVEMBER 7, 2008 UNDER MDR #2240869-2008-0025 AND ON NOVEMBER 11, 2008 UNDER MDR# 2240869-2008-0026. (B)(4). | 0
1872629 | THE STAFF REPORTED THAT THE OBTURATOR BROKE WHEN PLACED IN THE TROCAR ON TWO DIFFERENT PATIENTS, WITH TWO DIFFERENT SURGEONS, IN TWO DIFFERENT ROOMS FOR TWO DIFFERENT SURGERIES ON THE SAME DAY. IT BROKE NEAR THE SAME SPOT IN BOTH CASES. THE STAFF WERE ABLE TO RETRIEVE THE ITEM. THEY REPORTED THIS AS A RECOGNIZED DEFECT FOR THIS ITEM. THEY DID NOT FEEL THAT THE BREAK WAS RELATED TO SURGICAL TECHNIQUE. SIMILAR BREAKAGE HAS OCCURRED IN 5 OTHER OBTURATORS.======================HEALTH PROFESSIONAL'S IMPRESSION======================THERE APPEARS TO BE A COMMON AREA OF STRESS AND/OR WEAKNESS IN THE OBTURATOR. STAFF FEELS THAT THE PLASTIC IS NOT HOLDING UP TO REPEATED STEAM STERILIZATIONS. REPROCESSING HAS BEEN COMPLETED ACCORDING TO MANUFACTURER'S RECOMMENDATIONS. WE ARE PLANNING A SIMULATION TO SEE IF WE CAN RE-PRODUCE THE DEFECT TO EVALUATE PRODUCT AND TECHNIQUE. ======================MANUFACTURER RESPONSE PER SITE REPORTER======================THE REP THOUGHT IT MIGHT BE RELATED TO THE WAY WE PROCESS THE EQUIPMENT. HE WAS NOT ONSITE, NOR HAD HE SEEN THE ACTUAL DEFECT. MANUFACTURER IS COMING TO OUR FACILITY TO EVALUATE BREAKAGE.======================HEALTH PROFESSIONAL'S IMPRESSION======================THERE APPEARS TO BE A COMMON AREA OF STRESS AND/OR WEAKNESS IN THE OBTURATOR. WE ARE PLANNING A SIMULATION TO SEE IF WE CAN RE-PRODUCE THE DEFECT TO EVALUATE PRODUCT AND TECHNIQUE.======================MANUFACTURER RESPONSE PER SITE REPORTER======================THE REP THOUGHT IT MIGHT BE RELATED TO THE WAY WE PROCESS THE EQUIPMENT. HE WAS NOT ONSITE, NOR HAD HE SEEN THE ACTUAL DEFECT. | 0
1059859 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED FOR THE FALSE NEGATIVE RESULTS REPORTED BY THE CUSTOMER. THE LOG AND IMAGES REVIEWED BY OCD SECOND LEVEL SUPPORT REVEALED THAT THE ANALYZER CORRECTLY INTERPRETED THE PT'S RESULTS AS NEGATIVE. THE CUSTOMER REPORTED THAT THEY HAD PERFORMED ADD'L TESTING ON THE SAMPLE, AND CONCLUDED THAT THE SAMPLE HAS VARIABILITY. THE ISSUE WAS SAMPLE RELATED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4) >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER RESULTED A PREVIOUSLY KNOWN SAMPLE CONTAINING ANTI-K AS NEGATIVE. THE CUSTOMER REVIEWED THE PROCESSED GEL CARD AND VISUALLY CONFIRMED THE REACTION IN THE MICROTUBE WAS NEGATIVE. THE SAMPLE REACTED POSITIVE (1+) IN MANUAL GEL TEST USING THE SAME LOT OF REAGENTS. NO ERRONEOUS TEST RESULTS WERE REPORTED. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 1
1895777 | THE CUSTOMER REPORTED THAT THE PROVUE DILUTION CUP STARTED OVERFLOWING. NO ERRONEOUS RESULTS WERE REPORTED. >< AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND REPLACED THE WASTE BOTTLE FITTINGS, WASH STATION AND VALVES. THE FE PRIMED THE SYSTEM SEVERAL TIMES. THE FE CHECKED THE FUNCTIONAL OPERATION OF THE PROVUE. THE CUSTOMER RAN AND ACCEPTED QC. SERVICE DID RESOLVE PROBLEM. INSTRUMENT OPERATING AS EXPECTED. (B)(4). | 0
1588709 | THE CUSTOMER REPORTED THAT THE PROVUE PROBE IS DRIPPING. NO ERROR MESSAGES WERE POSTED. THE CUSTOMER INDICATED THAT THE PROVUE WAS TAKEN OUT OF USE. NO ERRONEOUS RESULTS WERE POSTED. >< AN OCD FIELD ENGINEER (FE) ARRIVED ON SITE AND OBSERVED THE BENT PROBE. THE FE REPLACED THE PROBE AND PERFORMED ALL REQUIRED ADJUSTMENTS. REPLACEMENT OF THE PROBE AND ADJUSTMENTS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION.(B)(4) | 0
1883775 | AN OCD FIELD ENGINEER ARRIVED AT THE CUSTOMER SITE AND CLEANED THE READER CAMERA. THE FE ALSO ADJUSTED THE READER CAMERA AND CREATED A NEW REFERENCE IMAGE. THE CUSTOMER RAN QC AND ACCEPTED THE RESULTS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4) >< THE CUSTOMER REPORTED THAT THE PROVUE INTERPRETED A GEL CARD AS NEGATIVE THAT WAS VISUALLY OBSERVED AS A WEAK POSITIVE. NO ERRONEOUS RESULTS WERE REPORTED. | 0
1345126 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER DRIPPED FLUID, RESULTING IN REAGENT CONTAMINATION. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS, WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< AN (B) (4) FIELD ENGINEER VISITED THE CUSTOMER SITE AND FOUND THE PROBE WAS BENT. THE FE REPLACED THE PROBE AND PERFORMED ADJUSTMENTS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. QC PASSED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). | 0
1060811 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THAT THE PROBE WAS BENT. REPLACEMENT OF THE PROBE HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). >< THE CUSTOMER REPORTED THAT THE PROBE WAS BENT AND DRIPPING ON THE ORTHO PROVUE ANALYZER. THE CUSTOMER ALSO OBSERVED FLUID ON TOP OF THE GEL CARD FOIL. THE CUSTOMER ABORTED TESTING AND DISCONTINUED THE USE OF THE INSTRUMENT. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/REAGENT, CARRY OVER AND/ OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1060820 | THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER WAS BENT AND LEAKING FLUID. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE/ REAGENT, CARRY OVER AND/OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND DETERMINED THAT THE PROBE WAS BENT. THE FE REPLACED THE PROBE AND PERFORMED THE APPROPRIATE ADJUSTMENTS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). | 0
1277335 | CUSTOMER REPORTED A PROBE DRIP. NO ERROR CODES WERE POSTED. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< AN OCD FIELD ENGINEER ARRIVED ON SITE AND REMOVED THE PROBE IN ORDER TO REMOVE A CLOT. THE VACUUM AND PRESSURE WERE CHECKED AND FOUND TO BE WITHIN SPECIFICATIONS. THE SYSTEM WAS PRIMED THREE TIMES AND NO LEAKING WAS OBSERVED. THE INSTRUMENT WAS RETURNED TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THE REPAIR. (B) (4). | 0
1353529 | ON (B) (6) 2009, A PRE-TRANSFUSION WORK-UP WAS PERFORMED ON A PATIENT SAMPLE WITH A HISTORY OF ANTI-E. ANTIBODY SCREEN WAS POSITIVE; PANEL TESTING PERFORMED IN A MANUAL GEL USING ORTHO REAGENTS CONFIRMED THE PRESENCE OF ANTI-E. THE CUSTOMER PERFORMED CROSSMATCH/COMPATIBILITY TESTING ON THE ORTHO PROVUE ANALYZER USING THE SAMPLE AND 2 DONOR UNITS OBTAINED COMPATIBLE/NEGATIVE RESULTS. THE PATIENT WAS TRANSFUSED WITH THE 2 UNITS. IT WAS LATER DETERMINED THAT THE LAB HAD FAILED TO PERFORM ANTIGEN TYPINGS ON THE UNITS CROSSMATCHED PRIOR TO TRANSFUSION. TESTING INDICATED THAT BOTH UNITS WERE E+. REPEAT TESTING PERFORMED ON THE PROVUE AND MANUEL GE. PROVUE YIELDED COMPATIBLE RESULTS. MANUAL GEN YIELDED 2+ INCOMPATIBLE RESULTS WITH BOTH UNITS. THE CUSTOMER INDICATED THAT THE PATIENT IS ASYMPTOMATIC; HAS BEEN DISCHARGED FROM HOSPITAL AND IS DOING WELL. OCD MEDICAL AFFAIRS HAS DETERMINED THAT A SERIOUS INJURY HAS NOT OCCURRED. >< AN OCD FIELD ENGINEER VISITED THE SITE AND PERFORMED REQUIRED MAINTENANCE. A REVIEW OF THE LOG FILES AND TIF IMAGES REVEALED THE CAMERA BRIGHTNESS WAS SET TO 130, WHICH WAS OUT OF SPECIFICATION. THE FE ADJUSTED THE CAMERA BRIGHTNESS TO 113, PERFORMED SYSTEM OPTIMIZATION AND 1 YEAR PREVENTATIVE MAINTENANCE TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. ON (B) (6) 2009, OCD SECOND LEVEL SUPPORT REVIEWED THE WADIANA LOGS AND INDICATED THAT THE ANALYZER DID NOT POST ANY ERROR MESSAGES DURING TESTING. THE TEMPERATURE WAS 36.9. ANALYSIS OF THE TIF IMAGES COULD NOT BE REVIEWED, SINCE THEY WERE NOT CAPTURED AT THE INITIAL TIME OF THE INCIDENT. BITMAP IMAGES WERE RECEIVED FOR THE BATCHES IN QUESTION, BUT WERE INCONCLUSIVE. DISCUSSED SAMPLE PREPARATION REQUIREMENTS PER ORTHO PROVUE USER'S GUIDE. THE CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). | 1
1659586 | AN OCD FIELD ENGINEER (FE) VISITED THE CUSTOMER SITE AND OBSERVED THE READER CAMERA COVER WAS BENT IN THE DOWN DIRECTION. THE FE STRAIGHTENED THE READER CAMERA COVER AND PERFORMED ALL READER CAMERA ADJUSTMENTS. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B) (4) >< THE CUSTOMER REPORTED THE ORTHO PROVUE ANALYZER MISREAD SAMPLE REACTIONS THAT CONTAINED MIXED FIELD REACTIVITY. THE ORTHO PROVUE DID NOT RESULT THE MIXED FIELD REACTIVITY AS DP (DOUBLE POPULATION) AS EXPECTED. NO ERRONEOUS RESULTS WERE REPORTED. A MISREAD TEST RESULT MAY LEAD TO ERRONEOUS RESULTS BEING REPORTED. | 1
1849976 | WE HAVE FOUND THAT THE PRODUCT FOR PHARMACIST CALLED SIEMENS PHARMACY AND MEDICATION ADMINISTRATION CHECK ALLOWED TWO FORMS OF POTENT STEROIDS TO BE DELIVERED FOR SIMULTANEOUS ADMINISTRATION BY INTRAVENOUS AND ORAL ROUTES, TWO FORMS OF IPRATROPIUM, AND FAILED TO SUFFICIENTLY WARN OF AN ALLERGY TO ANOTHER THERAPY - IT WAS ADMINISTERED - DESPITE THE ACCURATE LISTING IN THE PATIENT'S EMR. THE STEROID THERAPY WORSENED HYPERGLYCEMIA. | 1
1733204 | OUR REP REPORTS TO US THAT DURING AN ARTHROSCOPIC KNEE REPAIR, THE INFLOW WHEEL OF THE FMS PUMP KEPT SPINNING, CONTINUOUSLY ADDING FLUID INTO THE KNEE CAVITY. AT SOME POINT, THE SURGEON/STAFF DISCERNED EXCESSIVE JOINT SWELLING AND THE DEVICE WAS IMMEDIATELY POWERED OFF AND A REPLACEMENT SAME DEVICE WAS EMPLOYED TO COMPLETE THE REPAIR SUCCESSFULLY. THE PATIENT WAS KEPT FOR OBSERVATION FOR AN UNDISCLOSED AMOUNT OF TIME. IT IS REPORTED THAT THE SWELLING ABATED WITHIN A COUPLE OF HOURS INTO THE OBSERVATION. THE PATIENT WAS LATER RELEASED AND IS DOING FINE, NO PT CONSEQUENCES. >< THE DEVICE WAS RECEIVED AND WAS THOROUGHLY EVALUATED VISUALLY, ELECTRICALLY AND FUNCTIONALLY. VISUALLY, IT IS NOTED THAT THE COMPLAINT DEVICE ARRIVED WITH SOME MINOR DAMAGE: REPLACED THE DAMAGED TOP COVER, LEFT SAFETY COVER HINGES AND CHAMBER HOLDER. THE DEVICE WAS THEN TESTED FOR ELECTRICAL AND FUNCTIONAL CONFORMANCE; THE PUMP PASSED ALL DIAGNOSTIC AND FUNCTIONAL TESTS WELL WITHIN ITS DESIGNED AND MANUFACTURED PARAMETERS (THE UNIT WAS RUN CONTINUOUSLY FOR 2 DAYS). WE COULD FIND NO FAULT OR ANOMALIES WITH THE DEVICE. WE CANNOT DISCERN A ROOT CAUSE FOR THE REPORTED ISSUE. AT THIS POINT IN TIME, NO CORRECTIVE OR FURTHER ACTION IS WARRANTED. HOWEVER, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
1685882 | THE SALES REP HAS REPORTED THAT THERE WAS A PRESSURE ISSUE WITH THIS DEVICE. THE SURGEON DOES NOT BELIEVE THAT THE READINGS ON THE FMS PUMP DISPLAY WERE ACCURATE. THE PUMP WAS SET TO A CERTAIN PRESSURE, AND DISPLAYED THAT SETTING, HOWEVER, TOO WHATEVER DEGREE, THE SURGEON REPORTS THAT THE PT'S SHOULDER BECAME MORE SWOLLEN THAT HE WAS COMFORTABLE WITH. THEY STOPPED USE OF THE PUMP AND EMPLOYED ANOTHER FMS PUMP TO COMPLETE THE SURGERY WITHOUT FURTHER ISSUE OR HARM TO THE PT. OUTSIDE OF THE PUMP SWAP, NO OTHER INTERVENTION WAS TAKEN. >< ALTHOUGH THEY ARE REPORTING NO PT CONSEQUENCES AND NO EXTRAORDINARY MEASURES OF INTERVENTION WERE NEEDED FOR REMEDY, SOMETHING HAPPENED TO MAKE THE SURGEON UNCOMFORTABLE. BECAUSE OF THE EXPRESSION OF UNCOMFORTABILITY, A REPORT IS BEING FILED TO DOCUMENT THE EVENT. MITEK IS AT THIS POINT IN TIME IN THE INFO GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT. | 0
1910611 | RN SENT PHILIPS TELEMETRY PACK TO BIOMED BECAUSE IT PERSISTED IN DISPLAYING 'LOW BATTERY.' BIOMED CHANGED AA BATTERIES AND GOT THE SAME ISSUE OF 'LOW BATTERY' WITHIN 1-2 MINUTES. TECH OPENED BATTERY CASE AND ON REMOVING ONE OF THE TWO BATTERIES FOUND IT TO BE EXTREMELY HOT TO TOUCH. ON FURTHER INVESTIGATION IT WAS FOUND THAT THE NEGATIVE POLE SPRING (THE BATTERY TOWARDS THE OUT SIDE OF THE CASE) HAD SLIPPED OVER THE TOP OF RETAINING CLIP AND COME INTO CONTACT WITH WIRE FROM THE BATTERY CLOSER TO THE TOUCH PAD. THE LOAD ON OUTSIDE BATTERY WAS GREATLY INCREASED AND OVERHEATED. THE BATTERY INDICATOR WAS DISPLAYING 'LOW BATTERY' ON THE UNIT'S SCREEN. IN A SIMULATED SITUATION IT WAS DEMONSTRATED THAT THE OUTSIDE BATTERY WAS HEATING UP TO ~ 196.5 F DEGREES. THE PLASTIC RETAINING CLIP 'GRINDS DOWN' OVER TIME. THERE IS NO APPARENT WAY TO REPLACE THE CLIP. PATIENTS HAVE COMPLAINED THE PACK GETS HOT BUT WE HAD NEVER DISCOVERED THE CAUSE UNTIL NOW.HEALTH PROFESSIONAL'S IMPRESSION: WEAR AND TEAR ON A PLASTIC RETAINING CLIP THAT ALLOWS A NEGATIVE POLE BATTERY SPRING TO SLIP UP OVER TOP OF CLIP AND COME INTO CONTACT WITH SECOND AA BATTERY. THIS IS A SUBTLE PROBLEM THAT NURSING STAFF CANNOT EASILY VISUALIZE OR FIX. | 0
1156183 | REPORTER INDICATED THAT A (B) (4) COMPUTER WAS FREEZING DURING INTERROGATION. RESEATING THE FLASHCARD RESOLVED THE SCREEN FREEZING. IT WAS ALSO REPORTED THAT THE (B) (4) COMPUTER DID NOT STAY CHARGED, BUT THE REPORTER STATED IT MAY NOT HAVE BEEN KEPT PLUGGED IN AND CHARGING BETWEEN USES. ATTEMPTS TO THE REPORTER FOR ADDITIONAL INFORMATION AND IF THE (B) (4) COMPUTER WILL BE RETURNED HAVE BEEN UNSUCCESSFUL TO DATE. | 0
1156184 | REPORTER INDICATED A VNS (B) (4) COMPUTER WAS NOT PERFORMING AS INTENDED. TROUBLESHOOTING WITH ANOTHER VNS COMPUTER AND WAND REVEALED THE (B) (4) WAS NOT RECEIVING DATA. NO ERROR MESSAGES WERE NOTED OTHER THAN 'RETRY'. RESEATING THE FLASHCARD DID NOT RESOLVE THE ISSUE. THE (B) (4) WAS ALSO FULLY CHARGED. THE (B) (4) COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. >< DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1701623 | REPORTER INDICATED A VNS (B)(4) COMPUTER WITH VERSION 7.1 SOFTWARE WAS FREEZING DURING INTERROGATION. PERFORMING A SOFT RESET RESOLVED THE ISSUE, BUT THE REPORTER REQUESTED A REPLACEMENT. THE COMPUTER AND FLASHCARD WERE RETURNED FOR ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED COMPUTER AND THE ALLEGED SCREEN FREEZING IS A KNOWN EVENT THAT CAN OCCUR WITH THE COMBINATION OF THE (B)(4) COMPUTER AND 7.1 SOFTWARE. OTHER THAN THE SCREEN FREEZING EVENT, THE COMPUTER AND SOFTWARE PERFORMED SPECIFICATIONS. | 0
1509901 | REPORTER INDICATED A VNS (B) (4) COMPUTER WOULD NOT HOLD A CHARGE. THE COMPUTER AND FLASHCARD HAVE BEEN REQUESTED FOR RETURN BUT HAVE NOT BEEN RECEIVED TO DATE. >< DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1760897 | DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY. >< IT WAS REPORTED THAT THE SITE WAS RECEIVING FAILURE TO ESTABLISH COMMUNICATION ERROR MESSAGES USING THEIR PROGRAMMING SYSTEM. TROUBLESHOOTING WAS DONE OVER THE PHONE, AND IDENTIFIED THAT THE SERIAL ADAPTER CORD CONNECTION INTO THE HAND HELD WAS LOOSE AND WAS NOT ABLE TO BE TIGHTENED. THE NURSE MOVED THE CORD, AND WAS ABLE TO ESTABLISH COMMUNICATION WITH THE PT'S GENERATOR EVENTUALLY. THE SITE WAS SENT A NEW PROGRAMMING SYSTEM TO REPLACE THE NON-WORKING DEVICE. ATTEMPTS TO OBTAIN THE NON-WORKING HAND HELD AND SERIAL ADAPTER CABLE ARE UNDERWAY, HOWEVER, THE DEVICE HAS NOT BEEN RETURNED TO DATE. | 0
1638699 | IT WAS REPORTED THAT THE SITE WAS EXPERIENCING PROBLEMS WITH THEIR (B) (6) HAND HELD AND THAT IT WOULD NOT POWER ON. THE HAND HELD WAS PLUGGED INTO THE WALL OUTLET, AND THE CHARGING LIGHT WAS NOT ILLUMINATED, WHICH WOULD NORMALLY INDICATE THE DEVICE WAS CHARGING. A HARD RESET WAS PERFORMED, AND THE HAND HELD POWERED ON. THE BATTERY GAUGE WAS CHECKED AT THAT POINT, WHICH REVEALED THE HAND HELD HAD A FULLY CHARGED BATTERY. THE ADAPTER CABLES WERE INTERCHANGED WITH KNOWN WORKING CABLES, AND THE CHARGING LIGHT ILLUMINATED. TROUBLESHOOTING WAS NOT ABLE TO CONCLUSIVELY DETERMINE THE CAUSE OF THE CHARGING LIGHT NOT ILLUMINATING, AND BECAUSE THE HAND HELD WAS FULLY CHARGED, THE POWER CABLE AND/OR THE SERIAL CABLE ARE BELIEVED TO BE THE CAUSE OF THE LIGHT NOT ILLUMINATING. THE CABLES HAVE BEEN RETURNED TO MANUFACTURER, WHERE ANALYSIS IS UNDERWAY. GOOD FAITH ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE SITE REGARDING THE CURRENT FUNCTIONALITY OF THE HAND HELD HAVE BEEN MADE, BUT NO RESPONSE HAS BEEN RECEIVED TO DATE. | 0
1638698 | REPORTER INDICATED A VNS PATIENT WAS INTERROGATED ON (B) (6)2010 AND THE VNS WAS FOUND TO BE SET TO 0MA. THE PATIENT'S VNS SETTINGS WERE CHANGED BACK TO THE INTENDED SETTINGS. VNS PROGRAMMING HISTORY FOR THE PATIENT WAS OBTAINED AND REVIEWED, NOTING THE FOLLOWING EVENTS: ON (B) (6)2009-2:49:46 PM: INTERROGATION - 2.75 MA/15HZ/250PW/60 SEC ON/1.1 MIN OFF. ON (B) (6)2009-2:51:42 PM: NORMAL MODE TEST - OUTPUT STATUS = OK, LEAD IMPEDANCE = OK, DCDC CODE = 5, EOS = NO. ON (B) (6)2009-2:52:05 PM: FAULTED SYSTEMS TEST. ON (B) (6)2009-2:52:57 PM: SYSTEMS TEST - OUTPUT STATUS OK, LEAD IMPEDANCE = OK, DCDC CODE = 2, EOS = NO. NO FINAL INTERROGATION WAS PERFORMED ON (B) (6)2009 TO VERIFY INTENDED SETTINGS. WHEN THE PATIENT WAS SEEN AGAIN ON (B) (6)2010, THE SETTINGS WERE 0MA/20HZ/500PW/30 SEC ON/60 MIN OFF. THE FAULTED SYSTEMS DIAGNOSTICS TEST CAUSED THE SETTINGS TO CHANGE, BUT THIS WAS NOT NOTICED UNTIL THE NEXT OFFICE VISIT VNS INTERROGATION. THE REPORTER WAS APPRISED OF THE IMPORTANCE OF REINTERROGATING THE VNS TO ENSURE SETTINGS ARE AS INTENDED BEFORE THE PATIENT LEAVES THE OFFICE. >< ANALYSIS OF PROGRAMMING HISTORY. | 0
1371443 | REPORTER INDICATED THAT THERE WAS AN ISSUE WITH THE HANDHELD COMPUTER THAT WAS RESULTING IN THE MAGNET ACTIVATIONS SCREEN SHOWING MULTIPLE RESULTS PER LINE, WHEN ONLY 1 RESULT PER LINE SHOULD BE PRESENT. THE PRODUCT HAS BEEN RETURNED TO THE MFR, BUT ANALYSIS OF THE DEVICE IS NOT YET COMPLETED. | 0
1371441 | REPORTER INDICATED THAT THE HANDHELD COMPUTER WOULD NOT HOLD A CHARGE, EVEN WHEN PLUGGED IN, AND WOULD FREEZE DURING INTERROGATION. ATTEMPTS TO HAVE THE PRODUCT RETURNED FOR ANALYSIS HAVE BEEN UNSUCCESSFUL TO DATE. | 0
1577165 | IT WAS REPORTED THAT THE SITE'S HANDHELD WOULD CONTINUALLY FREEZE FOLLOWING SUCCESSFUL INTERROGATIONS AND OTHER SCREENS. IT IS UNK IF TROUBLESHOOTING WAS PERFORMED TO RESOLVE THE EVENT BUT THE SITE REQUESTED THAT A NEW DEVICE BE SENT TO REPLACE THE NON-WORKING DEVICE. THE NON-WORKING DEVICE HAS BEEN RETURNED TO THE MFR AND ANALYSIS IS UNDERWAY. HOWEVER, IT IS UNK THAT UNDER CERTAIN CONDITIONS THIS TYPE OF SCREEN FREEZE EVENT CAN OCCUR WITH THE (B)(4) HANDHELD COMPUTER AND CYBERONICS (B)(4) SOFTWARE. | 0
1418570 | DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. >< REPORTER INDICATED A VNS (B) (4) HANDHELD COMPUTER WAS FREEZING DURING INTERROGATION. THE COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. | 0
1460092 | REPORTER INDICATED A PATIENT'S VNS WAS FOUND TO BE SET TO 1MA OUTPUT CURRENT UPON INTERROGATION. THE PATIENT'S NORMAL SETTING IS 0.25MA OUTPUT CURRENT. THE PATIENT'S VNS WAS REPROGRAMMED TO THE INTENDED SETTINGS. THE REPORTER STATED THE PATIENT WAS SEEN BY ANOTHER PHYSICIAN JUST PRIOR TO THE REPORTER AND THAT THE OTHER PHYSICIAN HAD DISABLED THE VNS. IT IS BELIEVED THERE WAS AN INTERRUPTED DIAGNOSTICS TEST AT THE OTHER PHYSICIAN'S VISIT THAT CAUSED THE OUTPUT CURRENT TO CHANGE. ATTEMPTS FOR PROGRAMMING HISTORY FROM THE OTHER PHYSICIAN HAVE BEEN UNSUCCESSFUL TO DATE. | 0
1460094 | REPORTER INDICATED A VNS PATIENT WAS HAVING PAINFUL STIMULATION AT THE VNS ELECTRODE SITE ONCE EVERY HOUR. VNS DIAGNOSTICS WERE PERFORMED THE DAY BEFORE THE PATIENT BEGAN EXPERIENCING THE PAINFUL STIMULATION EVERY HOUR. IT IS BELIEVED THE VNS DIAGNOSTICS TEST WAS INTERRUPTED, CAUSING THE VNS SETTINGS TO CHANGE. ATTEMPTS FOR VNS PROGRAMMING HISTORY FROM THE REPORTER HAVE BEEN UNSUCCESSFUL TO DATE. | 0
1067463 | REPORTER INDICATED A VNS (B) (4) HANDHELD COMPUTER WAS FREEZING DURING INTERROGATION. THE HANDHELD COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. | 0
1674324 | IT WAS REPORTED TO MANUFACTURER THAT THE USER'S HANDHELD SCREEN WAS FREEZING ON THE SUCCESSFUL INTERROGATION SCREEN AND WOULD NOT ALLOW HIM TO PROCEED. IT IS UNKNOWN IF TROUBLESHOOTING WAS PERFORMED TO RESOLVE THE EVENT BUT THE SITE REQUESTED THAT A NEW DEVICE BE SENT TO REPLACE THE NON-WORKING DEVICE. THE NON-WORKING DEVICE HAS BEEN RETURNED TO THE MANUFACTURER AND ANALYSIS IS UNDERWAY. HOWEVER, IT IS KNOWN THAT UNDER CERTAIN CONDITIONS THIS TYPE OF SCREEN FREEZE EVENT CAN OCCUR WITH THE (B)(4) HANDHELD COMPUTER AND CYBERONICS 7.1 VERSION SOFTWARE. | 0
1336234 | REPORTER INDICATED A VNS (B) (4) HANDHELD COMPUTER WOULD NOT HOLD A CHARGE DESPITE BEING PROPERLY CHARGED. THE COMPUTER, FLASHCARD, AND PROGRAMMING WAND WERE RETURNED FOR ANALYSIS. ANALYSIS OF THE WAND WAS COMPLETED AND NOTED THE WAND BATTERY WAS DEPLETED. AFTER THE BATTERY WAS REPLACED, THE WAND FUNCTIONED NORMALLY. ANALYSIS OF THE COMPUTER AND FLASHCARD ARE PENDING. >< DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1336236 | REPORTER INDICATED A VNS (B) (4) COMPUTER SCREEN WAS FREEZING. THE HANDHELD COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. | 0
1693427 | IT WAS REPORTED THAT THE SITE WAS EXPERIENCING PROBLEMS WITH THEIR HP HAND HELD AND THAT THE SCREEN TOOK SEVERAL TAPS TO RESPOND TO STYLUS TAPS. THE MANUFACTURER REPRESENTATIVE WENT TO THE SITE TO TROUBLESHOOT WITH THE HAND HELD AND THE EVENT WAS DUPLICATED, AND IT TOOK 7 OR 8 TAPS FOR THE HAND HELD TO RESPOND. A HARD RESET WAS PERFORMED AND THE SCREEN WAS RE-ALIGNED IN THE PROCESS, HOWEVER, THE EVENT CONTINUED TO OCCUR. THE SITE WAS SENT A NEW HAND HELD IN REPLACEMENT AND THE NON-WORKING DEVICE HAS BEEN SENT BACK TO MANUFACTURER WHERE ANALYSIS IS UNDERWAY. | 0
1693426 | IT WAS REPORTED THAT THE PHYSICIAN'S (B)(6) HAND HELD SCREEN WAS FREEZING AT THE INTERROGATION SUCCESSFUL SCREEN WHILE USING THE VERSION 7.1 VNS SOFTWARE. THE PHYSICIAN IS ABLE TO RESOLVE THE PROBLEM WITH TROUBLESHOOTING, HOWEVER, THE EVENT CONTINUES TO OCCUR AND THE PHYSICIAN REQUESTED A NEW HAND HELD TO REPLACE THE (B)(6). THE HAND HELD AND SOFTWARE FLASHCARD HAVE BEEN RETURNED TO MANUFACTURER WHERE ANALYSIS IS UNDERWAY. | 0
1325543 | REPORTER INDICATED A HANDHELD VNS (B) (4) COMPUTER WAS NOT ABLE TO INTERROGATE PATIENTS. THE PROGRAMMING WAND WAS CONFIRMED TO BE WORKING PROPERLY. THE HANDHELD COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE CURRENTLY IN PRODUCT ANALYSIS. >< DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1426116 | REPORTER INDICATED THAT THE SITE'S HANDHELD COMPUTER WOULD NO LONGER HOLD A CHARGE. ALL ATTEMPTS FOR FURTHER INFO, AND TO HAVE THE HANDHELD COMPUTER RETURNED FOR ANALYSIS, HAVE BEEN UNSUCCESSFUL TO DATE. | 0
1507307 | MANUFACTURING COUNTRY REPRESENTATIVE REPORTED A SITE HAD A HANDHELD THAT WOULD NOT HOLD A CHARGE. GOOD FAITH ATTEMPTS HAVE BEEN MADE FOR THE PRODUCT TO BE RETURNED FOR PRODUCT ANALYSIS. THUS FAR THE PRODUCT HAS NOT BEEN RETURNED. | 0
1810781 | REPORTER INDICATED THAT A (B)(6) HANDHELD COMPUTER'S SCREEN FROZE DURING INTERROGATION. RESEATING THE FLASHCARD AND RESETTING THE DEVICE DID NOT RESOLVE THE ISSUE. PRODUCT IS EXPECTED TO BE RETURNED TO THE MFR. | 0
1731509 | IT WAS REPORTED THAT THE PHYSICIAN'S (B) (6) HAND HELD WAS FREEZING AT THE INTERROGATION SCREEN. THE PHYSICIAN WOULD PERFORM A SOFT RESET WHICH WOULD TEMPORARILY RESOLVE THE ISSUE, BUT THE PHYSICIAN IS REQUESTING A NEW HAND HELD AS THE EVENT CONTINUES TO OCCUR. IN ADDITION, THE PHYSICIAN REPORTED THAT THE HAND HELD IS NOT ALWAYS RESPONDING TO TAPS ON THE SCREEN USING THE STYLUS. THE MANUFACTURER REPRESENTATIVE ATTEMPTED TO RE-ALIGN THE SCREEN, BUT THE PROBLEM DID NOT RESOLVE. THE NON-WORKING HAND HELD AND SOFTWARE FLASHCARD HAVE BEEN RETURNED TO MANUFACTURER WHERE ANALYSIS IS UNDERWAY. | 0
1134784 | IT WAS REPORTED THAT A (B) (4) HANDHELD DEVICE FROZE AT THE 'INTERROGATION SUCCESSFUL' SCREEN. THE EVENT WAS RESOLVED BY REMOVING AND REINSERTING THE DEVICE'S FLASHCARD. NO ADDITIONAL EVENTS WERE REPORTED. | 0
1203129 | CONCLUSIONS: DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. >< RPTR INDICATED SQL ERRORS WERE NOTED WHILE ATTEMPTING TO UPLOAD FLASHCARD INFORMATION. AN ANALYSIS WAS PERFORMED ON THE FLASHCARD AND THE CAUSE FOR THE SQL ERROR IS ASSOCIATED WITH A CORRUPT DATABASE. AFTER THE ANALYSIS WAS COMPLETED, THE FLASHCARD WAS NOT UPDATED SO THE LAST KNOWN GOOD DATABASE COULD BE UPLOADED. AS A RESULT, THE VNS SOFTWARE ATTEMPTED TO READ THE CORRUPT CYBERONICS BU/CDB DATABASE. | 0
1514334 | THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE WAS UNABLE TO PACE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1901641 | IT HAS BEEN COMMUNICATED THAT THE DEVICE AND/OR LOT INFO IS NOT AVAILABLE FOR EVAL. WITHOUT THE DEVICE AND/OR LOT INFO IT IS NOT POSSIBLE FOR CODMAN TO CONDUCT A PROPER INVESTIGATION. IF AT SOME POINT THE DEVICE AND/OR LOT INFO DOES BECOME AVAILABLE, THIS COMPLAINT WILL BE RE-OPENED, EVALUATED AND A F/U REPORT WILL BE FILED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR COMPLAINTS. AT THE PRESENT TIME THIS COMPLAINT IS CONSIDERED CLOSED. >< AFFILIATE REPORTED THAT THE TIP OF THE SENSOR BROKE WHEN THE SENSOR WAS REMOVED. THE BROKEN TIP REMAINS IN THE PATIENT'S BODY. THE SURGEON COMMENTED THAT FRICTION WAS FELT WHEN THE SENSOR WAS REMOVED. THE SENSOR TIP GOT STUCK WITH THE TITANIUM PLATE WHICH WAS IMPLANTED NEAR THE SENSOR DURING REMOVAL. | 0
1709197 | EQUIPMENT INSTALLATION (B)(6) 2009 WAS PHASED TO 'GO LIVE' DURING (B)(6). (B)(6) THE HOSPITAL EXPERIENCED A NUMBER OF ISSUES. MULTIPLE INSTANCES OF DATABASE SERVER (DBS) REBOOTING CAUSING THE PATIENT INFORMATION CENTER (PICS) TO GO INTO LOCAL MODE AND FOR THE CLIENTS TO GO OFF LINE. MOUSE FAILURES WITH THE REMOTE CPU SOLUTION AND KEY BOARD, MONITOR, VIDEO (KVM) SWITCHES. THE PIC REBOOTS OR GOES BLANK RANDOMLY; PIC RECOVERS, HOWEVER DATA IS LOST DURING THE TIME OF THE 'BLANKING' OR REBOOTING. PRINTERS ARE GOING OFF LINE AND UNABLE TO PRINT. SYSTEM TESTED FOR HARDWARE FAILURES AND OR OBVIOUS CONFIGURATION ISSUES. SOME HARDWARE DEVICES KVM SWITCH AND MOUSE WERE REPLACED. THE RESULT OF THE PROCESS WAS THE PHILIPS LOCAL SERVICE ENGINEER WAS UNABLE TO ISOLATE ANY OBVIOUS SOURCE OF FAILURE. THERE WAS A FORMAL ESCALATION IN (B)(4) TO THE PHILIPS REGION SUPPORT MANAGER. ALL LOGS FROM SERVERS WERE SENT TO PHILIPS MONITORING HQ FOR ANALYSIS. DBS REBOOTS ERROR LOGS INDICATE THAT THE SYSTEM WAS THOUGHT THAT IT HAD RUN OUT OF AVAILABLE MEMORY AND CREATED A REBOOT TO DUMP ITS MEMORY CONTENT. PHILIPS ANALYZED THE SOFTWARE CODE AND DISCOVERED AN ERROR IN THE CODE THAT CREATED AN ERRONEOUS MEMORY CAP. A PATCH WAS CREATED IN (B)(4) WHICH CORRECTED THE PROBLEM. PRINTERS GOING OFF LINE WERE HOSPITAL PROCURED PRINTERS WHICH POSSESS 2 NIC CARDS INSTEAD OF ONE. THIS PROBLEM WAS MOST PREVALENT ON THE LEAST ACTIVE PRINTERS AND PHILIPS PROVIDED 2 DIVISION PRINTERS FOR THESE AREAS WHICH WERE INSTALLED IN (B)(4) WITH NO FURTHER ISSUES. PIC REBOOTS ANALYSIS OF THE ERROR LOGS INDICATED 2 TYPES OF ERROR GENERATED BY THE PICS WHEN THEY WOULD HAVE A DOCUMENTED A REBOOT OCCURRENCE LOSS OF COMMUNICATION WITH THE DBS AND NETWORK MESSAGING A 'DR WATSON' ERROR (A SOFTWARE ERROR THAT MAY OCCUR WHEN COPYING OR PASTING FILES USING WINDOWS EXPLORER). IN (B)(4) THE MEMORY LEAK PATCH WAS PERFORMED WHICH ELIMINATED THE REBOOTS ASSOCIATED WITH THE DBS REBOOT. REBOOTS GENERATING THE OTHER ERROR CONTINUE. REMOTE SERVER NETWORK (RSN) CONNECTION ESTABLISHED TO FACILITATE DOWNLOADS OF ERROR LOGS AND DEPLOYMENT OF REPORTING TOOLS THAT CAPTURE AND DRILL DOWN. MOUSE LOCK UPS IN THE REMOTE SOLUTION LED TO INVESTIGATION OF THE TECHNOLOGY USED TO COMMUNICATE BETWEEN THE REMOTE CPU AND KVM SWITCH. THE INITIAL TECHNOLOGY INSTALLED EXTENDED THE PS/2 MOUSE COMMUNICATION PROTOCOLS BETWEEN THE CPU AND MOUSE. THE LOCKUPS TEND TO OCCUR RANDOMLY AND HAVE HAPPENED IN ALL FOUR WORKSTATIONS, HOWEVER, MOST PREVALENT IN THE STATIONS ALONG THE WINDOW. REMEDY WAS TO CONVERT THE REMOTE SOLUTION FROM A PS/2 TO A USB, REPLACE HARDWARE COMPONENTS WITH ANOTHER STYLE, AND UPGRADE TO 'KAVOOM' KVM SWITCH SOFTWARE. THIS SEEMS TO BE WORKING AS THE MONITOR TECH CAN SCROLL THROUGH THE FOUR MONITORS WITHOUT LOCKING UP, HOWEVER, LIMITS THE MONITORING CAPABILITY OF EACH MONITOR TO 12 PATIENTS INSTEAD OF THE 16 PATIENTS THAT ARE ADVERTISED. | 1
1526433 | THE FOLLOWING WAS REPORTED: 'TUBING SNAPPED AT STOPCK AND SYRING LOCATION.' >< DEVICE EVALUATION: CUSTOMER REPORT WAS CONFIRMED. REC'D (1) SINGLE DPT - VAMP PLUS KIT IN OPEN PACKAGE. KIT APPEARED NOT USED. TUBING WAS COMPLETELY BROKEN FROM SOLVENT BOND JOINT WITH THE VAMP RESERVOIR STOPCOCK. ROUGH SURFACES WERE OBSERVED AT BROKEN TUBING ENDS. BONDING SOLVENT COMPLETELY FILLED UP THE AREA BETWEEN TUBING AND VAMP RESERVOIR TUBE HOUSING. BONDING SOLVENT WAS ALSO EVIDENT AT NON-BONDED SURFACE OF THE TUBING. (B) (4) | 0
1636906 | DEVICE EVALUATION.CUSTOMER REPORT WAS CONFIRMED. REC'D (1) DOUBLE DPT - VAMP ADULT KIT, MODEL PXVK1006, IN OPEN ORIGINAL PACKAGE. KIT WAS NOT USED. PRESSURE TUBING OF ARTERIAL LINE WAS COMPLETELY BROKEN OFF FROM SOLVENT BOND JOINT WITH THE FEMALE CONNECTOR (ATTACHED TO ZERO-STOPCOCK). BROKEN TUBING WAS BENT NEAR THE BOND JOINT. (B)(4) >< THE FOLLOWING WAS REPORTED: 'MODEL NUMBER REPORTED: PXVK1006. CUSTOMER STATED THAT THE TRANSDUCER IS BROKEN ON THE TUBING THAT IS NOT CONNECTED TO THE PATIENT SIDE OF TRANSDUCER.'. | 0
1636900 | CUSTOMER STATED THAT THE TRANSDUCER IS BROKEN ON THE TUBING THAT IS NOT CONNECTED TO THE PATIENT SIDE OF TRANSDUCER.. >< DEVICE EVALUATION.CUSTOMER REPORT WAS CONFIRMED. REC'D (1) DOUBLE DPT - VAMP ADULT KIT, MODEL PXVK1006, IN OPEN ORIGINAL PACKAGE. KIT WAS NOT USED. PRESSURE TUBING OF ARTERIAL LINE WAS COMPLETELY BROKEN OFF FROM SOLVENT BOND JOINT WITH THE FEMALE CONNECTOR (ATTACHED TO ZERO-STOPCOCK). BROKEN TUBING WAS BENT NEAR THE BOND JOINT. (B)(4) | 0
1517523 | THE DATA OF WITHDRAWN BLOOD TAKEN SOON AFTER THE PATIENT WAS MOVED FROM OPERATION ROOM TO WARD WAS ABNORMAL. THE KIT WAS USED SINCE THE OPERATION. CUSTOMER TOOK ANOTHER BLOOD SAMPLE FROM FEMORAL AREA AND FOUND LARGE DIFFERENCE BETWEEN TWO DATA. THE AMOUNT OF BLOOD CAN BE TAKEN FROM VAMP FLEX IS APPROX. 2.8 TIMES MORE THAN FROM FEMORAL AREA. BLOOD GAS MEASUREMENT DATA FROM THE KIT: HB-7.3, BE-5.7, NA-139.9, K-2.46, CL-118 BLOOD GAS MEASUREMENT DATA FROM FEMORAL: HB-10.3, BE-1.0, NA-136.7, K-3.56, CL-104 KIT RETURNED IS MISSING 30CM TUBING FROM THE TIP.. >< DEVICE EVALUATION: CUSTOMER REPORT WAS NOT CONFIRMED. ENTIRE THE KIT WITHOUT 30CM LONG TUBING FROM THE TIP WAS RETURNED (COMPARED WITH DRAWING). NO VISIBLE DEFECT WAS FOUND AT TUBING, STOPCOCKS AND CONNECTORS. LEAKAGE TESTING WAS PERFORMED UNDER 300MMHG IN WATER, NO LEAKAGE WAS OBSERVED. ATTACHED SYRINGE WAS ABLE TO ASPIRATE PROPERLY. | 0
1517521 | IT WAS REPORTED THAT UNKNOWN, BROWN MATERIAL WAS FOUND ON THE SPIKE OF IV SET. | 0
1636923 | THE FOLLOWING WAS REPORTED: 'THE TUBING AT THE BOTTOM OF THE IV SET CHAMBER WAS DETACHED WHEN THE LINE WAS PRESSURIZED AT 300MMHG AFTER PRIMING.' >< DEVICE EVALUATION. CUSTOMER REPORT WAS CONFIRMED. REC'D (1) SINGLE DPT KIT. IV TUBING WAS ABLE TO BE REMOVED FROM DRIP CHAMBER. THIS CONNECTION SITE IS HEAT SEALED SITE BUT NO INDICATION OF HEAT SEALING WAS OBSERVED AT THE SEALING SITE | 0
1636922 | THE FOLLOWING WAS REPORTED: 'THE LINE CONNECTED WITH 3 WAY PORT WAS CUT.' | 0
1636920 | THE FOLLOWING WAS REPORTED: 'THE LINE CONNECTED WITH 3 WAY PORT WAS CUT.' | 0
1881181 | EVALUATION IS PROGRESS, BUT NOT YET CONCLUDED. >< THE USER REPORTED THE CENTRAL CONTROL MONITOR DID NOT REMAIN IN A SET POSITION. THE USER REPORTED THAT THE BUSHING AND PIN THAT HOLD THE MONITOR SCREEN IN PLACE WERE WORN. THE DEVICE WAS USED FOR THE PROCEDURE. THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO A PT AS A RESULT OF THIS EVENT. | 0
1003483 | ONE PT URINE SAMPLE WITH DISCREPANT RESULTS FOR BLOOD AND LEUKOCYTES. INITIAL RESULTS FOR LEUKOCYTES AND BLOOD WERE NEGATIVE. VISUAL DIPSTICK SHOWED NEGATIVE FOR BLOOD AND LEUKOCYTES. URINE SAMPLES REPEATED WITH THE FOLLOWING RESULTS: LEUKOCYTES RESULTED 500/UL. VISUAL DIPSTICK SHOWED 3 +. BLOOD RESULTED 50/UL. VISUAL DIPSTICK SHOWED 3+. ERRONEOUS RESULTS NOT REPORTED. IF ADDITIONAL INFO IS RECEIVED, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
1480153 | THE CUSTOMER CALLED TO COMPLAIN THAT THEY HAD FAILED A CAP SURVEY DEALING WITH THE URINE MICROALBUMIN / CREATININE (UALB/CREAT) RATIO TEST. DISCORDANT UALB/CREAT RATIO PATIENT TEST RESULTS WERE REPORTED. MICROALBUMIN RESULTS WERE REPORTED WITH THE WRONG UNITS. THERE IS NO KNOWN REPORT OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCE DUE TO THE DISCORDANT UALB/CREAT RATIO RESULTS. >< A SIEMENS FIELD SERVICE ENGINEER (FSE) WORKED WITH THE CUSTOMER VIA TELEPHONE TO REVIEW THE REPORTING PARAMETERS FOR UALB, CREAT AND THE UALB/CREAT RATIO TEST. ANALYSIS OF THE INSTRUMENT AND INSTRUMENT DATA INDICATE THAT THE CAUSE FOR THE DISCORDANT UALB/CREAT RATIO TEST RESULTS WAS DUE TO THE CUSTOMER CHANGING THE UALB UNITS IN THEIR LIS. THE UALB RESULTS ARE CORRECTLY REPORTED IN MG/DL IN THE CENTRALINK DATA MANAGER AND UPLOADED TO THE LIS. THE CUSTOMER'S LIS WAS SETUP TO REPORT THE UALB IN UG/ML, BUT THE CONVERSION FACTOR TO SUPPORT THE LIS REPORTING UNITS WAS NOT FACTORED INTO THEIR FORMULA. THE CUSTOMER IS NOW RUNNING WITH THE CORRECT FORMULA. THE CENTRALINK IS PERFORMING WITHIN SPECIFICATIONS. NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED. | 1
1647359 | THE USER STATED THAT THERE WAS A REPORTED PROBLEM WITH THE PATIENT'S SPO2 AND THAT THERE WAS A 'LACK OF OXYGEN ISSUE'. DURING THE SCANNING OF A MULTIPLE TRAUMA PATIENT, IT WAS REPORTED THAT THE PATIENT'S SPO2 INITIALLY WAS 96% AND THEN DROPPED TO 84% AND THEN TO 74%. THE PATIENT WAS CYANOTIC WHEN REMOVED FROM THE MRI MACHINE. A CODE BLUE WAS CALLED AND THE PATIENT WAS TRANSFERRED BACK TO ICU. >< THE USER STATED THAT THEY DID NOT KNOW IF THE DEVICE MALFUNCTIONED BUT THEY WANTED THE DEVICE EVALUATED, DUE TO THE CODE BLUE. THE USER PERFORMED NUMEROUS TESTS ON THE DEVICE, INCLUDING ECG, SPO2, NIBP AND BATTERY TESTING, AND NO PROBLEMS WERE DETECTED. THE USER ALSO TESTED THE SPO2 ACCESSORIES AND FOUND NO ISSUES. FINALLY, THE USER PERFORMED PREVENTIVE MAINTENANCE ON THE DEVICE WITH NO PROBLEMS NOTED. THE DEVICE WAS RETURNED TO THE DEVICE MANUFACTURER FOR BENCH EVALUATION. ALL MANUFACTURER SUGGESTED TESTS AND FUNCTIONAL CHECKS PASSED WITH NO ISSUES. THE DEVICE WAS ALLOWED TO RUN OVERNIGHT TO OBSERVE FOR MALFUNCTIONS AND NONE WERE DETECTED. THERE IS NO INDICATION THAT THE DEVICE MALFUNCTIONED OR FAILED TO ALERT THE CLINICIANS TO A CHANGE IN PATIENT STATUS. AS A RESULT, CLINICIANS WERE ABLE TO APPLY AVAILABLE THERAPY TO THE PATIENT WITHOUT DELAY. ALSO, THE CHANGE IN THE PATIENT'S STATUS WAS NOT UNEXPECTED, GIVEN THE PATIENT'S PRE-EXISTING CONDITION (MULTIPLE TRAUMAS). NO ADDITIONAL INVESTIGATION OR ACTIONS BY THE DEVICE MANUFACTURER ARE WARRANTED. | 0
1027957 | PRELIMINARY TESTING OF THE DEVICE DOES NOT SUPPORT THAT THERE WAS A MALFUNCTION. PHILIPS IS IN THE PROCESS OF OBTAINING MORE INFO CONCERNING THIS EVENT, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SENT ONCE THE INVESTIGATION IS COMPLETED. >< THE PT IN BED 6 DIED. THE PT WAS BEING MONITORED WITH THE FOLLOWING PARAMETERS: ECG, NIBP, TEMP, INVASIVE BLOOD PRESSURE, SPO2. THERE WAS ONLY VISUAL ALARMING AT CENTRAL STATION AND NO SOUND. A NURSE RECOGNIZED THIS AND WENT TO THE ROOM. THE PT HAD NO BLOOD PRESSURE, BUT STILL HAD HEART ACTIVITY. THE BLOOD PRESSURE VALVE SHOWN AT THE MONITOR WAS 5/3 (SYSTOLIC/DIASTOLIC) AND THERE WAS NO ALARM AT THE BEDSIDE MONITOR. REANIMATION WAS NOT SUCCESSFUL. | 1
1421934 | UNDER SPECIALIZED CIRCUMSTANCES, FOR CERTAIN COMPONENTS THAT HAVE STRINGENT MFG TIME LIMITS, IF A COMPONENT WAS MANUFACTURED OUTSIDE OF SPECIFIED TIME LIMITS, THE SYSTEM MAY NOT PREVENT IT FROM BEING DISTRIBUTED. THIS PROBLEM CAN OCCUR IF BLOOD DONATION INFO FROM A MOBILE COLLECTION IS UPLOADED TO THE SOFTWARE AFTER THE COMPONENT HAS ALREADY BEEN CREATED IN THE SYSTEM. IF THE BLOOD DONATION INFO HAS BEEN INPUT INTO THE SYSTEM BEFORE THE COMPONENT IS CREATED IN THE SYSTEM, THEN THE SOFTWARE WILL FLAG THE COMPONENT AS LATE PRODUCTION AND PREVENT IT FROM BEING DISTRIBUTED. AS A WORKAROUND, THE USER FACILITIES HAVE BEEN INSTRUCTED TO ENTER DONATION INFO FROM MOBILE COLLECTIONS USING DIFFERENT REGISTRATION SCREEN(S) OF THE SOFTWARE. A SAFETY ADVISORY, WHICH INCLUDED THE WORKAROUND, WAS DISTRIBUTED TO CLIENTS. IN ADDITION, AN URGENT SOFTWARE FIX HAS BEEN MADE AVAILABLE. >< DEVICE EVAL: SOFTWARE PERFORMANCE EVAL TESTING WAS PERFORMED TO IDENTIFY THE SOURCE OF THE PROBLEM. THE PROBLEM WAS ISOLATED TO ONE BRANCH OF LOGIC IN THE SOURCE CODE IN A SINGLE PROGRAM UNIT. THE PROBLEM IS RELATED TO THE PROGRAM NOT USING THE CORRECT DATES TO DECIDE WHETHER A PRODUCT QUALIFIES FOR THE LATE PRODUCTION PROHIBITING FACTOR, I.E., THOSE PRODUCTS THAT MUST BE MANUFACTURED WITHIN A CERTAIN PERIOD OF HOURS OR DAYS. THE PROGRAM SHOULD PASS TWO DATES TO THE CONTROL LOGIC BUT THE SYSTEM WAS PASSING ONLY ONE DATE. A CHANGE WAS MADE TO THE LOGIC TO CORRECTLY PASS DATES TO THE COMPONENT LATE PRODUCTION PROCESS DURING THE UPLOAD DONOR LINK PROCESS. ALL OTHER POSSIBLE BRANCHES OF THE LOGIC WERE REVIEWED AND IT WAS DETERMINED THAT THE SOFTWARE IS PASSING THE CORRECT DATES. RISK ANALYSIS WAS CONDUCTED TO DETERMINE WHAT OTHER FUNCTIONS COULD BE IMPACTED. THE SOFTWARE MODIFICATION WAS VERIFIED AND VALIDATED. THE SCOPE OF THE TESTING INCLUDED THE SOFTWARE MODIFICATION ITSELF, ALL POTENTIALLY IMPACTED AREAS OF THE SOFTWARE, AND OTHER REGRESSION TESTS TO ENSURE THAT THE FIX WAS EFFECTIVE AND DID NOT ADVERSELY AFFECT THE SOFTWARE. THE TESTED SOFTWARE MODIFICATION WAS MADE AVAILABLE TO CLIENTS THE WEEK OF MARCH 23, 2009. | 1
1884340 | INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. BWI CONCOMITANT PRODUCTS USED DURING THE PROCEDURE: NAVISTAR THERMOCOOL DIAGNOSTIC/ABLATION DEFLECTABLE TIP CATHETER; MODEL #: D-1197-17-S; LOT #.: 15219285M. MANUFACTURER'S REF. NO.: (B)(4). >< IT WAS REPORTED THAT DURING A VT IDIOPATHIC PROCEDURE, AFTER DEFIBRILLATING A PATIENT, THE ECG SIGNALS (THE BS AND IC SIGNALS DISAPPEARED ON BOTH CARTO SYSTEM AND THE RECORDING SYSTEM. THE USER CHANGED THE CATHETER CABLE AND THE CATHETER WITHOUT RESOLUTION. THE USER FURTHER STATED THAT AFTER THE CATHETER CHANGE, SIGNALS RETURNED BUT ONLY INTERMITTENTLY AND WOULD DISAPPEAR AGAIN AFTER PACING (THE PACING WAS SCHEDULED/ NON-EMERGENCY PACING. THE PHYSICIAN WAS ATTEMPTING TO PACEMAP A VT). AFTER EXTENSIVE TROUBLE-SHOOTING, THE PIU WAS CHANGED AND THIS SEEMED TO INITIALLY RESOLVE THE ISSUE UNTIL THE USER BEGAN APPLYING ENERGY, THE SIGNALS DISAPPEARED AGAIN. THE USER NOTED A FLICKERING GREEN LIGHT ON THE FRONT OF THE PIU AND IT WAS DETERMINED THAT THE PIU POWER CABLE COULD POSSIBLY BE AT FAULT, AND WAS RECOMMENDED TO BE CHANGED. THE USER CHANGED THE PIU POWER CABLE AND THE FO CABLE, AS WELL AS REBOOTED THE SYSTEM AGAIN. THE PIU POWER LIGHT WAS NO LONGER ON. THE PHYSICIAN DID NOT ALLOW THE BWI REPRESENTATIVE PRESENT DURING THE CASE TO FURTHER CHECK FOR ANY OTHER SOLUTIONS. THE PHYSICIAN CONTINUED THE CASE WITHOUT THE USE OF CARTO AND COMPLETED THE PROCEDURE BY UTILIZING ONLY THE RECORDING SYSTEM FOR TRACINGS AND FLUOROSCOPY FOR CATHETER LOCATION. THE PHYSICIAN WAS NOT ABLATING AND PACING AT THE SAME TIME. THE PACING WAS DONE THROUGH THE PACING PORTS AND UTILIZING THE BLOOM STIMULATOR FOR PACING. THERE WAS NO PATIENT INJURY REPORTED. >< (B)(4). IT WAS REPORTED THAT DURING THE PROCEDURE THE ECG SIGNALS (BS AND IC) DISAPPEARED ON THE CARTO SYSTEM. PIU POWER LIGHT WAS NO LONGER ON. BIOSENSE WEBSTER FIELD SERVICE ENGINEER FOUND A DEFECTIVE PIU POWER CABLE. THE CABLE WAS REPLACED. SYSTEM IS TESTED AND READY FOR USE. IN ADDITION, THE DEVICE HISTORY REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THIS DEVICE. | 1
1817482 | IT WAS REPORTED THAT DURING THE PROCEDURE, THE PATIENT BECAME ASYSTOLIC AND THE CUSTOMER WAS TEMPORARILY NOT ABLE TO ENABLE PACING ON THE CARTO 3 SYSTEM: WHEN THEY CLICKED ON 'MAP CATH' TOO QUICKLY THE DROP-DOWN FOR SELECTING THE PACING CHANNEL DID NOT OPEN. THE CUSTOMER WAS THEN ABLE TO SELECT THE PACING CHANNEL BY HOVERING WITH THE MOUSE OVER THE 'MAP CATH' OPTION WITHOUT CLICKING ON IT, WHICH OPENED THE DROP-DOWN AND ENABLED THE 'MAP 1-2' TO BE SELECTED. THE PATIENT RECOVERED FROM THE ASYSTOLE AND THE CUSTOMER WAS ABLE TO COMPLETE THE CASE WITHOUT PATIENT CONSEQUENCES. NO ADDITIONAL PATIENT INFORMATION COULD BE OBTAINED FROM THE CUSTOMER. >< MANUFACTURER'S REFERENCE NUMBER: (B)(4).THE INVESTIGATION OF THIS COMPLAINT IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. >< MANUFACTURER REF (B)(4). DEAR MR. (B)(6), ON OCTOBER 19, 2010 BIOSENSE WEBSTER INC. (BWI) RECEIVED THE FDA FOLLOW-UP REQUEST, DATED OCTOBER 09, 2010 REGARDING REPORT #9681484-2010-00018: MANUFACTURER REF. (B)(4). THE REQUEST INCLUDED THE FOLLOWING QUESTIONS: THE FINAL RESULTS OF ANY FAILURE ANALYSIS OR LABORATORY TESTING OF THE DEVICE LISTED IN THE REPORT(S) INCLUDING: A COMPLETE DESCRIPTION OF THE METHODOLOGY(IES) USED, AN IDENTIFICATION OF THE FAILURE MODE(S) AND/OR MECHANISM(S) AND THE ASSOCIATED COMPONENT(S) INVOLVED, ANY CONCLUSIONS BASED ON THE FINAL FAILURE ANALYSIS OR LABORATORY TEST RESULTS. THE COMPLAINT, AS MENTIONED IN THIS REPORT, STATED THAT THE CUSTOMER WAS NOT ABLE TO ENABLE PACING ON THE BIOSENSE WEBSTER CARTO 3 SYSTEM DUE TO THE INABILITY TO SELECT THE CORRECT PACING ELECTRODES. THE EVENT OCCURRED WHEN THE PATIENT REQUIRED EMERGENCY PACING TO RESUSCITATE AS HE HAD BECOME ASYSTOLIC. IT WAS CLARIFIED THAT THE PATIENT BECOMING ASYSTOLIC WAS NOT CAUSED BY THE CARTO 3 SYSTEM. THE PHYSICIAN MANAGED TO OVERCOME THE DISABLED PACING QUICKLY AND WAS ABLE TO PERFORM THE PACING WITHOUT ANY CONSEQUENCES TO THE PATIENT. THE FAILURE ANALYSIS RESULTS INDICATED THAT THE PHYSICIAN DID NOT FOLLOW THE IFU RECOMMENDATION WHEN EMERGENCY PACING IS REQUIRED. WHEN EMERGENCY PACING IS REQUIRED, THE CARTO 3 IFU INDICATES THAT THE USER SHOULD USE THE DIRECT CONNECTION PORTS WHICH WILL ALLOW THE USER TO PACE DIRECTLY FROM THE STIMULATOR DEVICE. IN THIS EVENT, THE USER ATTEMPTED TO PERFORM THE EMERGENCY PACING THROUGH THE PORT THAT IS CONTROLLED BY THE CARTO 3 SYSTEM INSTEAD OF USING THE DIRECT CONNECTION PORTS AS REQUIRED IN THE IFU. BY ATTEMPTING TO USE THE CARTO 3 SYSTEM CONTROLLED PORT, THE USER EXPERIENCED A SOFTWARE FAILURE WHICH CAUSED PACING DELAY. THE SOFTWARE BUG THAT CAUSED THE SOFTWARE FAILURE WAS IDENTIFIED AND IT HAS BEEN CORRECTED IN THE NEW SOFTWARE VERSION (V.1.1.2.). HOWEVER, AS PREVIOUSLY STATED, THE MAIN CAUSE OF THE EVENT WAS IDENTIFIED AS USER RELATED SINCE THE PHYSICIAN DID NOT FOLLOW THE RECOMMENDATION IN THE IFU AND SHOULD NOT HAVE USED THE PORT CONTROLLED BY THE CARTO 3 SYSTEM FOR EMERGENCY PACING. PLEASE PROVIDE THE TOTAL NUMBER OF DEVICES MANUFACTURED AND DISTRIBUTED FOR THE LAST THREE YEARS BY YEAR FOR THE DEVICE INDICATED IN THE MEDICAL DEVICE REPORT. THE BIOSENSE WEBSTER CARTO 3 SYSTEM WAS 510(K) CLEARED BY THE FDA IN OCTOBER 2009. THE FOLLOWING TABLE INDICATES THE NUMBER OF SYSTEMS MANUFACTURED AND DISTRIBUTED SINCE ITS CLEARANCE: (B)(4). IF THE PATIENT OR DEVICE PROBLEMS (INCLUDING ANY FAILURE ANALYSES) NOTED IN THIS REPORT ARE PART OF A TREND ANALYSIS, PLEASE PROVIDE A COMPLETE LIST OF ALL RELATED MDR ACCESS NUMBERS, AS WELL AS THE BASIS FOR THEIR INCLUSION IN THE TREND, E.G. COMMON CLINICAL PRESENTATION/OBSERVATION, COMMON COMPONENT, COMMON MANUFACTURING PROCESS, KNOWN FAILURE MODE, ETC. MFG REPORT NO. 9681484-2010-00012 HAS BEEN SUBMITTED RELATED TO THIS EVENT WHERE THE CUSTOMER WAS NOT ABLE TO PACE USING THE CARTO 3. THE RISK RELATED TO THE SOFTWARE BUG IDENTIFIED DURING THIS EVENT WAS EVALUATED AND IT WAS DETERMINED THAT THE PROBABILITY OF INJURY TO THE PATIENT IS LOW (B)(4). PLEASE DESCRIBE YOUR DETERMINATION OF WHETHER THIS EVENT WAS DUE TO A SOFTWARE ERROR OR USER ERROR, INCLUDING YOUR RATIONALE. IF USER ERROR, PLEASE DESCRIBE WHETHER THIS ERROR IS REASONABLY LIKELY TO OCCUR AND WHY. PLEASE ALSO DESCRIBE THE REASONING BEHIND THE CURRENT SOFTWARE DESIGN OF THIS FEATURE. THE CARTO 3 SYSTEM IS DESIGNED WITH DIRECT CONNECTION PORTS WHICH ALLOWS PACING TO BE CONTROLLED BY THE STIMULATOR RATHER THAN BEING CONTROLLED BY THE CARTO 3 SYSTEM. IN ADDITION, THE SYSTEM PROVIDES A CONNECTION FOR THE STIMULATOR WHERE THE PACING IS CONTROLLED BY THE CARTO 3 SYSTEM. AS EXPLAINED, WHEN EMERGENCY PACING IS REQUIRED, THE CARTO 3 IFU INDICATES THAT THE USER SHOULD USE THE DIRECT CONNECTION PORTS WHICH WILL ALLOW THE USER TO PACE DIRECTLY FROM THE STIMULATOR DEVICE. IN THIS EVENT, THE USER ATTEMPTED TO PERFORM THE EMERGENCY PACING THROUGH THE PORT THAT IS CONTROLLED BY THE CARTO 3 SYSTEM INSTEAD OF USING THE DIRECT CONNECTION PORTS AS REQUIRED IN THE IFU. (B)(4). | 1
1568409 | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR INFO TECHNOLOGY MFR IN (B) (4) TO SUBMIT THIS EVENT. IN A RETROSPECTIVE REVIEW OF PHILIPS COMPLAINTS/CORRECTIVE ACTIONS THIS ISSUE IS NOW BEING REPORTED AS A MDR. A POTENTIAL PROBLEM EXISTS WHILE USING THE EASYVISION. THERE IS A POSSIBLE INCOMPATIBILITY WHEN IMAGES FROM SOME FULL-FIELD DIGITAL MAMMOGRAPHY (FFDM) UNITS ARE DISPLAYED ON SOME THIRD-PARTY PICTURE ARCHIVING AND COMMUNICATION SYSTEMS (PACS). THIS MAY RESULT WHEN USING THE PRINT FUNCTION THAT IT IS POSSIBLE UNDER CERTAIN CIRCUMSTANCES TO PRINT AN IMAGE WITH TEXT COVERING THE BREAST IMAGE. ALSO, IF THE IMAGE IS ROTATED DURING REVIEW, THE VIEW INFO DOES NOT FOLLOW, OR IF THE IMAGE IS IMPORTED TO THE PACS NOT HAVING THE AXILLA TOWARDS THE TOP OF THE VIEWPORT, THE VIEW INFO MAY BE INCORRECTLY DISPLAYED. THIS PROBLEM WAS DISCOVERED FROM A FDA NOTICE. >< THE ROOT CAUSE WAS A SOFTWARE ISSUE AND IS CORRECTED BY A FIELD CHANGE ORDER, (B) (4). | 1
1675649 | THE USER REPORTED THAT A PATIENT CODED IN THE MRI MAGNET ROOM AND THAT THEY WERE HAVING PROBLEMS IN GETTING ECG DURING SOME OF THE SCAN SEQUENCES. >< WE ARE CONSIDERING THAT THERE WAS A NEED FOR EMERGENT CARE BECAUSE THE PATIENT WAS REPORTED AS HAVING CODED. THEREFORE, THIS IS CONSIDERED AS A SERIOUS INJURY FOR A PATIENT WHO WAS BEING MONITORED USING AN INVIVO MONITOR. THE AVAILABLE INFORMATION SUPPORTS THAT THERE WAS NO DEVICE MALFUNCTION OR FAILURE TO PERFORM AS INTENDED. THE DEVICE MANUFACTURER IS CURRENTLY INVESTIGATING THIS EVENT AND WILL FILE A SUPPLEMENTAL REPORT AFTER THE INVESTIGATION IS COMPLETED. | 1
1715181 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E-6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 1
1220550 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER WAS CLEANING THE BARCODE READER ON THE ABBOTT TDXFLX ANALYZER AND HER HAND SLIPPED AND WAS SCRATCHED BY THE ANALYZER'S SAMPLING PROBE. THE CUSTOMER WAS WEARING GLOVES AT THE TIME OF THE EVENT. SHE WAS SEEN BY THE EMERGENCY ROOM PHYSICIAN. THE PHYSICIAN CLASSIFIED THE INJURY AS A SCRATCH NOT A PUNCTURE. THE CUSTOMER WAS GIVEN A TETANUS SHOT. BECAUSE THE INJURY WAS CLASSIFIED AS A SCRATCH NOT A PUNCTURE WOUND, THE CUSTOMER WAS TOLD SHE DID NOT HAVE TO HAVE PROPHYLACTIC TREATMENT. SHE WAS GIVEN THE OPTION TO HAVE THE SHOTS. THE CUSTOMER DECLINED THE TREATMENT. BLOOD WAS DRAWN FOR BASELINE TESTING. FOLLOWUP WITH THE CUSTOMER IN 2008 INDICATES THE CUSTOMER IS ATTENDING SAFETY TRAINING AND SHE IS FEELING FINE. A REVIEW OF THE SERVICE HISTORY RECORDS INDICATES THERE WERE NO OTHER INJURY COMPLAINTS SINCE THE INITIATION OF THIS COMPLAINT. THE CUSTOMER'S ISSUE IS ADDRESSED IN THE TDXFLX SYSTEM OPERATION MANUAL, SYSTEM DESCRIPTION, OPERATIONAL PRECAUTIONS AND LIMITATIONS, WHICH PROVIDES CAUTIONS AND WARNINGS WHEN PERFORMING MAINTENANCE AND RUNNING THE ANALYZER. SECTION 5, MAINTENANCE, PERIODIC HAZARDS, PROVIDES INSTRUCTIONS FOR CLEANING THE BARCODE READER. WITH THE AVAILABLE INFORMATION, A DEFICIENCY OF THE TDXFLX SYSTEM WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. >< (B)(4). DATE OF EVENT WAS CORRECTED FROM (B)(6) 2000 TO (B)(6) 2008. >< THE LAY USER/PATIENT CONTACTED LIFESCAN ALLEGING THAT THERE ARE BLUE AND GREEN MARKS ON THE METER'S DISPLAY. THE REPORTED ISSUE WAS UNRESOLVED WITH TROUBLESHOOTING. THERE WERE NO ALLEGATIONS OF HARM OR INJURY. REPLACEMENT PRODUCTS WERE SENT TO THE PATIENT. >< THE CUSTOMER STATES THAT WHILE CLEANING THE BARCODE READER ON THE TDXFLX ANALYZER HER HAND SLIPPED AND THE SAMPLE PROBE/FLUID SENSING ELECTRODE ASSEMBLY SCRATCHED HER HAND. THE CUSTOMER RECEIVED A TETANUS SHOT AND BLOOD WAS DRAWN FOR FURTHER WORK-UP. THERE IS NO FURTHER IMPACT TO THE OPERATOR REPORTED. | 0
1753065 | THIS EVENT IS NOT PATIENT SPECIFIC, RATHER IT IS A SOFTWARE INTERACTION PROBLEM INVOLVING PHILIPS INTELLIVIEW PATIENT MONITOR SYSTEM AND MCKESSON HORIZON CLINICAL CHARTING SYSTEM. ISSUE IS THAT READINGS OF NIBP DO NOT ALWAYS CHART THE CORRECT NIPB VALUE WHEN THE HORIZON CLINICAL APPLICATION IS SET TO FILTER OR CHART THE VALUE EVERY 15, 30, 60, ETC. MINUTES.OUR PROBLEM IS THAT THE NIBP VALUE THAT SHOWS UP IN MCKESSON HORIZON CLINICALS IS THE VALUE FROM A MEASUREMENT TAKEN 15 MINUTES EARLIER.EXAMPLE: CLINICIAN STARTS AN NIBP READING AFTER SETTING THE PHILIPS INTELLEVIEW MONITOR (IN THIS CASE AN MP70) TO TAKE AN NIBP MEASUREMENT EVERY 15 MINUTES. THE CLINICIAN PUSHING THE NIBP START BUTTON AT 11:55. THE INTELLEVIEW MONITOR STARTS THE INFLATION OF THE CUFF, THEN STARTS DEFLATING, THEN DISPLAYS THE VALUE. THE TIME NOW IS 11:57. AT 12:00 MCKESSON HORIZON CLINICALS PERFORMS A QUERY AND RECEIVES THE NIBP VALUE. EVERYTHING IS GOOD, VALUES ARE CORRECT.THE PROBLEM NOW STARTS. THE PHILIPS INTELLEVIEW MONITOR WILL, OVER TIME, ROUND UP THE NIBP START TIME TO THE QUARTER HOUR (THIS IS A BUILT IN FEATURE OF THE SOFTWARE THE USER CANNOT DISABLE). SO FOR EXAMPLE, AN HOUR LATER, INSTEAD OF THE NIBP STARTING AT SAY 12:55, IT WILL NOW START THE NIBP MEASUREMENT AT 1:00. ALSO AT 1:00, MCKESSON HORIZON CLINICALS IS SET TO QUERYING FOR NIBP DATA FROM PHILIPS WHICH AT THIS EXACT POINT IN TIME HAS NOT COMPLETED THE NIBP MEASUREMENT AND THUS STILL HAS THE VALUE FROM THE 12:45 NIBP MEASUREMENT. THE NIBP MEASUREMENT STARTS AT THE QUARTER HOUR BUT DOES NOT HAVE A VALUE UNTIL THE MEASUREMENT HAS COMPLETED WHICH MIGHT TAKE ANYWHERE FROM 45 SECONDS TO 3 MINUTES DEPENDING ON IF THE MONITOR HAS TO RETRY 2 OR 3 TIMES, ETC.IN CRITICAL CARE AREAS PATIENTS ARE MONITORED CLOSELY FOR THE EFFICACY OF MEDICATION ADMINISTRATION, ESPECIALLY VASOACTIVE DRUGS. A DELAYED READING OF BLOOD PRESSURE COULD LEAD THE BEDSIDE NURSE TO ADJUST THE DOSAGE OF MEDICATION INCORRECTLY, LEADING TO POSSIBLE SERIOUS INJURY OR EVEN DEATH.OUR BIOMED DEPT IS UNABLE TO FIND A SOLUTION TO THE DAS PROBLEM WORKING WITH EITHER MANUFACTURER, PHILIPS AND MCKESSON.THE FOLLOWING IS FROM AN E-MAIL COMMUNICATION WITH PHILIPS SUPPORT:CUSTOMER HAS CONCERNS ABOUT THE WAY PHILIPS SYSTEM SENDS NIBP DATA VIA HL7 TO MCKESSON CHARTING SYSTEM.HE HAS A LARGE SERVER NETWORK WITH 5 DBS AND 1 PATIENT LINK WITH MPX MONITORS. THE MPX MONITORS ARE CONFIGURED FOR AUTO NIBP MEASUREMENTS EVERY 15 MINUTES. THE MCKESSON SYSTEM QUERIES THE DBS OR PL ONCE EVERY MINUTE, HOWEVER IT APPLIES A FILTER TO THE NIBP DATA SO THAT THE NURSES VIEW IT ON THE CHARTING SYSTEM ONLY IN 15 MINUTE INTERVALS.THE PROBLEM IS THE MPX MONITOR STARTS THE NIBP CLOCK WHEN THE PUMP STARTS, BUT IT MAY SOMETIMES TAKE A MINUTE OR TWO TO ACTUALLY GET A READING (DUE TO PATIENT-RELATED OR APPLICATION ISSUES). DURING THAT TIME MCKESSON WOULD HAVE QUERIED FOR THE DATA AND RECEIVED IT BUT MAY NOT BE ABLE TO DISPLAY IN THE NURSES 15 MINUTE VIEW BECAUSE OF THAT 1-2 MINUTE LAG.AS A WORKAROUND, THE CUSTOMER HAS DISABLED THE EXTERNAL TIME SOURCE AT THE MASTER DBS AND HAS MANUALLY OFFSET THE TIME 1-2 MINUTES AHEAD AT THE MASTER DBS TO COMPENSATE FOR THE TIME DISCREPANCY OF THE NIBP MEASUREMENTS.THE CUSTOMER IS SUGGESTING THAT THE MPX MONITOR BE MODIFIED SO THAT ONCE IT DETECTS AN ACTIVE HL7 CONNECTION, IT SHOULD START THE NIBP PUMP 1-2 MINUTES EARLIER SO THAT THE DATA IS NOT MISSED ON THE DISPLAY VIEW AT THE MCKESSON CHARTING SYSTEM.THE CUSTOMER HAS ALSO RECEIVED FEEDBACK FROM CUSTOMER SUPPORT AT OTHER NON-PHILIPS SITES THAT (ANOTHER) VENDOR'S MONITORS DO NOT HAVE THIS PROBLEM BECAUSE THEY ALLOW YOU TO OFFSET THE NIBP START TIME A LITTLE BIT. | 1
1714081 | PATIENT 12 LEAD ECG RECORDS FOUND BY EDIT LAB STAFF IN THE MUSE UNCONFIRMED REPORT LISTING WITH INACCURATE NAMES ASSIGNED TO RECORDS BY AUTOMATIC PROCESS IN MUSE. -RECORDS TRANSMITTED TO MUSE AS PID: 999999999, NAME: ABSENT. -RECORDS FOUND IN UNCONFIRMED REPORT LISTING AS PID: 999999999, NAME: LASTNAME, FIRSTNAME ISSUE COMMUNICATED TO BIOMED BY THE HOSPITAL'S CARDIO ECG LAB STAFF IN MID 2008. BIOMED LATER INVESTIGATED AND COMPLETED ADDITIONAL TESTING TO CONFIRM THE REPORTED ISSUE. BIOMEDICAL / CLINICAL ENGINEERING: AS HOSPITAL MEDICAL AND INFORMATION TECHNOLOGY INTERFACE AND SHARE DATA, THE INTEGRITY OF EACH SYSTEM'S DATABASE BECOMES INCREASINGLY INTERDEPENDENT. IT IS NECESSARY TO DEVELOP AN UNDERSTANDING OF ALL DATA PROCESSES THAT OCCUR BEFORE, DURING AND FOLLOWING DATA ACQUISITION AND STORAGE IN ORDER TO VERIFY PATIENT RECORD FIDELITY. THIS REPORT DETAILS THE CHARACTERIZATION OF IMPROPER AUTOMATED HOSPITAL INFORMATION SYSTEMS DATA MERGES WITH REGARD TO PATIENT RECORDS IN A GE HEALTHCARE MUSE CLINICAL DATABASE. INITIAL DISCUSSIONS WITH GE: NO CORRECTIONS TO THE NORMALIZATION PROCESS WOULD BE MADE IN VERSION 5 - WILL CONSIDER AS A FEATURE REQUEST FOR VERSION 7. THE PROBLEM IS ONGOING. THE LATEST VERSION (7.1.1) DOES NOT BRING RESOLUTION TO THE ISSUE. GE'S RESPONSE INCLUDED RECOMMENDATIONS FOR PRACTICE CHANGE, BUT DID NOT SPECIFICALLY ADDRESS THE PROBLEM REPORTED IN ANY OF THE NEW VERSIONS (INCLUDING 7). GE HAS DECIDED TO MAKE THE CONCERN A FEATURE REQUEST. GE BASICALLY RECOMMENDED THAT USER ALWAYS ASSIGNS THE PROPER PID AND NAME BEFORE TRANSMISSION TO THE SYSTEM. DO NOT CONFIRM PATIENT RECORDS CONTAINING DEFAULT PID. ANY COMBINATION OF ALPHANUMERIC CHARACTERS WILL BE ACCEPTED AS A VALID PID. DETERMINE IF THERE ARE ANY EXISTING CONFIRMED RECORDS WITH PID SEQUENCES OF ALL 9S OR ANY OTHER DEFAULT PID AS DEFINED BY YOUR INSTITUTION. TO SET DEFAULTS PID SETTINGS FOR YOUR INSTITUTION CREATE A REPORT FOR EACH DEFAULT PID THAT CONTAINS THAT PID # AND ENTER 'UNKNOWN PATIENT' OR SIMILAR IN PATIENT NAME FIELD AND CONFIRM THE REPORT. ANY SUBSEQUENT RECORDS WITH A DEFAULT PID AND NO PATIENT NAME WILL NOW HAVE NAME FIELD POPULATED WITH UNKNOWN PATIENT. THIS WILL ALERT USER TO THE FACT THAT RECORD NEEDS TO BE ASSIGNED CORRECT PATIENT DEMOGRAPHICS. NEW LOGIC SHOULD REQUIRE A MINIMUM OF 2 PATIENT IDENTIFIERS FOR ANY MERGE TO TAKE PLACE. ADDITIONAL SAFEGUARDS SHOULD BE IN PLACE AT 12 LEAD CARTS AND BEDSIDE MONITORS THAT OPERATE AS SOURCES OF PATIENT DIAGNOSTIC INFORMATION. THESE SAFEGUARDS SHOULD LIMIT THE POSSIBILITY FOR RECORD EXPORT WITHOUT ADEQUATE PATIENT ID INFORMATION. | 1
1400450 | POSSIBLE FALSE NEGATIVE, DIAGNOSTIC CELLS OUTSIDE FIELD OF VIEW (FOV). CAS FELT THE CASE PRESENTED HAD TRIGGER CELLS IN THE 22 FOV THAT WOULD PROMPT AN AUTOSCAN. THE CUSTOMER DISAGREED. THE LAB IS WAITING FOR TISSUE CONFIRMATION. THE CUSTOMER DID NOT WANT TO SEND THE SLIDE IN FOR INVESTIGATION. SITE WILL BE MONITORED TO DETERMINE IF EXPECTED OCCURRENCES (AS DETERMINED BY (B) (6) STUDY) ARE EXCEEDED. CTS ON SITE WILL CONTINUE TO HOLD QUESTIONABLE CASES FOR CAS REVIEW ON THE NEXT VISIT. CAS HAS OFFERED TO REVIEW TRIGGER CELLS WITH THE CUSTOMER. THE CUSTOMER FELT A REVIEW WAS NOT NECESSARY AT THIS TIME. | 1
1400451 | POSSIBLE FALSE NEGATIVE. NO DELAY IN PATIENT DIAGNOSIS. ABNORMAL CELLS OUTSIDE FIELDS OF VIEW (FOV). TRIGGER CELLS NOT PRESENT IN THE FOV TO TRIGGER A FULL SCAN OF THE SLIDE, SLIDE WILL BE ANALYZED TO DETERMINE IF THE SLIDE REPRESENTS A 'RARE EVENT' AS EVIDENCED IN TABLE 8 OF THE THINPREP IMAGING SYSTEM OPERATION SUMMARY AND CLINICAL INFORMATION. | 1
1506825 | POSSIBLE FALSE NEGATIVE. ABNORMAL CELLS FOUND OUTSIDE FIELDS OF VIEW (FOV). NO TRIGGER CELLS PRESENT IN 22 FOVS TO PROMPT A FULL SCAN OF THE SLIDE. SLIDE WILL BE ANALYZED TO DETERMINE IF IT IS A 'RARE EVENT' AS DEFINED IN TABLE 8 OF THE THINPREP IMAGING SYSTEM OPERATION SUMMARY AND CLINICAL OPERATION. | 1
1657425 | RN B LOGGED ONTO HED (MCKESSON'S HORIZON EXPERT DOCUMENTATION) ON MOBILE COMPUTER AND CHARTED HER ASSESSMENTS ON HER 5 ASSIGNED PATIENTS. AT SOME POINT DURING HER HED CHARTING SHE NOTICED THAT A DIFFERENT NURSE'S NAME WAS NOW NOTED ON BOTTOM OF SCREEN. ALL OF HER CHARTING WAS CREDITED TO ANOTHER OFF-DUTY RN. THE HELPDESK CONTACTED HED RESOURCE RN AND THE CHARTING WAS ALL REMOVED. B THEN LOGGED IN AGAIN AND RE-DID ALL OF HER CHARTING. IT WAS FELT THAT THE OTHER NURSE HAD NOT LOGGED OFF PROPERLY, CREATING THE 'MIX-UP'. AT THE END OF HER SHIFT, B CAREFULLY LOGGED OFF, BELIEVING THAT THE PRIOR NURSE LOG OFF WAS SOMEHOW NOT DONE PROPERLY. RN C THEN ASSUMED CARE OF THOSE SAME PATIENTS AND SHE LOGGED ONTO THAT SAME COMPUTER. SHE BEGAN HER CHARTING ON HED AND ALSO NOTED SOME TIME INTO HER CHARTING THAT THE NAME AT THE BOTTOM OF THE SCREEN WAS NO LONGER HERS, BUT B. SHE CONTACTED B WHO WAS STILL IN-HOUSE AND NOTIFIED THE HELP DESK OF THE RECURRENT PROBLEM. A TECH WAS SENT TO CHECK OUT THE COMPUTER ON WHEELS. HE SEEMED TO KNOW IMMEDIATELY WHAT THE PROBLEM WAS. THE UNIT WAS PULLED FROM SERVICE UNTIL HE COULD 'RE-PROGRAM IT'.LESS THAN A MONTH AGO, A DIFFERENT RN ON A DIFFERENT UNIT EXPERIENCED THE SAME THING. SHE LOGGED OFF OF HER COMPUTER WHEN FINISHED CHARTING. NURSE TAKING OVER FOR HER CHARTED ON HED AFTER LOGGING IN AND NEVER NOTICED THAT THE FIRST NURSE'S NAME APPEARED THROUGHOUT THE PATIENT'S RECORD ON THE NEXT SHIFT. THE DAY SHIFT NURSE FOUND THE ERROR THE NEXT MORNING WHEN SHE RESUMED CARE OF THAT PATIENT. IT APPEARED ON THE HED THAT THE FIRST NURSE HAD PROVIDED ALL CARE AND MEDICATIONS FOR OVER 24 HOURS. THE HELP DESK WAS NOTIFIED AT THAT TIME AND THE CONCLUSION WAS AGAIN 'LOG-ON ERROR BY NURSE'. THIS SAME PROBLEM WAS NOTED SEVERAL MONTHS AGO WHILE INVESTIGATING A 'MISSED OPPORTUNITY' TO GIVE AN ANTIBIOTIC WITHIN 4 HOURS OF ARRIVAL. ALTHOUGH THE PATIENT DID NOT ARRIVE AT THE HOSPITAL UNTIL NOON, THE NURSE'S HED CHARTING INDICATED THAT THE ADMISSION ASSESSMENT WAS DONE AT 0700 WHICH IS WHEN THAT PARTICULAR NURSE HAD FIRST SIGNED ON FOR THAT SHIFT ON THAT COMPUTER. SINCE DISCOVERY OF THE ISSUES, IT NURSES HAVE BEEN ABLE TO REPRODUCE THE SCENARIO, WHICH CAN HAPPEN IF A CERTAIN SEQUENCE OF EVENTS TALKS PLACE. SHORT VERSION IS, THE SYSTEM IS WORKING AS DESIGNED.LET ME EXPLAIN THAT LAST STATEMENT:IT IS POSSIBLE FOR ONE TO LOG INTO THE SYSTEM AND THEN LOG OUT OF PART OF THE SYSTEM WITHOUT COMPLETELY LOGGING OUT OF ALL OPEN APPLICATIONS.SEQUENCE OF EVENTS TO DUPLICATE WHAT HAPPENED:NURSE LOGS INTO CARE ORGANIZER.NURSE SELECTS PATIENT.NURSE OPENS FLOW SHEET.FLOW SHEET OPENS IN A NEW WINDOW.AT THIS TIME BOTH CARE ORGANIZER AND THE FLOW SHEET HAVE THE SAME PERSON LOGGED IN.BECAUSE THERE ARE 2 WINDOWS YOU NEED TO SWITCH BETWEEN THEM TO SELECT OTHER PATIENTS, DO OTHER TASKS, ETC IN CARE ORGANIZER AND THEN SWITCH BACK TO THE FLOW SHEET. SOMETIMES THE FLOW SHEET (OR BOTH WINDOWS) IS MINIMIZED.NEXT, IT'S TIME TO LOG OUT.NURSE SELECTS CARE ORGANIZER.NURSE LOGS OUT.CARE ORGANIZER CLOSES BUT THE FLOW SHEET IS STILL OPEN AND LOGGED IN AS YOU. CARE ORGANIZER DOES NOT AUTOMATICALLY CLOSE THE OPEN FLOW SHEET WINDOW.THE OPEN FLOW SHEET MAY NOT BE NOTICED BECAUSE IT MAY BE MINIMIZED.NEXT NURSE COMES ALONG AND OPENS UP CARE ORGANIZER.NURSE LOGS INTO CARE ORGANIZER.NURSE SELECTS PATIENT.NURSE OPENS THE FLOW SHEET.MINIMIZED FLOW SHEET IS CHANGED TO THE PATIENT SELECTED BY THE NURSE.NURSE MAXIMIZES THE FLOW SHEET AND BEGINS TO CHART. HOWEVER, THE PRIOR NURSE IS STILL LOGGED INTO THE FLOW SHEET BECAUSE THE APPLICATION WAS NOT CLOSED BY THE FIRST NURSE.WE WERE ABLE TO REPRODUCE THE ABOVE SCENARIO WITH EASE.OK, SO WHAT DO WE DO ABOUT THIS ISSUE?FIRST, REMINDER TO STAFF TO LOG OUT OF ALL OPEN APPLICATIONS WHEN THEY LEAVE THEIR DEVICE.SECOND, WE ARE DEPLOYING A REPLACEMENT FOR THE CARE ORGANIZER WITH WHAT IS KNOW AS COMMON APPLICATION FRAMEWORK OR CAF. CAF REPLACES THE CURRENT CARE ORGANIZER WITH A CONSISTENT WAY OF DISPLAYING PATIENT DATA ACROSS MULTIPLE APPLICATIONS. MORE IMPORTANTLY, THE FLOW SHEET NO LONGER OPENS IN A SEPARATE WINDOW. IT IS NOW A TAB IN CAF. SO WHEN YOU LOG OUT OF CAF, YOU LOG OUT OF EVERYTHING. | 1
1891078 | THE CUSTOMER REPORTED THE PACHYMETRY PROBE IS NOT MEASURING. THE PT'S HAD TOPICAL ANESTHETIC. THE PROCEDURES WERE CANCELLED. THE PROCEDURES WERE COMPLETED WHEN THE NEW PROBE ARRIVED. NO PT INJURY WAS REPORTED. >< THE COMPANY SERVICE REP EXAMINED THE SYSTEM AND COULD NOT DUPLICATE THE PROBLEM REPORTED. THE SOFTWARE WAS UPDATED. THE SYSTEM WAS THEN TESTED AND MET ALL PRODUCT SPECIFICATIONS. THE PACHYMETRY PROBE WILL BE RETURNED FOR IN-HOUSE EVAL. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4). | 1
1890103 | (B)(4). OPERATORS NOTICED A VOLUME DISCREPANCY IN THE COLLECTED WASTE BAG, AS AN INSUFFICIENT VOLUME OF WASTE HAD BEEN COLLECTED COMPARED TO WHAT WAS EXPECTED. THE CUSTOMER STATED THAT THERE WAS A KINK BELOW THE CCM. THE DISPOSABLE SET WAS RETURNED FOR EVAL. DURING THE KIT INVESTIGATION A CLOSE EXAMINATION OF THE SEGMENT OF TUBING BETWEEN THE CCM AND 3-LUMEN CONNECTOR SHOWED NO EVIDENCE OF KINKING. THIS INDICATES THAT THE PRODUCT WAS NOT MIS-PACKAGED PRIOR TO STERILIZATION DUE TO THE ABSENCE OF A WELL-DEFINED KINK. THE PLASMA LINE KINKING MUST HAVE BEEN INDUCED BY THE SPECIFIC LOADING OF THE DISPOSABLE. UPON DHR INVESTIGATION FOR THIS LOT NUMBER NO EVIDENCE OF MFG ERRORS RELATED TO THIS FAILURE MODE WERE FOUND. >< THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. THE CUSTOMER CALLED AND STATED THAT THEY WERE PERFORMING A TPE PROCEDURE AND NOTICED THAT WHILE THEY HAD USED APPROXIMATELY 4L OF REPLACEMENT FLUID THERE IS ONLY ABOUT 1 L (VISUAL) IN THE PLASMA BAG. THEY HAVE DONE PROCEDURES FOR THIS PT BEFORE. ABOUT 35 MINS INTO THE TPE PROCEDURE, THE PT BEGIN TO COMPLAIN OF SOME PAIN AND DISCOMFORT, SOME 'BACK PRESSURE', SOME NUMBNESS AND TINGLING IN THE LIPS. THE STAFF THOUGH IT MIGHT BE DUE TO A CITRATE REACTION AND DECREASE THE INLET FLOW RATE, INCREASED HIS CALCIUM INFUSION, AND CHANGED THE RATIO FROM 15 TO 20 TO TRY TO ALLEVIATE THE SYMPTOMS. THE PT TOLD TRISTAN THAT HE HAS HAD SOME FEVERS AND CHILLS THE PREVIOUS EVENING. TRISTAN THEN THOUGHT IT MIGHT THAT THE PT MAY HAVE HAD AN INFECTION AROUND HIS PORT AND AS A RESULT MIGHT BECOMING SEPTIC. THEY HAD DRAWN LABS PRIOR TO BEGINNING PROCEDURE AND GOT RESULTS BACK DURING THE PROCEDURE. | 0
1879037 | ON THE MORNING OF THE EVENT, A CENTRAL MONITORING UNIT (CMU) TECH WAS REORDERING THE DISPLAY OF TELEMETRY PATIENTS ON THE SAME CARE UNIT. THE CMU TECH INADVERTENTLY SWITCHED TWO PATIENT ROOM NUMBERS, PLACING PATIENT 'A' IN ROOM (B)(6) AND PATIENT 'B' IN ROOM (B)(6) (IN ACTUALITY, THE PATIENTS WERE IN OPPOSITE ROOMS). THE CMU TECH WAS CHANGING PATIENT TELEMETRY SECTORS WHEN SHE SWAPPED THE PATIENTS BETWEEN SECTORS (B)(6). LATER IN THE MORNING, THE RHYTHM FOR PATIENT 'A' CHANGED AS HE BEGAN TO EXHIBIT SIGNIFICANT CARDIAC PAUSES. THUS, THE NURSING STAFF QUICKLY RESPONDED TO PATIENT 'B' IN ROOM (B)(6). AFTER NOTING THAT THE PATIENT NAME WAS DIFFERENT FROM THE TELEMETRY MONITOR DISPLAY, THE CLINICAL STAFF RACED TO ROOM (B)(6) AND ATTENDED TO PATIENT 'A', THEREBY AVERTING TREATMENT TO A WRONG PATIENT WHILE ANOTHER PATIENT WAS UNATTENDED FOR A POTENTIALLY LIFE-THREATENING EVENT. SHORTLY AFTER STABILIZING PATIENT 'A', A NURSE NOTICED THAT THE TELEMETRY PATIENT NAMES DID NOT CORRESPOND CORRECTLY FOR THE ROOM NUMBERS. CMU WAS NOTIFIED OF THE SITUATION AND THE ERROR WAS CORRECTED. PATIENT 'A' WAS TRANSFERRED TO CVICU. VERIFIED THAT A CMU TECH DID INDEED TRANSPOSE TWO PATIENT ASSIGNMENTS. LACKING AN ADT INTERFACE (CALLED AN HAI INTERFACE BY PHILIPS); THE CURRENT MANUAL DATA ENTRY METHOD DOES NOT HAVE A SAFEGUARD TO ENSURE DATA ACCURACY. HOSPITAL HAS BUDGETED FOR HAI INTERFACE AND IMPLEMENTED A PROCESS OF VERIFYING PATIENT AND ROOM LOCATIONS IN THE PHILIPS CENTRAL SYSTEM.HEALTH PROFESSIONAL'S IMPRESSION: PRODUCT ALLOWED FOR INCORRECT CORRELATION OF ROOM NUMBERS AND PATIENT NAMES. | 1
1516850 | THE INITIAL REPORTER CALLED BECAUSE ALL WORKSTATIONS WERE RUNNING SLOW. UPON FURTHER INVESTIGATION, IT WAS IDENTIFIED THAT THERE IS A POTENTIAL THAT GUARDIAN DATA BACKUPS WERE NOT OCCURRING COMPLETELY. >< THERE IS NO ESTABLISHED EXPIRATION DATE FOR THIS DEVICE. DEVICE MANUFACTURED DATE IS UNK AT THIS POINT. FURTHER ATTEMPTS WILL BE MADE FOR THIS INFO. EVALUATION SUMMARY - (B) (4) PERFORM TESTS WERE RUN TO SUBSTANTIATE PERFORMANCE ISSUES RELATED TO THE GUARDIAN BACKUP SERVICE RUNNING IN THE BACKGROUND. THE SITE WAS SET TO START THE GUARDIAN BACKUP SERVICE AT 8:00 P.M. AND STOP AT 6:00 A.M. THERE IS A POTENTIAL THAT SOME DATA WAS NOT COMPLETELY BACKED UP. FURTHER INVESTIGATION IS ONGOING. THIS IS NOT A SINGLE USE DEVICE. | 1
1882441 | (B)(4). FDA REGISTRATION NUMBER: REPORTING SITE (SPECIFICATION DEVELOPER): 2243969. MANUFACTURING SITE: 1049092. >< REPORTED BY THE COMPLAINANT AS FOLLOWS: FMS INSERTED (B)(6). ON (B)(6), A WOUND AT OPENING OF ANUS, LOCATED ALL AROUND THE TUBING (ALL AROUND ANAL OPENING) WAS NOTICED. CLINICAL EDUCATOR STATES: THIS IS THE WORSE WOUND SHE HAS SEEN DEVELOP FROM A FECAL MANAGEMENT DEVICE IN HER CAREER, AND WANTS THIS FACT KNOWN. SHE ALSO STATES THE WOUND WAS AT A STAGE OF NECROSIS. SHE CANNOT SAY IF STAFF HAD PERFORMED ALL WOUND PREVENTION MEASURES (TURNING PATIENT Q2HRS AS PER PROTOCOL, HAVE THERE BEEN LEAKS PRIOR, ETC¿). THEY DID NOT DISCONTINUE USE OF FMS, AS PATIENT WAS C-DIFF AND WITH DIARRHEA +++ PATIENT DECREASED (B)(6). KNOWN FOR SMALL CELL LUNG CA, C-DIFF COLITIS, PNEUMONIA, AS A RESULT OF PNEUMONIA, DEVELOPED PSEUDOMONAS, THEN WENT SEPTIC AND DIAGNOSIS WAS THAT HE EXPIRED DUE TO THE SEPSIS. WAS ON IV NOREPINEPHRINE. | 0
1728707 | REPORTED BY THE COMPLAINANT AS FOLLOWS: MALE END-USER ADMITTED TO MICU APPROXIMATELY THE 1ST WEEK OF (B)(6) 2010. FMS SIGNAL INSERTED FOR DIAGNOSIS OF C-DIFF AND INFECTED SACRAL WOUND. NURSE COULD NOT DETERMINE THE FMS SIGNAL INSERTION DATE. ADMITTING DIAGNOSIS OF POSSIBLE SEPSIS; METABOLIC ENCEPHALOPATHY; ANEMIA; SACRAL WOUND INFECTION. THE END-USER WAS MOVED OFF THE MICU ON 3/10 AND ADMITTED TO AN ACUTE REHAB UNIT AT A SELECT SPECIALTY HOSPITAL AND ON (B)(6) WAS ADMITTED BACK TO TRINITY MICU FOR A LOWER GI BLEED. IT WAS REPORTED THAT SOMETIME BETWEEN (B)(6) AND (B)(6) THE FMS SIGNAL WAS REMOVED. A COLONOSCOPY WAS PERFORMED WITH FINDING OF BLEEDING A RECTAL ULCERATION. INTERVENTIONS: TRANSFUSION OF 5 UNITS OF PRBC AND BLEED CAUTERIZED. END-USER WITH NO FURTHER BLEEDING. NURSE MANAGER; REPORTS NO INOTROPIC DRUGS; NO HISTORY OF ANAL OR PERI-RECTAL ABNORMALITIES SHE COULD LOCATE. ALSO REPORTS NO ANTICOAGULANTS OR ANTI-PLATELETS. REPORTS END-USER ON LOPRESSOR DRUGS. | 0
1667137 | (B) (4). (B) (4). (B) (4). (B) (4). >< REPORTED BY THE COMPLAINANT AS FOLLOWS: PT IS A WOMAN, AGED (B) (6). ILEAL CONDUIT WITH STENT. EVENT OCCURRED DURING NIGHT BETWEEN 6 AND 7 DAYS AFTER OPERATION. URINE DID NOT DRAIN FROM POUCH AND THAT CAUSED FEVER. PT WAS TAKEN CARE OF BY NURSE AND DOCTOR ON NIGHT DUTY. MILKING, WITHDRAWAL OF CV, RENAL PELVIC IRRIGATION. ANTIBIOTIC WAS ADMINISTERED. AFTER 2 DAYS, CONDITION BECAME STABLE. COMPLAINANT HOSPITAL CONCLUDES THAT ROOT CAUSE OF PT FEVER IS PRODUCT DEFECT. | 0
1809074 | THE NURSE WAS INSTRUCTED TO REMOVE THE RECTAL TUBE (FECAL MANAGEMENT SYSTEM). HOWEVER, AFTER SHE REMOVED THE FLUID, IT STILL WOULD NOT COME OUT. AFTER DOING A RECTAL EXAM, IT WAS FOUND THAT THE BALLOON OF THE FECAL MANAGEMENT SYSTEM WAS STILL INTACT AND IT DID FINALLY COME OUT. THE NURSE TRIED TO REMOVE THE FLUID AGAIN, BUT WAS UNSUCCESSFUL. | 0
1656460 | THE MEDICATION REVIEW SCREEN OF THE SUBJECT DEVICE DOES NOT SPECIFY THE EXACT DOSE IN MILLIGRAMS OF COMBINATION MEDICATIONS. FOR EXAMPLE, NARCOTICS ARE COMBINED WITH TYLENOL IN AT LEAST TWO STRENGTHS. LIQUID NARCOTIC TYLENOL-OXYCODONE COMBINATION IS REPORTED IN ML, NOT MG. THE EXACT DOSE OF TYLENOL IS NOT SPECIFIED AND REQUIRES KNOWLEDGE OF THE COMBINATION MEDICATION DOSE IN THE VOLUME SPECIFIED. CERTAIN FIELDS OF THE GRID DO NOT SPECIFY THE VOLUME, BUT RATHER STATE 'DATE/TIME' REQUIRING ANOTHER CLICK OR POP UP SCREEN. THE IMMEDIATE KNOWLEDGE OF TYLENOL DOSAGE IN MG IS DIRECTLY RELATED TO UNDERSTANDING AND PREVENTING EXCESSIVE DOSES. IN THE SUBJECT, 10 ML OF ACETAMINOPHEN-OXYCODONE IS INDICATED AS HAVING BEEN GIVEN 3 TIMES OVER 4 HOURS. THAT MEANS THAT 1950 MG OF TYLENOL WAS ADMINISTERED IN 4 HOURS WHILE THE PATIENT WAS IN A STATE OF STARVATION AND RECEIVING OTHER MEDICATION THAT INCREASE THE EFFECTS OF TYLENOL. THIS DOSE WOULD EQUATE TO 11,700 MG OF TYLENOL OVER 24 HOURS, NEARLY 3 TIMES THE MAXIMUM DAILY DOSE IN OTHERWISE HEALTH PEOPLE. IN THE ENSUING DAYS, THE PATIENT DEVELOPED ACUTE RENAL FAILURE, PRESUMABLY ACUTE TUBULAR NECROSIS, AND DIED. IN THE ABSENCE OF OTHER ETIOLOGY, THE EXCESS TYLENOL WAS THE CULPRIT. THIS WAS NOT CONSIDERED AS ETIOLOGY ANTE-MORTEM. THE COUNTERINTUITIVE SCREEN IMPAIRED THE PROFESSIONALS. THE PHARMACIST DID NOT RECOGNIZE AND STOP THE MEDICATION, THE NURSES ADMINISTERED IT, AND THE EXCESSIVE DOSE, CLINICALLY MEANINGLESSLY LISTED AS A VOLUME OF 10 ML -GIVEN 3 TIMES IN 4 HOURS- OF ACETAMINOPHEN-OXYCODONE, WAS MISSED BY THE PHYSICIANS. ADVERSE EVENTS HAVE BEEN ASCRIBED TO 'USER ERROR' BY VENDORS. THE DEVICE OFFERS A POTENT PROPENSITY TO LIFE ENDANGERING OVERSIGHTS. THERE ARE OTHER SCREENS ON THIS DEVICE WHICH PRESENT INFORMATION THAT INTERFERE WITH CLINICALLY USEFUL VISUALIZATION OF DATA. THE DATA DOES NOT FLOW TO THE PROFESSIONALS. IT IS NOT REPRESENTED IN A MEANINGFULLY USEFUL MANNER. THE PROFESSIONALS NEED TO HUNT FOR IT. AS SUCH, THE USER UNFRIENDLY SCREENS IMPAIR SAFE MEDICAL CARE CONSISTENT WITH THE IMPEDIMENT TO EXPEDIENT PROFESSIONAL UNDERSTANDING OF WHAT, EXACTLY, IS THE DOSE OF MEDICATION AND HOW MUCH WAS ADMINISTERED TO THE PATIENT. THIS SENTINEL CASE OF DEATH IS DIRECTLY ATTRIBUTED TO USER UNFRIENDLY SCREENS ON THIS DEVICE. | 1
1642178 | THIS TELEMETRY PATIENT WAS HOOKED UP TO A TRANSMITTER TELEMETRY PACK FOR THE PURPOSE OF CONTINUOUS CARDIAC MONITORING. AT ONE POINT THE MONITOR TECH NOTICED THAT THE PATIENT WAS OFF MONITOR. THE RN WENT TO CHECK TO SEE WHAT THE PROBLEM WAS. SHE CHECKED THE LEADS AND THE TELEMETRY BOX AND NOTICED THAT THE TELEMETRY BOX WAS VERY HOT. THE BOX WAS IN THE PATIENT'S GOWN POCKET. WHEN SHE REMOVED THE BOX SHE NOTICED THAT FLUID HAD LEAKED/DRIPPED OUT OF THE BATTERY COMPARTMENT. SHE ALSO NOTICED THAT THE MONITOR WAS SMOKING FROM THE BATTERY COMPARTMENT AND THE BATTERIES WERE HISSING. THE PATIENT WAS UNHARMED. THE PATIENT DENIED FEELING ANY HEAT RADIATING FROM THE MONITOR AT THAT TIME. THE MONITOR WAS IMMEDIATELY TAKEN OUT OF SERVICE AND EVALUATED BY BIO-ENGINEERING THE FOLLOWING DAY. ACCORDING TO THE BIOMEDICAL ENGINEER THERE WAS FLUID IN THE BATTERY COMPARTMENT WHICH COULD HAVE BEEN FROM THE BATTERY OR SOME OTHER UNKNOWN SOURCE. THE INSIDE OF THE TRANSMITTER BATTERY COMPARTMENT IS BURNT AND UNUSABLE. IT WAS RETURNED TO THE MANUFACTURER, PHILIPS FOR A REPLACEMENT. | 1
1663784 | THIS REPORT WAS FILED IN RESPONSE TO UFR (B)(4): THE REPLACED PATO SENSOR HAS BEEN TESTED AT MFG SITE AND IS WORKING TO SPECIFICATION. THE ERROR LOG OF APOLLO AND THE ERROR LOG OF THE GAS MEASUREMENT MODULE HAVE BEEN ANALYZED. EVERY 5 MINS DURING A CASE A SET OF ALL GAS CONCENTRATIONS ARE LOGGED BY THE APOLLO. THIS LOG INCLUDES THE ERROR LOG AND THE INFO LOG OF THE GAS MODULE WHICH CONTAINS ALL RELATED FAILURE MESSAGES OF THE GAS MEASURING SYSTEM. BOTH LOGS SHOW A COMMAND SEQUENCE SENT FROM THE APOLLO TO RESTART THE GAS MEASUREMENT MODULE AT TWO TIMES ON (B)(4) 2010: AT 9:26 AM AND THE SECOND AT 9:31 AM. A RESTART MEANS THAT DIRECTLY NO GAS READINGS ARE AVAILABLE FOR NEARLY TWO MINUTES AFTER THIS PHASE CO2 AND N20 ARE SHOWN IN REDUCED ACCURACY MODE AND AFTER 4 MINS THE WARM UP PHASE ENDS WITH A ZERO PROCEDURE - TAKING ANOTHER MINUTE. IF THE ZERO IS SUCCESSFUL AT FIRST TIME ALL GAS VALUES ARE DISPLAYED IN NORMAL ACCURACY. THIS RESTART OVERALL TAKES APPROX 5 MINUTES. NO REASON COULD BE IDENTIFIED TO EXPLAIN WHAT COULD HAVE TRIGGERED APOLLO TO SEND THE 'RESTART' COMMAND. POSSIBLY THE USE OF CELL PHONES OR OTHER ELECTRICAL EQUIPMENT MIGHT HAVE DISTURBED THE FUNCTION OF THE APOLLO. >< USER FACILITY REPORT REF NO. (B)(4). IT WAS REPORTED THAT DURING A CASE, THE AGENT, O2 AND N2O VALUES BLANKED OUT INTERMITTENTLY. THERE WAS NO PT INJURY REPORTED. A TECHNICIAN WAS DISPATCHED AND REPLACED THE PATO SENSOR. THE TECHNICIAN TESTED THE APOLLO AFTER COMPLETING THE REPAIR AND REPORTED THAT ALL TESTS PASSED TO DRAEGER SPECIFICATIONS. THE MACHINE HAS BEEN RETURNED TO USE. | 1
1784020 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< (B)(4). THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. | 0
1496155 | THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED, A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. >< THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR INFO SYSTEMS MANUFACTURER IN (B) (4) TO SUBMIT THIS EVENT THAT HAPPENED IN (B) (6). PROBLEM: CUSTOMER REPORTED THAT THE PT X-RAY IMAGE COULD NOT BE FOUND IN THIS VIEWFORUM WORKSTATION (PICTURE ARCHIVING COMMUNICATIONS SYSTEM) AFTER THE X-RAY SCANNING. HOWEVER, SINCE THE PT DATA IN THE X-RAY SYSTEM HAD ALREADY BEEN ERASED, THE PT IMAGE DATA WAS LOST. | 1
1653582 | CERNER'S CPOE SYSTEM HAS BEEN DEPLOYED. EVERY TEST IS ORDERED ELECTRONICALLY INCLUDING TESTS ON BLOOD SAMPLES AND RADIOLOGY TESTS IN ADDITION TO MEDICATIONS. IT HAS BEEN DISCOVERED THAT TESTS ARE SPURIOUSLY CANCELED ELECTRONICALLY AND NOT BY PHYSICIAN ORDER. THE STATED REASONS FOR CANCELLATION ARE ESTABLISHED BY THE EQUIPMENT WITHOUT NOTIFICATION OF THE PHYSICIANS. THIS CASE INVOLVES UNEXPLAINED BLOOD CLOTS IN A PT IN WHOM HYPERCOAGULABILITY TESTS WERE ORDERED, BUT NOT DONE. THEY WERE CANCELED BY THE SYSTEM 30 MINUTES AFTER THE ORDER WAS ENTERED WITH THE REPORTED REASON GENERATED BY COMPUTER BEING: INADEQUATE COLLECTION OF SPECIMEN. | 1
1302818 | TO DATE, THE DEVICE INVOLVED IN THE REPORTED INCIDENT HAS NOT BEEN RECEIVED FOR EVALUATION. AN INVESTIGATION HAS BEEN INITIATED BASED UPON THE REPORTED INFO. >< THE REPORTER STATE THAT A 1P2P WAS PLACED IN THE PT. AFTER PLACEMENT, THE PRODUCT BEGAN TO LEAK AND HAD TO BE REPLACED. THERE WERE NO RELATED ADVERSE EVENTS FOR THE PT. | 0
1119286 | THE PROBLEM WAS REPORTED FROM A CUSTOMER WHO WAS COMPARING MEASUREMENT DATA PROVIDED BY ISITE IN CORRELATION TO A KNOWN OBJECT. IN THIS EVENT ISITE REPORTED DIMENSIONAL DISCREPANCIES IN THE ORDER OF MAGNITUDE OF 1.5 X AND 2X. THE ROOT CAUSE OF THE PROBLEM STEMS FROM THE KNOWLEDGE THAT THE FUJI MODALITY WAS SENDING DICOM MEASUREMENT (B) (4), (B) (4) IN LIEU OF (B) (4), (B) (4), WHICH WAS RECOGNIZED BY ISITE. THE ISITE DICOM CONFORMANCE STATEMENT CLAIMS COMPLIANCE TO DICOM V. 3.0 FOR WHICH ISITE PACS 4.1 WAS DEVELOPED UNDER DICOM 3.0 RELEASE 2003. DURING THE PERIOD FROM 2003 TO 2008, DICOM DID NOT SPECIFY THE USE OF (B) (4), (B) (4) AND RECOMMENDED THE USE OF (B) (4), (B) (4). IN 2006, DICOM PUBLISHED A STATEMENT WHICH WARNED AGAINST THE USE OF (B) (4), (B) (4). THE ISITE PACS IFU STATES THAT MEASUREMENTS TAKEN USING ISITE ARE FOR REFERENCE ONLY AND THAT IF PRECISE MEASUREMENTS ARE REQUIRED THEN THE OPERATOR IS INSTRUCTED TO CALIBRATE ISITE USING A KNOWN FIDUCIARY. IN THE 2008 RELEASE OF DICOM V. 3.0, THE (B) (4), (B) (4) IS SPECIFIED FOR USE IN CONJUNCTION WITH (B) (4), (B) (4). THE POTENTIAL FOR FAILURE IS THE UNDERLYING FACT THAT MEASUREMENT DISCREPANCIES COULD RESULT IN PATIENT MISDIAGNOSIS FOR MISTREATMENT. >< DEVICE FROM RESERVE SAMPLE EVALUATED. | 1
1660121 | AN OCD FIELD ENGINEER INSPECTED THE INSTRUMENT. THE ISSUE WAS UNABLE TO BE REPRODUCED.(B)(4) >< THE CUSTOMER REPORTED THAT THE PROVUE PROBE IS DRIPPING. NO ERROR MESSAGES WERE POSTED. THE CUSTOMER INDICATED THAT THE PROVUE WAS TAKEN OUT OF USE. NO ERRONEOUS RESULTS WERE POSTED. | 0
1813613 | DETAILS WILL BE PROVIDED WITH THE ACTION THAT WILL BE CONDUCTED IN COOPERATION WITH THE U.S. FOOD AND DRUG ADMINISTRATION. (B)(4). >< A REMEDIAL ACTION HAS BEEN INITIATED TO NOTIFY THE CUSTOMER AND REMOVE AFFECTED PRODUCT FROM THE FIELD. ABBOTT POINT OF CARE WILL BE TAKING A MULTI-PHASE APPROACH IN REMOVING EXISTING GCX MOUNTABLE DOWNLOADER RECHARGERS AND REPLACING THEM WITH THE NEWEST PRODUCT DESIGN. THE NEW DESIGN INCLUDES IMPROVEMENTS TO: REDUCE THE ECG NOISE INTERFERENCE BETWEEN THE GCX MOUNTABLE DOWNLOADER RECHARGER AND THE PHILLIPS PT MONITORING SYSTEM THROUGH THE USE OF A NEW POWER SUPPLY, PROVIDE A STRAIN RELIEF FOR THE POWER CORD CONNECTOR AND ADDRESSES THE RECHARGING PIN FAILURES. THIS ACTION WILL BE CONDUCTED IN COOPERATION WITH THE U.S. FOOD AND DRUG ADMINISTRATION. | 0
1282697 | UNDER SPECIALIZED SITUATIONS, IF THE SOFTWARE IS SETUP TO ALLOW AN ELECTRONIC CROSSMATCH FOR A PATIENT THAT NEEDS TO HAVE AN ANTIBODY IDENTIFICATION TEST, THEN THE USER MAY NOT BE STOPPED FROM ISSUING THE BLOOD PRODUCT IF THE ANTIBODY IDENTIFICATION TEST HAS NOT BEEN COMPLETED. THE USER WILL RECEIVE AN ALERT THAT THE IDENTIFICATION TEST HAS NOT BEEN CONDUCTED. WHEN THE USER CLICKS THE CANCEL BUTTON ON THE WINDOW THAT DISPLAYS THE ALERT, AND THEN PERFORMS THE SAME ACTION THAT PROMPTED THE ALERT, THEN THE ALERT IS REMOVED. IF THE SOFTWARE IS CONFIGURED TO NOT ALLOW AN ELECTRONIC CROSSMATCH TO BE PERFORMED FOR A PATIENT THAT NEEDS TO HAVE AN ANTIBODY IDENTIFICATION TEST, OR ELECTRONIC CROSSMATCH IS NOT PERFORMED, THEN THE PROBLEM CANNOT OCCUR. AS A WORKAROUND, THE USER FACILITIES HAVE BEEN INSTRUCTED TO NOT ALLOW ELECTRONIC CROSSMATCHES FOR PATIENTS THAT HAVE HAD A POSITIVE ANTIBODY SCREEN. A SAFETY ADVISORY, WHICH INCLUDED THE WORKAROUND, WAS DISTRIBUTED TO ALL CLIENTS. IN ADDITION, A SOFTWARE FIX FOR THE PROBLEM HAS BEEN DISTRIBUTED TO ALL CLIENTS. >< SOFTWARE PERFORMANCE EVALUATION TESTING WAS PERFORMED TO IDENTIFY THE SOURCE OF THE PROBLEM. THE PROBLEM WAS ISOLATED TO ONE BRANCH OF LOGIC IN THE SOURCE CODE IN A SINGLE PROGRAM UNIT. IT WAS DETERMINED THAT WHEN THE ELECTRONIC CROSSMATCH WAS SET TO BE RUN FOR PATIENTS THAT HAD A POSITIVE ANTIBODY SCREEN, AND THE CANCEL BUTTON ON A SPECIFIC WINDOW WAS CLICKED, THEN THE LOGIC BRANCH WAS NOT RESET AS EXPECTED. A CHANGE WAS MADE TO RESET THE LOGIC BRANCH TO ITS CORRECT STATE. ALL POSSIBLE WAYS THAT THE BRANCH OF LOGIC WAS EXECUTED AND COULD BE RESET IN THE SOFTWARE WERE IDENTIFIED. RISK ANALYSIS WAS CONDUCTED TO DETERMINE WHAT OTHER FUNCTIONS COULD BE IMPACTED. THE SOFTWARE MODIFICATION WAS VERIFIED AND VALIDATED. THE SCOPE OF THE TESTING INCLUDED THE SOFTWARE MODIFICATION ITSELF, ALL POTENTIALLY IMPACTED AREAS OF THE SOFTWARE, AND OTHER REGRESSION TESTS TO ENSURE THAT THE FIX WAS EFFECTIVE AND DID NOT ADVERSELY AFFECT THE SOFTWARE. THE TESTED SOFTWARE MODIFICATION WAS DISTRIBUTED TO ALL CLIENTS THE WEEK OF NOVEMBER 10, 2008. | 1
1828148 | ABBOTT MOLECULAR INC. IS AWARE OF THE PROBLEM WITH THE VP 2000 PROCESSOR BASINS. AN URGENT PRODUCT RECALL (3005248192-06/18/2010-001-C) WAS DISTRIBUTED VIA (B)(6) SHIPMENT TO UNITED STATES CUSTOMERS BASED ON CUSTOMER RECEIPT OF VP 2000 PROCESSOR AND BASIN SPARE PART SINCE (B)(6) 2008. A MDR (3005248192-2010-0002) WAS ALSO FILED ON JULY 02, 2010. THREE POTENTIAL HAZARDS WERE IDENTIFIED: DELAY IN RESULTS. THERE IS NO HARM ASSOCIATED WITH THIS HAZARD CHEMICAL EXPOSURE. THERE IS NO HARM ASSOCIATED WITH THIS HAZARD AS THIS ISSUE DOES NOT INCREASE THE ALREADY EXPECTED EXPOSURE TO THESE REAGENTS. THEREFORE, THERE IS NO INTRODUCED OPERATOR HARM. SLIP HAZARD. EVALUATION DETERMINED THAT THE OVERALL HEALTH RISK WAS LOW, BASED ON THE FOLLOWING: A. THE OVERALL ESTIMATE FOR A LEAKING BASIN WAS 1.1% (BASED ON BASINS KNOWN TO BE RELATED TO THIS ISSUE) B. LEAKING RATE IS EXPECTED TO BE ABOUT 1 TO 2 DROPS PER MINUTE < 100 UL PER MIN.). WHEN A LOADED BASIN IS EITHER USED IN THE VP2000 OR STORED ON THE LAB BENCH, IN THE TYPICAL USAGE TIME < 8 HOURS), THE LEAKED LIQUID WILL BE LESS THAN 50 ML C. SLIP HAZARD WILL BE AVOIDED BY THE OPERATOR'S OBSERVANCE OF THE LIQUID ON THE FLOOR AND BY THE SMALL VOLUME OF LIQUID. I. IF A LEAK WERE TO OCCUR, IT WOULD OCCUR IN AN AREA THAT A LAB OPERATOR WOULD NOT NORMALLY WALK (IN THE VP2000) OR DURING REAGENT ADDITION (LAB BENCH). IN ADDITION, THE LEAK WOULD BE VISUALLY OBVIOUS TO THE OPERATOR. D. IF THE OPERATOR DID STEP ON THE LEAK, IT WOULD NOT NECESSARILY LEAD TO A FALL THAT WOULD LEAD TO A SERIOUS INJURY. >< THE ABBOTT VP 2000 PROCESSOR IS A DEVICE DESIGNED TO AUTOMATE AND STANDARDIZE SLIDE SPECIMEN PROCESSING AND ROUTINE SLIDE STAINING FOR THE LABORATORY. CUSTOMER CALLED IN REGARDS TO THE PRODUCT RECALL NOTIFICATION FOR POTENTIAL VP 2000 PROCESSOR LEAKING BASINS AND INDICATED THAT THEY HAVE LEAKING BASINS. THE CUSTOMER HAS 5 BASINS BETWEEN 2 MACHINES. THE NOTIFICATION ASKED FROM CUSTOMERS TO TAKE THE FOLLOWING ACTION: 'IF YOU NOTICE LEAKING OR WARPED BASINS OR AN INABILITY TO PLACE THE BASINS INTO THE VP 2000 PROCESSOR, PLEASE CALL YOUR LOCAL ABBOTT MOLECULAR REPRESENTATIVE FOR COORDINATION OF BASIN REPLACEMENT. AN ABBOTT MOLECULAR CUSTOMER SERVICE REPRESENTATIVE OR FIELD SERVICE REPRESENTATIVE WILL BE CALLING YOU TO COORDINATE REPLACEMENT OF YOUR VP 2000 HEATED REAGENT BASINS BEFORE (B)(6) 2010. PLEASE NOTE THAT ADDITIONAL MEASURES HAVE BEEN TAKEN TO ENSURE CURRENTLY AVAILABLE INVENTORY DOES NOT EXHIBIT THIS ISSUE.' ABBOTT MOLECULAR DISTRIBUTED AN URGENT PRODUCT RECALL (3005248192-06/18/2010-001-C) VIA (B)(6) SHIPMENT TO UNITED STATES CUSTOMERS BASED ON CUSTOMER RECEIPT OF VP 2000 PROCESSOR AND BASIN SPARE PART SINCE (B)(6) 2008. A MDR (3005248192-2010-0002) WAS ALSO FILED ON JULY 02, 2010. >< ON (B)(4) 2012, THROUGH AN INTERNAL AUDIT INVESTIGATION IT WAS DISCOVERED THAT THE DATE INCLUDED IN THE ORIGINAL MDR REPORT WAS INCORRECT. THE PURPOSE OF THIS MDR FOLLOW-UP REPORT #1 IS TO UPDATE THE ORIGINAL INCORRECT DATE FROM (B)(6) 2010 TO THE CORRECT DATE (B)(6) 2010. | 0
1467973 | THE USER REC'D DIFFERING LEUKOCYTE RESULTS FOR THREE METER READINGS AND ONE VISUAL READINGS OF TEST URINE STRIPS FOR ONE URINE SAMPLE. THE INITIAL METER READING WAS NEGATIVE, SECOND METER READING WAS '+' WHICH IS EQUIVALENT TO 75 LEU PER UL, THIRD METER READING WAS '++' WHICH IS EQUIVALENT TO 500 LEU PER UL. THE VISUAL READING PERFORMED BY THE OPERATOR WAS '++'. THE PT WAS NOT TREATED BASED ON THE METER OR VISUAL READINGS. | 0
1863409 | MEDICATION CART POWER SUPPLY BATTERY HARNESS CONNECTORS OVERHEAT AND FAIL.====================== MANUFACTURER RESPONSE FOR MEDICATION CART WITH COMPUTER AND UPS, LIONVILLE======================PROBLEM HAS OCCURRED ON SEVERAL MEDICATION CARTS. THE MANUFACTURER REPLACED BATTERIES AND POWER SUPPLIES ON SEVERAL. | 0
1774294 | IT WAS REPORTED IN A SERVICE REPORT THAT THE STRETCHER WOULD NOT RAISE AT THE HEAD END WITH PATIENT WEIGHT. NO ADVERSE CONSEQUENCES WERE ALLEGED. >< INVESTIGATION CONCLUDED THE ROCHE INSTALLER ADDED AN EXTRA SPACE IN A REGISTRY KEY WHICH CAUSED THE UPGRADE (FROM VERSION 8.01 TO 8.08) NOT TO COMPLETE SUCCESSFULLY AND THE DATA INNOVATIONS RULES TO NOT MIGRATE. THE ROCHE INSTALLER DID NOT VERIFY THE SOFTWARE UPGRADE ACCORDING TO PROCEDURES. INSTALLATION INSTRUCTIONS FOR ROCHE INSTALLERS WILL BE UPDATED TO CONTAIN APPROPRIATE VERIFICATION STEPS. PATIENT RESULTS WERE RELEASED BUT THEN THE LABORATORY DETERMINED SAMPLES SHOULD BE RERUN. NO PATIENT WAS TREATED WITH THE ORIGINAL RESULTS. >< FOLLOWING A SOFTWARE UPGRADE OF THE MIDDLEWARE BY A FIELD SERVICE REPRESENTATIVE, THE RULES, THAT CONTROL FUNCTIONS INCLUDING SAMPLE HANDLING, DID NOT GET UPDATED. SAMPLES DID NOT GET ARCHIVED AND THE SERUM INDICES TESTS DID NOT RUN BECAUSE THE RULES NO LONGER EXISTED AFTER THE UPGRADE. DUE TO THE LACK OF MIDDLEWARE RULES, AN INCORRECT SAMPLE TYPE WAS USED FOR TESTING ONE PATIENT SAMPLE. THE SAMPLE BYPASSED THE CENTRIFUGE AND WAS RUN AS WHOLE BLOOD INSTEAD OF SERUM. THE WHOLE BLOOD SAMPLE GAVE A SODIUM RESULT OF 134 MMOL/L AND A POTASSIUM RESULT OF 5.8 MMOL/L. THE RESULTS WERE REPORTED OUTSIDE THE LABORATORY. THE USER RECOGNIZED THE ERROR AND MANUALLY SPUN THE TUBE AND PLACED IT BACK ON THE ANALYZER. THE SERUM SODIUM RESULT WAS 136 MMOL/L AND THE POTASSIUM RESULT WAS 5 MMOL/L. THE SERUM RESULTS WERE REPORTED WITH A CORRECTIVE ACTION. NO TREATMENT WAS PROVIDED TO THE PATIENT BASED ON THE INCORRECT RESULTS. THE CURRENT CONDITION OF THE PATIENT IS 'FINE'. THE FIELD REPRESENTATIVE MANUALLY RESTORED THE MIDDLEWARE CONFIGURATION TO PUT THE RULES BACK IN PLACE AND CORRECTED THE EXISTING DATABASES. SHE VERIFIED THE USER IS RUNNING WITH ALL RULES PRESENT. >< . | 1
1644347 | SPACELABS CONTACTED THE HOSPITAL'S RISK MANAGER, WHO WAS ABLE TO OFFER LITTLE INFO ON THE ALLEGED INCIDENT. SHE DID STATE THAT THE HOSPITAL IS NOT REPORTING ANY MALFUNCTION, THE EQUIPMENT PERFORMED AS PROGRAMMED TO THE BEST OF HER KNOWLEDGE. NO SPECIFIC MODEL NUMBER OR SERIAL NUMBER INFO IS AVAILABLE FROM THE SITE. THE HOSPITAL'S CONCERN INVOLVES THE CONFIGURATION OF THE DEFAULT ALARM SETTINGS IN THE TELEMETRY SYSTEM, THE LOW BATTERY AUDIBLE ALARM HAD BEEN SET TO OFF. THE HOSPITAL HAS BEEN TOLD THAT (B) (4) REQUIRES ALL AUDIBLE ALARMS TO BE FUNCTIONAL. IN THIS CASE THE FACTORY SET DEFAULT FOR THE LOW BATTERY AUDIBLE ALARM HAD BEEN CHANGED BY THE HOSPITAL TO OFF. THE HOSPITAL HAS CONFIRMED THAT THE VISUAL, ON SCREEN LOW BATTERY ALARM WAS FUNCTIONAL. THE PRODUCT LABELING DOES NOT DESCRIBE PROPER USE OF THE CONFIGURABLE SETTINGS AND DOES RESTRICT ACCESS TO THESE SETTINGS. THE HOSPITAL BIOMED HAS ALREADY CORRECTED THIS ISSUE AND RETURNED THE LOW BATTERY AUDIBLE ALARM DEFAULT SETTING TO ON. NO FURTHER SERVICE WAS REQUIRED. ALL EQUIPMENT IS CURRENTLY IN USE MONITORING PATIENTS. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED. >< THE CUSTOMER REPORTED THAT THEIR TELEMETRY SYSTEM FAILED TO PROVIDE AN AUDIBLE ALARM FOR A LOW BATTERY CONDITION. THE STAFF ALLEGES THAT THE PT PASSED AWAY AS A RESULT. | 1
1627049 | THIS IS COMPUTER HARDWARE. THE PACS SOFTWARE RUNS ON THE HARDWARE. IN THIS CASE THE FIRMWARE WAS SUCCESSFULLY UPDATED AND THE SYSTEM WAS STABILIZED. THE PATIENT DATA WAS BACKED UP PRIOR THE FIRMWARE UPGRADE. THE LIKELY CAUSE OF THE MALFUNCTION IS THAT THE FIRMWARE VERSION, THE HARDWARE WAS NOT UP TO DATE. THIS IS NOT A SINGLE USE DEVICE. >< THE PACS SYSTEM'S SERVER REQUIRED A COMPLETE SYSTEM REBUILD DUE TO IBM FIRMWARE ISSUES. WHEN THIS OCCURS, THERE IS A POTENTIAL FOR PATIENT DATA LOSS OR DATA CORRUPTION ON THE SERVER. | 1
1381563 | THE USER ALLEGED THAT DURING A PROCEDURE THEY COULD NOT STOP OR HALT PACING. THE CATHETERS WERE REMOVED FROM THE PATIENT AND THERE WAS NO ADVERSE CONSEQUENCE TO THE PATIENT. THE USER STATED THAT HE REALIZED, AFTER THE FACT, THAT HE COULD HAVE STOPPED THE STIMULATION BY PRESSING THE TOUCH SCREEN 'HALT' BUTTON. >< (B)(4). THE FIELD SERVICE ENGINEER DETERMINED THAT THE COMPLAINT WAS RELATED TO A EP-WORKMATE COMPUTER SYSTEM CRASH. THE EP-WORKMATE SYSTEM WAS EVALUATED AT THE USER FACILITY. THE ROOT CAUSE OF THE EP-WORKMATE COMPLAINT COULD NOT BE DETERMINED. | 0
1031783 | IT WAS REPORTED THAT AN OPERATOR SUSTAINED A LEFT HAND INJURY WHEN TRYING TO ADJUST A PIPETTE TIP WHILE THE INSTRUMENT WAS OPERATING. THE INJURY REQUIRED 9 STITCHES, A TETANUS SHOT, AND MRI. NO INFORMATION HAS BEEN RECEIVED TO SUGGEST ANY ILL EFFECTS ASSOCIATED WITH THE INJURY. >< THE OLA2500 USER'S GUIDE WARNS (IN MULTIPLE LOCATIONS) THE USER OF INJURY FROM MOVING COMPONENTS IF THE SYSTEM IS NOT PLACED IN THE STOP MODE BEFORE PERFORMING SERVICE ACTIVITY, AND THE POSSIBILITY THAT THE SYSTEM COULD BEGIN MOVING WITHOUT WARNING IF THIS MEASURE IS NOT TAKEN. TRAINING REINFORCES THESE PRECAUTIONS AND INSTRUCTIONS IN ADDITION, A SAFETY WARNING STICKER IS FIXED TO THE COVER TO ALERT OPERATOR TO 'KEEP HANDS AWAY FROM MOVING PARTS.' THE ACCIDENT REPORT INDICATE THE OPERATOR DID NOT PUT THE SYSTEM IN STOP MODE PRIOR TO PLACING HIS HAND INSIDE THE PERFORM A SERVICE ACTIVITY. EMPLOYEE ADMITTED FAILURE TO FOLLOW PRECAUTIONS AND INSTRUCTIONS THAT CONTRIBUTED TO THE EVENT. SAFETY RELATED SITE PROCEDURES WERE ALSO DISREGARDED. THE SYSTEM WAS INSPECTED BY AN OLYMPUS FIELD SERVICE ENGINEER AND PERFORMED AS EXPECTED. | 0
1677717 | A 14 BEAT RUN OF VENTRICULAR TACHYCARDIA THAT DID NOT TRIGGER AN ALARM WAS OBSERVED. THE SOFTWARE DESCRIBED THE RUN AS ARTIFACT. WHEN THIS OCCURS, THE ALARM ON THE LEAD WITH THE ARTIFACT IS SUSPENDED WHILE THE SYSTEM SEARCHES FOR THE RETURN OF A GOOD RHYTHM IN THAT LEAD. DURING THE TIME THE ALARM IS SUSPENDED ONLY LETHAL ARRHYTHMIA ALARMS WILL BE ACTIVATED. THE LETHAL ALARMS INCLUDE ASYSTOLE, V FIB, V TACH, AND V FIB. A REVIEW OF THE 7 LEAD STRIP WAS REVIEWED BY THE NURSING STAFF, MEDICAL STAFF AND TECHNICIANS, AND AGREEMENT WAS UNANIMOUS THAT IT WAS A RUN OF VENTRICULAR TACHYCARDIA AND SHOULD HAVE TRIGGERED THE ALARM. NO PATIENT HARM RESULTED FROM THIS INCIDENT.====================== HEALTH PROFESSIONAL'S IMPRESSION======================THERE WAS NO PATIENT HARM AS A RESULT OF THIS INCIDENT ALTHOUGH IF NOT CLOSELY OBSERVED BY A CLINICIAN A MORE SUSTAINED RUN OF V TACH WITHOUT AN ALARM BEING TRIGGERED COULD LEAD TO PATIENT HARM.====================== MANUFACTURER RESPONSE FOR MONITOR, TELEMETRY, APEXPRO CH======================RESPONSE FROM THE MANUFACTURER IS THAT THE SOFTWARE DID NOT IDENTIFY V TACH BECAUSE THE COMPLEXES WERE NOT WIDE ENOUGH (>0.12 SECONDS FOR 6 CONSECUTIVE BEATS), THE QRS TOTAL HEIGHT WAS TOO LARGE AND CONSIDERED OUT OF RANGE AND A TACHY ALARM WAS NOT TRIGGERED BECAUSE THAT TYPE OF ALARM WAS SUPPRESSED BY THE 'ARTIFACT' ALARM. THEY SUGGEST CHANGING THE ELECTRODE CONFIGURATION AND MOVING THE ELECTRODES AWAY FROM THE HEART. | 0
1511841 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER RESULTED 3 DIFFERENT SAMPLES AS FALSE POSITIVE DURING ANTIBODY SCREEN. AN FE WENT TO THE SITE AND FOUND THAT WASH SOLUTION A CONTAINER/BOTTLE ONLY CONTAINED DI-IONIZED WATER WITH NO WASH SOLUTION A REAGENT IN THE BOTTLE. INCORRECT PREPARATION OF THE ABSENCE OF WAS SOLUTION MAY CAUSE CARRY-OVER, CROSS CONTAMINATION OR HEMOLYSIS OF REAGENTS OR SAMPLES, WHICH COULD RESULT IN ERRONEOUS TEST RESULTS. >< ON 9/25/2009, AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND FOUND A LEAK AT THE BOTTOM FITTING OF THE WAS STATION AND THE DISTRIBUTION BLOCK. THE FE REPLACED THE WASH STATION, PROBE AND TIGHTENED THE FITTING ON DISTRIBUTION BLOCK. QC PASSED WITHOUT ANY ERRORS. THE CUSTOMER CONTACTED OCD (CUSTOMER TECHNICAL SERVICES) TO REPORT THAT THE FALSE POSITIVE ISSUE WAS NOT RESOLVED. ON 9/26/2009, ANOTHER FE WENT TO THE SITE AND FOUND WHAT WASH SOLUTION A CONTAINER/BOTTLE ONLY CONTAINED DI-IONIZED WATER WITH NO WASH SOLUTION A REAGENT IN THE BOTTLE. THERE WERE NO RESULTS COMPROMISED. THE CUSTOMER RERAN THE SAMPLES MANUALLY BEFORE REPORTING RESULTS. THE CUSTOMER PREPARED A NEW WASH SOLUTION AND PERFORMED SEVERAL PRIMES AND FINAL WASH. QC PASSED. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. IF THE PROBE IS NOT WASH PROPERLY, IT COULD LEAD TO CARRY OVER, CROSS CONTAMINATION OR HEMOLYSIS OF REAGENT OR SAMPLES. USER ERROR COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B)(4). | 0
1900977 | THIS ADVERSE EVENT ENTAILS A CRITICALLY AND SUB ACUTELY ILL PT WITH MULTI ORGAN FAILURE AND DELIRIUM, HAVING SUSTAINED RESPIRATORY FAILURE REQUIRING TRACHEOSTOMY SEVERAL WEEKS EARLIER, BUT WAS CLINICALLY IMPROVING. WE HAVE FOUND THAT EMR AND CPOE ROUTINELY REQUIRE DETAILED ATTENTION FROM THE HEALTH CARE PROFESSIONAL USER. WHEN DEVICES ARE OF POOR USABILITY AND NOT SUBJECT TO AFTER MARKET SURVEILLANCE WHICH MIGHT ACTUALLY RESULT IN A FIX, THEY REQUIRE ADDITIONAL TIME OF THE HEALTH CARE PROFESSIONAL. WHEN THE FOCUS IS DEVOTED TO THE CPOE AND EMR, WE OBSERVE THAT PTS ARE NOT WATCHED AS CLOSELY AS THEY HAD BEEN PRIOR TO IMPLEMENTATION RESULTING IN NEGLECT AND DELAYS IN CARE. THIS PT PULLED OUT THE TRACHEOSTOMY TUBE WHEN ATTENTION OF HIS CARE TEAM WAS WITH THE EMR AND CPOE WITH THE ATTENDANT NEGLECT. DUE TO MULTIPLE ANATOMICAL ANOMALIES, IT COULD NOT BE REINSERTED. IN ASSOCIATION WITH THE PT'S ADVANCE DIRECTIVES, THE PT DIED WITHOUT FURTHER INTERVENTION. | 0
1899274 | AFTER USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED, A PROBLEM WITH TRANSFERRING FILES FROM THE SATELLITE PRINTED CIRCUIT TO THE FLASH DRIVE. THE ISSUE WAS SOLVED BY REBOOTING THE PRINTED CIRCUIT. SINCE THE EVENT OCCURRED AFTER USE OF THE DEVICE, THERE WAS NO PT INVOLVEMENT DURING THIS EVENT. | 1
1697706 | THE USER REPORTED WHILE TRYING TO QUERY INFO USING THE DATA MANAGEMENT SYSTEM, EXCEPTION ERRORS WERE OBSERVED. NO OTHER DETAILS REGARDING THE EVENT WERE PROVIDED. THE USER REPORTED THERE WAS NO ADVERSE CONSEQUENCE TO A PT AS A RESULT OF THIS EVENT. >< EVALUATION IN PROGRESS BUT NOT CONCLUDED. | 1
1611340 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REGARDING ERRONEOUSLY HIGH HEMOGLOBIN (HGB) RESULTS WITHOUT INSTRUMENT GENERATED FLAGS ON TWO (2) PATIENT SAMPLES: THE INITIAL RESULTS WERE 12.5 AND 10.0G/DL FOR PATIENTS A AND B, RESPECTIVELY. THE RESULTS WERE NOT REPORTED OUT OF THE LAB CUSTOMER RE-TESTED BOTH SPECIMENS TWICE AND OBTAINED LOWER HGB RESULTS. CUSTOMER CONSIDERED THE RERUN RESULTS TO BE CORRECT. BASED ON AVAILABLE INFORMATION, THERE WAS NO AFFECT TO PATIENT TREATMENT. >< QC WAS RUN BEFORE AND AFTER THE EVENT AND RESULTS RECOVERED WITHIN EXPECTED RANGE. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED: THE FSE REPLACED THE HGB CUVETTE ASSEMBLY DUE TO A SPILL OF DILUENT. THE FSE VERIFIED THE INSTRUMENT OPERATION. ALTHOUGH HARDWARE WAS ADDRESSED BY THE FSE, A CLEAR ROOT CAUSE FOR THIS EVENT HAS NOT BEEN DETERMINED. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. | 0
1835703 | FLEXIBLE DIGITAL URETEROSCOPE STOPPED WORKING DURING CASE. THE DISPLAY WENT BLACK.====================== HEALTH PROFESSIONAL'S IMPRESSION======================NO ADVERSE EVENT. | 0
1694749 | IT WAS REPORTED BY THE USER THAT THE CLINIC WAS HAVING ISSUE WITH THEIR SCOUT LINES AND LUMBAR POSITIONING. BASICALLY WHEN THEY VIEW THE IMAGES THEY ARE WRONG, BUT WHEN THEY BURN THEM TO A CD, THEY ARE CORRECT. >< THIS DEVICE WAS EVALUATED REMOTELY, WE WERE ABLE TO LOGIN TO THE CUSTOMER DATABASE. THIS IS SOFTWARE PRODUCT, NOT SINGE USE DEVICE. EVAL SUMMARY (B)(4): ON (B)(6) 2010, A CLINICAL ADMINISTRATOR OF (B)(6) JOINT AND SPINE CLINIC BROUGHT UP A SITUATION TO STRYKER TECH SUPPORT (B)(4) WHO WAS VISITING THE SITE WHERE THEY ARE HAVING ISSUES WITH THEIR SCOUT LINES ON LUMBAR MRI STUDIES. A LUMBAR TYPE OF AN MRI STUDY RELATES TO THE SPINE IN THE AREA OF THE BACK BETWEEN THE RIBS AND THE HIP BONES AND IT INCLUDES THE 5 LUMBAR VERTEBRAE FROM L1 TO L5. BASICALLY WHEN THEY VIEW THE SCOUT LINES ON THE SAGGITAL VIEWS (IMAGINARY VERTICAL PLANE SEPARATING THE BODY INTO LEFT AND RIGHT HALVES OF THE LUMBAR MRI STUDY IN THE OFFICEPACS POWER CLIENT) THEY ARE INCORRECTLY POSITIONED AS REFERENCED BY THE CORRESPONDING AXIAL SLICE VIEWS; THIS MEANS LOOKING AT THE BODY FROM TOP TO BOTTOM OF THE ANATOMICAL REGION REPRESENTED BY THE SCOUT LINES ON THE SAGGITAL VIEWS. HOWEVER, WHEN THEY BURN THE STUDY (OF THE LUMBAR MRI IMAGES) TO A CD, THE SCOUT LINES APPEAR CORRECT. ON (B)(6) 2010, FURTHER INVESTIGATION WAS CONDUCTED BY TESTING A SAMPLE MRI STUDY FROM THE CLINIC. IT WAS CONCLUDED THAT THE CALCULATION WHICH PLACES THE SCOUT LINES ON THE SAGGITAL VIEWS OF THE MRI IMAGES AS REPRESENTED BY THE CORRESPONDING AXIAL SLICE VIEWS OF THE SAME MRI IMAGES WAS SHIFTED BY A GIVEN BUT UNK MARGIN OF ERROR. THIS RESULTED IN THE ERROR WHERE THE SCOUT LIENS WERE ACTUALLY REPRESENTING A DIFFERENT ANATOMICAL AREA THAN WHAT THE AXIAL VIEWS WERE SHOWING. MORE INVESTIGATION SHOWED THAT THE LIKELY CAUSE OF THE CALCULATION ERROR WAS RELATED TO AN INCORRECT INSTALLATION OF OFFICEPACS POWER CLIENT EITHER INITIALLY (ON SITE SINCE THE CLINIC WAS ON OFFICEPACS 3.4 AND WAS UPGRADED ON SITE TO POWER) OR WHILE UPGRADING TO A LATER VERSION OF POWER. THIS SEEMS TO HAVE CAUSED AN UNSTABLE ENVIRONMENT ON THE CLIENT. THE CALCULATION OF THE SCOUT LINES POSITIONING TAKES PLACE ON THE CLIENT AND IT IS DEPENDENT ON THE '(B)(6)' FRAMEWORK INSTALLED ON THE CLIENT LOCALLY. WHEN EXPORTING THE MRI STUDY TO A CD, THE SCOUT LINE ISSUE WAS NOT PRESENT. THE CALCULATION USES A DIFFERENT MECHANISM WHICH USES LEADTOOLS. THIS IS UNAFFECTED BY (B)(6) OR THE LOCAL CLIENT ENVIRONMENT AND ARE INSTEAD PULLED FROM THE SERVER AND PLACED ON THE CD VIEWER WHICH HANDLES THE CALCULATION CORRECTLY. THIS INDICATES AND IS CONSISTENT WITH THE FACT THAT THE STUDY ITSELF WAS NOT FAILING AND THAT THE CLIENT'S ENVIRONMENT WAS THE MAIN CONTRIBUTING CAUSE TO THE ERROR IN THE SCOUT LINES CALCULATION. THE ISSUE ON THE ORIGINAL CLIENT APPEARS TO HAVE SUBSEQUENTLY BEEN PROPAGATED TO OTHER CLIENTS. THE CLINIC'S PERSONNEL UNKNOWINGLY REPLICATED THE CORRUPTED ENVIRONMENT BY TAKING AN 'IMAGE FILE' OF THE ORIGINAL CLIENT AND REPRODUCING IT TO THE REST OF THE CLIENTS BY IMAGING THEM WITH THE IMAGE OF THE CORRUPTED CLIENT. IT IS NOT CLEAR AT THIS POINT EXACTLY WHEN THE ISSUES WAS FIRST KNOWN. ACCORDING TO THE CLINICAL ADMINISTRATOR, THE DOCTOR KNEW OF THE ISSUE AND FACTORED IN THE ERROR OF THE SCOUT LINES WHEN BEFORE ANY PROCEDURES. HE FURTHER EXPRESSED THE CONCERN THAT DEPENDING ON THE MARGIN OF ERROR OF THE SCOUT LINES CALCULATION, IT IS POSSIBLE THAT A PHYSICIAN WILL NOT REALIZE THAT ERROR AND MAY POTENTIALLY OPERATE ON THE WRONG ANATOMICAL REGION. THE CLINIC HAS NOT REPORTED ANY PT HARM TO DATE, HOWEVER, THEY EXPRESSED THE CONCERN OF HAZARD AND POTENTIAL HARM. AS OF 4/30/2010, THE ISSUE HAS BEEN RESOLVED BY REINSTALLING THE (B)(6) FRAMEWORK ON THE CLIENT, AS WELL AS THE OFFICEPACS POWER CLIENT SOFTWARE AN ALL OF THEIR CLIENTS/WORKSTATIONS. THE ISSUE HAS NOT BEEN REPORTED AT OTHER SITES, AND IT IS NOT REPRODUCIBLE IN HOUSE. | 1
1486356 | THE USER REPORTED THAT MONITOR HAD NO AUDIBLE SOUND. >< METHOD AND RESULTS: THERE WAS NO ADVERSE EVENT AND THERE WAS NO PT INVOLVED. THE USER CLAIMED THAT THEY REPLACED THE DEVICE'S SPEAKER, BUT THIS DID NOT RECTIFY THE PROBLEM. THE DEVICE MFR ADVISED THE USER THAT THE NO SOUND CONDITION WAS LIKELY BEING CAUSED BY THE DEVICE'S MOTHERBOARD. THE USER WAS ALSO ADVISED THAT THE DEVICE NEEDED TO RETURN TO THE MFR FOR FACTORY REPAIR BECAUSE THE MOTHERBOARD REQUIRES PROGRAMMING AND CANNOT BE REPLACED IN THE FILED. THE USER REPORTED TO THE DEVICE MFR THAT THEY HAVE REMOVED THE DEVICE FROM SERVICE AND THAT THE DEVICE WILL REMAIN OUT OF SERVICE UNTIL IT IS REPAIRED, TESTED, AND DEEMED FULLY OPERATIONAL. AT THIS TIME, NO ADDITIONAL INVESTIGATION OR ACTION IS WARRANTED OR POSSIBLE. | 1
1051473 | IT WAS REPORTED THAT NIBP 'NO DETERMINATION' OCCURRED IN THE OPERATING THEATER DURING A CASE WITH SOLAR AND TRAM MODULE BEING USED IN CONJUNCTION WITH AVANCE ANESTHESIA MACHINE. THE USER WAS ALLEGEDLY UNABLE TO MEASURE NON-INVASIVE BLOOD PRESSURE REQUIRING PROLONGED ANESTHESIA AND SURGERY WITH DISATTACHMENT AND REATTACHMENT OF ROBOT REQUIRED. THE HOSPITAL HAS NOT PROVIDED INFO THAT A PT INJURY DID OCCUR. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1763127 | ONGOING ISSUES WITH ARTIFACT ON THE MRI IMAGES THAT APPEAR TO CORRESPOND WITH USE OF MEDRAD VERIS MRI MONITORING SYSTEM. A 'LINE' OR 'GRAINY' APPEARANCE HAS BEEN NOTED ON THE MRI IMAGES WHEN THE MONITORS ARE IN USE IN THE ROOM ONLY. CONCERNS FOR PATIENT SAFETY DUE TO ADDITIONAL SCANNING CAUSING DELAYS IN PATIENT CARE AND DECREASED IMAGE QUALITY.HEALTH PROFESSIONAL'S IMPRESSION: MULTIPLE TESTS HAVE BEEN DONE PER CLINICAL ENGINEERING THAT SEEM TO POINT TO MONITOR INTERFERENCE. THERE IS NO ARTIFACT IN EITHER MRI ROOMS UNLESS THE MEDRAD MONITOR IS IN USE IN THE ROOM. | 1
1887356 | SMARTMONITOR 2 WAS BEING USED DURING THE TRANSPORT OF AN INFANT WITH CENTRAL APNEA. DEVICE TURNED OFF WITHOUT NOTICE. THE DEVICE WAS FULLY CHARGED AT THE TIME OF THE FAILURE. THE DEVICE THEN TURNED BACK ON WITHOUT ANY ACTION BY THE USER. CHILD HAD AN APNEA EVENT WHILE THE DEVICE WAS NON-FUNCTIONAL. LOOKED UP THE SERIAL NUMBER AND FOUND THAT IT WAS IN THE GROUP OF MONITORS RECALLED BY RESPIRONICS. TOOK THE MONITOR BACK TO (B)(6) HEALTHCARE. RESPIRONICS AND (B)(6) CLAIM THAT THE DEVICE HAD BEEN RETURNED TO RESPIRONICS AFTER THE RECALL AND WAS 'REPAIRED'. AFTER THE 'REPAIR', THE RECALLED DEVICE WAS RETURNED TO SERVICE THROUGH (B)(6) HEALTHCARE. (B)(6) CLAIMS THAT THEY ARE GOING TO HAVE RESPIRONICS REPAIR THE DEVICE AGAIN. DATES OF USE: (B)(6) 2010. DIAGNOSIS OR REASON FOR USE: INFANT APNEA - CENTRAL-. | 1
1783662 | THE APNEA MONITOR WAS RETURNED TO THE MFR FOR EVAL AT THE REQUEST OF THE CUSTOMER. THE APNEA MONITOR WAS VISUALLY EXAMINED AND TESTED BY THE USING A SIMULATOR IN ACCORDANCE WITH THE SMART MONITOR 2 CHECKOUT PROCEDURE MANUAL ((B)(4)). THE APNEA MONITOR DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS AND PASSED ALL REQUIRED TESTING. THERE WERE NO OPERATIONAL ISSUES FOUND. THE APNEA MONITOR'S MEMORY DATA WAS DOWNLOADED AND ANALYZED BY TRAINED, CLINICAL ASSOCIATES. BASED ON THE DOWNLOADED MEMORY, THE MFR DETERMINED THAT THE APNEA MONITOR WAS IN USE IN THE HOME FROM (B)(6), 2010 TO (B)(6), 2010. THERE WERE FIFTY-ONE (51) PT EVENTS RECORDED DURING THAT TIME PERIOD, WITH THE LAST PT EVENT RECORDED ON (B)(6), 2010 AT 2:39:52 PM. ALL RECORDED PT EVENTS WERE ASSOCIATED WITH AUDIBLE AND VISUAL ALARMS. THE EQUIPMENT EVENTS SHOWED THE APNEA MONITOR WAS TURNED OFF ON (B)(6), 2010 AT 2:42:34 PM AND NOT TURNED ON AGAIN UNTIL (B)(6), 2010 AT 2:41:05 PM, WHEN THE MEMORY DATA WAS DOWNLOADED BY AN UNK HEALTHCARE PROFESSIONAL. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION ASSOCIATED WITH THE ADVERSE EVENT, AND NO EVIDENCE TO SUGGEST THE APNEA MONITOR IN ANY WAY CAUSED OR CONTRIBUTED TO THE EVENT. THE MONITOR PASSED ALL REQUIRED TESTING AND DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS. BASED ON ALL AVAILABLE INFO, THE MFR CONCLUDES THAT THE DEVICE DOES FUNCTION TO SPEC AND THAT NO FURTHER ACTION IS APPROPRIATE. >< A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER NOTIFIED THE MFR OF THE DEATH OF AN INFANT THAT OCCURRED ON (B)(6), 2010. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE DME WAS NOTIFIED BY THE POLICE DEPARTMENT OF THE DEATH ON (B)(6), 2010. THE INFANT EXPERIENCED AN ADVERSE OR LIFE THREATENING EVENT (ALTE) ON (B)(6), 2010, WAS TRANSPORTED TO A HOSPITAL, AND EXPIRED THE NEXT DAY WHILE ON LIFE SUPPORT. THE DME STATED THAT A CHECKOUT PROCEDURE WAS PERFORMED PRIOR TO THE MONITOR BEING PLACED INTO USE ON THE INFANT AND THE DEVICE PASSED ALL REQUIRED TESTING AT THAT TIME. | 0
1571438 | A DEFECT WAS DISCOVERED IN THE DEVICE THAT CONVERTS INPATIENT MEDICATIONS TO PRESCRIPTIONS. IT HAS TAKEN MONTHS TO RECOGNIZE AND CORRECT THIS POTENTIALLY HIGH RISK DEFECT. WHEN INPATIENT COMBINATION MEDICATION IS CONVERTED TO PRESCRIPTION, THE DEVICE DELETES THE STRENGTH OF THE MEDICATION. THE NUMBER OF PATIENTS AFFECTED IS UNKNOWN, INVOLVING CASES RECEIVING COMBINATION PILLS. | 1
1788175 | (B)(4)AS THE DATE OF ONSET OF THIS PERITONITIS EPISODE IS UNKNOWN, AND PATIENT DISCARDED SUPPLIES AFTER EACH THERAPY, THE SAMPLE WAS NOT REQUESTED. >< THIS IS A SPONTANEOUS CONSUMER REPORT BY A PATIENT'S FATHER FROM (B)(6) OF PERITONITIS IN A MALE PATIENT COINCIDENT WITH DIANEAL UNKNOWN THERAPY. ON AN UNKNOWN DATE, THE PATIENT BEGAN DIANEAL THERAPY. ON (B)(6)2010, THE PATIENT'S FATHER CONTACTED BAXTER (B)(4) TECHNICAL SERVICE CENTER AND STATED THAT THE PATIENT HAD BEEN HOSPITALIZED DUE TO PERITONITIS. [FOLLOW UP INFORMATION (20JUL2010) WAS GIVEN BY A PHYSICIAN. ON (B)(6)2010, HE DEVELOPED PERITONITIS BACTERIAL MANIFESTED BY PERITONEAL CLOUDY EFFLUENT, ABDOMINAL PAIN AND PYREXIA. ON (B)(6)2010, HE WAS HOSPITALIZED DUE TO THE EVENT. PERITONEAL EFFLUENT CULTURE WAS PERFORMED AND THE RESULT REVEALED (B)(6). VANCOMYCIN AND MODACIN WERE ADMINISTERED FOR THE EVENT. ON (B)(6)2010, PERITONEAL EFFLUENT WAS ALMOST CLEAR. AT THE TIME OF THIS REPORT, THE EVENT WAS RESOLVING AND PD THERAPY WAS CONTINUED UNCHANGED. THE PHYSICIAN BELIEVED THAT NONE OF THE EVENTS WERE RELATED TO PD THERAPY. PREVIOUS PERITONITIS WAS ALSO CAUSED BY AN EXIT SITE INFECTION. AS THE PATIENT WAS AT END STAGE, NO AGGRESSIVE MEDICAL TREATMENT HAD BEEN PERFORMED EXCEPT FOR THE PERITONITIS. | 0
1678405 | CUSTOMER REPORTED SPARKS COMING FROM THE PLUG ON THE LDX PRINTER. PRINTER STILL WORKS ON ALTERNATE POWER CORD. >< INVESTIGATION PENDING. | 1
1034380 | DR. (B) (6) AND (B) (6) REPORTED THAT THE (B) (4) TOOL REPORTS DIFFERENT VALUES (MAX SUV) FOR A SMALL TUMOR DEPENDING ON WHETHER THE PET IS OPENED ALONE OR WHETHER THE PET IS PARTICIPATING IN A FUSION WITH A CT. | 1
1630376 | DURING EVALUATION OF A RETURNED HOMECHOICE MACHINE, AN INCREASED INTRAPERITONEAL VOLUME (IIPV) SITUATION WAS IDENTIFIED WHICH OCCURRED ON (B)(6) 2010 DURING DRAIN CYCLE 3. THE PATIENTS FILL VOLUME WAS 2500 ML. THE DRAIN VOLUME WAS 4648 ML. WHICH MEETS IIPV CRITERIA. NO PATIENT INJURY OR MEDICAL INTERVENTION WAS REPORTED. ACCORDING TO THE NURSE THE HOME PATIENT (HP) EXPIRED ON (B)(6) 2010 DUE TO CARDIAC FAILURE AND STROKE. PER NURSE, THE DEATH WAS NOT RELATED TO THE BAXTER'S HC CYCLER. PERITONEAL DIALYSIS (PD) THERAPY WAS STOPPED SEVERAL WEEKS PRIOR TO THE PATIENT'S DEATH. >< (B)(4). PRODUCT ANALYSIS LAB (PAL) EVALUATED THE DEVICE AND NO FAILURE OR MALFUNCTION WAS IDENTIFIED THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED DIFFICULTY. ALL TESTING PERFORMED DURING THE HOMECHOICE RITE FUNCTIONAL TEST AND HOMECHOICE RITE ELECTRICAL SAFETY ANALYZER TEST PASSED. PER (B)(6), THE HOME PATIENT PASSED AWAY ON (B)(6) 2010. IIPVS PER THE LOGS ONLY: OCCURRENCE DATE: (B)(6) 2010 / CYCLE 3 / DRAIN VOLUME 4648 ML. PROBABLE CAUSES: INSUFFICIENT DRAIN. FALSE EMPTY DETECT AND USE ERROR. CLINICIAN INAPPROPRIATELY SET MINIMUM DRAIN VOLUME % SETTING TO LOW AND INSUFFICIENT DRAIN, FALSE EMPTY DETECT AND USE ERROR. INITIAL DRAIN ALARM SETTING INAPPROPRIATELY PROGRAMMED. THE DEVICE WAS ROUTED TO THE SERVICE AREA. (B)(4) IS CURRENTLY OPEN FOR THE REPORTABLE MALFUNCTION OF IIPV/OVER DELIVERY. | 0
1480688 | NO REPORT OF PT HARM OR INJURY RELATED TO THIS EVENT. CUSTOMER REPORTED THAT THE EXAM DATE/TIME ON THE PRINTED REPORT IS DIFFERENT BY TWO HRS FROM THE ACTUAL EXAM DATE/TIME. >< ISITE PACS IS A SOFTWARE PRODUCT. WE ARE ABLE TO EVALUATE THE PRODUCT AT THE CUSTOMER SITE BY REMOTE ACCESS, AS WELL AS, EVALUATE THE SAME PRODUCT VERSION IN-HOUSE. THE ISITE PACS SOFTWARE DISPLAYS CLINICAL REPORTS WITH THE STUDY PERFORMED DATE AND TIME IN GMT FORMAT, WHICH MAY NOT BE OBVIOUS TO THE END USER, WHO MAY EXPECT TO SEE THE LOCAL TIME IN THE REPORT. THIS MAY RESULT IN DISPLAYING A STUDY DATE AND TIME UP TO 12 HRS FORWARDS OR BACKWARDS IN TIME, AND IN SOME CASES MAY DISPLAY A SHIFT IN DATE UP TO ONE DAY. FURTHERMORE, IF NO IMAGES ARE ASSOCIATED TO THE REPORT, NO STUDY PERFORMED DATE/TIME STAMP WILL BE DISPLAYED. FIELD ACTION TAKEN FOR THIS ISSUE. (B)(6), RECALL AND EMERGENCY COORDINATOR, US FOOD AND DRUG ADMIN, (B)(6). CUSTOMERS WERE SENT A FIELD SAFETY NOTICE (COPY WAS PROVIDED IN ORIGINAL 806 REPORT). (B)(4). WHILE THE LOCAL EXAM DATE/TIME IS DISPLAYED IN ALL WORKLISTS AND ON THE IMAGES, THE DATE WHEN DISPLAYED IN THE CLINICAL REPORT IS PRESENTED IN GMT. (B)(4). | 1
1420080 | IN A CENTRICITY RIS IC/CENTRICITY PACS INTEGRATED SYSTEM, WHEN THE CUSTOMER IS OPERATING THEIR SYSTEM IN PACS MODE, THE OPPORTUNITY EXISTS TO GET THE TWO SYSTEM OUT OF SYNCH DISPLAYING TWO DIFFERENT PTS. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. >< A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. (B) (4). | 1
1836203 | THE PROVUE MISREAD A WEAK 1+ REACTION AS NEGATIVE IN THE TYPE AND SCREEN TESTING USING THE ABD/REV CARD FOR ABO TESTING. NO ERRONEOUS RESULTS WERE REPORTED. IF A SIGNIFICANT ANTIBODY IS NOT DETECTED DURING ANTIBODY SCREENING, OR IS NOT IDENTIFIED UPON FURTHER TESTING. THE PATIENT MAY BE TRANSFUSED WITH ANTIGEN-POSITIVE BLOOD AND SUFFER A HEMOLYTIC TRANSFUSION REACTION. >< AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND PERFORMED PREVENTATIVE MAINTENANCE PROCEDURE WHICH INCLUDES CLEANING, ADJUSTING, AND CREATING A NEW REFERENCE IMAGE FOR THE GEL CAMERA. THE PROVUE SUCCESSFULLY COMPLETED ALL ADJUSTMENTS AND DIAGNOSTIC TESTS. QC RAN WITHOUT ERROR AND WITHIN SPECIFICATION. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4) | 0
1282025 | UPON THE INITIAL CONTACT TO BAXA, THE CUSTOMER STATED THAT THEY WERE NOT AUTHORIZED TO DISCUSS THE DETAILS OF THIS INCIDENT, BUT DID STATE THAT IT WAS 'USER ERROR DURING ORDER ENTRY.' THE CUSTOMER WAS ABLE TO PROVIDE THEIR ABACUS DATABASE AND COMPOUNDER DATABASE FOR EVALUATION. AN EXAMINATION OF THE ABACUS DATABASE REVEALED THAT THE ORDER WAS ENTERED ON (B)(6) 2008. THE ORDER CONTAINED THE FOLLOWING INGREDIENTS: 40 MEQ/DL NA CHLORIDE, 40 MEQ/DL NA ACETATE, 30 MEQ/DL K CHLORIDE. THE TOTAL VOLUME OF THE BAG WAS 1000 ML. THE PATIENT'S PREVIOUS DAY'S ORDER HAD A UNIT VALUE OF MEQ/L RATHER THAN THE SUBJECT'S ORDER UNIT VALUE OF MEQ/DL. THEREFORE, BASED ON THE INFORMATION PROVIDED FOR ANALYSIS, IT CAN BE CONCLUDED THAT THE LIKELY CAUSE OF THE OVER-DELIVERY WAS THE SELECTION OF MEQ/DL, RATHER THAN MEQ/L BY THE USER COMPLETING ORDER ENTRY. THE ABACUS DATABASE SHOWED THAT THE SOFTWARE DISPLAYED WARNING LIMITS FOR ACETATE, CHLORIDE AND POTASSIUM. RATIONALE WAS GIVEN TO ALL THREE WARNING LIMITS AS 'OK' AND THE ORDER WAS COMPLETED. A SITE VISIT WAS CONDUCTED ON (B)(6) 2008, TO PROVIDE TECHNICAL ASSISTANCE. DURING THE SITE VISIT, EXTENSIVE ABACUS ORDER ENTRY TRAINING WAS CONDUCTED. ADDITIONAL HARD STOP WARNING LIMITS WERE IMPLEMENTED FOR SAFEGUARD, AND A WARNING LIMIT OVERRIDE SECTION WAS ADDED TO THEIR EXISTING ABACUS FORMULA LABEL. IF BAXA RECEIVES ADDITIONAL INFORMATION IN REGARDS TO THIS INCIDENT, A FOLLOW-UP MDR WILL BE SUBMITTED. >< ON (B)(6) 2008, BAXA WAS NOTIFIED BY THE CUSTOMER OF A PATIENT INVOLVED INCIDENT USING THE ABACUS V2.0 TPN CALCULATING SOFTWARE AND THE EXACTA-MIX 2400 COMPOUNDER FOR TPN PRODUCTION. IT WAS REPORTED THAT A PATIENT HAD RECEIVED A TPN BAG CONTAINING A HIGHER-THAN-ANTICIPATED VOLUME OF SODIUM CHLORIDE, SODIUM ACETATE AND POTASSIUM CHLORIDE. THREE ATTEMPTS HAVE BEEN MADE TO OBTAIN FURTHER INFORMATION FROM THE CUSTOMER ON THIS INCIDENT, HOWEVER, THE CUSTOMER HAS ELECTED TO NOT SHARE ANY FURTHER DETAILS. THE CURRENT CONDITION OF THE PATIENT IS UNKNOWN. | 1
1582401 | THIS ACTUAL BACK UP PRODUCT (SOFTWARE) DOES NOT HAVE AN ASSIGNED SERIAL NUMBER. THE DEVICE IS SOFTWARE AND WAS EVALUATED REMOTELY, BUT NOT RETURNED TO THE MFR. THE DEVICE WAS EVALUATED ON 12/14/2009 REMOTELY. BASED ON THE INVESTIGATION, TESTS WERE RUN TO SUBSTANTIATE PERFORMANCE ISSUES RELATED TO THE GUARDIAN BACKUP SERVICE RUNNING DURING CLINIC HOURS BECAUSE DATA BACKUP WAS NOT COMPLETED IN THE ALLOTTED TIMEFRAME. THE LIKELY CAUSE OF THE MALFUNCTION IS THAT THERE WAS TOO MUCH DATA FOR THE SYSTEM TO BACKUP. THERE IS A POTENTIAL THAT SOME DATA WAS NOT COMPLETELY BACKED UP. FURTHER INVESTIGATION IS ONGOING CAPA (B) (4) ADDRESSES THIS ISSUE. THIS IS NOT A SINGLE USE DEVICE. >< GUARDIAN BACKUPS RUNNING DURING CLINIC HOURS. INITIAL REPORTER CALLED AND REPORTED THAT HER INTERNET CONNECTION WAS RUNNING VERY SLOWLY. AFTER FURTHER INVESTIGATION, IT WAS DETERMINED THAT THE DATA BACK UP WAS RUNNING DURING CLINIC HOURS. | 1
1584674 | USER EXPERIENCED ISSUES WITH CONTROL AND PT RESULTS FOR URINE LEUKOCYTES. ONE PT URINE SAMPLE GENERATED DISCREPANT RESULTS DURING THIS TIME. THE SAMPLE WAS INITIALLY REPORTED OUT AS NEGATIVE. THE SAME SAMPLE WAS RETESTED MANUALLY USING A CHEMISTRIP 10 UA STRIP, GIVING A RESULT BETWEEN 1 TO 2 PLUS VISUAL READING FOR LEUKOCYTES AND WAS REPORTED AS 100 L PER UL TO THE DR. A MICROSCOPIC EXAMINATION OF THE SAME URINE SAMPLE SHOWED 11 TO 20 WBC PER HPF. A CORRECTED REPORT WAS ISSUED. THE USER STATED THE DR DID NOT INITIATE TREATMENT TO HER KNOWLEDGE. TO HER KNOWLEDGE, THE PT CONDITION REMAINS UNCHANGED FROM THE STATE SHE WAS IN PRIOR TO REPORTING THE RESULTS TO THE PHYSICIAN. THE FIELD SERVICE REP FOUND THE SAMPLE PROBE WAS OUT OF ALIGNMENT AND CLEANED, ALIGNED SAMPLE PROBE AND THE SAMPLE STRIP BASE. TESTS AND CONTROLS WERE RUN WHICH WERE ALL OKAY TO VERIFY THE ANALYZER PERFORMANCE. | 0
1573664 | (B)(4). >< BS ECG SIGNALS DISAPPEARED IN THE MIDDLE OF A CAE. ADVISED THAT THE ECG PATCHES & CABLE WERE CHANGED WITH NO EFFECT, AND SIGNALS ONLY RETURNED WHEN POWER TURNED OFF TO COM UNIT. THE PIU LED SHOWED IT WAS WORKING FINE AND NO ERRORS DISPLAYED ON THE WS. THIS IS NOT INDICATIVE OF A REPORTABLE EVENT. THE CASE WAS ABORTED AND RESCHEDULED TO A LATER DATE. THERE WAS NO FULL ANESTHESIA REQUIRED. THE PATIENT SUFFERED PSYCHOLOGICAL DAMAGE. | 1
1345276 | IT WAS REPORTED THAT A MAP SHIFT OF 10CM WAS NOTED AFTER THE MAP WAS CREATED, WHILE TRYING TO REACH THE POINT FOR ABLATION. NO PATIENT INJURY WAS REPORTED. >< (B) (4). | 1
1799366 | PATIENT WITH 1MM SKIN ABRASION TO THE RIGHT AUXILIARY ARM. THIS WAS THE LOCATION OF THE TEMPERATURE PROBE. ALL PACKAGING WAS SECURED AND LOT NUMBER.MANUFACTURER RESPONSE FOR PHYSIOLOGICAL MONITOR, TEMPERATURE MODULE, M1029A: VOICE MAIL TO PHILIPS. WILL PICK-UP THE UNIT NEXT DAY FOR EVALUATION.MANUFACTURER RESPONSE FOR TEMPERATURE PROBE, PATIENT, ER400-9 9FR: PHONE CALL TO THE PATIENT MONITORING ACCOUNT MANAGER FROM PHILIPS AND A PHONE CALL TO SMITHS MEDICAL. | 1
1802943 | IN THE MORNING, THE BIOMED RESPONDED TO A PROBLEM IN THE OR. AESTIVA VENTILATOR STOPPED. DR. WAS MANUALLY BAGGING PATIENT. AESTIVA VENTILATOR HAD A WARNING MESSAGE, PROMPTING USER TO PUSH KNOB TO ACKNOWLEDGE VENT HAD EXPERIENCED A PROBLEM. DOING THIS CLEARED THE MESSAGE AND ALLOWED THE USER TO RETURN TO VENTILATOR MODE, HOWEVER, THE PROBLEM REAPPEARED AFTER A FEW MINUTES.DURING THIS TIME, SEVERAL OTHER VENTILATORS, SOME IN OTHER BUILDINGS, REPORTED THE SAME PROBLEM. IT WAS QUICKLY LEARNED THAT PICIS SYSTEM (EMR ELECTRONIC MEDICAL RECORD INTERFACES WITH AESTIVA AND OTHER PATIENT MONITORING EQUIPMENT) HAD A SOFTWARE UPDATE (PUSH) EARLIER IN THE MORNING. IT WAS DETERMINED AND VERIFIED THAT PICIS WAS THE SOURCE OF THE PROBLEM THAT AFFECTED ALL OF THE VENTS IN VENTILATOR MODE. PICIS WAS DISCONNECTED FROM THE AESTIVA AND PROBLEM WENT AWAY. SPOKE WITH GE TECH SUPPORT. WENT INTO THE SERVICE ERROR LOG ON THE AESTIVA. ALL VENTILATORS HAD A 'WATCHDOG ERROR' IN THE LOG. INVESTIGATION REVEALED THAT THE PICIS UPDATE THAT WAS 'PUSHED' CONTAINED A FILE THAT SHOULD NOT HAVE BEEN PUSHED OUT, CAUSING THIS PROBLEM. IT WAS NOT DETECTED BY PICIS OR HIS HEALTH INFORMATION SYSTEM PERSONNEL PRIOR TO THE PUSH. NO PATIENT HARM EXPERIENCED DUE TO THIS EVENT, HOWEVER THERE WAS A POTENTIAL FOR A MASS FAILURE OF ALL VENTILATORS CONNECTED TO THE PICIS SYSTEM.======================HEALTH PROFESSIONAL'S IMPRESSION======================THE PICIS SYSTEM SOFTWARE ADVERSELY AFFECTED ALL GE AESTIVA VENTILATORS THAT WERE PUT INTO THE VENTILATOR MODE. AESTIVAS NOT IN VENT MODE (BAG MODE) WERE NOT EFFECTED UNLESS THE MODE WAS SWITCHED TO VENT. THE PICIS SOFTWARE PATCH ADVERSELY AFFECTED 5 AESTIVAS.A FILE WAS PUSHED OUT THAT WAS KNOWN TO BE INCOMPATIBLE WITH AESTIVA, BUT WAS INADVERTENTLY LEFT IN THE PATCH. IT SHOULD HAVE BEEN REMOVED PRIOR TO THE PUSH. THE PUSH WENT OUT TO ALL CONNECTED PICIS HARDWARE, INCLUDING BOXES MOUNTED ON THE ANESTHESIA MACHINE THAT CONNECTS TO THE DATA PORTS OF THE AESTIVA.IT IS BELIEVED THAT THIS FILE UPLOADED TO THE PICIS BOX CAUSED A WATCHDOG ERROR (CPU TIMING CONFLICT) IN THE AESTIVA. THE TIMING INCOMPATIBILITY LOADED UP THE AESTIVA'S CPU BUFFER AND CREATING AN ERROR REALIZED BY THE WATCHDOG SEQUENCE WHICH LED TO THE EVENTS DESCRIBED ABOVE. THE USER CAN CLEAR THE ERROR AND RETURN TO VENTILATOR MODE.AFTER THE PICIS COMPONENTS WERE REMOVED, THE VENTILATOR WERE TESTED AND FOUND TO BE OPERATING CORRECTLY. HIS (IT DEPT) REMOVED THE BAD FILE FROM PATCH AND RETESTED TO ENSURE THERE WERE NO COMPATIBILITY ISSUES. THERE HAVE BEEN NO PROBLEMS SINCE.======================MANUFACTURER RESPONSE FOR ANESTHESIA VENTILATOR AND EMR INTERFACE DEVICE, (BRAND NOT PROVIDED)======================GE RESPONSE: INITIALLY, BASED ON THE ERROR CODE, GE SAID TO REPLACE THE CPU. BUT UPON FURTHER RESEARCH BY THEIR TECH SUPPORT DEPT, IT WAS DECIDED THAT THEY COULD PERFORM INTERNAL CPU DIAGNOSTIC TESTS TO ENSURE THE CPU WAS OPERATING NORMALLY.PICIS RESPONSE: THEY CONFIRMED THEY KNEW THIS FILE SHOULD NOT HAVE BEEN INCLUDED IN THE PATCH. ATTRIBUTED TO HUMAN ERROR. | 1
1843551 | URINE CULTURE AND SENSITIVITY WAS ORDERED AND PERFORMED ON A PATIENT ILL WITH MULTIPLE CHRONIC ILLNESSES. THERE WAS NOT ANY RESULT ON THE EMR NOR INDICATION, IT WAS PENDING FOR TWO DAYS. ASSOCIATED WITH THE HOSPITAL'S AGGRESSIVE BED TURNOVER STRATEGIES, THE PATIENT WAS DISCHARGED. THE CULTURE RESULT CONSISTENT WITH A URINARY TRACT INFECTION WAS DEPOSITED ELECTRONICALLY AND SILENTLY IN THE EMR AFTER THE PATIENT WAS DISCHARGED. IT WENT UNNOTICED UNTIL THE PATIENT PRESENTED TO THE EMERGENCY WARD WITH A CLINICAL SYNDROME OF SEPSIS, REQUIRING HOSPITALIZATION. WE SEE THESE ISSUES COMMONLY WITH THE SILENT DEPOSITION OF DIGITIZED RESULTS IN EMRS. THIS FACILITATED NEGLECT IS THE GENESIS OF MALPRACTICE LAWSUITS. WE SEE FAILED RECONCILIATION BETWEEN ORDERS AND THE EXECUTION OF SAID ORDERS. WE SEE FAILED COMMUNICATION WITH THE PHYSICIANS OF THE STATUS OF THE ORDERS, IE WERE THEY CARRIED OUT, DID THE SPECIMEN ARRIVE AT THE LAB'? ETC. FOR CERTAIN TESTS, WE SEE THAT THE SOLE METHOD OF FINDING OUT THE STATUS OF SUCH TEST IS TO CALL THE LAB TO WHERE THE SPECIMEN WOULD HAVE BEEN SENT. EMRS ARE ILL DESIGNED FOR SUCH SAFETY AND EFFICACY. | 1
1749786 | THIS EVENT INVOLVES A (B)(6) MALE PT FOUND UNRESPONSIVE AND APNEIC LYING IN HIS HOSPITAL BED DESPITE CARDIAC MONITORING AND OXYGEN SATURATION MONITORING. HIS BEDSIDE MONITOR DID NOT ANNOUNCE/ALARM FOR A LETHAL ARRHYTHMIA - THE ARRHYTHMIA ALARMS WERE FOUND TO BE OFF AND THE BEDSIDE CARDIAC MONITOR ALARM AUDIO/VOLUME WAS FOUND TO BE OFF. THE PT WAS BEING MONITORED BY BEDSIDE CARDIAC MONITOR, WHICH WAS KNOWN TO BE OPERATIVE ON THE EVENING OF (B)(6) WHEN THE PT WAS ADMITTED TO THE UNIT. AT ROUGHLY 10:15 AM ON (B)(6) 2010, THE PT WAS FOUND PULSE LESS AND APNEIC IN HIS HOSP BED AND A CODE BLUE WAS CALLED. FOLLOWING ACLS PROTOCOL THE MEDICAL SENIOR RESIDENT RAN THE CODE, WHICH WAS UNSUCCESSFUL. IMMEDIATELY FOLLOWING THE CODE, NURSING STAFF AND CODE TEAM MEMBERS DISCOVERED THAT THE ARRHYTHMIA ALARMS WERE OFF FOR THIS PT. A BIOMEDICAL ENGINEERING MEMBER WAS CONSULTED IMMEDIATELY AND BEGAN INVESTIGATING THE CARDIAC AND CENTRAL MONITORS FOR INFO THAT DAY. A FULL INVESTIGATION WAS INITIATED FROM THE CENTER FOR QUALITY AND SAFETY THE FOLLOWING MORNING ON (B)(6) 2010. THE INTERNAL MONITORING LOG FILES WERE SENT TO THE MFR FOR REVIEW ON 01/26/2010 AND 01/27/2010. OUR INVESTIGATION REVEALED TWO DEVICE ISSUES THAT CONTRIBUTE TO THIS ADVERSE EVENT. FIRST, THE DEVICE HAD AN INTERNAL DESIGN SETTING THAT ALLOWED A USER TO DISABLE THE CENTRAL ALARM FROM THE BEDSIDE MONITOR WITHOUT WARNING. SECOND, THE DEVICE IS NOT DESIGNED TO BE INTERROGATED THOROUGHLY ENOUGH TO DETERMINE WHEN THE ALARM WAS TURNED OFF. | 1
1614302 | DEVICE EVALUATION: CUSTOMER REPORT WAS NOT CONFIRMED. FLOW RATE THROUGH FLOW RESTRICTOR WAS MEASURED 3.3 ML/HOUR. SPEC. IS 3-4 ML/HOUR PER DFU. NO VISIBLE DEFECT WAS NOTED FROM FLUSH DEVICE. >< THE FOLLOWING WAS REPORTED: 'THE FLOW RATE WAS ABNORMALLY HIGH AFTER PRIMING WHEN THE LINE WAS KEPT AT 300 MMHG PRESSURE DURING SET UP. EVENT DATE IS UNKNOWN. ' | 0
1624148 | IT WAS REPORTED BY THE INITIAL CALLER THAT SOME PTS HAD TO BE REIMAGED (X-RAYED) DUE TO A SOFTWARE SYSTEM FAILURE. >< ACCORDING TO THE INITIAL REPORTER, PTS HAD TO BE RE-IMAGED (X-RAYED). THE DEVICE IN QUESTION WAS EVALUATED REMOTELY BY STRYKER PERSONNEL. ON (B) (6) 2010, (B) (6) FROM THE CLINIC CALLED AND STATED THAT MORNING USERS IN THE OPERATING ROOM WERE HAVING PROBLEMS PULLING UP IMAGES FROM THEIR (B) (6) SERVER. HE MENTIONED THAT THOSE IMAGES WERE AVAILABLE ON THE PRODUCTION SERVER, BUT THE IMAGES DID NOT SYNCHRONIZE OVER THEIR BACKUP SERVER. OPERATING ROOMS ARE CONNECTED TO THE BACKUP SERVER TO PULL IMAGES FROM WITHIN THE OR. DUE TO THIS ISSUE, THEY HAD TO RE X-RAY A COUPLE PTS THAT MORNING. THE INITIAL INVESTIGATION WAS PERFORMED BY LOGGING INTO THE MAIN PRODUCTION SERVER REMOTELY WHICH REVEALED THAT THE SYNCHRONIZATION SERVICE HAD STALLED WHICH HAD PREVENTED THE IMAGE DATA FROM REPLICATING ACROSS TO THE BACKUP SERVER. SINCE THE OR WAS POINTED TO THE BACKUP SERVER AND THE SYNCHRONIZATION SERVICE HAD STALLED, THE USER WAS UNABLE TO ACCESS THE IMAGE DATA. ON 02/26/2010, FURTHER INVESTIGATION ON THIS CASE REVEALED THAT THE SITE HAD COMPLAINED ABOUT THIS ISSUE IN THE PAST. IT APPEARS THAT THE ROOT CAUSE OF THE MISSING IMAGES ON THE BACKUP SERVER IS RELATED TO A MIDDLEWARE CONFIGURATION INTERFACE WHICH HAD BEEN SENDING ADD'L INFO TO THE POWER SERVER CONSTRAINING THE (B) (6) DATA BASE. THIS INFO IS NOT USED OR DISPLAYED BY OFFICEPACS AND IS NOT NEEDED FOR THE FUNCTIONALITY OF THE SYSTEM. FURTHERMORE, THIS (B) (4) INFO, SUCH AS PT ADDRESSES HAD CONSTRAINED THE (B) (6) DATABASE TO THE POINT WHERE THE SYSTEM WAS FREEZING; THIS CAUSED THE SYNCHRONIZATION SERVICE TO FAIL. AFTER CONFIGURING THE (B) (4) INTERFACE BY INSTALLING A FILTER TO PREVENT THE INFO THAT IS NOT NEEDED FROM GETTING STORED BY THE (B) (4) DATABASE, NORMAL SYSTEM OPERATION RESUMED AND THE SYSTEM APPEARS TO BE WORKING WELL. PT INFO WAS NOT LOST OR CORRUPTED, BUT WAS NOT AVAILABLE ON THE BACKUP SERVER UNTIL THE SYNCHRONIZATION SERVICE WAS MANUALLY RESTARTED. THE USER REPORTED THAT THE PTS WERE REIMAGED ON (B) (6) 2010 BEFORE THE ISSUE WAS RESOLVED. THE ISSUE WAS RELATED TO AN (B) (4) CONSTRAINT CAUSED BY AN INCORRECT CONFIGURATION OF THE (B) (4) INTERFACE. THIS HAS NOT BEEN SEEN BEFORE. (B) (4) MESSAGES NEEDED TO BE FILTERED DUE TO THEIR LARGE VOLUME WHICH WERE NOT NEEDED OR USED BY OFFICEPACS POWER. | 1
1760079 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A RADIAL JAW 4 SINGLE USE BIOPSY FORCEPS LARGE CAPACITY WAS USED DURING A COLONOSCOPY PROCEDURE PERFORMED ON (B)(6), 2010 ACCORDING TO THE COMPLAINANT, DURING UNPACKING THE DEVICE WAS REMOVED FROM THE PRODUCT PACKAGING AND THE NEEDLE WAS FOUND BENT, INHIBITING THE JAWS/CUPS FROM CLOSING. THE PROCEDURE WAS COMPLETED WITH ANOTHER RADIAL JAW 4 SINGLE USE BIOPSY FORCEPS LARGE CAPACITY DEVICE. THERE WERE NO PATIENT COMPLICATIONS REPORTED AS A RESULT OF THIS EVENT. THE PATIENT'S CONDITION AT THE CONCLUSION OF THE PROCEDURE WAS REPORTED TO BE FINE. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< DURING TROUBLE SHOOTING, THE CELL-DYN ASPIRATION PROBE SLIPPED AND PUNCTURED AN ABBOTT FIELD SERVICE (FSR'S) THUMB WHILE THE FSR WAS TRYING TO TAKE IT OFF . THE FSR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT. THE FSR WASHED THE CUT WITH SOAP AND WENT TO THE EMERGENCY ROOM WHERE HIS BLOOD WAS DRAWN FOR (B) (6) TESTING (NO RESULTS AVAILABLE YET) AND A COMBO-VIRAL TREATMENT WAS STARTED. >< (B)(4). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. A REVIEW OF THE SYSTEM OPERATORS MANUAL INDICATED THAT INSTRUCTIONS ARE IN PLACE TO AVOID CONTACT WITH COMPONENTS SUCH AS PROBE AND VENT HEAD NEEDLE AND HANDLE THEM CAUTIOUSLY IN ORDER TO PREVENT INJURY. THERE WERE ALSO INSTRUCTIONS ON HOW TO REDUCE THE RISK OF SUCH AN INCIDENT. A REVIEW AND A LOOK BACK FOR COMPLAINTS WITH SIMILAR INCIDENTS DID NOT IDENTIFY COMPLAINTS WITH SIMILAR NATURE. BASED ON THE EVALUATION RESULTS, NO MALFUNCTION OR PRODUCT DEFICIENCY WERE IDENTIFIED RELATED TO THIS INCIDENT AND THE INCIDENT IS ADDRESSED IN THE PRODUCT LABELING ALONG WITH THE APPROPRIATE PREVENTIVE ACTIONS. | 0
1886196 | ADDITIONAL INFO FROM THE VOLUNTARY REPORT: CASE REPORT: (B)(6). THE CASE DOES NOT REPRESENT A TYPICAL COMPLAINT. NO CUSTOMER OR RESPONSIBLE CLINICIAN HAS APPROACHED NEOVENTA COMPLAINING OF AN UNEXPECTED BEHAVIOUR OF STAN. THE CASE HAS BECOME KNOWN TO NEOVENTA EMPLOYEES (B)(4). NEOVENTA'S MEDICAL ADVISOR, (B)(4) (NOT EMPLOYED IN NEOVENTA) GOT AWARE OF THE CASE IN (B)(4) 2007 (MORE THAN ONE YEAR AFTER THE INCIDENT) WHEN THE HOSPITAL APPROACHED HIM WITH A REQUEST TO ANALYZE THE RECORDINGS OF THE CASE. TRAINING HAS BEEN PERFORMED FIRST BY THREE USERS FROM THE CLINIC VISITING THE NEOVENTA OFFICE 2003. AFTERWARDS CONTINUOUS INTERNAL TRAINING AT CLINIC HAS BEEN PERFORMED. OUR CONCLUSION IS THAT: THERE WAS NO DEFICIENCY ON THE STAN MONITOR; THE STAN SIMPLIFIED CLINICAL GUIDELINES WERE NOT FOLLOWED, AS THEY RECOMMEND IMMEDIATE DELIVERY IN THE PRESENCE OF A PRETERMINAL TRACE. THE PROBLEM MAY BE POSED IN CLINICAL PRACTICE, WITH THE ASSESSMENT AND CLASSIFICATION OF THE FHR TRACE. THIS REPORT IS CLASSIFIED AS FINAL. INVESTIGATION HAS BEEN COMPLETED AND ROOT CAUSE HAS BEEN DETERMINED, TO THE EXTENT POSSIBLE. >< THE STAN CLINICAL GUIDELINES WERE NOT FOLLOWED IN THE PRESENCE OF AN ABNORMAL CTG. THE TRACING WAS PRETERMINAL AND IMMEDIATE DELIVERY WAS RECOMMENDED PER THE GUIDELINES. THE BABY WAS DELIVERED IN SERIOUS CONDITION. ADDITIONAL INFO: EVENT IN (B)(6). DATE OF INCIDENT: 2006-(B)(6). NEOVENTA GOT AWARE: 2007-(B)(6). SERIOUS INJURY FOR PT (BABY). USER ERROR; THE STAN CLINICAL GUIDELINES WERE NOT FOLLOWED, AS THEY RECOMMEND IMMEDIATE DELIVERY IN THE PRESENCE OF A 'PRETERMINAL' TRACE. ADD'L INFO FROM VOL REPORT: PT MONITORING SYSTEM, CARDIOGRAPHIC. | 1
1942177 | THE PROBLEM WAS WIDESPREAD DURING THE DAY. ABOUT 20 COMPUTERS RUNNING MOSAIQ SOFTWARE BEGAN TO LOCK-UP, ONE-BY-ONE. THE EX MOSAIQ COMPUTER LOCKED, BUT THE IX AND TX WAS WORKING. ELEKTA WAS MADE AWARE AND WAS AWAITING WORD FROM THE HOSPITAL TO SHUT-DOWN AND REBOOT THE SERVER. FINALLY, WHEN THE TX MOSAIQ COMPUTER LOCKED, WE FINISHED THE REMAINING PATIENT ON THE IX, AND THEN WE ORDERED ELEKTA TO REMOTELY RUN DIAGNOSTICS, THEN SHUT-DOWN THE SERVER, AND THEN REBOOT. WE WERE ASSISTED BY THE IS DEPARTMENT. THE PROCESS TOOK ABOUT AN HOUR. EVERYTHING CAME BACK NORMAL. THE SERVER FOR MOSAIQ WAS NOT ABLE TO BE REBOOTED BY OUR NETWORK APPS. WE HAD TO ALLOW IMPAC SUPPORT TO WEB INTO THE SERVER TO GET IT RUNNING AGAIN.THIS CAUSED A 2 HOUR DELAY TO THREE VARIAN LINAC RADIATION TREATMENT DEVICES. | 1
1637327 | RADIOGRAPHS AND PACS COMPUTERIZED LIBRARY OF RADIOGRAPHIC FILES RECORDED AN INCORRECT TIME OF ACQUISITION. MEDICAL RECORD WAS INACCURATE INTERFERING WITH THE TIME LINE OF MEDICAL CARE. THIS ERROR WAS REPLICATED IN MULTIPLE HOSPITALS. IT IS UP TO THE INVESTIGATORS TO DETERMINE HOW MANY HOSPITALS AND PT RECORDS WERE AFFECTED THROUGHOUT THE US. | 1
1721296 | QC WAS RUN BEFORE AND AFTER THE EVENT AND RECOVERED WITHIN SPECIFICATIONS. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED: THE FSE REPLACED THE BLOOD SAMPLE VALVE (BSV) ACTUATOR AND PERFORMED REPRODUCIBILITY AND MODE TO MODE TESTING. THE FSE RAN QC AND VERIFIED THE INSTRUMENT'S PERFORMANCE. HARDWARE ISSUE MAY HAVE CONTRIBUTED TO THIS EVENT. >< A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REGARDING ERRONEOUS LOW HEMOGLOBIN (HGB) RESULTS GENERATED BY THE COULTER LH500 ANALYZER. CUSTOMER INDICATED THE LOW HGB RESULTS WERE GENERATED ON ALL PATIENT SPECIMENS RUN IN SECONDARY MODE. RESULTS OBTAINED FROM THE PRIMARY MODE WERE CONSIDERED CORRECT. ALL SPECIMENS WERE RERUN ON A DIFFERENT INSTRUMENT TO CONFIRM THE PRIMARY MODE HGB RESULTS. PATIENT PRINTOUTS AND RAW DATA WERE REQUESTED BUT NOT PROVIDED. ERRONEOUS RESULTS WERE NOT REPORTED OUT OF THE LAB. NO DEATH OR SERIOUS INJURY, NO CHANGE TO PATIENT TREATMENT IS ASSOCIATED WITH THIS EVENT. | 0
1358431 | THE SPECIMENS WERE COLLECTED IN 3ML AND 4.7ML BD VACUTAINER TUBES. THE SAMPLES WERE LESS THAN 2 HOURS OLD AND STORED AT ROOM TEMPERATURE. QC WAS PERFORMED BEFORE AND AFTER THE EVENT AND RECOVERED WITHIN SPECIFICATIONS. BASED ON AVAILABLE INFORMATION, A FIELD SERVICE ENGINEER (FSE) CONTACTED THE CUSTOMER AND HAD THE CUSTOMER REPLACE. THE WASTE FILTER WHICH WAS ON BACKWARDS, BLEACHED THE BATHS, ADJUSTED THE HGB VACUUM, AND RAN REPRODUCIBILITY WHICH PASSED ALL RESULTS. A CLEAR ROOT CAUSE FOR THIS EVENT HAS NOT BEEN DETERMINED TO DATE. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. >< A PATIENT SAMPLE, DRAWN IN A 3ML TUBE, WAS TESTED FOR HEMOGLOBIN (HGB) AND PLATELETS (PLT) ON THE COULTER, AC'T DIFF 2 ANALYZER AND RESULTS OF 8.7G/DL AND 341X10^3CELLS/'L WERE OBTAINED, RESPECTIVELY.THE RESULTS WERE REPORTED OUT OF THE LAB. THE PHYSICIAN QUESTIONED THE RESULTS AND REQUESTED A REDRAW IN 4.7ML TUBE.THE REDRAWN SAMPLE RECOVERED HIGHER: A) HGB - 11.2G/DL, B) PLT - 477X10^3CELLS/'L.CUSTOMER THEN RE-RUN BOTH SAMPLES FOR HGB AND PLT , AND RESULTS WERE: A) ORIGINAL DRAW SAMPLE - HGB 9.6G/DL, PLT 389 X10^3CELLS/L, B) REDRAW SAMPLE- HGB 10.0G/DL, PLT 436 X10^3CELLS/L.NO DEATH OR SERIOUS INJURY, NO CHANGE TO PATIENT TREATMENT IS ASSOCIATED WITH THIS EVENT. | 0
1863305 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO OBTAINING ERRONEOUSLY HIGH HEMOGLOBIN (HGB) RESULTS OF 21.5 WITHOUT INSTRUMENT GENERATED FLAGS ON A PATIENT SPECIMEN, GENERATED BY THE COULTER ACT DIFF 2 ANALYZER. ERRONEOUS RESULTS WERE REPORTED OUTSIDE THE LABORATORY. THE PHYSICIAN QUESTIONED THE RESULT AND REQUESTED THAT THE PATIENT BE REDRAWN. THE REDRAWN SPECIMEN RECOVERED A LOWER HGB RESULT OF 15.5 AND A CORRECTED REPORT WAS SENT OUT. THE OPERATOR CONTACTED THE BCI HOTLINE AND WAS INSTRUCTED TO RERUN THE ORIGINAL SPECIMEN. THE RERUN OF THE INITIAL RESULT RECOVERED AN HGB RESULT OF 14.9. THE CUSTOMER INDICATED THAT THERE WAS NO CHANGE TO PATIENT TREATMENT, AND NO DEATH OR SERIOUS INJURY WAS ASSOCIATED WITH THIS ISSUE. >< THE PATIENT SPECIMEN WAS COLLECTED WITH A BUTTERFLY COLLECTION DEVICE INTO A 4ML BD VACUTAINER TUBE. THE CUSTOMER IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS (ACCURACY) AND REPRODUCIBILITY (PRECISION). CONTROLS WERE RUN BEFORE THIS INCIDENT AND RECOVERED WITHIN ASSAY RANGE. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND VERIFIED THE INSTRUMENT PERFORMANCE. FSE ALSO PERFORMED REPRODUCIBILITY AND CARRY OVER TESTS. | 0
1523383 | THE SAMPLE WAS A FINGER STICK COLLECTION METHOD. QC IS RUN DAILY AND IT WAS RUN BEFORE AND AFTER THE EVENT AND RECOVERED WITHIN LIMITS. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED: THE FSE CLEANED BATHS AND VERIFIED THE INSTRUMENT'S OPERATION. A CLEAR ROOT CAUSE FOR THIS EVENT HAS NOT BEEN DETERMINED TO DATE. A MALFUNCTION WILL BE ASSUMED FOR THE PURPOSE OF THIS REPORT. >< A PATIENT SAMPLE WAS TESTED FOR HEMOGLOBIN (HGB) AND PLATELETS (PLT) ON THE COULTER AC-T DIFF 2 ANALYZER AND RESULTS OF 5.1G/DL AND 121X10^3CELLS/UL WERE OBTAINED, RESPECTIVELY. THE RESULTS WERE NOT REPORTED OUT OF THE LAB. THE ORIGINAL SAMPLE WAS RE-TESTED AND REPEATED RESULTS WERE NORMAL: HGB - 13.1G/DL, PLT - 271X10^3CELLS/UL. BASED ON THE AVAILABLE INFORMATION THERE WAS NO AFFECT TO PATIENT OR USER. | 0
1799020 | QC IS WITHIN SPECIFICATION WITH RESPECT TO CONTROLS (ACCURACY) AND REPRODUCIBILITY (PRECISION). A BCI FIELD SERVICE ENGINEER (FSE) INSPECTED THE UNIT AND FOUND NO ISSUES. A CLEAR ROOT CAUSE CAN NOT BE DETERMINED FOR THIS EVENT. >< A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO ERRONEOUS HIGH RED BLOOD CELL (RBC), PLATELETS (PLT), AND RBC INDICES RESULTS GENERATED ON THE COULTER LH 500 WITHOUT GENERATED SUSPECT MESSAGES. THE CUSTOMER REPLACED THE DILUENT REAGENT AND RERAN THE SPECIMEN IN THE MANUAL MODE AND OBTAINED LOWER RESULTS. THE SPECIMEN WAS RERUN A SECOND TIME TO CONFIRM THE RESULTS. THE ERRONEOUS RESULTS WERE NOT REPORTED OUT OF THE LABORATORY. PATIENT TREATMENT WAS NOT IMPACTED BY THIS EVENT. | 0
1693189 | DURING PRIMING OF THE DEVICE IN PREPARATION FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE ARTERIAL MONITOR UNEXPECTEDLY STOPPED WORKING. AN ALTERNATE DEVICE WAS EMPLOYED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WAS NO ADVERSE CONSEQUENCE TO THE PT AS A RESULT OF THE EVENT. >< EVALUATION IN PROGRESS BUT NOT YET CONCLUDED. | 1
1900457 | DURING PREPARATION FOR USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE TEMPERATURE DISPLAY FAILED TO WORK. NO OTHER DETAILS REGARDING THE NATURE OF THE EVENT WERE PROVIDED. THE USER REPORTED THERE WERE NO ADVERSE CONSEQUENCES TO A PT AS A RESULT OF THIS EVENT. | 1
1837532 | THIS UNIT HAS THREE COMPONENTS, CONTROL, SLAVE MONITOR AND AMPLIFIER. IT IS USED AS A MAPPING SYSTEM ASSISTING THE PHYSICIANS IN LOCATING AREAS THAT NEED RADIOFREQUENCY ABLATING. THE PROCEDURE CAN BE PERFORMED WITHOUT THE SYSTEM, HOWEVER, IT IS MORE EFFICIENT WITH IT. IT WAS NOT PROPERLY IDENTIFYING THE REFERENCE CATHETERS. THE ST. JUDE REP WAS HERE TO SUPPORT THE UNIT UNTIL THE STAFF ARE TRAINED. ONCE HE IDENTIFIED THE PROBLEM WITH THE SYSTEM, HE CONTACTED THEIR TECHNICAL SERVICES AND THEY DETERMINED THE PROBLEM WAS WITH THE AMP. IT SEEMED TO BE MAPPING OKAY ONCE THE REP WORKED WITH IT A LITTLE. THERE WERE NO QUESTIONS ABOUT WHERE THE LESIONS WERE MADE. THE REFERENCE CATHETER WAS NOT SEEN VISUALLY ON THE SCREEN UNTIL YOU REDUCED THE MAGNIFICATION. BUT PER THE REP AND MY UNDERSTANDING, THE UNIT RECOGNIZED THE REFERENCE CATHETER AS BEING IN THE PROPER LOCATION. ONCE THE REP CONFIRMED WITH ST. JUDE'S TECH SERVICE THAT THE AMP WAS THE PROBLEM, THEY MADE ARRANGEMENTS TO SEND OUT A NEW AMP. THE AMP IS THE SAME TYPE AS THE AS THE ORIGINAL.====================== HEALTH PROFESSIONAL'S IMPRESSION======================THE AMPLIFIER PART OF SYSTEM WAS NOT WORKING.====================== MANUFACTURER RESPONSE FOR 3D MAPPING (NAVX), ST. JUDE======================COMPANY REPLACED EQUIPMENT. | 0
1034602 | ON 26 MARCH 2008, BAXA WAS NOTIFIED OF AN INCIDENT THAT HAPPENED IN THE EVENING OF THE (B) (6) 2008. CUSTOMER INDICATED THAT THERE WAS A ORDERING DISCREPANCY WHICH CAUSED A 'DOSAGE ERROR.' DURING THE INITIAL CONTACT, THE CUSTOMER INDICATED THAT DOCTORS FORWARD THEIR ORDERS TO THE PHARMACY AND INDICATE THEIR ORDERS 'PER LITER' AND THAT THE ABACUS TEMPLATES ARE SETUP FOR THE PHARMACISTS TO ORDER 'PER BAG'. CUSTOMER AT THE TIME OF INITIAL CONTACT DID NOT WANT TO PROVIDE ANY FURTHER DETAILS ON THIS ISSUE. >< UPON FOLLOW UP WITH THE CUSTOMER: VOICE MESSAGES WERE LEFT FOR THE INITIAL REPORTING CUSTOMER ON THE 27 AND 31 OF MARCH. BAXA WAS NOTIFIED THAT THE INITIAL REPORTER IS NO LONGER WITH THE FACILITY ON THE 1 APRIL. UPON SPEAKING WITH ANOTHER PHARMACIST ON 03 APRIL, THE CUSTOMER WAS UNAWARE OF ANY ADVERSE EVENT THAT HAPPENED ON THE (B) (6) 2008. CUSTOMER INDICATED THAT HE WOULD LOOK INTO THE MATTER TO SEE IF HE COULD DETERMINE WHAT HAD HAPPENED. UPON THE CUSTOMER'S SUBSEQUENT INVESTIGATION ( FOLLOW-UP CONTACT ON (B) (6) (VOICE MESSAGE) & (B) (6)). THE CUSTOMER WAS NOT MADE AWARE OF ANY PT INVOLVEMENT AND COULD NOT FIND ANY EVIDENCE OF A PT INVOLVED INCIDENT UPON INVESTIGATING THIS ISSUE FURTHER. THE CUSTOMER SPOKE TO THE PHARMACY TECHNICIAN WORKING IN THE EVENING OF THE (B) (6) 2008 AND THE TECH IS UNAWARE OF ANY PT EVENT. CUSTOMER INDICATED THAT HE FEELS LIKE THIS INVESTIGATION IS COMPLETE ON HIS END. IF IN THE FUTURE, THE CUSTOMER CONTACTS BAXA, A FOLLOW-UP MDR WILL BE SUBMITTED WITH ANY ADDITIONAL INFO. | 1
1619088 | THE PT WAS ADMITTED WITH ACUTE WHEEZE, ASTHMA, AND CONGENITAL CYANOTIC HEART DISEASE WITH A HISTORY OF PULMONARY HYPERTENSION, TRISOMY 21, STATUS POST EPICARDIAL PACEMAKER, STATUS POST BLALOCK-TAUSSIG SHUNT, ASTHMA, CONGENITAL CYANOTIC HEART DISEASE AND SEIZURE DISORDER. ON (B) (6) 2010, THE PT WAS FOUND UNRESPONSIVE AND WITHOUT A PULSE. A CODE WAS CALLED, BUT RESUSCITATIVE MEASURES WERE UNSUCCESSFUL. | 0
1754837 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< (B) (4). THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. | 0
1396686 | IT WAS REPORTED THAT A (B) (6) PATIENT RECEIVED A BURN INJURY FROM THE OXIMAX SPO2 PROBE AND TRAM MODULE. EXTENT OF THE PATIENT'S INJURY IS NOT KNOWN AT THIS TIME. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1674374 | (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. >< CUSTOMER REPORTS, DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. | 0
1674375 | CUSTOMER REPORTS, DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT HARM. >< (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 0
1674367 | CUSTOMER REPORTS, DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. >< (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: CUSTOMER DISCOVERED AND RESOLVED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. THIS JAM WAS DUE TO OVER-STUFFING OF THE DRAWERS. | 0
1674365 | (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. >< CUSTOMER REPORTS, DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. | 0
1586903 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT HARM. >< (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 0
1586904 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. >< (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 0
1544285 | CUSTOMER REPORTS DRAWER ON THE PYXIS ANESTHESIA SYSTEM FAILED DUE TO ITEM OBSTRUCTION. NO PT PRESENT. >< (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: CUSTOMER REMOVED PHYSICAL ITEM FROM BEHIND DRAWER THROUGH USE OF BACK PANEL KEYS. | 0
1544284 | (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM JAM. >< CUSTOMER REPORTS DRAWER ON THE PYXIS ANESTHESIA SYSTEM FAILED. NO PT HARM REPORTED. | 0
1775981 | (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. >< CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT PRESENT. | 0
1600811 | (B)(4) . ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. >< CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT PRESENT. | 0
1750191 | CARDIAC ENZYMES OF CPK AND TROPONIN WERE ORDERED STAT ON A PT WITH KNOWN CORONARY DISEASE WHEN FEVER AND TACHYCARDIA DEVELOPED WHILE HOSPITALIZED FOR ANOTHER PROBLEM. THE TEST HAS RESULTED APPROX 2 HOURS AFTER THE ORDER, IN THE LATE AFTERNOON. THE RESULTS WERE SENT ONLY TO THE ELECTRONIC RECORD IN THIS WANNABE PAPERLESS HOSPITAL. IN THE LATE EVENING HOURS, 6-7 HOURS AFTER THE CPK RESULT OF 23,000 -RESULT WAS NORMAL 2 DAYS EARLIER- HAD BEEN DETERMINED BY THE LABORATORY, AN ATTENDING CONSULTANT SERENDIPITOUSLY HAPPENED UPON THE EMR OF THIS PT AND DISCOVERED THE LIFE THREATENING ILLNESS, RHABDOMYOLYSIS, THAT WAS UNK TO ANYONE. THE RESULTS SILENTLY POPULATED THE EMR. WHO KNEW? THERAPY WAS INITIATED 7 HOURS LATE, AND POSSIBLE CULPRIT THERAPIES WERE DISCONTINUED. THERE ARE ALREADY REPORTS OF THIS TYPE OF DEFECT ON MAUDE, RESULTING IN DEATH. THIS PT MAY DIE FROM THE RENAL FAILURE AND OTHER COMPLICATIONS SUFFERED FROM THIS DISEASE. THE DEFECT INVOLVES A DEFICIENCY IN THE DEVICE, A FAILURE TO WARN THAT A NEW RESULT HAS BEEN SENT TO THE DEVICE. IT PUTS EVERY PT AT RISK OF ADVERSITY FROM UNEXPECTED ABNORMAL RESULTS. THE WORKFLOW, TO WORK AROUND THIS DEFECT, REQUIRES TRAINED EMPLOYEES TO CONTINUALLY SCAN EACH REPOSITORY OF THE EMR FOR ELECTRONICALLY TRANSMITTED DATA, IN ORDER TO DETERMINE IN A TIMELY MANNER, ABSOLUTE ABNORMAL RESULTS OR CLINICALLY SIGNIFICANT INTERVAL CHANGES. HOSPITALS THAT HAVE DEPLOYED EMRS AND HAVE 'ADVANCED' BEYOND PAPER RESULTS SENT TO THE WARD BY FACSIMILE OR VACUUM TUBE, DO NOT CHECK RESULTS IN AN ORGANIZED MANNER TO PROTECT PTS SIMILAR TO THE ONE IN THIS REPORT. | 1
1878252 | MEDWATCH, I AM WRITING THIS DUE TO A POTENTIAL PT SAFETY ISSUE. I WAS UNABLE TO FIND A GOOD MATCH FOR A DOCUMENT FORMAT ON YOUR WEBSITE FOR THIS PROBLEM AND I AM NOT SURE IF YOURS IS THE PROPER AGENCY TO CONTACT. IF NOT I WOULD APPRECIATE IT IF YOU WOULD PLEASE FORWARD THIS TO THE PROPER AGENCY. THE SAFETY ISSUE I AM CONCERNED ABOUT IS THE POTENTIAL FOR GIVING MULTIPLE MEDICATION DOSES TO PTS DUE TO A FLAW IN THE NEW ELECTRONIC MEDICAL RECORDS (EMR) SYSTEM THAT IS BEING USED AT (B)(6) URGENT CARE FACILITIES IN (B)(6) WHERE I WORK. I WORK AT (B)(6) URGENT CARE. IT IS A RELATIVELY FAST PACED, HIGH PT VOLUME ENVIRONMENT. NURSES DOING TREATMENT AT THE FACILITY MAY SHARE THE CARE OF INDIVIDUAL PTS. MULTITASKING IS THE RULE. PRIOR TO THE INITIATION OF USE OF EMR, A NURSE WOULD RECEIVE ORDERS FROM PROVIDERS ON PAPER. WHEN THE NURSE WOULD TAKE RESPONSIBILITY OF A GIVEN ORDER, SUCH AS GIVING A MEDICATION, STARTING AN IV, DOING A LAB, ETC., HE OR SHE WOULD WRITE THEIR INITIALS NEXT TO THE ORDER. THIS WOULD PREVENT OTHER NURSES FROM COMING AND DOING THE SAME ORDER. WITH THE USE OF THE NEW PAPERLESS EMR SYSTEM THIS METHOD OF INFORMING COWORKERS THAT YOU HAVE DONE AN ORDER NO LONGER EXISTS. I WILL ATTEMPT TO EXPLAIN. THERE IS A COMMON LIST OF PTS TO BE SEEN AND TREATED. ALL NURSES WORK FROM THIS LIST. TYPICALLY AT (B)(6) THERE WILL BE 2 OR 3 NURSES AND A MEDICAL ASSISTANT WORKING ON THE TREATMENT TEAM FOR A GIVEN SHIFT. THERE IS A COLOR CODED 'DOT' THAT APPEARS BEFORE THE NAME OF A PT INDICATING THAT A PROVIDER OR TRIAGE NURSE HAS ORDERED A PROCEDURE OR MEDICATION. A NURSE OR MA WILL SELECT THE PT FROM THE LIST, REMOVE THE DOT, AND GIVE THE MEDICATION OR DO THE PRESCRIBED PROCEDURE. AT THIS POINT, THERE IS NO PROBLEM. EVEN THOUGH THERE IS NO EVIDENCE THE MEDICATION WAS GIVEN IN THE LIST OF ORDERS FOR THAT PT, THE COLORED DOT IS GONE SO SUBSEQUENT NURSES LOOKING AT THE COMMON LIST OF PTS WOULD NOT BE PROMPTED TO GO IN AND TREAT THE PT. HOWEVER, IF A PROVIDER ORDERS ANOTHER MEDICATION OR PROCEDURE, THE DOT IN FRONT OF THAT PTS NAME REAPPEARS. A NURSE SEES THE DOT AND GOES INTO THE LIST OF ORDERS FOR THAT PT. THERE IS NO INDICATION OF WHICH ORDERS HAVE BEEN DONE, AND MOST IMPORTANTLY WHICH MEDICATIONS HAVE BEEN GIVEN. IT IS POSSIBLE FOR THE NURSE TO 'DATA MINE' OTHER SECTIONS OF THE CHART AND, IF THE PREVIOUS NURSE HAS HAD TIME TO ENTER A PROGRESS NOTE, DETERMINE IF THE MEDICATION HAS BEEN GIVEN. IN THIS FAST PACED ENVIRONMENT THIS CAN BE OVERLOOKED. MOSTLY NURSES TRY TO COMMUNICATE VERBALLY WITH OTHER MEMBERS OF THE TEAM TO TRY AND FIGURE OUT WHICH ORDERS HAVE BEEN DONE. THIS IS NOT ALWAYS POSSIBLE BECAUSE OTHER MEMBERS OF THE TREATMENT TEAM MAY BE IN CLOSED ROOMS WITH PTS. ON MY FIRST DAY WORKING TREATMENT USING EMR (I NORMALLY WORK TRIAGE) I POINTED OUT THIS FLAW TO MY SUPERVISOR. SHE IN TURN BUMPED IT UP THE CHAIN OF COMMAND. ACCORDING TO THE FEEDBACK I HAVE RECEIVED, CORRECTING THIS FLAW HAS BECOME A 'TOP PRIORITY' FOR (B)(6). I WAS TOLD A HOSP REP WAS SENT TO THE SOFTWARE VENDOR IN AN ATTEMPT TO GET THE COMP PROGRAM FIXED SO THAT WHEN A NURSE GIVES A MEDICATION IT IS INDICATED IN THE ORDER QUEUE. AS OF THIS DATE, 6 WEEKS AFTER THE INITIATION OF THE EMR PROGRAM, THE SOFTWARE FLAW REMAINS. A TEMPORARY FIX, OUTLINED IN THE ATTACHED EMAIL DATED 09/16/2010, WAS SUGGESTED. I HAVE NOT OBSERVED THIS TEMPORARY FIX USED BY ANYONE, NOR HAVE I SEEN ANY ATTEMPT TO ENFORCE ITS USE. IF I CAN CLARIFY ANY PART OF THIS LETTER PLEASE FEEL FREE TO CONTACT ME. EMR MEDICATION ERROR AVOIDANCE: MANY OF YOU AND YOUR STAFF HAVE EXPRESSED CONCERN OVER THE POSSIBILITY OF MEDICATION ERRORS IN LIGHT OF USING THE EMR. WE ARE WORKING WITH (B)(6) FOR A SYSTEM FIX, BUT IN THE MEANTIME, HAVE A MORE PEDESTRIAN PROCESS. AS SUGGESTIONS ARE MADE (WHICH WE WELCOME) WE WILL REVIEW AND IMPLEMENT APPROPRIATELY, BUT WOULD APPRECIATE IF ALL CLINICS WERE USING THE SAME PROCESS. IF THE PT MEETS ANTI-PYRETIC PROTOCOL AND IS GIVEN IBUPROFEN OR TYLENOL FROM TRIAGE - PUT A STICKY NOTE ON THE GREEN CHART AFTER ADMINISTRATION TO ALERT THE PERSON WHO ROOMS THE PT. WHEN THE PT IS ROOMED, WRITE 'TYLENOL OR MOTRIN GIVEN' ON THE WHITE BOARD ON THE PT'S ROOM SO THE MED IS NOT GIVEN A SECOND TIME. WHEN MEDICATIONS ARE ORDERED, THE NURSE WHO IS TAKING THE ORDER OFF SHOULD WRITE THE MEDICATION(S) ON THE WHITE BOARD BEFORE GIVING THE MED(S). AFTER THE POST-MED ASSESSMENT, THE MEDICATION SHOULD BE CROSSED OFF THE WHITE BOARD (NOT ERASED). THE MEDICATION SHOULD BE ERASED WHEN THE PT IS DISCHARGED FROM THE ROOM. IF A PT CHANGES ROOMS (I.E. COMES BACK FROM X-RAY TO A DIFFERENT ROOM) THE MEDS SHOULD BE WRITTEN ON THE BOARD AGAIN TO AVOID DUPLICATING MEDICATIONS. WE ARE WORKING HARD TO COME UP WITH OTHER SOLUTIONS AND WELCOME YOUR INPUT, BUT FEEL THAT THIS PROCESS WILL SERVE IN THE BEST INTEREST OF PT SAFETY IMMEDIATELY. | 1
1053470 | AN ADVERSE TREND FOR LIGHT STAINING AND INCONSISTENT STAINING WAS OBSERVED THROUGH MONITORING AND ANALYSIS OF INSTRUMENT SERVICE DATA. THE RISK ASSESSMENT PROCESS REQUESTED MEDICAL OPINION INDICATING THAT IF SAME SLIDE CONTROLS ARE NOT USED, THERE IS A POTENTIAL FOR A FALSE NEGATIVE RESULT. THE USE OF SAME SLIDE CONTROLS IS IN ALIGNMENT WITH THE (B) (4) AND THE (B) (4) AS STATED IN THE FEDERAL REGISTRY (SECTION 493. 1273; VOLUME 68, #16; 01/24/2003; PAGE 3709). THERE HAS BEEN NO REPORT OF MISDIAGNOSIS OR DELAY IN TREATMENT ASSOCIATED WITH THIS ISSUE. >< PRELIMINARY ANALYSIS INDICATES THE HEATER PADS WERE NOT OPERATING PROPERLY AND NEEDED TO BE REPLACED. THE INVESTIGATION IS ON-GOING. | 0
1637153 | PT INFO IS CONSIDERED CONFIDENTIAL, PER COUNTRY REGULATIONS. >< ACCORDING TO THE CUSTOMER, ON (B) (6) 2010 AT 10:10 AM AT ICU (B) (6), A PT HAD A RUN OF V-TACH, THEN V-FIB, WHILE BEING WASHED (THE ECG ELECTRODES WERE REPORTED AS BEING INTACT ON THE PT). PER THE HOSPITAL, THE V-TACH AND V-FIB EPISODE WAS NOT DETECTED BY THE B850 MONITOR BUT THE MONITOR ALLEGEDLY PROVIDED A HIGH HEART RATE ALARM, ALERTING THE DOCTOR OF THE EVENT. THE PT REPORTEDLY FAINTED, AND WHEN THE DOCTOR ARRIVED, THE DOCTOR VENTILATED THE PT WITH A MASK, AND THEN CONNECTED THE PT TO A VENTILATOR. THE PT'S IMPLANTABLE DEFIBRILLATOR FIRED DURING THE EVENT AND THE PT SPONTANEOUSLY RECOVERED FROM THE EVENT. PRIOR TO THE EVENT, THE PT WAS BEING TREATED FOR: DILATED CARDIOMYOPATHIES VIA ANTIBIOTIC THERAPY, OXYGEN, AND INOTROPIC DRUG THERAPY: DOPUTREX. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1571707 | TELEMETRY RHYTHM WAS NOT BEING PICKED UP BY THE MONITOR. LEADS WERE CHANGED AND THE TELEMETRY BOX WAS CHANGED AN HOUR LATER, THE PATIENT HAD A COURSE OF VFIB ALARMS. THE MONITOR DID NOT SOUND (ALARM), PATIENT CODED AND DIED. BIOMED LOOKED AT THE BOX (TRANSMITTER) AND COULD NOT FIND ANY PROBLEMS. | 1
1420098 | CUSTOMER ADVISORY (ATTACHED) WILL BE SENT TO ALL AFFECTED CUSTOMERS IN APRIL 2009. THE PROBLEM WILL BE RESOLVED IN RELEASE (B)(4), WHICH IS SCHEDULED TO BE AVAILABLE IN AUGUST 2009. >< IN OUR FOCAL WORKSTATION, THE SET UP REFERENCE DIALOG IS USED TO GENERATE SHIFTS FROM THE 'SCAN REFERENCE POINT' POSITION TO THE 'SETUP REFERENCE POINT'. IN RELEASE 4.34 THE ABILITY TO MANUALLY ENTER THE SHIFTS WAS INTRODUCED, BY ENTERING THESE SHIFTS THE 'SETUP REFERENCE' POINT IS MOVED AS DEFINED BY THE USER. A USER REPORTED THAT A PATIENT IS DEFINED AS PRONE AND THE SHIFTS ARE MANUALLY ENTERED AND THE 'LOCK SHIFT' BOX IS CHECKED, THE 'X' COORDINATE (LEFT/RIGHT) AND THE 'Z' COORDINATE (ANTERIOR/POSTERIOR) ARE REVERSED (LEFT TO RIGHT, ANTERIOR TO POSTERIOR). THIS COULD POTENTIALLY LEAD TO INCORRECT ISOCENTER DEFINITION FOR PRONE PATIENTS. ISOCENTER LOCATION SHOULD BE CHECKED WITH PORTAL IMAGING AND, IF PERFORMED, IT WILL BE EVIDENT TO THE USER THAT THIS ISOCENTER IS INCORRECT. THE USER REPORTED THAT THERE WERE NO PATIENTS INJURED AS A RESULTS OF THIS SOFTWARE BUG. | 1
1440202 | DEVICE EVALUATION - CUSTOMER REPORT WAS CONFIRMED. REC'D (1) VAMP JR. SYSTEM. SYSTEM APPEARED NOT USED. BOTH VENTED CAPS WERE STILL ATTACHED AT BOTH PROXIMAL AND DISTAL ENDS OF THE SYSTEM. PEDIATRIC TUBING WAS COMPLETELY BROKEN FROM SOLVENT BOND JOINT WITH SAMPLE SITE. ROUGH SURFACES WERE OBSERVED AT BROKEN TUBING ENDS. SITE. >< IT WAS REPORTED THAT THE PRODUCT CAME IN TWO PIECES, AND IT APPEARED THAT THE TUBING WAS BROKEN WITHIN THE PACKAGE.. >< IT WAS REPORTED THAT THE DEVICE WAS EXPLANTED AFTER AN IMPLANT DURATION OF APPROXIMATELY 119.2 MONTHS, DUE TO UNKNOWN REASONS. NO FURTHER DETAILS WERE PROVIDED. | 0
1592113 | THE CUSTOMER REPORTED THAT A PT DEATH OCCURRED AND THAT THE PT LAST HAD A WAVEFORM AN HOUR PRIOR TO BEING DISCOVERED DEAD BY STAFF. >< THE CUSTOMER REPORTED THAT A PT DEATH OCCURRED AND THAT THE PT LAST HAD A WAVEFORM AN HOUR PRIOR TO BE DISCOVERED DEAD BY STAFF. WE WILL CONSIDER THAT THE CONDITION OF THE PT WAS UNMONITORED FOR APPROX ONE HOUR AND THAT STAFF WERE NOT AWARE OF CHANGES IN THE PT'S CONDITIONS DURING THIS TIMEFRAME. THE INITIAL INFO INCLUDES THAT THE PT WAS NOT BEING MONITORED BY THE TELEMETRY SYSTEM AT THE TIME THAT THEY WERE DISCOVERED TO BE DEAD. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFO REGARDING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. (B) (4). | 1
1864353 | PT WAS ON TELEMETRY, BUT TELEMETRY WASN'T SHOWING RHYTHM. TOP OF TELEMETRY UNIT FOUND TAPED TO TELEMETRY CHANNEL. SCREW FOR BATTERY COVER HAD WORN AND HAD BEEN TAPED IN PLACE. TAPE HAD COME LOOSE AND BATTERIES DID NOT CONNECT PROPERLY. PT FOUND UNRESPONSIVE, NO RESPIRATION, WEAK PULSE. CODE BLUE CALLED, FAMILY REQUESTED CPR STOPPED. PT EXPIRED. | 1
1757991 | AFTER GE EMR SOFTWARE UPGRADE SIGNIFICANT PT INFORMATION IS OMITTED FROM FINAL REPORT OF PROCEDURE THAT AFFECTS DISCHARGE MEDICATIONS POST CORONARY DRUG ELUTING STENT PLACEMENT. THE GE MACLAB/CARDIOLAB XT UPGRADE INTERFACE TO THE GE CENTRICITY CARDIOLOGY DATA MGMT SYSTEM DOES NOT COMMUNICATE THE USE OF EVERILOMUS OR ZOTARILOMUS DRUG ELUTING STENTS IN THE FINAL REPORT OF PROCEDURE. WITHOUT THESE FINDINGS IN THE REPORT, PTS WITH THESE STENTS COULD BE DISCHARGED WITHOUT DUAL ANTIPLATELET THERAPY TO PREVENT SUBACUTE THROMBOSIS IN THE TREATED CORONARY ARTERY-S-. | 1
1593227 | THE CUSTOMER REPORTED THAT WHILE THE PANORAMA CENTRAL STATION WAS IN USE MONITORING PTS ALONG WITH AMBULATORY TELEPACKS, THE CENTRAL STATION DISPLAY BLACKED OUT. NO PT INJURY WAS REPORTED. >< THE COMPANY SERVICE REP REPLACED THE DISPLAY POWER SUPPLY. THE SYSTEM WAS TESTED TO FACTORY SPECS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. | 1
1593226 | THE CUSTOMER REPORTED THAT WHILE THE PANORAMA CENTRAL STATION WAS IN USE MONITORING PTS ALONG WITH AMBULATORY TELEPACKS, THE CENTRAL STATION DISPLAYS WENT OUT. NO PT INJURY WAS REPORTED. >< THE COMPANY SERVICE REP REPLACED THE SYS HARD DRIVE. THE SYSTEM WAS TESTED TO FACTORY SPECS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. | 1
1593229 | THE COMPANY SERVICE REP TALKED THE CUSTOMER THROUGH RESTING THE ELAN SWITCH OVER THE PHONE. NORMAL OPERATION WAS RESTORED. HE ORDERED A NEW SWITCH. ON (B)(6), HE RETURNED TO THE FACILITY AND REPLACED THE SWITCH. THE SYS WAS TESTED TO FACTORY SPECS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. >< THE CUSTOMER REPORTED THAT WHILE THE PANORAMA CENTRAL STATION WAS IN USE MONITORING PATIENTS ALONG WITH BEDSIDE MONITORS AND AMBULATORY TELEPACKS, THE CENTRAL STATION SHOWED COMMUNICATION LOSS ON EVERY PT CHANNEL. NO PT INJURY WAS REPORTED. | 1
1648738 | MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT. >< OUR REP IS REPORTING TO US THAT A PATIENT HAD AN ARTHROSCOPIC SHOULDER REPAIR. THE SURGEON WAS USING A LINVATEC HANDPIECE WITH AN FMS INTERFACE CABLE AND THE LINVATEC FOOT PEDAL CONTROLS. WHILE EXITING THE JOINT SPACE WITH THE HANDPIECE, THE BLADE WAS STILL ACTIVATED AND SOMEHOW THE SURGEON CAUGHT SOME TISSUE. [PER FOLLOW-UP PHONE CONVERSATION WITH THE REP: THE REP STATED THAT THE AFFECTED TISSUE HAD MINOR DAMAGE. THE SURGEON NICKED SOME OF THE BURSA AND SOME ROTATOR CUFF TISSUE MAKING THEM A LITTLE RAGGED. FOR REPAIR, THE SURGEON ABLATED AND CLEANED UP THE AREA; ADDED LESS THAN 20 SEC TO THE PROCEDURE; THERE WAS NO PATIENT HARM OR CONSEQUENCES. THE SURGEON WAS USING AN OLD STYLE, PREVIOUS GENERATION INTERFACE CABLE, WHICH MAY NOT HAVE BEEN COMPATIBLE WITH THE LINVATEK DEVICE MODEL. THE SURGEON SWITCHED TO THE PRESENT GENERATION INTERFACE CABLE, WHICH WORKED FLAWLESSLY, TO COMPLETE THE PROCEDURE WITHOUT FURTHER ISSUE OR HARM TO THE PATIENT. AFRIGAULT (B)(4) 2010 13:22:59. | 0
1741725 | INFO PROVIDED BY THE CUSTOMER INCLUDED A STATEMENT THAT NO ECG CABLE WAS CONNECTED TO THE TELEMETRY TRANSMITTER AT THE TIME OF THE EVENT. THE PRODUCT OPERATIONS MANUAL STATES THAT A VALID ECG SIGNAL IS REQUIRED IN ORDER TO HAVE THE SPO2 PARAMETER FULLY FUNCTIONAL. SPACELABS' TESTING OF THE TRANSMITTER, MONITOR, RECEIVER, AND ANTENNA SYSTEM AT THE CUSTOMER'S SITE CONFIRMED THAT THE TELEMETRY SYSTEM WAS OPERATING TO SPECIFICATION. SPACELABS HAS CONCLUDED THAT THERE WAS NO ALARM BECAUSE THE CUSTOMER MONITORED THE PT'S SPO2 WITHOUT CONNECTING THE ECG LEADS TO THE TRANSMITTER AS REQUIRED. THE CUSTOMER WAS GIVEN A COPY OF THE OPERATIONS MANUAL PAGE SPECIFYING THAT ECG USE IS MANDATORY FOR PROPER OPERATION OF THE SPO2 PARAMETER. THE CUSTOMER HAS ACKNOWLEDGED THIS REQUIREMENT. THIS INITIAL REPORT IS CONSIDERED FINAL AND THE ISSUE IS CLOSED. >< SOUTH SHORE HOSPITAL REPORTED THAT A SPACELABS TELEMETRY SYSTEM FAILED TO ALARM ON A PT EVENT (LOW SPO2). THE PT ULTIMATELY DIED. | 1
1581694 | A PRELIMINARY INVESTIGATION HAS DETERMINED THE REPORTED ERROR CONDITION IS INTERMITTENTLY BEING TRIGGERED WHEN THE USER ATTEMPTS TO CONNECT AN ADDITIONAL INFORMATION PUMP, WHILE IN THE APPLICATION, THE ORDER 'DOCUMENT AN ACTION' WINDOW IS OPEN. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION. >< IT WAS REPORTED THAT WHILE CONNECTING AN ADDITIONAL THIRD PARTY INFUSION PUMP TO THE CRITICAL CARE MGR (MEDICAL DEVICE DATA SYSTEM) APPLICATION, THE MANUAL DOCUMENTATION OF ANOTHER PUMP'S 'FLUID IN' DATA CHANGED FROM MANUAL TO AUTOMATIC AND RECORDED THE LAST KNOWN DATA VALUE CONTINUOUSLY. THE AUTOMATIC RECORDING OF DATA COULD NOT BE STOPPED. USERS COULD NOT MANUALLY UPDATE THE APPLICATION FOR NEW DOSES OF ADMINISTERED MEDICATION 'FLUID IN'. NO PT INJURIES HAVE BEEN REPORTED IN CONNECTION WITH THIS ERROR CONDITION. INFUSION OF FLUIDS AND/OR MEDICATION FLUIDS IS NOT INTERRUPTED AT THE INFUSION PUMP. THE ERROR IMPACTS RECORDING OF FLUID BALANCE DATA WITHIN THE CRITICAL CARE MGR APPLICATION. | 1
1205690 | CERNER HAS DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION OCTOBER 10, 2008 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE AND A QUERY THAT CLIENTS MAY USE TO DETERMINE IF SITES ARE AFFECTED BY THE ISSUE. A SOFTWARE MODIFICATION IS BEING DEVELOPED TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORP. WILL PROVIDE A FOLLOW-UP REPORT WHEN THE SOFTWARE MODIFICATION IS AVAILABLE. >< THE ISSUE INVOLVES THE ENTERPRISE MASTER PERSON INDEX (EMPI) FUNCTIONALITY WITHIN CERNER MILLENNIUM FOREIGN SYSTEM INTERFACES (FSI) AND AFFECTS SITES THAT HAVE THE BY FEED EMPI RECONCILE TURNED ON. WHEN AN INVALID CONTRIBUTOR SYSTEM IS INACTIVATED, A USER HAS THE ABILITY TO OVERRIDE THE SYSTEM INTEGRATION MANAGER (SI_MANAGER.EXE) WARNINGS AND MANUALLY REACTIVATE THE CONTRIBUTOR SYSTEM. SYSTEM INTEGRATION MANAGER UPDATES THE END_EFFECTIVE_DT_TM WHEN AN INVALID CONTRIBUTOR SYSTEM IS INACTIVATED. IF THE CONTRIBUTOR SYSTEM IS MANUALLY REACTIVATED, THE END_EFFECTIVE_DT_TM IS NOT UPDATED AGAIN. THE IMPACT ASSOCIATED WITH THIS ISSUE IS THAT A COMBINE OCCURS FOR PATIENTS THAT SHOULD NOT OCCUR. MULTIPLE PATIENTS THAT DO NOT HAVE THE SAME IDENTIFIERS ARE COMBINED INTO A SINGLE PATIENT. THIS COULD RESULT IN PATIENT'S PREVIOUS CLINICAL EVENTS AND DOCUMENTED OBSERVATIONS AND ORDERS NOT BEING AVAILABLE FOR ELECTRONIC DECISION SUPPORT. ADDITIONALLY, INACCURATE INFORMATION COULD COMBINE INTO A DISASSOCIATED PATIENT RECORD. CERNER HAS NOT RECEIVED COMMUNICATION OF AN ADVERSE PATIENT EVENT AS A RESULT OF THIS ISSUE. | 1
1267946 | THE ISSUE INVOLVES MEDICAL DEVICE INTEGRATION. THE +PR SOFTWARE FLAG ALLOWS FOR A NEW CAPABILITY TO APPEND A VALUE TO THE RESULT BASED ON A GIVEN ERROR CODE SENT BY THE DEVICE. THIS IS MANAGED BY THE INTERFACE CREATING A FLAT FILE ON THE NETWORK, WHICH DEFINES APPENDING THE < SIGN FOR AN ERROR FLAG OF 4 AND A > SIGN FOR AN ERROR FLAG OF 5. THIS IS NOT CORRECT. THE FLAT FILE SHOULD HAVE BEEN UPDATED TO APPEND THE > SIGN FOR AN ERROR FLAG 4 AND A < SIGN FOR AN ERROR FLAG OF 5, AND THE INTERFACE SHOULD HAVE BEEN CYCLED IN CERNER ENTERPRISE MANAGER. CERNER HAS NOT RECEIVED COMMUNICATION ON ANY ADVERSE PATIENT EVENTS AS A RESULT OF THIS ISSUE. >< CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON (B) (6), 2008 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE AND A SOFTWARE MODIFICATION TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORPORATION CONSIDERS THE ISSUE RESOLVED AND NO FURTHER NARRATIVE IS REQUIRED FOR FOLLOW-UP. | 1
1647149 | I HAD AUTOMATED LAMELLAR KERATOPLASTY, (B) (6) EYECARE, (B) (6), (B) (6) 1994 - AND (B) (6) 2008, PRK AT (B) (6) EYE INST., (B) (6). IMMEDIATELY AFTER THE FIRST, HAD SEVERE HALOS, WHICH SUBSIDED AFTER A COUPLE YEARS. THEREAFTER, EYES RECEDED GETTING PROGRESSIVELY MORE NEARSIGHTED, BACK ON GLASSES EST. 2000. AGAINST (B) (6) ADVICE STATED HAD THIN CORNEAS, SOUGHT PRK FROM (B) (6) . (B) (6) WENT OUT OF BUSINESS, NEVER HAD LAST CHECK UP. LEFT EYE FINE, RIGHT EYE HAS GHOSTING, DOUBLE VISION, SEEMS TO BE GETTING WORSE. EXTREMELY DIFFICULT TO READ IN DIM LIGHTING/NIGHT. NEED GLASSES. (B) (6) - UNDERSTAND PRESBYOPIA. AM ALSO NOW UNEMPLOYED, NO INSURANCE. CAN DO NOTHING ABOUT THIS. I'M NOT UPSET, BUT CONCERNED AND WORRIED ABOUT THE FUTURE, ABOUT OLD AGE, PERHAPS A NECESSITY FOR AN INTAKE OR CORNEAL TRANSPLANT OR NOT BEING ABLE TO UNDERGO CATARACT REMOVAL WHEN I'M OLD. OTHER THAN THAT, HAVE THE USUAL DRY EYES AND GLARE. ALSO HAD RADIAL KERATOTOMY IN ONE OF THE EYES -A TOUCH UP-, 1994, DON'T REMEMBER WHICH. | 0
1668478 | DURING PUBLICATION REVIEW, WE BECAME AWARE OF A LETTER TO EDITOR PUBLISHED IN RESUSCITATION (VOLUME 80, ISSUE 7, JULY 2009, PAGE 839; COPYRIGHT 2009 ELSEVIER IRELAND LTD). THE AUTHOR (CHRISTIAN VON BARY) DISCUSSES TWO INJURIES - LIVER LACERATIONS AND INTRA-ABDOMINAL BLEEDING NOTICED VIA A CT SCAN AFTER USE OF MECHANICAL CPR DEVICE. THE AUTHOR HAS DESCRIBED THE SITUATION AS PATIENT HAD UNDERGONE RADICAL PROSTATECTOMY. THE PATIENT DEVELOPED PULMONARY EMBOLISM AND WENT INTO CARDIAC ARREST, BECAUSE OF WHICH HE COLLAPSED IN THE BATHROOM. THE PATIENT WAS SUCCESSFULLY RESUSCITATED WITH A MECHANICAL CARDIAC COMPRESSION DEVICE. A CT SCAN CONFIRMED PULMONARY EMBOLISM AND ALSO REVEALED LIVER LACERATIONS AND INTRA-ABDOMINAL BLEEDING. THE AUTHOR HAS CITED SOME COMBINATION OF MULTIPLE POTENTIAL CAUSES FOR LIVER LACERATION, INCLUDING: RADICAL PROSTATECTOMY PERFORMED PRIOR TO CARDIAC ARREST AND MECHANICAL CPR; (INJURY RESULTING FROM) PATIENT'S COLLAPSE IN THE BATHROOM; AND IMPROPER PLACEMENT OF THE LOAD-DISTRIBUTING BAND USED WITH MECHANICAL CPR. THE AUTHOR HAS CITED SOME COMBINATION OF TWO DIFFERENT POTENTIAL CAUSES FOR INTRA-ABDOMINAL BLEED INCLUDING; THROMBOLYTIC THERAPY AND MECHANICAL COMPRESSION DEVICE. >< WHILE THE AUTHOR HAS NOT CONCLUSIVELY OFFERED EXACT CAUSE(S) FOR EITHER OF THE TWO INJURIES, IT CLEARS THAT THE USE OF THE MECHANICAL COMPRESSION DEVICE WAS SUCCESSFUL IN RESUSCITATING THE PATIENT. AS PER THE PUBLICATION OF 'THE SCIENCE AND PRACTICE OF RESUSCITATION MEDICINE' - 2ND EDITION (ATTACHMENT B) UNDER THE 'COMPLICATIONS OF CHEST COMPRESSIONS' SECTION, POSSIBLE COMPLICATIONS FROM CHEST COMPRESSIONS INCLUDE: RIB FRACTURE, STERNAL FRACTURE, FRACTURE OF CLAVICLES, SCAPULAE AND VERTEBRAL COLUMN, FLAIL CHEST, SKIN TRAUMA, SUBCUTANEOUS EMPHYSEMA, PULMONARY EDEMA, PULMONARY HEMATOMA, PNEUMOTHORAX, HEMOTHORAX, EMBOLIZATION, LUNG LACERATION, PNEUMOMEDIASTINUM, MEDIASTINITIS, CARDIAC CONTUSION/RUPTURE, PERICARDITIS LIVER LACERATION AND BOWEL INJURIES. PATIENT ANATOMICAL FEATURES AND PRE-EXISTING CONDITIONS COULD ALSO INCREASE THE POSSIBILITY OF SUCH INJURIES WHEN APPLYING EITHER MANUAL OR MECHANICAL CPR. SINCE THE ALTERNATIVE TO CPR IS DEATH, THE CLINICAL COMMUNITY AT LARGE DOES NOT CONSIDER SUCH INJURIES AS SIGNIFICANT. | 0
1747284 | AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND PERFORMED THE READER CAMERA ADJUSTMENT AND GENERATED A REFERENCE IMAGE.(B) (4). >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL.OCD MEDICAL DIRECTOR HAS CLASSIFIED THIS AS A SERIOUS INJURY. | 0
1541930 | THIS REPORT IS BEING FILED UPON NOTIFICATION FROM OUR INFO TECHNOLOGY MANUFACTURER IN (B)(4) TO SUBMIT THIS EVENT THAT HAPPENED IN (B)(6). PROBLEM: THE CUSTOMER COMPLAINTS OF INTERMITTENT LOSS OF PT IMAGES WHEN SENDING IMAGES FROM THE X-RAY SYSTEM TO THE VIEWFORUM WHICH IS A WORKSTATION, PICTURE ARCHIVING, AND COMMUNICATION SYSTEM. >< THE INVESTIGATION IS STILL ONGOING ON THIS EVENT. WHEN THE INVESTIGATION IS COMPLETED A FOLLOW-UP REPORT WILL BE SENT TO THE FDA. | 1
1610365 | HLA SOFTWARE SYSTEM WITHOUT ANY DOCUMENTATION THAT PERFORMS HLA MATCH CALCULATIONS AND REPORTING. UNABLE TO ACQUIRE USER MANUALS. TRANSPLANT SOFTWARE CALLED OTTR THAT PERFORMS CALCULATIONS AND RELATIVE DETERMINATIONS ON PT DATA AND USES ALOGRITHMS TO DETERMINE PT TREATMENT PLANS. UNABLE TO ACQUIRE DOCUMENTATION AND ADEQUATE USER MANUALS. | 0
1701409 | WE RECENTLY COMPLETED A CARDIOLOGY SYSTEM HARDWARE UPGRADE. THE VENDOR OPTED TO UPGRADE THE SYSTEM SOFTWARE AT THE SAME TIME. ABOUT 10 DAYS AFTER THE UPGRADE, IT WAS REPORTED THAT THE LOW DOSE STRESS REPORTS WERE NOT ABLE TO GENERATE A REPORT, AND PHYSICIANS WERE NOT ABLE TO COMPLETE AND SIGN REPORTS. THE SPECIFIC STATEMENT THAT WOULD NOT EMBED IN THE REPORT IS 'LOW DOSE: THERE WERE NO REGIONAL WALL MOTION ABNORMALITIES.' BECAUSE OF THIS, A REPORT WAS NOT GENERATED FOR A PATIENT AWAITING LIVER TRANSPLANTATION. ALTHOUGH THE DIAGNOSTIC TEST HAD BEEN COMPLETED, THE ANESTHESIOLOGIST WAS UNAWARE OF THIS BECAUSE OF THE SOFTWARE PROBLEM, AND AS A RESULT, THE PATIENT DID NOT RECEIVE A NEEDED ORGAN. THE VENDOR WAS NOTIFIED IMMEDIATELY AND RESOLVED THE SOFTWARE PROBLEM WITHIN TWO DAYS.====================== HEALTH PROFESSIONAL'S IMPRESSION======================GLITCH CAUSED BY SOFTWARE UPGRADE====================== MANUFACTURER RESPONSE FOR SOFTWARE, VERICIS======================THEY HAVE INDICATED THEY WILL BE REPORTING TO MEDWATCH | 1
1075722 | IT WAS REPORTED THAT ON (B) (4) 2008 AT AROUND 8:30 AM APPROXIMATELY 4 MONITORS REBOOTED AT THE SAME TIME FOR ABOUT 5 MINUTES. AFTER FIVE MINUTES, THE MONITORS BOOTED UP NORMALLY. NO PATIENT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1610402 | TELEMETRY BOX FOUND MALFUNCTIONING, TRACING WENT OFF SCREEN. UPON NURSES INSPECTION, THERE WAS A HOLE ABOUT THE SIZE OF A DIME THAT MELTED TO THE EXTERIOR OF THE UNIT. BOX WAS HOT TO TOUCH AND IT SMELLED LIKE BURNING PLASTIC. THERE WAS NO HARM TO THE PT. | 0
1701313 | IT WAS REPORTED THAT DURING CLEANING, AN AS/3 MONITOR MOUNTED ON A CEILING MOUNT BRACKET BROKE AWAY FROM THE BRACKET AND FELL TO THE GROUND. THE ACTUAL BASE OF THE F-CU8 CASE BROKE AROUND THE BRASS INSERTS THAT SECURE THE BRACKET TO THE MONITOR. THE MONITOR HAS BEEN REMOVED FROM SERVICE FOR EVALUATION. NO STAFF OR PATIENT INJURY WAS REPORTED IN THIS INCIDENT. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. >< THE DEVICE MANUFACTURE DATE IS UNKNOWN. THE DEVICE WAS SHIPPED ON (B)(4)1993. | 0
1495626 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< (B) (4). THE METER DID NOT CONFIRM FOR INCORRECT DATE AND TIME. HOWEVER, THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT DATE AND TIME ISSUES WHEN THE DATA IS UPLOADED TO A COMPUTER. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO PRECISION LINK. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1778117 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE DEVICE MANUFACTURE DATE IS UNKNOWN AS THE METER SERIAL NUMBER WAS NOT PROVIDED. | 0
1645851 | A CUSTOMER REPORTED THE TIME LOSES MINUTE ON THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THE METER WAS CONFIRMED TO EXHIBIT A DATE AND TIME ISSUE. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT DATE AND TIME ISSUES WHEN THE DATA IS UPLOADED TO A COMPUTER. THIS OCCURS WHEN RESULTS OBTAINED ON A METER WITH INCORRECT DATE AND TIME ARE UPLOADED TO PRECISION LINK. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1832189 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1671263 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1512662 | (B) (4). THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1745246 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1636899 | A CUSTOMER REPORTED HE RECEIVED AN ERROR-6 DISPLAY MESSAGE ON HIS BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN HIS METER WERE NOT PROPERLY SET, AND HE REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE MANUFACTURE DATE FOR THE REPORTED METER IS UNKNOWN. | 0
1438140 | (B) (4). THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. >< A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1784007 | (B)(4). THE CUSTOMER'S PRODUCT HAS BEEN REQUESTED BACK FOR AN INVESTIGATION. A F/U REPORT WILL BE FILED ONCE INVESTIGATION RESULTS ARE AVAILABLE. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE ADC FA16MAY2006 LETTER. >< A CUSTOMER REPORTED THAT THE UNIT OF MEASUREMENT SETTING OF THEIR FREESTYLE BLOOD GLUCOSE MONITOR CHANGED FROM MG/DL TO MMOL/L. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1577200 | A CUSTOMER REPORTED RECEIVING AN ERROR 6 DISPLAYING ON THEIR BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE DEVICE MANUFACTURE DATE IS UNKNOWN. | 0
1577201 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE DEVICE MANUFACTURE DATE IS UNKNOWN. >< A CUSTOMER REPORTED RECEIVING AN ERROR MESSAGE ON THE DISPLAY OF THEIR BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 0
1808431 | A CUSTOMER REPORTED INCORRECT TIME/DATE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER.THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE TEST STRIP LOT REPORTED BY THE CUSTOMER IS EXPIRED. EXPIRATION DATE WAS NOV 2008. | 0
1817390 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1437649 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< (B) (4). THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
1629701 | IT WAS REPORTED THAT THE PT EXPERIENCED A HYPOGLYCEMIC EVENT AND SEIZURE. >< THE DEVICE IS NOT BEING RETURNED TO ANIMAS FOR EVAL. IT WAS REPORTED THAT THE PT DELIVERED MULTIPLE BOLUSES IN A SHORT PERIOD OF TIME IN ORDER TO TREAT ELEVATED BG LEVELS. THE PT TREATED FOR LOW BG BY INGESTING GLUCOSE TABS. THE PUMP CONTAINS A FEATURE WHICH ALLOWS THE USER TO SET DELIVERY LIMITS. SINCE SETTING THESE LIMITS ON THE PUMP, THE PT HAS CONTINUED TO USE THE PUMP WITHOUT FURTHER SUBSEQUENT EVENTS. | 0
1369310 | LABELING ADEQUATELY WARNS OF EYE CONTACT. RARE BUT KNOWN REACTION. LAB CONTACTED AND FOLLOW UP MADE TO MONITOR PATIENT PROGRESS. PATIENT HAS NOT COMMUNICATED WITH LAB. CURRENT CONDITION UNKNOWN FOR SURE, BUT LIKELY RECOVERED WITH NO FURTHER ACTION REQUIRED. >< FEMALE PATIENT UNDERWENT SLEEP STUDY. WOKE UP WITH RED EYE. TEN20 PASTE GOT INTO THE EYE. ACCORDING TO TECHNOLOGIST, PATIENT REFUSED AN 'EMERGENCY EYE WASH'. PATIENT HAS BLURRED VISION IN THE ONE EYE, HAS A SCRATCHED CORNEA AND IS TAKING ANTIBIOTICS AND USING PRESCRIPTION EYE DROPS. PATIENT HAS SEEN A PHYSICIAN, BUT HAS NOT RELEASE MEDICAL INFORMATION TO THE SLEEP LAB. PATIENT HAS MISSED SOME DAYS OF WORK AND IS SAID TO BE WEARING AN EYE PATCH. UPON A SECOND AND THIRD COMMUNICATION WITH LAB, THEY HAVE HEARD NO MORE FROM THE PATIENT AND THEY ASSUME THE PATIENT HAS RECOVERED. | 0
1879600 | PATIENT SUSTAINED A BURN TO HIS RIGHT INDEX FINGER FROM THE FINGER PROBE OF TELEMETRY/PULSE OXIMETER UNIT. FINGER WAS NOTED TO BE 'COVERED IN SOOT' WHERE THE PROBE WAS LOCATED. PATIENT INDICATED IT WAS 'SMOKING AND SHOCKED HIM.' | 0
1548201 | DEVICE EVALUATION: CUSTOMER REPORT WAS CONFIRMED. KIT APPEARED NOT USED. PROTECTION CAP AT DISTAL END WAS LOST, COMPARED WITH ITS DRAWING. THE CAP WAS NOT RETURNED. >< THE FOLLOWING WAS REPORTED: 'FO CAP AT DISTAL SIDE WAS FOUND DETACHED WHEN PACKAGE WAS OPENED.' | 0
1674047 | THE DEVICE HAS BEEN RECEIVED AND WILL BE PROVIDING A FOLLOW-UP REPORTED WHEN OUR INVESTIGATION IS COMPLETE. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1894727 | NURSE SCANS A MEDICATION THAT IS NOT ON PATIENT PROFILE. SCANNER BEEPS WITH THE SAME AUDIBLE TONE AS A CORRECT MED, BUT A MESSAGE POPS UP ON THE MEDITECH BMV (BEDSIDE MEDICATION VERIFICATION) SCREEN ON THE PC WITH AN 'OK' BUTTON THAT IS HIGHLIGHTED. THE NURSE CONTINUES TO SCAN THE NEXT MED WITHOUT INTERACTING WITH THE SCREEN AND THE ERROR MESSAGE IS CLEARED. WHAT SHOULD HAPPEN IS THAT A DIFFERENT AUDIBLE TONE WOULD SOUND THAT WOULD ALERT THE NURSE AND THE ERROR WOULD NOT CLEAR UNLESS THE NURSE INTERACTED WITH THE SCREEN, SELECTING 'OK' THEN 'EXIT' (F11) - THE ERROR MESSAGE WOULD NOT CLEAR JUST BY THE NURSE SCANNING ANOTHER MED.MEDITECH TESTED SCANNING AND RECEIVED APPROPRIATE ERROR MESSAGES FOR MEDICATION NOT ON PATIENT PROFILE. ALSO REPORTED THAT THERE ISN'T A TWO-WAY INTERFACE BETWEEN MEDITECH AND THE SCANNER. INFORMATION IS SENT FROM THE SCANNER TO MEDITECH AND THERE ISN'T A MEANS FOR MEDITECH TO SEND A RESPONSE BACK TO THE SCANNER TO GENERATE A DIFFERENT TONE BASED ON THE SCANNING ACTIVITY. | 1
1654763 | THE ACCOUNT STATED THAT THE ARCHITECT I2000 ANALYZER GENERATED A FALSE POSITIVE TOXO IGG RESULT OF 26 IU/ML ON A PATIENT KNOWN TO BE NEGATIVE. THE PATIENT WAS TESTED WITH ANOTHER TECHNIQUE IN ANOTHER LAB AND WAS NEGATIVE (NO DATA PROVIDED). THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1584339 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.014/0.011/0.061 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1772818 | THE CUSTOMER STATED FALSELY ELEVATED RESULTS WERE GENERATED ON THE ARCHITECT STAT TROPONIN-I ASSAY. A PATIENT GENERATED A RESULT OF 0.047 NG/ML AND UPON REPEAT, THE RESULT WAS 0.006 NG/ML. THE CUSTOMER IS USING A CUT-OFF OF 0.028 NG/ML. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT DUE TO THIS ISSUE. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1772816 | THE CUSTOMER STATED FALSELY ELEVATED RESULTS WERE GENERATED ON THE ARCHITECT STAT TROPONIN-I ASSAY. A PATIENT PRESENTED TO THE ER WITH DIZZINESS AND HIGH BLOOD PRESSURE GENERATED A TROPONIN-I RESULT OF 0.255 NG/ML ON THE ARCHITECT I2000SR ANALYZER. THE PHYSICIAN ORDERED A CK-MB TEST ON THE SAMPLE AND ANOTHER TROPONIN WAS ALSO RUN INADVERTENTLY WITH A RESULT OF 0.013 NG/ML. THE CUSTOMER TESTED THE SAMPLE A THIRD TIME, YIELDING A TROPONIN-I RESULT OF 0.005 NG/ML. THE FIRST RESULT HAD BEEN REPORTED, BUT A CORRECTED REPORT OF <0.021 NG/ML WAS ISSUED. THE CUSTOMER'S CUT-OFF IS 0.028 NG/ML. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED DUE TO THIS ISSUE. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1598981 | (B) (4). THE CUSTOMER REPLACED AND CALIBRATED THE ARCHITECT STAT PROBE. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE. THE ARCHITECT SYSTEM OPERATIONS MANUAL ((B) (4)) MAY, 2008) PROVIDES INFORMATION REGARDING OPERATIONAL PRECAUTIONS AND LIMITATIONS AND HAZARDS - OPERATOR RESPONSIBILITY. THE CUSTOMER WAS RE-TRAINED ON THE NEED TO USE PERSONAL PROTECTIVE EQUIPMENT AS PER THE SYSTEM OPERATIONS MANUAL. BASED ON THE AVAILABLE INFORMATION, A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. >< THE CUSTOMER STATES THAT WHILE TROUBLESHOOTING A SAMPLE PIPETTING ERROR ON THE ARCHITECT I2000SR ANALYZER THE FLUSH FLUIDS PROCEDURE WAS RECOMMENDED BY THE ABBOTT CUSTOMER TECHNICAL ADVOCATE. DURING THIS PROCEDURE, THE STAT PROBE CRASHED AGAINST THE WASH STATION AND BECAME BENT, SPRAYING ARCHITECT WASH BUFFER INTO THE CUSTOMER'S FACE AND EYES. THE CUSTOMER WAS NOT WEARING ANY PERSONAL PROTECTIVE EQUIPMENT EXCEPT FOR A LAB COAT. THE CUSTOMER RINSED THEIR EYES WITH WATER AND NO MEDICAL ATTENTION WAS SOUGHT. THE CUSTOMER CONTINUED WORKING THAT DAY. THERE WAS NO REPORTED IMPACT TO PATIENT MANAGEMENT DUE TO THIS ISSUE. | 0
1612923 | (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETED. >< THE ACCOUNT GENERATED INCONSISTENT ARCHITECT ESTRADIOL RESULTS FOR A PATIENT. THE INCONSISTENT ARCHITECT ESTRADIOL RESULTS OCCURRED BETWEEN NEAT VERSUS DILUTED PROCESSING. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. DATA PROVIDED: PATIENT 2 ARCHITECT ESTRADIOL = 500 PG/ML (NEAT), 500 PG/ML (AUTODILUTION), >1000 (MANUAL DILUTION). | 0
1584343 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< (B)(4). THE INVESTIGATION BEGAN WITH A REVIEW OF THE SYSTEM LOGS FOR THE ARCHITECT I2000SR ANALYZER. THE REVIEW WAS NOT ABLE TO IDENTIFY A CAUSE FOR THE DISCREPANT RESULTS. THE COMPLAINT TEXT INDICATES THAT THE CUSTOMER IS NOW CENTRIFUGING ALL LITHIUM HEPARIN SAMPLES FOR 15 MINUTES, BUT THE CENTRIFUGE SPEED IS NOT INDICATED. THE CUSTOMER ALSO USES TWO DIFFERENT SIZED CENTRIFUGES FOR PROCESSING SAMPLES. THE CUSTOMER ALSO RETURNED PRESSURE MONITORING (PM) AND LIQUID LEVEL SENSE (LLS) LOGS BUT THESE LOGS ONLY CONTAINED INFORMATION FROM THE CLINICAL CHEMISTRY ANALYZER AND NOT THE IMMUNOASSAY ANALYZER, WHICH IS WHERE THE ISSUE OCCURRED. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-106) (B)(6), 2009 AND THE ARCHITECT TROPONIN ASSAY PACKAGE INSERT ((B)(4)) CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. NO ADVERSE TRENDS WERE IDENTIFIED RELATED TO THE ISSUE, AND A REVIEW OF THE INSTRUMENT'S SERVICE HISTORY DID NOT DETECT ANY REPEATS OF THE ISSUE. SAMPLE COLLECTION AND/OR PREPARATION WERE NOT ABLE TO BE RULED OUT AS POTENTIAL CAUSES OF THE ISSUE THE CUSTOMER OBSERVED. BASED ON THE AVAILABLE INFORMATION, A MALFUNCTION OF THE SYSTEM WAS NOT IDENTIFIED. >< THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.018/0.015/0.105/0.197/0.162/0.032/0.000/0.002 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1584342 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER: AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.000/0.000/0.304/0.000/0.077/0.009/ 0.163/0.005 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1478249 | THE CUSTOMER STATES THAT THEY ARE OBSERVING REPRODUCIBILITY ISSUES WHEN PROCESSING PATIENT SPECIMENS USING THE ARCHITECT I1000 ANALYZER WITH THE TROPONIN-I ASSAY. THE INITIAL RESULTS APPEAR TO BE FALSELY ELEVATED COMPARED TO RESULTS OBTAINED UPON REPEAT TESTING. THE CUSTOMER PROVIDED THE FOLLOWING PATIENT DATA. RESULTS (NG/ML) INITIAL RESULT: 0.93 RETEST RESULTS: 0.01 & 0.01. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. >< (B)(4). INVESTIGATION SUMMARY: AN EVALUATION WAS PERFORMED IN RESPONSE TO TROPONIN REPRODUCIBILITY ISSUES ON AN ARCHITECT I1000SR ANALYZER. THE EVALUATION CONSISTED OF A REVIEW OF THE COMPLAINT TEXT, INSTRUMENT SERVICE HISTORY, AND A REVIEW OF THE ARCHITECT SYSTEM OPERATIONS MANUAL. THE ABBOTT CUSTOMER SUPPORT SPECIALIST (CSS) ASSISTED THE CUSTOMER WITH REPLACING THE WASH ZONE PROBES. THE ARCHITECT OPERATIONS MANUAL PROVIDES MULTIPLE PROBABLE CAUSES, WHICH INCLUDE WASH ZONE PROBES ARE BENT OR DAMAGED, AND CORRECTIVE ACTIONS TO ASSIST WITH RESOLVING ERRATIC RESULTS. A MALFUNCTION OF THE SYSTEM WAS IDENTIFIED. BASED ON A REVIEW OF THE INSTRUMENT'S SERVICE HISTORY RECORD, SINCE THE CUSTOMER REPLACED ALL 3 WASH ZONE PROBES, NO REPEATS OF THE ISSUE HAVE BEEN REPORTED. NO ADVERSE TRENDS OR PRODUCT DEFICIENCY WERE IDENTIFIED RELATED TO THE ISSUE. THIS IS THE FINAL REPORT. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1871748 | (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). THE ARCHITECT TOTAL B-HCG PACKAGE INSERT (49-3364\R3; JUNE 2010) AND THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-108) CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. NO PRODUCT DEFICIENCY OF THE ARCHITECT I2000SR INSTRUMENT WAS FOUND. BASED UPON THE DATA AVAILABLE AND THE RESULTS OF THIS EVALUATION OF THE ARCHITECT I2000SR SYSTEM, A SINGLE DEFINITIVE CAUSE FOR THE CUSTOMER'S ISSUE COULD NOT BE DETERMINED. REVIEW OF COMPLAINT TRACKING AND TRENDING. >< (B)(4). AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSE) VISITED THE CUSTOMER SITE AND REPLACED THE V1/V3 VALVES ON THE TRIGGER/PRETRIGGER MANIFOLD WITH NEW 104 VALVES. THE SAMPLE PROBE IS LESS THAN THREE MONTHS OLD AND THE WASH ZONE WAS NOT LEAKING. THE EVALUATION FAILED TO DETERMINE A SINGLE DEFINITIVE CAUSE OF THE CUSTOMER'S CURRENT ISSUE. REVIEW OF DOCUMENTATION IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. A DEFICIENCY WAS IDENTIFIED, POSSIBLY DUE TO SAMPLE OR REAGENT CARRY-OVER. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. >< THE CUSTOMER STATES THAT ONE FEMALE PATIENT GENERATED AN ARCHITECT TOTAL B-HCG ASSAY RESULT OF 450 MIU/ML THAT WAS REPORTED FROM THE LAB. THE FOLLOWING DAY A SECOND SAMPLE WAS COLLECTED FROM THIS SAME PATIENT AND GENERATED AN ARCHITECT TOTAL B-HCG ASSAY RESULT OF LESS THAN 2 MIU/ML. BOTH SAMPLES FROM THIS PATIENT WERE THEN RETESTED AND BOTH GENERATED RESULTS OF LESS THAN 2 MIU/ML. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1612940 | THE ACCOUNT GENERATED INCONSISTENT ARCHITECT ESTRADIOL RESULTS FOR A PATIENT. THE INCONSISTENT ARCHITECT ESTRADIOL RESULTS OCCURRED BETWEEN NEAT VERSUS DILUTED PROCESSING. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. DATA PROVIDED: PATIENT 1 - RECEIVING TREATMENT FOR OVARIAN STIMULATION. (B)(6) 2010, ARCHITECT ESTRADIOL = >1000 PG/ML (NEAT), 2434 PG/ML (AUTODILUTION). (B)(6) 2010, ARCHITECT ESTRADIOL = 977 PG/ML (NEAT), 2800 PG/ML(1:10 MANUAL DILUTION), 1500 PG/ML (1:5 AUTODILUTION). (B)(6) 2010, ARCHITECT ESTRADIOL = >1000 PG/ML (NEAT), 2212 PG/ML (AUTODILUTION). >< (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETED. | 0
1836534 | (B)(4), EVALUATION - (THE FIELD SERVICE REPRESENTATIVE ( FSR) DID NOT EXAMINE THE DAMAGED UPS AS IT WAS REMOVED FROM THE LABORATORY). - (THE FSR DID NOT DISCOVER ANY ISSUE WITH THE AXSYM ANALYZER). THE ONEAC UPS HAD BEEN INSTALLED WITH THE AXSYM 1 YEAR AGO, AND THE UPS WAS NOT NEW WHEN IT WAS INSTALLED. A FIELD SERVICE REPRESENTATIVE (FSR) REPLACED THE UPS, AND PERFORMED A SUCCESSFUL VOLTAGE CHECK ON THE AXSYM. THE FSR PERFORMED AN AXSYM START UP SUCCESSFULLY, AND THE CUSTOMER RAN ASSAYS TO VERIFY THE PERFORMANCE OF THE AXSYM. THE FSR DID NOT EXAMINE THE UPS AS IT WAS REMOVED FROM THE LAB BY THE BIO-MEDICAL DEPARTMENT. ON 09/01/10, THE TECHNICAL APPLICATION SPECIALIST (TAS) ASKED THE CUSTOMER TO TAKE PHOTOS OF THE DAMAGED UPS AND FORWARD THEM TO ABBOTT. THE PHOTOS HAVE NOT BEEN RECEIVED AT ABBOTT. THE ABBOTT AXSYM SYSTEM OPERATIONS MANUAL CONTAINS ADEQUATE INFORMATION ON THE AXSYM ANALYZER POTENTIAL HAZARDS, OPERATIONAL PRECAUTIONS AND LIMITATIONS. THE MANUAL NOTES THE AXSYM SYSTEM POSES NO UNCOMMON ELECTRICAL SHOCK HAZARD TO OPERATORS, AND BASIC ELECTRICAL HAZARD AWARENESS IS ESSENTIAL TO THE SAFE OPERATION OF ANY SYSTEM. THE ABBOTT SERVICE AND SUPPORT MANUAL CONTAIN ELECTRICAL SPECIFICATIONS FOR THE AXSYM ANALYZER. A SERVICE HISTORY REVIEW FOUND NO ADDITIONAL CUSTOMER COMPLAINTS HAVE BEEN INITIATED ON AXSYM (B)(4) SINCE THE FSR SERVICED THE ANALYZER, AND REPLACED THE DAMAGED UPS. THE SERVICE REVIEW ALSO FOUND NO ISSUES WITH THE AXSYM HAD BEEN REPORTED BY THE CUSTOMER PRIOR TO THE UPS FIRE. THE AUGUST 2010 ABBOTT METRICS FOR THE AXSYM ANALYZER DID NOT IDENTIFY ANY ADVERSE TRENDS DUE TO THE ISSUE BEING EVALUATED. THE SAFETY HAZARDS MEMO CONTAINS INFORMATION NOTING ABBOTT DIAGNOSTIC EQUIPMENT AND ACCESSORIES ARE CERTIFIED TO THE APPROPRIATE SAFETY STANDARDS, AND ADEQUATE PROTECTION IS PROVIDED FOR THE OPERATOR AGAINST ELECTRICAL SHOCK OR BURN, EXCESSIVE TEMPERATURE, AND SPREAD OF FIRE FROM THE EQUIPMENT. BASED ON THE AVAILABLE INFORMATION AND THIS EVALUATION, NO DEFICIENCY WAS IDENTIFIED FOR THE AXSYM ANALYZER LIST NUMBER 07A83-95 FOR THE ISSUE UNDER EVALUATION. >< THE CUSTOMER STATED THAT AN ONEAC UNINTERRUPTABLE POWER SYSTEM (UPS) USED WITH THE AXSYM ANALYZER CAUGHT ON FIRE AT 1:40 AM WHEN NO ONE WAS IN THE LABORATORY. THE FIRE COMPANY PUT OUT THE UPS FIRE USING FOAM SPRAY. THE HOSPITAL DID NOT NEED TO BE EVACUATED, AND NO INJURIES OR BURNS OCCURRED. THE CUSTOMER REQUESTED FIELD SERVICE TO VERIFY IF THE AXSYM WAS AFFECTED. THE CUSTOMER STATED THE AXSYM AND UPS WERE BOTH FUNCTIONING FINE PRIOR TO THE FIRE. | 0
1584312 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.026/0.015/0.017/0.023/0.041 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1564714 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATED A PREGNANT PATIENT THAT PREVIOUSLY TESTED NEGATIVE FOR CMV IGM ON ARCHITECT (INDEX VALUE OF 0.67) AND VIDAS, GENERATED A LOW REACTIVE RESULT (INDEX VALUE OF 1) ON THE ARCHITECT I2000SR ANALYZER WITH THE USE OF REAGENT LOT NUMBER 82373LF00. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< (B)(4) ERRATIC CMV IGM RESULTS. THE INSTRUMENT LOGS WERE NOT AVAILABLE FOR RETURN. SCHEDULED MAINTENANCE WAS DUE ON (B)(6), 2009. AFTER MAINTENANCE PROCEDURES WERE PERFORMED ON (B)(6), 2009, THE CUSTOMER ISSUE WAS RESOLVED. THE CURRENT RATE FOR THE ARCHITECT I2000SR ERRATIC OCCURRENCES/MILLION TESTS AND COMPLAINTS/MILLION TESTS ARE BELOW THE INTERNAL RATE DOCUMENTED AT THE LAUNCH FOR THE ARCHITECT I2000SR. A REVIEW OF QUALITY METRICS DID NOT IDENTIFY ANY ADVERSE TRENDS RELATED TO THIS ISSUE. THE SERVICE HISTORY REVIEW FOUND NO ADDITIONAL INCIDENTS OF THE ARCHITECT I2000SR, (B)(6), GENERATING DISCREPANT RESULTS OR IMPRECISE CONTROLS. ARCHITECT SYSTEM OPERATIONS MANUAL (LIST NUMBER 201837-105) PROVIDES ADEQUATE INFORMATION ABOUT TROUBLESHOOTING FOR ERRATIC RESULTS, OPERATIONAL PRECAUTIONS, AND LIMITATIONS. IN THE ARCHITECT CYTOMEGALOVIRUS (CMV) IGM REAGENT PACKAGE INSERT (56-7560/R1), LITERATURE IS PROVIDED IN DESCRIBING SUITABLE SPECIMENS, RESULTS, AND LIMITATIONS OF THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, NO PRODUCT DEFICIENCY WAS DETERMINED. AFTER MAINTENANCE PROCEDURES WERE PERFORMED, THE CUSTOMER ISSUE WAS RESOLVED. | 0
1814369 | (B)(4). BASED ON THE AVAILABLE INFORMATION, THE CAUSE OF THE LEAKAGE COULD NOT BE DETERMINED. AN EVALUATION WAS PERFORMED IN RESPONSE TO TEMPERATURE ERRORS AND SMOKE BEING OBSERVED ON THE ARCHITECT I2000 ANALYZER SERIAL # (B)(4). THE EVALUATION CONSISTED OF A REVIEW OF COMPLAINT TEXT, A COMPLAINT SEARCH AND A REVIEW OF CURRENT ARCHITECT LABELING. AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED AND INDICATED THAT LEAKAGE WAS VISIBLE EVERYWHERE INSIDE THE INSTRUMENT, INCLUDING, BUT NOT LIMITED TO THE GUTTER. THE FSR WAS UNABLE TO REPAIR THE ANALYZER. THE ARCHITECT I2000 ANALYZER SERIAL # (B)(4) WAS DE-INSTALLED FROM THE CUSTOMER SITE ON (B)(4) 2010. A REPLACEMENT ARCHITECT I2000 ANALYZER SERIAL # (B)(4) WAS INSTALLED AT THE CUSTOMER SITE ON (B)(4) 2010. A REVIEW OF TRENDING FOR THE ARCHITECT ISYSTEMS AND A SEARCH FOR SIMILAR COMPLAINTS FROM (B)(4) 2010 THROUGH (B)(4) 2010, DID NOT IDENTIFY A PRODUCT ISSUE OR AN ADVERSE TREND RELATED TO THE ISSUE IN THIS COMPLAINT. ABBOTT DIAGNOSTIC EQUIPMENT AND ACCESSORIES ARE CERTIFIED TO THE APPROPRIATE SAFETY STANDARDS, AND ADEQUATE PROTECTION IS PROVIDED FOR THE OPERATOR AGAINST ELECTRICAL SHOCK OR BURN, EXCESSIVE TEMPERATURE, AND SPREAD OF FIRE FROM THE EQUIPMENT. THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES ADEQUATE INSTRUCTIONS REGARDING ELECTRICAL HAZARDS AND AN EMERGENCY SHUTDOWN PROCEDURE. BASED ON THE AVAILABLE INFORMATION, A DEFICIENCY IN THE SYSTEM WAS NOT IDENTIFIED. >< THE CUSTOMER STATED THERE IS TEMPERATURE INSTABILITY ON ALL CHANNELS ON THE ARCHITECT I2000 ANALYZER RESULTING IN TEMPERATURE ALARMS (ERROR CODE 7003) BEING GENERATED. IN ADDITION, SMOKE WAS OBSERVED WHILE PERFORMING DAILY MAINTENANCE. THE INSTRUMENT WAS TURNED OFF. NO INJURY WAS REPORTED. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1637346 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT STATED THAT THE ARCHITECT ANALYZER HAS GENERATED ERRATIC RESULTS ON ANOTHER PATIENT SAMPLE. ON (B)(6) 2010, THE INITIAL RESULT WAS 126 PG/ML. THE ACCOUNT THEN RE-CALIBRATED THE ANALYZER AND THE RESULT WAS 94.2 PG/ML. ON (B)(6) 2010, THE SAMPLE WAS TESTED AGAIN AND THE RESULT WAS 89.2 PG/ML. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1481832 | THE CUSTOMER STATED AN EMERGENCY ROOM DOCTOR QUESTIONED AN ELEVATED TROPONIN-I RESULT OF 0.63 NG/ML (CUSTOMER'S NORMAL RANGE 0.0 - 0.03 NG/ML) GENERATED ON THE AXSYM ANALYZER. CONTROLS HAD PREVIOUSLY BEEN WITHIN RANGE EARLIER THAT MORNING, BUT WHEN RETESTED, CONTROL VALUES WERE OUT OF RANGE. THE CUSTOMER REPLACED THE SAMPLING PROBE, REPLACED THE BULK SOLUTIONS, AND USED NEW LOTS OF REAGENT TO TROUBLESHOOT THE CONTROL ISSUE. CONTROLS WERE RUN AND WERE WITHIN RANGE. THE PATIENT SAMPLE WAS RETESTED YIELDING A VALUE OF <0.02 NG/ML AND A CORRECTED REPORT WAS ISSUED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< (B) (4) THE CUSTOMER REPLACED THE SAMPLING PROBE, REPLACED THE BULK SOLUTIONS, AND USED NEW LOTS OF REAGENT TO TROUBLESHOOT THE CONTROLS OUT OF RANGE RESULTS. THE CUSTOMER RAN A SUCCESSFUL FLUIDICS CHECK AND MEIA VERIFICATION. AFTER TROUBLESHOOTING, THE CUSTOMER RAN CONTROLS AND RESULTS WERE WITHIN EXPECTED RANGES. THE CUSTOMER RERAN PATIENT SAMPLES AND FOUND DISCREPANT RESULTS HAD BEEN GENERATED. THE CUSTOMER REQUESTED SERVICE BECAUSE THEY FOUND THE AXSYM WAS LEAKING AND THERE WERE PUDDLES ON THE FLOOR UNDER THE AXSYM. THE FIELD SERVICE REPRESENTATIVE (FSR) FOUND THE SAMPLING PROBE WAS LOOSE AND IN NEED OF CALIBRATION. THE FSR TIGHTENED AND CALIBRATED THE SAMPLING PROBE. THE FSR FOUND THE WASTE DRAIN WAS OVERFLOWING. THE FSR REPLACED THE QUICK DISCONNECT KIT AND THE WASTE BOTTLE ASSEMBLY TO RESOLVE THE LEAKAGE PROBLEM. THE FSR PERFORMED A STARTUP AND FLUSHED FLUIDS TO VERIFY THE LEAKAGE WAS RESOLVED. THE AXSYM SYSTEM OPERATION MANUAL CONTAINS ADEQUATE INFORMATION ON THE PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR INCONSISTENT RESULTS. THE PROBABLE CAUSES INCLUDE DIRTY, DAMAGED, MISALIGNED OR MALFUNCTIONING SAMPLING PROBE. THE TROPONIN-I ADV PACKAGE INSERT WAS FOUND TO CONTAIN ADEQUATE INFORMATION ON SPECIMEN COLLECTION AND PREPARATION, CAUTIONS AND LIMITATIONS OF THE PROCEDURE. THE INSERT NOTES IT IS IMPORTANT TO FOLLOW THE ROUTINE MAINTENANCE PROCEDURES DEFINED IN THE AXSYM SYSTEM OPERATIONS MANUAL FOR OPTIMAL AXSYM PERFORMANCE. A COMPLAINT SERVICE HISTORY REVIEW FOUND NO ADDITIONAL DISCREPANT RESULT COMPLAINTS HAVE BEEN DOCUMENTED FOR AXSYM (B) (4) SINCE THE SAMPLING PROBE WAS INSTALLED CORRECTLY AND CALIBRATED. A REVIEW OF QUALITY METRICS FOR THE AXSYM ANALYZER DID NOT IDENTIFY ANY ADVERSE TRENDS FOR TROPONIN ASSAY DISCREPANT RESULTS GENERATION. THE CURRENT RATE FOR AXSYM ERRATIC OCCURRENCES/ MILLION TESTS AND COMPLAINTS/ MILLION TESTS ARE WITHIN EXPECTED ACTION LIMITS. BASED ON THE AVAILABLE INFORMATION AND THIS EVALUATION, NO DEFICIENCY WAS IDENTIFIED FOR THE AXSYM ANALYZER LIST NO. 07A83-01 RELATED TO THE ISSUE UNDER INVESTIGATION. | 0
1487128 | (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED THERE WAS CORROSION PRESENT ON THE ARCHITECT I2000SR WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. | 0
1487125 | THE CUSTOMER STATED THERE WAS CORROSION PRESENT ON THE ARCHITECT I2000SR WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF (B)(4) AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). | 0
1481802 | (B) (4) THE CUSTOMER REPLACED THE SAMPLING PROBE, REPLACED THE BULK SOLUTIONS, AND USED NEW LOTS OF REAGENT TO TROUBLESHOOT THE CONTROLS OUT OF RANGE RESULTS. THE CUSTOMER RAN A SUCCESSFUL FLUIDICS CHECK AND MEIA VERIFICATION. AFTER TROUBLESHOOTING, THE CUSTOMER RAN CONTROLS AND RESULTS WERE WITHIN EXPECTED RANGES. THE CUSTOMER RERAN PATIENT SAMPLES AND FOUND DISCREPANT RESULTS HAD BEEN GENERATED. THE CUSTOMER REQUESTED SERVICE BECAUSE THEY FOUND THE AXSYM WAS LEAKING AND THERE WERE PUDDLES ON THE FLOOR UNDER THE AXSYM. THE FIELD SERVICE REPRESENTATIVE (FSR) FOUND THE SAMPLING PROBE WAS LOOSE AND IN NEED OF CALIBRATION. THE FSR TIGHTENED AND CALIBRATED THE SAMPLING PROBE. THE FSR FOUND THE WASTE DRAIN WAS OVERFLOWING. THE FSR REPLACED THE QUICK DISCONNECT KIT AND THE WASTE BOTTLE ASSEMBLY TO RESOLVE THE LEAKAGE PROBLEM. THE FSR PERFORMED A STARTUP AND FLUSHED FLUIDS TO VERIFY THE LEAKAGE WAS RESOLVED. THE AXSYM SYSTEM OPERATION MANUAL CONTAINS ADEQUATE INFORMATION ON THE PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR INCONSISTENT RESULTS. THE PROBABLE CAUSES INCLUDE DIRTY, DAMAGED, MISALIGNED OR MALFUNCTIONING SAMPLING PROBE. THE TROPONIN-I ADV PACKAGE INSERT WAS FOUND TO CONTAIN ADEQUATE INFORMATION ON SPECIMEN COLLECTION AND PREPARATION, CAUTIONS AND LIMITATIONS OF THE PROCEDURE. THE INSERT NOTES IT IS IMPORTANT TO FOLLOW THE ROUTINE MAINTENANCE PROCEDURES DEFINED IN THE AXSYM SYSTEM OPERATIONS MANUAL FOR OPTIMAL AXSYM PERFORMANCE. A COMPLAINT SERVICE HISTORY REVIEW FOUND NO ADDITIONAL DISCREPANT RESULT COMPLAINTS HAVE BEEN DOCUMENTED FOR AXSYM (B) (4) SINCE THE SAMPLING PROBE WAS INSTALLED CORRECTLY AND CALIBRATED. A REVIEW OF QUALITY METRICS FOR THE AXSYM ANALYZER DID NOT IDENTIFY ANY ADVERSE TRENDS FOR TROPONIN ASSAY DISCREPANT RESULTS GENERATION. THE CURRENT RATE FOR AXSYM ERRATIC OCCURRENCES/ MILLION TESTS AND COMPLAINTS/ MILLION TESTS ARE WITHIN EXPECTED ACTION LIMITS. BASED ON THE AVAILABLE INFORMATION AND THIS EVALUATION, NO DEFICIENCY WAS IDENTIFIED FOR THE AXSYM ANALYZER LIST NO. 07A83-01 RELATED TO THE ISSUE UNDER INVESTIGATION. >< THE CUSTOMER STATED AN EMERGENCY ROOM DOCTOR QUESTIONED AN ELEVATED TROPONIN-I RESULT OF 0.47 NG/ML (CUSTOMER'S NORMAL RANGE 0.0 - 0.03 NG/ML) GENERATED ON THE AXSYM ANALYZER. CONTROLS HAD PREVIOUSLY BEEN WITHIN RANGE EARLIER THAT MORNING, BUT WHEN RETESTED, CONTROL VALUES WERE OUT OF RANGE. THE CUSTOMER REPLACED THE SAMPLING PROBE, REPLACED THE BULK SOLUTIONS, AND USED NEW LOTS OF REAGENT TO TROUBLESHOOT THE CONTROL ISSUE. CONTROLS WERE RUN AND WERE WITHIN RANGE. THE PATIENT SAMPLE WAS RETESTED YIELDING A VALUE OF <0.02 NG/ML AND A CORRECTED REPORT WAS ISSUED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
1487105 | (B)(4). EVALUATION - UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE FIELD SERVICE ENGINEER STATED THE ARCHITECT I2000SR WASH ZONE GROUND STRAP HAD CORRODED. THERE WAS NO REPORT OF INCIDENT OR INJURY INVOLVED WITH THIS ISSUE. | 0
1487103 | (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED THERE WAS CORROSION PRESENT ON THE ARCHITECT I2000SR WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B)(4). EVALUATION - UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4). IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). | 0
1828001 | (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED SPARKS WERE OBSERVED AT THE POWER SUPPLY SWITCH OF THE AXSYM ANALYZER. THE POWER SUPPLY WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO FIRE OR SMOKE OBSERVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT OR HARM TO LABORATORY PERSONNEL REPORTED. | 0
1546963 | (B)(4). ARCH SAMPLE PROBE LN: 8C94-42; ARCH CA19-9 XR ASSAY LN: 2K91-37 LOT: 75371M200. AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE CUSTOMER SITE AND REPLACED THE SAMPLE AND REAGENT 1 PROBES. THE FSR INDICATED THAT THE CAUSE OF THE DISCREPANT RESULTS WAS POOR PIPETTING FROM THE REAGENT 1 PROBE. THE CUSTOMER RETURNED INSTRUMENT LOGS WERE REVIEWED FOR CONTENT. THE LOG CONTAINS CA 19-9 ASSAY RESULTS FROM THE DATE THE SAMPLE IN QUESTION WAS ANALYZED. THE SAMPLE IDENTIFICATION (SID) WAS NOT REPORTED, THEREFORE IT IS UNKNOWN IF THE SAMPLE RESULTS WITHIN THE RESULTS LOG ARE THE SPECIFIC RESULTS ASSOCIATED WITH THE COMPLAINT CURRENTLY UNDER EVALUATION. THE REVIEW OF THE RETURNED SYSTEM LOGS DID NOT DETERMINE THE CAUSE OF THE ISSUE CURRENTLY UNDER EVALUATION. A REVIEW OF THE COMPLAINT TRACKING DATABASE AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-106) JANUARY, 2009 AND THE ARCHITECT CA 19-9 XR ASSAY PACKAGE INSERT ((B)(4)) PROVIDE INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. BASED ON THE AVAILABLE INFORMATION, A SINGLE DEFINITIVE CAUSE FOR THE INCONSISTENT RESULTS OBSERVED BY THE CUSTOMER COULD NOT BE DETERMINED. A MALFUNCTION IN THE SYSTEM WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATES THAT CA19-9 RESULTS ARE TOO LOW WHEN TESTED WITH THE AUTO-DILUTION MODE ON ONE OF THE ARCHITECT I2000 ANALYZERS IN THE LAB. ONE SAMPLE GENERATED A NEAT RESULT OF >1200 U/ML. THE SAMPLE TESTED AT A VALUE OF 450 U/ML WHEN TESTED WITH THE AUTO-DILUTION MODE (1:10). THE CUSTOMER THEN MADE A MANUAL 1:11 DILUTION OF THE SAMPLE AND GENERATED A FINAL CA19-9 RESULT OF 1600 U/L. THE SAMPLE WAS THEN TESTED ON ANOTHER I2000 ANALYZER WITH ITS AUTO-DILUTION MODE AND A RESULT OF 1400 U/ML WAS GENERATED. CONTROLS WERE WITHIN SPECIFICATIONS FOR ALL RUNS. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. THE CUSTOMER REPLACED THE SAMPLE PROBE BUT THE ISSUE WAS NOT RESOLVED. A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1622260 | (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT A PATIENT SAMPLE GENERATED A FALSELY LOW ARCHITECT TOTAL PSA RESULT. AN INITIAL RESULT OF >100 NG/ML WAS OBTAINED. UPON DILUTION, A RESULT OF 884 NG/ML WAS OBTAINED THAT WAS CONSISTENT WITH THE PATIENT CLINICAL HISTORY. WHEN THE DILUTED SAMPLE WAS REPEATED A RESULT OF <0.8 NG/ML WAS OBTAINED. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. | 0
1595928 | (B)(4). AXSYM CA 125 ASSAY, LN: 7K55-20; AXSYM CA 15-3, LN: 7K56-20. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT FALSELY DECREASED RESULTS HAVE BEEN GENERATED FOR SEVERAL TUMOR MARKER ASSAYS ON THE AXSYM ANALYZER. THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED THE FOLLOWING RESULTS: CEA = 1.9 NG/ML; CA 15-3 = 0.5 U/ML AND CA 125 = 8.6 U/ML. THESE RESULTS WERE QUESTIONED BY THE PATIENT'S PHYSICIAN AS THE PATIENT IN A KNOWN CANCER PATIENT. THE CUSTOMER RETESTED THE SAMPLES FROM THE ORIGINAL SAMPLE TUBE AND GENERATED THE FOLLOWING RESULTS: CEA = 357 NG/ML; CA 15-3 = 470 U/ML AND CA 125 = 1880 U/ML. THE CUSTOMER STATES THAT THE SECOND SET OF RESULTS MATCHES THIS PATIENT'S PREVIOUS RESULTS. THE CUSTOMER THEN RETESTED THE SAMPLES FROM THE SAMPLE CUPS THAT INITIAL RESULTED IN THE LOW VALUES AND ALL RETESTED RESULTS WERE ABOVE THE LINEAR RANGE OF EACH ASSAY. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. A SERVICE CALL WAS INITIATED. | 0
1595915 | (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP MDR WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< THE ACCOUNT STATED THAT THE ARCHITECT I2000SR ANALYZER HAS GENERATED A FALSE POSITIVE B-HCG REAGENT ON A MALE PATIENT. THE ACCOUNT THEN FOUND 3 STATIC SPIKE READINGS THAT DID NOT GENERATE ANY EXCEPTIONS IN THE LOGS. THE INITIAL RESULT WAS POSITIVE AT 535.98 IU/ML. THE SAMPLE WAS RETESTED AND THE RESULT WAS NEGATIVE AT < 1.2 IU/ML. THE QC WAS IN RANGE. ANOTHER SAMPLE WAS DRAWN FROM THE PATIENT AND THE RESULT WAS CONFIRMED TO BE NEGATIVE. THE ACCOUNT IS CONCERNED BECAUSE THE ANALYZER DID NOT GENERATE ERRORS DUE TO THE STATIC READINGS. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1557665 | (B)(4). EVALUATION: ERRATIC TOTAL B-HCG RESULT. 400: V-WHEEL OF PROCESS CAROUSEL. ERRATIC TOTAL B-HCG RESULT. AN INVESTIGATION WAS CONDUCTED IN RESPONSE TO THE COMPLAINT. A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO CHECK THE INSTRUMENT. THE CUSTOMER CONFIRMED THAT THERE WAS NO VISIBLE TRACE OF FIBRIN IN THE SAMPLE COLLECTION TUBES, THE QUALITY CONTROLS WERE ALL WITHIN THE EXPECTED RANGES AND THE PREVENTATIVE MAINTENANCE OF THE ANALYZER WAS UP TO DATE. THE ONLY ERRORS OBSERVED ON THE ANALYZER OCCURRED ON THE VACUTAINER DRY COLLECTION TUBES AND WERE NOT OBSERVED ON OTHER TYPES OF TUBES. THE VACUTAINER GENERATED POOR RESULTS OCCASIONALLY. THE DECONTAMINATION OF THE SAMPLE TUBES WAS PERFORMED BY THE CUSTOMER DURING THE MAINTENANCE PROCEDURE. THE FSR CLEANED THE PROCESS PATH, PROCESS CAROUSEL, PUMP 1,2 AND 3 AND HE ALSO CLEANED THE WASTE CUPS OF THE PROBE AND REPLACED THE V-WHEELS OF THE PROCESS CAROUSEL. AFTER THE INSTRUMENT TROUBLE SHOOTING WAS PERFORMED, THE ANALYZER WAS PERFORMING AS EXPECTED. A REVIEW OF THE COMPLAINT HISTORY FOR THIS ANALYZER (B)(4) OVER A PERIOD OF TWO MONTHS DID NOT IDENTIFY ANY OTHER COMPLAINTS RELATED TO THIS ISSUE. IT WAS MENTIONED IN THE AXSYM SYSTEM OPERATIONS MANUAL THAT DETERMINATION OF THE AXSYM RESULTS MUST BE USED IN CONJUNCTION WITH OTHER CLINICAL DATA (SYMPTOMS, RESULTS OF OTHER TESTS AND PATIENTS HISTORY). THE ASSAY PACKAGE INSERT ADDRESSED THE PROBABLE CAUSES ALONG WITH THE CORRECTIVE ACTIONS FOR THIS ISSUE UNDER THE SECTION OF TROUBLE SHOOTING AND DIAGNOSTICS. BASED ON THE AVAILABLE INFORMATION, THE CAUSE OF THIS ISSUE WAS NOT IDENTIFIED NOR WAS ANY DEFICIENCY IDENTIFIED FOR THIS COMPLAINT. THIS IS A FINAL REPORT. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED THAT THE AXSYM ANALYZER GENERATED DISCREPANT PATIENT RESULTS FROM SAMPLES BELONG TO THREE DIFFERENT PATIENTS. A SAMPLE FROM THE THIRD PATIENT GENERATED AN INITIAL RESULT OF TOTAL PSA OF 6.17 NG/ML THAT WAS REPEATED TWICE AT 3.98 NG/ML AND 3.89 NG/ML. NONE OF THE RESULTS WERE REPORTED OUT OF THE LAB AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1557661 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATED THAT THE AXSYM ANALYZER GENERATED DISCREPANT PATIENT RESULTS FROM SAMPLES BELONGING TO THREE DIFFERENT PATIENTS. A SAMPLE FROM THE FIRST PATIENT GENERATED AN INITIAL B-HCG RESULT OF 36.7 MIU/ML (POSITIVE) WHICH WAS REPEATED TWICE AT LESS THAN 5 MIU/ML (NEGATIVE). NONE OF THE RESULTS WERE REPORTED OUT OF THE LAB AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). EVALUATION: ERRATIC TOTAL B-HCG RESULT. V-WHEEL OF PROCESS CAROUSEL. ERRATIC TOTAL B-HCG RESULT. AN INVESTIGATION WAS CONDUCTED IN RESPONSE TO THE COMPLAINT. A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO CHECK THE INSTRUMENT. THE CUSTOMER CONFIRMED THAT THERE WAS NO VISIBLE TRACE OF FIBRIN IN THE SAMPLE COLLECTION TUBES, THE QUALITY CONTROLS WERE ALL WITHIN THE EXPECTED RANGES AND THE PREVENTATIVE MAINTENANCE OF THE ANALYZER WAS UP TO DATE. THE ONLY ERRORS OBSERVED ON THE ANALYZER OCCURRED ON THE VACUTAINER DRY COLLECTION TUBES AND WERE NOT OBSERVED ON OTHER TYPES OF TUBES. THE VACUTAINER GENERATED POOR RESULTS OCCASIONALLY. THE DECONTAMINATION OF THE SAMPLE TUBES WAS PERFORMED BY THE CUSTOMER DURING THE MAINTENANCE PROCEDURE. THE FSR CLEANED THE PROCESS PATH, PROCESS CAROUSEL, PUMP 1,2 AND 3 AND HE ALSO CLEANED THE WASTE CUPS OF THE PROBE AND REPLACED THE V-WHEELS OF THE PROCESS CAROUSEL. AFTER THE INSTRUMENT TROUBLE SHOOTING WAS PERFORMED, THE ANALYZER WAS PERFORMING AS EXPECTED. A REVIEW OF THE COMPLAINT HISTORY FOR THIS ANALYZER (B)(4) OVER A PERIOD OF TWO MONTHS DID NOT IDENTIFY ANY OTHER COMPLAINTS RELATED TO THIS ISSUE. IT WAS MENTIONED IN THE AXSYM SYSTEM OPERATIONS MANUAL THAT DETERMINATION OF THE AXSYM RESULTS MUST BE USED IN CONJUNCTION WITH OTHER CLINICAL DATA (SYMPTOMS, RESULTS OF OTHER TESTS AND PATIENTS HISTORY). THE ASSAY PACKAGE INSERT ADDRESSED THE PROBABLE CAUSES ALONG WITH THE CORRECTIVE ACTIONS FOR THIS ISSUE UNDER THE SECTION OF TROUBLE SHOOTING AND DIAGNOSTICS. BASED ON THE AVAILABLE INFORMATION, THE CAUSE OF THIS ISSUE WAS NOT IDENTIFIED NOR WAS ANY DEFICIENCY IDENTIFIED FOR THIS COMPLAINT. THIS IS A FINAL REPORT. | 0
1462693 | (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF (B)(4) AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). >< WHILE PERFORMING A TECHNICAL SERVICE BULLETIN, THE ABBOTT GLOBAL ENGINEER SERVICE REPRESENTATIVE NOTICED THE ARCHITECT WASH ZONE GROUND STRAPS WERE CORRODED. NO INCIDENTS OR INJURY OCCURRED. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1565210 | (B)(4). THIS IS AN INITIAL REPORT. A FINAL REPORT WILL BE SUBMITTED UPON THE COMPLETION OF THE INVESTIGATION. AN INVESTIGATION IS IN PROCESS. >< (B)(4). AN ABBOTT FIELD SERVICE ENGINEER (FSE) ARRIVED AT THE CUSTOMER SITE AND CONFIRMED THAT SMOKE WAS EMITTED FROM THE SYSTEM CONTROL CENTER POWER SUPPLY. THE POWER SUPPLY WAS REPLACED AND SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE POWER SUPPLY ONLY WAS RETURNED FOR THIS EVALUATION. A VISUAL INSPECTION FOR DAMAGE WAS PERFORMED ON THE POWER SUPPLY. A BURNT SMELL LINGERED WITH DUST AND SMOKE RESIDUE ACCUMULATED ON THE INSIDE OF THE METAL CASING. IT WAS DETERMINED THAT THE COUNT SWITCHER ASSEMBLY SHORTED ON THE CIRCUIT BOARD OF THE POWER SUPPLY CAUSING THE EMISSION OF SMOKE. THE ARCHITECT I2000SR IS CERTIFIED TO THE APPROPRIATE US, (B)(4), AND INTERNATIONAL SAFETY STANDARDS THAT APPLY TO ELECTRICAL EQUIPMENT FOR MEASUREMENT, CONTROL, AND LABORATORY USE. IN ORDER TO MINIMIZE THE RISK OF SPREAD OF FIRE, THE POWER SUPPLY IS ENCLOSED IN A TOOL ACCESSIBLE METAL CAGE. THE POWER SUPPLY MEETS THE REQUIREMENTS OF THE SAFETY STANDARDS. A REVIEW FOR SIMILAR COMPLAINTS FOR THE PAST SIX MONTHS DID NOT RETURN ANY OTHER COMPLAINTS WHICH WERE SIMILAR TO THE ISSUE DESCRIBED IN THIS COMPLAINT FOR THE SCC POWER SUPPLY. THE ARCHITECT SYSTEM OPERATIONS MANUAL ((B)(4) 2008) PROVIDES INFORMATION IN REGARDS TO POTENTIAL SAFETY ISSUES AND ELECTRICAL HAZARDS. A MALFUNCTION OF THE SYSTEM WAS NOT IDENTIFIED. THE DEVICE WAS PERFORMING AS INTENDED BY CONTAINING THE DAMAGE TO THE POWER SUPPLY AND NOT SPREADING THE ISSUE TO THE REST OF THE ANALYZER OR THE SURROUNDING LABORATORY ENVIRONMENT. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT DURING A NORMAL SAMPLE RUN ON THE ARCHITECT I2000SR ANALYZER, A LOUD NOISE ACCOMPANIED BY A LITTLE BIT OF SMOKE CAME FROM THE SYSTEM CONTROL CENTER (SCC). THE MONITOR WENT BLANK AND THE SCC FAILED TO RESPOND. A SERVICE CALL WAS INITIATED. THERE WERE NO INJURIES REPORTED NOR DAMAGE TO THE SURROUNDING LABORATORY ENVIRONMENT. | 0
1575491 | THE ACCOUNT STATED THAT A FALSE POSITIVE B-HCG RESULT WAS GENERATED BY THE ARCHITECT ANALYZER. AN INITIAL RESULT OF 1017 MIU/ML (POSITIVE) WAS GENERATED AND REPORTED. THE PHYSICIAN QUESTIONED THE RESULT AS THE PATIENT WAS NOT PREGNANT. A NEW SAMPLE WAS OBTAINED AND A B-HCG RESULT OF < 1.20 MIU/ML (NEGATIVE) WAS GENERATED. THE ORIGINAL SAMPLE WAS REPEATED AND A RESULT OF 5.58 MIU/ML WAS GENERATED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. >< (B) (4):THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR DISCOVERED CRYSTALLIZATION ON REAGENT 1 AND 2 PROBES. THE FSR REPLACED BOTH PROBES AND PERFORMED A REPRODUCIBILITY TEST. THE ANALYZER WAS DETERMINED TO BE PERFORMING WITHIN SPECIFICATIONS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE ARCHITECT (B)(4) PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. THIS IS THE FINAL REPORT. | 0
1579881 | (B)(4). ARCH STAT TROPONIN-I LN:2K41 LOT: 34609UN09 EXPIRES: 31 AUGUST 2010. TO INVESTIGATE THE CURRENT CUSTOMER ISSUE, A REVIEW WAS PERFORMED OF THE COMPLAINT TEXT, INSTRUMENT HISTORY, AND THE ARCHITECT SYSTEM LABELING. ERRONEOUS RESULTS GENERATED FOR ONE OF THE SAMPLES IN QUESTION MAY HAVE BEEN CAUSED BY A HIGH FIBRIN CONTENT OBSERVED IN THE SAMPLE. THE CUSTOMER WAS NOT FOLLOWING THE ARCHITECT OPERATIONS MANUAL INSTRUCTIONS ON USING 0.5% HYPOCHLORITE SOLUTION FOR THE DAILY MAINTENANCE PROCEDURE. THE CUSTOMER WAS USING A 1% SOLUTION INSTEAD. A REVIEW OF THE SERVICE HISTORY FOR THE ARCHITECT I2000SR ANALYZER, (B)(4), ALONG WITH THE COMPLAINT DATABASES IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (201837-105) AND THE STAT TROPONIN-I (TNI II) REAGENT PACKAGE INSERT (840638/R7) BOTH CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S ISSUE. BASED ON THE AVAILABLE INFORMATION FOR THIS CURRENT INVESTIGATION, A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. >< (B)(4). ARCH STAT TROPONIN-I: LN:2K41, LOT: 34609UN09, EXPIRES: 31 AUGUST 2010. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT ONE PATIENT GENERATED THE FOLLOWING ARCHITECT STAT TROPONIN-I ASSAY RESULTS: (B)(6) 2010: 13:40, 0.05 NMOL/L; (B)(6) 2010: 18:00, 0.06 NMOL/L; (B)(6) 2010: 00.05, 0.01 NMOL/L (ISTAT METHODOLOGY); (B)(6) 2010: 0.13 NMOL/L; ISTAT RESULTS:0.01 / 0.00 NMOL/L. THE (B)(4) 2010 SAMPLE WAS RECENTRIFUGED AND RETESTED AND GENERATED AN ARCHITECT STAT TROPONIN-I ASSAY RESULT OF 0.05 NMOL/L. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. CONTROLS HAVE REMAINED WITHIN SPECIFICATIONS. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1579882 | THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED AN ARCHITECT STAT TROPONIN-I ASSAY FALSE POSITIVE RESULT OF 0.12 NMOL/L. THE SAMPLE RETESTED AT LESS THAN 0.03 NMOL/L. THE SAMPLE WAS AGAIN RETESTED WITH RESULTS OF 0.03 AND 0.04 NMOL/L. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. CONTROLS HAVE REMAINED WITHIN SPECIFICATIONS. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). ARCH STAT TROPONIN-I LN:2K41-38 LOT: 34609UN09 EXPIRES: 31 AUGUST 2010.THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< (B)(4). ARCH STAT TROPONIN-I LN:2K41 LOT: 34609UN09 EXPIRES: 31 AUGUST 2010. TO INVESTIGATE THE CURRENT CUSTOMER ISSUE, A REVIEW WAS PERFORMED OF THE COMPLAINT TEXT, INSTRUMENT HISTORY, AND THE ARCHITECT SYSTEM LABELING. ERRONEOUS RESULTS GENERATED FOR ONE OF THE SAMPLES IN QUESTION MAY HAVE BEEN CAUSED BY A HIGH FIBRIN CONTENT OBSERVED IN THE SAMPLE. THE CUSTOMER WAS NOT FOLLOWING THE ARCHITECT OPERATIONS MANUAL INSTRUCTIONS ON USING 0.5% HYPOCHLORITE SOLUTION FOR THE DAILY MAINTENANCE PROCEDURE. THE CUSTOMER WAS USING A 1% SOLUTION INSTEAD. A REVIEW OF THE SERVICE HISTORY FOR THE ARCHITECT I2000SR ANALYZER, SERIAL NUMBER (B)(4), ALONG WITH THE COMPLAINT DATABASES IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (201837-105) AND THE STAT TROPONIN-I (TNI II) REAGENT PACKAGE INSERT ((B)(4)) BOTH CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S ISSUE. BASED ON THE AVAILABLE INFORMATION FOR THIS CURRENT INVESTIGATION, A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. | 0
1462651 | THE FIELD SERVICE REPRESENTATIVE WAS PERFORMING A TECHNICAL SERVICE BULLETIN AND FOUND THE ARCHITECT WASH ZONE GROUND STRAPS WERE CORRODED. NO INCIDENTS OR INJURY OCCURRED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(6) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A STAINLESS STEEL ALLOY. A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW STAINLESS STEEL STRAP. IN ADDITION TO THE IMPLEMENTATION OF A STAINLESS STEEL WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010, INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT STAINLESS STEEL ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(6). | 0
1462634 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< WHILE REPLACING THE ARCHITECT WASH ZONE GROUND STRAP, THE FIELD SERVICE REPRESENTATIVE NOTICED CORROSION WAS PRESENT ON BOTH WASH ZONE 1 AND WASH ZONE 2 GROUND STRAPS. THE GROUND STRAPS WERE REPLACED WITHOUT INCIDENT OR INJURY. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. (B)(4) WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A STAINLESS STEEL ALLOY. A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW STAINLESS STEEL STRAP. IN ADDITION TO THE IMPLEMENTATION OF A STAINLESS STEEL WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM 31 JUL 2009 TO 30 JULY 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT STAINLESS STEEL ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. | 0
1462649 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE FIELD SERVICE REPRESENTATIVE DISCOVERED THE ARCHITECT WASH ZONE GROUND STRAP WAS CORRODED. THE GROUND STRAP WAS REPLACED WITHOUT INCIDENT OR INJURY. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(6) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A STAINLESS STEEL ALLOY. A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW STAINLESS STEEL STRAP. IN ADDITION TO THE IMPLEMENTATION OF A STAINLESS STEEL WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT STAINLESS STEEL ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(6). | 0
1462617 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< WHILE PERFORMING A TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE FOUND THE ARCHITECT WASH ZONE GROUND STRAP WAS CORRODED. THE GROUND STRAP WAS REPLACED WITHOUT INCIDENT OR INJURY. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF (B)(4) AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). | 0
1584375 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER: AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.020/0.114/0.023/0.018 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1702195 | (B)(4). CORRECTION THE B-HCG ASSAY WAS INADVERTENTLY OMITTED FROM THE TEXT IN THE INITIAL MDR. BASED UPON THE AVAILABLE INFORMATION; A SINGLE DEFINITIVE CAUSE FROM THE INSTRUMENT PERSPECTIVE COULD NOT BE IDENTIFIED. THE COMPLAINT TEXT INDICATES THAT FALSE POSITIVE B-HCG RESULTS WERE GENERATED ON AN ARCHITECT I1000SR ANALYZER, (B)(4). THE CUSTOMER STATED THAT THE SAMPLE PROBE WAS REPLACED EIGHT DAYS PRIOR TO RECEIVING THE DISCREPANT B-HCG RESULTS. THE CUSTOMER ALSO STATED THAT THEY CHANGED AND CALIBRATED THE SAMPLE PROBE, BUT CHANGING THE PROBE DIDN'T RESOLVE THE ISSUE, AS ERRATIC B-HCG RESULTS WERE STILL BEING RECEIVED. FURTHER REVIEW OF THE COMPLAINT TEXT ALSO INDICATED WHEN THE CUSTOMER WAS AT THE END OF THE REAGENT KIT (12 TESTS REMAINING), LOW CONTROL RESULTS WERE RECEIVED, AND THE CUSTOMER SWITCHED TO A NEW REAGENT PACK WITH THE SAME LOT. THE COMPLAINT EVALUATION FOR THE ARCHITECT I1000SR INSTRUMENT WAS PERFORMED BY REVIEWING THE COMPLAINT TEXT AND INSTRUMENT SERVICE HISTORY FOR (B)(4). BASED UPON THE AVAILABLE INFORMATION; A SINGLE DEFINITIVE CAUSE FROM THE INSTRUMENT PERSPECTIVE COULD NOT BE IDENTIFIED. THE INVESTIGATION FOR THE ASSAY WAS ALSO CONDUCTED AND THE RESULTS ARE AS FOLLOWS; DURING THIS INVESTIGATION THE CONTROL AND PANEL VALUES OBTAINED DURING FINAL PRODUCT RELEASE TESTING WERE REVIEWED FOR ARCHITECT TOTAL B-HCG REAGENT KIT, LIST NUMBER 7K78-20, LOT NUMBER 8919JN00. THE REVIEW DEMONSTRATED THAT ALL CONTROL AND PANEL VALUES WERE WITHIN SPECIFICATION AND NO ADVERSE TRENDS WERE OBSERVED. A REVIEW OF THE STABILITY DATA GENERATED ILLUSTRATED THAT ALL CONTROLS AND ASSAYS PARAMETERS WERE WITHIN SPECIFICATIONS FOR EACH OF THE TIME POINTS ASSESSED. IN ADDITION, EXTENSIVE TESTING HAS BEEN PERFORMED ON SITE AT ABBOTT LONGFORD WHEREBY INVESTIGATION PROTOCOLS ADDRESSED THE ABILITY OF ARCHITECT TOTAL B-HCG REAGENT KIT IN QUESTION TO ACCURATELY DETECT B-HCG CONCENTRATIONS IN ABBOTT ARCHITECT TOTAL B-HCG CONTROLS PANELS AND PATIENT SAMPLE. THE INVESTIGATION DEMONSTRATED THAT ARCHITECT TOTAL B-HCG REAGENT LOT 81919JN00 IS PERFORMING ACCEPTABLY AND WITHIN LABEL CLAIMS. FURTHERMORE, ARCHITECT TOTAL B-HCG REAGENT LOT 81919JN00 WAS USED IN (B)(4) TO EVALUATE THE PERFORMANCE OF THE ARCHITECT TOTAL B-HCG ASSAY IN RELATION TO ITS ABILITY TO RELIABLY RECOVER B-HCG IN PATIENT SAMPLES. ALL CONTROL LEVELS WERE WITHIN ABBOTT ARCHITECT TOTAL B-HCG CONTROL SPECIFICATION LIMITS. FROM THE INVESTIGATION PERFORMED IT CAN BE CONCLUDED THAT SAFETY, EFFECTIVENESS AND LABEL CLAIMS WERE MET FOR THE ARCHITECT TOTAL B-HCG REAGENT LOT IN QUESTION. >< (B)(4). A PRODUCT DEFICIENCY HAS BEEN IDENTIFIED. WITH A UNIQUE COMBINATION OF INSTRUMENT AND REAGENT CONDITIONS THERE IS A POTENTIAL FOR SAMPLE CARRYOVER WHICH CAN RESULT IN FALSELY ELEVATED B-HCG CONCENTRATIONS ON THE ARCHITECT I1000SR SYSTEM. NEGATIVE SAMPLES HAVE RETURNED BOTH POSITIVE AND GREY ZONE B-HCG RESULTS WHEN A HIGH CONCENTRATION B-HCG SAMPLE IS ASSESSED PRIOR TO THE NEGATIVE SAMPLE WITH ARCHITECT TOTAL B-HCG (LIST NUMBER 7K78/6C21) ON AN I1000SR INSTRUMENT WHICH ALSO UTILIZES THE ARCHITECT RUBELLA IGG (LIST NUMBER 6C17) ASSAY. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. AN INVESTIGATION IS IN PROCESS. >< THE ACCOUNT STATED THAT THE ARCHITECT I1000SR ANALYZER HAS GENERATED DISCREPANT RESULTS. THE INITIAL RESULT WAS 55.7 MIU/ML, THE SAMPLE WAS RETESTED TWICE AND RESULTS WERE 13.3 AND <0.12 MIU/ML. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE SUSPECT MEDICAL DEVICE HAS CHANGED FROM THE (B)(4). MFR # 3005094123-2011-00589 HAS BEEN SUBMITTED TO ADDRESS THIS ERROR. | 0
1702194 | THE ACCOUNT STATED THAT THE ARCHITECT I2000SR ANALYZER HAS GENERATED A FALSE REACTIVE RESULT. THE INITIAL RESULT WAS 36.0 MIU/ML, THE SAMPLE WAS RETESTED AND THE RESULT WAS <0.12 MIU/ML. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1584399 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.132/0.011/0.012 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1462608 | (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF (B)(4) AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< UPON INSPECTING THE ARCHITECT WASH ZONE GROUND STRAP, THE FIELD SERVICE REPRESENTATIVE OBSERVED CORROSION ON BOTH THE WASH ZONE 1 AND WASH ZONE 2 GROUND STRAPS. THE STRAPS WERE REPLACED WITHOUT INCIDENT OR INJURY. | 0
1584380 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.009/0.006/0.092/0.012/0.256 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1636883 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT STATED THAT THE ARCHITECT ANALYZER HAS GENERATED ERRATIC RESULTS ON A PATIENT SAMPLE. THE INITIAL RESULT WAS 84.8 PG/ML, THE RETESTS RESULTS WERE BOTH 48 PG/ML. THE 48 RESULT WAS REPORTED. THE SAMPLE WAS THEN RETESTED THE FOLLOWING DAY AND THE RESULTS WERE 68.1 AND 68.9 PG/ML. THE SAMPLE WAS RETESTED AGAIN THE VERY NEXT DAY AND THE RESULTS WERE 55.0 AND 56.2 PG/ML. FIELD SERVICE WAS REQUESTED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1584385 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.021/0.076/0.012/0.053 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1585541 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.047/0.073/0.089/0.012/0.038/0.012/0.012/0.416/0.232 NG/ML. THE SAMPLE GENERATED A RESULT OF 0.022 NG/ML ON AN ARCHITECT PLATFORM AT ANOTHER FACILITY. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1477076 | (B)(4). OTHER (B)(4): ERRATIC CYTOMEGALOVIRUS (CMV) IGG, (B)(6), AND RUBELLA IGG RESULTS. INVESTIGATION PLAN SUMMARY: A REVIEW OF THE COMPLAINT TEXT INDICATES THE CUSTOMER OBSERVED RUBELLA IGG AND (B)(6) CONTROLS FAILING. THE CUSTOMER IS ALSO EXPERIENCING LOW REACTIVE CYTOMEGALOVIRUS (CMV) IGG PATIENT RESULTS THAT ARE RETESTED NEGATIVE WHEN SENT FOR CONFIRMATION. EVEN AFTER REPLACING THE SAMPLE PROBE, THE ISSUE PERSISTED. THE CUSTOMER ALSO NOTICED EXCESSIVE CRYSTAL BUILD UP AROUND THE SAMPLE PROBE WASH CUP, WHICH CONTINUOUSLY RETURNED AFTER CLEANING THE AREA. THE FIELD SERVICE ENGINEER (FSE) INSTALLED TECHNICAL SERVICE BULLETIN (TSB) (B)(4) ARCHITECT SOFTWARE V5.10 INSTALLATION ONTO THE ANALYZER ON (B)(6) 2009. HE REPLACED AND CALIBRATED THE REAGENT 1, REAGENT 2, AND SAMPLE PROBES. THE FSE REMOVED THE SYRINGES TO CHECK AND CLEAN ALL THE PIPETTOR AND TUBING CONNECTIONS. NO LEAKS WERE OBSERVED IN THE PUMP BAY AREA. HE TIGHTENED ALL OF THE WASH ZONE, PRETRIGGER, TRIGGER VALVES. THE WASH ZONE 1 AND 2 ASPIRATION, OPTICS BACKGROUND, TRIGGER, AND PRETRIGGER GRAVIMETRIC TESTS WERE PERFORMED. ALL THESE TESTS PASSED SPECIFICATIONS. THE QUALITY CONTROL VALUES WERE WITHIN ACCEPTABLE RANGES. NO SAMPLE CARRYOVER ISSUES WERE OBSERVED DURING THE CARRYOVER STUDY. THE PRECISION OF RUBELLA IGG, (B)(6), AND CMV IGM WERE WITHIN NORMAL RANGES. REVIEW OF COMPLAINT METRICS IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. BASED ON THE AVAILABLE INFORMATION, NO DEFICIENCY IDENTIFIED FOR THIS COMPLAINT. A MALFUNCTION OF THE SYSTEM WAS NOT IDENTIFIED. AFTER THE FIELD SERVICE ENGINEER PERFORMED A MULTIPOINT INSTRUMENT CHECK AND TROUBLESHOOTING OF THE INSTRUMENT, THE CUSTOMER ISSUE HAS NOT REOCCURRED, THE INSTRUMENT HAS BEEN PERFORMING AS INTENDED. THIS IS THE FINAL REPORT. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT PATIENT SAMPLES PROCESSED USING THE ARCHITECT CMV IGM ASSAY ARE GENERATING FALSELY POSITIVE RESULTS. THE CUSTOMER STATES THAT RESULTS ARE NEGATIVE WHEN SAMPLES ARE SENT OUT OF THE LAB FOR CONFIRMATION TESTING. NO SPECIFIC PATIENT DATA WAS PROVIDED. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. | 0
1585544 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.101/0.027/0.029 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1628351 | A FIELD SERVICE ENGINEER (FSR) REPORTED THAT WHILE PERFORMING A SAMPLE PROBE CALIBRATION ON THE AXSYM ANALYZER HE ACCIDENTALLY PLACED HIS HAND IN THE PATH OF THE MOVING SAMPLE PROBE PIERCING THE NAIL OF HIS MIDDLE FINGER. THE FSR WAS WEARING DISPOSABLE GLOVES. A DOCTOR AT A HOSPITAL EMERGENCY ROOM CLEANED THE WOUND AND ADMINISTERED A TETANUS VACCINATION. THE DOCTOR ALSO PRESCRIBED MEDICATION (PILLS) CONTAINING LAMIVUDINE, ZIDAVUDINE, LOPINAVIN AND RITONAVIN. BLOOD WAS DRAWN FOR (POLYMERASE CHAIN REACTION) PCR TESTING AND LIVER FUNCTION. NO PATIENT RESULTS WERE IMPACTED DUE TO THIS ISSUE. >< (B) (4). PRODUCT LABELING WAS REVIEWED AND WAS FOUND TO CONTAIN INFORMATION ON THE CORRECT PROCEDURE FOR PERFORMING A SAMPLING PROBE CALIBRATION, OPERATIONAL PRECAUTIONS, AND HAZARDS OF THE AXSYM SYSTEM. THE AXSYM SYSTEM IS EQUIPPED WITH PROTECTIVE DOORS. THE SYSTEM STATUS MUST BE PAUSED, READY, OR STOPPED WHEN PERFORMING MAINTENANCE PROCEDURES. LABELING WARNS THAT PROBES ARE SHARP AND POTENTIALLY CONTAMINATED WITH INFECTIOUS MATERIALS. AVOID CONTACT WITH THESE COMPONENTS AND HANDLE THEM CAUTIOUSLY IN ORDER TO PREVENT INJURY. NO ADVERSE TRENDS DUE TO OPERATORS BEING INJURED BY SAMPLING PROBES ON THE AXSYM ANALYZER WERE IDENTIFIED. BASED ON THE AVAILABLE INFORMATION, PRODUCT LABELING WAS DETERMINED TO BE ADEQUATE AND NO DEFICIENCY WAS IDENTIFIED RELATED TO THE ISSUE UNDER EVALUATION. | 0
1462809 | WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE NOTICED THE ARCHITECT WASH ZONE GROUND STRAP WAS CORRODED. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION - UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. | 0
1462825 | REPORTEDLY, THE RING WAS EXPLANTED AT IMPLANT DUE TO REGURGITATION. PER THE CER: 'THAT THE RING WAS EXPLANTED AT IMPLANT BECAUSE OF THE REGURGITATION. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE WILL BE NOT RETURNED (DISCARDED).' THE DATE OF EXPLANT IS UNKNOWN; THEREFORE THE AWARE DATE IS BEING USED AS THE OCCURRENCE DATE. INFORMATION WAS LEARNED FROM THE IMPLANT PATIENT REGISTRY. >< (B)(4). THE SELECTION OF SERIOUS INJURY WAS MADE IN ERROR. NO INJURY WAS INVOLVED WITH THIS CUSTOMER ISSUE. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION: UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND. THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE (B)(4). IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT STAINLESS STEEL ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. >< WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE OBSERVED CORROSION ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1462828 | (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION: UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY (B)(4) GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. >< UPON INSPECTING THE ARCHITECT WASH ZONE GROUND STRAP, THE FIELD SERVICE ENGINEER OBSERVED BUILD UP OF 'DEBRIS' THAT APPEARED GREEN IN COLOR AND 'SALT SEPARATION' MATERIAL AROUND THE TOP OF THE GROUND STRAP. THERE WAS NO REPORT OF INJURY OR INCIDENT. | 0
1462822 | WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE OBSERVED CORROSION ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B)(4), EVALUATION: UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY (B)(4) GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4). IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. >< (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1585536 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.017/0.037/0.022/0.023 NG/ML. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1639002 | THE ASSISTANT TECHNICIAN FOR THE DISTRIBUTOR ((B) (4)) WAS VISITING THE CUSTOMER SITE AND ACCIDENTALLY PUNCTURED THE PALM OF HIS LEFT HAND AFTER REMOVING THE PROCESSING PROBE FROM THE AXSYM ANALYZER. HE SOUGHT MEDICAL ATTENTION AND WAS PRESCRIBED ZIDOVUDINE (300 MG) AND LAMIVUDINE (150 MG). HEMOGRAM, UREA, CREATININE, TGO/TGP AND VIRAL MARKER TESTING WAS ALSO PERFORMED. VIRAL MARKER TEST RESULTS WERE NEGATIVE. >< (B) (4). THE AXSYM SYSTEM OPERATION MANUAL WAS REVIEWED AND WAS FOUND TO CONTAIN ADEQUATE INFORMATION ON THE CORRECT PROCEDURE FOR MAINTAINING AND REPLACING PROCESSING PROBES, OPERATIONAL PRECAUTIONS, AND HAZARDS OF THE AXSYM SYSTEM. THE TECHNICIAN WAS RETRAINED TO FOLLOW PROCEDURES ACCORDING TO THE MANUAL. A COMPLAINT REVIEW WAS PERFORMED AND NO ADVERSE TRENDS DUE TO OPERATORS BEING INJURED BY THE AXSYM PROCESSING PROBE WERE IDENTIFIED. BASED ON THE AVAILABLE INFORMATION AND THIS EVALUATION, NO DEFICIENCY WAS IDENTIFIED FOR THE AXSYM ANALYZER LIST NO. 07A83-01 OR THE AXSYM PROBE LIST NO. 09A59-01 RELATED TO THE ISSUE UNDER EVALUATION. | 0
1585539 | THE CUSTOMER STATES THAT FALSE POSITIVE RESULTS ARE BEING GENERATED BY THE ARCHITECT STAT TROPONIN-I ASSAY. ONE PATIENT SAMPLE GENERATED RESULTS OF 0.071/0.283/0.032/0.021/0.027/0.023 NG/ML AND 0.023/0.020 NG/ML ON AN ARCHITECT PLATFORM AT ANOTHER FACILITY. THE CUSTOMER USES A CUT-OFF VALUE OF 0.03 NG/ML. THE SAMPLE GENERATED AN I-STAT RESULT OF 0.1 NG/ML. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB AND THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1462819 | (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE FIELD SERVICE REPRESENTATIVE WAS PERFORMING A TECHNICAL SERVICE BULLETIN AND NOTICED CORROSION WAS PRESENT ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. | 0
1462803 | (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE OBSERVED CORROSION ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO NIOSH, THE NATIONAL INSTITUTE FOR OCCUPATIONAL SAFETY AND HEALTH. | 0
1462806 | (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE NOTICED THERE WAS CORROSION PRESENT ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. >< (B)(4). EVALUATION - UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF STAINLESS STEEL AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF STEEL AND TIN AS REPLACEMENT FOR COPPER. THE STUDY INDICATED THERE ARE NO KNOWN EXPLOSIVE AZIDES FORMED IF STEEL, WITH A SIGNIFICANT AMOUNT OF NICKEL AND CHROMIUM, COMES IN CONTACT WITH THE BUFFER SOLUTION. A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4). IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) WZ MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). | 0
1462800 | (B)(4). UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM CORROSION WHICH CAN FURTHER DEVELOP INTO AN EXPLOSIVE COMPOUND; THE COPPER PRESENT IN THE GROUND STRAP WAS EXPOSED TO SODIUM AZIDE FOR A LONG ENOUGH PERIOD OF TIME TO FORM CORROSION. AN INVESTIGATION WAS CONDUCTED AND CONCLUDED THAT UNDER CERTAIN CONDITIONS, SODIUM AZIDE CAN REACT WITH UNPROTECTED COPPER TO FORM AN EXPLOSIVE COMPOUND. INDEPENDENT STUDIES WERE PERFORMED TO DETERMINE THE CHEMICAL COMPOSITION OF DEPOSITS FOUND ON FIELD RETURNED GROUND STRAPS, POTENTIAL HAZARDS OF COPPER AZIDE DEPOSITS, FORMATION RATES OF COPPER AZIDE, THE FEASIBILITY OF (B)(4) AS A POTENTIAL GROUND STRAP MATERIAL AND OTHER POTENTIAL HAZARDS. ONE STUDY REVEALED THERE ARE OTHER METALS, POTENTIALLY UTILIZED IN INSTRUMENT DESIGNS, THAT MAY FORM AN UNSTABLE AZIDE COMPOUND ARE GOLD, SILVER, COPPER, BRASS AND TIN. THE (B)(4) PERFORMED A STUDY TO DETERMINE THE VIABILITY OF (B)(4) AS REPLACEMENT FOR COPPER. (B)(4). A GLOBAL ASSESSMENT OF SODIUM AZIDE WAS CONDUCTED. OTHER PARTS IDENTIFIED AS CONTAINING COPPER WILL BE REPLACED WITH OTHER MATERIAL WITH NO KNOWN POTENTIAL VOLATILE METAL FORMATIONS. INSTRUMENT DESIGNS WILL AVOID THE USE OF BARE COPPER OR TIN COATED COPPER IN AREAS WHERE SIGNIFICANT EXPOSURE TO SODIUM AZIDE MIGHT BE EXPECTED. CORRECTIVE ACTIONS WERE IMPLEMENTED IN RESPONSE TO THIS ISSUE. THE MATERIAL OF THE GROUND STRAP WAS CHANGED TO A (B)(4). A MANDATORY TECHNICAL SERVICE BULLETIN REQUIRED THAT EVERY COPPER GROUND STRAP BE REPLACED BY FIELD SERVICE TO THE NEW (B)(4) STRAP. IN ADDITION TO THE IMPLEMENTATION OF A (B)(4) MECHANISM GROUND STRAP, INSTRUCTIONS ON THE REMOVAL AND PACKAGING OF POTENTIALLY CORRODED GROUP STRAPS WERE INCORPORATED INTO LABELING. A REVIEW OF COMPLAINT TRACKING AND TRENDING FROM (B)(4) 2009 TO (B)(4) 2010 INDICATED THAT THERE WERE NO COMPLAINTS WITH RESPECT TO CORROSION OR ADVERSE EVENTS IN CONJUNCTIONS WITH THE REPLACEMENT (B)(4) ARCHITECT WASH ZONE MECHANISM GROUND STRAP. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE POTENTIAL HAZARDS INVOLVED WITH PRODUCTS CONTAINING SODIUM AZIDE; HOWEVER, LABELING WAS ENHANCED TO INCLUDE REFERENCE TO (B)(4). >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< WHILE PERFORMING THE TECHNICAL SERVICE BULLETIN, THE FIELD SERVICE REPRESENTATIVE NOTICED CORROSION ON THE ARCHITECT WASH ZONE GROUND STRAP. THERE WAS NO REPORT OF INCIDENT OR INJURY. | 0
1106572 | DURING INTERNAL ROUTINE TESTING, IT WAS FOUND THAT UNDER CERTAIN CIRCUMSTANCES, EDITING TEXT IN THE REPORT MAY PREVENT NEW PHRASES FROM POPULATING THE REPORT. THIS OCCURS WHEN REPORT TEST IS MANUALLY MODIFIED AND ADDITIONAL PHRASES ARE SUBSEQUENTLY ADDED AS A RESULT OF WORKSHEET INTERACTIONS. AS A CONSEQUENCE, SOME PHRASES THAT FOLLOW IN THE SPECIFIC SECTION OF THE MANUALLY EDITED TEXT WILL NOT POPULATE THE REPORT. THIS BEHAVIOR WAS NOT REPORTED FROM A CUSTOMER SITE AND THEREFORE WE ARE UNAWARE OF ANY PATIENT INJURY ASSOCIATED WITH THIS ISSUE. THIS ISSUE HAS THE POTENTIAL TO CREATE A SITUATION WHERE CRITICAL REPORT INFO FOR TREATMENT PLANNING IS NOT AVAILABLE IF THE REPORTING PHYSICIAN DOES NOT VERIFY THE REPORT CONTENT. A CUSTOMER SAFETY ADVISORY NOTICE WAS CREATED AND WILL BE DISTRIBUTED TO ALL AFFECTED CUSTOMERS. FURTHER ON A SOFTWARE PATCH WILL BE MADE AVAILABLE FOR UPGRADE OF THE SOFTWARE VERSION. >< (B) (4). (B) (4). | 0
1732176 | ON (B)(6)2010, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT THEY SMELLED SMOKE FROM AN I-STAT DOWNLOADER RECHARGER AND THE UNIT WAS WARM TO TOUCH. THERE WAS NO REPORT OF ANY INJURY ASSOCIATED WITH THE EVENT. | 0
1630790 | THE CUSTOMER STATED A CELL-DYN RUBY ANALYZER GENERATED MULTIPLE SAMPLING ERRORS AND A FALSELY DECREASED HEMOGLOBIN RESULT FOR ONE PATIENT SAMPLE. THE ANALYZER GENERATED AN INITIAL HEMOGLOBIN RESULT OF 7.1 G/DL. THE PATIENT WAS ADMITTED FOR A BLOOD TRANSFUSION AND A SECOND SAMPLE WAS DRAWN. THE SECOND SAMPLE YIELDED A NORMAL HEMOGLOBIN RESULT BETWEEN 10 G/DL AND 11 G/DL. THE PATIENT WAS NOT TRANSFUSED AND THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1868169 | (B)(4). EVALUATION (B)(4) (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. >< THE CUSTOMER'S CELL-DYN SAPPHIRE ASPIRATION BOTTOM SENSOR WAS PROACTIVELY REPLACED PER FA05APR2010. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. | 0
1611602 | THE CUSTOMER REPORTED A BROKEN WASTE SENSOR OF THE CELL-DYN ANALYZER THAT WAS REPLACED TO RESOLVE THE ISSUE. NO EXPOSURE OR INJURIES WERE REPORTED. NO IMPACT TO USER SAFETY OR PATIENT RESULTS WAS REPORTED. >< (B)(4). OTHER, PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. BASED ON THE INTENDED USE, THIS PART WILL WEAR OUT AND NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. NO PREVENTIVE MAINTENANCE SCHEDULE WAS IN PLACE. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY AND SAFETY ISSUES WERE DETERMINED TO BE WITHIN ESTABLISHED FREQUENCY. A HEALTH HAZARD ASSESSMENT (HHA) DETERMINED A LOW HEALTH RISK, WHICH IS CONSISTENT WITH THE RISK MANAGEMENT FILE ASSOCIATED WITH THE PRODUCT ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXIST A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THE WASTE OUTLET TUBE ASSEMBLY PART LIFE EXPECTANCY OF 2 YEARS IS NOT CURRENTLY IN OUR PRODUCT LABELING. THE PREVENTIVE ACTION WILL INVOLVE A LABEL CHANGE FOR THE WASTE OUTLET TUBE ASSEMBLY. WE ARE CHANGING OUR LABELING TO HAVE A PART REPLACEMENT SCHEDULE OF 6 MONTHS SO PARTS ARE REPLACED PRIOR TO THE END OF THE 2-YEAR LIFE FAILURE. THIS WILL HELP PREVENT OVERFLOW ISSUES CAUSED BY END OF LIFE PART FAILURE. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1611604 | THE CUSTOMER STATED THAT THE CELL-DYN WASTE CONTAINER HAS OVERFLOWED SEVERAL TIMES DUE TO THE WASTE SENSOR NOT DEMONSTRATING THE ALARM MESSAGE FOR A FULL WASTE. NO EXPOSURE OR INJURIES WERE REPORTED. NO IMPACT TO USER SAFETY OR PATIENT RESULTS WAS REPORTED. THE CUSTOMER ORDERED A REPLACEMENT FOR THE WASTE LINE ASSEMBLY PER FA30NOV2009 LETTER SINCE THE WASTE LINE ASSEMBLY (THEY HAVE) HAS BEEN ON SINCE INSTALLATION. >< (B)(4). EVALUATION RESULTS: PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. BASED ON THE INTENDED USE, THIS PART WILL WEAR OUT AND NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. NO PREVENTIVE MAINTENANCE SCHEDULE WAS IN PLACE. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY AND SAFETY ISSUES WERE DETERMINED TO BE WITHIN ESTABLISHED FREQUENCY. A HEALTH HAZARD ASSESSMENT (HHA) DETERMINED A LOW HEALTH RISK, WHICH IS CONSISTENT WITH THE RISK MANAGEMENT FILE ASSOCIATED WITH THE PRODUCT ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THE WASTE OUTLET TUBE ASSEMBLY PART LIFE EXPECTANCY OF 2 YEARS IS NOT CURRENTLY IN OUR PRODUCT LABELING. THE PREVENTIVE ACTION WILL INVOLVE A LABEL CHANGE FOR THE WASTE OUTLET TUBE ASSEMBLY. WE ARE CHANGING OUR LABELING TO HAVE A PART REPLACEMENT SCHEDULE OF 6 MONTHS, SO PARTS ARE REPLACED PRIOR TO THE END OF THE 2-YEAR LIFE FAILURE. THIS WILL HELP PREVENT OVERFLOW ISSUES CAUSED BY END OF LIFE PART FAILURE. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1611605 | (B)(4). RESULTS: PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. BASED ON THE INTENDED USE, THIS PART WILL WEAR OUT AND NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. NO PREVENTIVE MAINTENANCE SCHEDULE WAS IN PLACE. A REVIEW OF THE CELL-DYN SYSTEM OPERATOR'S MANUAL SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY AND SAFETY ISSUES WERE DETERMINED TO BE WITHIN ESTABLISHED FREQUENCY. A HEALTH HAZARD ASSESSMENT (HHA) DETERMINED A LOW HEALTH RISK, WHICH IS CONSISTENT WITH THE RISK MANAGEMENT FILE ASSOCIATED WITH THE PRODUCT ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXIST A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THE WASTE OUTLET TUBE ASSEMBLY PART LIFE EXPECTANCY OF 2 YEARS IS NOT CURRENTLY IN OUR PRODUCT LABELING. THE PREVENTIVE ACTION WILL INVOLVE A LABEL CHANGE FOR THE WASTE OUTLET TUBE ASSEMBLY. WE ARE CHANGING OUR LABELING TO HAVE A PART REPLACEMENT SCHEDULE OF 6 MONTHS SO PARTS ARE REPLACED PRIOR TO THE END OF THE 2-YEAR LIFE FAILURE. THIS WILL HELP PREVENT OVERFLOW ISSUES CAUSED BY END OF LIFE PART FAILURE. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< AFTER RECEIVING THE CUSTOMER LETTER OF FA30NOV2009, THE CUSTOMER REPORTED THAT THE CELL-DYN WASTE SENSOR HAS BEEN FAILING TO DEMONSTRATE THE ALARM OF A FULL WASTE. THE CUSTOMER ORDERED A REPLACEMENT OF THE WASTE LINE ASSEMBLY SINCE THE ONE THEY HAVE IS FIVE YEARS OLD. NO EXPOSURE OR INJURIES WERE REPORTED. NO IMPACT TO USER SAFETY OR PATIENT RESULTS WAS REPORTED. | 0
1778174 | (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE ACCOUNT STATED THAT THE CD3200 SL ANALYZER GENERATED AN INITIAL HEMOGLOBIN (HGB) RESULT OF 4.4 G/DL. THE SAMPLE WAS REPEATED 2 TIMES WITH RESULTS OF 11.8 AND 11.8 G/DL. THE ACCOUNT STATED THAT THE INITIAL RESULT HAD RBC MORPH AND DFLT FLAGS. THE INITIAL RESULT WAS NOT REPORTED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. | 0
1551400 | THE ACCOUNT STATED THAT THE CELL-DYN 3200 ANALYZER'S SENSOR HAS A FAULTY WASTE SENSOR. THE WASTE SENSOR WAS REPLACED. THE ANALYZER IS NOW PERFORMING WITHIN SPECIFICATION. THERE WAS NO DIRECT FLUID CONTACT OR INJURIES REPORTED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1551401 | THE ACCOUNT STATED THAT CELL-DYN 3500 ANALYZER'S SENSOR WAS BROKEN. THE SENSOR WAS REPLACED AND IS NOW WORKING LIKE IT SHOULD. THERE WAS NO DIRECT FLUID CONTACT OR INJURIES REPORTED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1551402 | (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< THE ACCOUNT STATED THAT THE CELL-DYN RUBY ANALYZER IS NOT WORKING PROPERLY. THE ACCOUNT WILL MONITOR THE WASTE TO ASSURE IT WILL NOT OVERFLOW. THERE WAS NO DIRECT FLUID CONTACT OR INJURIES REPORTED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1551403 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, (B)(4), WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< THE ACCOUNT STATED THAT THE CELL-DYN 3700 ANALYZER'S WASTE SENSOR IS BROKEN. A NEW WASTE SENSOR WAS ORDERED. THERE WAS NO DIRECT FLUID CONTACT OR INJURIES REPORTED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1478206 | THE CUSTOMER STATES THAT THE PRINTER ATTACHED TO THE CELL-DYN 3500 ANALYZER IS NOT PRINTING RESULTS AS SHOWN ON THE SCREEN OF THE CD 3500 ANALYZER. THE CUSTOMER STATES THAT SOME LETTERS AND NUMBERS ARE REPLACED WITH DIFFERENT NUMBERS OR CHARACTERS. EXAMPLE: THE #9 IS PRINTING AS THE #1 AND THE #8 IS PRINTING AS THE #0. THE CUSTOMER NOTICED THE ERROR AND IS NOT RELEASING RESULTS FROM THE PRINTER, BUT IS CHOOSING TO RELEASE RESULTS FROM THE LABORATORY INFORMATION SYSTEM INSTEAD. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED RELATED TO THIS ISSUE. >< (B)(4). ADAPTER. INVESTIGATION SUMMARY: THE CUSTOMER REPORTED THAT THE PRINTER WAS NOT PRINTING OUT RESULTS AS SHOWN ON THE SCREEN. THE CUSTOMER REPORTED THAT THE ISSUE HAD BEEN RESOLVED; SERVICE WAS NO LONGER REQUIRED. THE CUSTOMER STATED THAT THE ISSUE WAS RELATED TO TOO MANY PAGES BEING PRINTED AT THE SAME TIME. PRINTING ONE (1) PAGE AT A TIME RESOLVED THE ISSUE. THE CUSTOMER THEN STATED THAT THE PRINTER WAS PRINTING UNREADABLE REPORTS. THE CUSTOMER REQUESTED FIELD SERVICE FOR ISSUE RESOLUTION. A FIELD SERVICE REPRESENTATIVE (FSR) REPLACED MEDIPRINT ADAPTER AND VERIFIED PROPER OPERATION OF PRINTER; ISSUE WAS RESOLVED. THE INSTRUMENT WAS PERFORMING WITHIN SPECIFICATIONS. A NON-STATISTICAL TREND (NST) REVIEW WAS PERFORMED AND NO NST WAS IDENTIFIED. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 3500 INSTRUMENT FOR THE REPORTED ISSUE. THERE IS NO SYSTEMATIC ISSUE WITH THE CELL-DYN 3500 PRODUCT LINE. THIS IS THE FINAL REPORT. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1649356 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649355 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. H3 OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649352 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649353 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649350 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER WERE ADJUSTED TO MATCH THE STANDARD SETTING PER FA23MAR2010. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649351 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. H3 OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649345 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. H3 OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED A CELL-DYN EMERALD GENERATED DISCREPANT RBC, HEMOGLOBIN AND MCHC RESULTS COMPARED THE RESULTS GENERATED BY A REFERENCE LABORATORY. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649343 | THE CUSTOMER STATED A CELL-DYN EMERALD ANALYZER GENERATED FALSELY LOW MCHC RESULTS FOR MULTIPLE PATIENT SAMPLES. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. H3 OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649344 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS. | 0
1649376 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649377 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649374 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649372 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1868189 | THE CUSTOMER'S CELL-DYN SAPPHIRE BOTTOM SENSOR BASE BOARD WAS REPLACED PER FA05APR2010. THE PART WAS RETURNED TO BE EVALUATED. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. >< (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. >< (B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1649368 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649364 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649362 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649360 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1868170 | (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. >< IT WAS REPORTED TO BAXTER THAT ONE (1) CE INFUSOR LV2 DEVICE WAS OBSERVED TO HAVE THE TUBING BROKEN OFF AT THE JUNCTION OF THE BLUE WINGED CAP AND LUER LOCK. THERE WAS NO PATIENT INVOLVEMENT. THIS WAS OBSERVED BEFORE USE. NO ADDITIONAL INFORMATION IS AVAILABLE. >< (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< THE CUSTOMER'S CELL-DYN SAPPHIRE BOTTOM SENSOR WAS REPLACED PER FA05APR2010. IT WAS NOTICED THAT THE SENSOR TOGGLES WHEN THE CABLE IS MOVED. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. | 0
1649391 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1174883 | THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< EVALUATION RESULTS: OTHER, PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THE WASTE OUTLET TUBE ASSEMBLY HAS A LONG HISTORY OF BEING RELIABLE. THERE HAVE BEEN NO TRENDS IN THE PERFORMANCE OF THIS PART FROM A RELIABILITY PERSPECTIVE. IT HAS CONSISTENTLY MET RELIABILITY EXPECTATIONS. THE WASTE OUTLET TUBE ASSEMBLY HAS BEEN ON THE MARKET FOR OVER 10 YEARS, AND MEETS RELIABILITY EXPECTATIONS FOR THE CELL-DYN 1800 INSTRUMENT. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL, SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY WAS DETERMINED TO BE ACCEPTABLE, AND THE PRODUCT IS MEETING EXPECTED SAFETY PERFORMANCE AS ESTABLISHED IN THE PRODUCT RISK MANAGEMENT FILE. ALTHOUGH THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THIS IS THE FINAL REPORT. >< THE ACCOUNT STATED THAT THE CELL-DYN 1700 ANALYZER'S WASTE CONTAINER OVERFLOWED. THE ACCOUNT STATED THAT THE OPERATOR DID NOT HEAR THE WASTE FULL ALARM AND IS NOT SURE IF THE ALARM IS WORKING PROPERLY. THERE WERE NO INJURIES OR DIRECT FLUID CONTACT REPORTED. THE ACCOUNT WILL ORDER A REPLACEMENT WASTE ASSEMBLY AND WILL MONITOR THE WASTE UNTIL THE REPLACEMENT ARRIVES. >< (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< (B)(4). EVALUATION, RESULTS (B)(4): OTHER, PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. BASED ON THE INTENDED USE, THIS PART WILL WEAR OUT AND NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. NO PREVENTIVE MAINTENANCE SCHEDULE WAS IN PLACE. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY AND SAFETY ISSUES WERE DETERMINED TO BE WITHIN ESTABLISHED FREQUENCY. A HEALTH HAZARD ASSESSMENT (HHA) DETERMINED A LOW HEALTH RISK, WHICH IS CONSISTENT WITH THE RISK MANAGEMENT FILE ASSOCIATED WITH THE PRODUCT ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXIST A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THE WASTE OUTLET TUBE ASSEMBLY PART LIFE EXPECTANCY OF 2 YEARS IS NOT CURRENTLY IN OUR PRODUCT LABELING. THE PREVENTIVE ACTION WILL INVOLVE A LABEL CHANGE FOR THE WASTE OUTLET TUBE ASSEMBLY. WE ARE CHANGING OUR LABELING TO HAVE A PART REPLACEMENT SCHEDULE OF 6 MONTHS, SO PARTS ARE REPLACED PRIOR TO THE END OF THE 2-YEAR LIFE FAILURE. THIS WILL HELP PREVENT OVERFLOW ISSUES CAUSED BY END OF LIFE PART FAILURE. THIS IS THE FINAL REPORT. >< (B)(4). EVALUATION RESULTS (PRODUCT MEETS RELIABILITY DATA). AFTER FURTHER REVIEW, IT WAS DETERMINED THAT THIS COMPLAINT IS ASSOCIATED WITH FA30NOV2009, RCR NUMBER OF 2919069-12/1/09-005-C. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1848599 | (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED A DILUENT EMPTY ERROR MESSAGE WAS GENERATED ON THE CELL-DYN 3200 AND WHITE SMOKE WAS SEEN COMING FROM THE INSIDE OF THE INSTRUMENT. THE INSTRUMENT WAS TURNED OFF AND FIELD SERVICE WAS DISPATCHED. NO INJURY WAS REPORTED DUE TO THIS ISSUE. >< (B)(4). THE INSTRUMENT WAS INSPECTED BY A FIELD SERVICE REPRESENTATIVE (FSR). THE FSR FOUND DILUENT REAGENT INSIDE THE INSTRUMENT. THE VACUUM/PRESSURE MODULE WAS TAKEN OUT AND THE INSIDE OF THE INSTRUMENT WAS CLEANED. THE ELECTRONIC CARDS AND POWER SUPPLIES WERE CHECKED AND NO APPARENT DAMAGE WAS FOUND IN THE COMPONENTS. THE VACUUM ACCUMULATOR #1 WAS FOUND TO BE DAMAGED, AND THE VACUUM PUMP BECAME WET WITH DILUENT REAGENT; BOTH PARTS WERE REPLACED. UPON INSTRUMENT INITIALIZATION, IT WAS NOTICED THAT THE MONITOR WOULD NOT DISPLAY ANY IMAGE. THE INSTRUMENT WAS TESTED WITH A GENERIC MONITOR, WHICH WORKED. THE MONITOR WAS REPLACED AND THE INSTRUMENT WAS RETURNED INTO AN OPERABLE STATUS. THE FSR STATED THAT THE MONITOR WAS THE PROBABLE SOURCE OF THE SMOKE DURING THE EVENT, SINCE IT DID NOT COME ON AFTER INSTRUMENT INITIALIZATION. NO FURTHER INVESTIGATION WAS REQUIRED. THE CELL-DYN 3200 RISK ANALYSIS INDICATES THAT SYSTEM COMPONENTS ARE POTENTIAL SOURCE OF BURNS, FIRE, OR SMOKE BUT CONTROLS IN PLACE TO REDUCE RISK, INCLUDE CURRENT PROTECTION (FUSING), COVERS, AND INSTRUCTIONS IN THE OPERATOR'S MANUAL, SAFETY ICONS, FUSE LABELING, AND HAZARD WARNING LABELS. THE CELL-DYN 3200 SYSTEM OPERATOR'S MANUAL, LIST NUMBER 06H60-01, INDICATES THAT IN CASE OF AN EMERGENCY SITUATION, TURN POWER OFF IN ANY ORDER AS QUICKLY AS POSSIBLE. A NON-STATISTICAL TREND (NST) REVIEW WAS PERFORMED FOR THE REPORTED ISSUE FROM (B)(4) 2010 THROUGH (B)(4) 2010. NO NST WAS IDENTIFIED DURING THE SEARCHED PERIOD. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 3200 RELATED TO THE REPORTED ISSUE. | 0
1649389 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649385 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649386 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< (B)(4). THE CORRECT AWARE DATE IS 3/25/2010. DURING INSTALLATION OF A CELL-DYN EMERALD AT A CUSTOMER SITE, AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) DISCOVERED THAT GAIN SETTINGS FOR THE STANDARD AND NON-STANDARD SPECIMEN TYPES DID NOT MATCH. BY DEFAULT, GAIN SETTINGS SHOULD MATCH FOR ALL SPECIMEN TYPES. CONSEQUENTLY, PATIENT SAMPLES RUN IN A NON-STANDARD SPECIMEN TYPE DID NOT MATCH RESULTS FROM THE SAME PATIENT SAMPLE RUN IN STANDARD SPECIMEN TYPE. ALL PARAMETERS CAN BE AFFECTED BY GAIN SETTING DIFFERENCES, WITH THE EXCEPTION OF HGB. GAIN SETTINGS CANNOT BE MODIFIED BY THE CUSTOMER, ONLY IN THE FACTORY OR BY AN FSR. QUALITY CONTROL, CALIBRATION AND GAIN ADJUSTMENT ALL USE THE STANDARD SPECIMEN TYPE ONLY. IN THE FACTORY, THERE ARE ASSEMBLY INSTRUCTIONS TO ADJUST THE GAIN SETTINGS FOR THE STANDARD SPECIMEN TYPE AND APPLY THE ADJUSTMENTS TO ALL THE NON-STANDARD TYPES; HOWEVER, THERE IS NO VERIFICATION OF THE GAIN SETTINGS FOLLOWING ANY MODIFICATION. INVESTIGATION INTO THIS ISSUE, INCLUDING A SUPPLIER CORRECTION ACTION REPORT TO OEM, IS NOW COMPLETE. WE IDENTIFIED THE ROOT CAUSE OF THE ISSUE TO BE RELATED TO OEM AND ABBOTT HEMATOLOGY DOCUMENTATION AND PROCESS ERRORS, WHERE CRITICAL STEPS WERE INCOMPLETE OR INACCURATE REGARDING INSTRUCTIONS FOR THE GAIN SETTING VALUES ON THE NON-STANDARD TYPES. A PRODUCT CORRECTION LETTER, FA23MAR2010, WAS ISSUED TO ALL CURRENTLY ACTIVE CELL-DYN EMERALD CUSTOMERS, INSTRUCTING THEM TO PERFORM A REVIEW OF GAIN SETTINGS AND ADJUSTMENT WITH ASSISTANCE FROM ABBOTT CUSTOMER SERVICE AND SUPPORT, IF NECESSARY. ADDITIONALLY, ON 01 JULY 2010, A FOLLOW-UP PRODUCT CORRECTION LETTER, FA01JUL2010, WAS ISSUED TO ALL ACTIVE CELL-DYN EMERALD CUSTOMERS WITH A NEW SOFTWARE VERSION, WHICH ELIMINATES THE POSSIBILITY OF THE ISSUE OCCURRING. | 0
1649383 | THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS | 0
1649384 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649381 | (B)(4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649382 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1649380 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS >< THE CUSTOMER STATED THE GAIN SETTINGS ON A CELL-DYN EMERALD ANALYZER DID NOT MATCH THE STANDARD SETTING. ABBOTT CUSTOMER SERVICE INSTRUCTED THE CUSTOMER ON HOW TO ADJUST THE GAIN SETTINGS AND THE ISSUE WAS RESOLVED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1098665 | THE ACCOUNT STATED THAT THE CELL-DYN 1800 ANALYZER'S WASTE SENSOR IS NOT WORKING PROPERLY AND THE WASTE CONTAINER ALMOST OVERFLOWED. THE ACCOUNT VERIFIED THAT THE WASTE TUBING IS BROKEN. THE ISSUE WAS RESOLVED BY REPLACING THE WASTE OUTLET ASSEMBLY. THERE WAS NO DIRECT FLUID CONTACT OR INJURY REPORTED. >< THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION RESULTS (PRODUCT MEETS RELIABILITY DATA) AFTER FURTHER REVIEW, IT WAS DETERMINED THAT THIS COMPLAINT IS ASSOCIATED WITH FA30NOV2009, RCR NUMBER OF 2919069-12/1/09-005-C. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION RESULTS: PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THE WASTE OUTLET TUBE ASSEMBLY (LN 92161-02) HAS A LONG HISTORY OF BEING RELIABLE. THERE HAVE BEEN NO TRENDS IN THE PERFORMANCE OF THIS PART FROM A RELIABILITY PERSPECTIVE. IT HAS CONSISTENTLY MET RELIABILITY EXPECTATIONS. THE WASTE OUTLET TUBE ASSEMBLY HAS BEEN ON THE MARKET FOR OVER 10 YEARS, AND MEETS RELIABILITY EXPECTATIONS FOR THE CELL-DYN 1800 INSTRUMENT. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL, LIST NUMBER 07H80-01, REVISION E, SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY WAS DETERMINED TO BE ACCEPTABLE, AND THE PRODUCT IS MEETING EXPECTED SAFETY PERFORMANCE AS ESTABLISHED IN THE PRODUCT RISK MANAGEMENT FILE. ALTHOUGH THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THIS IS THE FINAL REPORT. >< (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. >< (B)(4). EVALUATION, RESULTS (B)(4): OTHER, PRODUCT MEETS RELIABILITY DATA. ABBOTT LABORATORIES RECEIVED AN INITIAL CUSTOMER COMPLAINT ALLEGING THE CELL-DYN INSTRUMENT DID NOT SIGNAL A WASTE FULL ALARM. ALTHOUGH THE ORIGINAL PART FOR THIS EVENT COULD NOT BE INVESTIGATED, THE INVESTIGATION OF THE REPORTED ISSUE DETERMINED THAT THE PART WAS IN USE GREATER THAN 2 YEARS. THE INVESTIGATION INTO RETURNED PARTS FROM THE FIELD SHOWED THAT ALL HAVE BEEN IN USE GREATER THAN 2 YEARS. IN-HOUSE INVESTIGATION OF THE RETURNED PARTS COULD NOT REPRODUCE THE OCCURRENCE OF THE INITIAL COMPLAINT. THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. BASED ON THE INTENDED USE, THIS PART WILL WEAR OUT AND NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. IT CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. NO PREVENTIVE MAINTENANCE SCHEDULE WAS IN PLACE. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. THE WASTE OUTLET TUBE ASSEMBLY IS PERFORMING AS DESIGNED. THE RELIABILITY AND SAFETY ISSUES WERE DETERMINED TO BE WITHIN ESTABLISHED FREQUENCY. A HEALTH HAZARD ASSESSMENT (HHA) DETERMINED A LOW HEALTH RISK, WHICH IS CONSISTENT WITH THE RISK MANAGEMENT FILE ASSOCIATED WITH THE PRODUCT ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THERE EXIST A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE IF THE PRODUCT FAILS AT THE END OF ITS LIFE EXPECTANCY. THE WASTE OUTLET TUBE ASSEMBLY PART LIFE EXPECTANCY OF 2 YEARS IS NOT CURRENTLY IN OUR PRODUCT LABELING. THE PREVENTIVE ACTION WILL INVOLVE A LABEL CHANGE FOR THE WASTE OUTLET TUBE ASSEMBLY. WE ARE CHANGING OUR LABELING TO HAVE A PART REPLACEMENT SCHEDULE OF 6 MONTHS SO PARTS ARE REPLACED PRIOR TO THE END OF THE 2-YEAR LIFE FAILURE. THIS WILL HELP PREVENT OVERFLOW ISSUES CAUSED BY END OF LIFE PART FAILURE. THIS IS THE FINAL REPORT. | 0
1661861 | THE PERFORMANCE OF THE ASPIRATION PROBE DETECTOR PRINTED CIRCUIT BOARD (ALSO CALLED THE ASPIRATION BOTTOM SENSOR) USED ON THE CELL-DYN SAPPHIRE AUTOMATED HEMATOLOGY ANALYZER MAY BE IMPACTED DUE TO END-OF-LIFE FAILURE. DISCREPANT (LOWER) PATIENT RESULTS MIGHT BE GENERATED DUE TO THIS ISSUE WHEN SAMPLES ARE RUN IN THE CLOSED MODE OF OPERATION. SPECIMENS CAN STILL BE RUN IN THE OPEN MODE OF OPERATION WITH NO IMPACT OR DELAY IN REPORTING RESULTS. NO IMPACT TO USER SAFETY OR PATIENT MANAGEMENT WAS REPORTED RELATED TO THIS ISSUE. A PRODUCT CORRECTION WAS ISSUED AND REPORTED UNDER 21 CFR806 TO THE FDA (B)(4) DISTRICT ON APRIL, 5, 2010. >< IT WAS REPORTED THAT DURING A RADICAL PROSTATECTOMY PROCEDURE, THE JAW WOULD NO LONGER OPEN. THE JAW DID NOT CLAMP THE TISSUE. ANOTHER DEVICE WAS USED TO COMPLETE THE PROCEDURE. THERE WERE NO ADVERSE CONSEQUENCES FOR THE PATIENT. >< (B)(4), EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP FIELD ACTION (B)(4) INCLUDED INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH (B)(4) INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< (B)(4), CONCLUSION (ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES | 0
1553296 | THE CUSTOMER STATES THAT A BURNING SMELL WAS COMING FROM THE CD 3200 SL ANALYZER. THE ANALYZER WAS THEN TURNED OFF AND SERVICE WAS DISPATCHED. THE FSR OBSERVED CHARRING INSIDE THE ANALYZER INCLUDING BOARDS. THE FSR ALSO STATES THAT IT APPEARS BLOOD SPRAYED AND POOLED INSIDE OF THE ANALYZER. NO ACTUAL FIRE WAS OBSERVED AND NO DAMAGE WAS REPORTED OUTSIDE OF THE ANALYZER. THERE WAS NO REPORT OF INJURY OR EXPOSURE REPORTED RELATED TO THIS ISSUE. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1553292 | (B)(4), RETURN OF INSTRUMENT WAS REQUESTED BUT NOT RECEIVED. INVESTIGATION SUMMARY: THE CUSTOMER TECHNICAL ADVOCATE (CTA) CONFIRMED THAT THE INSTRUMENT MAINTENANCE WAS CURRENT, HGB REFERENCE WAS WITHIN SPECIFICATIONS, BACKGROUNDS WERE WITHIN SPECIFICATIONS, AND THE CONTROLS RECOVERED ON TARGET. THE CTA CONFIRMED THAT THE CUSTOMER FOLLOWED MIXING/HANDLING/STORAGE PER RECOMMENDATIONS. FIELD SERVICE REPRESENTATIVE (FSR) WAS SENT TO SERVICE THE INSTRUMENT. AFTER SERVICE, THE INSTRUMENT WAS WITHIN SPECIFICATION, HOWEVER THE ISSUE RECURRED. FIELD SERVICE MANAGER REQUESTED ARRANGEMENTS BE MADE TO RETURN THE INSTRUMENT SINCE THE ISSUE CAN NOT BE RESOLVED BY FIELD SERVICE. THE INSTRUMENT HAS NOT BEEN RETURNED FOR INVESTIGATION. THE CUSTOMER HAS RECEIVED A NEW INSTRUMENT (CELL-DYN EMERALD) TO RESOLVE THE ISSUE. COMPLAINT TRENDING ANALYSIS DID NOT INDICATE ANY ADVERSE TREND FOR THE CELL-DYN 1800, LIST BASE NUMBER 07H77-01, RELATED TO THE REPORTED ISSUE. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 1800, LIST NUMBER 07H77-01, FOR THE REPORTED ISSUE. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT THEY ARE NOTED HEMOGLOBIN RESULT DISCREPANCIES FOR PATIENT SAMPLES PROCESSED USING A CELL-DYN 1800 ANALYZER. AN INITIAL HGB=6.2 G/DL AND HCT=37.6% WAS OBTAINED AND WHEN REPEATED YIELDED A RESULT OF HGB=12.2 G/DL AND HCT=38.8%. A FLAG WAS GENERATED TO ALERT THE USER OF A RESULT BELOW THE ESTABLISHED PATIENT RANGE. THE LOW RESULT WAS NOT REPORTED OUT OF THE LAB WITH NO KNOWN IMPACT TO PATIENT MANAGEMENT. | 0
1711090 | (B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< THE ASPIRATION BOTTOM SENSOR OF THE CELL-DYN SAPPHIRE ANALYZER WAS REPLACED AT THE CUSTOMER'S SITE PER FA05APR2010. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1713693 | (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. OTHER TEXT : AN INVESTIGATION IS IN PROCESS. >< THE CUSTOMER STATED A CELL-DYN SAPPHIRE ANALYZER GENERATED FALSELY DECREASED WBC RESULTS FOR 20 TO 25 PATIENT SAMPLES. ONE OF THE FALSELY DECREASED WBC RESULTS WAS REPORTED OUTSIDE THE LABORATORY. THE CELL-DYN SAPPHIRE ANALYZER GENERATED AN INITIAL WBC RESULT OF 0.4. UPON REPEAT A WBC VALUE OF 10 WAS GENERATED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. | 0
1728708 | (B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< THE CUSTOMER REPORTED INTERMITTENT MIX HEAD FAILED TO HOME ERRORS FOR THE PAST TWO WEEKS. THE FIELD SERVICE REPRESENTATIVE INSPECTED THE CELL-DYN SAPPHIRE ANALYZER AND REPLACED THE ASPIRATION SENSOR SWITCH. NO IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT WAS REPORTED. >< (B)(4). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1661882 | THE CUSTOMER STATED THAT ALL THE HEMATOLOGY PATIENT RESULTS GENERATED ON THE CELL-DYN SAPPHIRE ANALYZER SHIFTED 10% LOWER THAN EXPECTED. THE FIELD SERVICE REPRESENTATIVE RESOLVED THE ISSUE BY REPLACING A BROKEN WIRE OF THE ASPIRATION BOTTOM SENSOR PROBE. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. >< A PATIENT WITH MASS AT APPENDIX WITH ATYPICAL CELLS INDICATING POSSIBLE CANCER; PRESENTED FOR RESECTION AND UNDERWENT LAP-ASSISTED COLECTOMY. AFTER COMPLETE MOBILIZATION, A SMALL INCISION WAS MADE RIGHT AT THE UMBILICUS IN A TRANSVERSE MANNER AND THE MUSCLE WAS DIVIDED WITH A CAUTERY. MOBILIZED BOWEL EASILY PROLAPSED UP INTO THE FIELD. A GIA WAS USED TO COME ACROSS THE BOWEL AT 10 CM PROXIMAL TO THE ILEOCECAL VALVE. THE MESENTERY WAS TAKEN DOWN USING SEQUENTIAL CLAMPS. HIGH LIGATION WAS PERFORMED OF THE ILEOCOLIC (RIGHT COLIC) ARTERY USING A SUTURE LIGATURE OF 0 VICRYL WITH A REINFORCING 2-0 VICRYL TIE. THE RIGHT BRANCH OF THE MIDDLE COLIC WAS TAKEN USING A 2-0 VICRYL SUTURE LIGATURE. A GIA WAS THEN USED TO COME ACROSS THE PROXIMAL TRANSVERSE COLON. THE SPECIMEN WAS REMOVED. A SIDE-TO-SIDE FUNCTIONAL AND END-TO-END ANASTOMOSIS WAS THEN CREATED USING 80-MM GIA. THERE WAS NO BLEEDING FROM THE STAPLE LINE. THE TA 60 WAS USED TO CLOSE THE ENTEROTOMY FOR THE GIA. A CROTCH STITCH WAS PLACED USING 3-0 VICRYL AT THE OTHER END OF THE ANASTOMOSIS. THE BOWEL WAS THEN PLACED BACK IN THE ABDOMEN. ALL SKIN INCISIONS WERE CLOSED. HE MADE SLOW PROGRESSION IN POSTOPERATIVE RECOVERY. SEVEN DAYS LATER, A CT SCAN SHOWED AN ILEUS, BUT NO EVIDENCE OF ANY LEAK OR ABSCESS. TWO DAYS AFTER THE CT, THE PT HAD RAPID A-FIB, ELEVATED WBC AND SIGNS OF SEPTIC SHOCK; WORK-UP INCLUDED ABDOMINAL (ABD) CT WHICH REVEALED PNEUMOBILIA (MOST LIKELY SOURCE PNEUMOBILIA IS ISCHEMIC BOWEL OR OTHER INJURY TO THE BOWEL). ON RE-OPERATION, ANASTOMOTIC LEAK WAS IDENTIFIED AT THE PREVIOUS ANASTOMOSIS AND REPAIRED. THE PATIENT CONTINUED TO DETERIORATE POST-OPERATIVELY DESPITE AGGRESSIVE RESUSCITATION EFFORTS AND PASSED AWAY PRESUMABLY FROM SEPTIC SHOCK (NO AUTOPSY PER FAMILY REQUEST).HEALTH PROFESSIONAL'S IMPRESSION: ANASTOMOTIC LEAK AT SITE WHERE GIA80 STAPLER WAS USED. NO OBVIOUS PROBLEMS WERE SEEN DURING THE INITIAL SURGERY. >< (B)(4), EVALUATION CONCLUSION(ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE) AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. >< (B)(4), EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL (B)(4) WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH (B)(4) INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1728705 | (B)(4). A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. >< A PATIENT WHO UNDERWENT A LOW ANTERIOR RESECTION PROCEDURE WITH ANASTOMOSIS CREATED WITH THE CAR DEVICE DUE TO LOW RECTAL CANCER, RETURNED ON POD 4 WITH ISCHEMIC DISTAL COLON. RE-OPERATION WAS CONDUCTED AND THE INITIAL ANASTOMOSIS WITH THE DEVICE WAS FOUND TO BE INTACT WITH NO LEAKAGE, NO ABSCESS AND NO PERFORATION. THE PATIENT TOLERATED THE PROCEDURE WELL WITHOUT ANY COMPLICATIONS. >< (B)(4). EVALUATION (OTHER): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. >< (B)(4). THIS MDR IS BEING FILED TO DISASSOCIATE THIS COMPLAINT FROM CORRECTION AND REMOVAL (B)(4), SINCE IT WAS PREVIOUSLY REPORTED UNDER MDR NUMBER 2919069-2010-00312. >< DURING A REVIEW OF THE PUMP'S EVENT HISTORY BY BAXTER PERSONNEL, A COLLEAGUE INFUSION PUMP WAS FOUND TO HAVE EXPERIENCED FAILURE CODE 808:02, WHICH INTERRUPTED DELIVERY. THERE WAS NO REPORT OF PATIENT INJURY, MEDICAL INTERVENTION NECESSARY OR ADVERSE REACTION IN ASSOCIATION WITH THIS EVENT. THE USER INTERFACE MODULE MASTER SOFTWARE VERSION IS 5.09.90, WHICH IS CLASSIFIED AS REMEDIATED. NO ADDITIONAL INFORMATION IS AVAILABLE. >< THE CUSTOMER REPORTED A DEFECTIVE ASPIRATION BOTTOM SENSOR OF THE CELL-DYN SAPPHIRE ANALYZER. THE PART WAS REPLACED PER FA05APR2010 TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT WAS REPORTED. | 0
1637320 | THE ACCOUNT GENERATED INCONSISTENT AND FALSELY ELEVATED WBC RESULTS WHEN PROCESSING A PATIENT SPECIMEN ON THE CELL-DYN SAPPHIRE ANALYZER. ADDITIONALLY, THE ACCOUNT NOTICED THE SCATTERGRAM POPULATIONS WERE SHIFTED FOR THIS PATIENT SPECIMEN. THE INCONSISTANT AND FALSELY ELEVATED WBC RESULTS WERE NOT REPORTED OUSTIDE OF THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. DATA PROVIDED:(NO UNITS OF MEASUREMENT WERE PROVIDED);FIRST RUN: WBC = 3.26, NORMAL MODE WITH RST RBC ERROR MESSAGE;SECOND RUN: 16.1, LYSIS PLUS MODE;THIRD RUN: 44.2, LYSIS PLUS MODE; FOURTH RUN: 33.1 LYSIS PLUS MODE;FIFTH RUN: 19.3, NORMAL MODE; MALASSEZ TEST (MANUAL TEST FOR COUNTING CELLS) = 16.8. >< (B)(4) AN EVALUATION IS IN PROCESS. A FOLLOW-UP MDR WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETED. | 0
1639152 | IT WAS REPORTED THAT A PT DEATH OCCURRED WHILE BEING MONITORED ON THE PATIENTNET TELEMETRY SYSTEM. A 'NO TELEM' ALERT WAS ALLEGEDLY PROVIDED; HOWEVER, THE SYSTEM ALARM SETTING FOR THE 'NO TELEM' ALERT WAS SET TO ADVISORY AT THE TIME OF THE EVENT. >< PER THE USER FACILITY REPORT, THE CUSTOMER COMPLETED A LOG REVIEW WITH THE FOLLOWING RESULTS: 'PRIOR TO 3:32:53 THE FULL DISCLOSURE STRIP SHOWS THAT THE ECG SIGNALS FROM THE TRANSMITTER WERE OF GOOD QUALITY WITH NO BREAKS IN THE SIGNAL. THEN AT 03:32:53, THE SIGNAL ABRUPTLY CEASES AND THE 'NO TELEM' MESSAGE IS PRINTED ON THE FULL DISCLOSURE STRIP. AT THE TIME OF THE INCIDENT, THE TRANSMITTER WAS NOT QUARANTINED. THE TRANSMITTER WAS REISSUED TO ANOTHER PT. WHEN THE TRANSMITTER WAS RETURNED TO TELEMETRY, THE BATTERY DOOR HINGE WAS BROKEN.' IN F/U DISCUSSIONS WITH THE CUSTOMER, THE CUSTOMER STATED THEY BELIEVED THE BATTERY BECAME DISLODGED WHEN THE PT INVOLVED IN THE REPORTED INCIDENT FELL DURING THE EVENT. THE BIOMED REPORTEDLY QUARANTINED THE EQUIPMENT AFTER THE BATTERY DOOR HINGE WAS NOTED TO BROKEN. THE BIOMED TESTED THE TRANSMITTER AND STATED THAT IT PERFORMED AS EXPECTED. AT THE TIME OF THE EVENT, THE CIC WAS CONFIGURED SO THAT LEADS FAIL WAS IN SYSTEM ADVISORY LEVEL. SINCE THIS EVENT, IT HAS BEEN CHANGED TO SYSTEM WARNING LEVEL. THE CUSTOMER DOES NOT WANT GE TO INVESTIGATE AT THIS TIME. THE CUSTOMER HAS STATED, THEY WILL CONTACT GE IF THEY DECIDE A LOG REVIEW BY GE OR FURTHER INVESTIGATION OF THE TRANSMITTER IS NEEDED. A F/U REPORT WILL BE FILED IF ADD'L INFO IS RECEIVED. | 1
1867563 | THE PROBE HAS BEEN RECEIVED AND IN-HOUSE TESTING IS IN PROGRESS. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4). >< A HOSPITAL REPORTED DURING AN OPHTHALMIC EXAMINATION, THE PROBE WAS DAMAGED, CAUSING A CORNEAL ABRASION IN BOTH EYES OF THE PT. THE EVENT RESOLVED WITH AN UNK TREATMENT. ADD'L INFO HAS BEEN REQUESTED REGARDING WHAT TREATMENT WAS USED FOR THE PT. | 0
1759821 | DEVICE EVAL ANTICIPATED, BUT NOT YET BEGUN. >< DURING NON-CLINICAL ACTIVITY OF THE DEVICE, THE USER REPORTED THE ARTERIAL MONITOR WAS UNABLE TO DISPLAY TEMPERATURE READINGS. NO OTHER DETAILS REGARDING THE NATURE OF THE EVENT WERE PROVIDED. SINCE THE EVENT OCCURRED DURING NON-CLINICAL ACTIVITY, THERE WAS NO PT INVOLVEMENT DURING THIS EVENT. | 1
1893115 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE MONITOR WAS UNABLE TO DISPLAY PRESSURE READINGS. NO OTHER DETAILS REGARDING THE NATURE OF THE EVENT WERE PROVIDED. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY AND THERE WERE NO ADVERSE CONSEQUENCES TO A PATIENT AS A RESULT OF THIS EVENT. | 1
1575418 | UPON RECEIPT OF THE DEVICE, THE USER REPORTED, THE SAFETY MONITOR DID NOT STOP THE ROLLER PUMPS WHEN A PRESSURE ALARM OCCURRED AS EXPECTED. THE DEVICE IS BEING RETURNED FOR INVESTIGATION. SINCE THE EVENT OCCURRED UPON RECEIPT OF THE DEVICE, THERE WAS NO PATIENT INVOLVEMENT DURING THIS EVENT. >< EVAL ANTICIPATED, BUT NOT YET BEGUN. | 0
1816185 | THE CUSTOMER REPORTS SEVERAL REQUESTS THAT FAILED TO BE PROCESSED THROUGH THE PHONE LIST DUE TO EITHER THE T-TYPE RQUEUE RECORD NOT HAVING BEEN QUEUED OR DUE TO THE RQUEUE RECORD BEING DELETED. PANIC RESULTS SHOULD BE SENT TO THE PHONE LIST BASED ON EITHER STAT_ ROUTE (PU=PLOC+PLIST OR P=PLOC+PLIST) OR STAT_PRINT = 3 IN PANEL MASTER. TO DATE, GE HAS BEEN UNABLE TO REPRODUCE THE ISSUE. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. >< A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1882021 | A PT WITH DIALYSIS DEPENDENT RENAL FAILURE COMPLICATING A STROKE AND HEART ATTACK HAD MULTIPLE SPECIALISTS IN TEAM WITH THE PCP MANAGING CARE AND ENTERING ORDERS ON THE (B)(6) COMPUTER. ON THE PT'S ACTIVE 'PRN MED ORDERS' WERE SEVEN SEPARATE ORDERS FOR POTASSIUM -2 BY MOUTH, 7 BY PARENTERAL-. ON THE PT'S ACTIVE 'PRN MED ORDERS' WERE FOUR SEPARATE ORDERS FOR PHOSPHATE -ALL PARENTERAL-. POTASSIUM PUTS MOST PTS WITH RENAL FAILURE AT RISK. THE NURSE HAS MANY CHOICES FOR THERAPY, BUT WHICH ONE WILL THE NURSE ADMINISTER? THIS PT WAS NOT OVERDOSED, BUT IS AT RISK FOR EXCESS POTASSIUM. AN ONGOING HAZARD, THE DEVICE ENABLES EASE IN ORDERING DUPLICATE MEDICATIONS AND USER UNFRIENDLINESS FOR THE PHYSICIAN TO ELIMINATE THOSE WHICH ARE NOT BEING USED. | 1
1062297 | OUR EVAL CONFIRMED THAT THE TOP COVER OF THE DEVICE ACTUALLY SEPARATED FROM THE BASE AS REPORTED. THIS IS A KNOWN SUPPLIER ISSUE AND WE HAVE IMPLEMENTED SUPPLIER CORRECTIVE ACTION AND INITIATED A FIELD CORRECTION. THE DEVICE IDENTIFIED IN THIS INCIDENT FALLS IN THE RANGE OF THE POTENTIALLY DEFECTIVE DEVICES OF THE FIELD CORRECTION NUMBER: (B) (4). >< WHEN TAKING THE PT MONITOR INFINITY DELTA FROM THE INFINITY DOCKING STATION (IDS) IT FELL DOWN. THE TOP COVER OF THE IDS SEPARATED FROM THE BOTTOM (THE NUTS WERE TORN OFF). NO PERSONS WERE INJURED. (B) (4). | 0
1425249 | DRAEGER COMPLETED A COMPLETE EVAL OF THE REPORTED INCIDENT. THE ACTUAL COMPLAINT WAS RELATED TO THE BRACKET USED TO SECURE THE DRAEGER DELTA XL PT MONITOR AND NOT THE ACTUAL MONITOR. THE CITED MOUNTING ARM WITH PLATE WAS INSTALLED FOUR WEEKS PRIOR TO THE INCIDENT. NO PROBLEMS WERE FOUND BY THE CUSTOMER WITH ANY OF THE MOUNTING ARMS/PLATES THAT WERE INSTALLED AT THE SAME TIME OF THE CITED MOUNTING ARM/PLATE. THE CUSTOMER HAS REINSTALLED THE MONITOR AND ITS MOUNTING ARM ON THE MOUNTING WALL RACK AFTER DETERMINING THAT OPTICALLY NO PROBLEM COULD BE DETERMINED CONCERNING THE MONITOR MOUNTING PLATES AND THEIR FIXATION. THERE WAS NO RETURN OF THE MOUNTING ARM/PLATE TO DRAEGER MEDICAL FOR FURTHER EVAL PRIOR TO THIS REINSTALL. THE ROOT CAUSE FOR THE LOOSENING OF THE MONITOR'S MOUNTING ARM FROM THE MOUNTING'S WALL RACK APPEARS TO BE DUE TO USER ERROR (NOT CORRECTLY SECURING THE ARM TO THE WALL RACK). NO CORRECTIVE ACTIONS HAVE BEEN IDENTIFIED. WE WILL CONTINUE TO MONITOR FOR SIMILAR REPORTS OF THIS CONDITION. >< IT WAS REPORTED THAT THE MONITOR, INCLUDING ITS MOUNTING, FELL DOWN NEAR THE PT'S HEAD. NO PT INJURY WAS REPORTED. (B) (4). | 0
1794768 | PATIENT HAD A 12 BEAT RUN OF VTACH WHICH DID NOT ALARM ON THE MONITORING SYSTEM. THE EVENT WAS OBSERVED BY THE CARDIAC ARRHYTHMIA TECHNICIAN. THE PHYSICIAN WAS NOTIFIED, LABS DRAWN, AND TREATMENT ADMINISTERED AS ORDERED WITHOUT FURTHER INCIDENT.MANUFACTURER RESPONSE: MONITOR SYSTEM, TELEMETRY, APEXPRO.GE IS RESPONSIVE AND ANALYZING STRIPS. CONFERENCE CALL PLANNED. | 1
1674379 | A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED FOR THE INDICATED PACKAGING LOT FOR ANY DEVIATIONS RELATED TO THE REPORTED DEFECT OF THE COMPLAINT. THE REVIEW CONFIRMS THAT THE LOT MET ALL MATERIAL, ASSEMBLY AND PERFORMANCE SPECIFICATIONS. A REVIEW OF THE (B)(6) 2010 NAVILYST MEDICAL COMPLAINT REPORT DID NOT NOTE ANY ADVERSE TRENDS FOR THE PRODUCT FAMILY OF CONTRAST MANAGEMENT AND THE FAILURE MODE OF 'AIR BUBBLES'. THE HOSPITAL DID NOT RETURN A USED DEVICE, BUT DID RETURN 18 UNUSED UNITS FROM THE REPORTED DEVICE LOT. NO VISUAL DEFECTS WERE NOTED. AN AQL SAMPLING OF THE RETURNED DEVICES WAS CONDUCTED IN WHICH MULTIPLE ASPIRATIONS WERE PERFORMED WITH A SYRINGE AFTER THE CONTRAST SETS HAD BEEN ATTACHED TO MANIFOLDS AND A CONTRAST MEDIA BAG. IN EACH CASE, AFTER THE SYSTEM WAS PRIMED AND DE-BUBBLED, NO AIR BUBBLES WERE NOTED. THE TESTS WERE ALSO PERFORMED USING SALINE. THE DIRECTIONS FOR USE PACKAGED WITH THIS DEVICE CONTAINS THE FOLLOWING CAUTION: ENSURE THAT YOU ARE MAKING SECURE CONNECTIONS WHEN USING THIS DEVICE TO PREVENT THE INTRODUCTION OF AIR INTO THE SYSTEM THAT COULD RESULT IN EMBOLISM AND IN RARE INSTANCES DEATH. ALL CONNECTIONS SHOULD BE FINGER TIGHTENED. OVER TIGHTENING CAN CAUSE CRACKS AND LEAKS TO OCCUR THAT COULD RESULT IN EMBOLISM AND IN RARE INSTANCES DEATH. ALL CONNECTIONS SHOULD BE FINGER TIGHTENED. OVER TIGHTENING CAN CAUSE CRACKS AND LEAKS TO OCCUR THAT COULD RESULT IN EMBOLISM AND OR EXPOSURE TO BIOHAZARDS. EXAMINE PRODUCT CAREFULLY FOR ENTRAPPED AIR AND FULLY DEBUBBLE PRIOR TO INJECTION TO MINIMIZE THE POTENTIAL FOR EMBOLISM AND IN RARE INSTANCES DEATH. MANUFACTURING PROCESS CONTROLS IN PLACE FOR THIS DEVICE INCLUDE MULTIPLE VISUAL EXAMINATIONS TO INSURE THAT THE CHECK VALVES HAVE BEEN BONDED CORRECTLY TO THE TUBING. THIS COMPLAINT IS UNABLE TO BE CONFIRMED AS THE RETURNED SAMPLES WERE FOUND TO BE WITHIN SPECIFICATION AND FUNCTIONED AS INTENDED. THE REPORTED ISSUE WAS UNABLE TO BE DUPLICATED WHEN THE DEVICES WERE TESTED. (B)(4). >< AS REPORTED, HOSPITAL WAS EXPERIENCING AIR IN THE TUBING OF THE CONTRAST MANAGEMENT DEVICE. THIS OCCURRED DURING PREPARATION FOR A PROCEDURE WHEN CONTRAST WAS ASPIRATED BY A SYRINGE THROUGH A MANIFOLD TO WHICH THE CONTRAST LINE WAS ATTACHED. THE CONTRAST LINE WAS REPLACED WITH ANOTHER FROM A DIFFERENT LOT, AND NO FURTHER DIFFICULTY WAS ENCOUNTERED. NO AIR WAS INJECTED AND THERE WAS NO PT INJURY. IT WAS INDICATED THAT THE USED DEVICE WOULD BE RETURNED. | 0
1753446 | SPECIMEN LABEL FROM THE BRIDGE SPECIMEN COLLECTION SYSTEM PRINTED LABEL WITH PATIENT A AND CPI, BUT WITH A BAR CODED ACCESSION NUMBER AND TEST THAT BELONGED TO A PREVIOUSLY PRINTED PATIENT B. THE LABORATORY AUTOMATION LINE CAUGHT AND REJECTED THE ERRONEOUS SPECIMEN LABEL AND NO ERRONEOUS RESULTS WERE REPORTED ON THE EITHER PATIENT. ====================== MANUFACTURER RESPONSE FOR CERNER'S BRIDGE SPECIMEN COLLECTION SYSTEMCERNER SUGGESTED A PATCH THAT 'CLEANS UP' LABEL FILES WHEN EVER THE APPLICATION IS 'ABRUPTLY INTERRUPTED', SUCH AS LOSS OF WIRELESS CONNECTIVITY, ETC. THEY BELIEVE THIS WILL PREVENT THE 'INCOMPLETE LABEL RECORD' FROM 'PATIENT A' FROM EVER BEING AVAILABLE FOR 'PATIENT B'. WE MOVED THAT CODE INTO PRODUCTION; NO NEW INCIDENTS HAVE BEEN REPORTED. | 1
1886086 | REFERENCE: MDR 1220063-2010-00040. PER TELEPHONE COMMUNICATION WITH DRAEGER MEDICAL REGARDING THE SUBMITTAL OF ADDITIONAL MDRS FOR EACH SERIAL NUMBER LISTED ON THE INITIAL MDR SUBMITTED, DRAEGER IS SUBMITTING THIS REPORT. IT WAS REPORTED THAT THE DISPLAY OF THE INFINITY DELTA PRESENTS ALEATORY DISFUNCTIONS. IT WAS ALSO REPORTED THAT ONCE THE MONITOR IS RUNNING, THE 'MENU' AND 'STAND BY' WINDOW SUDDENLY APPEAR AND THERE IS NO WAY TO GET OUT OF IT BY PRESSING ANY OTHER KEY UNTIL IT DISAPPEARS BY ITSELF. THERE ARE NO ERRORS FOUND (IN THE ERROR LIST) THAT COULD BE ASSOCIATED TO THIS IRREGULARITY PRESENT BY THE EQUIPMENT. THERE WAS NO PT INJURY REPORTED. (B)(4). >< WHEN DRAEGER INITIALLY RECEIVED THIS REPORT, IT WAS DEEMED NOT REPORTABLE. HOWEVER, THE RESULTS OF THE INVESTIGATION DETERMINED THAT THE EVENTS INVOLVE REPORTS THAT SOME OF THE MONITOR'S FIXED KEYS BECOME INOPERATIVE OR ACTIVATE WITHOUT USER INTERACTION. THE WORST CASE SCENARIO OF THIS TEMPORARY MALFUNCTION COULD BE THAT THE MONITOR MAY DISCHARGE A PT AUTOMATICALLY. THEREFORE, IT WAS DETERMINED THAT THIS COMPLAINT IS REPORTABLE. AN INVESTIGATION WAS INITIATED WITH THE SUPPLIER OF THE SWITCH PANEL. IT WAS DETERMINED THAT THIS CONDITION IS DUE TO A CONTAMINATION PRESENT IN THE FIXED KEY PANEL MATERIAL. A RISK ANALYSIS WAS PERFORMED. DRAEGER MEDICAL IS INITIATING A CORRECTIVE ACTION TO ADDRESS THIS PROBLEM. NO PT DEATHS OR INJURIES HAVE BEEN REPORTED AS A RESULT OF THIS CONDITION. | 1
1473160 | NO PT INJURIES HAVE BEEN REPORTED IN ASSOCIATION WITH THIS EVENT. AS PART OF THE INVESTIGATION, THE REPORTED ERROR COULD BE REPLICATED INTERNALLY USING THE SOFTWARE VERSION IN QUESTION (B) (4). THE COMPANY HAS DEVELOPED A FIX FOR THIS SOFTWARE ANOMALY AND IS IN THE PROCESS OF PERFORMING FIELD CORRECTIONS OF THE AFFECTED SOFTWARE. >< IF PT REPORTS ARE CREATED CONSECUTIVELY FROM UNIQUE PT RAW DATA FILES, WITHOUT CLOSING THE ILUMAVISION APPLICATION IN BETWEEN EACH REPORT CREATION, THE PT IDENTIFIER THAT SHOWS ON THE FIRST CREATED REPORT WILL APPEAR ON ALL SUBSEQUENT CREATED REPORTS. AS A CONSEQUENCE, A REPORT (ELECTRONIC FILE OR PRINTED HARDCOPY) MAY HAVE THE INCORRECT PT IDENTIFIER AND MAY RESULT IN INACCURATE OR INAPPROPRIATE TREATMENT DECISIONS. THE RAW DATA IMAGE FILES USED TO CREATE THE PT REPORTS ARE NOT AFFECTED. | 1
1282747 | A TECHNICIAN TRIED TO REMOVE A SAMPLE FROM THE SYSTEM BEFORE IT WAS SAMPLED, WITHOUT THE SYSTEM BEING IN THE READY STATE. THE DPP SAMPLE DILUTION PROBE SWUNG OVER TO ASPIRATE THE SAMPLE AND PUNCTURED THE TECHNICIAN'S HAND, DRAWING BLOOD. THE INJURY THAT THE TECHNICIAN SUSTAINED WAS DUE TO OPERATOR ERROR. THE TECHNICIAN IMMEDIATELY SOUGHT TREATMENT AT THE ER AND FOLLOWED UP WITH EMPLOYEE HEALTH SERVICES. IT IS UNK IF ANY FURTHER TREATMENT WAS PROVIDED. >< THIS EVENT IS BEING REPORTED BECAUSE THE INJURY OCCURRED IN A BIOHAZARDOUS ENVIRONMENT. ANALYSIS OF THE INFO INDICATES THAT THE CAUSE FOR THE INCIDENT WAS DIRECTLY RELATED TO THE FAILURE OF THE USER TO FOLLOW DOCUMENTED PROCEDURES FOR OPERATION OF THE INSTRUMENT. THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. NO FURTHER EVAL OF THE DEVICE IS REQUIRED. | 0
1879524 | THE LOW BATTERY AUDIBLE ALARM ACTIVATED 10 HOURS AFTER BATTERY EXCHANGE. THE RN SAW 2 OF 5 POWER BARS. THE BATTERY WAS REPLACED. THE PATIENT'S CARDIAC RHYTHM WAS NOT LOST. ROUTINELY, THIS FACILITY EXCHANGES THE BATTERIES EVERY 12 HOURS. | 1
1574909 | THIS CPOE PRODUCT ALLOWS DOSES TO BE ORDERED THAT ARE NOT A MULTIPLE-S- OF THE PILL SIZE. A REPRESENTATIVE EXAMPLE IS 20MG ATENOLOL PO DAILY. PILL SIZE IS 25MG. THIS FEATURE FACILITATES DOSING ERRORS AND ENABLES DANGEROUSLY ABSURD DOSING WITHOUT WARNING. WHEN THE REPORTS SUCH DOSE HAVING BEEN GIVEN, THE CLINICIAN IS LEFT TO WONDER, WHAT DOSE WAS GIVEN, ACTUALLY. | 1
1834713 | PATIENT PRESENTED WITH SEVERE ABDOMINAL COLIC AND HYPERTENSIVE URGENCY. THE SYSTOLIC BLOOD PRESSURE WAS 225 MMHG. BECAUSE OF CLINICAL CIRCUMSTANCES, AND EVERY 3 HOUR CHECK OF BLOOD PRESSURE AND EVERY 3 HOUR ANTI-HYPERTENSIVE THERAPY WAS RECOMMENDED. IT COULD NOT BE ORDERED BECAUSE EVERY 3 HOURS WAS NOT AN OPTION FOR 'MANAGEMENT FREQUENCY' ON THE DROP DOWN MENU NOR COULD IT BE MANUALLY ENTERED. AN ORDER FOR PARENTERAL LABETALOL 10 MG EVERY 3 HOURS PRN SYSTOLIC BLOOD PRESSURE OF GREATER THAN 160 MM HG, THE INDIVIDUALIZED THERAPEUTIC OPTION, COULD NOT BE ORDERED. THE DROP DOWN MENU FOR DOSING INTERVALS HAD APPROXIMATELY 245 OPTIONS, BUT NOT 'Q 3 HOURS', NOR COULD IT BE ENTERED MANUALLY. ALTHOUGH OTHER ORDERING OPTIONS ARE AVAILABLE, WE HAVE FOUND THE RIGIDITY OF THIS SYSTEM TO LIMIT, DIRECT, AND CONTROL THE CARE TO BE ADMINISTERED TO ANY INDIVIDUAL PATIENT, CONVEYING UNNECESSARY RISK. | 1
1834723 | THIS POST OPERATIVE PATIENT SUSTAINED BLOOD LOSS ANEMIA -HCT-23%-. TWO UNITS OF PACKED RED BLOOD CELLS WERE ELECTRONICALLY ORDERED. A POST TRANSFUSION HEMATOCRIT/HEMOGLOBIN LEVEL WAS ELECTRONICALLY ORDERED TO BE DRAWN AT TWO HOURS AFTER THE SECOND UNIT OF RED BLOOD CELLS WAS TRANSFUSED. UNBEKNOWNST TO THE CLINICAL STAFF CARING FOR THE PATIENT, THE CPOE SYSTEM CANCELED THIS BLOOD DRAW. HOURS LATER, THE MORNING BLOOD TESTS WERE DRAWN AND RESULTED. THE HEMATOCRIT/HEMOGLOBIN LEVEL WAS CALLED IN A 'PANIC' -HCT 25%- DESPITE THE TRANSFUSION. THE SYSTOLIC BLOOD PRESSURE INDICATED INCIPIENT SHOCK -75 MM HG- AND THERE WERE ISCHEMIC CHANGES ON THE ELECTROCARDIOGRAM, NECESSITATING VIGOROUS ADDITIONAL CARE, TRANSFUSION, AND INTENSIVE CARE UNIT TRANSFER. THE CANCELLATION OF THE BLOOD DRAW BY THE CPOE EQUIPMENT DELAYED TREATMENT OF THIS ELDERLY PATIENT'S BLOOD LOSS ANEMIA CAUSING LIFE THREATENING COMPLICATIONS. WE HAVE REPEATEDLY FOUND THAT THE CPOE EQUIPMENT CAPRICIOUSLY CANCELS TESTS WITHOUT A DOCTOR'S ORDER SUBJECTING PATIENTS TO UNDUE RISK. THE INCIDENCE HAS NOT BEEN DETERMINED TO THE BEST OF OUR KNOWLEDGE. | 1
1828319 | MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT. >< WE RECEIVED AN ADVERSE EVENT REPORT VIA THE (B)(6) THAT WAS AUTHORED BY THE USER FACILITY. THE REPORT DETAILS A PROBLEM WITH AN 'ATLANTECH REVOLUTION DUAL PUMP' DISTRIBUTED BY (B)(4). THIS DEVICE WAS MANUFACTURED BY FMS UNTIL SOMETIME IN 2006, AND IS CE LICENSED BY FMS, THE MANUFACTURER. THE REPORT STATES THAT DURING AN ARTHROSCOPIC SHOULDER REPAIR (DATE UNDEFINED), THAT WITHIN 5 MINUTES OF THE START OF THE PROCEDURE, THE DEVICE STARTED TO 'MAKE LOUD NOISES'; THE SURGEON ORDERED THAT THE DEVICE BE SHUT OFF AS 2 LITERS OF FLUID HAD ENTERED THE PATIENT'S BODY / JOINT SPACE, WHICH CAUSED THE SHOULDER TO BECOME OVERINFLATED AND ALSO WHAT APPEARS TO BE 'EXTRAVASATIONS' TO THE CHEST AND TRACHEAL AREA; THEY REPORT 'INCREASED PRESSURE TO THE CHEST AND TRACHEAL SHIFT'. AT THIS POINT, THE PROCEDURE WAS ABANDONED. THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND TREATED WITH ANTIBIOTICS. THE PATIENT UNDERWENT A SUCCESSFUL RE-SURGERY FOR REMEDY ON (B)(6) 2010 AND IS REPORTED TO BE DOING FINE. TO DATE THIS IS ALL OF THE INFORMATION AVAILABLE; THERE ARE QUESTIONS OUT FOR FURTHER INFORMATION AND DETAIL. | 0
1711538 | THE VENCOR PROTOUCH COMPUTERIZED ELECTRONIC CARE RECORD AND ORDERING SYSTEM IS DEPLOYED AT ALL LONG TERM ACUTE CARE HOSPITALS UNDER THE OWNERSHIP OF A PUBLICLY TRADED CORPORATION. IT IS SAID TO BE A 'HOME GROWN' SYSTEM WITH EXCEPTIONAL FEATURES HIGHLIGHTED BY ANTIQUATED PLATFORMS. NURSES DESCRIBE CLUNKINESS TO CHART PT APPROPRIATE DESCRIPTION WITH LIMITATION TO ACCURACY. THE EQUIPMENT OFTEN FAILS TO TRANSFER ORDERS AND ASSOCIATED COMMUNICATIONS TO THE RECIPIENT WHETHER THAT BE LABORATORY, CONSULTANTS, NURSES, REHABILITATION, AND NUTRITION, RESULTING IN FAILURE TO EXECUTE ORDERS ACCURATELY, IF AT ALL, AND HAS RESULTED IN CASES IN WHICH ERRONEOUS RECORDS BELONGING TO AN ALTERNATIVE PT ARE ELECTRONICALLY APPENDED TO THE INCORRECT PT. THIS SOFTWARE FAILURE IN THE ELECTRONIC ORDERING SYSTEM POSES RISK TO EACH PT THROUGHOUT THE NATIONAL NETWORK OF HOSPITALS AND SAID RISK IS INSUFFICIENTLY REDUCED. A LARGELY OVERLOOKED SOURCE OF ERROR OF ALL ELECTRONIC ORDERING SYSTEMS IS THE COMPUTERIZED TRANSFORMATION OF THE ORDER OR COMMUNICATION BETWEEN PARTIES SUCH THAT THE RECIPIENT GETS THE COMPUTERIZED ALTERED VERSION RATHER THAN THE ORIGINAL DESCRIPTION WITH NEITHER PARTIES SEEING THE SAME SCREEN. | 1
1846122 | A CUSTOMER REPORTED THAT THE SYSTEM DID NOT PROVIDE THE EXPECTED ALARM FOR A PT EVENT THAT RESULTED IN THE PT'S AICD (AUTOMATIC IMPLANTABLE CARDIOVERTER DEFIBRILLATOR) FIRING. THE CUSTOMER STATES THAT THE SYSTEM ALARMED FOR ACC VENT (ACCELERATED VENTRICULAR RHYTHM) 14 SECONDS PRIOR TO THE EVENT. THE PT REPORTEDLY RECOVERED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE FILED ONCE INVESTIGATION RESULTS ARE AVAILABLE. >< THE DEVICE MFR DATE IS UNK. THE DEVICE WAS SHIPPED ON 01/10/2002. FOLLOWING THE EVENT, THE DISTRIBUTOR'S TECHNICAL ENGINEER PERFORMED A CHECK OF THE LIGHT MONITOR. UPON ARRIVAL, IT WAS DISCOVERED THAT THE MONITOR WAS SHUT DOWN AND WAS STORED IN A BOX. THE TECHNICAL ENGINEER RESTARTED THE MONITOR AND PERFORMED A TEST WITH THE PROBE ON HIS FINGER. THE LIGHT MONITOR AND THE CENTRAL STATION REPORTEDLY ALARMED ACCORDING TO SPECIFICATIONS. DUE TO THE MONITOR BEING SHUT DOWN PRIOR TO THE TESTING, NO LOG OR TREND DATA IS AVAILABLE FROM THE LIGHT MONITOR FOR INVESTIGATION. THE ICENTRAL (CENTRAL STATION) LOGS WERE HOWEVER OBTAINED FOR EVAL. THE LOGS INDICATE THAT ECG AND NIBP WERE NOT MONITORED. 'NIBP CUFF LOOSE', AND 'ECG LEADS OFF' NOTIFICATIONS WERE OBSERVED IN THE LOGS. AN 'SPO2 LOW' ALARM WAS ALSO TRIGGERED. THE LIGHT MONITOR IS REPORTEDLY QUARANTINED BY THE LOCAL POLICE AT THIS TIME, AND THEREFORE IS NOT AVAILABLE FOR GE HEALTHCARE'S INVESTIGATION. WITHOUT THE MONITOR AND ADDITIONAL DETAILS OF THE PT EVENTS BETWEEN 12:00 AND 15:00, NO FURTHER INVESTIGATION CAN BE COMPLETED. GE HEALTHCARE WILL FILE A SUPPLEMENTAL REPORT SHOULD ADDITIONAL INFO BECOME AVAILABLE FOR INVESTIGATION. | 0
1638480 | A CUSTOMER REPORTED EXPERIENCING TELEMETRY DROPOUT INTERMITTENTLY, FOR A PERIOD OF TIME BETWEEN ONE AND THIRTY MINUTES, SEVERAL TIMES PER DAY. NO ADVERSE EVENTS WERE REPORTED, HOWEVER, THE CUSTOMER STATED THE ISSUE DELAYED PATIENT TREATMENT. NO ADDITIONAL PATIENT INFORMATION HAS BEEN PROVIDED AT THIS TIME. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
1697680 | THE CUSTOMER REPORTED A NETWORK OUTAGE RESULTING IN A LOSS OF COMMUNICATION FOR APPROX 98 TELEMETRY BEDS AND 10 HARDWIRE BEDS FOR 6 HOURS. NO DEATH OR SERIOUS INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1940095 | INVERTER BOARD USED IN JACO BATTERY PACK ON COMPUTER CART (WORKSTATION ON WHEELS) LOCATED IN THE BURN INTENSIVE CARE UNIT, CAUGHT FIRE NECESSITATING SUPPRESSION, FIRE DEPARTMENT RESPONSE, AND EVACUATION/RELOCATION OF PTS, VISITORS AND STAFF. NO INJURIES WERE SUSTAINED. | 0
1589914 | ACCOUNT ALLEGED THAT THE BED WILL DRIFT DOWN. >< USER RECEIVED INACCURATE RESULTS FOR ONE PATIENT'S URINE SAMPLE. URINE SAMPLE INITIALLY RESULTED NEGATIVE FOR ERYTHROCYTES (BLOOD) AND NEGATIVE FOR PROTEIN. THE SAMPLE WAS REPEATED GIVING 250 ERY/UL FOR ERYTHROCYTES AND 150 MG/DL FOR PROTEIN. MICROSCOPIC EXAMINATION OF THE SAME SAMPLE SHOWED > 100 RBC/HPF. THE SAMPLE WAS ADDITIONALLY REPEATED ON A DIFFERENT ANALYZER (OLYMPUS), WHICH GAVE A TOTAL PROTEIN URINE RESULT OF 349.5 MG/DL. USER SAID THE RERUN RESULTS WERE REPORTED, NOT THE ORIGINAL URINE RESULTS AND NO PATIENTS WERE ADVERSELY AFFECTED. THE EVENT WAS INVESTIGATED AND A CAUSE COULD NOT BE DETERMINED. ALL CONTROL SAMPLES WERE WITHIN SPECIFICATION. A PRECISION CHECK WAS COMPLETED WITH ALL VALUES WITHIN SPECIFIED RANGES. SERVICE TESTING PASSED WITHOUT ERROR. >< . >< ONE VIAL OF TEST STRIPS FROM THE REPORTED LOT NUMBER WAS RETURNED FOR INVESTIGATION. TESTS WERE PERFORMED USING THE RETURNED AND RETENTION MATERIALS. ALL RESULTS MET SPECIFICATIONS AND NO MALFUNCTION WAS OBSERVED. NO PATIENTS WERE ADVERSELY AFFECTED. | 0
1667005 | THE PATIENT WAS BEING AV PACED BY A TEMPORARY PACEMAKER. THE PATIENT EXPERIENCED APPARENT ASYSTOLE OF ABOUT 13 SECONDS AFTER THE TEMPORARY EXTERNAL PACEMAKER LOST POWER, FOLLOWED BY ONLY OCCASIONAL HEARTBEATS FOR THE NEXT MINUTE UNTIL VENTRICULAR RHYTHM STABILIZED. THE PATIENT'S ECG MONITOR ALARMED FOR ASYSTOLE, BUT THE EXPECTED NOTIFICATION FROM THE CENTRAL MONITORING OFFICE DID NOT OCCUR. THE BEDSIDE MONITOR'S AUDIBLE ALARM IS BELIEVED TO HAVE OPERATED, BUT IT WAS REPORTED THAT THE EVENT PRODUCED NO ALARM OR AUTOMATED PRINTING IN THE CENTRAL MONITORING OFFICE. THE PATIENT'S OXYGEN SATURATION VALUES DID NOT DROP DURING THE EVENT. THE CODE TEAM RESPONDED AND CONNECTED THE SPARE PACEMAKER ALREADY PRESENT AT THE BEDSIDE. THE PATIENT WAS AWAKE AND RESPONSIVE DURING EVENT AND THERE WAS NO CHANGE IN PATIENT CONDITION DUE TO THE EVENT. AFTER THE EVENT, IT WAS CONFIRMED THAT THE AUDIBLE ALARM IN THE PICU WAS HEARD AND THE STAFF RESPONDED APPROPRIATELY. THE CENTRAL MONITOR DID ALARM, BUT RECORDS INDICATE THE RECORDING WAS STOPPED AND THE ALARM SILENCED. IT WAS DETERMINED THAT ALL MONITORING EQUIPMENT WAS FUNCTIONING PROPERLY. IT WAS ALSO DETERMINED THAT THE TEMPORARY PACEMAKER BATTERY HAD DEPLETED. THE STAFF ROUTINELY START OUT WITH A NEW BATTERY ON EACH PATIENT AND HOSPITAL POLICY DIRECTS STAFF TO CHANGE BATTERIES EVERY 2-3 DAYS. ====================== HEALTH PROFESSIONAL'S IMPRESSION ====================== A DEPLETED BATTERY MAY HAVE CAUSED LOSS OF POWER TO PACEMAKER, BUT THIS WAS NOT CONFIRMED. THE PATIENT'S OWN UNDERLYING VENTRICULAR RHYTHM RESUMED AND SUSTAINED THE PATIENT UNTIL THE REPLACEMENT PACEMAKER WAS CONNECTED. | 1
1216359 | IT WAS REPORTED THAT THE CENTRICITY WEB (CWEB) EXPERIENCED A FORCED LOG OUT, WHILE BEING USED IN SURGERY. THE FORCED LOG OUT OCCURRED DUE TO A DEFECT CREATING THE CWEB FORCE OUT, RATHER THAN CWEB SESSION ENDING AT A USER PRE-SET TIME. CENTRICITY WEB IS INTENDED FOR IMAGE REVIEW AND NOT FOR DIAGNOSTIC USE. | 1
1750152 | FAILURE OF CMV RESULT INTERFACE FROM PRODUCT MANAGEMENT APPLICATION -(B)(4)- TO DONOR MANAGEMENT APPLICATION -(B)(4)-. CMV REACTIVE TEST RESULTS DID NOT INTERFACE FROM THE LAB APPLICATION TO UPDATE THE DONORS CMV STATUS. THIRTEEN DONORS -13- WERE IDENTIFIED THAT HAD SUBSEQUENT DONATIONS WHICH WERE TESTED FOR CMV AND FOUND TO BE NON-REACTIVE FOR CMV ANTIBODIES AND PRODUCTS WERE LABELED AS CMV NEGATIVE. SEVENTY-THREE -73- PRODUCT RECALL LETTERS WERE SENT TO (B)(6) HOSPITALS. AS OF (B)(6) 2010, NO ADVERSE AFFECTS TO RECIPIENTS HAVE BEEN REPORTED TO (B)(6). THE INTERFACE FAILURE WAS DISCOVERED ON (B)(6) 2010 AND REPORTED TO THE MANUFACTURER OF THE SOFTWARE APPLICATION -(B)(4)-. | 0
1931377 | WE PRESENT THIS REPORT AS AN EXAMPLE OF ERROR FACILITATED BY DEFICIENT HUMAN FACTOR DEVELOPMENT OF MEDICAL DEVICE DATA SYSTEM SCREENS THAT DISPLAY RESULTS OF LAB TESTS ON AN EHR. THE SPREAD SHEET LISTING THE TESTS BY LINE ON THE LEFT HAND COLUMN Y AXIS HAS THE DATES DISPLAYED ACROSS THE TOP X AXIS. THE ERROR PROMOTING LAYOUT DEFICIENCY IS THAT A COLUMN DOES NOT REPRESENT THE SAME DATE FROM THE TEST LISTED FIRST, AT THE TOP, TO THE TEST LISTED LAST, AT THE BOTTOM. IF A PARTICULAR TEST WAS NOT PERFORMED ON A CERTAIN DAY, THERE IS NOT ANY VISUAL REPRESENTATION, SUCH AS A BLANK SPACE FOR THAT DATE, THAT IT WAS NOT DONE ON THAT DATE. FOR THE PT, HOSPITALIZED FOR FAILURE TO THRIVE, SUBJECT TO THIS COMPLAINT WHO WAS PRESCRIBED COUMADIN FOR THROMBOPHILIA SYNDROME, THE MOST RECENT INR TEST WAS APPROX 2 WEEKS OLD AND WAS AN OUTPATIENT TEST, BUT APPEARED IN THE COLUMN FOR THE DATE OF THE MOST RECENT IN HOSPITAL BLOOD CHEMISTRY RESULT, GIVING THE ILLUSION ON QUICK GLANCE THAT IT WAS A RECENT RESULT. THE COUMADIN WAS PRESCRIBED BASED ON AN OUTDATED TEST AND RESULTED IN AN INR THAT BECAME DANGEROUSLY HIGH, TRIGGERING THE NEED FOR PARENTERAL VITAMIN K. THE PT DID NOT BLEED AND HAS LIVED HAPPILY EVER AFTER AN EXTRA DAY IN THE HOSPITAL. WE BELIEVE THAT THE DEFECT FACILITATING OBFUSCATION OF THE DATE A TEST WAS DONE IN RELATION TO ALL OTHER LAB TESTS, AND THE DEFECT COMINGLING OR FAILING TO IDENTIFY AN OUTPATIENT TEST APART FROM AN INPATIENT TEST, GENERATES INNUMERABLE ERRORS. THERE IS NOT ANY RECORD BEING KEPT OF THIS, TO OUR KNOWLEDGE. | 1
1908838 | CENTRAL MONITOR IN THE ICU FAILED. MD ATTEMPTING TO LOOK AT PREVIOUS RHYTHM STRIPS USING THE FULL DISCLOSURE APPLICATION. SHE MANAGED TO LOOK AT THE RHYTHM USING THE ARRHYTHMIA RECALL. HOWEVER SHE STILL WANTED TO USE FULL DISCLOSURE SO SHE CALLED THE MANUFACTURER WHO WALKED HER THROUGH USING THE MOUSE. MD STILL HAVING PROBLEMS SO SHE CONTACTED ANOTHER REPRESENTATIVE WHO CONTINUED TO GUIDE HER THROUGH THE PROCESS TO DISCONNECT AND CONNECT THE CABLES. AT THAT TIME THE CENTRAL STATION WENT BLANK AND THE PATIENTS' DATA DISAPPEARED AND 'NOT CONNECTED' KEYS KEPT BLINKING ON THE MONITOR/DISPLAY. THE MANUFACTURER WAS UNABLE TO RECONNECT AND IS SENDING A LOANER. | 1
1711484 | TESTS OF BLOOD WERE ORDERED TO BE PERFORMED AND WERE PERFORMED IN THE EARLY MORNING HOURS. THE RESULTS WERE SENT ELECTRONICALLY TO THE COMPUTERIZED STORAGE SECTION ON THE RECORD. THE CRITICAL ABNORMALITY WAS IN THE INTERVAL CHANGE OVER THE 8 HOURS FROM THE EARLIER TEST, RATHER THAN THE ABSOLUTE RESULT OF THE TEST THAT WAS SILENTLY SENT TO THE ELECTRONIC REPOSITORY. NOT ONE HEALTH PROFESSIONAL KNEW IT WAS THERE UNTIL APPROX 6 HOURS LATER. THE NEW RESULT RETURNED SHOWING AN ABSOLUTE LEVEL LIFE THREATENING ANEMIA. THE CRITICAL TEST REFLECTING THE INTERVAL CHANGE WAS BURIED IN THE ELECTRONIC REPOSITORIES. THE PT DIED FROM BLOOD LOSS. THE ELECTRONIC DEVICE DID NOT WARN THERE WAS A NEW RESULT. THIS DEATH ASSOCIATED PROBLEM, THE FAILURE TO NOTIFY OF A NEW RESULT, IS A FLAW OF ALL ELECTRONIC REPOSITORIES OF PT'S RESULTS. | 1
1665473 | IT HAS BEEN DISCOVERED THAT THE MEDICATION REVIEW SCREEN CONSISTING OF A LIST OF THE MEDICATIONS ON THE LEFT MOST COLUMN AND A GRID INDICATING THE NUMBER OF DOSES ADMINISTERED IN EACH TIME PERIOD. THE DEFECT MANIFESTS ITSELF IN THE REPRESENTATION OF EXACTLY WHAT MEDICATION WAS ADMINISTERED. FOR PTS WHO ARE TAKING TWO OR MORE FORMS OF THE SAME PHARMACOLOGICAL AGENT, EG NITROGLYCERIN IN THE PASTE FORM, SUB LINGUAL FORM, AND PILL FORM, THE NAME OF THE MEDICATION IS REPRESENTED AS 'NITROGLYCERIN' ON THE LIST. THE NUMBER OF TIMES IT WAS ADMINISTERED IS CONFUSINGLY LISTED AS A NUMBER IN PARENTHESES WITHOUT CLARITY ON WHICH FORM OF THE MEDICATION WAS ADMINISTERED, CREATING CONFUSION IN THE CLINICIAN AND RISK TO THE PT. | 1
1726923 | THE COMPUTERIZED PHYSICIAN ORDER ENTRY INSTRUMENTS WERE DEPLOYED. PLEURAL FLUID SPECIMEN WERE OBTAINED. CHEMICAL AND CYTOLOGICAL ANALYSIS WAS ORDERED ON THE SPECIMEN USING THE CPOE INSTRUMENTS. THE SPECIMEN WAS NOT ANALYZED AS ORDERED. CYTOLOGY RESULTS DID NOT BECOME AVAILABLE. SIMILAR DYSFUNCTION APPEARED IN MORE THAN 6 PATIENTS. THE HOSPITAL HAS RESORTED TO PAPER FORMS TO CLARIFY THE DYSFUNCTION OF THIS CPOE INSTRUMENT. THE PATIENTS MAY HAVE CANCER CAUSING THE FLUID BUILD UP. THIS WILL NOT BE KNOWN UNTIL TIME PASSES. THERE MAY BE RESULTANT DELAY IN DIAGNOSIS OF CHEST ORGAN CANCER. | 1
1713653 | THE CUSTOMER REPORTED A PT INCIDENT INVOLVING PICIS' ED (EMERGENCY DEPT.) ELECTRONIC HEALTH RECORD APPLICATION RELATED TO ALLERGY DATA RECEIVED VIA INTERFACE FROM THEIR ENTERPRISE INFORMATION SYSTEM, WHICH WAS 'CODED' USING A 3RD PARTY DRUG DATABASE TERMINOLOGY NOT SUPPORTED BY PICIS. THE NUMERIC CODE REPRESENTING 'SULFONAMIDES' IN THE ENTERPRISE SYSTEM REPRESENTS A DIFFERENT ALLERGY CONCEPT (METHYLXANTHINES) IN PICIS' SUPPORTED THIRD PARTY DRUG DATABASE. THE SCREEN DISPLAYED THE CORRECT ALLERGY NAME OF SULFONAMIDES TO THE PHYSICIAN, BUT THE BACKGROUND ALLERGY CHECKING LOGIC (DETERMINED BY THE NUMERIC CODE SENT FROM THE INTERFACE) DID NOT NOTIFY THE PHYSICIAN OF A POTENTIAL ALLERGY INTERACTION, WHEN A PRESCRIPTION FOR BACTRIM WAS WRITTEN. THE PT WAS DISCHARGED FROM THE ED WITH THE WRITTEN PRESCRIPTION WHICH WAS LATER FILLED BY A RETAIL PHARMACIST. THE PT MANIFESTED SYMPTOMS OF AN ALLERGIC REACTION WITHIN 24 HOURS, WAS SEEN TWICE BY PHYSICIANS AND SENT HOME WITH OUTPATIENT TREATMENT BUT ULTIMATELY REQUIRED 48 HOURS OF INPATIENT OBSERVATION, BEFORE BEING DISCHARGED HOME WITH NO LONG TERM EFFECTS. >< OUR INVESTIGATION FOUND THAT THE REASON FOR NO ALLERGY INTERACTION NOTIFICATION WAS A RESULT OF RECEIVING CODED ALLERGIES FROM A NON-SUPPORTED EXTERNAL DRUG DATABASE SOURCE. THE ALLERGY DATA IN QUESTION WAS RECEIVED BY PICIS' ED HEALTH RECORD APPLICATION IN 2006. IT IS CUSTOMARY IN HEALTHCARE APPLICATIONS FOR ALLERGY INFORMATION TO 'CARRY THROUGH' FROM VISIT TO VISIT. THE IMPORTANCE OF THIS IS THAT IN THE COURSE OF OUR INVESTIGATION, WE CONCLUDED THAT THIS CLIENT IS NO LONGER SENDING NON-SUPPORTED ALLERGY CODES AND HAS NOT DONE SO SINCE LATE 2008. PICIS IS WORKING WITH THE CUSTOMER TO REVIEW ALLERGY CODES SENT PRIOR TO 2008 TO ASCERTAIN IF THERE ARE ANY ADDITIONAL ERRONEOUS ALLERGY CODES STORED IN THE DATABASE. ADDITIONALLY, DURING THE COURSE OF OUR INVESTIGATION, WE IDENTIFIED TWO SIMILAR CUSTOMER SITE CONFIGURATIONS WITH NON-SUPPORTED DRUG DATABASES. THERE WERE NO ASSOCIATED PT INCIDENT REPORTS. HOWEVER, WE ARE WORKING WITH THESE CUSTOMERS TO PREVENT UNSUPPORTED ALLERGY CODES FROM BEING SENT TO THE PICIS APPLICATION. PICIS IS ALSO EXPLORING THE POTENTIAL FOR SIMILAR SYSTEM ADMINISTRATION ERROR IN OTHER SITES. WE WILL EVALUATE THE NEED FOR CORRECTIVE ACTION AND IMPLEMENT ACCORDINGLY. | 1
1887065 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO OBTAINING A BLOOD SAMPLE RESULT THAT WAS TYPED AS 'O NEGATIVE' ON THE PK7300 AUTOMATED MICROPLATE SYSTEM. THIS RESULT WAS DISCREPANT WITH THE DONOR HISTORY, WHICH WAS TYPED AS 'O POSITIVE' FOR MULTIPLE PREVIOUS DONATIONS. THE SAMPLE WAS RETESTED AS 'O POSITIVE' ON THE SAME INSTRUMENT AND CORRECTED REPORTS WERE SENT OUT. THERE WAS NO DEATH, INJURY, OR AFFECT TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. >< AN INSPECTION WAS PERFORMED BY A BCI FIELD SERVICE ENGINEER (FSE). ALL FUNCTIONS, ALIGNMENTS AND DISPENSE VOLUMES WERE CHECKED. NO OUT OF SPECIFICATION CONDITIONS OR FAILURES WERE IDENTIFIED. REAGENT DISPENSE CHECKS WERE PERFORMED ON (B)(6) 2010 AND WERE ALL WITHIN RANGE. BASED ON THE INVESTIGATION, TESTING, DATA, AND BATCH RECORD REVIEW THE REAGENTS WERE MANUFACTURED APPROPRIATELY AND ARE PERFORMING AS EXPECTED. INVESTIGATION FOR THE PK7300 INSTRUMENT IS STILL ONGOING. | 0
1430168 | AN OCD FIELD ENGINEER CONTACTED THE CUSTOMER VIA PHONE REGARDING THE UNEXPECTED VOLUME ERRORS. THE CUSTOMER INDICATED THAT THE ISSUE WAS RESOLVED AFTER CHECKING THE WASH AND WASTE BOTTLE CAPS. SERVICE WAS NOT NEEDED TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. PRODUCT LABELING INSTRUCTS THE CUSTOMER THAT 'WHILE IT IS POSSIBLE TO MODIFY THE '~' RESULT, THIS SYMBOL INDICATES A POTENTIAL VOLUME ISSUE IN THE MICROTUBE AND SHOULD BE REVIEWED BY THE TECHNOLOGIST. THE SAMPLE SHOULD BE RETESTED'. OCD CUSTOMER TECHNICAL SERVICES (CTS) INFORMED THE CUSTOMER REGARDING THE PROPER PROCEDURE FOR THE ERROR MESSAGES AS PER THE PROVUE'S IFU. THE CUSTOMER INDICATED THAT ALL SAMPLES IN QUESTIONS WOULD BE RETESTED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B)(4). >< THE CUSTOMER REPORTED PERFORMING TESTING ON THE ORTHO PROVUE ANALYZER. THE PROVUE FUNCTIONED AS EXPECTED, POSTED CONDITION CODES AND BROUGHT THE GEL CARDS TO THE SERVICE RACK FOR MANUAL REVIEW DUE TO THE UNEXPECTED VOLUME ISSUES. HOWEVER, THE CUSTOMER REVIEWED, RESULTED THE GEL CARDS AND ACCEPTED THE TEST RESULTS WITHOUT RETESTING THE SAMPLES. THIS INCIDENT IS BEING REPORTED BASED ON USER ERROR. PROVUE MALFUNCTION DID NOT OCCUR. FALSE RESULTS MAY BE REPORTED THAT SHOULD HAVE BEEN ABORTED OR INVALIDATED. | 0
1459982 | DEVICE EVALUATION ANTICIPATED, BUT NOT YET BEGUN. THE SUSPECT DEVICE HAS BEEN RETURNED FOR EVALUATION; HOWEVER, THE INVESTIGATION IS NOT COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN ADDITIONAL INFO IS AVAILABLE. >< THE COMPLAINANT REPORTED THAT DURING AN ANGIOGRAPHY, THE ROTATOR OF THE TUBING BROKE. NO HARM OR INJURY TO THE PT WAS REPORTED. | 0
1551301 | THE COMPLAINANT REPORTED THE ROTATOR BROKE DURING A LEFT VENTRICULOGRAM PROCEDURE. THE DEVICE WAS DISCARDED AT HOSPITAL. >< DEVICE EVALUATION: THE SUBJECT DEVICE WAS NOT RETURNED FOR EVALUATION. THEREFORE THE COMPLAINT COULD NOT BE CONFIRMED. THE DEVICE HISTORY RECORD WAS REVIEWED WITH NO RELATED EXCEPTIONS NOTED. IN COMPLAINTS SIMILAR TO THIS ONE, THE ROOT CAUSE OF THE MALFUNCTION WAS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. POTENTIAL CORRECTIVE ACTIONS ARE BEING EVALUATED. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. METHOD: ACTUAL SUSPECT DEVICE WAS NOT RETURNED. | 0
1551302 | THE COMPLAINANT REPORTED THE ROTATOR BROKE AT PRESSURE LIMIT OF 500 PSI. >< DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED; THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE ROOT CAUSE OF THE MALFUNCTION IS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. POTENTIAL CORRECTIVE ACTIONS ARE BEING EVALUATED. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. DEVICE HISTORY RECORD WAS REVIEWED. | 0
1551304 | DEVICE EVALUATION: THE SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED; THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE ROOT CAUSE OF THE MALFUNCTION IS ATTRIBUTED TO FAILURE OF THE ADHESIVE BOND BETWEEN THE ROTATOR HOUSING AND THE ROTATOR COLLAR. POTENTIAL CORRECTIVE ACTIONS ARE BEING EVALUATED. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. DEVICE HISTORY RECORD WAS REVIEWED. >< COMPLAINANT REPORTED THE ROTATOR BROKE DURING A LEFT VENTRICULOGRAM PROCEDURE. | 0
1707006 | DEVICE EVALUATION: THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. EVALUATION METHOD: DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. >< DURING A LEFT CORONARY VENTRICULOGRAM THE ROTATOR BROKE. THERE WAS NO REPORT OF HARM OR INJURY. THE CUSTOMER REPORTED THEY HAD NINE DEFECTIVE DEVICES BUT HAVE NOT PROVIDED ANY ADDITIONAL INFO OR CLINICAL DETAILS FOR THOSE DEVICES. CUSTOMER WILL NOT BE RETURNING ANY OF THE DEVICES FOR EVALUATION. | 0
1701703 | THE ROTATOR BROKE DURING USE. NO HARM OR INJURY REPORTED. >< DEVICE EVALUATION: DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. COMPLAINT DATABASE REVIEW FOUND ON SIMILAR COMPLAINTS FOR THIS LOT NUMBER. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. METHOD: OTHER, DEVICE HISTORY RECORD WAS REVIEWED. COMPLAINT DATABASE WAS REVIEWED. CONCLUSIONS: OTHER, A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. | 0
1511236 | CONCLUSIONS: OTHER, A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. >< ROTATOR BROKE DURING ANGIOGRAPHY USING A MICRO CATHETER WHICH OCCURRED TWICE IN A ROW. ACTUAL PRESSURE 920 PSI. (THIS REPORT IS TWO OF TWO FOR REPORTED DEVICE. THE FIRST DEVICE NOTED IN THIS EVENT IS REPORTED ON MDR NUMBER 1721504-2009-00111). | 0
1733376 | THE ROTATOR BROKE DURING USE. NO HARM OR INJURY REPORTED. >< ONE USED SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS CONFIRMED. THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. COMPLAINT DATA BASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. THE ROOT CAUSES OF THE FAILURE ARE ATTRIBUTED TO THE MANUFACTURING BONDING PROCESS. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. THIS LOT WAS BUILT PRIOR TO IMPLEMENTATION OF THE NOTED CORRECTIVE ACTIONS. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. METHOD OTHER: DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. | 0
1053199 | THE CUSTOMER INDICATED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER WAS NOT DISPENSING THE CORRECT AMOUNT OF FLUID INTO THE MTS GEL CARDS. NO ERRONEOUS RESULTS WERE REPORTED. INCORRECT OR NO DISPENSE CAN LEAD TO ERRONEOUS TEST RESULTS AND TRANSFUSION OF INCOMPATIBLE BLOOD. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THAT THE PRESSURE WAS OUT OF SPECIFICATIONS AND THE WASH STATION WAS CRACKED. REPLACEMENT OF THE PROBE, WASH STATION AND THE APPROPRIATE ADJUSTMENTS HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). | 0
1645762 | A CUSTOMER REPORTED THAT ARTERIAL BLOOD PRESSURE ALARM LIMITS WERE OUT OF NORMAL RANGE, AND THE MONITOR DID NOT ALARM. THERE WAS NO REPORTED DEATH OR SERIOUS INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1813844 | THE SERIAL NUMBER OF THE PATIENT DATA MODULE IS NOT AVAILABLE AT THIS TIME. (B)(4). >< A CUSTOMER REPORTED THAT NIBP (NON-INVASIVE BLOOD PRESSURE) READINGS WERE INACCURATE, DELAYING VASOPRESSOR THERAPY TO THE PATIENT. AT 12:30, THERE WAS A NIBP OF 160/70 FOR SEVERAL READINGS. WHEN THE NURSE VERIFIED THE NIBP READING WITH A MANUAL BLOOD PRESSURE, THEY OBTAINED A BLOOD PRESSURE OF 80/SYSTOLIC. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1691454 | WE USE A GE CENTRICITY PACS. ONE OF THE RADIOLOGY GROUPS THAT READS FOR THE HOSPITAL HAS ASKED THAT THE FILTERS THAT IDENTIFY 'UNREAD' STUDIES BE CREATED WITH THE LOGIC THAT REMOVES CASES THAT ARE BEING VIEWED (A.K.A. 'LOCKED') FROM THE UNREAD LIST - THEY DON'T WANT TO OPEN THE CASES TO SEE THAT ANOTHER USER IS VIEWING THE STUDY.RECENTLY, STUDIES WERE FOUND TO BE LOCKED THAT WERE NOT BEING DISPLAYED - AS LOCKED, THERE WAS A DELAY IN THE INTERPRETATION AS THE STUDIES WERE NOT ON AN UNREAD WORKLIST.IT IS UNKNOWN WHETHER THE ISSUE IS CAUSED BY A BUG IN THE SOFTWARE. THE SAME VERSION OF THE CLIENT WAS RE-INSTALLED AND THE PROBLEM HAS NOT CONTINUED. | 1
1660718 | CASE WAS IN SESSION. A LOUD BANG WAS HEARD. TEE SCREEN WAS BLANK. THE PROBE COULD NOT BE REMOVED FROM THE PATIENT AT THE TIME; IT WAS REMOVED AT THE END OF THE CASE WITHOUT HARM TO THE PATIENT. | 1
1633693 | A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). >< THE CUSTOMER REPORTS AN INCORRECT RESULTS. NUMERIC RESULTS WAS SENT WITH A COMMA FROM THE PERFORMING LAB. ULTRA ONLY DISPLAYED THE DIGIT BEFORE THE COMMA. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. | 1
1482943 | CUSTOMER REPORTS THAT THE LAB INSTRUMENT, CURRENTLY CONFIGURED TO PERFORM PT, PTT, INR AND FIBRINOGEN TESTS. RESULTS IN A 0 (ZERO). THE 0 RESULT WAS CONVERTED TO '>200' BY A LIBRARY CODE AND AUTO-ACCEPTED. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. >< A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1759847 | PHYSICIAN WROTE A NOTE IN THE EHR AND AN EMPLOYEE IN MEDICAL RECORDS WAS ABLE TO PUT THE DOCUMENTATION 'IN ERROR'. THERE WAS NO ADVERSE REACTION TO THE PATIENT. HOWEVER, THIS PHYSICIAN CHECKED DOCUMENTATION ON HER NEXT SHIFT TO ENSURE THAT THE PATIENT WAS ADMITTED AND NOTED THAT HER DOCUMENTATION WAS NOT IN THE EHR DUE TO THE DOCUMENTATION BEING ALTERED. THE PATIENT WAS ADMITTED WITHOUT INFORMATION FOR HOSPITALISTS BECAUSE THE DOCUMENTATION HAD BEEN PUT 'IN ERROR'.THE PURPOSE OF THE 'IN ERROR' FUNCTION IS AS FOLLOWS: IT IS USED TO CORRECT A MISTAKE IN THE MEDICAL RECORD. FOR EXAMPLE, LET'S SAY I WRITE A NOTE ON PATIENT A, BUT MADE A MISTAKE AND SHOULD HAVE BEEN REPORTING ON PATIENT B. THE ORIGINAL DOCUMENT SECTION CAN BE COPIED AND PASTED AND THE DOCUMENTATION ON THE WRONG PATIENT CAN BE PUT 'IN ERROR' SO THAT THE INFORMATION IS NOT READILY VISIBLE, YET IS NOT REMOVED FROM THE ELECTRONIC MEDICAL RECORD. THE PRACTITIONER WHO REPORTED THIS INCIDENT IS CONCERNED THAT OTHER AUTHORS ARE ABLE TO 'IN ERROR' THE ORIGINAL AUTHOR'S REPORTS WITHOUT NOTIFICATION. OUR MEDICAL RECORDS DEPARTMENT 'IN ERROR''S, AND COPIES AND PASTES DOCUMENTS SEVERAL TIMES A DAY TO CORRECT NOTE TYPOS, DOCUMENTATION ERRORS, ETC. IT IS FELT THAT IT WOULD BE UNMANAGEABLE TO RELY ON THE AUTHORS TO CORRECT THE ERRORS IN CERNER. THE DOCUMENT CLEARLY STATES AT THE TOP WHO MADE THE 'IN ERROR' AND WHY. WHEN DOCUMENTATION IS PUT 'IN ERROR', THE INFORMATION IS NOT SEEN BECAUSE THERE IS AN 'IN ERROR' FILTER WHICH BLOCKS THE ORIGINAL DOCUMENTATION. HOWEVER, IF THE PRACTITIONER OR PERSON VIEWING THE ELECTRONIC MEDICAL RECORD IS AWARE OF THE 'IN ERROR FILTER', AND REMOVES IT, THEN THE VIEWER CAN SEE THE DOCUMENT EXISTS 'IN ERROR'. THE DOCUMENT IS NOT COMPLETELY REMOVED FROM THE ELECTRONIC MEDICAL RECORD - IT IS JUST NOT READILY VISIBLE.====================== HEALTH PROFESSIONAL'S IMPRESSION======================THE PHYSICIAN IS CONCERNED THAT ORIGINAL DOCUMENTATION CAN BE PUT 'IN ERROR' WITHOUT CONSULTING THE ORIGINAL AUTHOR OF THE REPORTS. | 1
1620210 | PATIENT FOUND DEAD IN HOME. MEDICAL ALERT NECKLACE CAUGHT ON BACK OF CHAIR. PATIENT STRANGLED BY NECKLACE. | 0
1784288 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< (B)(4). THE PRODUCT HAS BEEN RETURNED AND AN INVESTIGATION IS IN PROCESS. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE ADC FA16MAY2006 LETTER. | 0
1510098 | THE COMPANY SERVICE REP EXAMINED THE SYSTEM AND PERFORMED MEASUREMENTS WITH THE MODEL EYE USING THE CUSTOMER'S PROBE AND A NEW PROBE. THE MEASUREMENTS WERE VERY CLOSE WITH BOTH PROBES. TROUBLESHOOTING WAS PERFORMED BY REPLACING THE CPU CARD AND THE MEASUREMENTS DID NOT CHANGE. THE COMPANY SERVICE REP CONCLUDED THE BIOMETRY PROBE MOST LIKELY CONTRIBUTED TO THE EVENT REPORTED. THE BIOMETRY PROBE WILL BE SENT IN-HOUSE FOR EVAL. THE SYSTEM WAS THEN TESTED AND MET ALL PRODUCT SPECS. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFO BECOMES AVAILABLE. (B) (4). POWER, MISCALCULATION OF. (B) (4). (B) (4). >< THE SURGEON REPORTED THE PT PRESENTED WITH MYOPIA AFTER SURGERY. MULTIPLE ATTEMPTS WERE MADE FOR MORE INFO AND PT STATUS WITH NO RESPONSE TO DATE. | 0
1917703 | A COORDINATING ADMINISTRATOR REPORTED THAT AN ULTRASOUND BIOMETRY MACHINE WAS NOT GIVING CONSISTENT RESULTS WHEN USED ON A PT UNDER TOPICAL ANESTHESIA. THE PT WAS SENT HOME, AND RETURNED ON ANOTHER DAY TO HAVE THE DIAGNOSTIC PROCEDURE COMPLETED. IT WAS REPORTED THAT THERE WAS NO PT HARM. >< FDA EVAL CODES; METHOD: USE OF DEVICE WAS EVALUATED VIA PHONE SUPPORT. ADDITIONAL INFO FROM THE CUSTOMER BY PHONE REVEALED THAT THERE WAS A PARAMETER SET ERROR. THE CUSTOMER WAS TRYING TO PERFORM THE EXAM WITH THE IMMERSION TECHNIQUE, BUT THE EQUIPMENT WAS SET ACCORDING TO THE CONTACT TECHNIQUE. NO SVC VISIT WAS PERFORMED AND NO SAMPLES WERE RETURNED TO THE MFR. (B)(4). | 1
1358053 | USER FACILITY REPORTED TO THE MANUFACTURER THAT THE USER WAS NOT RECEIVING ANY ALARM TONES FROM THE DEVICE. >< THE MANUFACTURER IS WAITING FOR THE USER FACILITY TO RETURN THE DEVICE FOR EVALUATION. UPON COMPLETION OF THE EVALUATION OF THE DEVICE IN QUESTION, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
1836169 | I HAVE TWO AUTOMATIC BLOOD PRESSURE MONITORS, ONE MADE BY MICROLIFE AND ONE MADE BY LIFESOURCE. BOTH THESE MONITORS, AFTER INFLATING TO ABOVE SYSTOLIC PRESSURE, RELEASE PRESSURE SO SLOWLY THAT THE READING BECOMES VERY PAINFUL AND THE PROLONGED HIGH PRESSURE CAUSES THE CAPILLARIES IN MY HAND TO BURST. MY LEFT HAND IS NOW VERY SPLOTCHY. I HAVE TAKEN PHOTOS OF MY HAND WHICH I CAN FORWARD TO YOU. TODAY MY BLOOD PRESSURE WAS 190/78, SO THESE UNITS SHOULD BE ABLE TO READ THOSE VALUES WITHOUT CAUSING DAMAGE OR PAIN. THESE UNITS DO NOT FUNCTION EFFECTIVELY AND SHOULD BE REMOVED FROM THE MARKET. DIAGNOSIS OR REASON FOR USE: #1 & #2 - HYPERTENSION. EVENT ABATED AFTER USE: #1 & #2 - YES. EVENT REAPPEARED AFTER REINTRODUCTION: #1 & #2 - YES. THE DEVICE IS ALSO SOLD AT (B)(6) PHARMACIES. | 0
1702186 | EVAL IN PROGRESS, BUT NOT YET CONCLUDED. >< DURING SERVICING OF THE ON-LINE BLOOD GAS MONITOR IN THE REPAIR CENTER, THE BLOOD PRESSURE MONITOR GENERATED A 'CHECKSUM' FAILURE DURING THE FINAL TEST. A NEW ARTERIAL PRINTED CIRCUIT CARD WAS TRIED WHICH DID NOT RESOLVE THE PROBLEM. A NEW (B)(4) PRINTED CIRCUIT BOARD WAS ALSO TRIED AND DID NOT RESOLVE THE PROBLEM. A NEW BLOOD PRESSURE MONITOR WAS USED WHICH RESOLVED THE ISSUE AND PASSED ALL THE TESTING. SINCE THE EVENT OCCURRED DURING SERVICING OF THE DEVICE, THERE WAS NO PT INVOLVEMENT DURING THIS EVENT. | 0
1893116 | DURING PREVENTATIVE MAINTENANCE OF THE DEVICE, THE USER REPORTED THE MOUNTING ARM THAT HOLDS THE CENTRAL CONTROL MONITOR MADE NOISE WHEN THE SCREEN WAS POSITIONED. NO OTHER DETAILS REGARDING THE NATURE OF THIS EVENT WERE PROVIDED. SINCE THE EVENT OCCURRED DURING PREVENTATIVE MAINTENANCE OF THE DEVICE, THERE WAS NO PATIENT INVOLVEMENT DURING THIS EVENT. >< EVALUATION IN PROGRESS, BUT NOT YET CONCLUDED. | 1
1638681 | EVAL IN PROGRESS BUT NOT YET CONCLUDED. >< DURING PREPARATION FOR USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED WHEN THE CENTRAL CONTROL MONITOR WAS POWERED ON, A WHITE BLANK SCREEN DISPLAYED. THIS ISSUE OCCURRED TWICE. THE THIRD TIME, THE USER REBOOTED THE SYSTEM WHICH RESOLVED THE ISSUE. AN ALTERNATE DEVICE WAS EMPLOYED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WAS NO ADVERSE CONSEQUENCE TO THE PT AS A RESULT OF THIS EVENT. | 1
1638692 | DURING ROUTINE TESTING OF THE DEVICE, THE USER REPORTED THE DISPLAY ON THE CENTRAL CONTROL MONITOR FROZE AFTER APPROX TEN HOURS OF TESTING. THE USER REBOOTED THE MONITOR, RESOLVING THE ISSUE. THE DEVICE WAS RETURNED FOR FURTHER INVESTIGATION. SINCE THE EVENT OCCURRED DURING ROUTINE TESTING OF THE DEVICE, THERE WAS NO PT INVOLVEMENT DURING THIS EVENT. >< EVAL IN PROGRESS BUT NOT YET CONCLUDED. | 1
1702181 | EVAL IN PROGRESS, BUT NOT YET CONCLUDED. >< AFTER USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE TOUCH SCREEN ON THE CENTRAL CONTROL MONITOR WOULD NOT RESPOND TO TOUCHING OF THE SCREEN AND THE CURSOR WAS NOT VISIBLE. NO ERROR MESSAGES WERE OBSERVED. THE USER ATTEMPTED TO CALIBRATE THE SCREEN THREE TIMES WITH THE SAME RESULT. THE USER ALSO CHECKED THE CONNECTION FOR EACH CABLE AND PRINTED CIRCUIT BOARD INSIDE THE MONITOR. ALL CONNECTIONS WERE FINE. THE DEVICE WAS RETURNED FOR INVESTIGATION AND REPAIR. SINCE THE EVENT OCCURRED AFTER A CARDIOPULMONARY BYPASS PROCEDURE, THERE WAS NO PT INVOLVEMENT DURING THIS EVENT. | 1
1733327 | DURING PREPARATION FOR USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE CLOCK ON THE CENTRAL CONTROL MONITOR DID NOT FUNCTION AS EXPECTED. THE USER REPORTED THE CLOCK ADVANCES FURTHER THAN IT SHOULD. THE DEVICE WAS USED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WAS NO ADVERSE CONSEQUENCES TO THE PT AS A RESULT OF THIS EVENT. >< EVALUATION IN PROGRESS, BUT NOT YET CONCLUDED. | 1
1660120 | THE CUSTOMER REPORTED THAT THE PROVUE PROBE IS DRIPPING ADN THE DILUTION CUP OVERFLOWING. NO ERROR MESSAGES WERE POSTED. THE CUSTOMER INDICATED THAT THE PROVUE WAS TAKEN OUT OF USE. NO ERRONEOUS RESULTS WERE POSTED. >< AN OCD FIELD ENGINEER REPLACED THE NECESSARY PARTS AND RETURNED THE INSTRUMENT TO EXPECTED OPERATION.(B)(4) | 0
1760697 | ON AT LEAST 5 SEPARATE OCCASIONS IN THE NEONATAL INTENSIVE CARE UNIT WHEN USING A PHILLIPS MODEL 3001A MULTI MEASUREMENT SERVER -MMS- WITH SPO2 TYPE OPTION AO1 IN AN MP-70 BEDSIDE MONITOR FOR PULSE OXIMETRY MEASUREMENT ON A DARK-SKINNED NEONATAL INFANT, MEASUREMENT DATA BECAME INTERMITTENT AND THE MONITOR DISPLAYED 'SENSOR MALFUNCTION.' REPLACING THE SENSOR DID NOT HELP. USING THE SAME SENSOR BUT SWITCHING TO A MODEL M1020B SPO2 MODULE IN THE SAME MONITOR SHOWED IMPROVED AND CONSISTENT MEASUREMENT DATA. IN THE M3001A, WHEN THE PROBLEM OCCURRED, THE SPO2 CABLE WAS REMOVED BUT THE 'SENSOR MALFUNCTION' MESSAGE DID NOT CLEAR EVEN WHEN 'SILENCE' WAS SELECTED. A POWER REBOOT WAS REQUIRED TO CLEAR THE MESSAGE. THESE PROBLEMS HAVE NOT RESULTED IN ANY HARM TO THE PTS. A PHILIPS MEDICAL SERVICE REP HAS OBSERVED THE PROBLEM AND HAS INFORMED US THAT PHILIPS IS INVESTIGATING THE ISSUE. | 1
1429815 | BOSTON SCIENTIFIC CORP (BSC) WAS INFORMED BY THE USER FACILITY THAT THE INTRAVASCULAR ULTRASOUND IMAGING SYSTEM'S (ILAB) IMAGING PC, SITUATED IN THE CONTROL ROOM OF THE CATHETER LABORATORY, WAS REPORTED TO HAVE SMOKE COMING OUT OF IT. THE ILAB SYSTEM WAS NOT IN USE AT THE TIME OF THE EVENT, BUT WAS IN AN IDLE STATE. THERE WERE NO PT OR USER INJURIES/COMPLICATIONS REPORTED DUE TO THIS EVENT. ACCORDING TO THE USER FACILITY THE ROOM WAS EVACUATED AND THE FIRE DEPT WAS CALLED PER THE REQUIREMENTS OF THE USER FACILITY'S SAFETY PROCEDURES. >< (B) (4). | 1
1584485 | PATIENT WAS ADMITTED TO ROOM A JUST AFTER HIS ROOMMATE WAS ADMITTED TO B (SAME ROOM). THE TELEMETRY BOXES HAD BEEN REVERSED WHEN THE ROOM WAS SET UP AND THE BOXES (TRANSMITTERS) WERE NOT CHECKED AS THEY WERE PLACED ON THE PATIENTS. THE NURSES DID NOT CALL OUT TO MONITORS (CENTRAL STATION) AFTER PLACING THE MONITORS (TRANSMITTERS) ON THE PATIENTS. BOTH PATIENTS WERE HOOKED UP (CONNECTED TO THEIR TRANSMITTER) WITHIN A SHORT PERIOD OF TIME. THE ISSUE WAS DISCOVERED WHEN ONE PATIENT WAS TAKEN OFF THE MONITOR IN THE MORNING FOR A BATH. MONITOR STRIPS AND COMPUTER CHARTING WERE CORRECTED. THIS OCCURRENCE WAS EXPLAINED TO THE PATIENTS AND FAMILY MEMBERS. SPOKE WITH BOTH NURSES ABOUT CALLING OUT TO MONITORS (CENTRAL STATION) WHEN HOOKING UP TELEMETRY. | 1
983680 | (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATED THAT A NURSE HAD VISITED THE DOMICILE OF A PT FOR THE PURPOSE OF OBTAINING BLOOD SAMPLES FOR ANALYSIS. ONE EDTA AND ONE CITRATE SAMPLE ANTICOAGULANT TUBES WERE DRAWN FROM THE PT. THE EDTA SAMPLE TUBE GENERATED A HEMOGLOBIN RESULT OF 5 G/DL USING A CELL-DYN 3200 SL ANALYZER. THE RESULT WAS REPORTED TO A PHYSICIAN THAT QUESTIONED THE RESULT AND THE PT WAS HOSPITALIZED. A NEW SAMPLE WAS DRAWN AT THE HOSPITAL AND TESTED (METHOD NOT PROVIDED), YIELDING A HEMOGLOBIN RESULT OF 12 G/DL. THE CITRATE SAMPLE WAS THEN TESTED OBTAINING A HEMOGLOBIN RESULT OF 16 G/DL. NO FURTHER IMPACT TO PT MGMT WAS REPORTED. | 0
1784041 | (B)(4). THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
1782509 | (B)(4). THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW-UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
1624888 | PT WAS UNDERDOSED. A 40CGY WAS DELIVERED AND 400CGY PLANNED. ONCENTRA BRACHY DID NOT UPDATE THE DWELL TIMES AFTER THE USER SELECTED 5MM BOX AND CHANGED THE F-FACTOR. WHEN THE PROBLEM WAS DISCOVERED 40CGY X 2 WAS DELIVERED. | 0
1686631 | THE PHARMACY MEDICATION ALERT SYSTEM WAS UPDATED AND THE ALERTS WERE INADVERTENTLY TURNED OFF. THE ALERTS WERE DOWN FOR APPROXIMATELY 7 DAYS. | 1
1847459 | IT WAS REPORTED THAT THE ARMS OF THE HOLDER CANNOT WITHSTAND THE MECHANICAL LOAD OF THE DEVICE. THERE WAS NO PT INJURY REPORTED. (B)(4). >< DRAGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 0
1876802 | PATIENT WEARING A TELEMETRY UNIT (TRANSMITTER) STANDING IN THE HALLWAY LEANING AGAINST A DOOR FRAME AND ALLEGED TO HAVE RECEIVED A 'SHOCK' / 'BURN' AT THE SITE OF THE LEFT UPPER CHEST AT THE TELEMETRY LEAD. TWO DAYS POST EVENT THE PATIENT WAS NOTED TO HAVE A SMALL REDDENED CIRCULAR AREA AT THE POINT OF CONTACT FROM THE TELEMETRY GEL PAD (ELECTRODE). HEALTH PROFESSIONAL'S IMPRESSION: UNKNOWN, POSSIBLE CONDUCTER OF STATIC ELECTRICITY FROM THE FLOOR. | 0
1455795 | IT WAS REPORTED THAT THE METAL LOCKING TABS OF SIX DASH PORT 2 DOCKING STATIONS BROKE OFF OF THE MONITOR CLAMP ON THE PIVOT BRACKET. IN THIS SCENARIO THE DASH MONITOR WOULD NOT BE CONNECTED TO THE DASH PORT 2, POTENTIALLY LEADING TO A SITUATION WHERE THE MONITOR COULD FALL ON A PT OR HEALTH CARE PROVIDER. THIS REPORT IS FOR THE FIRST OF SIX AFFECTED DEVICES. THERE WAS NO REPORT OF PT OR USER INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
1557681 | DEVICE EVALUATION: CUSTOMER REPORT WAS CONFIRMED. REC'D (1) SINGLE DPT - VAMP JR. KIT IN OPEN ORIGINAL PACKAGE. KIT WAS NOT USED. PEDIATRIC TUBING WAS COMPLETELY BROKEN FROM SOLVENT BOND JOINT WITH SAMPLE SITE. CROSS SURFACES OF BROKEN TUBING ENDS WERE UNEVEN AND ROUGH. (B) (4). >< THE FOLLOWING WAS REPORTED: 'VISUALIZED IN THE PACKAGE AS BROKEN TUBING JUST PROXIMAL TO THE RESERVOIR.' | 0
1487266 | THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PT, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PT DUE TO THE REPORTED MALFUNCTION. | 0
1511842 | THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER FAILED TO PIPETTE REAGENTS AND RED CELLS INTO GEL CARDS. CUSTOMER TECHNICAL SERVICES PERFORMED TROUBLESHOOTING WITH THE CUSTOMER BY ASKING THE CUSTOMER TO PERFORM A PRIME CYCLE. THE CUSTOMER OBSERVED DRIPPING FROM THE HANDLER SYSTEM. THE CUSTOMER WAS NOT ABLE TO TELL IF ANY ERROR CODE WAS POSTED. FLUID LEAK CAN LEAD TO DILUTION OR CONTAMINATION OF REAGENTS / SAMPLES AND ERRONEOUS TEST RESULTS. >< THE CUSTOMER INDICATED THAT THE REAGENTS ON BOARD WERE CONTAMINATED. THE CUSTOMER NOTICED DROPLETS ON THE REAGENT CAROUSEL AND CENTRIFUGE. ALL TESTING WAS ABORTED. NO ERRONEOUS RESULTS WERE REPORTED. CONTAMINATED REAGENTS WERE DISCARDED. THE CUSTOMER REMOVED THE INSTRUMENT FROM OPERATION. A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND FOUND THAT THE SYRINGE WAS LEAKING. THE FE REPLACED THE SYRINGE, WASH STATION ASSAY, CLEANED THE ANALYZER AND CORROSION IN THE DILUTER. PIPETTING TEST WAS ACCEPTABLE AND THE FLUIDIC SYSTEM WAS WITHIN SPECIFICATIONS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). | 0
1079971 | (B) (4). (B) (4). FAILURE INVESTIGATION IN PROGRESS. WILL SUBMIT A FOLLOW UP REPORT UPON COMPLETION. >< INITIAL REPORT: ON (B) (6) 2008, THE CUSTOMER REPORTED AN ISSUE WITH THE PYXIS ANESTHESIA SYSTEM 3500 SERVICE RELEASE 2 (PAS 3500 SR2) HAVING LOCKED UP WHEN ACCESSING MULTIPLE RAPID ACCESS MATRIX DRAWERS. THE DRAWERS WERE BEING ACCESSED BY THE ANESTHESIOLOGIST DURING SETUP PRIOR TO THE START OF THE CASE. THE FAILURE WAS DESCRIBED AS A SYSTEM FREEZE IN WHICH THE STATION'S CLOCK WAS NO LONGER ADVANCING AND THE SYSTEM HAD TO BE REBOOTED. IN THIS SCENARIO, THE RAPID ACCESS DRAWERS REMAINED UNLOCKED AND THE CONTENTS WITHIN THE DRAWERS WERE ACCESSIBLE TO THE USER. THE MINIDRAWER REMAINED LOCKED AND CONTENTS WITHIN THE DRAWER WERE NOT ACCESSIBLE TO THE USER. | 0
1756852 | (B) (4). THE METER HAS BEEN CONFIRMED TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE ADC FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER. UPON INVESTIGATION, THE METER WAS FOUND TO EXHIBIT THE MEMORY OVERWRITE MALFUNCTION. IN ADDITION, THE CUSTOMER COMPLAINED THAT HE WAS UNABLE TO TEST AND WAS FEELING 'DISORIENTED AND FEELING BAD' HE SELF-TREATED BY TAKING A GLUCOSE TABLET. THERE WAS NO REPORT OF DEATH, THIRD PARTY MEDICAL INTERVENTION OR SERIOUS INJURY. | 0
1416822 | IT WAS REPORTED THAT THE MONITOR DID NOT ALARM AS EXPECTED WHEN THE PT HAD VENTRICULAR TACHYCARDIA (VT). IT WAS ALSO REPORTED THAT THE MONITOR ALARMED CORRECTLY WHEN THE PT HAD VENTRICULAR FIBRILLATION (VF). IT WAS REPORTED THAT A PT WAS INJURED, BUT THE EXTENT OF THE INJURY IS UNK. (B) (4). >< THERE HAVE BEEN NO SIMILAR COMPLAINTS FILED WITHIN THE COMPLAINT DATABASE. THE ALARM SETTINGS THAT WERE SUBMITTED BY THE COMPLAINANT ARE DATED FROM (B) (6) (FOUR DAYS AFTER THE EVENT). THE CUSTOMER STATES THE MONITOR HAS BEEN SEQUESTERED SINCE THE DATE OF THE EVENT WHEN THESE PRINTOUTS WERE TAKEN. THE ALARMS SETTING FROM (B) (6)2009, INDICATE THAT THE VT ALARM WAS SET UNDER THE PARAMETER OF A LIFE-THREATENING ALARM. THERE WERE NO ALARM LOGS CAPTURED RELATED TO THE DATE OF THE EVENT. THE ROOT CAUSE FOR A MONITOR TO POST WITH AN ALARM LEVEL OTHER THAN WHAT IT WAS INTENDED AS CONFIGURED BY THE USER IS UNDETERMINED. ADD'L INFO TO COMPLETE AN INVESTIGATION WAS NOT MADE AVAILABLE. ACCORDING TO THE COMPLAINANT, THE CITED MONITOR WAS TESTED AND FUNCTIONS AS PER MFR'S SPECS. THE MONITOR HAS BEEN RETURNED TO THE CUSTOMER AND IS FUNCTIONING WITH NO FURTHER RELATED ISSUES REPORTED. NO ACTIONS ARE PLANNED BY THE MFR AT THIS TIME. WE WILL CONTINUE TO MONITOR FOR SIMILAR REPORTS OF THIS CONDITION. | 1
1416821 | IT WAS REPORTED THAT AN SC9000XL MONITOR NORMALLY DOCKED ON A DOCKING STATION, FELL ON PT LOCATED IN A BED. (B) (4). >< DRAEGER IS STILL INVESTIGATING THIS REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION IS COMPLETED. | 0
1938584 | A SITE REPORTED THAT A PT HAD A BRADYCARDIC EPISODE AND WENT INTO ASYSTOLE. RESUSCITATION MEASURES WERE UNSUCCESSFUL. THE PT EXPIRED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A F/U REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
1510903 | DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. >< IT WAS REPORTED THAT THERE WAS AN UNEXPLAINED INCREASE IN BIS VALUES DURING A PROCEDURE. THERE WAS NO PATIENT INJURY REPORTED. (B) (4). | 1
1510917 | IT WAS REPORTED THAT A DELTA XL MONITOR WENT BLACK AND THEN DISPLAYED QUESTIONABLE VALUES. THERE WAS A PATIENT INJURY REPORTED. (B) (4). >< DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1779492 | IT WAS REPORTED THAT THE PT TOOK OFF THE ECG ELECTRODES OR THE SPO2 FINGER CLIP REPEATEDLY. THE MONITOR REPORTEDLY DID NOT GIVE WHITE (ADVISORY) ALARMS TO ALERT THE USER TO RESPOND TO THE PT'S SITUATION. IT WAS REPORTED THAT THE PT DIED. THE MONITOR WAS VERIFIED BY DRÄGER SERVICE AND THE IN-HOUSE BIOMEDICAL DEPT. NO INDICATION OF A MALFUNCTION OF THE MONITOR OR THE CENTRAL REGARDING THE ALARMING COULD BE FOUND. (B)(4). >< DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1636026 | IT WAS REPORTED THAT THE ICS CENTRAL STATION REBOOTED SEVERAL TIMES WHILE MONITORING PATIENTS. THERE WAS NO PATIENT INJURY REPORTED ASSOCIATED WITH THIS EVENT. THE ICS HAS REPORTEDLY REMAINED IN USE. (B)(4). >< DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1593222 | THE CUSTOMER REPORTED THAT WHILE THE MONITOR WAS IN USE WITH A PT, THE MONITOR DISPLAYED A MESSAGE THAT IT NEEDED TO BE PLUGGED IN, EVEN THOUGH IT WAS ALREADY PLUGGED IN. WHEN THEY ATTEMPTED TO PLUG IT IN WITH MORE FORCE, IT SPARKED. NO INJURIES WERE REPORTED. >< UPON RECEIPT OF THE LINE CORD AT THE FACTORY, (B)(4) ENGINEERING OBSERVED A BURN INSIDE THE WALL CONNECTOR WHERE THE WHITE CONNECTOR MEETS THE WALL PRONG OR BLADE. THE LINE CORD WAS RETURNED TO THE VENDOR, ELECTRI-CORD, WITH A REQUEST FOR CORRECTIVE/PREVENTIVE ACTION (B)(4). ELECTRI-CORD RESPONDED: ROOT CAUSE ANALYSIS: THE MFG DATE CODE OF THE CORD WAS 0519 (10TH WEEK OF 2005), WEEK OF FEB 5, 2005. THE MOLDED PLUG WAS ECM CAT. MC-10H WITH TALLER BLADES AND GROUND PIN. THE LINE BLADES BROKE INSIDE THE MOLDED PLUG FROM FORCES THAT WERE SUFFICIENT TO BREAK THE LINE BLADES OF THE PLUG. IT IS UNK THE SOURCE OF THESE FORCES OR MIS-USE THAT CAUSED THE BROKEN BLADES. THE BROKEN BLADES CAUSED ARCING TO OCCUR, THUS THE VISUAL SPARKING. CONTAINMENT PLAN: THERE ARE NONE OF THESE PARTS WITH TALLER BLADES/GROUND PINS IN STOCK AT ECM. CORRECTION ACTION: ELECTRI-CORD MFG. HAS DISCONTINUED THE USE OF THE CRIMPED STYLE TALLER BLADE THAT WAS USED IN THE PRODUCTION OF THE RETURNED CORD IN ALL OF THEIR MOLDED PLUGS. THESE HAVE BEEN REPLACED BY SOLDER TYPE BLADES/GROUND PINS THAT ARE MORE RESISTANT TO BREAKAGE. THE DISCONTINUATION OF THE USE OF THE TALLER BLADES/GROUND PIN IN THEIR CORDS WAS EFFECTIVE IN MID AUGUST, 2009. ALL PRODUCTION NOW USES ALTERNATE STYLE BLADES WHICH ARE MORE RESISTANT TO MIS-USE. ELECTRI-CORD REPORTS A 0.0012% FAILURE RATE FOR THESE CORDS; THEREFORE, NO ADDITIONAL CORRECTIVE ACTION IS PLANNED BY (B)(4) AT THIS TIME. REGARDING THIS PARTICULAR CUSTOMER, WE ARE REPLACING THEIR DAMAGED LINECORD. WE SUGGESTED THAT THEY INSPECT THEIR LINECORDS, AND WE WILL REPLACE ANY DAMAGED LINECORDS THAT THEY HAVE. | 1
1400346 | IT HAS BEEN IDENTIFIED THAT THE DRIVER SOFTWARE OF THE X-RAY UNIT COULD PLACE A SCANNED X-RAY IMAGE INTO A TEMP FOLDER, THEN PRESENT THAT SCANNED IMAGE DURING A LATER ACQUISITION OF A NEW IMAGE. THE IMAGE FROM THE TEMP FOLDER MAY THEN BE PRESENTED TO A NEW PATIENT FILE. >< AN X-RAY IMAGE WAS TAKEN ON A NEW PATIENT WITH THEIR ORTHORALIX 8500DDE UNIT BUT DURING THE IMAGE ACQUISITION, THE PROGRAM PULLED UP IMAGES FROM ANOTHER PATIENT'S FILE AND THE NEW IMAGES COULD NOT BE FOUND. THE X-RAY IMAGES FOR THE PATIENT HAD TO BE RE-TAKEN. | 1
1586811 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE GEL CAMERA IS READING FALSE POSITIVES IN THE FORWARD AND REVERSE GROUPING. NO ERRONEOUS RESULTS WERE REPORTED. A FALSE POSITIVE RESULT MAY LEAD TO MISCLASSIFICATION OF A PT'S ABO GROUP, WITH THE POTENTIAL FOR TRANSFUSION OF ABO INCOMPATIBLE BLOOD. >< SHIPPED ON 11/2004. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND REPLACED ILLUMINATION CARDS, READER PCB AND CLEANED READER CAMERA. FE RE-ADJUSTED READER CAMERA AND CREATED REFERENCE IMAGE. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THE REPAIR. (B)(4). | 0
1463089 | AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE. THE FE CHECKED THE READER CAMERA OPTICS FINE ADJUSTMENTS, HANDLER PROBE, WASH BLOCK, DIFFUSER PLATE, PRESSURE AND VACUUM, WHICH WERE ALL WITHIN SPECIFICATIONS. OCD SECOND LEVEL SUPPORT (TAC) REVIEWED THE LOG FILES SUBMITTED BY THE FE. REVIEW OF THE GEL CARD AND BITMAP IMAGE BOTH SHOWED A VERY WEAK REACTION. THE REFERENCE IMAGE PRE ADJUSTMENT AND POST ADJUSTMENT AS WELL AS THE TIF IMAGE SHOWED THAT THERE MAY BE TOO MUCH LIGHT ON THE BOTTOM HALF OF THE CARD. THE FE CLEANED THE LOWER AND UPPER LED BANKS. THE UPPER BANK WAS FOUND TO BE DIRTY. READER ADJUSTMENTS WERE COMPLETED AND NEW REFERENCE IMAGE WAS CREATED. WAD DIAGNOSTIC TEST WAS SUCCESSFULLY AND QC TESTING WAS ACCEPTABLE. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER INTERPRETED A POSITIVE (1+) REACTION WITH SCREEN CELL #1 AS NEGATIVE. THE CUSTOMER VISUALLY INSPECTED THE RESULTS AND CONFIRMED THAT THE REACTIONS IN THE MICROTUBE WERE POSITIVE (1+). THE CUSTOMER PERFORMED PANEL TESTING WITH THE SAMPLE IN MANUAL GEL METHOD AND IDENTIFIED ANTI-K ANTIBODY. THE INCIDENT OCCURRED ON (B) (6)2009. NO ERRONEOUS RESULTS WERE REPORTED. THE CUSTOMER INDICATED THAT THEY HAVE NOT OBSERVED THE ISSUE WITH ANY OTHER SAMPLES AND THIS WAS AN ISOLATED INCIDENT. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1693408 | OUR ANESTHESIA DEPARTMENT IS THE (B) (6) IN THE UNITED STATES TO GO-LIVE WITH AN ANESTHESIA EMR USING EPIC SOFTWARE - (B) (6). ONE OF THE KEY COMPONENTS OF THE EMR IS THE ABILITY TO AUTOMATICALLY CAPTURE ANESTHESIA MACHINE AND PHYSIOLOGIC MONITOR DATA VIA DEVICE INTERFACES. HOWEVER, SHORTLY AFTER GO-LIVE WE NOTICED AN INACCURACY WITH THE DATA COMING FROM THE GE AESTIVA ANESTHESIA MACHINE. AESTIVAS ARE CONFIGURED TO VENTILATE IN ONE OF FOUR MODES -UNIQUE IDENTIFIER IN PARENTHESES-: VOLUME CONTROL -V- PRESSURE CONTROL -P- PSV-PRO -O- SIMV-VC -S-. EACH MODE OF VENTILATION IS SUPPOSED TO SEND A UNIQUE IDENTIFIER AS DATA OUTPUT. THIS IDENTIFIER IS THEN SENT TO THE CAPSULE DATACAPTOR SERVER WHERE THE INTERFACE TRANSLATES THE DATA AND SENDS IT TO THE EMR THROUGH INTERNAL TESTING -WITH THE HELP OF CAPSULE-, AND FINALLY VIA DIRECT CONFIRMATION FROM GE SENIOR PRODUCT MGMT, WE HAVE DISCOVERED THAT AESTIVAS ARE SENDING WRONG INFO ON VENTILATION MODE THROUGH THEIR SERIAL PORTS. VOLUME CONTROL AND PRESSURE CONTROL BOTH COME OVER AS 'V' AND 'P' RESPECTIVELY - THESE ARE CORRECT. HOWEVER, PSV-PRO IS SENT AS 'P' INSTEAD OF 'O' AND SIMV-VC IS SENT AS 'V' INSTEAD OF 'S'. IN ADDITION TO THE DATA BEING ERRONEOUS, THE PROTOCOL/SUPPORT DOCUMENTATION IS WRONG AS WELL. THERE IS A SIGNIFICANT CLINICAL DIFFERENCE BETWEEN THESE FOUR MODES OF VENTILATION AND IT IS IMPORTANT TO ENSURE THAT THIS DATA IS ACCURATE. OUR FINDINGS HAVE BEEN COMMUNICATED DIRECTLY WITH (B) (6), SENIOR PRODUCT MANAGER GE, AND (B) (6), ENGINEER GE. THIS WEEK WE MET WITH A REP OF GE AND WE WERE INFORMED THAT THEY KNOW ABOUT THE PROBLEM, BUT THERE ARE NO PLANS TO FIX IT. AS MORE HOSPITALS AND ANESTHESIA DEPTS ACROSS THE NATION BEGIN TO DEVELOP AND IMPLEMENT EMRS, ACCURATE ELECTRONIC DATA INTERFACE WILL BE CRITICAL. WE HAVE GIVEN GE MANY OPPORTUNITIES TO RESPOND AND CORRECT THE PROBLEM, BUT TO NO AVAIL. FRANKLY, WE ARE STUNNED BY GE'S LACK OF INTEREST AND RESPONSE. THIS PROBLEM IS THE DIRECT RESPONSIBILITY OF GE AND WILL BE A PROBLEM NO MATTER WHAT EMR IS CHOSEN BY HEALTHCARE INSTITUTIONS. WE URGE YOU TO WARN ANESTHESIA DEPTS AND HOSP IT STAFF THAT THIS IS A PROBLEM. CAPSULE DATACAPTOR AND EPIC - (B) (6)- EMR ANESTHESIA SOFTWARE -ALTHOUGH THIS WOULD HAPPEN WITH ANY EMR-. | 1
1723735 | IT WAS REPORTED THAT WHEN A WEIGHT BASED MEDICATION IS ENTERED INTO THE SYSTEM, THE PT'S ADMISSION WEIGHT OVERRIDES THE PT'S DAILY WEIGHT VALUES THAT HAVE BEEN ENTERED INTO THE SYSTEM. THERE WAS NO PT INJURY REPORTED. (B)(4). >< DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1631725 | THE CENTRAL MONITORING STATION FAILED IN CCU, BLANK SCREEN. SUPERVISOR CONTACTED, BIOMED CONTACTED. THIS MEANS WE CAN NOT SEE THE CENTRAL STATION DATA FOR ANY CRITICAL CARE PATIENT OR PRINT OFF INFORMATION SUCH AS CLINICAL VITAL SIGNS, PRESSURE AND RHYTHM STRIPS. CLINICAL ENGINEERING FOUND THAT THE HARD DRIVE IN THE CENTRAL SERVER HAD MALFUNCTIONED. HARD DRIVE WAS REPLACED, INSTALLED BOOT DISK AND SOFTWARE. | 1
1506833 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE POWERED ON WITHOUT DISPLAY. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1787061 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE INAPPROPRIATELY SHUT DOWN. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. | 0
1787062 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1747822 | THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. >< COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PT, THE DEVICE WAS UNABLE TO PACE. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PT DUE TO THE REPORTED MALFUNCTION. | 0
1564756 | THE DEVICE HAS NOT BEEN RECEIVED FOR EVAL AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE DISPLAYED A 'PREAMP INTERNAL REFERENCE FAILURE' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1564759 | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO CARDIOVERT, A PEDIATRIC PT, THE DEVICE FAILED TO DELIVER SHOCK. COMPLAINANT INDICATED THAT THEY DID NOT HAVE A PEDI ENERGY REDUCER ADAPTER CABLE. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PT DUE TO THE REPORTED MALFUNCTION. >< THE COMPLAINANT INDICATED THAT THE DEVICE WILL NOT BE SENT IN FOR EVAL. | 0
1564758 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVAL AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1632522 | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS BEEN RECEIVED AND WILL BE PROVIDING A FOLLOW-UP REPORT WHEN OUR INVESTIGATION IS COMPLETED. | 0
1747818 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE, THE UNIT TOOK A REPORTED TWO MINUTES TO CHARGE ENERGY. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE CUSTOMER HAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR REPAIR. | 0
1747817 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVAL, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1641492 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE DISPLAYED A 'PACER COMM ERROR' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
1823272 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN REC'D FOR EVAL AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1559341 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETE. | 0
1419035 | THE PRODUCT HAS NOT BEEN RECEIVED FOR EVALUATION AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE RETURNED A 'NO SHOCK ADVISED' DETERMINATION FOR A SIMULATED SHOCKABLE RHYTHM. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1739394 | THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE POWERED ON WITHOUT DISPLAY. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1594582 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE FAILED TO CHARGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1493030 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE FAILED TO CHARGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1493031 | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PT, THE DEVICE FAILED TO DISCHARGE. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PT DUE TO THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION, AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1713689 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED FOR EVAL AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1487118 | THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING THE DEVICE DISPLAYED 'PACER COMM ERROR' AN 'DEFIB COMM ERROR' MESSAGES. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1487115 | THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE DISPLAYED A 'LEAD FAULT' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1627103 | THE PRODUCT HAS NOT BEEN RECEIVED FOR EVAL AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE FAILED TO POWER ON. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1847245 | THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE'S DISPLAY WENT BLANK AND THE DEVICE WOULD NOT POWER UP. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1589132 | COMPLAINANT ALLEGED THAT THE DEVICE WOULD INTERMITTENTLY NOT POWER UP. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETE. | 0
1589134 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE POWERED ON WITHOUT DISPLAY. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1589133 | THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE FAILED SELF-TEST. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1509057 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE DISPLAYED A 'LEAD FAULT' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS BEEN RECEIVED, AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
1787048 | THE DEVICE HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WAS UNABLE TO DETECT THE ATTACHED PADDLES. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1484040 | COMPLAINANT ALLEGED THAT THE DEVICE PROMPTED AN 'ATTACH PADS' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS NOT BEEN RECEIVED AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
1566586 | THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE PROVIDED WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR A (B) (6) FEMALE PT FOR A FEVER, THE DEVICE DISPLAYED A 'LEAD FAULT' MESSAGE. COMPLAINANT INDICATED THAT CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PT. COMPLAINANT INDICATED THAT THERE WAS NO ADVERSE EFFECT TO THE PT DUE TO THE REPORTED MALFUNCTION. | 0
1694780 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE DISPLAYED A 'LEAD FAULT' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1660512 | THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE WAS UNABLE TO PACE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1660510 | COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE FAILED TO POWER ON. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1868906 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE DISPLAYED A 'DEFIB ERROR 11' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. >< THE DEVICE HAS BEEN RECEIVED, AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN OUR INVESTIGATION IS COMPLETED. | 0
1506091 | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO MONITOR A PATIENT, THE DEVICE'S DISPLAY STARTED TO ROLL DOWN AND THE CLINICIAN WAS UNABLE TO SEE THE ECG WAVEFORM. COMPLAINANT INDICATED THAT THERE WAS NO ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION. >< THE DEVICE HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
1791156 | THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE DISPLAYED A 'LEAD FAULT' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1588761 | THE DEVICE HAS BEEN RECEIVED, AND WE WILL BE PROVIDING A FOLLOW-UP REPORT WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE'S DISPLAY BLANKED OUT. COMPLAINANT INDICATED THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1726168 | THE DEVICE HAS BEEN RECEIVED AND WILL BE PROVIDING A FOLLOW-UP REPORT WHEN OUR INVESTIGATION IS COMPLETED. >< COMPLAINANT ALLEGED THAT DURING BIOMED TESTING, THE DEVICE DISPLAYED A 'DEFIB ERROR CODE 11' MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. | 0
1726167 | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO DEFIBRILLATE A (B) (6) FEMALE PATIENT, THE DEVICE INAPPROPRIATELY SHUT DOWN. COMPLAINANT INDICATED THAT THE CLINICIAN OBTAINED ANOTHER DEVICE TO CONTINUE TREATING THE PATIENT. COMPLAINANT DID NOT INDICATE THAT THERE WAS ANY ADVERSE EFFECT TO THE PATIENT DUE TO THE REPORTED MALFUNCTION. >< THE PRODUCT HAS NOT BEEN RECEIVED FOR EVALUATION AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
1589673 | DURING A HYSTEROSCOPY CASE VISIBILITY WAS POOR. THE PHYSICIAN STATES THERE WAS INABILITY TO MAINTAIN UTERINE PRESSURE TO ALLOW VISUALIZATION. FLUID WAS FOUND ON THE FLOOR NEXT TO OPERATIVE SITE. THE PROCEDURE HAD TO BE STOPPED SO PT MAY GO THROUGH AN ADDITIONAL PROCEDURE. >< DEVICE IS NOT BEING RETURNED FOR EVALUATION. (B)(4). | 0
1511232 | CUSTOMER REPORTED THE MPA PIPETTED THE WRONG SAMPLE INTO THE INCORRECT CUPS. CUSTOMER STATES HE HAD THREE SPECIMENS THAT GAVE WRONG RESULTS COMPARED TO THE ORIGINAL TUBE. THE MPA PIPETTED THE SAMPLES AND THEY WERE SENT TO THE INSTRUMENT AND RUN, THE RESULTS DID NOT MATCH RESULTS FROM THE ORIGINAL TUBES. (THE TECH NOTICED THE RESULTS WERE HIGH. HE RAN THE SAMPLES FROM THE ORIGINAL TUBES AND GOT DIFFERENT RESULTS). HE CORRECTED THE RESULTS AND CALLED THE ER DOCTOR WITH THE RIGHT RESULTS, HE SAID THEY HAVE NO WAY OF KNOWING IF THE PATIENTS WERE TREATED. THE FOLLOWING RESULTS WERE PROVIDED, CUSTOMER WAS UNSURE WHICH COBAS ANALYZER SOME OF THESE REPEATS WERE TESTED ON. THREE DIFFERENT TYPES OF SAMPLE CONTAINERS ALL FROM SAME SAMPLE: SAMPLE CONTAINER 1 (MPA CUP, COBAS 2) INITIAL RESULTS, BUN 201 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG); CALCIUM 14.7 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG); CO2 15.6 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG); CREATININE 18.1 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG); SODIUM 191 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG); POTASSIUM 8.94 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE CONTAINER 2 (ORIGINAL TUBE, UNKNOWN COBAS) INITIAL RESULTS, BUN 10 MG PER DL; CALCIUM 9.3 MG PER DL: CO2 24.1 MMOL PER L; CREATINE 1.0 MG PER DL; SODIUM 134 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG); POTASSIUM 5.4 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE CONTAINER 1 (ORIGINAL MPA CUP, COBAS 1) REPEAT RESULTS, BUN RESULT 220 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG); CALCIUM 12.4 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG), CO2 10.6 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG), CREATININE 18.6 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG), SODIUM 193 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG), POTASSIUM 8.9 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE CONTAINER 1 (ORIGINAL MPA CUP, COBAS 2) REPEAT RESULTS, BUN 219 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG), CREATININE 20.1 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE CONTAINER 3 (ORIGINAL TUBE PUT IN NEW CUP, UNKNOWN COBAS) INITIAL RESULTS, BUN 11 MG PER DL, CREATININE 1.0 MG PER DL. >< ADDITIONAL INFORMATION WILL BE PROVIDED IN SUPPLEMENTAL MEDWATCH REPORT. >< ADDITIONAL INFORMATION WILL BE PROVIDED IN SUPPLEMENTAL MEDWATCH REPORT. >< CUSTOMER REPORTED THE MPA PIPETTED THE WRONG SAMPLE INTO THE INCORRECT CUPS. CUSTOMER STATES HE HAD THREE SPECIMENS THAT GAVE WRONG RESULTS COMPARED TO THE ORIGINAL TUBE. THE MPA PIPETTED THE SAMPLES AND THEY WERE SENT TO THE INSTRUMENT AND RUN, THE RESULTS DID NOT MATCH RESULTS FROM THE ORIGINAL TUBES. (THE TECH NOTICED THE RESULTS WERE HIGH. HE RAN THE SAMPLES FROM THE ORIGINAL TUBES AND GOT DIFFERENT RESULTS). HE CORRECTED THE RESULTS AND CALLED THE ER DOCTOR WITH THE RIGHT RESULTS, HE SAID THEY HAVE NO WAY OF KNOWING IF THE PATIENTS WERE TREATED. THE FOLLOWING RESULTS WERE PROVIDED, CUSTOMER WAS UNSURE WHICH COBAS ANALYZER SOME OF THESE REPEATS WERE TESTED ON. TWO DIFFERENT TYPES OF SAMPLE CONTAINERS ALL FROM SAME SAMPLE: SAMPLE 1 (MPA CUP, UNKNOWN COBAS), SODIUM 117 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG), POTASSIUM 6.38 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE 2 (ORIGINAL TUBE, UNKNOWN COBAS), SODIUM 137 MMOL PER L, POTASSIUM 3.4 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG). >< ADDITIONAL INFORMATION WILL BE PROVIDED IN SUPPLEMENTAL MEDWATCH REPORT. >< CUSTOMER REPORTED THE MPA PIPETTED THE WRONG SAMPLE INTO THE INCORRECT CUPS. CUSTOMER STATES HE HAD THREE SPECIMENS THAT GAVE WRONG RESULTS COMPARED TO THE ORIGINAL TUBE. THE MPA PIPETTED THE SAMPLES AND THEY WERE SENT TO THE INSTRUMENT AND RUN, THE RESULTS DID NOT MATCH RESULTS FROM THE ORIGINAL TUBES. (THE TECH NOTICED THE RESULTS WERE HIGH. HE RAN THE SAMPLES FROM THE ORIGINAL TUBES AND GOT DIFFERENT RESULTS). HE CORRECTED THE RESULTS AND CALLED THE ER DOCTOR WITH THE RIGHT RESULTS, HE SAID THEY HAVE NO WAY OF KNOWING IF THE PATIENTS WERE TREATED. THE FOLLOWING RESULTS WERE PROVIDED, CUSTOMER WAS UNSURE WHICH COBAS ANALYZER SOME OF THESE REPEATS WERE TESTED ON. TWO DIFFERENT TYPES OF SAMPLE CONTAINERS ALL FROM SAME SAMPLE: SAMPLE CONTAINER 1 (MPA CUP, UNKNOWN COBAS), BUN 91 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG), CREATININE 13.6 MG PER DL (ACCOMPANIED BY HIGH DATA FLAG), POTASSIUM 8.6 MMOL PER L (ACCOMPANIED BY HIGH DATA FLAG). SAMPLE CONTAINER 2 (ORIGINAL TUBE, UNKNOWN COBAS), BUN 9 MG PER DL, CREATININE 0.6 MG PER DL (ACCOMPANIED BY LOW DATA FLAG), POTASSIUM 3.2 MMOL PER L (ACCOMPANIED BY LOW DATA FLAG). | 0
1828047 | ON 08/25/2010, THE REVIEW OF A DATA CARD, WHICH WAS SENT TO AN END CUSTOMER TO ASSIST WITH A CUSTOMER SERVICE REQUEST, IDENTIFIED THAT THE CARD HAD BEEN FORMATTED INCORRECTLY. THE CARD CONTAINED A TEST FILE WHICH, IF INSTALLED ONTO THE AED, WOULD CAUSE THE DEVICE TO BE IN A TEST MODE AND WOULD PREVENT THE DELIVERY OF THERAPY, IF USED IN A RESCUE ATTEMPT. THERE WAS NO PT INVOLVEMENT. >< FURTHER INVESTIGATION DETERMINED THAT THE DATA CARD CONTAINED A TEST FILE WHICH, IF INSTALLED ONTO THE AED, WOULD CAUSE THE DEVICE TO BE IN A TEST MODE AND, IF USED DURING A RESCUE ATTEMPT, WOULD PREVENT THE DELIVERY OF THERAPY. THE ISSUE WAS IDENTIFIED IMMEDIATELY UPON REVIEW OF THE RETURNED DATA CARD BY DEFIBTECH. DEFIBTECH IS FILING THE MDR AS THE INITIAL REPORTER IN THIS REPORT. THE DATA CARD AND AED WERE REMOVED FROM SERVICE AND REPLACED. | 0
1941468 | AFTER ENTERING THE PT INFORMATION INTO THE NEW MOTILITY COMPUTER, FOLLOWING PROTOCOL, THE CORRECT PROMPTS WERE NOT AVAILABLE WHEN THE PROCEDURE STARTED. THE MOTILITY CATHETER WAS PLACED DURING AN UPPER ENDOSCOPY. THE STUDY WAS STARTED BUT THE USUAL PROMPTS WERE NOT AVAILABLE. THE STUDY WAS CONTINUED BECAUSE IT WAS RECORDING. WE WERE ABLE TO END AND SAVE AS IS NORMALLY DONE. THE PT INFO WAS NOT FOUND IN THE ANALYSIS PORTION. | 1
1887064 | THE ALTERNATE INSTRUMENT'S IMAGE ANALYSIS MEASUREMENTS (IAM) WAS SUGGESTIVE OF A WEAK REACTION AS THE SPC IS NEAR THE THRESHOLD LOW SETTING. FOLLOWING THE CONCLUSION OF THE INVESTIGATION BCI WAS INFORMED BY THE CUSTOMER TECHNICAL SUPPORT (CTS) THAT THE DONOR UNIT WAS TYPED MANUALLY AS C POSITIVE. A REVIEW OF ALL BCI COMPLAINT RECORDS, IN RELEVANCE TO THIS EVENT, DID NOT REVEAL ANY OTHER REPORTS OF FALSE NEGATIVE RESULTS. NO EVIDENCE OF ANALYZER MALFUNCTION WAS OBSERVED IN TEST DATA PROVIDED. THE ROOT CAUSE FOR THIS EVENT IS UNDETERMINED. >< A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO OBTAINING A BLOOD SAMPLE RESULT TYPED AS 'C NEGATIVE' ON THE PK7300 AUTOMATED MICROPLATE SYSTEM. THE BLOOD DONOR HAD PREVIOUSLY TESTED AS 'C POSITIVE' BY AN ALTERNATE INSTRUMENT IN 2009. THE ERRONEOUS RESULT WAS REPORTED OUTSIDE THE LABORATORY BUT A CORRECTED REPORT WAS AMENDED, WHICH WAS TYPED MANUALLY AS 'C POSITIVE.' NO PATIENTS WERE HARMED IN THIS INCIDENT. CUSTOMER TECHNICAL SUPPORT (CTS) CONFIRMED THAT THE DONOR UNIT WAS NOT TRANSFUSED FOR THE PURPOSE OF FILLING AN ORDER FOR 'C NEGATIVE' BLOOD. | 0
1779675 | ONE PATIENT SAMPLE WITH FALSE NEGATIVE READING FOR URINE ERYTHROCYTES. INITIAL READING BY METER WAS NEGATIVE. SAME STRIP READ VISUALLY WAS POSITIVE. IF ADDITIONAL INFORMATION IS RECEIVED, APPROPRIATE NOTIFICATION WILL BE PROVIDED. | 0
1774590 | DEVICE EVALUATION: OTHER, THE DEVICE EVALUATION HAS NOT BEEN COMPLETED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. COMPLAINT DATABASE REVIEW FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. EVALUATION METHOD: DEVICE HISTORY RECORD WAS REVIEWED. COMPLAINT DATA BASE WAS REVIEWED. CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. >< THE ROTATOR BROKE WHEN DISCONNECTING IT FROM THE CATHETER HUB. IT OCCURRED AFTER A LEFT VENTRICULOGRAM PROCEDURE WAS COMPLETED. INJECTOR SETTINGS: FLOW-10ML/SEC, VOLUME-30, PRESSURE-1050 PSI. THERE WAS NO REPORT OF HARM OR INJURY. | 0
1918239 | THE DEVICE WAS NOT AVAILABLE FOR EVALUATION IMMEDIATELY FOLLOWING THE SURGERY; IT WAS RETURNED TO NUVASIVE 19 MONTHS AFTER THE REPORTED EVENT. THE DEVICE WAS TESTED UPON RETURN AND YIELDED NO INFORMATION SUGGESTING IT FAILED TO OPERATE PROPERLY BEFORE, ON, OR AFTER THAT DATE. IT IS UNKNOWN IF NEUROMUSCULAR PARALYTIC AGENTS OR RELAXANTS WERE IN USE AT THE TIME OF REPORTED EVENT. REVIEW OF LABELING NOTES: 'CHRONICALLY COMPRESSED NERVES, OR SEVERELY COMPRESSED NERVES IN AN ACUTE SETTING, ARE KNOWN TO BE LESS SENSITIVE TO DEPOLARIZATION CURRENTS (I.E., HAVE SIGNIFICANTLY HIGHER DEPOLARIZATION CURRENT VALUES). THEY ARE ALSO LESS LIKELY TO DEMONSTRATE SIGNIFICANT CHANGES IN THEIR THRESHOLD DEPOLARIZATION CURRENT VALUES IMMEDIATELY FOLLOWING NERVE DECOMPRESSION. UNDER SUCH CIRCUMSTANCES, EXERCISE CAUTION IN INTERPRETING DISPLAYED DATA.' 'THE NEUROVISION JJB SYSTEM MAY NOT BE EFFECTIVE, AND IS NOT INTENDED FOR USE, WHEN MUSCLE RELAXANTS OR EPIDURAL BLOCKS HAVE BEEN USED FOR, OR IN CONJUNCTION WITH, ANESTHESIA.' >< IT WAS REPORTED THAT DURING AN XLIF SURGERY FOLLOWED BY PERCUTANEOUS PEDICLE SCREW PLACEMENT AT SPINE LEVELS L2-L5 ON (B)(6) 2008 THE NEUROVISION JJB SYSTEM IN USE DID NOT FUNCTION PROPERLY. PER THE PHYSICIAN'S REPORT THE DEVICE READINGS FOR EACH SCREW INDICATED > 20 MA DURING THE CASE WITH NO INDICATION OF PROXIMITY TO NERVE TISSUE. IT WAS ALLEGED THAT FOLLOWING SURGERY THERE WAS SIGNIFICANT PAIN AND NEUROLOGICAL DEFICITS. POST-OPERATIVE CT SCAN ON (B)(6) 2008 NOTED PEDICLE SCREW MALPOSITIONING ON THE LEFT L2 VERTEBRAL BODY. REVISION SURGERY OCCURRED ON (B)(6) 2008 TO REMOVE THE LEFT SIDE PEDICLE SCREWS AND ROD. A SECOND CT SCAN ON (B)(6) 2008 CONFIRMED THE L2 MALPOSITION, AS WELL AS RIGHT-SIDED L2-L4 PEDICLE SCREW MALPOSITION. THE RIGHT SIDE PEDICLE SCREW CONSTRUCT WAS REMOVED ON (B)(6) 2009. THE PATIENT'S PAIN AND NEUROLOGICAL DEFICITS WERE REPORTED TO HAVE CONTINUED TO THE PRESENT DATE. | 0
1848927 | THREE STUDIES THAT WERE NOT SENT TO PACS THAT ARE NOW NO LONGER AVAILABLE ON THAT DEVICE.ACCORDING TO OUR INTERNAL BIOMED PERSONNEL, THE REASON THE CR (COMPUTED RADIOGRAPHY) WORKSTATION COULD NOT SEND THE IMAGES WAS THAT THEY WERE SOMEHOW CORRUPTED. THE DOCTOR CAME AND READ THEM OFF THE WORKSTATION AND THEN WE HAD TO INITIALIZE THE PATIENT AND IMAGE DATABASE TO FREE UP THE UNIT TO START SENDING IMAGES TO PACS AGAIN. THIS UNIT IS UP AND FULLY FUNCTIONAL AT THIS TIME, BUT THE IMAGES FROM THAT ONE PATIENT ARE NOT IN THE DATABASE ANY MORE. ALSO I HAVE PUT IN A WORKORDER WITH FUJI TO REPAIR THE PRINTER. THE FMDPL PRINTER HAS A PRINT SERVER THAT SITS IN AN OFFICE DOWN THE HALL. THE HARD DRIVE CRASHED ON THAT AND WE CANNOT PRINT ANYTHING STRAIGHT TO THE PRINTER. THEY HAVE TO USE THE PRINT FUNCTION FROM THE PACS SYSTEM UNTIL WE CAN SEE IF THE OLD PRINT SERVER CAN BE REPAIRED.BIOMED HAS TALKED WITH THE FUJI REPRESENTATIVE. WE VERIFIED WITH HIM THAT THE ONLY WAY TO GET THE SYSTEM BACK UP DUE TO CORRUPTED IMAGES WAS TO INITIALIZE THE DATABASE. HE DID RECOMMEND THAT ALL OF THE WORKSTATIONS BE TURNED OFF AND RE-BOOTED ONCE A DAY TO HELP THE PERFORMANCE OF THESE UNITS AND TO REDUCE THE POSSIBILITIES OF DATA CORRUPTION. HE LISTED ANOTHER POSSIBILITY THAT WOULD CAUSE THIS ISSUE, AND THAT IS IF THE NETWORK DROPPED OUT OR HAD SOME SORT OF ISSUE DURING THE SENDING OF THE IMAGE. BIOMED WILL TALK TO FUJI ABOUT REPLACING THE NETWORK CARD IN THAT UNIT JUST TO BE SURE IT IS NOT PART OF THE PROBLEM. | 1
1759844 | MCKESSON HAS IDENTIFIED THAT WHEN REPORTING ECG FILES FROM MORTARA ECG CARTS USING THE DICOM INTERFACE, THE HORIZON CARDIOLOGY ECG MANAGEMENT SYSTEM DOES NOT DISPLAY THE AMPLITUDES OF THE WAVEFORM PROPERLY.====================== HEALTH PROFESSIONAL'S IMPRESSION======================THE ECG WAVEFORM AMPLITUDE IS 20% HIGHER OR LOWER THAN IT ACTUALLY IS IN COMPARISON TO THE CART PRINTOUT, AFFECTING THE WAVEFORM ANALYSIS ALGORITHM IN THE ECG MANAGEMENT SYSTEM. THIS INACCURACY MAY RESULT IN FALSE NEGATIVE OR FALSE POSITIVE DIAGNOSIS. CARDIOLOGIST'S BELIEVE PATIENT SAFETY IS LOW.======================MANUFACTURER RESPONSE (ACCORDING TO REPORTER) FOR ECG MANAGEMENT, HORIZON CARDIOLOGY======================A PATCH IS IN TEST ENVIRONMENT. | 1
1876106 | VITEK 2 REPORTED INCORRECTLY THAT A C. FREUNDII ISOLATE WAS AMOXICILLIN/CLAVULANATE SUSCEPTIBLE. THIS ORGANISM IS ROUTINELY AMOXICILLIN/CLAVULANATE RESISTANT. THE PT WAS GIVEN ANTIBIOTIC THERAPY BASED ON THIS INCORRECT SUSCEPTIBILITY DATA. | 0
1642078 | PATIENT MONITOR BROKE OFF ITS MOUNT ADAPTER ON THE GCX ROLLING MONITOR STAND. THE MOUNT BROKE ON THE STAND, CAUSING THE MONITOR TO FALL. THE MONITOR WAS DAMAGED AS A RESULT. | 1
1907551 | PLEASE NOTE THAT THE OPERATORS MANUAL INDICATES A WARNING AND A CAUTION REGARDING THE PRIMING OF THE KIT. WARNING: OPENING THE RETURN LINE ROLLER CLAMP AFTER VENIPUNCTURE AND BEFORE THE BLOOD SAMPLING POUCH IS FILLED MAY RESULT IN AN AIR EMBOLISM. CAUTION: THE KIT SHOULD BE ALMOST COMPLETELY FULL OF FLUID (SALINE AND ACD) AT THE END OF PRIME. THE CONTAINERS ON THE FRONT SCALE HOOKS SHOULD HAVE LITTLE OR NO FLUID IN THEM EXCEPT FOR THE CONTAINER ON THE FRONT LEFT SCALE HOOK ON A DOUBLE NEEDLE OR MNC KIT, OR THE CONTAINER ON THE FRONT RIGHT SCALE HOOK ON A SINGLE NEEDLE KIT. SOME AIR SHOULD BE EXPECTED NEAR THE FISTULA (S). IF THERE ARE LARGE AMOUNTS OF AIR IN THE KIT IMMEDIATELY FOLLOWING PRIME, DO NOT BEGIN COLLECTION. REMOVE THE KIT AND TRY INSTALLING AND PRIMING ANOTHER KIT. >< A CUSTOMER REPORTED THAT WHILE SETTING UP FOR AN AMICUS PROCEDURE, THEY PERFORMED A RETURN LINE VENIPUNCTURE WITHOUT PRIMING THE RETURN LINE. THE ROLLER CLAMP WAS THEN OPENED AND THE DONOR RECEIVED THE TUBING FULL OF AIR. THE OPERATOR OF THE DEVICE CONTACTED FENWAL INQUIRING IF IT WAS OK TO DO ANOTHER PROCEDURE. THE FENWAL REPRESENTATIVE INDICATED THEY SHOULD IMMEDIATELY NOTIFY THEIR MEDICAL DIRECTOR OF THE EVENT. THE OPERATOR ALSO WANTED TO KNOW APPROXIMATELY HOW MANY MLS OF AIR IS IN THE TUBING. A FOLLOW UP CALL TO THE CUSTOMER WAS MADE AND THE CUSTOMER INDICATED THAT THEY HAD CALLED AN AMBULANCE BECAUSE THE DONOR WAS COUGHING AND THE MEDICAL DIRECTOR HAD COME TO EVALUATE THE DONOR. FENWAL INFORMED THE CUSTOMER THAT THERE IS APPROXIMATELY .2MLS IN EACH INCH OF TUBING. SEVERAL ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL DETAILS AND THE FACILITY WOULD NOT PROVIDE ANY DETAILS ON THE DONOR OR THE DONOR'S CONDITION. | 0
1837909 | A MARK WAS LEFT ON THE PT'S LEG 2 CM BY .5 CM DURING THE DELIVERY. IT BLISTERED. THE MARK CAME FROM A LEG PLATE. (B)(6) WAS NOTIFIED WHO IN TURN NOTIFIED (B)(4), THE STUDY SPONSOR. PT HAD A C-SECTION ON 9/9. A PHOTO OF THE MARK WENT TO (B)(6) WHO SENT IT TO (B)(4). | 0
1644470 | THE ADVERSE EVENT WAS CAUSED BY USER ERRORS AND DEVICE EVAL IS NOT APPLICABLE. (B) (4). VARIAN MEDICAL SYSTEMS, INC. BECAME AWARE OF THE ADVERSE EVENT THROUGH POST MARKET SURVEILLANCE ACTIVITIES WHEN THE EVENT WAS REPORTED IN THE (B) (4) ON MARCH 4, 2010. THE CUSTOMER HAD NOT PREVIOUSLY CONTACTED VARIAN MEDICAL SYSTEMS. >< ON (B) (6) 2010, A PT WAS TREATED BY BRACHYTHERAPY FOR PROSTATE CANCER WITH SIXTY FIVE (65) IODINE-125 (I-125) SEEDS. THE I-125 SEEDS WERE IMPLANTED USING REAL TIME DOSIMETRY UNDER ULTRASONIC GUIDANCE. THE WRITTEN DIRECTIVE CALLED FOR A THERAPEUTIC RADIATION DOSE OF 145 GRAY (GY) [14,500 RAD] TO THE PROSTATE VOLUME, PLUS 5 MILLIMETERS OF MARGIN, USING INTEROPERATIVE PLANNING. ON (B) (6) 2010, THE PT RETURNED FOR A 30 DAY POST-IMPLANT CT. THE SCAN SHOWED THAT THE IMPLANTED SEEDS WERE 'IN AN APPROPRIATE PATTERN', BUT OUTSIDE THE INTENDED TARGET. THE INITIAL (B) (6) REPORT WAS MADE BY PHONE TO THE STATE DEP ON (B) (6). THE DOCTORS PERFORMING THE PROCEDURE - WHO ARE NOT NAMED - PLANNED TO DELIVERY RADIATION TO THE ENTIRE PROSTATE PLUS A FIVE-MILLIMETER MARGIN OUTSIDE THE GLAND. THERE IS NO INFO ON THE STATUS OF THE PT. THE SITE WAS USING VARISEED TREATMENT PLANNING SYSTEM. HOSPITAL CONTACT HAS STATED THAT VARISEED DID NOT CAUSE OR CONTRIBUTE TO THE ADVERSE EVENT. THE USER SITE ORIGINALLY REPORTED THERE HAD BEEN A MALFUNCTION OF THE ULTRASOUND MACHINE, BUT HAVE SINCE RETRACTED THAT CLAIM. | 1
1576965 | IT WAS REPORTED THAT A HARD DRIVE FAILURE OCCURRED ON THE CDT LAN, RESULTING IN TOWER ERRORS AND LOSS OF COMMUNICATION. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1487633 | THE USER NOTICED THE PROVIDED COORDINATES WERE NOT AS INTENDED AND DID NOT APPLY THE SURGERY WITH THOSE COORDINATES. THE USER USED AN IMAGING DATA SET THAT WAS NOT INTENDED FOR THE SURGERY. CORRECTIVE AND PREVENTIVE ACTIONS: SINCE THERE WAS NO QUALITY OR PERFORMANCE ISSUE OR MALFUNCTION OF BRAINLAB EQUIPMENT, THERE ARE NO REMEDIAL ACTIONS INTENDED BY BRAINLAB AT THIS POINT OF TIME. >< THE USER USED AN IMAGING DATA SET THAT WAS NOT INTENDED FOR SURGERY AND REPORTED THAT INCORRECT COORDINATES FOR A BRAIN TRAJECTORY USING A STEREOTACTIC ARC HAD BEEN PROVIDED BY THE BRAINLAB PLANNING SOFTWARE. THE USER NOTICED THE PROVIDED COORDINATES WERE NOT AS INTENDED AND DID NOT APPLY THE SURGERY WITH THOSE COORDINATES. THERE WAS NO QUALITY OR PERFORMANCE ISSUE OR MALFUNCTION OF BRAINLAB EQUIPMENT. | 1
1832153 | THE CUSTOMER STATES THAT THEY HAVE BEEN GENERATING DISCREPANT RESULTS FOR MULTIPLE ASSAYS ON THE ARCHITECT C16000 ANALYZER. THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED AN INITIAL SODIUM ASSAY RESULT OF 136 MMOL/L THAT RETESTED AT 180 MMOL/L. NO SUSPECT RESULTS HAVE BEEN REPORTED FROM THE LAB. A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). AN ABBOTT FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE ON SEVERAL OCCASIONS TO TRY AND RESOLVE THE ISSUE OF ERRATIC RESULTS. DURING THE PROCESS OF TROUBLESHOOTING, MANY OF THE PRIMARY OPTICS, DISPENSE, ASPIRATION, AND WASH COMPONENTS WERE REPLACED, CLEANED, AND/OR ALIGNED. RESOLUTION WAS ACHIEVED WHEN THE FSE REPLACED THE SYRINGE SEAL 1, SEAL 2 AND O-RING FOR THE R2A SYRINGE. IT WAS ALSO NOTED THAT THE 1B PROBE WAS LEAKING WATER ON THE REAGENT CAROUSEL COVER. THE FSE THEN REPLACED THE PROBE WASH BELLOW AND POPPET VALVES, PERFORMED A SYSTEM FLUID FLUSH AND THE DAILY MAINTENANCE PROCEDURE. THIS APPEARS TO SUGGEST (BUT NOT CONFIRM) THAT THE POSSIBLE CAUSE IS RELATED TO THE LEAKING R2A SYRINGE AND/OR A LEAKING INTERNAL PROBE WASH PUMP. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (201837-108) CONTAINS INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. BECAUSE MULTIPLE COMPONENT/PART REPLACEMENT WAS REQUIRED TO RESOLVE THE CUSTOMER ISSUE ASSOCIATED WITH THIS COMPLAINT, A SINGLE DEFINITIVE CAUSE WAS NOT ESTABLISHED. BASED ON THE INFORMATION PROVIDED, A MALFUNCTION CANNOT BE IDENTIFIED. METHOD: REVIEW OF COMPLAINT TRACKING AND TRENDING. | 0
1757205 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER HAS BEEN GENERATING SPORADIC CALCIUM RESULTS OF <0.5 MMOL/L. UPON REPEAT THE RESULTS ARE IN THE NORMAL RANGE. THEY HAVE NOT SEEN THIS PROBLEM ON THEIR SECOND C8000 ANALYZER. NO SPECIFIC PATIENT DATA WAS PROVIDED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1509085 | (B) (4) THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) ARRIVED AT THE CUSTOMER SITE AND INSPECTED THE ANALYZER. THE TUBING FROM THE ICT (INTEGRATED CHIP TECHNOLOGY) MODULE WAS DISCOLORED AND THE ICT PROBE WAS LOOSE. THE DISCOLORED TUBING AND ICT PROBE WERE REPLACED AS WELL AS THE SAMPLE PROBE. THE FSR VERIFIED THE ANALYZER'S PERFORMANCE WITH ACCEPTABLE PRECISION RUNS AND INDICATED THE PROBABLE CAUSE FOR THE CUSTOMER'S ISSUE WAS A LOOSE ICT PROBE. THE CUSTOMER WAS MADE AWARE THAT THE ICT PROBE MUST BE TIGHTENED IN POSITION WHEN RUNNING. NO SYSTEM LOGS WERE AVAILABLE FOR REVIEW FOR THIS EVALUATION. COMPLAINT ACTIVITY WAS REVIEWED AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL ((B)(4), (B)(9) 2009) AND ICT SAMPLE DILUENT PACKAGE INSERT ((B)(4)) PROVIDE SUFFICIENT INFORMATION REGARDING SCHEDULED MAINTENANCE, COMPONENT REPLACEMENT, SAMPLE HANDLING, INTERPRETING RESULTS AND TROUBLESHOOTING THE CUSTOMER'S CURRENT ISSUE. BASED ON THE AVAILABLE INFORMATION, THE DISCREPANT RESULTS WERE PROBABLY CAUSED BY THE LOOSE ICT PROBE; HOWEVER, THE 1.0 ML SYRINGES REPLACED BY THE CUSTOMER AND THE DISCOLORED ICT TUBING AND SAMPLE PROBE REPLACED BY THE FSR ARE ALSO POSSIBLE CAUSES. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT DISCREPANT RESULTS WERE REPORTED FROM THE LAB THAT WERE GENERATED ON THE ARCHITECT C8000 ANALYZER FOR THE CHLORIDE ASSAY. THE CUSTOMER STATES THAT CONTROLS HAD BEEN WITHIN SPECIFICATION FROM 9 AM THROUGH 1 PM, BUT BY 3 PM CONTROL VALUES HAD DROPPED. THE CUSTOMER RETESTED SUSPECT SAMPLES ON THE ARCHITECT C16000 ANALYZER AND ISSUED CORRECTED REPORTS. NO PATIENT MANAGEMENT WAS AFFECTED BY THE REPORTED RESULTS. THIS PATIENT GENERATED AN INITIAL RESULT OF 106 MMOL/L THAT RETESTED AT 116 MMOL/L. A SERVICE CALL WAS INITIATED. | 0
1594513 | THE CUSTOMER STATED CONTINUOUS IMPRECISION WAS OCCURRING ON THE ARCHITECT C8000 ANALYZER. ONE PATIENT RESULT GENERATED A FALSELY ELEVATED ARCHITECT CO2 RESULT OF 41 MEQ/L. WHEN REPEATED, THE SAMPLE GENERATED A CO2 RESULT OF 25.3 MEQ/L. THE CUSTOMER VERIFIED THE RESULT ON ANOTHER ARCHITECT C8000 OBTAINING A RESULT OF 25 MEQ/L. THE CUSTOMER CHECKED THE SAMPLE INTEGRITY AND NOTED THE SAMPLE WAS SHORT; HOWEVER, THE ANALYZER DID NOT GENERATE AN ASPIRATION ERROR. THERE WERE NO RESULTS REPORTED OUTSIDE OF THE LABORATORY AND NO IMPACT TO PATIENT MANAGEMENT OCCURRED. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW UP MDR WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. OTHER: AN INVESTIGATION IS IN PROCESS. | 0
1557640 | (B)(4). THE CUSTOMER FOUND THAT NOZZLE 1 ON THE CUVETTE WASHER WAS DRIPPING AND CHANGED THE POPPET VALVE TO RESOLVE THE ISSUE. AN INVESTIGATION WAS INITIATED TO FURTHER INVESTIGATE THE CUSTOMER ISSUE. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND A PRODUCT SAFETY REVIEW. TRACKING AND TRENDING DID NOT INDICATE ANY ADVERSE TREND FOR THE ARCHITECT C8000 FOR THE REPORTED COMPLAINT ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE ARCHITECT C8000 FOR FALSELY DECREASED RESULTS. >< THE CUSTOMER STATED AN ARCHITECT C8000 ANALYZER GENERATED FALSELY DECREASED SODIUM, POTASSIUM AND CHLORIDE RESULTS FOR ONE PATIENT SAMPLE. THE SODIUM RESULTS OF 117, POTASSIUM RESULTS OF 3.3 AND CHLORIDE RESULTS OF 88 WERE QUESTIONED BY THE HEALTHCARE PROVIDER. UPON REPEAT, SODIUM RESULTS OF 140, POTASSIUM RESULTS OF 4.2 AND CHLORIDE RESULTS OF 99 WERE GENERATED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1838277 | (B)(4). THE CUSTOMER REPORTED VISIBLE SMOKE COMING FROM THE SYSTEM CONTROL CENTER (SCC) CENTRAL PROCESSING UNIT (CPU), BUT NO SPARKS OR FIRE, AND NO INJURIES. THE CUSTOMER UNPLUGGED THE CPU. THE FSR REPLACED THE POWER SUPPLY FOR THE SCC CPU AND VERIFIED THAT THE INSTRUMENT WAS OPERATIONAL. THE CUSTOMER'S COMPLAINT HISTORY DOES NOT INCLUDE A RECURRENCE OF THE ISSUE. THE SUSPECT POWER SUPPLY UNIT WAS NOT RETURNED FOR EXAMINATION DUE TO IT BEING DISCARDED BY THE CUSTOMER; THEREFORE, NO IDENTIFICATION CAN BE MADE AS TO THE EXACT CAUSE OF THE EVENT. A REVIEW OF DOCUMENTATION DID NOT IDENTIFY A PRODUCT ISSUE OR AN ADVERSE TREND RELATED TO THE ISSUE IN THIS COMPLAINT. THE ARCHITECT SYSTEM OPERATIONS MANUAL PROVIDES ADEQUATE INSTRUCTIONS REGARDING ELECTRICAL HAZARDS AND AN EMERGENCY SHUTDOWN PROCEDURE. THE C8000 IS CERTIFIED TO THE APPROPRIATE US, (B)(6) AND INTERNATIONAL SAFETY STANDARDS TO ENSURE THAT THE DESIGN PROVIDES ADEQUATE PROTECTION AGAINST SPREAD OF FIRE FROM THE EQUIPMENT. BASED ON THE COMPLAINT DOCUMENTATION THE SCC DID NOT VIOLATE THE MAIN OBJECTIVES OF THE STANDARDS FOR SAFETY. BASED ON THE INVESTIGATION FINDINGS, A SPECIFIC CAUSE WAS NOT IDENTIFIED AND THERE IS NO INDICATION OF A DEFICIENCY OF THE ARCHITECT C8000 SYSTEM. >< THE CUSTOMER STATES THAT SMOKE WAS OBSERVED COMING FROM THE POWER SUPPLY UNIT (PSU) OF THE ARCHITECT C8000 ANALYZER. THERE WAS NO REPORT OF IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT. NO INJURIES OR DAMAGE TO THE SURROUNDING ENVIRONMENT WERE REPORTED RELATED TO THIS ISSUE. | 0
1660551 | THE CUSTOMER STATES THAT THEY OBSERVED THE ARCHITECT C8000 ANALYZER SAMPLING FROM AN EMPTY/NEARLY EMPTY TOTAL BILIRUBIN REAGENT CARTRIDGE. ONE PATIENT SAMPLE TESTED DURING THIS TIME GENERATED A RESULT OF 3 MG/DL THAT WAS REPORTED FROM THE LAB. THE SAMPLE RETESTED AT 12 MG/DL AND A CORRECTED REPORT WAS ISSUED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. A SERVICE CALL WAS INITIATED. >< (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. CC BILIRUBIN CALIBRATOR, LN: 1E66-04, LOT#: 73372M100. | 0
1639983 | THE CUSTOMER WAS RETRIEVING A BOTTLE OF DETERGENT B. SHE SHOOK THE CONTAINER WHICH WAS UNCAPPED, CAUSING THE DETERGENT TO SPLASH INTO HER EYE. SHE IMMEDIATELY FLUSHED HER EYE WITH WATER. THE CUSTOMER WENT TO SEE AN OPHTHALMOLOGIST AND WAS PRESCRIBED LACRIFILM OINTMENT AND TYLENOL. NO HARM OCCURRED TO THE EYE BUT THE EYELID WAS VERY IRRITATED. NO INFORMATION WAS PROVIDED REGARDING WHETHER THE TECHNICIAN WAS WEARING EYE PROTECTION AT THE TIME OF THE INCIDENT. >< (B) (4). THE ARCHITECT SYSTEM OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY DESCRIBE THE HAZARDS AND PRECAUTIONS FOR HANDLING DETERGENT B. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED RELATED TO THE COMPLAINT ISSUE. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. | 0
1553915 | (B)(4). SAMPLE PROBE A LIST#: 9D46-02. AFTER RECEIVING LOWER THAN EXPECTED RESULTS FOR THE SODIUM AND POTASSIUM ASSAYS ON THE AEROSET ANALYZER, THE CUSTOMER NOTICED THAT THE ANGLE OF SAMPLE PROBE A WAS ALMOST FLAT. THE CUSTOMER REPLACED THE DAMAGED PROBE AND SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE. A REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT IDENTIFY ANY OTHER ISSUES WITH THE SYSTEM. THE INVESTIGATION CONSISTED OF A REVIEW OF COMPLAINT ACTIVITY FROM (B)(6) 2009 THROUGH (B)(6) 2009 AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE AEROSET OPERATIONS MANUAL, LIST NUMBER 9D06-05, CONTAINS INFORMATION ADDRESSING THE CURRENT CUSTOMER ISSUE. BASED ON THE AVAILABLE INFORMATION, A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. >< THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED INITIAL SODIUM RESULTS OF 108 MMOL/L AND POTASSIUM RESULTS OF 3.28 MMOL/L ON THE AEROSET ANALYZER. THE SAMPLE WAS RETESTED ON THE OTHER AEROSET ANALYZER IN THE LAB AND GENERATED SODIUM RESULTS OF 141 MMOL/L AND POTASSIUM RESULTS OF 4.18 MMOL/L, WHICH THE CUSTOMER STATES ARE CORRECT. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. UPON INSPECTION OF THE SUSPECT ANALYZER, THE CUSTOMER FOUND THAT THE TIP OF SAMPLE PROBE A WAS ALMOST FLAT. THE CUSTOMER REPLACED SAMPLE PROBE A. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). AEROSET SAMPLE PROBE A LN: 9D46-02. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1553921 | (B)(4). AEROSET SAMPLE PROBE A LN: 9D46-02. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED INITIAL SODIUM RESULTS OF 110 MMOL/L AND POTASSIUM RESULTS OF 3.18 MMOL/L ON THE AEROSET ANALYZER. THE SAMPLE WAS RETESTED ON THE OTHER AEROSET ANALYZER IN THE LAB AND GENERATED SODIUM RESULT OF 141 MMOL/L AND POTASSIUM RESULT OF 4.03 MMOL/L, WHICH THE CUSTOMER STATES ARE CORRECT. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. UPON INSPECTION OF THE SUSPECT ANALYZER, THE CUSTOMER FOUND THAT THE TIP OF SAMPLE PROBE A WAS ALMOST FLAT. THE CUSTOMER REPLACED SAMPLE PROBE A. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. >< (B)(4). SAMPLE PROBE A LIST#: 9D46-02. AFTER RECEIVING LOWER THAN EXPECTED RESULTS FOR THE SODIUM AND POTASSIUM ASSAYS ON THE AEROSET ANALYZER, THE CUSTOMER NOTICED THAT THE ANGLE OF SAMPLE PROBE A WAS ALMOST FLAT. THE CUSTOMER REPLACED THE DAMAGED PROBE AND SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE CUSTOMER REQUIRED NO FURTHER ASSISTANCE. A REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT IDENTIFY ANY OTHER ISSUES WITH THE SYSTEM. THE INVESTIGATION CONSISTED OF A REVIEW OF COMPLAINT ACTIVITY FROM (B)(6) 2009 AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE AEROSET OPERATIONS MANUAL, LIST NUMBER 9D06-05, CONTAINS INFORMATION ADDRESSING THE CURRENT CUSTOMER ISSUE. BASED ON THE AVAILABLE INFORMATION, A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. THIS IS A FINAL REPORT. | 0
1483631 | (B) (4), CONCOMITANT MEDICAL PRODUCTS: ANALYZER LIST NUMBERS 8C89 AND 3M74, PRESSURE MONITORING LIQUID LEVEL SENSE KIT, U10 CHIP AND EJECTORS, PART NUMBER 7-200788-01, AND PRESSURE MONITORING ALGORITHM LIQUID LEVEL SENSE BOARD, PART NUMBER 7-94255-01. INVESTIGATION IN PROCESS, NO METHOD, RESULTS OR CONCLUSION CODE CAN BE CHOSEN AT THIS TIME. THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. >< (B)(4). EVALUATION METHOD (B)(4) OTHER, A REVIEW OF THE SOFTWARE CODE, BRAIN STORMING MEETINGS, HUMAN ERROR AND FIVE WHYS ANALYSIS WAS PERFORMED TO DETERMINE ROOT CAUSE OF THE EVENT. THE ROOT CAUSE OF THIS ISSUE WAS IDENTIFIED AS FAILURE TO DOCUMENT THAT ARCHITECT I2000 AND I2000SR INSTRUMENTS WOULD REQUIRE A CHANGE TO THE SLOPE SCORE THRESHOLDS WHEN SOFTWARE V5.00 OR HIGHER IS USED IN CONJUNCTION WITH LLS/PM BOARD FIRMWARE V15 OR 17. THIS ISSUE WAS IDENTIFIED DURING SOFTWARE DEVELOPMENT OF V5.00, BUT NOT CLEARLY DEFINED IN THE SOFTWARE RELEASE ACTION PLAN. IN RESPONSE TO THIS ISSUE TWO MANDATORY TECHNICAL SERVICE BULLETINS WERE ISSUED TO ABBOTT FIELD SERVICE TO CHANGE THE SLOPE SCORE THRESHOLDS ON ALL AFFECTED ANALYZERS. ALL AFFECTED ANALYZERS HAVE BEEN SERVICED. IN ADDITION, ARCHITECT SYSTEM SOFTWARE V7.0 WILL INCLUDE THIS RESOLUTION. >< ABBOTT ARCHITECT SYSTEM SOFTWARE (B) (4) AND (B) (4) ARE NOT COMPATIBLE WITH THE ENHANCED PRESSURE MONITORING (PM) ALGORITHM LIQUID LEVEL SENSE (LLS) BOARD OR PRESSURE MONITORING LIQUID LEVEL SENSE KIT, U10 CHIP AND EJECTORS. SOFTWARE (B) (4) AND (B) (4) USE INCORRECT SLOPE SCORE THRESHOLDS FOR THE ENHANCED LLS BOARD. THIS CAUSED AN INCREASE IN PRESSURE MONITORING ERRORS FOR THOSE CUSTOMERS WHO HAVE SOFTWARE (B) (4) OR (B) (4) AND THE ENHANCED LLS BOARD BECAUSE THE THRESHOLD RANGE IS TOO NARROW. THE ISSUE MAY OCCUR ON THE I2000 AND I2000SR MODULES ONLY, AND MAY RESULT IN A DELAY OF RESULTS. A PRODUCT CORRECTION HAS BEEN ISSUED AND REPORTED UNDER 21 CFR806 TO THE FDA, (B) (4), UNDER THE REMEDIAL ACTION (B) (4). NO ADVERSE IMPACT TO PATIENT MANAGEMENT HAS BEEN REPORTED RELATED TO THIS ISSUE. | 1
1609397 | AROUND 3:30 AM ON (B) (6) 2010, I WOKE UP WITH SOME CHEST PAIN. HAVING DECIDED TO MEASURE BLOOD PRESSURE, I PUT A MONITOR CUFF ON. THE MICROLIFE BLOOD PRESSURE MONITOR IS AN AUTOMATED UNIT THAT TAKES 3 CONSEQUENT MEASUREMENTS. IT PUMPED AIR ONCE, RELEASED THE PRESSURE, THEN PUMPED THE SECOND TIME AND RELEASED THE PRESSURE. AFTER THE THIRD PUMP, THE UNIT MALFUNCTIONED. THE AIR PUMP BROKE WHILE IT WAS SQUEEZING MY ARM WITHOUT RELEASING AIR. THIS HAPPENED DEEP AT NIGHT AFTER WAKING UP, AND I DID NOT THINK ABOUT TAKING THE CUFF OFF, I JUST LOOKED AT THE UNIT DISORIENTED AND THEN, I LOST CONSCIOUSNESS. MY BODY FELL ON THE GROUND. THIS WOKE UP MY DAUGHTER WHO CALLED AN AMBULANCE. I WAS TAKEN TO THE EMERGENCY ROOM AND THEN HOSPITALIZED FOR OBSERVATIONS. DATES OF USE: 3 YEARS. | 0
1046541 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER INTERPRETED A POSITIVE ANTIBODY SCREEN RESULT AS NEGATIVE. VISUAL INSPECTION OF THE GEL CARD SHOWED THE REACTION WAS POSITIVE (1+). NO ERRONEOUS RESULTS WERE REPORTED. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND PERFORMED ADJUSTMENTS TO THE CAMERA, HANDLER AND CENTRIFUGE. THE FE ALSO REPLACED THE CRACKED WASHED BLOCK AND OTHER COMPONENTS. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). | 0
1329082 | THE CUSTOMER REPORTED WHEN A SERIES WITH A LARGE NUMBER OF IMAGES (160-180) AND A SERIES NUMBER OF 400 OR 500 RECONSTRUCTS AND IS LOADED INTO A VIEWING WINDOW ON THE RA1000 WORKSTATION, AND AN ATTEMPT IS MADE TO FLIP THE IMAGES IN STACK MODE, AND NOT ALL IMAGES WILL FLIP. >< THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. (B)(4). | 1
1456022 | THE CUSTOMER REPORTED THE ORTHO PROVUE ANALYZER RESULTED A SAMPLE CONTAINING ANTI-E ANTIBODY AS NEGATIVE USING 0.8% RESOLVE PANEL A LOT#VRA128. SINCE ANTIBODY SCREEN RESULTS WERE INITIALLY POSITIVE WITH THE SAMPLE, THE CUSTOMER PERFORMED A FULL CROSSMATCH ON TWO UNITS AND OBTAINED NEGATIVE/COMPATIBLE RESULTS. BOTH UNITS WERE TRANSFUSED SUCCESSFULLY WITH NO REPORTS OF ANY ADVERSE REACTIONS. TESTING OCCURRED ON (B) (6) 2009. THE UNITS TRANSFUSED WERE CONFIRMED TO BE E NEGATIVE. REPEAT TESTING PERFORMED ON THE PROVUE USING THE SAMPLE AND ALTERNATE LOT OF REAGENTS AND ALSO LOT#VRA128 SHOWED POSITIVE RESULTS WITH E+ DONORS. THE CUSTOMER WAS UNABLE TO DUPLICATE THE FALSE RESULTS ON THE PROVUE. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED REGARDING THE FALSE NEGATIVE ANTIBODY IDENTIFICATION RESULTS INITIALLY OBTAINED ON THE PROVUE WITH THE PANEL CELLS. THE CUSTOMER REPORTED THAT THEY PERFORMED ADDITIONAL TESTING WITH 0.8% RESOLVE PANEL B CELLS AND OBTAINED POSITIVE (1+) REACTIVITY WITH E+ HOMOZYGOUS CELLS (E+E-), EXCEPT FOR CELL #20 (E+E+) WHICH REACTED NEGATIVE WITH THE SAMPLE. THE CUSTOMER FURTHER REPEATED THE SAMPLE ON THE ANALYSER USING THE SAME 0.8% RESOLVE PANEL A LOT# VRA128 AND OBTAINED A 1+ REACTION WITH BOTH E+ DONORS ON THE PANEL (CELL #3 E+, E- AND 6# E+E+) WHICH WAS INITIALLY NEGATIVE. THE CUSTOMER WAS UNABLE TO DUPLICATE THE NEGATIVE REACTIVITY USING LOT# VRA128 WITH THE SAMPLE. ALL TESTING WAS ACCEPTABLE UPON REPEAT TESTING. THEREFORE, SERVICE WAS NOT REQUIRED,. INSTRUMENT MALFUNCTION COULD NOT BE IDENTIFIED. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4) | 0
1420029 | THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER WAS DRIPPING. INFORMATION WAS NOT PROVIDED REGARDING POSSIBLE RESULT CODES. NO ERRONEOUS RESULTS WERE REPORTED. PROBE ISSUES MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS, WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND DETERMINED THE PROBE WAS BENT. THE FE REPLACED THE PROBE AND WASH STATION AND PERFORMED FINE PROBE ADJUSTMENTS. THE FE ALSO PERFORMED WAD DIAGNOSTICS LEVEL SENSING, REAGENT, SAMPLE IDENTIFICATION AND PIPETTING. REPLACEMENT OF THE PROBE AND WASH STATION AND THE APPROPRIATE ADJUSTMENTS HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THE CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). | 0
1481854 | AN OCD FIELD ENGINEER WAS AT THE CUSTOMER SITE AND VERIFIED THAT ALL CAMERA ADJUSTMENTS WERE FINE. THE CUSTOMER PERFORMED TESTING IN MANUAL GEL, AND WEAK TO 1+ REACTIONS WERE OBTAINED FOR SAMPLES THAT THE PROVUE RESULTED AS NEGATIVE. ALTHOUGH THE TESTING IN MANUAL GEL YIELDED POSITIVE RESULTS, THE TESTING WAS INCONCLUSIVE DUE TO THE INCONSISTENT RESULTS BETWEEN TWO TECHNOLOGISTS. THE CUSTOMER REPORTED THAT THIS OCCURRED WITH FOUR OTHER PTS, HOWEVER, NO FURTHER INFO OR DETAILS WERE GIVEN. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4) >< THE CUSTOMER REPORTED THAT NEGATIVE REACTIONS WERE OBSERVED ON THE PROVUE ANALYZER, COMPARED TO WEAK TO 1 + REACTIONS OBTAINED IN MANUAL GEL. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1481855 | THE CUSTOMER REPORTED THAT NEGATIVE REACTIONS WERE OBSERVED ON THE PROVUE ANALYZER, COMPARED TO WEAK TO 1 + REACTIONS OBTAINED IN MANUAL GEL. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< AN OCD FIELD ENGINEER WAS AT THE CUSTOMER SITE AND VERIFIED THAT ALL CAMERA ADJUSTMENTS WERE FINE. THE CUSTOMER PERFORMED TESTING IN MANUAL GEL, AND WEAK TO 1+ REACTIONS WERE OBTAINED FOR SAMPLES THAT THE PROVUE RESULTED AS NEGATIVE. ALTHOUGH THE TESTING IN MANUAL GEL YIELDED POSITIVE RESULTS, THE TESTING WAS INCONCLUSIVE DUE TO THE INCONSISTENT RESULTS BETWEEN TWO TECHNOLOGISTS. THE CUSTOMER REPORTED THAT SIMILAR ISSUE OCCURRED WITH FOUR OTHER PTS, HOWEVER, NO FURTHER INFO OR DETAILS WERE GIVEN. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4) | 0
1079429 | NO DEFINITIVE ROOT CAUSE WAS DETERMINED REGARDING HOW THE REAGENT CELLS BECAME HEMOLYZED. THE CUSTOMER DID NOT OBSERVE ANY ISSUES WITH THE PROBE OR VOLUME. SERVICE WAS NOT NEEDED SINCE THE CUSTOMER INDICATED THAT THE ANALYZER WAS FUNCTIONING AS EXPECTED AFTER THE INCIDENT. INSTRUMENT MALFUNCTION WAS NOT IDENTIFIED. USER ERROR COULD NOT BE RULED OUT. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THE INCIDENT. (B) (4). >< THE CUSTOMER REPORTED THAT THEY PERFORMED TESTING ON THE ORTHO PROVUE ANALYZER USING DISCOLORED (BROWN COLORED) REAGENTS AFTER OBSERVING THE ISSUE. TESTING WAS DISCONTINUED WHEN HEMOLYSIS WAS NOTED IN THE REAGENT VIALS. THE CUSTOMER INDICATED THAT A SET OF VIALS WAS OPENED ON JUNE 13TH AND THE HEMOLYSIS/BROWN COLORING WAS NOTED ON JUNE 15TH. THE CUSTOMER INDICATED THAT THEY DID NOT HAVE ANY ISSUES WITH QC, AND IS CONFIDENT THAT NO ERRONEOUS RESULTS WERE REPORTED. THE USE OF HEMOLYZED REAGENT RED CELLS MAY AFFECT THE INTERPRETATION IN DETECTING ANTIBODY-INDUCED HEMOLYSIS. | 0
1059891 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE READ A SAMPLE BARCODE ID INCORRECTLY. THE INCIDENT OCCURRED ON (B) (6) 2008. NO ERRONEOUS RESULTS WERE REPORTED. SAMPLE MISIDENTIFICATION MAY LEAD TO THE REPORTING OF ERRONEOUS TEST RESULTS. >< NO DEFINITIVE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND VERIFIED THE APPROPRIATE SETTINGS TO RETURN THE INSTRUMENT TO EXPECTED OPERATIONS. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4) | 0
1472084 | A POSSIBLE ROOT CAUSE, TORN/DAMAGED BARCODE LABEL, WAS DETERMINED REGARDING THE BARCODE MISREAD ISSUE. DURING OCD CUSTOMER TECHNICAL SERVICES (CTS) TROUBLESHOOTING ASSISTANCE WITH THE CUSTOMER, THE CUSTOMER INSPECTED THE SAMPLE BARCODE LABEL AND DETERMINED THAT PART OF THE LABEL WAS TORN/ DAMAGED. THIS RESULTED IN THE PROVUE'S INCORRECT READ OF THE BARCODE. THE PROVUE AND THE HAND HELD SCANNER WERE ABLE TO READ THE SAME LABEL CORRECTLY WHEN THE UNTORN SIDE OF THE LABEL WAS READ. THE CUSTOMER REPLACED THE TORN LABEL WITH A NEWLY PRINTED ONE AND RETESTED THE SAMPLE ON THE PROVUE. THE PROVUE IDENTIFIED THE CORRECT SAMPLE BARCODE ID NUMBER. PROVUE MALFUNCTION WAS NOT IDENTIFIED. USER ERROR COULD NOT BE RULED OUT AS A CONTRIBUTING FACTOR. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS INSTRUMENT SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE MISREAD A SAMPLE BARCODE ID NUMBER. THE CUSTOMER INDICATED THAT THEY REVIEWED THE RESULTS AFTER THE COMPLETION OF THE RUN AND NOTICED THE DIFFERENCE IN SAMPLE NUMBERS. THE PROVUE IS NOT LINKED TO THE LIS SYSTEM. THE SAME SAMPLE TUBE WAS PLACED ON THE SAMPLE CAROUSEL A SECOND TIME AND THE SAMPLE BARCODE WAS CORRECTLY IDENTIFIED BY THE PROVUE. NO ERRONEOUS RESULTS WERE REPORTED. SAMPLE MISIDENTIFICATION MAY LEAD TO THE REPORTING OF ERRONEOUS TEST RESULTS. | 0
1325573 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND THE CUSTOMER INDICATED THAT THE ROBOTIC HANDLER WAS DRIPPING. THE FFE DETERMINED THE PROBE WAS BENT. THE FE REPLACED THE PROBE AND PERFORMED ADJUSTMENTS. THE REPLACEMENT OF THE PROBE AND REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). >< THE CUSTOMER REPORTED THAT THE WASH DILUTION CUP OF THE ORTHO PROVUE ANALYZER OVERFLOWED CONTAMINATING THE REAGENTS. NO ERRONEOUS RESULTS WERE REPORTED. FLUID LEAK CAN LEAD TO DILUTION OF REAGENT/SAMPLE, CARRY OVER/CROSS CONTAMINATION FROM MICROTUBE TO MICROTUBE, ERRONEOUS TEST RESULTS AND TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1418567 | AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND PERFORMED CAMERA ADJUSTMENTS. THE CUSTOMER RAN CONTROLS OF THE PATIENT SAMPLE. THE CONTROL RESULTS WERE ACCEPTABLE, HOWEVER, THE PATIENT'S RESULT WAS STILL NEGATIVE. VISUAL INSPECTION OF THE PROCESSED CARD WAS NEGATIVE. THE CUSTOMER CONFIRMED THAT THE INSTRUMENT WAS FUNCTIONING AS EXPECTED. AN OCD LAB SPECIALIST FOLLOWED UP WITH THE CUSTOMER AND FOUND THE SAMPLE CONTAINS ANTI-M ANTIBODY REACTING INCONSISTENTLY IN MANUAL GEL AND ON THE PROVUE. THE CUSTOMER ALSO CONFIRMED THAT THE PATIENT HAD ANTI-M ANTIBODY USING THE TUBE METHOD. THE CUSTOMER WAS UNABLE TO DETERMINE WHETHER THE ANTI-M ANTIBODY WAS HOMOZYGOUS OR HETEROZYGOUS. INCIDENT APPEARS TO BE SAMPLE RELATED SINCE THE SAMPLE WAS INCONSISTENTLY REACTING IN ALL THE DIFFERENT METHODOLOGIES. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). >< THE CUSTOMER STATED THAT A SAMPLE CONTAINING ANTI-M AND POSSIBLY ANOTHER UNIDENTIFIED ANTIBODY WAS REACTING INCONSISTENTLY ON THE ORTHO PROVUE ANALYZER. THE SAMPLE ALSO REACTED INCONSISTENTLY IN GEL AND ALTERNATIVE TUBE METHOD. THE CUSTOMER WAS UNABLE TO DETERMINE WHETHER THE SAMPLE WAS HOMOZYGOUS OR HETEROZYGOUS FOR ANTI-M USING THE TUBE METHOD. NO ERRONEOUS RESULTS WERE REPORTED. FALSE NEGATIVE TEST RESULTS CAN LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1373102 | THE CUSTOMER REPORTED THAT THEY WERE CONFIDENT THE INCIDENT WAS ISOLATED AND RELATED TO THE WASTE TUBING NOT BEING PROPERLY CONNECTED. THIS MAY HAVE RESULTED IN THE LEAKAGE. THE PROPER CONNECTION TO THE WASTE BOTTLE HAS RETURNED THE ANALYZER TO EXPECTED OPERATION. NO REPAIRS WERE REQUIRED. THE CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS INSTRUMENT SINCE THIS INCIDENT. (B)(4). >< THE CUSTOMER REPORTED THAT THE PROBE LEAKED FLUID AND THE DILUTION CUP OVERFLOWED ON THE ORTHO PROVUE ANALYZER. ALL TESTING WAS ABORTED. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1363789 | OCD SECOND LEVEL SUPPORT REVIEWED THE PROVUE LOG FILES REGARDING THE ISSUE. INVESTIGATION REVEALED THAT THE CUSTOMER HAD SELECTED THE OPTION OF 'PRESENCE VERIFICATION ONLY' FOR THE REAGENT VIALS. THE PROVUE CONFIGURATION WAS ALTERED, DISABLING THE PSID IDENTIFICATION OF THE REAGENT CELLS AND THE REAGENT BARCODE ID WAS DISABLED. THE LOG FILES ALSO REVEALED THAT THE REAGENT VIALS WERE PLACED INCORRECTLY ON THE PROVUE, RESULTING IN THE QC FAILURE. AN OCD FIELD ENGINEER CORRECTED THE CONFIGURATION SETTINGS ON THE INSTRUMENT AND DISCUSSED THE ISSUE WITH THE CUSTOMER. QC WAS PROCESSED AND ACCEPTED. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B)(4). >< THE CUSTOMER REPORTED THAT DURING THE VALIDATION PROCESS OF THE ORTHO PROVUE ANALYZER TO THE LAB INFO SYSTEM, THE PROVUE WAS INFECTED WITH A VIRUS THAT CORRUPTED THE SOFTWARE SYSTEM. THE CUSTOMER FEELS THAT SINCE (B)(6) 2009 THE PROVUE SOFTWARE IS SWITCHING THE REACTIONS FOR THE REVERSE A AND B CELLS; THE SCREEN 1 AND 2 MICROTUBES AND VICE VERSA. THE ISSUE WAS AFFECTING ALL BLOOD TYPES. THIS ISSUE WAS OCCURRING WITH ALL SAMPLES INCLUDING THE QC SAMPLES. THE PROVUE WAS TAKEN OUT OF SERVICE. ALL TESTING WAS REPEATED MANUALLY. NO ERRONEOUS RESULTS WERE REPORTED. | 0
1036352 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED FLUID AND THE DILUTION CUP OVERFLOWED. THE CUSTOMER INDICATED THAT THE REAGENTS WERE CONTAMINATED. TESTING WAS ABORTED. NO ERRONEOUS RESULTS WERE REPORTED. PROBE DRIP MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER WENT TO THE CUSTOMER SITE AND DETERMINED THE PROBE WAS BENT. REPLACEMENT OF THE PROBE AND THE APPROPRIATE ADJUSTMENTS HAS RETURNED THE INSTRUMENT TO EXPECTED OPERATION. THE CUSTOMER HAS REPORTED ONE ADDITIONAL COMPLAINT SINCE THIS INCIDENT FOR A BENT PROBE COMPLAINT WITHOUT A FLUID DRIP. (B) (4) | 0
1369922 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE MISREAD SAMPLE BARCODE ID NUMBERS. NO ERRONEOUS RESULTS WERE REPORTED. SAMPLE MISIDENTIFICATION MAY LEAD TO THE REPORTING OF ERRONEOUS TEST RESULTS. >< AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND DETERMINED THAT THE CUSTOMER'S BARCODE LABEL PRINTER WAS NOT PRINTING GOOD QUALITY LABELS. THERE WERE INCONSISTENCIES IN THE LABELS, WHICH WAS CAUSED BY THE LABEL PRINTER SKIPPING. THE PROVUE WAS ABLE TO CORRECTLY READ LABELS PRINTED FROM A DIFFERENT LABEL PRINTER. THE FE PERFORMED SAMPLE CAMERA ADJUSTMENTS AND ADVISED CUSTOMER TO MONITOR BARCODE LABEL PRINTER AND REPLACE IT IF NEEDED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS INSTRUMENT SINCE THIS INSTRUMENT. INCIDENT IS ISOLATED. (B)(4). | 0
1369921 | THE CUSTOMER INADVERTENTLY FILLED THE WRONG WASH SOLUTION INTO THE WRONG CONTAINER. WASH SOLUTION A CONTAINS SODIUM AZIDE AND SALINE USED TO WASH AND RINSE THE INSIDE OF THE PROBE. WASH SOLUTION B IS USED TO WASH THE OUTSIDE OF THE PROBE. WASH SOLUTION B CONTAINS A DETERGENT, WHICH COULD ALSO HAVE A LYSING EFFECT, DESTROYING THE RED CELLS IN A REAGENT PRODUCT. THE CUSTOMER REPORTED THAT THE CONTAINER IN QUESTION WAS CLEANED AND REPLENISHED WITH THE APPROPRIATE WASH SOLUTION. THE CUSTOMER INDICATED THAT QC TESTING PASSED WITH ACCEPTABLE RESULTS. THE CUSTOMER REPEATED TESTING OF ALL THE SAMPLES THAT MAY HAVE BEEN AFFECTED AND REPORTED NO DISCREPANCIES WERE NOTED. INCIDENT OCCURRED AS A USER ERROR. THE CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS INSTRUMENT SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B)(4). >< THE CUSTOMER REPORTED THAT THEY INADVERTENTLY FILLED WASH SOLUTION A AND B CONTAINERS WITH WASH SOLUTION A FLUID AND PERFORMED PATIENT TESTING ON THE ORTHO PROVUE ANALYZER. SAMPLES WERE PROCESSED WITH ONLY WASH SOLUTION A/ THE CUSTOMER REPORTS THAT THE NIGHT SHIFT TECH DID PERFORM QC TESTING ON THE PROVUE SYSTEM AND NO DISCREPANCIES WERE NOTED, AND QC PASSED. THE CUSTOMER DID NOT HAVE SPECIFIC DETAILS IF ANY INCORRECT OR ERRONEOUS RESULTS WERE REPORTED. AN INCORRECT WASH SOLUTION USED DURING TESTING MAY CAUSE CARRY-OVER, CROSS CONTAMINATION OR HEMOLYSIS OF REAGENTS OR SAMPLES, WHICH COULD RESULT IN ERRONEOUS TEST RESULTS. | 0
1458583 | AN OCD FIELD ENGINEER (FSE) VISITED THE CUSTOMER SITE, REPLACED THE PROBE, WASH STATION AND PERFORMED THE NECESSARY ADJUSTMENTS AND VERIFICATIONS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. THE CUSTOMER PROCESSED SAMPLES WITHOUT ANY ISSUES. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). >< THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER DRIPPED AND THE PROBE WAS PIERCING THE GEL CARDS OFF CENTER. THE CUSTOMER DID NOT KNOW IF THE REAGENTS OR SAMPLES WERE CONTAMINATED OR WHETHER THE ANALYZER POSTED ANY ERROR MESSAGES FOR THE PROBE DRIP. THE CUSTOMER DISCONTINUED THE USE OF THE INSTRUMENT. NO ERRONEOUS RESULTS WERE REPORTED. PROBE ISSUES MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1458584 | THE CUSTOMER PERFORMED TROUBLESHOOTING REGARDING THE FALSE INTERPRETATIONS OF THE TEST RESULTS AND FOUND THAT OBSTRUCTION OF A GEL CARD NEAR THE PROVUE GEL CAMERA MAY HAVE RESULTED REACTION MISREADS. SERVICE WAS NOT REQUIRED SINCE THE CUSTOMER RESOLVED THE ISSUE. THE CUSTOMER INDICATED THAT THE INSTRUMENT WAS PERFORMING AS EXPECTED. THE CUSTOMER REPEATED THE SAMPLES IN QUESTION AND OBTAINED ACCEPTABLE RESULTS. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4). >< THE CUSTOMER REPORTED THAT THE ORTHO PROVUE ANALYZER MISREAD REACTIONS IN THE FORWARD ABO TYPING. THE PROVUE DROPPED THE CARD INTO THE WASTE BIN. HOWEVER, THE CARD GOT STUCK NEAR THE CAMERA. THIS MAY HAVE RESULTED IN THE MISREADS. THE INCIDENT OCCURRED ON (B) (6) 2009. HOWEVER, THE ISSUE WAS DISCOVERED ON (B) (6) 2009. THE CUSTOMER INDICATED THAT THERE WERE NO ERRONEOUS RESULTS REPORTED. A FALSE POSITIVE RESULT MAY LEAD TO THE MISCLASSIFICATION OF A PT'S ABO GROUP, WITH THE POTENTIAL FOR TRANSFUSION OF THE ABO INCOMPATIBLE BLOOD. | 0
1458585 | THE CUSTOMER REPORTED THAT THE PROBE ON THE ORTHO PROVUE ANALYZER DRIPPED FLUID. THE REAGENTS AND SAMPLES WERE CONTAMINATED. NO ERRONEOUS RESULTS WERE REPORTED. PROBE ISSUES MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. >< AN OCD FIELD ENGINEER VISITED THE CUSTOMER SITE AND FOUND THE PRESSURE REGULATOR WAS CLOGGED RESULTING IN THE PROBE DRIP. THE FE CLEANED AND ADJUSTED THE PRESSURE REGULATOR, REPLACED THE PROBE AND WASH BLOCK WHICH WAS CLOGGED AND CRACKED. THE CAVRO-DILUTER SYRINGE WAS REPLACED DUE TO DRIED FLUID LEAKING FROM BOTTOM OF THE SYRINGE CYLINDER. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. QC PASSED. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. INCIDENT IS ISOLATED. (B) (4). | 0
1501079 | A POSSIBLE ROOT CAUSE WAS DETERMINED. AN OCD FIELD ENGINEER ARRIVED AT THE SITE AND DETERMINED THE PRESSURE REGULATOR WAS CLOGGED. THE FE REMOVED THE REGULATOR AND CLEANED AND ADJUSTED IT RETURNING THE INSTRUMENT TO EXPECTED OPERATION. A SEARCH WAS PERFORMED CONCERNING COMPLAINTS LOGGED BY THIS CUSTOMER. THIS CUSTOMER HAS NOT LOGGED ANY COMPLAINTS AGAINST THIS ANALYZER SINCE THIS INCIDENT. (B) (4) >< THE CUSTOMER REPORTED THAT THE PROBE OF THE ORTHO PROVUE ANALYZER IS DRIPPING. THE CUSTOMER STATED THAT NO ERROR MESSAGES WERE POSTED. THE PROVUE WAS TAKEN OUT OF USE. NO ERRONEOUS RESULTS WERE REPORTED. PROBE ISSUES MAY LEAD TO DILUTION OF SAMPLE / REAGENT, CARRY OVER AND / OR CROSS CONTAMINATION AND ERRONEOUS RESULTS WHICH COULD LEAD TO TRANSFUSION OF INCOMPATIBLE BLOOD. | 0
1628078 | WE HAVE A CERNER EMR SYSTEM THAT IS CONNECTED TO A PHILIPS PATIENT MONITORING SYSTEM TO CHART VITAL SIGNS (VS) SEMI-AUTOMATICALLY. THE SYSTEM REQUIRES THE NURSE TO REQUEST THE VS BE DOWNLOADED, VERIFY THE VS AND HAVE THEM POST TO THE PATIENTS CHART. SEVERAL TIMES THE STAFF REPORTS THAT VITAL SIGNS HAVE BEEN DOWNLOADED, VIEWED, APPROVED AND HAVE SUBSEQUENTLY DISAPPEARED FROM THE SYSTEM. CERNER APPARENTLY WAS AWARE OF THE PROBLEM AND INSTALLED A 'PATCH' TO FIX IT. THE FREQUENCY OF THE PROBLEM HAD SIGNIFICANTLY DECREASED BUT THE PROBLEM CONTINUES TO PRESENT. THIS IS A COMPLEX PROBLEM AND TROUBLESHOOTING REQUIRES COLLABORATION BETWEEN IN-HOUSE BIOMED, THE HOSPITAL'S CORPORATE IT EMR TEAM, AND THE CERNER TROUBLESHOOTING TEAM. | 1
1601322 | THE HOSPITAL USES A NIC VUE SYSTEM CONSISTING OF EEG MACHINE WHICH MAY BE USED ANYWHERE IN THE HOSPITAL, AND AN EEG WORKSTATION ON WHICH EEG STUDIES ARE STORED AND REVIEWED. THE EEG MACHINE IS DESIGNED TO COMMUNICATE WITH THE WORKSTATION VIA THE HOSPITAL'S LAN. THE POLICY OF THE HOSPITAL'S IT DEPARTMENT IS THAT ANY SYSTEM CONNECTING TO THE HOSPITAL NETWORK MUST HAVE THE CORPORATE ANTI-VIRUS SOFTWARE (MCAFEE) LOADED ON IT. USING INFORMATION PROVIDED BY NICOLET, THE HOSPITAL IT GROUP WAS ABLE TO LOAD THIS SOFTWARE PROTECTION ON THE EEG WORKSTATION, AND CONFIGURE IT ACCORDING THE THE NICOLET RECOMMENDATIONS. HOWEVER THE HOSPITAL'S IT DEPARTMENT CONSIDERS THE EEG MACHINE ITSELF TO BE A MEDICAL DEVICE, AND THEIR POLICY PROHIBITS THEM FROM INSTALLING SUCH ANTIVIRUS SOFTWARE ON A MEDICAL DEVICE.THE BIOMEDICAL ENGINEERING DEPARTMENT HAS TESTED A COMMERCIALLY PURCHASED VERSION OF MCAFEE ANTIVIRUS SOFTWARE ON THE EEG MACHINE, AND FOUND THAT IT BLOCKED COMMUNICATION WITH THE NETWORKED WORKSTATION. THE HOSPITAL'S IT GROUP MAY HAVE THE EXPERTISE TO PROPERLY CONFIGURE ANTIVIRUS SOFTWARE ON THE EEG MACHINE, AS THEY DID THE WORKSTATION, BUT FEELS THAT IT IS NOT APPROPRIATE FOR THEM TO INSTALL OR CONFIGURE SOFTWARE ON A MEDICAL DEVICE.AT THIS TIME THE EEG MACHINE IS BEING USED, UNPROTECTED, ON THE HOSPITAL NETWORK, IN ORDER THAT IT CAN OPERATE PROPERLY AND COMMUNICATE WITH THE WORKSTATION. WE BELIEVE THAT OTHER FACILITIES MAY ENCOUNTER SIMILAR ISSUES AS THE NUMBER OF MEDICAL DEVICES CONNECTING WITH HOSPITAL LAN'S CONTINUES TO INCREASE. | 1
1414861 | ON (B)(6) 2009, BAXA WAS NOTIFIED BY THE CUSTOMER OF A PATIENT INVOLVED INCIDENT USING THE ABACUS V2.0 TPN CALCULATING SOFTWARE FOR TPN PRODUCTION. IT WAS REPORTED THAT THIS PATIENT RECEIVED A TPN BAG CONTAINING A LESS-THAN-DESIRED PROTEIN LEVEL. NO ADDITIONAL MEDICAL INTERVENTION WAS NEEDED AND NO APPARENT LONG-TERM INJURY OR ILLNESS RESULTED FROM THIS ISSUE. >< IN ORDER TO GAIN ADDITIONAL INFORMATION AND OBTAIN THE DATA FILES NEEDED REGARDING THIS INCIDENT, THE CUSTOMER WAS CONTACTED ON (B)(6) 2009. IN EACH INSTANCE, THE CUSTOMER STATED THAT THEY WERE TOO BUSY AT THE TIME AND ELECTED TO NOT SHARE ANY FURTHER DETAILS ABOUT THIS INCIDENT. NO DATA FILES FOR THE SUBJECT ORDER HAVE BEEN PROVIDED TO BAXA FOR EVALUATION. THROUGH OUR INVESTIGATION SO FAR WE DO NOT HAVE INFORMATION SUGGESTING THAT THE ABACUS CALCULATING SOFTWARE CAUSED OR CONTRIBUTED TO THIS ISSUE AS A RESULT OF FAILURE, MALFUNCTION, IMPROPER OR INADEQUATE DESIGN, MANUFACTURE, OR LABELING OF THE PRODUCT. IF BAXA RECEIVES ADDITIONAL INFORMATION IN REGARDS TO THIS INCIDENT, A FOLLOW-UP MDR WILL BE SUBMITTED. | 1
1545161 | (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN'S INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. >< CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT HARM. | 0
1942560 | FLOOR PEDAL WOULD NOT WORK DURING AN ACUTE MYOCARDIAL INFARCTION. FLOOR PEDAL CHARGED ALL NIGHT AND ALL MORNING, LED INDICATOR LIGHT WAS GREEN WHEN FLOOR PEDAL UNPLUGGED. FLOOR PEDAL WAS UNPLUGGED FOR LESS THAN ONE HOUR WHEN WE ATTEMPTED TO USE IT AND THERE WAS NO FUNCTION.SERVICE TECH CAME ONSITE AND RECOMMENDED REPLACING THE PARTS. | 0
1617027 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. >< THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 1
1679255 | WHILE MONITORING THE ECG WAVEFORM ON THE MP90 MONITOR, FOR NO APPARENT REASON THE QRS COMPLEX FLAT LINES AND THE MONITOR ALARMS ASYSTOLE. WHILE ALARMING ASYSTOLE THE WAVEFORM WILL PRODUCE WHAT APPEARS TO BE A SQUARE WAVE (AN OVERLOAD OF INPUT TO ECG MODULE) AND THEN BACK TO FLAT LINE AND SOME TIME LATER CORRECTS ITSELF. DURING ONE OF THE EARLY INCIDENTS A NURSE HAD A PATIENT ON A DRIP TO SLOW DOWN THE HEART RATE. WHEN SHE SAW ASYSTOLE ALARM SHE IMMEDIATELY SHUT DOWN THE DRIP. AS SHE CONTINUED TO REVIEW DATA ON THE MONITOR SHE REALIZED THE ALARM AND WAVEFORM WERE INCORRECT AND THAT THE ARTERIAL BLOOD PRESSURE (ABP) AND CENTRAL VENOUS PRESSURE (CVP) WERE CORRECTLY DISPLAYING GOOD PLETH WAVEFORMS. NO ADVERSE OUTCOME, BUT THE MONITOR CAUSED A RESPONSE FROM THE NURSE THAT STOPPED TREATMENT. THIS PROBLEM HAS OCCURRED IN ALL THREE ICU'S AND AT RANDOM BEDS AT RANDOM TIMES AND ONLY ONE MONITOR AT A TIME. DURING INVESTIGATION THERE WERE EXTERNAL DEVICES USED; A WIRELESS LAPTOP WAS IN USE IN THE BED NEXT TO THE AFFECTED BED. BY TURNING ON AND OFF THE LAPTOP WE COULD REPRODUCE THE PROBLEM. A BLACKBERRY PHONE IN THE AREA WOULD REPEATEDLY CAUSE THE MONITOR TO ALARM ASYSTOLE. IN A DIFFERENT ICU 8 FLOORS AWAY A VENTILATOR WITH HEATED CIRCUIT CAUSED ASYSTOLE ALARM.====================== HEALTH PROFESSIONAL'S IMPRESSION======================PHILIPS FIELD SERVICE ENGINEER QUERIED FELLOW FIELD SERVICE ENGINEERS (FSE) AND LEARNED OUR MMS ECG FIRMWARE VERSION D.02.02 NEEDED TO BE UPGRADED. FSE UPGRADED ALL MMS ECG FIRMWARE (IN MP90 MONITORS) TO D.02.05 AND X2 MODEL MONITORS TO RELEASE G.01.76. THIS WAS DONE TO PREVENT THE POSSIBLE INTERFERENCE CAUSED BY NEPTUNE HEATERS (HUMIDIFIERS) IN THE PATIENT VENTILATORS. HUMIDIFIER RESEARCH ENGINEERS CONTACTED THIS REPORTER REGARDING THE INCIDENT AT THIS FACILITY AND A SISTER FACILITY. THEY STATED THERE IS AN ISSUE WITH ANOTHER COMPANY'S MONITORS AND THE NEPTUNE HEATED CIRCUITS AND WERE CONCERNED ABOUT PHILIPS MONITORS. THEY INFORMED US THAT THEY ARE IN THE PROCESS OF A SOFTWARE UPGRADE TO ELIMINATE POTENTIAL INTERFERENCE WITH PHYSIOLOGICAL MONITORS BY RAISING THE FREQUENCY USED IN THE HEATED CIRCUITS.====================== MANUFACTURER RESPONSE FOR PHYSIOLOGICAL MONITOR, MODULE, INTELLIVUE MMS======================FIELD SERVICE ENGINEER UPGRADED ALL MMS ECG FIRMWARE TO RELEASE D.02.05 IN AN ATTEMPT TO REDUCE OR ELIMINATE THE ECGINTERFERENCE THEY WERE EXPERIENCING. THE ISSUE COULD BE REPLICATED BY USE OF A CELLULAR PHONE AND A LAPTOP COMPUTER USING A WIRELESS CARD. AFTER THE UPGRADE, NO INTERFERENCE WAS OBSERVED. | 1
1586063 | THIS IS A REMOTE DATA BACKUP SERVICE AND THERE IS/WAS NO OPERATOR. THE DEVICE IS SOFTWARE AND WAS EVALUATED REMOTELY, BUT WAS NOT RETURNED TO THE MFR. TESTS WERE RUN TO SUBSTANTIATE ISSUES RELATED TO THE GUARDIAN BACKUP SERVICE RUNNING DURING CLINIC HOURS BECAUSE DATA BACKUP WAS NOT COMPLETED IN THE ALLOTTED TIMEFRAME. THERE IS A POTENTIAL THAT SOME DATA WAS NOT COMPLETELY BACKED UP. THE DEVICE WAS EVALUATED REMOTELY ON. THIS IS NOT A SINGLE USE DEVICE. DEVICE WAS EVALUATED REMOTELY. >< THE INITIAL REPORTER CALLED AND STATED THAT THEY HAVE BEEN EXPERIENCING SLOW PERFORMANCE DURING CLINIC HOURS, AND THEY HAVE FOUND IN MANY CASES THAT THE BACKUPS WERE RUNNING DURING CLINIC HOURS. | 1
1497496 | PREVIOUSLY REPORTED IN MDR #2916710-2009-00016: THE INVESTIGATION REVEALED A DESIGN DEFICIENCY CAUSING PLAN EDIT SETTINGS TO MALFUNCTION UNDER SPECIFIC CIRCUMSTANCES, WHERE THE USER SELECTS TO USE A CONE BEAM CT SET UP FIELD. IN THIS CIRCUMSTANCE, THE USER IS ABLE TO ACQUIRE THE EXISTING GANTRY ANGLE, AND IS ABLE TO APPLY THAT GANTRY ANGLE TO ALL TREATMENT FIELDS IN THE PLAN. THERE IS A WARNING MESSAGE THAT THE DOSE DISTRIBUTION MAY BE ALTERED BY THE CHANGE, BUT, HAVING PASSED THAT MESSAGE, THERE IS NOTHING THAT PREVENTS THE USER FROM PROCEEDING TO TREAT WITH THE INCORRECT GANTRY ANGLE. THE POTENTIAL RISK IS HAVING A MULTIPLE FIELD IMRT PLAN WHERE THE GANTRY ANGLES FOR SEVERAL BEAMS ARE CHANGED TO A COMMON VALUE FOR A SESSION. THIS WOULD LEAD TO SEVERAL FIELDS OF ONE FRACTION BEING TREATED INCORRECTLY, POSSIBLY EXPOSING CRITICAL RADIOSENSITIVE TISSUES TO UNINTENDED RADIATION. THE ISSUE IS LIMITED TO ONE TREATMENT FRACTION BECAUSE THE PLANT WOULD BE UNAPPROVED FOR TREATMENT THE SECOND TIME. NO MISADMINISTRATION OCCURRED IN THIS CASE. HOWEVER, FURTHER ACTIONS ARE ADDRESSED IN VARIAN'S CAPA PROCESS. IN ADDITION, AN URGENT MEDICAL DEVICE CORRECTION NOTIFICATION WILL BE SENT TO AFFECTED CUSTOMERS. ALL CORRECTIVE ACTION WILL BE REPORTED UNDER THE GUIDELINES OF (B) (5). NO ADD'L F/U TO THIS MDR IS EXPECTED. >< DURING POST PLAN MODIFICATION, THE USER WAS ABLE TO ACQUIRE GANTRY ANGLES IN AN UNEXPECTED AND UNINTENDED MANNER. HOWEVER, THE USER SHOULD NOT HAVE BEEN ABLE TO ACQUIRE THE GANTRY ANGLE BECAUSE THE USER DID NOT HAVE RIGHTS TO THAT PRESCRIPTION MODIFICATION PARAMETER. | 1
1357078 | THE INVESTIGATION CONFIRMED THE CUSTOMER'S ALLEGATION. A SEQUENCE OF ACTIONS CAN LEAD TO THE POSSIBILITY OF MULTIPLE TREATMENTS OF A RETIRED PRESCRIPTION: USER PERFORMS ONLY THE IMAGE PORTION OF A TREATMENT, THEN RETIRES THE PLAN IN FAVOR OF AN EDITED PLAN. USER FAILS TO APPROVE THE NEW PLAN FOR TREATMENT. USER CALLS UP PT FOR TREATMENT AND RECEIVES MESSAGE 'ONE OR MORE PLANS ARE UNAPPROVED/PLANNING APPROVED FOR THIS PT. ONLY THE TREATMENT APPROVED PLANS WILL BE LOADED FOR TREATMENT' AND LOADS THE RETIRED PLAN (BY DESIGN - SINCE THIS PLAN HAS A PARTIALLY TREATED FRACTION, COMPLETING THAT FRACTION IS ALLOWED BY THE SYSTEM). USER TREATS THE RETIRED PLAN. DUE TO A MALFUNCTION WHERE THE PLAN IS NOT AVAILABLE FOR DOSE ACCUMULATION, THE SYSTEM FAILS TO SAVE THE TREATMENT RECORD AT THE END OF THE TREATMENT, AND PRESENTS TWO ERROR MESSAGES TO THE USER; 'WARNING: SAVE OF SESSION DATA AND/OR TREATMENT RECORD FAILED', FOLLOWED BY 'SESSION SAVE FAILED: SESSION DATA AND/OR TREATMENT RECORDS WILL NOT BE STORED ... AND WILL NEED TO BE IMPORTED FROM THE DICOM BACKUP LOCATION LATER' - THIS MESSAGE MUST BE SIGNED BY AN AUTHORIZED USER. USER FAILS TO IMPORT THE TREATMENT RECORD PRIOR TO THE NEXT TREATMENT SESSION, AND, SINCE THE RETIRED PLAN IS STILL PARTIALLY TREATED, REPEATS STEPS 3 TO 5. ALTHOUGH NO INJURY WAS REPORTED IN THIS CASE, THE CHANGE TO THE PLAN COULD BE A SHIFT IN POSITION OR FIELD SHAPE, POTENTIALLY LEADING TO A SERIOUS INJURY. VARIAN HAS ISSUED A CUSTOMER TECHNICAL BULLETIN TO DISCUSS ISSUES AROUND RETIRED PLANS AND WILL REFER THE MALFUNCTION HERE TO THE CAPA SYSTEM. NO FURTHER REPORT IS ANTICIPATED. >< THE PHYSICIST STATED THAT A PLAN WAS GENERATED ON (B) (6) 2008 AND WAS IMAGED. A REVISION WAS THEN MADE: HOWEVER, THE RETIRED PLAN WAS BROUGHT UP ON THE 4DITC AND TREATED ON BOTH (B) (6) 2008 AND (B) (6) 2008. HE NOTICED THAT THESE TREATMENTS WERE NOT SAVED BACK INTO HISTORY SO WENT IN TO DO A MANUAL TREATMENT AND THAT IS WHEN HE NOTICED THAT THE RETIRED PLAN HAD BEEN TREATED TWICE (AFTER IT HAD BEEN RETIRED). THERE WAS NO MISADMINISTRATION TO THE PT AS THE ONLY DIFFERENCE BETWEEN THE RETIRED AND REVISED PLAN WAS A SETUP FIELD. | 1
1357083 | NO PT INJURY OCCURRED IN THIS INSTANCE. THIS MATTER REMAINS UNDER INVESTIGATION, BUT AS A RESULT OF INVESTIGATIONS OF OTHER COMPLAINTS WITH A SIMILAR SYMPTOM, BUT DIFFERENT CAUSE, VARIAN HAS DETERMINED THAT, IF FOUND DUE TO MALFUNCTION, MULTIPLE TREATMENTS OF A RETIRED PLAN COULD REPRESENT A POTENTIAL SERIOUS INJURY. THE FOLLOW-UP REPORT WILL BE COMPLETED WHEN MORE DETAIL BECOMES AVAILABLE FROM THE INVESTIGATION. >< THE CUSTOMER WAS PERFORMING CHART CHECKS AND NOTICED A PLAN REVISION THAT OCCURRED AT THE 4D CONSOLE. AFTER LOOKING INTO THE ISSUE THE CUSTOMER NOTICED THAT THE ORIGINAL PLAN WAS TREATED THE FIRST DAY THEN RETIRED. THE SECOND DAY THE REVISED PLAN WAS TREATED. ON DAY 3-8 THE RETIRED PLAN WAS TREATED. THE CUSTOMER DOES NOT KNOW WHERE THE REVISION CAME FROM BUT DOES REMEMBER HAVING PORTAL IMAGER ISSUES THAT SAME DAY. CUSTOMER CREATED A NEW PLAN TO CORRECT CHART. | 1
1806748 | MONOPOLAR CUT STILL ENGAGED AFTER SURGEON TOOK FOOT OFF PEDAL. | 0
1789313 | A PHYSICIAN REPORTED HE FELT THE BLOOD PRESSURES TAKEN BY SOME OF THESE VITAL SIGN MONITORS WERE NOT ACCURATE. THE PHYSICIAN FELT THE MONITORS WERE DISPLAYING A FALSE HIGH BLOOD PRESSURE. OUR FACILITY CONTACTED THE DISTRIBUTOR TO REPORT THE CONCERN, AND WE HAD OUR BIOMEDIC CONTRACT COMPANY CHECK THE ACCURACY OF THESE VITAL SIGN MONITORS AND THE RESULT WAS: THE VITAL SIGN MONITORS WERE DISPLAYING SOME FALSE HIGH READINGS. THE DISTRIBUTOR IS IN THE PROCESS OF REPLACING ALL THESE VITAL SIGN MONITORS WITH NEW ONES. THESE MONITORS WERE PURCHASED NEW APPROXIMATELY 3-4 MONTHS AGO. NOTE: FOUR MORE VITAL SIGN MONITORS WERE INVOLVED WITH THIS PRODUCT APPARENT MALFUNCTION. DATES OF USE: SEVERAL MONTHS DURING 2010. DIAGNOSIS OR REASON FOR USE: MONITORING VITAL SIGNS IN THE ED AND ON MEDICAL SURGICAL UNIT. | 1
1456033 | OLYMPUS WAS INFORMED, THAT USER FACILITY PERSONNEL REPORTED SMELLING SOMETHING BURNING AROUND THE WORKSTATION. SERVICE STAFF FROM OLYMPUS VISITED THE FACILITY AND CHECKED THE WORKSTATION. THE SOURCE OF THE ODOR WAS ISOLATED TO THE TRANSFORMER, WHICH WAS SAID TO BE PARTIALLY BURNT. >< THE SUBJECT CART IS A CLASS 1, 510 (K) - EXEMPT DEVICE. THE SUBJECT ENDOSCOPY CART IS INTENDED TO PROVIDE MOBILE STORAGE FOR ENDOSCOPY EQUIPMENT SUCH AS VIDEO MONITORS, KEYBOARDS, LIGHT SOURCES, PRINTERS, ETC., AND ARE USED IN ENDOSCOPY SUITES/ UNITS. THE ORIGINAL EQUIPMENT MANUFACTURER (OEM) HAS DETERMINED THAT THE FUSE-HOLDER ON THE TRANSFORMER FOR THE SUBJECT MODEL ENDOSCOPY CART, WM-NP1, AS WELL AS SIX MORE OTHER MODELS, WM-P1, WM-260, WM-SC, WM-WP1, WM-DP1, WM-W260, WM-M260 CAN OVERHEAT, IF THE FUSE CARRIER IS NOT SUFFICIENTLY TIGHTENED. THE OEM'S RECOMMENDED TORQUE SETTING FOR THE FUSE HOLDER IS 0.5 NM (NEWTON METER). THE OEM REPORTED THAT THE CORRECT TORQUE SETTING ON THE TRANSFORMERS WERE IMPLEMENTED ON ALL TRANSFORMERS STARTING WITH A (B)(4) AND UP. ALL TRANSFORMERS MANUFACTURED AFTER THIS SERIAL NUMBER HAVE BEEN TORQUED CORRECTLY. A TOTAL OF 3,371 CARTS HAVE BEEN DISTRIBUTED IN THE US WITH THE AFFECTED TRANSFORMERS TO 1,810. WHILE OLYMPUS BELIEVES THE LIKELIHOOD OF OCCURRENCE FOR THIS PHENOMENON IS QUITE LOW, THE COMPANY WILL BE NOTIFYING CUSTOMERS TO INSPECT THE FUSE HOLDER AND TIGHTEN, TO REDUCED LIKELIHOOD OF OCCURRENCE. NO INJURIES HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. (B)(4). | 1
1890077 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. ON (B)(6) 2009, AT 15 MINUTES AFTER AUTOMATED BLOOD COLLECTION ON THE TRIMA ACCEL SYSTEM, THE DONOR SUFFERED A MYOCARDIAL INFARCTION. EMERGENCY MEASURE WERE TAKEN AND THE DONOR WAS ADMITTED TO THE HOSPITAL. THE INITIAL COMMUNICATION WAS RECEIVED BY CARIDIANBCT ON (B)(6) AND CARIDIANBCT NOTIFIED THE FRENCH REGULATORY AUTHORITIES THAT SAME DAY. THE FINS REPORTED DATE IS (B)(6) AS THIS WAS THE DATE WE HAD RECEIVED SUFFICIENT INFO TO REGISTER A COMPLAINT. IN THE OPINION OF EFS, TRIMA IS NOT RESPONSIBLE. >< (B)(4). THE DONOR WAS HOSPITALIZED FROM (B)(6) 2009 TO (B)(6) 2009 IN CARDIOLOGY FOR COMPLICATED MYOCARDIAL INFARCTION OF THROMBOSIS AND IMPORTANT LEFT VENTRICULAR DYSFUNCTION. THIS INFARCTION OCCURRED IMMEDIATELY AFTER A TRIMA DONATION. THE DONOR LEFT HOSPITAL WITH A TREATMENT AND A NEW ASSESSMENT TO OCCUR IN 6 MONTHS. TRIMA EXTRACORPOREAL VOLUME FOR THIS DISPOSABLE IS APPROXIMATELY 196ML, LESS THAN A SINGLE WHOLE BLOOD DONATION UNIT. THE SYSTEM DID NOT COLLECT RED BLOOD CELLS (RBC) DURING THIS PROCEDURE AND RINSEBACK WAS PERFORMED. THE RESIDUAL PACKED RBC VOLUME REMAINING IN THE DISPOSABLE FOLLOWING THIS PROCEDURE IS APPROXIMATELY 30ML. THIS WOULD RESULT IN AN IMPERCEPTIBLE CHANGE IN HEMATOCRIT FOR THE DONOR. ON (B)(6), SERVICE TECHNICIAN (B)(6) HAS BEEN TO THE EFS CENTER: (B)(6). HE HAS CHECKED THE TRIMA 1T04137, EVERYTHING IS FINE; (B)(4). SINCE THIS PROCEDURE, THEY HAVE DONE 49 PROCEDURES WITH NO PROBLEMS. ENCLOSED IS THE RUN DATA FOR THE PROCEDURE. RDF AVAILABLE - 1T04137_20090907_002_000703. DLOG. RUN DATA FILE (RDF) ANALYSIS OF THE PROCEDURE WAS PERFORMED ON (B)(6) 2009 BY A MEMBER OF CARIDIANBCT GLOBAL SYSTEMS PERFORMANCE TEAM. RDF ANALYSIS INDICATES THE RUN WAS UNREMARKABLE WITH NOTHING UNUSUAL NOTED AND THERE WAS NO FAILURE OF THE TRIMA SYSTEM. IT MET THE MANUFACTURING SPECIFICATIONS. | 0
1890078 | (B)(4). THE MD WAS INFORMED OF THE REACTION AND ATTENDED THE DONOR. THE PROCEDURE WAS PAUSED AND THE MD GAVE BENADRYL AND HYDROCORTISONE TO THE DONER. AFTER THE MD OBSERVED THE DONOR FOR FURTHER REACTION AND NONE WAS OBSERVED, THE PROCEDURE WAS RESUMED AND COMPLETED WITHOUT FURTHER PROBLEMS. THE DONOR WAS STABLE AND TOLERATED THE REST OF THE PROCEDURE. IV CALCIUM WAS BEING USED DURING PROCEDURE. THIS DISPOSABLE SET WAS NOT AVAILABLE FOR A SPECIFIC ROOT CAUSE ANALYSIS. A DHR REVIEW WAS CONDUCTED BASED ON THE LOT NUMBER OF THE SET PROVIDED BY THE CUSTOMER. THE REVIEW SHOWED NO MANUFACTURING OR STERILIZATION ANOMALIES OCCURRED DURING PRODUCTION AND NOTHING OF THIS TYPE OF INCIDENT WAS REPORTED FROM THIS LOT. THE DONOR WAS LATER FOUND TO BE ALLERGIC TO HYDROCODONE, BUT THE SOURCE OF THE REACTION WAS UNKNOWN. THE COBE SPECTRA APHERESIS SYSTEM ESSENTIAL GUIDE STATES 'THE COBE SPECTRA SYSTEM HAS MANY SAFETY FEATURES. A DONOR AND/OR PATIENT REACTION CAN OCCUR RAPIDLY, HOWEVER. IT IS IMPERATIVE THAT THE OPERATOR CONTINUOUSLY MONITOR THE COBE SPECTRA SYSTEM AND THE DONOR AND/OR PATIENT.' THE CUSTOMER WAS REMINDED OF THE ALTERNATIVE SINGLE PASS PRIME PROCEDURE IN THE EVENT, THE MD SUSPECTS A HYPERSENSITIVITY OF THE PATIENT/DONOR TO ETO RESIDUALS. >< THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. CUSTOMER REPORTED A DONOR REACTION WHILE GOING THROUGH STEM CELL COLLECTION PROCEDURE ON (B)(6) YEAR OLD MALE. THE REACTION OCCURRED AT 20 MINUTES INTO THE PROCEDURE AND THE PATIENT DEVELOPED ITCHING AND HIVES. | 0
1890076 | (B)(4). FIELD DIAGNOSTIC/CORRECTION: COMPLETE PROCEDURAL DETAILS ARE UNK. INVESTIGATION: THIS DISPOSABLE SET WAS UNAVAILABLE FOR RETURN FOR INVESTIGATION. TRENDS SUGGEST THAT THE EVENT REPORTED IS RANDOM AND ISOLATED ON A GENERAL BASIS. IF THE KIT IS LATER RECEIVED AND FINDINGS DIFFER FROM THE INFO PRESENTED IN THIS RECORD, AN ADDENDUM WILL BE ADDED TO CORRECT/UPDATE THE INVESTIGATION FINDINGS. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. CORRECTIVE/PREVENTIVE ACTION: THE TRIMA ACCEL OPERATOR'S MANUAL CAUTIONS USERS, 'THE TRIMA ACCEL AUTOMATED BLOOD COLLECTION SYSTEM HAS MANY SAFETY FEATURES. HOWEVER, A DONOR REACTION CAN OCCUR RAPIDLY. THEREFORE, IT IS IMPERATIVE THAT THE TRIMA ACCEL COLLECT SYSTEM AND THE DONOR BE MONITORED THROUGHOUT THE PROCEDURE' (P/N 777821-126). NO CORRECTIVE AND PREVENTIVE ACTIONS BY CARIDIANBCT QUALITY ASSURANCE WILL OCCUR AT THIS TIME. TRENDS ARE REGULARLY MONITORED TO DETERMINE APPROPRIATE CORRECTIVE ACTIONS BY MANUFACTURING OR ENGINEERING. >< THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. THE CUSTOMER CALLED INVESTIGATING AN ALLERGIC TYPE REACTION A DONOR HAD AFTER DONATING ON TRIMA ACCEL. THE DONOR APPARENTLY HAD REACTIONS AFTER THE LAST FOUR DONATIONS ON (B)(6) 2009, (B)(6) 2009, (B)(6) 2009, AND (B)(6) 2009. THE DONOR DID NOT INFORM THE OPERATORS UNTIL THE LAST INCIDENT. THE SYMPTOMS WERE AS FOLLOWS: RASH ON HANDS AND FEET (SWOLLEN), STIFF JOINTS, OCCASIONALLY FEVER AT NIGHT, RED BUMPS ON NECK, HEADACHE, AND SORE TONGUE. THERE WERE DIFFERENT TUBING SET LOT#S AND TRIMAS INVOLVED. NO UNUSUAL EVENTS WERE REPORTED DURING THE PROCEDURES. PER FOLLOW-UP, DONOR IS DOING FINE. | 0
1470095 | AN IMAGING DYNAMICS COMPANY (IDC) X3C 2200 DIGITAL RADIOGRAPHIC SYSTEM WAS IN USE AT (B)(6). AT APPROX 9:00 PM AN IMAGE OF THE PT TAKEN BY THE RADIOLOGY TECHNICIAN APPEARED ON THE SYSTEM PREVIEW SCREEN, BUT SUBSEQUENTLY DID NOT APPEAR IN THE STUDY ON THE IDC UNIT. THE SUBJECT REQUIRED AN ADDITIONAL EXPOSURE TO COMPENSATE FOR THE MISSING IMAGE. UPON FURTHER INVESTIGATION, IT WAS DETERMINED THE COMPUTER WAS NOT CONFIGURED AS REQUIRED BY ICD, AND A LACK OF ADEQUATE AVAILABLE COMPUTER MEMORY CAUSED A LONG IMAGE PROCESSING TIME TO OCCUR. DURING THE INITIAL IMAGE PROCESSING, THE RADIOLOGY TECHNICIAN ATTEMPTED TO TAKE ANOTHER EXPOSURE CAUSING THE INITIAL IMAGE TO BE LOST. THERE WAS NO INJURY TO ANY HEALTH PROFESSIONAL. THERE WAS NO ACCIDENTAL RADIATION EXPOSURE. | 1
1617187 | EXTERIOR EAR PIECES OF HEARING AIDS ARE TOO SHORT AND ARE EASILY DISLODGED BY EYEGLASS SIDE PIECES-THAT OCCURRED IN A HOSPITAL IN (B) (6), WHERE I ACCOMPANIED AN ACCIDENT VICTIM, (MY HUSBAND), TO THE EMERGENCY ROOM OF A HOSPITAL. HOSPITAL PERSONNEL AND I SEARCHED DILIGENTLY BUT COULD NOT FIND THE RIGHT EAR AID. ON A NUMBER OF SUBSEQUENT OCCASIONS, THE HEARING AIDS WERE CAUGHT JUST AS THEY/IT STARTED A REPEAT. SEVERAL LETTERS TO RESOUND, THE MAKER, REQUESTING CORRECTION OF THE HARDWARE PROBLEM HAVE GONE UNANSWERED-EVEN UNACKNOWLEDGED. SAME LETTERS TO MY PROFESSIONAL WHO RECOMMENDED THESE - HAVE ALSO BEEN IGNORED. | 0
1083974 | IT WAS REPORTED THAT A PT HAD A VFIB/VTACH-LIKE RHYTHM THAT NEVER ALARMED AS VFIB, VTACH OR ASYSTOLE. STAFF DISCONNECTED THE PT FROM TELEMETRY, CONNECTED A ROVING DASH TO VERIFY DATA AT THE BEDSIDE, AND DEFIBRILLATED THE PT. NO PT INJURY REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1587832 | THE CUSTOMER REPORTED THAT THE MONITOR WAS NOT PROVIDING AUDIBLE ALARMS. >< THE CUSTOMER CLAIMED THAT THE DEVICE IS NOT GIVING ANY AUDIBLE ALARMS. THE CUSTOMER WAS ADVISED BY TECHNICAL SUPPORT TO CONNECT AN OSCILLOSCOPE TO THE SPEAKER CONNECTIONS AND THEN INITIATE AN ALARM CONDITION. IT THERE IS A SIGNAL ON THE OSCILLOSCOPE, THEN THE PROBLEM IS THE SPEAKER. IF THERE IS NO SIGNAL, THEN THE PROBLEM IS WITH A PCB. THE CALL NOTES INDICATE THAT THE CUSTOMER WAS GOING TO TRY THIS AND CALL BACK IF THE PROBLEM PERSISTED. THE CUSTOMER DID NOT CALL BACK. THE INVESTIGATOR CALLED THE CUSTOMER TO INQUIRE ABOUT THE OUTCOME OF THE REPORTED PROBLEM. THE CUSTOMER REPORTED THAT THEY FOUND A LOOSE CONNECTION INSIDE THE DEVICE AND WHEN THE CONNECTION WAS RESTORED, THE AUDIBLE ALARMS BEGAN WORKING AGAIN. WHEN ASKED HOW THE LOOSE CONNECTION OCCURRED, THE CUSTOMER CLAIMED THAT HE DID NOT KNOW, BUT SUSPECTED THAT IT BECAME LOOSE AS THE RESULT OF DAILY USE. THE CUSTOMER DID NOT REQUIRE ANY ADDITIONAL ASSISTANCE FROM INVIVO REGARDING THIS REPORTED PROBLEM. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED. (B)(4). | 1
1832373 | THE PT WITH BACKGROUND LUNG DISEASE WAS ADMITTED TO THE HOSPITAL WITH WHEEZING. FROM TWO DIFFERENT SITES IN THE HOSPITAL, TWO DIFFERENT HIGH DOSE PARENTERAL STEROID THERAPIES WERE ORDERED BY TWO DIFFERENT DRS AND ENTERED INTO THE RECORD. THE PT RECEIVED BOTH POTENT STEROIDS FOR 24 HOURS CAUSING ELEVATED BLOOD SUGAR BUT NO OTHER, AS YET, UNTOWARD EVENTS. THE DOSE CHECKING AND RECONCILIATION APPS OF THE EHR, THE PHARMACY ELECTRONIC RECORDS AND CONTROL SYSTEM, THE HUMAN RESOURCES WORKING THE PHARMACY, THE NURSES ADMINISTERING THE MEDICATION, AND THE EXPERIENCED DRS ORDERING AND TREATING THE PT FAILED TO DETECT THE ERROR. NB. DUPLICATE MEDICATIONS HAVE BEEN ADMINISTERED TO PTS WITH ELECTRONIC RECORDS. THE INCIDENCE AND ADVERSE EVENTS CAUSED BY THIS TYPE OF SERIOUS ERROR HAVE NOT BEEN REPORTED IN PEER REVIEWED LITERATURE TO THE BEST OF OUR KNOWLEDGE. | 1
1634538 | A CUSTOMER REPORTED THAT THREE ICENTRAL (CENTRAL STATIONS) SHUT DOWN DURING THE NIGHT. AN ERROR MESSAGE WAS DISPLAYED. AFTER A SHORT PERIOD OF TIME, THE ICENTRAL(S) RESTARTED THEMSELVES. THIS IS THE THIRD OF THREE REPORTS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1633628 | THE TREATMENT UNIT WAS THE NUCLETRON HDR V3. THE EVENT WAS AN UNDER DOSE FOR 2 FRACTIONS (BID). THE DOSE PRESCRIPTION PER FRACTION WAS 4 GY. THE DOSE DELIVERED PER FRACTION WAS 0.4 GY. THE TOTAL UNDER DOSE WAS 7.2 GY. THE EVENT WAS CAUSED BY A SOFTWARE BUG THAT UPDATED THE ISODOSE LINES UPON CHANGING THE NORMALIZATION, BUT DID NOT UPDATE THE DWELL TIMES. THIS WAS THE FIRST TIME WE HAD USED THIS PARTICULAR NORMALIZATION METHOD WITH THIS PLANNING SYSTEM SO OTHER PATIENTS WOULD NOT HAVE BEEN AFFECTED BY THIS BUG. THE EVENT WAS DISCOVERED DURING OUR QA CHECK FOR THE PATIENT. PATIENT QA INVOLVES RELOADING THE PATIENT'S PLAN EACH DAY AND RECALCULATING THE DWELL TIMES. THE DWELL TIMES WERE UPDATED AFTER THE PLAN WAS RELOADED AND THE ERROR WAS DISCOVERED. THE ATTENDING WAS NOTIFIED AND THE PRESCRIPTION MODIFIED TO ACCOUNT FOR THE DEFICIT. NUCLETRON WAS CONTACTED TO REPORT THE SOFTWARE BUG. | 1
1430587 | SAME PERSON ID FOR TWO PTS. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. >< A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4) | 1
1424464 | THE COMPANY REP REPLACED THE HARD DRIVE. THE SYSTEM WAS TESTED TO FACTORY SPECS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO THE CUSTOMER. >< THE CUSTOMER REPORTED THAT WHILE THE CENTRAL STATION WAS IN USE MONITORING PTS, THE DISPLAYS AT THE NURSE'S STATION WENT BLANK. NO PT INJURY WAS REPORTED. | 1
1869239 | THE DEVICE MFR DATE IS UNK. THE DEVICE WAS SHIPPED ON 01/10/2002. FOLLOWING THE EVENT, THE DISTRIBUTOR'S TECHNICAL ENGINEER PERFORMED A CHECK OF THE LIGHT MONITOR. UPON ARRIVAL, IT WAS DISCOVERED THAT THE MONITOR WAS SHUT DOWN AND WAS STORED IN A BOX. THE TECHNICAL ENGINEER RESTARTED THE MONITOR AND PERFORMED A TEST WITH THE PROBE ON HIS FINGER. THE LIGHT MONITOR AND THE CENTRAL STATION REPORTEDLY ALARMED ACCORDING TO SPECIFICATIONS. DUE TO THE MONITOR BEING SHUT DOWN PRIOR TO THE TESTING, NO LOG OR TREND DATA IS AVAILABLE FROM THE LIGHT MONITOR FOR INVESTIGATION. THE ICENTRAL (CENTRAL STATION) LOGS WERE HOWEVER OBTAINED FOR EVAL. THE LOGS INDICATE THAT ECG AND NIBP WERE NOT MONITORED. 'NIBP CUFF LOOSE', AND 'ECG LEADS OFF' NOTIFICATIONS WERE OBSERVED IN THE LOGS. AN 'SPO2 LOW' ALARM WAS ALSO TRIGGERED. THE LIGHT MONITOR IS REPORTEDLY QUARANTINED BY THE LOCAL POLICE AT THIS TIME, AND THEREFORE IS NOT AVAILABLE FOR GE HEALTHCARE'S INVESTIGATION. WITHOUT THE MONITOR AND ADDITIONAL DETAILS OF THE PT EVENTS BETWEEN 12:00 AND 15:00, NO FURTHER INVESTIGATION CAN BE COMPLETED. GE HEALTHCARE WILL FILE A SUPPLEMENTAL REPORT SHOULD ADDITIONAL INFO BECOME AVAILABLE FOR INVESTIGATION. >< IT WAS REPORTED THAT THE LIGHT MONITOR DID NOT ALARM FOR A PT IN CARDIAC ARREST. THE PT'S SPO2 AND NIBP WERE REPORTEDLY BEING MONITORED AT THE TIME. THE NURSE STATES THAT THE PT WAS FOUND IN COMPLETE CYANOSIS, HOWEVER THE LIGHT MONITOR WAS NOT ALARMING. THE MESSAGE 'PULSE SEARCH' WAS NOTED ON THE MONITOR. THE PT WAS SUCCESSFULLY RESUSCITATED FOLLOWING THE EVENT, HOWEVER SUBSEQUENTLY DIED ON (B)(6)2010. | 1
1712696 | PHYSICIAN USING DEVICE -DIGITAL URETEROSCOPE- DURING PROCEDURE NOTED 'COBWEB-LIKE' MATERIAL IN PT URETER AFTER LASER FRAGMENTATION OF KIDNEY STONE WITH DEVICE. AFTER REMOVAL OF DEVICE PHYSICIAN USED BASKET TO RETRIEVE 'SMALL BALL OF MONOFILAMENT MATERIAL' FROM PT URETER. ADD'L SIMILAR MATERIAL WAS FOUND TO BE PROTRUDING FROM TIP OF DEVICE AFTER REMOVAL FROM PT, APPROX 8-9 CM IN LENGTH. PER PHYSICIAN, MATERIAL APPEARED TO HAVE ORIGINATED FROM THE WORKING CHANNEL OF THE URETEROSCOPE. MATERIAL HAS BEEN SAVED AND SEQUESTERED. MATERIAL WAS EXAMINED BY PATHOLOGIST AND DETERMINED TO BE SYNTHETIC, NOT BIOLOGICAL. PER PHYSICIAN, MATERIAL APPEARS TO BE NYLON MONOFILAMENTS. THIS WAS SCHEDULED, OUTPATIENT PROCEDURE. PT EVIDENCED NO ILL OR ADVERSE AFFECTS DURING RECOVERY PERIOD AND WAS DISCHARGED LATER SAME DAY TO HOME IN STABLE CONDITION. DATES OF USE: (B) (6) 2010--(B) (6) 2010. DIAGNOSIS OR REASON FOR USE: KIDNEY STONE REMOVAL. | 0
1549920 | THE INITIAL REPORTER CALLED AND REPORTED THAT OFFICEPACS WILL NOT LAUNCH ON THE WORKSTATIONS BECAUSE GUARDIAN IS RUNNING ON THE SERVER. THEY WERE RECEIVING AN ERROR STATING 'OFFICEPACS HAS ENCOUNTERED AN ERROR AND NEEDS TO BE RESTARTED.' UPON FURTHER INVESTIGATION, IT WAS IDENTIFIED THAT GUARDIAN WAS RUNNING DURING BUSINESS HOURS AND POTENTIALLY CAUSING THE ERROR. THE TIME SCOPE OF THE DATA BACKUPS WAS CHANGED TO FACILITATE THE OFFICE HOURS OF THE CLINIC. AS A RESULT OF THIS CHANGE, THERE IS POTENTIAL THAT GUARDIAN DATA BACKUPS WERE NOT OCCURRING COMPLETELY. >< THERE IS NO ESTABLISHED EXPIRATION DATE FOR THIS DEVICE. DEVICE MANUFACTURED DATE IS UNK AT THIS POINT. FURTHER ATTEMPTS WILL BE MADE FOR THIS INFORMATION. EVALUATION SUMMARY - WIN PERFORM TESTS WERE RUN TO SUBSTANTIATE PERFORMANCE ISSUES RELATED TO THE GUARDIAN BACKUP SERVICE RUNNING IN THE BACKGROUND. THE SITE WAS SET TO HAVE GUARDIAN FUNCTIONAL OUTSIDE OF BUSINESS HOURS. THERE IS A POTENTIAL THAT SOME DATA WAS NOT COMPLETELY BACKED UP. FURTHER INVESTIGATION IS ONGOING. THIS IS NOT A SINGLE-USE DEVICE. | 1
1781517 | INSTANCE 1: ON (B)(6) 2010, A RADIOLOGIST WAS USING POWERSCRIBE TO DICTATE A SERIES OF REPORTS FOR A LIST OF PTS 'OPENED' IN ISITE WHEN THE WORKSTATION 'CRASHED' FOR UNK REASONS. SUCH CRASHES ARE NOT INFREQUENT. FOLLOWING REBOOTING OF THE WORKSTATION, ALL THE PTS WERE 'LOCKED' IN ISITE (AS IS TYPICAL FOLLOWING A CRASH) AND WERE THEN UNLOCKED BY THE ADMINISTRATOR. SHORTLY THEREAFTER, THE RADIOLOGIST OPENED A PREVIOUSLY DICTATED PT AND DISCOVERED A MISMATCH: THE POWERSCRIBE REPORT TEXT AGREED WITH THE IMAGES AND THE IDENTIFIERS DISPLAYED IN ISITE BUT IT DID NOT AGREE WITH THE NAME DISPLAYED IN THE POWERSCRIBE PT NAME FIELD. THE POWERSCRIBE PT NAME FIELD INDICATED A DIFFERENT PT, ONE WHOSE DICTATION HAD BEEN PERFORMED EARLIER IN THE DAY. THE RADIOLOGIST LOGGED OUT OF POWERSCRIBE AND ISITE, LOGGED BACK IN TO BOTH, AND SAW THE THIS MISMATCH PERSISTED. THE WORKSTATION WAS THEN REBOOTED, THOUGH IT OTHERWISE APPEARED TO BE FUNCTIONING NORMALLY, AND THE MISMATCH WAS NO LONGER SEEN. INSTANCE 2: ON THE FOLLOWING DATE, A RADIOLOGIST DICTATED A REPORT WITH NO MORE THAN SIX PTS 'OPENED' IN ISITE. THIS WAS DONE ON A DIFFERENT WORKSTATION THAT IN INSTANCE 1, ABOVE. HE SAVED THE REPORT (USING THE 'SAVE' BUTTON), DID NOT SIGN THE REPORT, AND SUBSEQUENTLY DISCOVERED THE DICTATION UNDER A DIFFERENT PT'S NAME AS HE BEGAN TO REVIEW THIS DIFFERENT PT'S IMAGES. THE DIFFERENT PT SHOULD HAVE HAD NO DICTATION AT ALL. ADDITIONALLY, NO DICTATION WAS PRESENT UNDER THE PROPER PT'S NAME. THIS INSTANCE, UNLIKE THE FIRST, INVOLVED NO WORKSTATION CRASH OR REBOOTING. HAD THESE MISMATCHES OF PT'S NAMES NOT BEEN DETECTED BY THE RADIOLOGIST, WRONG-NAME REPORTS WOULD HAVE BEEN GENERATED. BECAUSE THESE INSTANCES REPRESENT SIGNIFICANT POTENTIAL PT SAFETY ISSUES, BOTH POWERSCRIBE (DICTAPHONE) AND ISITE (PHILIPS) WERE NOTIFIED AND ENGAGED PROMPTLY. DATES OF USE: (B)(6) 2010. DIAGNOSIS OR REASON FOR USE: DICTATION. PACS SYSTEM IS PHILIPS ISITE VERSION (B)(4) DICTATION SOFTWARE IS NUANCE HEALTHCARE'S POWERSCRIBE V. | 1
1707861 | THE VIGILANCE MONITOR, MODEL VGS, (B) (4), WAS REPORTED BY CUSTOMER AS THE UNIT SMOKED. 'SMOKE EXITED THE UNIT WHEN TESTING CARDIAC OUTPUT FUNCTION IN THE BIOMED SHOP. CCO NOT RESPONDING/NO BOLUS'. AN EDWARDS ELECTRONICS TECHNICIAN EVALUATED THE MONITOR AND THE COMPLAINT WAS CONFIRMED, CARDIAC OUTPUT NOT WORKING. THE U16 ON THE CICO BOARD WAS FOUND TO BE BURNT OUT. THE DEFECTIVE CICO BOARD WAS REMOVED AND REPLACED. THE SERVICE HISTORY RECORD WAS REVIEWED AND ALL MANUFACTURING REQUIREMENTS WERE MET. >< REPORTEDLY, THE UNIT SMOKED. 'SMOKE EXITED THE UNIT WHEN TESTING CARDIAC OUTPUT FUNCTION IN THE BIOMED SHOP. CCO NOT RESPONDING/NO BOLUS'. NO PATIENT INJURY REPORTED. | 0
1420026 | ON (B) (6) 2009, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED A MARTEL PRINTER FOR THE I-STAT1 ANALYZER WAS HOT TO TOUCH. THERE WAS NO REPORT OF ANY INJURY ASSOCIATED WITH THE COMPLAINT. | 1
1765576 | THE ABBOTT VP 2000 PROCESSOR IS A DEVICE DESIGNED TO AUTOMATE AND STANDARDIZE SLIDE SPECIMEN PROCESSING ROUTINE SLIDE STAINING FOR THE LABORATORY. THE FOLLOWING DEFICIENCIES WERE IDENTIFIED IN THE COMPLAINT INVESTIGATION (TWO CUSTOMER COMPLAINTS FROM ONE CUSTOMER SITE) THAT COULD LEAD TO LEAKAGE FAILURES OR THE INABILITY TO RUN THE INSTRUMENT: THE BASINS EXHIBIT A CRACK ALONG THE BOTTOM SEAM AND NON-UNIFORM COATING IN THE INTERIOR CORNERS. THE BASINS EXHIBIT SEVERE WARPING OF THE ALUMINUM HOUSING. THE BASINS EXHIBIT INCORRECT DIMENSIONS WHEN PERFORMING SQUARENESS TEST. THERE IS NO HEALTH HAZARD ASSOCIATED WITH A DELAY OF RESULTS, HOWEVER, LEAKING COULD RESULT IN A SLIP HAZARD. THE OVERALL ESTIMATE FOR THIS FAILURE IS 1.1% (BASED ON BASINS KNOWN TO BE RELATED TO THIS ISSUE). LEAKING RATE IS EXPECTED TO BE ABOUT 1 TO 2 DROPS PER MINUTE (< 100 UL PER MIN). WHEN A LOADED BASIN IS EITHER USED IN THE VP2000 OR STORED ON THE LAB BENCH, IN THE TYPICAL USAGE TIME (< 8 HRS), THE LEAKED LIQUID WILL BE LESS THAN 50 ML. THE OVERALL RISK IS LOW. >< ON (B)(4), 2012, THROUGH AN INTERNAL AUDIT INVESTIGATION IT WAS DISCOVERED THAT SOME DATES INCLUDED IN THE ORIGINAL MDR REPORT WERE INCORRECT. (B)(4). >< ABBOTT MOLECULAR TO DISTRIBUTE URGENT PRODUCT RECALL (3005248192-06/18/2010-001-C) VIA (B)(6) SHIPMENT TO UNITED STATES CUSTOMERS BASED ON CUSTOMER RECEIPT OF VP 2000 PROCESSOR AND BASIN SPARE PART SINCE (B)(6) 2008. TECHNICAL SERVICE BULLETIN WILL BE MADE AVAILABLE TO INSTRUCT FIELD SERVICE ENGINEERS TO INSPECT AND, IF NECESSARY, REPLACE HEATED BASINS. ADDITIONALLY, AS SPARE PART REPLACEMENT INVENTORY BECOMES AVAILABLE, (B)(4) TECHNICAL SERVICES MAY BEGIN EXECUTING TELEPHONE PROTOCOL AND PROVIDE REPLACEMENT BASINS SUFFICIENT FOR CUSTOMER TO DESTROY ALL CURRENT HEATED BASINS AND REPLACE WITH THOSE PROVIDED. TECHNICAL SERVICES WILL PROVIDE CUSTOMER REPLY FORM TO BE COMPLETED BY CUSTOMER AT TIME OF RECEIPT OF REPLACEMENT BASINS. | 0
1758738 | TRAUMA PATIENT BEING TRANSPORTED TO CT SCANNER ROOM IN THE ED ON TRANSPORT MONITOR THAT FAILED DUE TO SUSPECTED BATTERY ISSUE. MONITOR WAS FLASHING A YELLOW LIGHT. PATIENT WAS NOT ABLE TO BE MONITORED AND EXPIRED AT SOME POINT DURING TRANSPORT, OR DURING CT PROCEDURE.HEALTH PROFESSIONAL'S IMPRESSION: FAULTY MONITOR. | 1
1674302 | SPACELABS' FIELD SERVICE ENGINEER TESTED THE DEVICE WITH A SIMULATOR IN THE HOSPITAL. SPACELABS COULD NOT REPRODUCE THE CUSTOMER'S COMPLAINT AND FOUND THE DEVICE WORKING TO SPECIFICATIONS. AS A PRECAUTIONARY MEASURE, BECAUSE OF THE AGE OF THE DEVICE (MORE THAN 10-YEARS OLD), SPACELABS' FSE REPLACED THE MONITOR'S MAIN PRINTED CIRCUIT BOARD ASSEMBLY. THE HOSPITAL HAS RETURNED THE EQUIPMENT TO FULL SERVICE AND IS CURRENTLY USING THE DEVICE TO MONITOR PATIENTS. THERE HAVE BEEN NO REPORTS OF SIMILAR ISSUES. WE HAVE CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AND CONSIDER THIS ISSUE CLOSED. >< CUSTOMER REPORTED THAT A SPACELABS PATIENT MONITOR RESET DURING A PATIENT CODE. | 1
1714343 | ON DAY OF INCIDENT THE TOSHIBA CUSTOMER ENGINEER WAS CALLED TO THE SITE AND HAD THE CUSTOMER LOCK UP ROOMS AND PUT UP CAUTION AROUND MRI SECTION AND ORDER THE SYSTEM TO BE RAMPED DOWN. THE CLEANING PERSON WAS TAKEN TO THE HOSPITAL EMERGENCY ROOM FOR TREATMENT. ON 5/7/10 THE MRI SYSTEM WAS RAMPED DOWN AND THE FLOOR BUFFER WAS REMOVED. VERY LITTLE DAMAGE WAS NOTED AND ALL DAMAGE TO THE SYSTEM WAS FIXED. AN EVAL OF THE SYSTEM WAS CONDUCTED AFTER IT WAS RAMPED UP AND THE SYSTEM IS 100% OPERATIONAL. NO IMAGE QUALITY ISSUES. UNABLE TO REATTACH THUMB OF HOSPITAL CLEANING PERSON, BUT WERE ABLE TO DO A SKIN GRAFT. AT THE TIME OF THE EVENT WARNING SIGNS WERE POSTED OUTSIDE THE MR ROOM. >< A FLOOR BUFFER WAS TAKEN INTO THE MAGNET ROOM BY A HOSPITAL CLEANING PERSON. THE BUFFER WAS PULLED INTO THE MAGNET FIELD AND STUCK TO THE REAR OF THE MAGNET. THE CLEANING PERSON'S LEFT THUMB WAS CAUGHT BETWEEN THE BUFFER HANDLE AND THE REAR COVER OF THE MAGNET AND WAS PARTIALLY AMPUTATED. | 0
1713463 | CALLER REPORTED DISCREPANT RESULTS FOR COC AND APAP ON TRIAGE TOX DRUG SCREEN WITH MTD. CUSTOMER WAS NOTIFIED BY TECHNICIANS THAT THEY HAVE SEEN DISCREPANT RESULTS ON COC AND APAP FOR TOX DRUG SCREEN IN THE EMERGENCY ROOM WHEN THE SAMPLE IS REPEATED ON A NEW DEVICE. ONE SAMPLE WENT FROM POSITIVE TO NEGATIVE COC. ONE SAMPLE WENT FROM NEGATIVE TO POSITIVE COC. ONE SAMPLE HAD DISCREPANT RESULTS FOR APAP, BUT THE CALLER IS NOT SURE IF IT WENT FROM POSITIVE TO NEGATIVE OR VICE VERSA. CUSTOMER HAS NO INFO ON PT'S (MEDICAL HISTORY, MEDICATION, ETC.). >< INVESTIGATION PENDING. | 0
1699089 | WHEN A CERNER SERVICE PACK WAS LOADED INTO OUR SYSTEM IN MARCH , THE CONTENTS OF THE PACKAGE ADVERSELY AFFECTED THE BLOOD BANK'S ABILITY TO USE THE 'PATIENT PRODUCT INQUIRY' FUNCTION (PPI). PPI IS USED BY THE BLOOD BANK TO PERFORM A NUMBER OF PROCESSES, MOST IMPORTANTLY A LISTING OF PATIENT'S TRANSFUSION HISTORY. WITH THIS DEFECT ,WHEN PPI IS SELECTED, THE COMPUTER FREEZES UP AND EVENTUALLY GIVES US AN ERROR MESSAGE. NOT HAVING THE ABILITY TO USE THIS FUNCTION LIMITS THE BLOOD BANK'S ABILITY TO VIEW RECENT TRANSFUSION HISTORY, AS WELL AS BLOOD TYPE, ANTIBODIES, AND OTHER CRUCIAL BLOOD BANK INFORMATION. SOME, BUT NOT ALL, OF THIS INFORMATION CAN BE LOCATED USING OTHER FUNCTIONS, WITH THE MAJOR EXCEPTION OF RECENT TRANSFUSION HISTORY. WHILE THE SCREEN IS FROZEN, WE ARE UNABLE TO USE OTHER CERNER BLOOD BANK FUCNTIONS. IN ADDITION, AT TIMES, WE HAVE HAD MULTIPLE COMPUTER TERMINALS TIED UP AND UNAVAILABLE FOR USE UNTIL THESE SCRIPT ERRORS TIME OUT. | 1
1781891 | IT WAS DISCOVERED THAT MULTIPLE ECGS WERE RECONCILED TO ONE ORDER IN OUR ELECTRONIC MEDICAL RECORD (EMR). THE ECG REQUEST IS PLACED IN THE EPIC EMR, SENT TO THE PHILIPS TRACE MASTERVUE. THE ECG IS RECORDED ON AN ELECTROCARDIOGRAPHY CART AND THEN RECONCILED IN THE TRACE MASTERVUE TO THE EPIC ORDER NUMBER. WHEN MORE THAN ONE ECG IS RECONCILED TO A PARTICULAR ORDER, THE LATEST ECG RECONCILED TO THE ORDER WILL LINK TO THAT ORDER IN THE EMR. SPECIFICALLY, IF A PATIENT HAS AN ECG AT 8 AM (AND IT IS RECONCILED TO ORDER #1) AND THEN HAS A SECOND ECG THAT DAY AT 12 P.M., IT SHOULD BE RECONCILED TO ORDER #2 (GIVEN THE CLINICIAN ENTERED A SECOND ORDER FOR THE PATIENT.) IF THE SECOND ECG IS RECONCILED TO ORDER #1, IT WILL LINK TO THAT ORDER AND REPLACE THE FIRST ECG. A SECOND ISSUE RECOGNIZED IS THAT ON OCCASION ECGS FROM DIFFERENT PATIENTS WERE RECONCILED TO THE SAME ORDER CAUSING PATIENT MISMATCHES. THERE WERE 10 CASES IDENTIFIED, ALL WERE INVESTIGATED AND RESULTED IN NO PATIENT HARM.HEALTH PROFESSIONAL'S IMPRESSION: MANAGEMENT SYSTEM AND ELECTRONIC MEDICAL RECORD (EMR).MANUFACTURER RESPONSE FOR ECG MANAGEMENT SYSTEM, TRACE MASTERVUE:MANUFACTURER IS AWARE OF PATIENT MISMATCH WHEN ORDER IS RE-USED BUT DO NOT HAVE A SOLUTION. | 1
1548220 | DEVICE EVALUATION: CUSTOMER REPORT WAS CONFIRMED. CONTAMINATION SHIELD OF VAMP FLEX WAS DETACHED. CONTAMINATION SHIELD WAS EASY TO REATTACH AND TAKE OFF. NO VISIBLE DEFECT WAS FOUND ON ITS BODY OR CONTAMINATION SHIELD. >< THE FOLLOWING WAS REPORTED: 'CONTAMINATION SHIELD OF VAMP FLEX WAS DETACHED DURING THE FIRST BLOOD SAMPLING. CONTAMINATION SHIELD WAS NEW TYPE.' | 0
1832387 | WE FOUND THAT THIS DIABETIC PT WITH SUSPECTED PNEUMONIA WAS RECEIVING HIGH DOSE STEROIDS AND 2 LITERS OF D5 NORMAL SALINE INTRAVENOUS FLUID EVERY DAY OF A 10 DAY HOSPITALIZATION. EVERY DAY, THERE WAS AN ELABORATE MORE THAN 4 PAGE -REPETITIVE/CUT AND PASTE- PROGRESS NOTE ELECTRONICALLY GENERATED BY THE ATTENDING PHYSICIAN AND OTHERS FROM CONSULTANT-S. SCRUTINY OF THE EHR AND ORDERS REVEALED ONE WEIGHT RECORDED -ON ADMISSION- AND NONE AFTER 10 DAYS, NO RECORD OF FLUID BALANCE, MULTIPLE INTERVENTIONS FOR LABILE GLUCOSE LEVELS -ON STEROIDS AND SUGAR WATER, INCREASING EDEMA, AND LITTLE SIGN OF CLINICAL IMPROVEMENT OF PULMONARY STATUS. YET, DESPITE ELEVATED SERUM GLUCOSE, GLUCOSE CONTAINING INTRAVENOUS FLUIDS WERE CONTINUOUSLY GIVEN AT 75 CC PER HOUR AND DESPITE ADEQUATE ORAL INTAKE, CAUSING LABILE GLUCOSE LEVELS AND HEART FAILURE WITH FLUID BUILD UP. OBSCURED AMIDST THE MONOTONOUS PAGES OF LINES OF MOSTLY CLINICALLY IRRELEVANT UNUSABLE INFO WERE ESSENTIAL CLINICAL INFORMATION. RESULTANT IMPACT WAS FLUID IN THE CHEST CAVITY AND EDEMA OF THE EXTREMITIES AND FAILURE TO PROGRESS, REQUIRING LONGER HOSPITALIZATION. AS HAS BEEN SEEN AND REPORTED ELSEWHERE, THE USERS OF THIS DEVICE MADE ERRORS IN INTERNAL MEDICINE 101, IMPAIRED BY THE POOR USABILITY OF THE DEVICE ITSELF WHICH OBFUSCATED KEY RECORDS AND INFO. | 1
1751075 | REVIEW OF RESULTS IMAGES: INITIAL TESTING: (A11 TO H11) THE SAMPLE TYPED AS A POSITIVE AS EXPECTED. REPEAT TESTING: (A10 TO H10) THE SAMPLE APPEARS TO HAVE A FIBRIN CLOT IN WELL B10. THE SAMPLE TYPED AS AB POSITIVE. REPEAT TESTING: (A11 TO H11) THE SAMPLE TYPED AS A POSITIVE AS EXPECTED. A REVIEW OF IMAGES INDICATED THERE WAS A FIBRIN CLOT IN THE ANTI-B WELL THAT GAVE THE POSITIVE RESULTS. PER THE GALILEO OPERATOR MANUAL, 'VERY LARGE RED CELL CLOTS MAY NOT BE DETECTED BY THE GALILEO. CLOTS THAT INCLUDE GREATER THAN 50% OF THE SAMPLE RED BLOOD CELLS MAY NOT BE DETECTED BY THE INSTRUMENT. CLOTTED SAMPLES SHOULD NOT BE TESTED ON THE GALILEO.' >< CUSTOMER REPORTED AN ABO DISCREPANCY WHEN TESTING A PATIENT SAMPLE ON THE GALILEO. THE SAMPLE TYPED AS A POSITIVE ON THE INSTRUMENT. THE SAMPLE WAS RE-TESTED AND RESULTED AS AB POSITIVE. THE SAMPLE WAS TESTED A THIRD TIME ON THE GALILEO AND RESULTED AS A POSITIVE. | 0
1751078 | A SERVICE CALL WAS MADE. THE ISSUE WAS CORRECTED BY REPLACING THE WASHER MODULE. THE CUSTOMER RE-TESTED THE SAMPLES; THE INSTRUMENT IS OPERATING AS EXPECTED. >< CUSTOMER REPORTED UNEXPECTED NEGATIVE RESULTS WITH CAPTURE-R READY SCREEN 3 (CRRS 3) WHEN TESTING A SAMPLE ON THE GALILEO. | 0
1767412 | REVIEW OF INSTRUMENT RESULTS: ANTI A- 96 (AGGLUTINATION). ANTI B-51 (AGGLUTINATION). ANTI D-93 (AGGLUTINATION). MONO CTRL-10 (NO AGGLUTINATION). MIX WELL - -1. INT- AB POS. NOTE ERROR MESSAGE ON PLATE. 'IMAGE ANALYSIS EMPTY OR TOO DENSE WELLS DETECTED.'; ALSO NOTED IN THE LOG LIST REVIEW OF INFORMATION. THERE IS A NOTED THE PIPETTOR IS HALTED (LEFT ARM CLOT DETED AFFECTED NEEDLE 2, SAMPLE ADDITION (PIPETTOR) (B)(4) - CLOT DETECTED' UNABLE TO VERIFY IF THE CLOT WAS REMOVED PRIOR TO CONTINUED TESTING ON THE INSTRUMENT. THE REPORTED ABO DISCREPANCY SAMPLE WAS TESTED USING PROBE #2. THE FOLLOWING GALILEO LIMITATIONS ARE LISTED IN THE OPERATOR MANUAL: -FORWARD ONLY ABO-RH TESTING HAS A HIGHER RISK OF MISTYPE DUE TO THE ABSENCE OF THE REVERSE TYPE RESULTS. HAZARDOUS MISTYPES MAY OCCUR, SUCH AS AN A SAMPLE BEING INTERPRETED AS GROUP AB, OR AN RH (D) NEGATIVE SAMPLE BEING INTERPRETED AS RH (D) POSITIVE. FOR THIS REASON, ABO-RH RESULTS SHOULD ALWAYS BE COMPARED TO THE PATIENT OR DONORS' HISTORY. VERY LARGE RED CELL CLOTS MAY NOT BE DETECTED BY THE GALILEO. CLOTS THAT INCLUDE GREATER THAN 50% OF THE SAMPLE RED BLOOD CELLS MAY NOT BE DETECTED BY THE INSTRUMENT. CLOTTED SAMPLES SHOULD NOT BE TESTED ON THE GALILEO. INSTRUMENT IS OPERATING AS EXPECTED. >< CUSTOMER REPORTED AN ABO DISCREPANCY WHEN TESTING A DONOR SAMPLE ON THE GALILEO. THE DONOR BLOOD UNIT IS LABELED AS A POS. INITIAL TESTING ON THE GALILEO RESULT AS AB POS. REPEAT TESTING ON THE GALILEO RESULTED IN A POS. MANUAL TUBE BENCH METHOD RESULTED AS A POS. | 0
1594564 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE RESULT WHEN TESTING A SAMPLE ON THE ECHO. THE PATIENT HAS A HISTORY OF ANTI-K. THE SAMPLE WAS TESTED ON 2 ECHO INSTRUMENTS AND BOTH RESULTED NEGATIVE, BUT THE IMAGES APPEAR POSITIVE. >< REVIEW OF CAMERA IMAGES SHOW THAT CELL 3 RESULTED NEGATIVE ON TESTING WITH BOTH ECHO INSTRUMENTS, BUT APPEARS POSITIVE. VISUALLY THE IMAGE APPEAR AS A 1+ REACTION, BUT RESULTED NEGATIVE. ISSUE COMMUNICATED TO CUSTOMERS IN TECHNICAL COMMUNICATION CC-09-042-02 ON (B)(6) 2009. ADVISED THE CUSTOMER TO PERFORM A VISUAL VERIFICATION OF NEGATIVE REACTIONS BEFORE FINAL RELEASE OF THE WELL RESULTS. | 0
1747878 | REVIEW OF INSTRUMENT RESULTS: INITIAL TESTING: A10:99 B10:75 C10:99 D10:6 B10 WELL VISIBLE AGGLUTINATION, HOWEVER, NOT CLEAN WITH MIX FIELD BACKGROUND UPON FURTHER INSPECTION OF THE IMAGE, IT APPEARED THAT B10 WELL SUGGESTED POSSIBLE FIBRIN. PER THE GALILEO OPERATOR MANUAL, 'VERY LARGE RED CELL CLOTS MAY NOT BE DETECTED BY THE GALILEO. CLOTS THAT INCLUDE GREATER THAN 50% OF THE SAMPLE RED BLOOD CELLS MAY NOT BE DETECTED BY THE INSTRUMENT. CLOTTED SAMPLES SHOULD NOT BE TESTED ON THE GALILEO.' >< CUSTOMER REPORTED AN ABO DISCREPANCY WHEN TESTING A DONOR SAMPLE ON THE GALILEO. A KNOWN A POSITIVE SAMPLE RESULTED AS AB POSITIVE ON THE GALILEO. REPEAT TESTING ON A NEW SEGMENT FROM THE SAME DONOR UNIT RESULTED A POSITIVE AS EXPECTED. | 0
1627910 | CUSTOMER DID NOT RETURN SAMPLE FOR SEROLOGICAL INVESTIGATION. >< CUSTOMER REPORTED A PHENOTYPING DISCREPANCY ON THE ECHO WHEN TESTING A PATIENT SAMPLE WITH A KNOWN ANTI-C. TWO TUBES WITH THE SAME PATIENT SAMPLE RESULTED DIFFERENTLY (ONE POSITIVE AND ONE NEGATIVE). THE SAMPLE WAS RE-TESTED ON THE ECHO AND BOTH RESULTED POSITIVE. | 0
1759939 | REVIEW OF INSTRUMENT IMAGES SHOWS NO SAMPLE WAS PIPETTED INTO THE FIRST 8 CRRID WELLS UPON INITIAL TESTING, POSSIBLY DUE TO A BUBBLE OR CLOT IN THE SAMPLE. >< CUSTOMER REPORTED AN UNEXPECTED NEGATIVE RESULT WHEN TESTING A PATIENT SAMPLE WITH CAPTURE-R READY ID (CRRID) ON THE ECHO. THE PATIENT HAS A HISTORY OF ANTI-C AND ANTI-E. THE CUSTOMER REPEATED TESTING AND THE RESULTS WERE POSITIVE (3 TO 4+ REACTIONS). | 0
1747903 | A SYNOPSIS OF THE FINAL IMAGE FOR SCREENING ASSAY IS AS FOLLOWS:CELL#/WELL# RX STRENGTH TYPING VISIBLE INTERP. 1 / E1 10 E=E+ VISUALLY NEGATIVE WELL; 2 / F1 12 E+E= VISUALLY NEGATIVE WELL;A SYNOPSIS OF THE FINAL IMAGE FOR REPEAT SCREENING ASSAY IS AS FOLLOWS:CELL#/WELL# RX STRENGTH TYPING VISIBLE INTERP. 1 / E1 4 E=E+ VISUALLY NEGATIVE WELL; 2 / F1 19 E+E= VISUALLY NEGATIVE WELL; VERY SIGHT FUZZY BACKGROUND;A SYNOPSIS OF THE FINAL IMAGE FOR PLATE ANTIBODY ID PANEL IS AS FOLLOWS:CELL#/WELL# RX STRENGTH TYPING VISIBLE INTERP; 1 / A1 13 E+E+ VISUALLY NEGATIVE WELL; 2 / B1 8 E=E+ VISUALLY NEGATIVE WELL; 3 / C1 23 E+E= SLIGHT RED CELL ADHERENCE; 4 / D1 8 E=E+ VISUALLY NEGATIVE WELL; 5 / E1 7 E=E+ VISUALLY NEGATIVE WELL; 6 / F1 8 E+E+ VISUALLY NEGATIVE WELL; 7 / G1 6 E=E+ VISUALLY NEGATIVE WELL; 8 / H1 6 E=E+ VISUALLY NEGATIVE WELL; 9 / A2 8 E=E+ VISUALLY NEGATIVE WELL; 10 / B2 9 E=E+ VISUALLY NEGATIVE WELL; 11 / C2 9 E=E+ VISUALLY NEGATIVE WELL; 12 / D2 8 E=E+ VISUALLY NEGATIVE WELL; 13 / E2 7 E=E+ VISUALLY NEGATIVE WELL; 14 / F2 7 E=E+ VISUALLY NEGATIVE WELL.COULD NOT RULE OUT A SAMPLE AS THE CAUSE OF THE UNEXPECTED REACTIVITY. THE CUSTOMER DOES NOT HAVE SAMPLE AVAILABLE FOR FURTHER SEROLOGICAL INVESTIGATION. >< CUSTOMER REPORTED AN UNEXPECTED NEGATIVE RESULT WHEN TESTING A SAMPLE BY 2 CELL SCREENING ASSAY AND ANTIBODY ID PANEL ON THE GALILEO. REPEAT TESTING ALSO RESULTED NEGATIVE. NO ADVERSE EVENT OCCURRED AS A RESULT OF THE UNEXPECTED NEGATIVE REACTIVITY. | 0
1667997 | A SERVICE CALL WAS MADE. REAGENT CARRYOVER SUSPECTED. CHANGED REAGENT PROBE AND CARRIED OUT QC TESTS. REAGENT QC SUCCESSFUL. >< CUSTOMER REPORTED ABO DISCREPANCIES ON THE GALILEO. A POSITIVE DONOR SAMPLES ARE RESULTING AS AB POSITIVE WITH 4+ REACTIONS. | 0
1156304 | WHEN UNITS FOR WHICH WORK HAS BEEN PERFORMED WITH THE STATUS OF AVAILABLE, CROSSMATCHED, ISSUED, SHIPPED, QUARANTINED, PENDING UNIT DESTRUCTION AND DESTROYED ARE DELETED FROM THE COMPONENT PROCESSING WORKSHEET, THE AUDIT HISTORY INFORMATION IS LOST. TABLES TB_UNIT (WHICH HOLDS UNIT STATUS AND BLOOD TYPE) AND TB_COMP_CREATE (LINKS A CHILD UNIT TO ITS PARENT) ARE DELETED THROUGH THIS PROCESS. AFTER DELETING A UNIT, TBAU_UNIT IS THE ONLY TABLE WHICH HOLDS ANY HISTORY RECORDS THAT A UNIT_ID WAS LINKED TO A DONATION ID (UNIT_SOURCE_ID) AND POSSIBLY THERE ARE STATUS CHANGES RECORDED IN TBAU_UNIT_TRAIL. WHEN YOU DELETE COMPONENTS FROM THE COMPONENT PROCESSING WORKSHEET, YOU ACTUALLY CASCADE DELETE THEM FROM EVERYWHERE IN THE DATABASE THAT IS LINKED. | 1
1743541 | THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. >< (B)(4). THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. | 0
1798623 | (B)(4). THE PRODUCT HAS BEEN REQUESTED FOR INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF THE METER IS RETURNED. CUSTOMERS AND RETAILERS HAVE BEEN INFORMED THROUGH THE FA16MAY2006 LETTER. >< THE CUSTOMER REPORTED AN ISSUE WITH THEIR METER WHICH SUGGESTS THE MEMORY OVERWRITE MALFUNCTION HAS OCCURRED. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. | 0
1725713 | DURING A LAPAROSCOPIC SIGMOID RESECTION, MD WAS UNABLE TO TAKE PICTURES DURING THE PROCEDURE FOR THE PATIENT'S RECORD DUE TO A PAPER JAM IN THE PRINTER. THE BIOMEDICAL ENGINEER ON CALL WAS PAGED. STEPS GIVEN OVER THE PHONE WERE UNSUCCESSFUL IN REMOVING THE PAPER JAM. | 1
1581253 | IT WAS FOUND THAT SOME TESTS THAT WERE ELECTRONICALLY ENTERED WERE CAPRICIOUSLY DISCONTINUED OR CANCELLED BY THE DEVICE PRIOR TO THE TEST HAVING BEEN COMPLETED. THIS CAUSED DELAYS AND DISRUPTIONS OF CARE WITH VARYING DEGREES OF ADVERSITY FOR THE PTS AND DISTRUST FOR THE FUNCTIONALITY OF THE SYSTEM BY THOSE USING IT. THE ABSOLUTE INCIDENCE OF THIS OCCURRING HAS NOT BEEN DETERMINED BUT IS FREQUENT. | 1
1868123 | THE MPA PLUS IS RESPONSIBLE FOR SAMPLE PRE-PROCESSING, INCLUDING CENTRIFUGING, ALIQUOTING AND BARCODE LABELING OF SPECIMENS PRIOR TO ANALYSIS ON THE COBAS 6000 ANALYZERS. THE CENTRIFUGE ON THE MPA AT THE CUSTOMER SITE WAS NOT SPINNING SPECIMENS CORRECTLY SENDING UN-SPUN SAMPLES TO THE COBAS 6000 ANALYZERS FOR ANALYSIS. THE USER RECEIVED QUESTIONABLE RESULTS FOR SODIUM, POTASSIUM AND CHLORIDE ON TWO DIFFERENT COBAS 6000 ANALYZERS SERIAL NUMBERS (B)(4) AND (B)(4). THE EVENT INVOLVED FIVE PATIENT SAMPLES. TWO PATIENT SAMPLES GAVE DISCREPANT RESULTS FOR POTASSIUM. PATIENT SAMPLE 1, ON COBAS 6000 S/N (B)(4) THE ORIGINAL POTASSIUM RESULT GAVE 5.8 MMOL/L. THIS RESULT WAS REPORTED OUTSIDE THE LABORATORY. THE SAMPLE WAS REPEATED WHICH RESULTED AT 3.98 MMOL/L. THE REPEAT POTASSIUM RESULT OF 3.98 MMOL/L WAS REPORTED OUTSIDE THE LABORATORY AS THE FINAL RESULT FOR THIS PATIENT SAMPLE. PATIENT SAMPLE 2, ON COBAS 6000 S/N (B)(4) THE ORIGINAL POTASSIUM RESULT GAVE 7.9 MMOL/L. THIS RESULT WAS REPORTED OUTSIDE THE LABORATORY. THE SAMPLE WAS REPEATED WHICH RESULTED AT 5.2 MMOL/L. THE REPEAT POTASSIUM RESULT OF 5.2 MMOL/L WAS REPORTED OUTSIDE THE LABORATORY AS THE FINAL RESULT FOR THIS PATIENT SAMPLE. NO ADVERSE EVENT WAS ALLEGED REGARDING THIS EVENT. THE FIELD SERVICE REPRESENTATIVE FOUND THE MPA CENTRIFUGE MISADJUSTED. HE ADJUSTED THE TEMPERATURE SETTING AND CLEANED CENTRIFUGE SAMPLE BUCKETS. PERFORMANCE TESTS WERE RAN WITH NO ERRORS AND WERE ACCEPTABLE. >< . | 0
1460751 | A CUSTOMER REPORTED THAT REPEAT TESTING ON A NEW SAMPLE THAT HAD INITIALLY TESTED NEGATIVE ON THE PROVUE ON (B) (6)2009, WAS TESTING NEGATIVE AGAIN ON (B) (6)2009. THE CUSTOMER TESTED A NEW SAMPLE FROM THE SAME PT ON (B) (6)2009 AND OBTAINED NEGATIVE RESULTS ON THE PROVUE USING VSS236 AND GEL CARD LOT 013009001-21. THE CUSTOMER VISUALLY INSPECTED THE GEL CARDS THAT WERE TESTED ON THE PROVUE AND CONFIRMED THAT ALL REACTIONS RESULTED BY THE PROVUE WERE NEGATIVE. THE SAME SAMPLE WAS TESTED USING MANUAL GEL WITH THE SAME SCREENING CELLS AND GEL CARDS. CELL 1 WAS WEAKLY POSITIVE (W- 1+), CELL 2 AND CELL 3 WERE NEGATIVE. THE SAMPLE WAS THEN TESTED USING MANUAL GEL WITH 0.8% RESOLVE PANEL A, VRA 129, AND THE SAME LOT OF GEL CARDS. ALL K+ CELLS OF THE PANEL REACTED WEAK-1+, DEMONSTRATING THE ANTI-K. >< AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE. THE FE CHECKED THE TEMPERATURE REGULATION OF THE INCUBATOR, PRESSURE AND VACUUM ADJUSTMENTS AND CENTRIFUGE RPMS, WHICH WERE ALL WITHIN SPECIFICATIONS. WAD DIAGNOSTICS TESTING WAS ACCEPTABLE. THE FE INDICATED THAT THE CUSTOMER SENT SOME OF THE SAMPLES IN QUESTION TO A REFERENCE LAB FOR TESTING. THE REFERENCE LAB OBTAINED NEGATIVE RESULTS USING TUBE METHOD FOR ANTIBODY IDENTIFICATION. THE FE SUBMITTED THE LOG FILES TO OCD SECOND LEVEL SUPPORT (TAC) FOR INVESTIGATION SINCE THE ISSUE WAS NOT RESOLVED. TAC REVIEWED THE LOG FILES AND INDICATED THAT THE DATA FILES ONLY CONTAINED THE MOST RECENT (B) (6) FILE WHICH BEGAN ON (B) (6)2009. BASED ON THE DATA OBTAINED FROM THE CUSTOMER REGARDING THE SAMPLE TESTED ON (B) (6)2009, THE TIFF IMAGES SHOWED NO REACTIONS IN THE MICROTUBES. THE PROVUE READ THE REACTIONS AS NEGATIVE. BITMAP IMAGES WERE REVIEWED BY 2ND LEVEL SUPPORT FOR BATCH 318 (SAMPLE FROM (B) (6)) AND ALL WELLS WERE CLEARLY NEGATIVE. PROVUE READ REACTION CORRECTLY AS NEGATIVE. THE READER CAMERA BRILLO SETTING FROM THE CALIB VALUES FILE DATED 05/29/09 WAS AT 117 WHICH IS WITHIN SPECIFICATIONS. THIS CUSTOMER HAS NOT LOGGED ANY SIMILAR COMPLAINTS AGAINST THIS ANALYZER SINCE (B) (6)2009. (B) (4). REFER TO MEDWATCH # MFR 1056600-2009-00137 REGARDING THE INCIDENT THAT OCCURRED ON (B) (6)2009. | 0
1560140 | EVAL OF THE INCIDENT DEVICE CONFIRMED THE DEVICE WOULD NOT POWER ON. THE PROBLEM WAS IDENTIFIED AS A FAULTY MAIN CPU PCBA. THE FAULTY PCBA WAS REPLACED AND THE MODULE WORKED PER SPECIFICATIONS. REVIEW OF AVAILABLE DATA DOES NOT INDICATE THAT THIS FAILURE IS PART OF A TREND. WE CONSIDER THIS FAILURE TO BE AN ISOLATED EVENT WITH NO FURTHER CORRECTIVE ACTION OR INVESTIGATION NEEDED. WE HAVE CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AND CONSIDER THIS ISSUE CLOSED. >< SPACELABS RECEIVED A REPORT THAT A PORTABLE MONITOR WENT BLANK DURING A PROCEDURE. | 1
1776062 | DURING THE COURSE OF CERVICAL SPINE SURGERY, THE O-ARM OF THE MEDTRONICS IMAGING SYSTEM LOST CONNECTIVITY AND FAILED TO RESPOND. THE PATIENT REQUIRED PHYSICAL REMOVAL FROM THE O-ARM WHEN THE MANUAL OPENING FAILED AS WELL. THE SURGERY WAS EVENTUALLY COMPLETED; NO DISCERNABLE HARM NOTED TO THE PATIENT.====================== HEALTH PROFESSIONAL'S IMPRESSION======================LOSS OF POWER AND SUBSEQUENT MANUAL FAILURE PREVENTED EASY ACCESS/POSITIONING OF THE PATIENT INTRAOPERATIVELY. | 1
1882015 | THE CENTRAL MONITORING UNIT (CMU) INFORMED THE ADMINISTRATIVE NURSE MANAGER (ANM) OF MULTIPLE TELEMETRY FAILURES (16), AS WELL AS PAGING THE ON-CALL BIOMED TECH. THE ANM DETERMINED WHICH PATIENTS WERE HAVING ARRHYTHMIAS ON AFFECTED NURSING UNITS AND CONTACTED RESPECTIVE NURSES TO EXECUTE DOWNTIME PROCEDURE (UTILIZE FREE-STANDING MONITORS WITH AUDIBLE ALARMS ON VULNERABLE PATIENTS, WITH CONTINUOUS ROUNDING FOR ASSESSMENTS). | 1
1898455 | RYMED INVISION-PLUS NEURTAL WITH A LOT NUMBER OF R00420 WAS FOUND TO HAVE A BROKEN PIECE OF CENTER CORE IN MICRO IV TUBING AFTER RUNNING IV POTASSIUM VIA A CENTRAL LINE. EQUIPMENT WAS KEPT AND WILL BE SENT TO THE MFR. | 0
1892314 | UNABLE TO PRINT MONITOR STRIPS ON THIS PARTICULAR PATIENT IN THE ROOM AND THEREFORE NOT ABLE TO DOCUMENT. PATIENT NEEDED TO GO BACK TO SURGERY AND DEFINITELY NEEDED PROPER DOCUMENTATION. | 1
1592119 | 'IVENT' VENTILATOR ON ELDERLY PATIENT WITH HISTORY OF LUNG CANCER AND RESPIRATORY FAILURE STARTED TO ALARM 'HIGH O2.' VENTILATOR WAS SWITCHED TO 760 VENTILATOR. THERE WAS NO HARM TO THE PATIENT. | 0
1883512 | ANESTHESIA DELIVERY UNIT STOPPED VENTILATING PATIENT, CRNA CERTIFIED REGISTERED NURSE ANETHESTHETIST AND SRNA STUDENT REGISTERED NURSE ANESTHETIST HAD TO HAND VENTILATE UNTIL NEW MACHINE COULD BE BROUGHT INTO ROOM. CRNA BOARD RUNNER LOCATED ANOTHER MACHINE AND BROUGHT INTO ROOM. PATIENT REMAINED HEMODYNAMICALLY STABLE AND AMNESTIC THROUGHOUT THE CHANGE IN ANESTHESIA MACHINE. VENTILATOR THOUGHT TO HAVE FAILED WHILE CHANGING BETWEEN PRESSURE CONTROL VENTILATION AND VOLUME CONTROL. BIOMED REVIEWED ERROR CODES ON MACHINE. CONTACTED MANUFACTURER AND DETERMINED THAT THE INSPIRATIONAL PROPORTIONAL VALVE NEEDS TO BE REPLACED. WILL REPAIR AND RETURN TO SERVICE.======================HEALTH PROFESSIONAL'S IMPRESSION======================ADU MALFUNCTIONED | 0
1358029 | IT WAS REPORTED THAT THE AC/DC BOARD BURNED OUT ON THE CARDIOCAP 5 BETWEEN CASES. THE STAFF REPORTED A LIGHT SCENT OF SOMETHING BURNING IN THE ROOM. THERE WAS NO PT BEING MONITORED ON THE DEVICE AT THE TIME OF THE EVENT. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1362534 | (B) (6) PRIVACY LAW PROHIBITS RELEASING PATIENT INFORMATION. GE HEALTHCARE INVESTIGATED THE NMF CPU MOTHERBOARD AND F-CPU POWER BOARD. BOTH BOARDS WERE BADLY BURNED. AS A RESULT, THE BOARDS WERE UNABLE TO BE REPAIRED OR TESTED. VISUAL INSPECTION REVEALED THAT A CAPACITOR AND PRINTED CIRCUIT BOARD AT THE SAME LOCATION BENEATH THE CAPACITOR WAS BURNED, RESULTING IN A HOLE IN THE PRINTED CIRCUIT BOARD ON THE NMF CPU MOTHERBOARD. SEVERAL COMPONENTS WERE BURNED ON THE F-CPU POWER BOARD, RESULTING IN A HOLE IN THE PRINTED CIRCUIT BOARD AS WELL. DUE TO THE EXTENT OF THE DAMAGE TO BOTH BOARDS, THE ROOT CAUSE FOR THE COMPONENT FAILURES REMAINS UNKNOWN, AS NEITHER BOARD COULD BE TESTED. THE DEFECTIVE BOARDS WERE REPLACED AND THE MONITOR WAS PLACED BACK INTO USE. >< IT WAS REPORTED THAT PARTS OF THE POWER BOARD AND CPU MOTHERBOARD FAILED ON THE S/5 CRITICAL CARE MONITOR, RESULTING IN A LOSS OF MONITORING WITHOUT AN ALARM. NO INJURY WAS REPORTED. | 1
1083972 | IT WAS REPORTED THAT UNITY NETWORK ARRHYTHMIA ALARMS (CRISIS AND/OR WARNING LEVEL) ARE NOT GENERATED ON ICENTRAL. NO PT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1908033 | A SITE INITIALLY REPORTED TO GE HEALTHCARE ON (B)(6) 2008 THAT A NURSE FELT A TINGLING SENSATION ON HER FINGERTIPS FROM TYPING IN PT INFO WHILE PERFORMING AN EKG (ELECTROCARDIOGRAPH) ON A PT. THE NURSE'S FINGER TIPS REPORTEDLY TURNED WHITE ON THE SURFACE, HOWEVER NO SERIOUS INJURY WAS REPORTED AT THIS TIME. ON (B)(6) 2010, THE SITE INFORMED GE THAT THE NURSE INVOLVED IN THE REPORTED EVENT HAD SUSTAINED SECOND DEGREE BURNS ON THE WRIST AND HAND. THE NURSE WAS EVALUATED BY A PHYSICIAN AND WAS TREATED WITH BACITRACIN AND BANDAGES. THE NURSE WAS REPORTEDLY ABLE TO RETURN TO WORK THE SAME DAY OF THE EVENT. >< THE MFG SITE RECEIVED THE UNIT FOR INVESTIGATION. THE KEYPAD WAS VIEWED UNDER A 40X MICROSCOPE TO OBSERVE ANY PUNCTURE IN THE KEYPAD. NO ABNORMALITY WAS OBSERVED. THE UNIT ALSO UNDERWENT SAFETY TESTS CONFIRMING THAT THE UNIT'S LEAKAGE CURRENT WAS WITHIN SPECIFICATION. THE UNIT WAS THEN TESTED WHILE SWITCHED OFF AND WHILE SWITCHED ON, BOTH ON MAINS AND BATTERY POWER. NO FAILURES WERE OBSERVED. THE REPORTED ISSUE COULD NOT BE REPLICATED. BASED ON THE RESULTS OF THE INVESTIGATION, AND DUE TO THE CONSTRUCTION OF THE KEYPAD, LOW VOLTAGE AND MINIMAL CURRENT OF THE DEVICE, IT IS UNLIKELY THAT THE UNIT'S KEYPAD CONTRIBUTED TO THE PT BURN. THE ROOT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED. AS A PREVENTIVE MEASURE, THE UNIT WAS REPLACED AT THE CUSTOMER SITE. | 0
1593049 | ON (B)(6) 2009, MCKESSON RECEIVED A REPORT FROM A CLINICAL SITE, WHERE THE HORIZON CARDIOLOGY HEMO SYSTEM IN THE CATHETERIZATION LAB PRESENTED UNUSUAL ECG WAVEFORMS ON THE REAL TIME MONITOR. THE HOSP REPORTED THAT BASED ON THE PHYSICIAN'S INTERPRETATION OF THE ECG WAVEFORMS, AFTER THE PROCEDURES WERE COMPLETE, TWO PTS WERE TRANSFERRED TO ANOTHER AREA OF THE HOSP FOR ADDITIONAL OBSERVATION. DURING THE ADDITIONAL OBSERVATION, NO UNUSUAL ECG RHYTHMS WERE DETECTED USING OTHER PHYSIOLOGICAL MONITORING EQUIPMENT. THE HOSP ALSO REPORTED THAT NO FURTHER TREATMENT WAS PROVIDED. >< MCKESSON PERFORMED AN INVESTIGATION TO DETERMINE THE ROOT CAUSE OF THE EVENT. IT WAS FOUND THAT IF SOUND PROMPTS ARE CONFIGURED ON THE HORIZON CARDIOLOGY HEMO SYSTEM, A CONFLICT BETWEEN THE INTEGRATED SOUND COMPONENT ON THE REAL-TIME MONITORING COMPUTER AND THE PROCESS OF ACQUIRING PT'S PHYSIOLOGICAL PARAMETERS RESULTED IN LOSS OF PARTS OF THE ECG WAVEFORMS WHEN THERE WAS ACTIVITY USING THE SOUND COMPONENT. THE INTEGRATED SOUND COMPONENT WAS DISABLED ON (B)(6) 2010 AND WAS PERMANENTLY REPLACED WITH AN ADDITIONAL SOUND ADAPTER ON (B)(6) 2010. THE RESOLUTION OF THE ISSUE WAS VALIDATED WITH THE CLINICAL STAFF ON-SITE. NO OTHER EVENTS WERE REPORTED SINCE DISABLING THE INTEGRATED SOUND COMPONENT ON (B)(6) 2010. >< MCKESSON PERFORMED AN INVESTIGATION TO DETERMINE THE ROOT CAUSE OF THE EVENT. IT WAS FOUND THAT IF SOUND PROMPTS ARE CONFIGURED ON THE HORIZON CARDIOLOGY HEMO SYSTEM, A CONFLICT BETWEEN THE INTEGRATED SOUND COMPONENT ON THE REAL-TIME MONITORING COMPUTER AND THE PROCESS OF ACQUIRING PT'S PHYSIOLOGICAL PARAMETERS RESULTED IN LOSS OF PARTS OF THE ECG WAVEFORMS WHEN THERE WAS ACTIVITY USING THE SOUND COMPONENT. THE INTEGRATED SOUND COMPONENT WAS DISABLED ON (B)(6) 2010 AND WAS PERMANENTLY REPLACED WITH AN ADDITIONAL SOUND ADAPTER ON (B)(6) 2010. THE RESOLUTION OF THE ISSUE WAS VALIDATED WITH THE CLINICAL STAFF ON-SITE. NO OTHER EVENTS WERE REPORTED SINCE DISABLING THE INTEGRATED SOUND COMPONENT ON (B)(6) 2010. >< ON (B)(6) 2009, MCKESSON RECEIVED A REPORT FROM A CLINICAL SITE, WHERE THE HORIZON CARDIOLOGY HEMO SYSTEM IN THE CATHETERIZATION LAB PRESENTED UNUSUAL ECG WAVEFORMS ON THE REAL TIME MONITOR. THE HOSP REPORTED THAT BASED ON THE PHYSICIAN'S INTERPRETATION OF THE ECG WAVEFORMS, AFTER THE PROCEDURES WERE COMPLETE, TWO PTS WERE TRANSFERRED TO ANOTHER AREA OF THE HOSP FOR ADDITIONAL OBSERVATION. DURING THE ADDITIONAL OBSERVATION, NO UNUSUAL ECG RHYTHMS WERE DETECTED USING OTHER PHYSIOLOGICAL MONITORING EQUIPMENT. THE HOSP ALSO REPORTED THAT NO FURTHER TREATMENT WAS PROVIDED. | 1
1935211 | THE ON/OFF TAB BROKE OFF WHEN THE PATIENT ATTEMPTED TO SUCTION HIS MOUTH. REPORTER INDICATED THE PLASTIC SLIDING ON/OFF BUTTON BROKE. THE PATIENT WAS NOT HARM AND A NEW DEVICE OF THE SAME MODEL WAS USED TO COMPLETE THE PROCEDURE. | 0
1358031 | IT WAS REPORTED THAT THE PT'S WEIGHT ON PT DATA DEMOGRAPHIC WENT DOWN INCREMENTALLY TO (B) (6) LBS. THE EVENT IS RELATED TO AN ISSUE WHERE THE ANESTHESIA MACHINE OR VENTILATOR DOES NOT RECEIVE THE PT WEIGHT THAT IS CHANGED OR ENTERED IN THE PT MONITOR. THE ANESTHESIA MACHINE OR VENTILATOR SENDS ITS OWN PT WEIGHT TO THE PT MONITOR THAT OVERRIDES THE PT WEIGHT CHANGED OR ENTERED IN THE PT MONITOR, RESULTING IN WRONG PT WEIGHT AND CALCULATED BSA. INCORRECT CALCULATED BSA VALUE MAY FURTHER RESULT IN INCORRECT INDEXED HEMODYNAMIC, OXYGENATION AND VENTILATION ON THE PT MONITOR. WRONG PT WEIGHT MAY AFFECT DRUG CALCULATION VALUES. NO PT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1784728 | ON (B)(6) 2010, THE DATEX-OHMEDA S/5 AVANCE ANESTHESIA SYSTEM CHECK WAS PERFORMED PRIOR TO THE FIRST CASE OF THE DAY. THE MACHINE, SYSTEM, CIRCUIT AND LOW P LEAK TEST WERE COMPLETED. THE LOW P LEAK TEST WAS COMPLETED WITH THE ORANGE RUBBER DOUBLE ENDED STOPPER IN PLACE. THE ORANGE STOPPER WAS NOT REMOVED PRIOR TO THE START OF THE FIRST ANESTHESIA CASE AND THE BREATHING CIRCUIT WAS CONNECTED TO THE ORANGE STOPPER. FOLLOWING THE LOW P LEAK TEST THE NURSE, WHO PERFORMED THE SYSTEM CHECKS, ENTERED THROUGH THE SCREENS TO EXIT THE CHECK. THE ANESTHESIOLOGIST ENTERED THE ROOM TO FIND A START CASE SCREEN AND PROCEEDED TO COMPLETE A CIRCUIT CHECK. THE CASE WAS STARTED AND THE ANESTHESIOLOGIST NOTICED THAT SHE WAS NOT ABLE TO VENTILATE THE PT. IT WAS DISCOVERED THAT THE ORANGE STOPPER REMAINED IN THE ANESTHESIA MACHINE. AT THAT TIME, THE STOPPER WAS REMOVED IMMEDIATELY AND THE BREATHING CIRCUIT WAS ATTACHED TO THE INSPIRATORY FLOW SENSOR PORT. THERE IS NOT A HARD STOP IN PLACE IN THIS ANESTHESIA MACHINE TO PREVENT THIS TYPE OF EVENT TO OCCUR WHEN THE STOPPER REMAINS IN THE CHECK PORT. A SIMPLE REDESIGN OF THE STOPPER WOULD PREVENT THIS FROM HAPPENING AGAIN. | 0
1881903 | PT HAVING A PERMANENT PACEMAKER IMPLANTED AND ECG ON MONITOR CUT OUT. PORTABLE MONITORING SYSTEM UTILIZED. RISK TO SAFETY OF PT. | 0
1813797 | ERROR IDENTIFIED IN (B)(4) -BARCODING- DATABASE. ON LOADING A NEW PRODUCT IN OUR FORMULARY, THE DATABASE AUTOPOPULATED THE 'DESCRIPTION' FIELD WITH 'EPINEPHRINE' FOR THE PHARMEDIUM PRODUCTS FOR NOREPINEPHRINE, (B)(4) AND (B)(4). OUR CERNER BRIDGE VERSION IS (B)(4) WITH DATABASE VERSION (B)(4). | 1
1888547 | DURING A PROCEDURE, THE ANESTHESIA MACHINE STOPPED WORKING AND GAVE A MESSAGE 'TRY ANOTHER CASSETTE, SCHEDULE SERVICE.' THIS HAS OCCURRED BEFORE WITH NO PROBLEMS FOUND IN THE LOGS OR VAP TEST. MACHINE PASSED ALL CHECKS. THE MACHINE WAS TURNED OFF, THEN ON AGAIN TO CLEAR AND BEGAN WORKING FOR THE REMAINDER OF THE CASE. THE MACHINE WAS CHANGED OUT OF THE OR SUITE AND REPLACED. THERE WAS NO INJURY TO THE PATIENTS.======================HEALTH PROFESSIONAL'S IMPRESSION======================CIRCUIT BOARD FAILURE. | 0
1619928 | A GE REP HAS NOT YET EVALUATED THE SYSTEM DUE TO CUSTOMER SCHEDULE. >< CUSTOMER REPORTED THE SYSTEM IS DISPLAYING INTERLOCK FAILURE MESSAGES. NO PT INJURY WAS REPORTED. | 0
1825089 | ELDERLY FEMALE PATIENT TAKEN TO OR FOR TOTAL KNEE PROCEDURE. IT WAS REPORTED BY THE ANESTHESIA TECH THAT THE DATEX OHMEDA AESTIVA-5 ANESTHESIA MACHINE HAD ENCOUNTERED PROBLEMS. IT WAS NOTED THE BELLOWS ON THE MACHINE WERE NOT RISING APPROPRIATELY. BACK UP MACHINE WAS OBTAINED AND SWITCHED OUT WITH ORIGINAL MACHINE WITH NO FURTHER REPORTED ISSUES WITH EITHER THE 2ND MACHINE OR THE PATIENT. ANESTHESIA TECH REPORTS UPON TROUBLESHOOTING THE ISSUE IT WAS NOTED THE MACHINE HAD AN EXCESSIVE AMOUNT OF DUST ORIGINATING FROM THE CO2 ABSORBENT CANISTER. WE HAVE REPORTED THIS PROBLEM WITH THIS PRODUCT PREVIOUSLY BUT THIS IS THE FIRST TIME IT HAS DIRECTLY AFFECTED PATIENT/PROCEDURE. OUR HOSPITAL HAD DETERMINED WE WOULD NOT BE USING THIS PRODUCT AS SOON AS THE REPLACEMENT PRODUCT WE HAD FOUND (AMSORB PLUS) ARRIVED AT OUR HOSPITAL. OUR FACILITY HAS NOW JUST LEARNED THE SUPPLIER OF SODASORB HAS NOW DETERMINED THEY WILL NOT BE ORDERING FROM THIS PARTICULAR VENDOR ANYMORE AND WILL BE RETURNING TO THEIR ORIGINAL PRODUCT (MEDISORB) WHICH HAD BEEN DISCONTINUED. NO INJURY TO PATIENT, JUST A DELAY IN PROCEDURE WITH HAVING TO SWITCH TO AN ALTERNATE MACHINE AND AN IMPACT OF TIME AS IT TOOK OUR ANESTHESIA STAFF TIME TO CLEAN LARGE AMOUNT OF DUST OUT OF MACHINE AND RUN MACHINE THROUGH ITS PACES BEFORE FEELING SECURE ENOUGH TO RETURN IT TO SERVICE.======================MANUFACTURER RESPONSE FOR PREPACKAGED CO2 ABSORBENT CARTRIDGE, SODASORB======================MANUFACTURER IS AWARE OUR HOSPITAL HAS BEEN VERY UNHAPPY WITH THE SODASORB PRODUCT. WE HAD USED MEDISORB FOR QUITE SOME TIME WITH NO ISSUES AND THEN SUDDENLY SODSASORB JUST STARTED BEING DELIVERED IN ITS PLACE. WE EXPRESSED CONCERN AND STARTED EVALUATING OTHER PRODUCTS FROM OTHER MANUFACTURERS. WE HAVE NOW LEARNED THE SODASORB PRODUCT WILL NOW BE DISCONTINUED AND GE WILL BE RETURNING TO THE ORIGINAL MEDISORB PRODUCT. | 0
1603924 | A 100% OF ACQUIRED MEASUREMENTS/CALCULATIONS DID NOT TRANSFER OVER TO THE STRUCTURED REPORT OF MCKESSON HORIZON CARDIOLOGY FROM A (B) (4) SEQUOIA ULTRASOUND SYSTEM. MISSING DATA: ALL MEASUREMENTS PERFORMED ON PT: 2D AND DOPPLER. | 1
1798744 | I AM REPORTING ON THE ADVERSE REACTIONS OF THE MERCURY AMALGAM FILLINGS PLACED IN MY MOUTH WHEN I WAS STILL AN ADOLESCENT. I WAS WHAT ONE WOULD CALL AN ALLERGIC AND SENSITIVE CHILD. AFTER HAVING THE AMALGAMS PLACED, MY HEALTH BEGAN TO EXHIBIT STRANGE SYMPTOMS GEOGRAPHIC TONGUE, TINEA VERSICOLOR, AUTOIMMUNE ISSUES, ETC. MORE IMPORTANTLY, MY MOTHER IS A MEDICAL PROFESSIONAL, AND HAD SHE KNOWN THAT MERCURY WAS A 50 PERCENT COMPONENT OF AMALGAM FILLINGS, SHE NEVER WOULD HAVE LET A DENTIST NEAR ANYONE IN OUR FAMILY. RECENTLY, IN ADDITION TO MY OTHER DIAGNOSES, I WAS DIAGNOSED WITH ADHD AND HEAVY METAL TOXICITY. I HAD MY MERCURY AMALGAM FILLINGS REMOVED IN (B)(6)2010. AT THIS POINT, MY SYMPTOMS RANGED FROM GENERAL ERETHISM, SPONTANEOUS LACTATION, AGORAPHOBIA, MISCARRIAGES, MIGRAINES, BRAIN FOG, FIBROMYALGIA, CFS, INABILITY TO LEARN NEW THINGS, DIFFICULTY READING, FITS OF RAGE, MOOD SWINGS, MULTIPLE FOOD ALLERGIES -INCLUDING GLUTEN AND DAIRY- ETC. AND JUST TO BE CLEAR, MY AGORAPHOBIA WAS SO DEBILITATING, I HAD TO WITHDRAW FROM A MAJOR UNIVERSITY UNDER FULL ACADEMIC SCHOLARSHIP IN 2000. THE TWO DAYS POST-REMOVAL, I HAD NOT FELT BETTER THAN I HAD FELT IN YEARS. IT WAS INSTANTANEOUS, BUT NOT LONG LASTING. I WILL NOT SAY HOW, FOR FEAR THAT THE FDA WILL BAN ANY TREATMENT NOT DIRECTLY PROFITING THE PHARMACEUTICAL INDUSTRY, BUT I WILL SAY THAT PEOPLE ARE GETTING BETTER REGARDLESS OF WHAT THE FDA CHOOSES TO ALLOW IN THESE ARCHAIC DENTAL DEVICES. I WILL SUBMIT WITH REQUEST UNDER SUPERVISION OF MY ATTORNEY. | 0
1773657 | TESTING INDICATED A VACUUM INSULATION FAILURE IN THE CRYOPROBE. FROSTING OF THE SHAFT WAS CONFIRMED. NO PT INJURY WAS REPORTED. >< DURING SURGERY THE SHAFT BEGAN FREEZING UP SHAFT, PROBE WAS REPLACED AND THE SURGERY WAS COMPLETED. NO PT INJURY REPORTED. | 0
1855377 | NURSE USING SAGE ORAL CARE KIT TO CLEAN ICU PT'S MOUTH. PT SEDATED AND ON MECHANICAL VENTILATION. PT BIT OFF GREEN SPONGE AT THE TIP OF THE SWAB. NURSE HAD TO RETRIEVE BROKEN OFF PORTION FROM PT'S MOUTH. POTENTIAL FOR CHOKING. PT HAD NO SIGNIFICANT ADVERSE OUTCOME. | 0
1582407 | VENTILATOR ALARM SOUNDED, DISPLAY SHOWING 'LOW OXYGEN.' PATIENT'S OXYGEN SATURATION SHOWED BETWEEN LOW 80S- 90%. PATIENT WAS MANUALLY VENTILATED, OXYGEN SATURATION IMPROVED TO 99-100%. VENTILATOR WAS CHANGED AND THERE WAS NO FURTHER EPISODES OF HYPOXIA. | 0
1474595 | REPORTED OBSERVATION: PATIENT DEVELOPED CANALICULITIS IN THE LEFT EYE AFTER INSERTION OF AN OASIS MEDICAL FORM FIT HYDROGEL CANALICULAR PLUG. (B)(4). PRODUCT LOT NUMBER: LH0409F. PRODUCT INSERTED ON: (B)(6) 2009. DATE OF COMPLICATION: (B)(6) 2009. REPORTED TO (B)(4) ON (B)(6) 2009. >< DEVICE WAS NOT RETURNED FOR EVALUATION. A REVIEW OF BATCH PRODUCTION AND STERILIZATION RECORDS SHOWS THAT THERE WERE NO REMARKABLE EVENTS, ANOMALIES OR DEVIATIONS DURING PRODUCTION OF FORM FIT HYDROGEL CANALICULAR PLUG, (B)(4), PRODUCT LOT NUMBER LH0409F. A TOTAL OF (B)(4) FORM FIT HYDROGEL CANALICULAR PLUGS, (B)(4), PRODUCT LOT NUMBER LH0409F WERE RELEASED AND DISTRIBUTED. THERE HAVE BEEN NO OTHER ADVERSE EVENTS REPORTED BY ANY OTHER USER/CUSTOMER INVOLVING PLUGS FROM (B)(4), LOT LH0409F. TWO EVENTS WERE REPORTED BY THIS USER/CUSTOMER, THE OTHER REPORT IS DOCUMENTED ON MFG. REPORT #2083373-2009-00002. THE PLUG WAS IRRITATED FROM THE CANALICULAR, THE PATIENT WAS PRESCRIBED ANTIBIOTIC DROPS. THE ISSUE WAS RESOLVED FOR THE PATIENT, NO FURTHER TREATMENT WAS REQUIRED. (B)(4). CANNOT DEFINITIVELY DETERMINE THAT THE FORM FIT HYDROGEL CANALICULAR PLUG, (B)(4), FROM PRODUCT LOT NUMBER LH0409F WAS DIRECTLY OR INDIRECTLY RESPONSIBLE FOR THE REPORTED OBSERVATION OF CANALICULITIS. | 0
1353608 | INVESTIGATION CONFIRMS THAT EACH TIME A TX03 TRANSACTION IS RECEIVED, IT IS OVERWRITING THE FIRST AND SECOND ADDENDUMS. THERE WAS NO PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. (B) (4). >< THE THIRD ADDENDUM IS OVERWRITING THE FIRST AND SECOND ADDENDUMS, LEAVING ONLY THE ORIGINAL REPORT DISPLAYED IN THE REPORT PALETTE. THE EXPECTED BEHAVIOR WOULD BE THAT ALL THREE ADDENDUMS PLUS THE ORIGINAL REPORT WOULD BE DISPLAYED IN REVERSE CHRONOLOGICAL ORDER, I.E. NEWEST ADDENDUM FIRST FOLLOWED BY NEXT NEWEST ADDENDUM, ETC, WITH THE ORIGINAL REPORT BEING LAST. | 1
1353607 | THIS MDR IS BEING SUBMITTED LATE AS THIS IS PART OF OUR ONGOING QUALITY IMPROVEMENT ACTIVITIES. WE HAVE UPDATED THE RISK ASSESSMENT CONTROL RECORD AND THE UPDATE HAS RESULTED IN THE ASSESSMENT THAT THIS REPORTED COMPLAINT MEETS THE REPORTABILITY CRITERIA AS OUTLINED IN THE RISK ASSESMENT CONTROL RECORD FOR GE HEALTHCARE. THE SITE NETWORK CONFIGURATION WAS NOT IN COMPLIANCE WITH GE SPECIFICATIONS STATED IN CENTRICITY WORKSTATION SERVICE MANUAL DOCUMENT NUMBER 2014206_001R3 AND CENTRICITY WORKSTATION INSTALLATION MANUAL DOCUMENT NUMBER 2014205_001R3. THERE WAS NO PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. (B) (4). >< AFTER A COMPLETED DIAGNOSTIC, THE IMAGES ON THE REPORTED DEVICE ARE AUTO CLOSED BY THE REPORT CLIENT. WHEN MORE THAN ONE PATIENT VIEWING WINDOW IS DISPLAYED, IMAGES OF THE PREVIOUS PATIENT REMAIN. | 1
1515311 | ON (B)(6) 2009, MCKESSON (B)(4) RECEIVED A COPY OF A MEDWATCH USER FACILITY REPORT FILED WITH FDA ON MARCH 13, 2009 BY (B)(6) MEDICAL CENTER, (B)(6). (B)(4). MCKESSON IS SUBMITTING THE ENCLOSED MEDWATCH REPORT TO CLARIFY INFO SUBMITTED BY THE USER FACILITY IN MEDWATCH REPORT (B)(4) AND TO PROVIDE EVIDENCE SUPPORTING MCKESSON (B)(4) DECISION THAT THE REPORTED EVENT DOES NOT MEET THE CRITERIA FOR A MANDATORY MEDICAL DEVICE REPORTABLE EVENT UNDER 21 CFR 803. ON FEBRUARY 23, 2009 MCKESSON (B)(4) RECEIVED INFORMATION FROM (B)(6) MEDICAL CENTER INDICATING THAT THE INFORMATION SYSTEM MODULE OF HORIZON CARDIOLOGY HEMO FROZE FOR SIX MINUTES. DISPLAY OF UPDATED OPERATOR CHARTING OF SOME CATHETERIZATION PROCEDURE DATA FOR THE PATIENT DID NOT OCCUR DURING THIS TIME. HOWEVER, VALUES FOR ALL THE PHYSIOLOGICAL PARAMETERS MONITORED DURING THE PROCEDURE CONTINUED TO BE DISPLAYED IN REAL-TIME AND ALL DATA COLLECTED DURING THE PROCEDURE WAS STORED ON THE SYSTEM. MCKESSON'S ROOT CAUSE ANALYSIS OF THE EVENT DETERMINED THAT THE PROBLEM AFFECTED ONLY HORIZON CARDIOLOGY HEMO VERSION 11.0.0.4.236 WHEN CARDIAC OUTPUT MEASUREMENT PROCESS CANNOT OBTAIN STABILIZED TEMPERATURE AND CHARTING PROCESS OF VITALS SNAPSHOT (SAMPLE) ARE BEING PERFORMED SIMULTANEOUSLY. THE DISPLAY OF UPDATED GRAPHS OF SOME CATHETERIZATION PROCEDURE DATA WOULD RESUME WHEN A WATCHDOG MECHANISM EMBEDDED IN THE SOFTWARE DETECTS THE SCENARIO. MCKESSON'S ANALYSIS ALSO DETERMINED THAT FOR THE EVENT AND THE DISCOVERED PROBLEM, ALL THE PHYSIOLOGICAL PARAMETERS MONITORED DURING A PROCEDURE CONTINUED TO BE/WOULD CONTINUE TO BE DISPLAYED IN REAL-TIME AND ALL DATA COLLECTED DURING A PROCEDURE WAS STORED/WOULD BE STORED ON THE SYSTEM. THUS, THE LIKELIHOOD OF SERIOUS INJURY TO A PATIENT AS A RESULT OF THE PROBLEM IS EXTREMELY UNLIKELY AND AS A RESULT IT DID NOT/DOES NOT MEET THE CRITERIA FOR A MANDATORY MEDICAL DEVICE REPORTABLE EVENT. A SOFTWARE UPDATE WAS DEVELOPED TO ELIMINATE THE PROBLEM AND ON MARCH 11, 2009, IT WAS INSTALLED ON THE SITE'S SIX AFFECTED HORIZON CARDIOLOGY HEMO VERSION 11.0.0.4.236 SYSTEMS. MCKESSON ALSO VERIFIED THAT THE SOFTWARE UPDATE WAS IMPLEMENTED ON ADDITIONAL TWO SITES THAT UTILIZED HORIZON CARDIOLOGY HEMO VERSION 11.0.4.236. >< DURING A CARDIAC CATHETERIZATION PROCEDURE, THE INFORMATION SYSTEM MODULE OF THE HORIZON CARDIOLOGY HEMO SYSTEM COLLECTED BUT DIDN'T ALLOW CHARTING OF PRESSURES AND CARDIAC OUTPUT PARAMETERS FOR A PERIOD OF 6 MINUTES. REAL-TIME VALUES FOR ALL THE PHYSIOLOGICAL PARAMETERS MONITORED DURING THE PROCEDURE CONTINUED TO BE AVAILABLE DURING THE ENTIRE PROCEDURE AND ALL DATA COLLECTED DURING THE CASE WAS STORED ON THE SYSTEM. THERE WAS NO HARM TO THE PATIENT. | 1
1869497 | A NURSE REPORTEDLY DISCOVERED TOWARDS THE END OF THE CASE THAT THE PT'S FACE WAS BLUE. THE ANESTHESIOLOGIST WAS NOTIFIED AND PROCEEDED TO ADMINISTER MEDICATION (E.G. EPINEPHRINE), INTUBATED THE PT, CONNECTED THE PT TO THE ANESTHESIA MACHINE AND VENTILATED THE PT BOTH MANUALLY AND MECHANICALLY. RESUSCITATION MEASURES ARE SAID TO HAVE TAKEN APPROX 45 MINS. IT IS ALLEGED THAT THE CARDIOCAP 5 MONITOR DID NOT ALARM DURING THE EVENT. FOLLOWING THE EVENT, THE PT REPORTEDLY REMAINED UNCONSCIOUS AND WAS TRANSFERRED TO ANOTHER MEDICAL CENTER, WHERE SHE SAID TO BE IMPROVING. >< THE DEVICE MANUFACTURE DATE IS UNK. THE DEVICE WAS SHIPPED ON 12/11/2005. PRIOR TO GE HEALTHCARE'S ARRIVAL FOR THE ONSITE INVESTIGATION, THE MONITOR HAD BEEN TURNED OFF BY HOSPITAL STAFF, THEREFORE, NO TREND DATA WAS AVAILABLE FOR ANALYSIS. DURING GE HEALTHCARE'S ONSITE INVESTIGATION, INITIAL INSPECTION OF THE MONITOR REVEALED THE ALARM VOLUME WAS SET TO THE DEFAULT SETTING OF 2. TESTING WAS CONDUCTED ON THE MONITOR, INCLUDING SIMULATED ALARM CONDITIONS, AND THE UNIT WAS FOUND TO OPERATE ACCORDING TO MFR'S SPECIFICATIONS. ADDITIONALLY, THE BLOOD PRESSURE CUFF WAS CONNECTED TO A HOSPITAL STAFF MEMBER, WITH THE MONITOR SET TO TAKE A BLOOD PRESSURE EVERY MIN. THIS TESTING LASTED 10 MINS, AND THE MONITOR WAS FOUND TO OPERATE AS DESIGNED. SUBSEQUENT TO GE HEALTHCARE'S INVESTIGATION, THE HOSPITAL REPORTED THEY WILL REQUEST AN INDEPENDENT PARTY PERFORM A CHECKOUT ON THE MONITOR. UNTIL THAT OCCURS, THE HOSPITAL REPORTED, THE EQUIPMENT AND OPERATING ROOM WILL REMAIN IN QUARANTINE. GE HEALTHCARE'S INVESTIGATION OF THE MONITOR FOUND THE UNIT TO FUNCTION ACCORDING TO THE MFG SPECIFICATIONS. THE REPORTED COMPLAINT COULD NOT BE DUPLICATED. AS HOSPITAL STAFF TURNED THE MONITOR OFF FOLLOWING THE CASE, THE TREND DATA FOR THE CASE CANNOT BE EVALUATED. | 1
1931327 | INFUSION DEVICE UTILIZED TO DELIVER VANCOMYCIN 2.5GM/NS 240ML CONTINUOUSLY AT 10ML/HR OVER 24 HOURS. MEDICATION INFUSED IN OVER 15 HOURS INSTEAD OF 24 HOURS PER CAREGIVER'S REPORT. | 1
1656442 | OUTPUT FROM VAPORIZER 3.5% HIGHER THAN ACTUAL SETTING. | 0
1883558 | THE ANESTHESIA MACHINE SYSTEM CHECK-OUT PROCEDURE WAS PASSED IN AM PRIOR TO START OF THE DAY'S CASES. DURING THE FIRST CASE, THE VENTILATOR TURNED ON, BUT THE CRNA CERTIFIED REGISTERED NURSE ANESTHETIST NOTICED THAT BELLOWS WERE NOT CYCLING. THE CASE PROCEEDED WITH MANUAL VENTILATION AND THE MACHINE WAS CHANGED-OUT AFTER THE CASE. THE PATIENT SUFFERED NO ADVERSE EFFECTS.======================MANUFACTURER RESPONSE FOR ANESTHESIA UNIT, DATEX-OHMEDA ADU======================RECOMMEND THE ON-SITE BIOMED TECH TO CHECK DIAGNOSTICS FOR BOTH FGCFRESH GAS CONTROL AND VENT. MONITOR DIGIPOWER VOLTAGES AND INSPECT PROPORTIONAL VALVE. IF NO PROBLEMS FOUND, DO FUNCTIONAL CHECK AND RETURN TO SERVICE. | 0
1936053 | ADDITIONAL INFO REQUESTED FROM PHYSICIAN ON (B)(4) 2010. >< THE PT'S 5CM MASS ON THE LEFT KIDNEY WAS TREATED WITH CRYOTHERAPY USING THREE R3.8L CRYO PROBES AND CONSISTED OF THREE FREEZE CYCLES. DURING PASSIVE THAW OF THE FIRST FREEZE CYCLE, INSUFFLATION WAS LOST. ONCE REGAINED, THE URETER WAS FOUND TO BE STUCK TO THE ICEBALL AND IRRIGATION WAS USED TO FREE IT. PROCEEDED TO A SECOND FREEZE CYCLE AND PASSIVE THAW WAS COMPLETED. DUE TO THE SIZE OF THE MASS, CRYO PROBES WERE THEN INDIVIDUALLY THAWED WITH A 1CM PULLBACK PRIOR TO THE THIRD FREEZE CYCLE AND FOLLOWED BY A SHORT THAW TO REMOVE PROBES. ONCE PROCEDURE WAS COMPLETED PT WAS TAKEN TO RECOVERY. A SHORT TIME LATER, THE PT WAS RETURNED BACK TO THE OPERATING ROOM DUE TO EXCESSIVE BLEEDING WHICH RESULTED IN THE LEFT KIDNEY BEING REMOVED. | 0
1740053 | AFTER 3 MINUTES AND 31 SECONDS INTO THE FIRST FREEZE, IT WAS NOTICED THAT THE PROBE WAS FREEZING UP THE SHAFT AND PROBE HANDLE. PROBE WAS REMOVED FROM THE FIELD AND REPLACED. NO INJURY. >< TESTING INDICATED A VACUUM INSULATION FAILURE IN THE CRYOPROBE. FROSTING OF THE SHAFT WAS CONFIRMED. NO PATIENT INJURY WAS REPORTED. | 0
1783766 | A CUSTOMER REPORTED THAT AT APPROX 8 PM, ONE ECG ELECTRODE WAS DISCONNECTED FROM A PT WHEN THEY BECAME RESTLESS. THE CARDIOCAP 5 ALLEGEDLY DID NOT ACTIVATE AN ALARM. THE PT REPORTEDLY EXPIRED. FOLLOWING THE EVENT, A GE FIELD ENGINEER TESTED THE MONITOR IN QUESTION WITH AN ECG SIMULATOR AND WAS FOUND TO FUNCTION CORRECTLY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A F/U REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1059610 | THE EVENT OCCURRED DUE TO THE FAILURE TO FOLLOW CLINICALLY PRESCRIBED PROCEDURE WHICH INCLUDES TRACING EACH INDIVIDUAL LEADWIRE FROM ITS COLORED IDENTIFIED CONNECTOR BACK TO THE ACQUISITION MODULE LABEL TO INSURE THAT IT IS MATCHED TO THE CORRECT LABEL LOCATION. INSTRUCTIONS FOR USE WERE REVIEWED FOR BOTH THE ECG DEVICES AND THE LEADWIRES. THE ECG DEVICE INSTRUCTIONS CONTAINS THE FOLLOWING WARNING STATEMENT: 'CAUTION. TRACE EACH INDIVIDUAL LEADWIRE FROM ITS COLORED ID'ED CONNECTOR BACK TO THE ACQUISITION MODULE LABEL TO INSURE THAT IT IS MATCHED TO THE CORRECT LABEL LOCATION. IMPROPER CONNECTION WILL CAUSE INACCURACIES IN THE ECG.' THE LEADWIRE INSTRUCTIONS FOR USE WILL BE MODIFIED TO INCLUDE THE SAME STATEMENT. >< THE CUSTOMER REPORTED THAT THE SOCKET CONNECTORS IN THE MAC 1200 10 LEAD TRUNK CABLE MAKE IT POSSIBLE FOR THE USER TO INCORRECTLY PUT THE CHEST AND LIMB LEADS IN THE WRONG SIDE OF THE TRUNK CABLE. NO PT INJURY WAS REPORTED. | 1
1908921 | UNABLE TO ADVANCE OR REMOVE TRANSESOPHAGEAL ECHOCARDIOGRAM PROBE AS USUAL. REQUIRED INTERVENTION. PROBE EVALUATED BY MANUFACTURER. | 0
1570737 | THERE IS A SERIOUS DESIGN FLAW WITH THE GE INNOVA 4100 IQ EQUIPMENT, ALLOWING THE PROPYLENE GLYCOL FLUID TO SHORT THE DETECTOR ELECTRONICS WHEN THE PLASTIC CONNECTORS FAIL. THIS IS AN OPERATOR AND PT SAFETY HAZARD. THE PROPYLENE GLYCOL COULD LEAK ONTO THE PT DURING THE PROCEDURE. ONE OF THE TECHS SLIPPED IN THE PUDDLE WHILE PREPPING THE TABLE. THE FLUID IS CLEAR, AND IT SHOULD BE COLORED. THE EQUIPMENT WAS INSTALLED IN 2007. THE COUPLER FAILED AND COOLANT LEAKED INTO ELECTRONICS OF A VERY EXPENSIVE PART. IT APPEARS THIS WAS DUE TO AN ERROR DURING THE ORIGINAL INSTALLATION. GE IS ESCALATING THIS ISSUE INTERNALLY WITHIN GE HEALTHCARE. GE SHOULD CHECK ALL INNOVA SYSTEMS WITHIN OUR 5 HOSPITALS' NETWORK THAT ARE UNDER CONTRACT TO ENSURE THERE ARE NO MORE ISSUES. THIS ISSUE WAS IDENTIFIED IN THE PERIPHERAL VASCULAR LABS AT TWO HOSPITALS. | 0
1705636 | CARDIOLAB SYSTEM CLOSED OUT OF THE APPLICATION AND RETURNED TO THE WINDOWS DESKTOP DURING A PROCEDURE. COULD NOT RESTART CASE. AFTER CONSULTING WITH GE TECH SUPPORT, REPLACED THE REMOVABLE DVD DISK AND STARTED A NEW CASE.TECH SUPPORT RECOMMENDED REPLACING THE REMOVABLE DVD DISK. THE ENGINEER ALSO STATED THAT THE APPLICATION WILL CLOSE OUT OF THE PROGRAM IF IT CAN NOT WRITE TO THE DVD DISK, EVEN IN THE MIDST OF A PROCEDURE. FIELD ENGINEER WAS DISPATCHED, CONFIRMED THE ENGINEERS STATEMENTS AND DOWNLOAD THE APPLICATIONS AND SYSTEMS LOGS A LONG WITH THE PATIENT STUDY. | 1
1807553 | 1) THE PATIENT WAS IN PLACE ON THE ERCP (ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY) PROCEDURE TABLE WITH THE SAFETY STRAP ON. HE HAD BEEN GIVEN A MILD SEDATIVE. HE WAS LYING PRONE AND HAD LIMB EKG LEADS ON.2) THE TABLETOP WAS POSITIONED IN THE 'TOE' DIRECTION WITH APPROXIMATELY 6 TO 8 INCHES TO MAKE A LITTLE 'SHELF' SPACE AT THE HEAD END FOR VARIOUS PROCEDURAL SUPPLIES. THE FOOT END, THEREFORE, EXTENDED PAST THE TABLE BASE BY THE SAME AMOUNT.3) THREE STAFF WERE IN THE ROOM: ANESTHESIOLOGIST, NURSE AND RESIDENT.4) THE RESIDENT THINKS THE TABLE STARTED TO TILT WHEN HE GRABBED THE HAND GRIP OF THE IMAGING CONSOLE AND STARTED TO PULL IT OVER THE PATIENT. THE ANESTHESIOLOGIST THINKS THE RESIDENT GRABBED THE CONSOLE'S HANDGRIP AND POSITIONED THE CONSOLE OVER THE PATIENT TO TAKE A BASELINE PICTURE. HE THEN LET GO AND TURNED AWAY AND THE TILTING BEGAN. THE NURSE THINKS THAT NOBODY WAS DOING/TOUCHING ANYTHING WHEN THE TILTING BEGAN. THE ANESTHESIOLOGIST THINKS THE LEVER SWITCH ON THE CONSOLE BECAME STUCK, PERHAPS NOT MECHANICALLY BUT ELECTRICALLY.5) THE NURSE REMEMBERS THAT THE TILTING WAS FAST. THE ANESTHESIOLOGIST REMEMBERS THAT THE TILTING WAS FAST, TAKING 5 TO 7 SECONDS TO REACH ITS MAXIMUM POSITION. IT DID NOT KEEP DRIVING ONCE IT REACHED ITS HIGHEST POSITION. IT CUT OFF OR WAS CUT OFF. THE CONSOLE AND TABLE ARE BOLTED TOGETHER SO BOTH TILT TOGETHER.6) THE FOOT END OF THE TABLE WENT UPWARD, THE HEAD END DOWNWARD. IT WENT ALL THE WAY (APPROXIMATELY 75-80¿) UNTIL THE FOOTBOARD STRUCK THE PLASTIC COVER (OWING TO ITS BEING EXTENDED AS DESCRIBED IN #2, ABOVE) OF THE MAVIG MONITOR BOOM CEILING MOUNT AND BROKE PART OF IT. THE PATIENT'S FEET WENT UP AND HEAD WENT DOWN. THE PATIENT'S LEGS BENT AT THE KNEE. THE ANESTHESIOLOGIST SUPPORTED THE PATIENT'S SHOULDERS AND KNEELED DOWN AS THE HEAD END WENT LOWER. THE PATIENT'S IV WAS FINE.7) MEANWHILE, ON THE FLOOR, THE BOTTOM EDGE OF THE TABLE FRAME AND THE FRONT COVER CAME IN CONTACT WITH THE CONMED ESU FOOTSWITCH. THIS BROKE A PIECE OF THE TABLE'S FRONT COVER OFF AND DEFORMED THE FOOTSWITCH ENOUGH SO THAT IT COULD NO LONGER BE USED AND HAS TO BE REPLACED.8) IT PUSHED THE MONITOR BOOM IN THE HEAD DIRECTION UNTIL THE BOOM CAME IN CONTACT WITH A SHORT PARTITION NEAR THE CEILING THAT MARKS THE SEPARATION BETWEEN THE PROCEDURE ROOM AND A SORT OF BACK ROOM. THE BOOM STRUCK THE SHEETROCK OF THE PARTITION AND MADE A DENT 8 TO 10 INCHES IN DIAMETER AND DAMAGED THE COVER OF THE SWIVEL OF THE BOOM THAT THE MONITORS HANG FROM. 9) THE MAIN EMERGENCY POWER SWITCH ON THE WALL BETWEEN THE PROCEDURE ROOM AND THE BACK ROOM WAS PUSHED, SHUTTING OFF POWER TO THE WHOLE ERCP SYSTEM.10) THE ANESTHESIA MACHINE AND PHYSIOLOGICAL MONITOR DID NOT EXPERIENCE ANY KIND OF POWER INTERRUPTION.11) THE PATIENT WAS CAREFULLY UNSTRAPPED FROM THE TABLE AND, TOGETHER WITH THE MATTRESS, WAS SLID ONTO THE FLOOR WHERE HE WAS ABLE TO STAND UP. AT LEAST FOUR AND AS MANY AS SIX PEOPLE HELPED WITH THIS. THE PATIENT HAD HIS SCHEDULED PROCEDURE IN THE OTHER ERCP ROOM.12) GE SERVICE WAS CALLED AND WAS ON-SITE IN THE AFTERNOON. THE GE SERVICE REPRESENTATIVE WAS ABLE TURN THE POWER BACK ON AND TILT THE TABLE BACK DOWN TO ITS NORMAL HORIZONTAL POSITION USING THE MANUAL POSITIONING CONTROLS ON THE TABLE'S FRONT COVER. A SECOND GE SERVICE REP ARRIVED A LITTLE LATER.FIELD SERVICE WAS ON SITE LATER THAT EVENING TO ASSESS DAMAGE. ONE REP WAS AVAILABLE THE NEXT DAY TO CONFER WITH HOSPITAL STAFF. HE ALSO ORDERED REPLACEMENT PARTS. SYSTEM WAS OPERATIONAL AGAIN THE FOLLOWING BUSINESS DAY. THEY DID NOT OFFER ANY EXPLANATION FOR HOW THIS HAD HAPPENED. | 0
1601422 | RECEIVED PATIENT AND UPON FIRST ASSESSMENT ONE HOUR LATER FOUND LIGHT ON OMNI BED TO BE UNDER THE RADIANT WARMER. THE LIGHT WAS MELTING AND BUBBLING. THE LIGHT WAS REMOVED FROM RADIANT WARMER'S PATH. THERE WAS NO INJURY TO THE PATIENT. THE LIGHT IS STILL IN WORKING ORDER. | 0
1713905 | PATIENT WAS BROUGHT TO CATH LAB FOR EMERGENT TEMPORARY PACEMAKER. MONITOR TECH INITIATED THE HORIZON CARDIOLOGY PROGRAM FOR MONITORING THE PATIENT AND DOCUMENTING THE CASE. SYSTEM FAILED TO INITIALIZE. AN ERROR MESSAGE APPEARED 'INTERNAL'. IS WAS CONTACTED, BUT THEY WERE UNABLE TO RESOLVE ISSUE. MCKESSON SUPPORT WAS ALSO CONTACTED FOR ASSISTANCE. SITUATION NOT RESOLVED UNTIL CASE COMPLETED. ALL DOCUMENTATION AND VITAL SIGNS DONE MANUALLY. NO HARM TO THE PATIENT. | 1
1929306 | THE TIP OF THE ORAL CARE DEVICE (PLASTIC AND SPONGE) BROKE OFF IN PATIENT'S MOUTH WHILE NURSE WAS BRUSHING TEETH. IT WAS DIFFICULT TO REMOVE FROM PATIENT'S MOUTH. THE TIP WAS REMOVED WITH NO ADVERSE OUTCOME.====================== HEALTH PROFESSIONAL'S IMPRESSION======================UNKNOWN | 0
1668019 | TOOTHETTE SUCTION ORAL SWAB WAS USED BY THE NURSE TO PROVIDE ORAL CARE TO AN UNRESPONSIVE PATIENT WHO HAD A TRACHEOSTOMY. THE PATIENT BIT THE TIP OFF COMPLETELY, AND SWALLOWED IT. ENT VISUALIZED THE PIECE BEHIND THE UVULA. SEDATION WAS GIVEN, AND THE PIECE WAS REMOVED USING A LARYNGOSCOPE. THERE WAS NO CHANGE IN THE PATIENT'S CONDITION RELATED TO THIS EVENT. | 0
1837899 | MCKESSON EKG MONITORING NON-FUNCTIONING ON HEMODYNAMIC SYSTEM. HAD TO PLACE THE PT ON PORTABLE MONITOR. | 1
1598487 | PATIENT WAS AN INPATIENT HAVING AN INTERVENTIONAL PROCEDURE. THE PATIENT WAS PLACED ON THE HEMODYNAMIC HEART MONITOR DURING THE CASE. THE MONITOR SHOWED WHAT APPEARED TO BE AN IRREGULAR RHYTHM WHICH WAS THOUGHT TO BE ATRIAL FIBRILLATION. THIS IS AN ABNORMAL RHYTHM THAT NEEDS TO BE EVALUATED. IN REVIEWING THE PRE WORKUP EKG, THE PATIENT WAS NOT IN AN IRREGULAR RHYTHM OR ATRIAL FIBRILLATION. PATIENT WAS ORDERED FOR A CARDIOLOGY CONSULT AND PLACED ON TELEMETRY POST PROCEDURE WHICH WAS CANCELLED AS SOON AS WE DISCOVERED THE PROBLEM WAS THE MONITORING SYSTEM AT THE END OF THE DAY. PATIENT WAS CRITICAL AND ALREADY HAD MANY CONSULTS. EXPLAINED TO PATIENT'S PHYSICIAN THE PROBLEM AND THAT THE PATIENT WAS NOT IN ATRIAL FIB. | 1
1592114 | 'IVENT' VENTILATOR BEGAN ALARMING ON PATIENT WITH READING OF 'LOW MINUTE ALARM'. IVENT WAS IMMEDIATELY SWITCHED OUT TO A PURITAN BENNETT 760 VENTILATOR. THERE WAS NO HARM TO THE PATIENT. | 0
1803659 | OUR HOSPITAL HAD USED A PRODUCT FROM GE NAMED 'MEDISORB' FOR A LONG PERIOD OF TIME, WITH NO ISSUES OR CONCERNS. THIS PRODUCT IS A CO2 ABSORBENT USED IN ALL OF OUR ANESTHESIA MACHINES. IN EARLY JULY IT WAS NOTED THAT OUR FACILITY STARTED RECEIVING A DIFFERENT PRODUCT FROM GE. INITIALLY WE RETURNED THE PRODUCT THINKING IT WAS AN ERROR IN ORDERING. WE AGAIN RECEIVED AND RETURNED THE DIFFERENT PRODUCT. UPON QUESTIONING, GE STATED THEY HAD DISCONTINUED MAKING MEDISORB AND THAT THEY HAD SENT A LETTER WHICH OUR MATERIALS MANAGEMENT DEPARTMENT DID NOT RECEIVE. WE HAVE MULTIPLE CONCERNS WITH THE REPLACEMENT PRODUCT: 1. THE LID TO THE CANISTERS CONTAINING THE SODASORB ARE NOT GROOVED IN ANY WAY, THEY JUST SIT ON THE BASE WHICH ALLOWS THE LID TO 'POP OFF' EASILY LETTING THE TOXIC PELLETS INTO THE OPERATING ROOMS. 2. THE TOP EDGE OF THE BASE PORTION OF THE CANISTER HAS CRACKED IN SEVERAL INSTANCES WHICH ALLOWS PRESSURE LEAKS AND EXCESSIVE AMOUNT OF DUST IN THE OR'S. IN THE PAST 2 WEEKS OUR ANESTHESIA MACHINES HAVE HAD TO BE REPAIRED/CLEANED TWICE DUE TO EXCESSIVE DUST. THE DUST CAUSES MACHINE LEAKS, GETS EVERYWHERE, ON THE MACHINE SEALS AND IN THE PAST IT HAS EVEN MADE IT THROUGH THE MACHINE TO THE PATIENT CIRCUIT. 3. THE REPLACEMENT PRODUCT IS NOT LABELED WITH A PRODUCT NAME. THE CANISTER JUST STATES PREPACKAGED CARTRIDGE, ORIGINAL, DISPOSABLE. THERE ALSO APPEAR TO BE 2 LOT NUMBERS ON THE PACKAGING WITH NEITHER OF THEM BEING LOCATED WHERE THE CANISTER SAYS 'LOT'. ONE IS PRINTED IN SMALL LETTERS RUNNING HORIZONTALLY ON CANISTER...THE OTHER IS AROUND THE UPPER LIP OF CANISTER. 4. IT IS FELT BY ANESTHESIOLOGY STAFF THAT THE SODASORB HEATS TO A MUCH HIGHER TEMPERATURE THAN THE ORIGINAL MEDISORB AND APPEARS TO NEED CHANGING WITH HIGHER FREQUENCY. IT IS ALSO TAKING STAFF A LARGE AMOUNT OF TIME TROUBLESHOOTING ISSUES, SAFELY REMOVING EXCESS DUST OUT OF THE PRODUCT WITHOUT COMPACTING THE GRANULES IN AN EFFORT TO MAINTAIN THE UTMOST CO2 ABSORPTION. 5. WE HAVE BEEN TOLD OTHER FACILITIES HAVE STARTED VOICING SIMILAR COMPLAINTS AND GE HAS BEEN IN CONTACT WITH THE MANUFACTURER AND IT HAS BEEN SUGGESTED THAT A MOVEMENT FOR CHANGE MAY BE STARTING. OUR FACILITY WILL NOT BE USING THIS PRODUCT. WE FEEL IT IS UNSAFE TO OUR PATIENTS, OUR STAFF AND OUR EQUIPMENT. BOTH CANISTERS ARE FROM SAME LOT NUMBER. WE ARE NOW TRIALING AN ALTERNATE PRODUCT FROM A DIFFERENT VENDOR.======================MANUFACTURER RESPONSE FOR CO2 ABSORBENT CANISTER ======================MULTIPLE EMAILS AND MEETINGS WITH PRODUCT REP. | 0

## 2011
1214830 | DURING AN INVESTIGATION OF AN INCIDENT REPORTED BY ONE OF INTELERAD'S FIELD PERSONNEL IN (B)(4), INTELERAD MEDICAL IDENTIFIED A SOFTWARE DEFECT IN THE SYSTEM THAT WOULD IN CERTAIN CIRCUMSTANCES PREVENT CERTAIN IMAGES FROM BEING DISPLAYED WITH NO WARNING OR ALARM TO NOTIFY THE USER OF A FAILURE. WHEN RETRIEVING IMAGES USING THE DICOM SERVICE, AND WHEN THESE IMAGES ARE PROCESSED IN A BATCH, IT IS POSSIBLE THAT THE INTELEVIEWER DATABASE ON THE WORKSTATION MAY NOT BE UPDATED AND SOME IMAGES MAY NOT BE RETRIEVED. DEPENDING ON HOW THE IMAGES ARE LOADED, THIS MAY RESULT IN EITHER THE IMAGES NOT BEING SHOWN AT ALL (I.E. INTELEVIEWER REPORTING THAT THERE ARE FEWER IMAGES IN A SERIES THAN WERE SENT BY THE SERVER), OR A DELAY IN SHOWING THE IMAGES BECAUSE THEY HAVE TO BE LOADED AGAIN. BASED ON A REVIEW OF THE POSSIBLE HEALTH HAZARDS WITH A LOCAL CUSTOMER, THERE WAS A CONCLUSION THAT SINCE THE PROBLEM COULD NOT ALWAYS BE DETECTED, THERE WAS A RISK OF A POSSIBLE MISDIAGNOSIS. THE INTELEVIEWER MODULE IS IN BOTH THE INTELEVIEWER WORKSTATION AND THE INTELEPACS PRODUCTS. ADDITIONAL BRAND NAME: INTELEVIEWER WORKSTATION. ADDITIONAL OTHER# V.3.5.2 PI TO P56. DEVICE MANUFACTURE DATE IS THE DATE THE INTELEVIEWER MODULE WAS RELEASED FOR INCORPORATION INTO THE INTELEPACS AND INTELEVIEWER WORKSTATION PRODUCTS IN THE FIELD. WE DO NOT "MANUFACTURE" SOFTWARE SINCE WE DEPLOY REMOTELY; THEREFORE, THE DATE OF THE SOFTWARE RELEASE TO THE CUSTOMER IS THE DATE OF MANUFACTURE. THERE WAS NO PATIENT EVENT IN THIS CASE AND NONE HAVE BEEN REPORTED TO DATE, HOWEVER, THE SOFTWARE ISSUE COULD POTENTIALLY LEAD TO AN ADVERSE EVENT WHICH IS WHY THIS MDR IS BEING SUBMITTED. THE ISSUE WAS IDENTIFIED DURING A TRAINING ACTIVITY ON SITE WHERE THE SYSTEM WAS BEING USED IN A MODE UNDER WHICH THE SOFTWARE BUG MANIFESTS ITSELF. | 1
1406002 | ASPECT REPRESENTATIVE WAS INFORMED ON (B)(6), 2009 BY AN ATTENDING ANESTHESIOLOGIST OF A CASE OF APPARENT INTRAOPERATIVE AWARENESS. THE AWARENESS CASE HAD BEEN REPORTED TO THE ANESTHESIOLOGIST BY A RESIDENT PHYSICIAN, AND HAD OCCURRED ON (B)(6), 2009. THE CASE INVOLVED A NINE YEAR-OLD CHILD WHO COMPLAINED OF BEING AWAKE, FEELING INITIAL SKIN INCISION AND DEVICE IMPLANTATION DURING A COCHLEAR IMPLANT PROCEDURE. FOLLOWING THE CASE, THE CHILD EXPERIENCED BAD DREAMS AND NIGHTMARES; THEREFORE, WAS REFERRED TO AND TREATED BY A PSYCHOLOGIST. ACCORDING TO THE ANESTHESIOLOGIST, THE CHILD IS CURRENTLY DOING VERY WELL. THE ANESTHESIOLOGIST PROVIDED A COPY OF THE AUTOMATED ANESTHESIA RECORD (WHICH INCLUDED BIS VALUES) ANNOTATED WITH THE TIME PERIODS OF REPORTED AWARENESS. THE ELECTRONIC BIS MONITOR TREND WAS ALSO REVIEWED. THE ANESTHESIA MANAGEMENT INCLUDED INHALATION INDUCTION WITH SEVOFLURANE, FOLLOWED BY ENDOTRACHEAL INTUBATION, AND MAINTENANCE WITH PROPOFOL AND REMIFENTANIL. INTRAOPERATIVE DRUG DOSING AND HEMODYNAMIC PARAMETERS WERE UNREMARKABLE AROUND THE TIME PERIODS OF AWARENESS. SIMILARLY, THE AVERAGE BIS VALUE WAS APPROXIMATELY 45 AROUND THESE TIME PERIODS. USE OF BIS MONITORING TO HELP GUIDE ANESTHETIC ADMINISTRATION MAY BE ASSOCIATED WITH THE REDUCTION OF INCIDENCE OF AWARENESS WITH RECALL IN ADULTS DURING GENERAL ANESTHESIA AND SEDATION. DEVICE LABELING STATES CLINICAL JUDGMENT SHOULD ALWAYS BE USED WHEN INTERPRETING THE BIS IN CONJUNCTION WITH OTHER AVAILABLE CLINICAL SIGNS. RELIANCE ON THE BIS ALONE FOR INTRAOPERATIVE ANESTHETIC MANAGEMENT IS NOT RECOMMENDED. THERE APPEARED TO BE NO DISCREPANCY IN THE INFORMATION PROVIDED FROM THE BIS MONITOR WITH THE OTHER AVAILABLE CLINICAL SIGNS AND ANESTHESIA MANAGEMENT. DEVICE FUNCTIONED AS INTENDED AND DID NOT MALFUNCTION. BASED ON OUR INVESTIGATION, WE HAVE CONCLUDED BIS MONITORING DID NOT CAUSE OR CONTRIBUTE TO A DEATH, AND DID NOT MALFUNCTION. OUR INVESTIGATION CONCLUDES THAT THIS INCIDENT DID NOT RESULT IN PERMANENT DAMAGE TO A BODY STRUCTURE OR PERMANENT IMPAIRMENT OF A BODY FUNCTION, DID NOT CAUSE OR CONTRIBUTE TO A DEATH, AND DID NOT MALFUNCTION. HOWEVER, THIS PATIENT WHOSE ANESTHESIA CARE INCLUDED BIS MONITORING REQUIRED TREATMENT BY A PSYCHOLOGIST, CONSIDERED TO BE A MEDICAL INTERVENTION TO PRECLUDE SERIOUS INJURY. THEREFORE, AN MDR IS REQUIRED. | 0
1409316 | THE PATIENT MONITOR MANUFACTURER RECEIVED A LETTER FROM THE CUSTOMER THAT STATED THE FOLLOWING: "THE INVIVO MONITOR MOUNTED ON THE INTERVENTIONAL TABLE PATIENT SLIDE IS MAGNETIC. IT HAS NO CAUTION OR WARNING SIGNS ON THE MONITOR ITSELF. WHEN THE CUSTOMER UNDOCKS THE TABLE, THE TABLE NEEDS TO MAKE A 90 DEGREE TURN TO EXIT THE OUTER DOUBLE DOORS. WHEN THE TABLE TURNS, IT TURNS TOWARD THE FRONT OF THE MAGNET AND THE PULL OF THE MAGNETIC FIELD FOR THE MONITOR IS FELT. IF THE CUSTOMER IS UNAWARE AND INATTENTIVE, THE TABLE COULD MOVE ON ITS OWN TOWARDS THE MAGNET." THERE WAS NO MALFUNCTION OF THE DEVICE. THE MRI SUITE WERE THE DEVICE IS LOCATED WAS STILL UNDER CONSTRUCTION ADN THE DEVICE WAS NOT IN USE AT THE TIME OF THE INCIDENT. THE USER FACILITY REPORTED THAT THE PATIENT MONITOR HAS NO CAUTION OR WARNING SIGNS ON THE MONITOR ITSELF. THE DEVICE DOES NOT HAVE WARNING LABELS ON IT, WHICH ARE NOT OBSCURED BY THE MOUNTED CONFIGURATION AND FULLY VISIBLE TO THE USER. IN ADDITION TO THE MR CONDITIONAL LABELS ON BOTH THE FRONT AND BACK OF THE DEVICE, THE DEVICE ALSO HAS A WARNING LABEL DEPICTING THE DISTANCE THE DEVICE SHOULD REMAIN AWAY FROM THE MRI MAGNET (OUTSIDE THE 2,000 GAUSS (0.2T) FIELD LINE). THESE WARNINGS ARE ALSO PRESENT IN THE INSTRUCTIONS FOR USE THAT ACCOMPANY THE DEVICE. THE PRESENCE OF THESE LABELS ON THE DEVICE AT THE USER FACILITY WAS CONFIRMED BY BOTH THE SALES REPRESENTATIVE AND THE DEVICE MANUFACTURER'S FIELD SERVICE STAFF. IT IS BELIEVED THAT THE USERS WERE EITHER MOVING THE PATIENT TABLE/DEVICE MOUNT FOR CONSTRUCTION PURPOSES OR TRYING TO MIMIC USE OF THE PATIENT TABLE/DEVICE MOUNT IN PREPARATION FOR WHEN THE MRI SUITE WAS FULLY OPERATIONAL. IN DOING SO AND BY MOVING THE PATIENT TABLE/DEVICE MOUNT, THE USERS VIOLATED THE MR CONDITIONAL PROVISIONS OF THE DEVICE AS STATED ON THE DEVICE'S LABELING. AS INDICATED ABOVE, THE MRI SUITE WHERE THE DEVICE WAS LOCATED WAS STILL UNDER CONSTRUCTION AT THE TIME OF THE INCIDENT. ADDITIONALLY, THE INCIDENT OCCURRED AFTER THE DEVICE WAS INITIALLY INSTALLED, BUT BEFORE THE USERS WERE TRAINED ON THE DEVICE. THE DATE THE USER TRAINING OCCURRED WAS DICTATED BY THE USER AND COINCIDED WITH THE FINAL STAGES OF THE CONSTRUCTION OF THE MRI SUITE. TRAINING THE USERS, WHICH OCCURRED (B)(6) 2009, COVERED GENERAL USE OF THE DEVICE AND INCLUDED DETAINED DISCUSSIONS ON THE MR CONDITIONAL ASPECTS OF THE DEVICE. ADDITIONAL FOLLOW-UP WITH THE USER WAS COMPLETED BY THE SALES REPRESENTATIVE ON (B)(4) 2009, AND THE USER INDICATED THAT THEY HAVE DONE SEVERAL CASES SINCE THE COMPLETION OF THE MRI SUITE AND HAVE NOT HAD ANY MORE INCIDENTS WITH THE DEVICE. NO ADDITIONAL ACTION IS WARRANTED OR PLANNED. | 0
1416625 | THE ISSUE INVOLVES CERNER MILLENNIUM FIRSTNET AND AFFECTS USERS THAT UTILIZE THE FILTER OPTION WITHIN THE TRACKING LIST TO LIMIT THE PATIENTS THAT ARE DISPLAYED IN THE LIST BASED ON SPECIFIC DISPLAY CRITERIA. IN CERNER MILLENNIUM FIRSTNET, WHEN THE USER SELECTS A PATIENT ON A FILTERED LIST AND THEN WAITS FOR THE SYSTEM TO AUTOMATICALLY REFRESH, A DIFFERENT PATIENT MAY ALSO BE HIGHLIGHTED AND A ROW INDICATOR DISPLAYED FOR THE NEW PATIENT IN ADDITION TO THE PATIENT, THE USER WAS ATTEMPTING TO SELECT. PATIENT CARE COULD BE ADVERSELY AFFECTED, AS CLINICIANS COULD INADVERTENTLY SELECT THE INCORRECT PATIENT AND PERFORM DOCUMENTATION ON THAT PATIENT RATHER THAN THE INTENDED PATIENT. THIS ISSUE COULD RESULT IN MEDICATION ORDERS AND CLINICAL DOCUMENTATION BEING ASSOCIATED WITH THE INCORRECT PATIENT. CERNER DISTRIBUTED PRIORITY REVIEW FLASH NOTIFICATION (B)(4) ON (B)(4), 2008, TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE FLASH INCLUDES A DESCRIPTION OF THE ISSUE AND NOTIFIES CLIENTS OF THE SOFTWARE CORRECTIONS AVAILABLE TO ADDRESS THE ISSUE (EXCEPTION PACKAGES (B)(4)). ADDITIONALLY, THE FLASH NOTIFIES CLIENTS THAT THEY CAN REMOVE FILTERS FROM TRACKING LISTS TO PREVENT THE ISSUE FROM OCCURRING UNTIL THE SOFTWARE CORRECTION CAN BE INSTALLED IN THEIR PRODUCTION ENVIRONMENT. CERNER CORPORATION CONSIDERS THIS ISSUE TO BE RESOLVED AND NO FURTHER NARRATIVE IS REQUIRED FOR FOLLOW-UP. | 1
1416630 | A CLIENT SITE REPORTED AN EVENT ON (B)(6) 2009 WHICH WAS CAUSED BY A SEQUENCE OF OPERATOR ERRORS. THE MEDICAL DEVICE WYNDGATE'S SAFETRACE TX REBRANDED AS HORIZON BLOOD BANK DID NOT MALFUNCTION. THE SCENARIO AT THE REPORTING SITE INVOLVED TWO PATIENTS (PATIENT A AND PATIENT B FOR DESCRIPTIVE REASONS) AND THEIR ASSOCIATED BLOOD SPECIMENS. AS A CONSEQUENCE OF SEQUENTIAL USER ERRORS IN ENTERING THE PATIENT'S SPECIMEN ACCESSION NUMBER, THE TESTING RESULTS FROM AN INCORRECTLY IDENTIFIED SPECIMEN WERE LINKED TO THE WRONG PATIENT. INFORMATIONAL BULLETINS HAVE BEEN DISTRIBUTED BY WYNDGATE TO ALL SITES THAT CURRENTLY USING EITHER HORIZON BLOOD BANK OR SAFETRACE TX PRODUCTS. ADDITIONAL MANUFACTURER NARRATIVE IS INCLUDED WITH THIS REPORT AS AN ATTACHMENT. CONCLUSION - THE MEDICAL DEVICE PERFORMED AS INTENDED AND DID NOT MALFUNCTION. THE REPORTED INCIDENT OCCURRED SOLELY BECAUSE OF A UNIQUE SET OF SEQUENTIAL OPERATOR ERRORS. THIS SCENARIO IS UNIQUE BECAUSE THE OPERATOR ENTERED THE SAME ERRONEOUS SPECIMEN NUMBER ON TWO CRITICAL WINDOWS; ON THE RESULT ENTRY WINDOW, IF THE OPERATOR HAD ENTERED ANY SPECIMEN NUMBER OTHER THAN THE SAME ONE THAT WAS INCORRECTLY LINKED PREVIOUSLY, THE MEDICAL DEVICE WOULD HAVE DETECTED THE OPERATOR ERROR. WYNDGATE RECOMMENDS THAT THE SITE RETRAIN ITS EMPLOYEES ON THE PROPER USE OF WYNDGATE'S SOFTWARE. IT SHOULD ALSO IMPLEMENT ADDITIONAL OPERATOR CONTROLS ON ITS PROCESS LINKING THE SPECIMEN TO THE PATIENT. THE MEDICAL DEVICE IS DEPENDENT ON THE OPERATOR CONSISTENTLY AND ACCURATELY ENTERING CORRECT DATA FOR THE PROPER FUNCTIONING OF THIS SYSTEM. THE SOFTWARE CAN ONLY REACT TO THE DATA THAT IS ENTERED. IN THIS CASE, WHERE THE WRONG DATA WAS ENTERED TWICE, HHB HAD NO WAY TO CONTROL THE OUTCOME. THE LINKAGE OF THE SPECIMEN TO THE PATIENT IS A CRITICAL ENTRY POINT WHICH FURTHER DEVICE PROCESSING OCCURS. A REMINDER WILL BE SENT TO ALL CLIENTS TO REITERATE THE IMPORTANCE OF PROPER SPECIMEN LINKAGE AND DATA ENTRY. ACTION TAKEN - AS A PRECAUTION, WYNDGATE SENT A PRODUCT BULLETIN TO ALL OF ITS CLIENTS USING THE HBB SYSTEM REITERATING THE IMPORTANCE OF PROPER SPECIMEN LINKAGE AND DATA ENTRY. THIS SAME NOTICE WILL ALSO EMPHASIZE THE NEED TO THOROUGHLY TRAIN EMPLOYEES ON THE PROPER USE OF THE HBB SYSTEM. | 0
1418516 | NO REPORT OF PT HARM OR INJURY RELATED TO THIS EVENT. CUSTOMER REPORTED THAT MEASUREMENT IS INCONSISTENT WITH MEASUREMENT PERFORMED ON MODALITY. ISITE PACS IS A SOFTWARE PRODUCT. WE ARE ABLE TO EVALUATE THE PRODUCT AT THE CUSTOMER SITE BY REMOTE ACCESS, AS WELL AS EVALUATE THE SAME PRODUCT VERSION IN-HOUSE. PER INVESTIGATION AT TIME OF EVENT: ISITE PACS VERSION 3.3 TAKES CONSISTENT APPROACH OF USING IMAGER PIXEL SPACING FOR MEASUREMENT PERFORMED ON PROJECTION RADIOGRAPHY IMAGES. IMAGER PIXEL SPACING IS REQUIRED TO BE PRESENT IN PROJECTION RADIOGRAPHY IMAGES. THIRD PARTY MODALITY IS USING AN OPTIMAL DICOM TAG WHICH MAY OR MAY NOT BE PRESENT IN A GIVEN IMAGE. THE OPTIONAL DICOM TAG IN USE BY THE THIRD PARTY MODALITY ESTIMATES CONVERSION FROM IMAGER PIXEL SPACING INTO THE PT PIXEL SPACING. THE VALUE OF ESTIMATED RADIOGRAPHIC MAGNIFICATION FACTOR DOES NOT NECESSARILY RELATE TO THE PLANE IN WHICH MEASUREMENT IS PERFORMED. THE VALUE OF ESTIMATED RADIOGRAPHIC MAGNIFICATION FACTOR MAY OR MAY NOT BE PRESENT, DEPENDENT ON MODALITY CONFIGURATION. THE ONLY ACCURATE WAY TO GET PT MEASUREMENTS ON PROJECTION IMAGING IS TO USE FIDUCIAL MARKERS AND CALIBRATE MEASUREMENTS. THIS CAPABILITY IS SUPPORTED IN ISITE PACS. THE 3.3 IFU INCLUDES WARNINGS OF MODALITY DEPENDENT MEASUREMENTS AND THE STRONG RECOMMENDATION TO USE FIDUCIAL MARKERS FOR DETERMINING MAGNIFICATION FACTORS. | 1
1455031 | A CUSTOMER REPORTED A LOSS OF MONITORING FOR MULTIPLE PTS THROUGHOUT THE HOSPITAL NETWORK. ATS TOWERS WERE DOWN, TELEMETRY WAS DROPPING OUT, AND BESIDE MONITORS WERE NOT COMMUNICATING. TOTAL NETWORK DOWNTIME UNTIL THE ISSUE WAS COMPLETELY RESOLVED WAS 10 HOURS FROM THE TIME OF REPORT. NO PT INJURY OR DEATH REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1474593 | REPORTED OBSERVATION: PATIENT DEVELOPED (B)(6) IN ONE EYE AFTER INSERTION OF AN OASIS MEDICAL FORM FIT HYDROGEL CANALICULAR PLUG. (B)(4). PRODUCT LOT NUMBER: LH0209B. PRODUCT INSERTED ON: (B)(6) 2009. DATE OF COMPLICATION: (B)(6) 2009. REPORTED TO (B)(4) ON (B)(6) 2009. DEVICE WAS NOT RETURNED FOR EVALUATION. A REVIEW OF BATCH PRODUCTION AND STERILIZATION RECORDS SHOWS THAT THERE WERE NO REMARKABLE EVENTS, ANOMALIES OR DEVIATIONS DURING PRODUCTION OF FORM FIT HYDROGEL CANALICULAR PLUG, (B)(4), PRODUCT LOT NUMBER LH0209B. A TOTAL OF (B)(4) FORM FIT HYDROGEL CANALICULAR PLUGS, (B)(4), PRODUCT LOT NUMBER LH0209B WERE RELEASED AND DISTRIBUTED. THERE HAVE BEEN NO OTHER ADVERSE EVENTS REPORTED BY ANY OTHER USER/CUSTOMER INVOLVING THIS LOT OF PRODUCT. THE PLUG WAS IRRIGATED FROM THE CANALICULAR, THE PATIENT WAS PRESCRIBED ANTIBIOTIC DROPS. THE ISSUE WAS RESOLVED FOR THE PATIENT, NO FURTHER TREATMENT WAS REQUIRED. (B)(4) CANNOT DEFINITIVELY DETERMINE THAT THE FORM FIT HYDROGEL CANALICULAR PLUG, (B)(4), FROM PRODUCT LOT NUMBER LH0409F WAS DIRECTLY OR INDIRECTLY RESPONSIBLE FOR THE REPORTED OBSERVATION OF (B)(6). | 0
1489953 | USERS WITH WRITE/UPDATE RIGHTS TO NETWORK FOLDERS ARE ABLE TO ACCESS ANATOMIC PATHOLOGY RESULTS AND POTENTIALLY EDIT/DELETE THEM. PRODUCTS AND RELEASES AFFECTED: SUNRISE LABORATORY VERSION 648C, 650 AND 3.0. BACKGROUND: CURRENTLY, ANATOMIC PATHOLOGY RESULTS ALONG WITH LAB DATA/TEMPLATES ARE STORED IN WORD FILES OUTSIDE OF THE (B)(6) LABORATORY DATABASE IN A NETWORK DIRECTORY/FOLDER. READ/WRITE PERMISSIONS ARE USUALLY GRANTED TO ALL (B)(6) LABORATORY USERS SO THEY CAN ACCESS THESE TEMPLATES AND OTHER REQUIRED FOLDERS. HOWEVER, USERS WITH WRITE/UPDATE RIGHTS FOR THIS NETWORK DIRECTORY/FOLDER AND KNOWLEDGE OF THIS FOLDER LOCATION CAN ACCESS ANY FILE IN THIS DIRECTORY AND POTENTIALLY EDIT/DELETE THESE RESULT FILES. THE DEVICE IS SOFTWARE AND CANNOT BE PHYSICALLY RETURNED. WE USED OUR COPY OF THE SOFTWARE FOR TESTING. WORKAROUND: IF ANY MANUAL EDITS OUTSIDE OF THE SYSTEM DO OCCUR, THEY CAN BE TRACKED BY STANDARD OPERATING SYSTEM DIRECTORY FEATURES LIKE "DATE CREATED", "DATE MODIFIED", "AUTHOR". RESOLUTION: CUSTOMERS WERE NOTIFIED BY PRODUCT ADVISORY COMMUNICATION. THIS ISSUE IS BEING ADDRESSED IN SERVICE PACK 2 FOR VERSION 3.0 SCHEDULED FOR RELEASE ON (B)(6) 2010. CUSTOMERS WILL BE ADVISED TO IMPLEMENT THE SERVICE PACK. | 1
1494572 | A ROUTINE RENAL CRYO ABLATION PROCEDURE WAS SUCCESSFULLY PERFORMED AND APPROX 3-4 HOURS POST PROCEDURE, THE ATTENDING (B)(6) TECHNOLOGIST WAS NOTIFIED BY THE HOSPITAL (B)(6), THAT THE PT HAD EXPERIENCED A STROKE AND THAT THEY HAD PULLED THE USED CRYOPROBE FROM THEIR SHARPS CONTAINER. (B)(4) CONTACTED THE HOSPITAL AND PROVIDED (B)(6) INFO ON THE EQUIPMENT USED AND ASKED IF ANY FURTHER INFO WAS REQUIRED. NONE WAS REQUESTED AND NO FURTHER CONTACT HAS BEEN INITIATED BY THE HOSPITAL. AN INCIDENT REPORT FROM THE CRYO TECH PRESENT DURING THE PROCEDURE: "WE USED ONE R.3 8L RENAL PROBE FOR THE CASE. LOT# 08-0371 THE START TIME WAS 9:23 AM. THE DOCTOR WANTED TO USE A R.3 8L PROBE SO THE CIRCULATOR OPENED IT UP. I HAD THE SCRUB TECH (B)(6) PLACE IT INTO A BASIN OF WATER AND WE USED THE PRETEST MODE ON THE CS. THE DOCTOR, ASSISTANT, AND THE SCRUB TECH ALL WATCHED THE PRETEST. EVERYTHING WAS FINE. NO AIR BUBBLES APPEARED AND THE PROBE APPEARED TO BE WORKING CORRECTLY. THE TEST RAN FOR 30 SECONDS THAW THEN 30 SECONDS FREEZE, AND THEN ANOTHER 30 SECOND THAW. THE DOCTOR PLACED THE PROBE IN THE TUMOR (VIA LAPAROSCOPIC TECHNIQUE) AND WE DID A 10 MINUTE FREEZE FOLLOWED BY A PASSIVE THAW. THE DOCTOR THEN HAD ME DO ANOTHER 10 MINUTE FREEZE FOLLOWED THIS TIME BY AN ACTIVE THAW (ABOUT 2 MINUTES). THEY REMOVED THE RENAL PROBE AND HANDED IT OFF TO ME. I UNPLUGGED THE CS (CRYOSURGICAL CS SYSTEM) AND CUT THE LINE. I THEN PLACED IT INTO THE SHARPS CONTAINER. I PUT DOWN MY STOP WATCH AT 10:50. THIS IS AFTER I HAD MOST OF MY EQUIPMENT OUT OF THE ROOM AND IT LOOKED LIKE THEY WERE GETTING READY TO TRANSPORT THE PT. I RECEIVED A CALL FROM THE HOSPITAL (B)(6) AT APPROX 2:00 PM. SHE SAID THAT THE PT HAD WHAT APPEARED TO BE AN AIR EMBOLISM AND HAD A STROKE. SHE WANTED TO GIVE ME A HEADS UP. SHE SAID THAT THEY GOT THE RENAL PROBE FROM THE SHARPS CONTAINER AND THEY WERE GOING TO HOLD ONTO IT IN CASE IT NEEDED TO BE TESTED. AFTER I SPOKE TO (B)(6) I CALLED (B)(4) IMMEDIATELY." F/U WITH THE ABOVE REPORTING TECH ON (B)(6) 2009 YIELDED NO NEW INFO OR PT STATUS. | 0
1544220 | A THIRD PARTY (NON AYCAN) DICOM-MODALITY-WORKLIST-SERVER SOLUTION, DESIGNED TO GENERATE STUDY-INSTANCE-UID'S IN CONNECTION WITH PT DATA FOR VARIOUS MODALITIES WAS REINSTALLED BY A THIRD PARTY COMPANY AT ONE OF AYCAN'S CUSTOMER LOCATIONS. AFTER A SYSTEM MANIPULATION BY A TECHNICIAN OF THIS THIRD PARTY COMPANY, THE SYSTEM ABOVE STARTED TO ALLOCATE STUDY-INSTANCE-UID'S, WHICH HAVE ALREADY BEEN USED FOR PRIOR PT DATA INFO. THIS MALFUNCTION CAUSED THE SYSTEM TO GENERATE IMAGE DATA SETS WITH NOT UNIQUE STUDY INSTANCE UIDS. THESE INCORRECT DATA SETS WERE SENT, FOR DIAGNOSIS, TO AYCAN'S WORKSTATION OSIRIX PRO AND OTHER WORKSTATIONS. AS A RESULT OF THIS THIRD PARTY COMPANY'S SYSTEM MALFUNCTION INCORRECTLY IDENTIFIED IMAGE DATA WAS DISPLAYED WITH INCORRECT PT DATA. DUE TO THE FACT THAT THE CUSTOMER IMMEDIATELY NOTICED THE MALFUNCTION OF THE SYSTEM NO PT OR OTHERS WERE HARMED IN ANY WAY. BECAUSE THE ABOVE DESCRIBED OCCURRENCE IS NOT CAUSED BY OUR SYSTEM, WE BELIEVE THIS MIGHT NOT BE AN OCCURRENCE ACCORDING 21 CFR 803. BUT BECAUSE THIS MALFUNCTION OF THIS THIRD PARTY SOLUTION IS CAUSING OUR SYSTEM TO DISPLAY INCORRECT INFO WE DECIDED TO REPORT THIS IN ORDER TO REDUCE THE RISK OF MISDIAGNOSIS. AS THIS ISSUE ONLY HAPPENS IN COMBINATION OF THE FOLLOWING SITUATIONS WE SEE THE RISK OF REOCCURRENCE TO BE MARGINAL. CONFIG MISTAKE OF THIRD PARTY COMPANY'S TECH. THIRD PARTY SOFTWARE DOESN'T DETECT CONFIGURATION ERRORS. MODALITY DOES NOT DETECT DOUBLE UIDS. WORKSTATION WHERE NEW DATA IS BEING DISPLAYED (WITH DOUBLE UIDS) HAS "OLD" DATA (WHERE UIDS WERE ALLOCATED THE FIRST TIME) STILL AVAILABLE. PLANNED CORRECTIVE ACTION: SECURITY ALERT NOTIFICATION AT ALL CUSTOMERS AND RESELLERS (COPY TO MEDWATCH). SECURITY ALERT NOTIFICATION TO LEADERSHIP TEAM OF THE FREE OPEN-SOURCE PROJECT (B)(4) (WHICH "AYCAN'S (B)(4) PRO" IS BUILT ON). SECURITY ALERT NOTIFICATION TO THE ENTIRE USER GROUP (MAILING LIST) OF THE OPEN SOURCE SOFTWARE. SOFTWARE UPDATE TO HELP FURTHER DISCOVER INCORRECTLY IDENTIFIED DATA SENT TO WORKSTATION (AS MUCH POSSIBLE WITH AVAILABLE DATA). INDICATION: AYCAN IS NOT THE MFR/ SUPPLIER OF THE FREE AVAILABLE OPEN-SOURCE SOFTWARE (B)(4). | 1
1619770 | CUSTOMER REPORTS THE IMAGE IS DISPLAYED WITH A 180 DEGREE HORIZONTAL FLIP ON WEB-DX. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1628620 | A (B)(6) MALE, UNDERGOING A CYSTOSCOPY, LEFT URETEROSCOPY AND LEFT URETERAL STENT EXCHANGE WHEN THE UROLOGIST NOTICED A SMALL CRACK IN THE EXTERNAL CASING OF THE GYRUS FLEXIBLE DIGITAL URETEROSCOPE AND THERE WAS SOME EXPOSED METAL (POSSIBLY FIBER OPTIC FIBERS) NOTED AS WELL. THE UROLOGIST EXAMINED THE PT'S URETHRA AND BLADDER, AS THIS IS WHERE THE CRACK WOULD HAVE COME IN CONTACT WITH TISSUE, THERE WAS NO EVIDENCE OF INJURY. "MILD ERYTHEMA" WAS NOTED IN THE MUCOSA OF THE URETHRA, BUT THE UROLOGIST FELT HIS WAS NOT ATYPICAL OF WHERE A PASSAGE OF A RIGID SCOPE WOULD HAVE BEEN. NO INJURY TO THE PT NOTED. SCOPE WAS SEQUESTERED AND SENT TO SPC FOR CLEANING. PT WAS SENT TO PACU FOR RECOVERY AND WAS DISCHARGED LATER THAT SAME DAY. FINDINGS OF THE EVAL - IMAGE IS GOOD, MAJOR LOSS OF DEFLECTION DUE TO TWIST WITH CUT IN THE SHAFT FROM 43.2CM TO 43.6CM AND AT 63.8CM WITH SHAFT WIRES PROTRUDING THROUGH THE COVER. FAILS LEAK TEST DUE TO THE TWIST AND LASER BURN AT DISTAL FORCEPS TUBE. THE TWIST OCCURRED DUE TO EXCESSIVE FORCE APPLIED TO THE SHAFT AND HOW THE INSTRUMENT WAS BEING MANIPULATED BY THE USER. | 0
1629900 | THE CUSTOMER REPORTED A BROKEN ELECTRICAL CONNECTION CABLE OF THE CELL-DYN ANALYZER WASTE SENSOR. THE PART IS TO BE REPLACED IN ORDER TO RESOLVE THE ISSUE. NO WASTE EXPOSURE WAS REPORTED. NO IMPACT TO USER SAFETY OR TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1632390 | THE ACCOUNT STATED THAT THE ARCHITECT ANALYZER HAD BEEN GENERATING 1007 ERRORS AND IT WAS DETERMINED THAT THE PRE-TRIGGER SOLUTION WAS NOT BEING DISPENSED. THE SAMPLES THAT WERE RUN AT THAT TIME WERE REPEATED AND DISCREPANCIES WERE DISCOVERED. AN INITIAL ARCHITECT B12 RESULT OF 538 PG/ML WAS REPORTED. UPON REPEAT, THE RESULT WAS 147 PG/ML. A CORRECTED REPORT WAS SUBMITTED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1632393 | THE ACCOUNT STATED THAT THE ARCHITECT ANALYZER HAD BEEN GENERATING 1007 ERRORS AND IT WAS DETERMINED THAT THE PRE-TRIGGER SOLUTION WAS NOT BEING DISPENSED. THE SAMPLES THAT WERE RUN AT THAT TIME WERE REPEATED AND DISCREPANCIES WERE DISCOVERED. AN INITIAL ARCHITECT TSH RESULT OF 0.34 UIU/ML WAS REPORTED. UPON REPEAT, THE RESULT WAS 1.67 UIU/ML. A CORRECTED REPORT WAS SUBMITTED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1639024 | AS REPORTED: TRANSDUCER INTERMITTENTLY LOSING WAVEFORM WHILST ON PTS. AS A RESULT PTS HAVE BEEN TREATED ACCORDINGLY AND THEREFORE TREATED INCORRECTLY. | 0
1650326 | IT WAS REPORTED THAT WHEN A USER INITIATED A TELEMETRY PATIENT TRANSFER FROM A BED ON ONE ICS TO A BED ON ANOTHER ICS, THE ICS FROM WHICH THE PATIENT WAS BEING TRANSFERRED BLANKED OUT TO A BLANK GREEN SCREEN AND APPEARED TO REBOOT. AFTER THE APPARENT REBOOT, NORMAL MONITORING FUNCTION WAS REPORTEDLY RESTORED. THERE WAS NO PATIENT INJURY REPORTED. THE ICS CENTRAL STATION HAS REPORTEDLY REMAINED IN USE. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1659201 | THE CUSTOMER STATED A PATIENT SAMPLE GENERATED A HEMOGLOBIN RESULT OF 13 G/DL ON THE CELL-DYN 3200 ANALYZER. THE SAMPLE WAS RETESTED, YIELDING A HEMOGLOBIN RESULT OF 4 G/DL. AFTER PERFORMING CLEANING PROCEDURES, THE CUSTOMER OBTAINED AN ACCUMULATOR WET ERROR AND NOTED THAT WASTE CHAMBER #1 AND THE VENT ACCUMULATOR WERE NOT DRAINING. FIELD SERVICE WAS DISPATCHED TO INSPECT AND SERVICE THE INSTRUMENT. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
1664837 | THE CUSTOMER UPDATED THE GAINS SETTINGS ON A CELL-DYN EMERALD ANALYZER PER THE FA23MAR2010 CUSTOMER COMMUNICATION. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
1664982 | THE CUSTOMER STATED THAT THE CELL-DYN SAPPHIRE ANALYZER GENERATED ERROR MESSAGES TO FLAG HEMATOLOGY QUALITY CONTROLS AND PATIENT RESULTS WHICH WERE OUTSIDE THE ANALYZER MEASUREMENT LIMITS. THE FIELD SERVICE REPRESENTATIVE REPLACED PARTS OF THE ASPIRATION BOTTOM SENSOR IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED.><(B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1664983 | ABBOTT FIELD SERVICE REPRESENTATIVE REPLACED A WORN-OUT DILUENT SYRINGE BELT AND THE ASPIRATION BOTTOM SENSOR OF THE CELL-DYN SAPPHIRE ANALYZER. THE RELATION OF THE HEMOGLOBIN/HEMATOCRIT WAS IMPACTED BY THE WORN-OUT PARTS. NO IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT WAS REPORTED. (B)(4). A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED.><(B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1665006 | THE CUSTOMER STATED THAT THE CELL-DYN SAPPHIRE ANALYZER HAS BEEN GENERATING LOWER THAN EXPECTED PATIENT HEMOGLOBIN RESULTS AND THEREFORE MISMATCHED HEMOGLOBIN/HEMATOCRIT RATIO. NO SPECIFIC PATIENT DATA WAS PROVIDED. QUALITY CONTROLS WERE ALSO OUT OF RANGE LOW. TROUBLE SHOOTING AND MAINTENANCE PERFORMED BY THE CUSTOMER DID NOT RESOLVE THE ISSUE. THE ISSUE WAS RESOLVED BY THE FIELD SERVICE REPLACING THE ASPIRATION BOTTOM SENSOR DUE TO BROKEN WIRE THAT WAS WORN-OUT FROM NORMAL USE. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES.><(B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1665058 | THE CUSTOMER STATED THAT THE ASPIRATION PROBE OF THE CELL-DYN SAPPHIRE ANALYZER IS NOT MOVING CORRECTLY. THE FIELD SERVICE REPRESENTATIVE REPLACED BROKEN NEEDLES AND OTHER PARTS OF THE PROBE DETECTOR ASSEMBLY IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT WAS REPORTED. (B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES.><(B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1666963 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT HARM. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1672111 | A SPACELABS CUSTOMER REPORTED THAT WHEN REMOTE VIEWING A TELEMETRY PATIENT AT A BEDSIDE MONITOR, THE ECG WAVEFORM AND HEART RATE INFORMATION DISPLAYED WAS DIFFERENT THAN THAT DISPLAYED AT THE CENTRAL STATION. THE HOSPITAL BIOMED REPORTED THAT ALL WAVEFORMS, NUMERIC AND PRINTED INFORMATION AT THE CENTRAL STATION WAS ACCURATE. ALL INFORMATION IN CLINICAL ACCESS, LONG TERM STORAGE / FULL DISCLOSER SYSTEM, WAS ACCURATE. THE DISCREPANCY WAS ONLY OBSERVED AT THE REMOTE VIEW, BEDSIDE MONITOR. THE HOSPITAL STAFF DISABLED THE REMOTE VIEW FUNCTION, AFTER THAT, THE ANOMALY WAS NO LONGER OBSERVED AND IT WAS NOT REPEATABLE, EVENT AFTER RE-ESTABLISHING A REMOTE VIEW. SPACELABS AND THE HOSPITAL HAVE BEEN UNABLE TO DUPLICATE THESE SYMPTOMS AND ARE CONTINUING THE EVALUATION OF THE DEVICE AND THE TELEMETRY ENVIRONMENT. ADDITIONAL INFORMATION WILL BE PROVIDED AS SOON AS IT BECOMES AVAILABLE. | 1
1675207 | A LOSS OF TELEMETRY MONITORING WAS REPORTED FOR 12 PTS IN THE MEDICAL-SURGICAL UNIT. NO ADVERSE PT OUTCOME WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1675757 | UPON START UP OF THE ARCHITECT I2000K ANALYZER, THE OPERATOR TRIED TO IMPROVE THE CONNECTION OF THE CABLE WITH HIS FEET, CAUSING THE WASTE TO BE OVERTURNED, AND AS A RESULT, WASTE SPILLED ONTO THE ENTIRE ELECTRICAL PART OF THE PUMP, BURNING IT. FIRE IN THE EXTERNAL WASTE PUMP WAS PRESENT. THERE WAS NO DEATH OR INJURY TO PATIENT OR OPERATOR DIRECTLY OR INDIRECTLY. THERE WAS NO DAMAGE TO THE FACILITY. THE FIRE WAS CAUSED BY USER ERROR. (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1689402 | IT WAS REPORTED THAT THE CARTO 3 SW VERSION 1.1 WAS HAVING PACING ISSUES. THE USER WOULD ENABLE A CATHETER TO PACE, IT WOULD DISAPPEAR ON ACL, AND THEN IF THE USER SWITCHED POLES FOR PACING (DELAY OF PACING RE-ROUTING); IT WOULD DELAY THE CHANGES AND THEN PACED AT A HIGH OUTPUT. THE HIGH OUTPUT CAUSED A PT TO GO INTO ATRIAL FIBRILLATION (AFIB). ATRIAL FIBRILLATION OCCURRED ALMOST IMMEDIATELY AFTER THE PACING. THE PACING WAS THEN TERMINATED UPON RECOGNITION OF THE RHYTHM. IT WAS STATED THAT THE PT WENT INTO AN ARRHYTHMIA (AFIB) AND SELF-TERMINATED WITHIN 5-10 MINUTES AFTER INITIATION. THERE WAS NO PROCEDURE DONE TO TERMINATE THE ARRHYTHMIA; THE EVENT WAS SELF-LIMITING. ACCORDING TO THE HOSP, THE PT WAS UNDERGOING GENERAL ANESTHESIA AT THE TIME OF THE EVENT. THERE WAS NO SIGNIFICANT CHANGES IN VITAL SIGNS OR LEVEL OF CONSCIOUSNESS WERE REPORTED DURING OR AT THE END OF THE ABLATION PROCEDURE. THE PT WAS FULLY RECOVERED AND SATISFIED WITH THE PROGNOSIS. (B)(4). | 0
1693471 | THE CUSTOMER REPORTED THAT WHILE THE PANORAMA CENTRAL STATION WAS IN USE MONITORING PTS ALONG WITH BEDSIDE MONITORS AND AMBULATORY TELEPACKS, THE CENTRAL STATION DISPLAY WENT OUT. NO PT INJURY WAS REPORTED. THE CUSTOMER'S BIOMED REPLACED THE DISPLAY WITH A GENERIC LCD DISPLAY. WHEN THE COMPANY SERVICE REP ARRIVED AT THE SITE, HE REPLACED THE GENERIC DISPLAY WITH A FACTORY APPROVED FROM THE SERVICE SPARE EQUIPMENT POOL. HE RETURNED THE ORIGINAL DISPLAY TO THE FACTORY FOR REPAIR, WHERE THE POWER BOARD AND TOUCHSCREEN WERE REPLACED. THE DISPLAY WAS RETURNED TO THE CUSTOMER AND INSTALLED BY THE SERVICE REP. THE SYSTEM WAS TESTED TO FACTORY SPECIFICATIONS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. (B)(4). | 1
1693477 | IT WAS REPORTED THAT THE APL (ADJUSTABLE PRESSURE LIMITING) VALVE LOCKED UP AT THE END OF A CASE. THE CAREGIVERS ELECTED TO EMERGE THE PT FROM ANESTHESIA ON AN AMBU BAG. THE PT REQUIRED REINTUBATION DUE TO CO2 RETENTION. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
1695186 | THE CUSTOMER STATED THE QUALITY CONTROL WAS OUT OF RANGE ON THE CELL-DYN 1800 ANALYZER. WHILE SERVICING THE ANALYZER, A FIELD SERVICE REPRESENTATIVE FOUND THE DILUENT SYRINGE WAS LEAKING AND REPLACED IT WITH THE CORRECT REPLACEMENT SYRINGE. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). NEW DIE-CAST DRIVES WERE DESIGNED AND RELEASED FOR USAGE ON THE CELL-DYN 1800 AND CELL-DYN 1700 INSTRUMENTS IN THE YEAR 2005. THE 10 ML DILUENT SYRINGE, LIST NUMBER 04H36-01, WAS INCORRECTLY LISTED AS THE REPLACEMENT SYRINGE WITH THE NEW DIE-CAST DRIVES. CUSTOMERS WERE UNABLE TO ORDER THE PROPER SYRINGE WITH THE THICKER SPACER TO PREVENT THE TOLERANCE STACK-UP ISSUE AS A LIST NUMBER HAD NOT ESTABLISHED FOR THE SYRINGE. EFFECTIVE 29 MAY 2009, THE 10 ML DILUENT SYRINGE DRIVE WITH THE 0.25-INCH THICK WASHER WAS RELEASED UNDER LIST NUMBER 09H35-01 FOR DISTRIBUTION AS A DROP-IN REPLACEMENT FOR THE SYRINGE AND A PRODUCT INFORMATION LETTER WAS ISSUED TO INFORM NEW CUSTOMERS OF THE CORRECT PART NUMBER TO ORDER. A FIELD COMMUNICATION WAS ISSUED ON 08 JUNE 2009 TO ALL AFFECTED CUSTOMER TO PROVIDE INFORMATION REGARDING THE ISSUE AND THE CORRECT SYRINGE PART NUMBER TO ORDER. THE INVESTIGATION CONCLUDED THAT THE ISSUE OCCURRED DUE TO AN OPEN LOOP PROCESS, WHICH DID NOT INCLUDE ADEQUATE CHECKS AND BALANCES TO ENSURE ITEMS IDENTIFIED ON A CHANGE IMPACT ASSESSMENT FORM WERE ADDRESSED VIA AN ENGINEERING CHANGE ORDER. THE CURRENT PROCESS HAS BEEN REVISED TO PREVENT THIS ISSUE FROM RECURRING | 0
1695875 | ABBOTT FIELD SERVICE REPRESENTATIVE REPLACED THE CELL-DYN SAPPHIRE BOTTOM SENSOR PER FA05APR2010 WHILE PERFORMING A PERIODIC MAINTENANCE FOR THE CUSTOMER. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES.><(B)(4). A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1695923 | ABBOTT CUSTOMER SUPPORT ASSISTED THE CUSTOMER IN UPDATING THE GAINS SETTINGS ON A CELL-DYN EMERALD ANALYZER PER THE FA23MAR2010 CUSTOMER COMMUNICATION. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
1697816 | A SOFTWARE DEFECT IN THE PRELUDE MODULE OF SYMPHONY DOES NOT ALLOW THE ASPIRIN FLAG OF THE ASPIRIN QUESTION ON THE DONOR QUESTIONNAIRE TO CLONE PROPERLY. WHEN THE ASPIRIN QUESTION IS CLONED, A SCREEN WILL DISPLAY THE CLONED INFORMATION AND THE ASPIRIN FLAG IS RESET TO 'NO' INSTEAD OF COPYING THE ORIGINAL VALUE. EVALUATION SUMMARY: ONE CUSTOMER COMPLAINT WAS RECEIVED ON (B)(6) 2010, FROM (B)(6), NOTING THAT AN UNEXPECTED CHANGE WAS MADE ON THE DONOR QUESTIONNAIRE THAT WAS FOUND IN THE PRODUCTION ENVIRONMENT. HAEMONETICS CONDUCTED AN INVESTIGATION AND IT WAS DETERMINED THAT WHEN PERFORMING A CHANGE IN THE QUESTIONNAIRE MODULE BY HAEMONETICS STAFF RESULTED IN AN ERROR ON THE QUESTIONNAIRE IN PRELUDE. MORE SPECIFICALLY, THE ASPIRIN QUESTION WAS CLONED FOR THE CUSTOMER AND THE ASPIRIN FLAG PROPERTIES OF THAT QUESTION DID NOT COPY OVER. THE VERIFICATION OF THE CLONED QUESTION WAS MISSED AND PLACED INTO THE VALIDATION SITE FOR THE CUSTOMER. CUSTOMER VALIDATION DID NOT DETECT THE CHANGE TO THE ASPIRIN FLAG FOR THE CLONED QUESTION. FURTHER INVESTIGATION FROM THE ENGINEERING DEPARTMENT FOUND THAT THERE IS A SOFTWARE DEFECT THAT DOES NOT ALLOW THE ASPIRIN FLAG OF THE QUESTION TO CLONE PROPERLY. ONCE THE ASPIRIN QUESTION IS CLONED, A SCREEN WILL DISPLAY THE CLONED INFORMATION. THE ASPIRIN FLAG IS RESET TO 'NO' DURING THE CLONING PROCESS. ALL INFORMATION IS TO BE REVIEWED PRIOR TO PROMOTING THE QUESTION. VALIDATION IS TO BE PERFORMED TO MAKE SURE THAT ALL THE INFORMATION THAT WAS CLONED IS ACCURATE. ALL CUSTOMERS WERE NOTIFIED BY TELEPHONE OF THIS DEFECT ON (B)(4) 2010, AND BY OFFICIAL NUMBERED LETTER ON (B)(4) 2010. ANALYSIS OF THE CUSTOMER'S DATABASES REVEALED NO INSTANCES OF UNSAFE PRODUCT BEING RELEASED FOR TRANSFUSION BECAUSE OF THIS DEFECT. | 1
1699753 | THE CUSTOMER'S CELL-DYN SAPPHIRE ASPIRATION BOTTOM SENSOR WAS REPLACED PER FA05APR2010. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED.><(B)(4). ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1701176 | CUSTOMER REPORTS PALETTE IS NOT SYNCHRONIZED WITH THE PATIENT JACKET. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1701614 | THE CUSTOMER REPORTS THAT THE RIS-IC IS NOT OPENING AND CLOSING EXAMS CORRECTLY IN PACS. THE SYSTEM WAS QUERYING THE DATABASE BY ORDER NUMBER, BUT PULLING THE EXAM WITH THE ACCESSION NUMBER INSTEAD, RESULTING IN RETRIEVAL OF THE INCORRECT EXAM/PATIENT. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1708026 | THE INITIAL RPTR SENT A DIGITAL COPY OF A RECORDING TO NEOVENTA, TOGETHER WITH SOME BACKGROUND INFO, AND THE FOLLOWING QUESTION: "TIME 15:12, THE BASELINE OF THE HEART FREQUENCY DECLINES FOR 4 MINUTES AND THE FECG COMPLEX IS NOT RECORDED AT ALL - THE DECISION OF THE EMERGENCY-SECTION IS MADE. THE RESULT: PERFECT BABY WAS BORN - THAT IS GOOD! (PH FROM THE UMBILICAL CORD WAS 7,31) (THE BABY WAS EXAMINED CAREFULLY LATER IN ORDER TO ELIMINATE THE POSSIBLE HEART DISEASES. NOTHING SHOWED UP - LUCKILY.) THE SCALP ELECTRODE WAS FIRMLY ATTACHED EVEN AFTER THE BABY WAS BORN. NOTICED THAT THERE WAS SAME KIND OF DECELERATION (SHORTER - THOUGH) AT 12:53 O'CLOCK. THE SIGNAL QUALITY SEEMED TO BE POOR IN BOTH SITUATIONS, AND NO T/QRS RATIOS WERE SHOWN AT THAT TIME EITHER. MY QUESTIONS ARE: IS IT POSSIBLE THAT THE SYSTEM SOFTWARE COULD HAVE CAUSED THE DECELERATION? COULD THERE BE ANY EFFECT (SORT OF INTERFERENCE) IF THE TOCO TRANSDUCER IS LAID BESIDE THE MOTHER - NOT REGISTRATING THE CONTRACTIONS AT THAT TIME?" THE EXAMINATION OF THE DIGITAL STAN DEVICE TRACING SUGGESTS THAT THE SKIN ELECTRODE WAS LOOSE (SEE BLUE FIELD IN FIGURE ABOVE) AND THEREFORE, NO CORRECT RECORDING OF FETAL HEART RATE OR ST INFO COULD BE PERFORMED BY THE STAN DEVICE. THE SITUATION OCCURRED WHEN THE PT WAS PLACED ON THE SIDE FOR AN EPIDURAL ANESTHESIA PLACEMENT. IT MAY INDICATE THAT THE PT WAS IN PAIN, RESULTING IN SWEATING, MOIST SKIN AND THE LOOSENING SKIN ELECTRODE. EVAL SUMMARY: THE DEVICE ITSELF HAS NOT YET BEEN EVALUATED, BUT DIGITAL DATA FROM THE RECORDING WAS SUBMITTED AS PART OF INITIAL REPORT. BASED ON THIS DATA, WHICH ALSO INCLUDES IMPEDANCE VALUES FROM ELECTRODES, THE CONCLUSION IS DRAWN THAT SKIN ELECTRODE WAS MOST LIKELY LOOSE THAT LAST 4 MINUTES OF THE RECORDING. IT CAN BE SEEN FROM DIGITAL DATA, THAT DURING THIS PERIOD, THE SYSTEM APPEARS TO HAVE PICKED UP SOME SORT OF PERIODIC SIGNAL INTERFERENCE IN THE SPIRAL-TO-SCALP LEAD, WITH AN AMPLITUDE FOUR TIMES STRONGER THAN THE BABY'S QRS AMPLITUDE, INTERFERENCE TO THE RIGHT, FETAL ECG QRS-COMPLEX TO THE LEFT. THIS TYPE OF RAW DATA SEGMENT ARE ONLY STORED ONCE EVERY FIVE MINUTES, AND THEREFORE, ONLY ONE OF THESE PEAKS ARE IDENTIFIED DURING THE FOUR-MINUTE INTERFERENCE. THIS MOST LIKELY CAUSED THE FHR DETECTION ALGORITHM TO STOP TRIGGING ON THE BABY QRS, AND STARTED TO PRESENT AN FHR TRACE BASED ON THIS STRONGER SIGNAL INSTEAD. (SIMILAR BEHAVIOUR WAS ACTUALLY SEEN FOR ABOUT 30 SECONDS EARLIER IN THE RECORDING AS WELL, WHEN THE SKIN ELECTRODE ALSO BECAME LOOSE.) THE SOURCE FOR THIS INTERFERENCE HAS NOT BEEN POSSIBLE TO IDENTIFY. AS THE DIGITAL DATA SHOWS VERY GOOD SIGNAL QUALITY AND PERFORMANCE APART FROM THE PERIODS OF HIGH ELECTRODE IMPEDANCE, THE DEVICE IS CONSIDERED TO WORK ACCORDING TO SPECIFICATIONS. BUT TO ENSURE THIS, FRONT-END ELECTRONICS (ELECTRODE CABLE AND INPUT AMPLIFIER) HAVE BEEN REQUESTED FOR EVAL, AS WELL AS COMPLEMENTARY INFO OF ELECTRODE BRANDS USED. WHEN SKIN OR SCALP ELECTRODE IS NOT PROPERLY ATTACHED THIS IS CLEARLY INDICATED BY THE DEVICE BY GIVING AN ALARM. INSTRUCTIONS TO CHECK SKIN AND SCALP ELECTRODES ARE ALREADY INCLUDED IN THE USER MANUALS AND IT IS ALSO PART OF TROUBLESHOOTING OF SIGNAL QUALITY PROBLEMS IN TRAINING MATERIAL. NO FURTHER FIELD ACTIONS ARE JUSTIFIED. | 1
1708811 | IT WAS REPORTED THAT THE MAIN SCREEN KEY IS STUCK AND IT IS IMPOSSIBLE TO START THE MONITOR. THERE WAS NO PT INJURY REPORTED. DRAEGER REFERENCE NUMBER: (B)(4). WHEN DRAEGER INITIALLY RECEIVED THIS REPORT, IT WAS DEEMED NOT REPORTABLE. HOWEVER, THE RESULTS OF THE INVESTIGATION DETERMINED THAT THE EVENTS INVOLVE REPORTS THAT SOME OF THE MONITOR'S FIXED KEYS BECOME INOPERATIVE OR ACTIVATE WITHOUT USER INTERACTION. THE WORST CASE SCENARIO OF THIS TEMPORARY MALFUNCTION COULD BE THAT THE MONITOR MAY DISCHARGE A PT AUTOMATICALLY. THEREFORE, IT WAS DETERMINED THAT THIS COMPLAINT IS REPORTABLE. AN INVESTIGATION WAS INITIATED WITH THE SUPPLIER OF THE SWITCH PANEL. IT WAS DETERMINED THAT THIS CONDITION IS DUE TO A CONTAMINATION PRESENT IN THE FIXED KEY PANEL MATERIAL. A RISK ANALYSIS WAS PERFORMED. DRAGER MEDICAL IS INITIATING A CORRECTIVE ACTION TO ADDRESS THIS PROBLEM. NO PT DEATHS OR INJURIES HAVE BEEN REPORTED AS A RESULT OF THIS CONDITION. | 1
1709382 | A CUSTOMER REPORTED THAT THEY HAVE OBSERVED CARRYOVER IN TESTS IMMEDIATELY FOLLOWING TESTING PERFORMED WITH ONE PATIENT SAMPLE CONTAINING ANTI-E, -C AND AN AUTOANTIBODY. CUSTOMER HAS NOT OBSERVED THIS ISSUE WITH ANY OTHER PATIENT SAMPLE. NO ERRONEOUS RESULTS WERE REPORTED. FE VERIFIED THAT THE CUSTOMER IS MEASURING DI WATER WHEN MAKING SOLUTIONS A AND B .THE BOTTLES ARE IN THE CORRECT POSITION. NO TUBING IS MIX WITHER BETWEEN THE BOTTLES OR ON THE WASH BLOCK. CUSTOMER DOES NOT PROP UP THE SOLUTION BOTTLES THEY ARE DOING THE FINAL WASH EVERY DAY. THEY ARE FOLLOWING THE INSTRUCTIONS IN THE USERS GUIDE. FE CHECKED FOR CRACKS IN THE WASH BLOCK AND FOUND NONE. WASH BLOCK WAS DIRTY SO FE CLEANED. FE CHECKED THE PROBE AND IT WAS NOT BENT. FE CHECKED THE VACUUM/PRESSURE IN THE BOTTLES. VACUUM AND PRESSURE WERE WITHIN SPECIFICATIONS. PROVUE IS OPERATING AS EXPECTED. (B)(4) | 0
1722242 | IT WAS REPORTED THAT DURING A PROCEDURE, THE SYSTEM GAVE AN ERROR MESSAGE FROM WINDOWS STATING IT HAD ENCOUNTERED A PROBLEM AND HAD TO BE SHUT DOWN. THE COMPUTER SHUT ITSELF DOWN. THE PHYSICIAN RESTARTED IT AND ENCOUNTERED THE SAME MESSAGE AT THE LOGIN SCREEN. THEY DID A HARD SHUTDOWN AND RESTARTED, AGAIN ENCOUNTERING THE SAME MESSAGE. THE SUPERDIMENSION PORTION OF THE CASE WAS THEN ABORTED WITH THE PATIENT UNDER GENERAL ANESTHESIA. A SUPERDIMENSION REPRESENTATIVE WENT TO THE SITE TO EVALUATE THE SYSTEM AND IT WAS DISCOVERED THAT THE MOUSE THAT WAS BEING USED WAS NOT THE MOUSE THAT WAS INITIALLY INSTALLED WITH THE SYSTEM. THE MOUSE THAT WAS BEING USED WAS INCOMPATIBLE WITH THE PC. THE CASE WAS NOT COMPLETED WITH THE SUPERDIMENSION SYSTEM WITH THE PATIENT UNDER GENERAL ANESTHESIA. THERE WAS NO HARM OR INJURY TO THE PATIENT, BUT IN AN ABUNDANCE OF CAUTION WE ARE FILING THIS MDR BECAUSE OF THE ADDITIONAL RISK ASSOCIATED WITH MULTIPLE EXPOSURES TO GENERAL ANESTHESIA. | 1
1722389 | ORTHO PROVUE SOFTWARE VERSION 3.1: THE ISSUE WAS DISCOVERED INTERNALLY AND FOLLOWING SCENARIO OCCURRED: WHILE PERFORMING IAT XM TESTING, THE SYSTEM FAILED TO DISPENSE PLASMA AFTER CLOT DETECTION (MAQERRCODE4 ). ALTHOUGH NO SAMPLE WAS DISPENSED, NO ERROR CODE WAS POSTED AND TEST RESULTS WERE REPORTED BY THE SYSTEM. OCD INITIATED A FAILURE INVESTIGATION UNDER (B)(4). | 0
1724069 | PER (B)(4), ABBOTT FIELD SERVICE REPRESENTATIVE REPLACED THE ASPIRATION BOTTOM SENSOR OF THE CUSTOMER'S CELL-DYN SAPPHIRE ANALYZER. NO IMPACT TO PATIENT RESULTS OR PATIENT MANAGEMENT WAS REPORTED. (B)(4). A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED.><(B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY OR MALFUNCTION WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. | 0
1726973 | AN AGFA INSTALLATION/CONFIGURATION ERROR DURING A REPORT CUSTOMIZATION RESULTED IN 26 REPORTS CONTAINING ERRONEOUS ANATOMIC SEGMENT LOCATORS. A REVIEW OF THE CUSTOMER'S DATABASE FOUND TWENTY-SIX STUDY REPORTS WITH ANATOMIC MODIFIER ERRORS. THERE WERE NO REPORT(S) OF PATIENT HARM. THIS MDR REPORT IS DUE TO POTENTIAL HARM THAT COULD HAVE BEEN CAUSED AS A RESULT OF THE PROBLEM. ONLY ONE CUSTOMER SITE WAS AFFECTED AND WE HAVE TAKEN INTERIM MEASURES THAT PREVENT THE PROBLEM FROM OCCURRING PENDING FINAL CORRECTIVE ACTION. | 1
1731250 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT HARM. (B)(4). ADDITIONAL DATE/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1736780 | REPORTEDLY, WHILE WASHING A BONE MARROW HARVEST, DURING THE CENTRIFUGATION PHASE, THE KIT GOT TWISTED AT THE ROTATING JOINT OF THE CENTRAL STEM TUBE (ROTATING SEAL). PARTIAL LOSS OF BONE MARROW (RTD = 53%). THIS REPORT IS BEING MADE DUE TO THE POTENTIAL FOR LOSS OF THE STEM CELLS CAUSED BY A PRODUCT PROBLEM DISCOVERED AS A RESULT OF THE REVIEW OF THE INVESTIGATION OF THIS REPORT ON 06/03/2010. (B)(4). THE 2991 ROUND BAG AND CERAMIC SEAL WERE RETURNED FOR INVESTIGATION. UPON RECEIPT CARIDIANBCT INVESTIGATORS CONFIRMED THAT THE CERAMIC SEAL DID NOT FREELY ROTATE. BLOOD RESIDUE WAS NOTED ABOVE THE CERAMIC SEAL AND MINIMAL RESIDUE IN THE ROUND BAG. THE TUBE BELOW THE CERAMIC SEAL WAS SEVERELY TWISTED AND HAD DETACHED FROM THE SEAL. THE ROOT CAUSE FOR THE FAILURE APPEARS TO BE RELATED TO A VENDOR MANUFACTURING DEFECT. INVESTIGATION FOUND THAT DURING THE MANUFACTURING PROCESS AT THE CERAMIC SUPPLIER, THE CORRECT PROCEDURE WAS NOT FOLLOWED, RESULTING IN ORGANIC SALTS REMAINING ON THE SMOOTH MATING SURFACE OF THE CERAMIC SEAL, WHICH THEN ACTS AS A HOLDING AGENT. THE INTERFERENCE OF THE HOLDING AGENT CAUSED THE SEAL TO NOT PERFORM AS INTENDED INCLUDING TWISTING THE POSSIBLE BREAKAGE DURING PROCESSING RESULTING IN THE POTENTIAL FOR LOSS OF PRODUCT. CORRECTIVE/PREVENTIVE ACTION: THE SUPPLIER'S MANUFACTURING SUPERVISORS WERE NOTIFIED OF THE ISSUE AND RETRAINING WAS REQUESTED. AN ADDITIONAL INSPECTION STATION HAS ALSO BEEN ADDED TO THE MANUFACTURING PROCESS TO ENSURE THE ROTATING SEAL IS NOT SEIZED AND FUNCTIONS PROPERLY PRIOR TO PACKAGING. IN AUGUST 2009, AN ADDITIONAL WASH CYCLE WAS INCORPORATED AT THE CERAMIC SUPPLIER TO PRODUCE CLEANER PARTS. RISK OF BONE MARROW LOSS MAY BE MITIGATED BY SPLITTING THE BONE MARROW INTO TWO BATCHES THAT ARE KEPT SEPARATE UNTIL ALL PROCESSING IS COMPLETE. (B)(4). WE WILL CONTINUE TO MONITOR. THE CUSTOMER IS NOT ALLEGING ANY NEGATIVE IMPACT TO THE BONE MARROW RECIPIENT AS THE CELLS WERE RECOVERED AND PROVIDED TO THE PT. THE COMPANY BECAME AWARE OF THE REPORTABILITY OF THIS EVENT ON 6/3/2010. | 0
1736801 | REPORTEDLY, WHILE WASHING A BONE MARROW HARVEST, DURING THE CENTRIFUGATION PHASE, THE KIT GOT TWISTED AT THE ROTATING JOINT OF THE CENTRAL STEM TUBE (ROTATING SEAL). ALL CELLS WERE ABLE TO BE RECOVERED IN A STERILE FASHION AND WERE GIVEN TO THE PT. THIS REPORT IS BEING MADE DUE TO THE POTENTIAL FOR LOSS OF THE STEM CELLS CAUSED BY A PRODUCT PROBLEM DISCOVERED AS A RESULT OF THE REVIEW OF THE INVESTIGATION OF THIS REPORT ON (B)(6) 2010. (B)(4). THE (B)(4) ROUND BAG AND CERAMIC SEAL WERE RETURNED FOR INVESTIGATION. UPON RECEIPT CARIDIANBCT INVESTIGATORS CONFIRMED THAT THE CERAMIC SEAL DID NOT FREELY ROTATE. BLOOD RESIDUE WAS NOTED ABOVE THE CERAMIC SEAL AND MINIMAL RESIDUE IN THE ROUND BAG. THE TUBE BELOW THE CERAMIC SEAL WAS SEVERELY TWISTED AND HAD DETACHED FROM THE SEAL. THE ROOT CAUSE FOR THE FAILURE APPEARS TO BE RELATED TO A VENDOR MANUFACTURING DEFECT. INVESTIGATION FOUND THAT DURING THE MANUFACTURING PROCESS AT THE CERAMIC SUPPLIER, THE CORRECT PROCEDURE WAS NOT FOLLOWED, RESULTING IN ORGANIC SALTS REMAINING ON THE SMOOTH MATING SURFACE OF THE CERAMIC SEAL, WHICH THEN ACTS AS A HOLDING AGENT. THE INTERFERENCE OF THE HOLDING AGENT CAUSED THE SEAL TO NOT PERFORM AS INTENDED INCLUDING TWISTING THE POSSIBLE BREAKAGE DURING PROCESSING RESULTING IN THE POTENTIAL FOR LOSS OF PRODUCT. CORRECTIVE/PREVENTIVE ACTION: THE SUPPLIER'S MANUFACTURING SUPERVISORS WERE NOTIFIED OF THE ISSUE AND RETRAINING WAS REQUESTED. AN ADDITIONAL INSPECTION STATION HAS ALSO BEEN ADDED TO THE MANUFACTURING PROCESS TO ENSURE THE ROTATING SEAL IS NOT SEIZED AND FUNCTIONS PROPERLY PRIOR TO PACKAGING. IN (B)(6) 2009, AN ADDITIONAL WASH CYCLE WAS INCORPORATED AT THE CERAMIC SUPPLIER TO PRODUCE CLEANER PARTS. RISK OF BONE MARROW LOSS MAY BE MITIGATED BY SPLITTING THE BONE MARROW INTO TWO BATCHES THAT ARE KEPT SEPARATE UNTIL ALL PROCESSING IS COMPLETE. THE OCCURRENCE RATE FOR THIS EVENT TYPE IS (B)(4). WE WILL CONTINUE TO MONITOR. THE CUSTOMER IS NOT ALLEGING ANY NEGATIVE IMPACT TO THE BONE MARROW RECIPIENT AS THE CELLS WERE RECOVERED AND PROVIDED TO THE PT. THE COMPANY BECAME AWARE OF THE REPORTABILITY OF THIS EVENT ON (B)(6) 2010. | 0
1737399 | THE ACCOUNT STATED THAT A FALSE POSITIVE AXSYM TOXO IGG RESULT WAS GENERATED. AN INITIAL NEGATIVE RESULT OF 0.1 IU/ML WAS GENERATED. THE SAMPLE WAS REPEATED AND A POSITIVE RESULT OF 23.3 IU/ML WAS GENERATED. THE SAMPLE WAS RUN ON THE ARCHITECT AND A NEGATIVE RESULT OF 1.4 IU/ML WAS GENERATED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. THIS REPORT IS BEING FILED ON AN INTERNATIONAL PRODUCT, LIST NUMBER 9K08 THAT HAS A SIMILAR PRODUCT DISTRIBUTED IN THE US, LIST NUMBER 3B22. | 0
1739861 | DURING THE FREEZE TESTING OF THE PROBE PRIOR TO CONTACT WITH THE PATIENT PROBE 1 FROZE UP THE ENTIRE SHAFT. IT WAS REPLACED. TESTING INDICATED A VACUUM INSULATION FAILURE IS THE CRYOPROBE. FROSTING OF THE SHAFT WAS CONFIRMED. NO PATIENT INJURY WAS REPORTED. | 0
1739878 | ON (B)(6) 2010, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THE POWER CORD WAS HOT TO TOUCH ON THE (B)(4) PRINTER, FOR THE I-STAT 1 ANALYZER. THERE WAS NO REPORT OF ANY INJURY ASSOCIATED WITH THE COMPLAINT. | 0
1740039 | PROBE #1 STARTED FROSTING DURING THE PROCEDURE. CASE WAS STOPPED AND PROBE #1 WAS REPLACED. NO INJURY. TESTING INDICATED A VACUUM INSULATION FAILURE IN THE CRYOPROBE. FROSTING OF THE SHAFT WAS CONFIRMED. NO PT INJURY WAS REPORTED. | 0
1740824 | THE LAB SUPERVISOR CALLED TO REPORT THAT HER TECHS HAD FORGOTTEN TO PLACE THE SOLUTION A AND SOLUTION B FLUIDS INTO THEIR RESPECTIVE CONTAINERS WHEN THEY LOADED THE SYSTEM UP FOR TESTING A RUN. THEY PROCEEDED TO TEST WITH ONLY DEIONIZED WATER LOADED IN BOTH WASH CONTAINERS. THE TECH PROCEEDED TO RUN AND FINISH TESTING BEFORE THE SUPERVISOR REALIZED THE ERROR. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4). PATIENT RESULTS WERE NOT COMPROMISED. LABORATORY SUPERVISOR ALERTED THE USER OF THEIR MISTAKE. RESULTS WERE REPEATED WITH THE APPROPRIATE WASH SOLUTIONS ADDED TO THE INSTRUMENT. | 0
1743152 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1749473 | ON (B)(6) 2010, THE CUSTOMER CALLED AND STATED THAT SHE COULD NOT FIND A SPECIFIC STUDY FOR A PATIENT. PRODUCT WAS EVALUATED REMOTELY AFTER THE COMPLAINT WAS TAKEN. THE DEVICE WAS EVALUATED REMOTELY, WE ACCESSED THE CUSTOMERS DATABASE.: SUMMARY OF EVALUATION: ON (B)(6) 2010, (B)(6), RAD TECH FROM THE CLINIC, CALLED AND STATED THAT THEY ARE UNABLE TO FIND THE STUDIES FOR 2 SPECIFIC PATIENTS. STRYKER WAS UNABLE TO LOCATE THE STUDIES ON THE SYSTEM WHICH APPEARED TO HAVE BEEN LOST. ON (B)(6) 2010, FURTHER INVESTIGATION INTO THE ISSUE REVEALED THE FOLLOWING: DURING THE WEEK OF (B)(6) 2010, STRYKER REPRESENTATIVES ATTEMPTED TO UPGRADE THE CUSTOMER AT (B)(6) ORTHOPEDIC SURGEONS ((B)(6)) FROM AN OLDER VERSION OF THE OFFICE PACS SOFTWARE APPLICATION TO A NEWER VERSION BY SETTING UP DUAL SERVERS. ONE FOR LEGACY (3.4D) STUDIES AND ONE FOR NEW (POWER 4.1 SP3) STUDIES. THERE WAS ALSO AN UPGRADE EFFORT FOR THE EXTERNAL RAID ENCLOSURE HARDWARE TO GO FROM AN OUTDATED INLINE SYSTEM TO A NEW RAID STORAGE SYSTEM MADE BY IBM. THIS INVESTIGATION REVEALED THAT DURING THE UPGRADE A NUMBER OF STUDIES WERE LOST FROM THE CUSTOMER'S DATA SET. THIS OCCURRED BETWEEN THE DATES OF (B)(6) 2010 THROUGH (B)(6) 2010 DURING WHICH THE CLINIC GENERATED PATIENT STUDIES. THE STUDIES WERE LOCALLY GENERATED AT THE CLINIC AND CONTAINED CR TYPE X-RAY IMAGES. AT THIS TIME IT IS UNCLEAR WHETHER THE STUDY LOSS WAS DUE TO INCOMPLETE/INCORRECT TRANSFER OF OLD DATA OR DUE TO AN INVALID CONFIGURATION OF THE DUAL SYSTEM SETUP AT THE TIME. THE MOST PROBABLE ROOT CAUSE APPEARS TO BE RELATED TO INSUFFICIENT DATA TRANSFER PROCEDURES, THE LACK OF A DOCUMENTED PROCESS, AND INSTRUCTIONS FOR SYSTEM CONFIGURATION DURING UPGRADE EFFORTS IN ORDER TO ENSURE THE PRESERVATION AND FIDELITY OF CUSTOMER DATA. THE CLINIC HAS NOT REPORTED ANY PATIENT HARM AS A RESULT OF THE PATIENT DATA LOSS. THE LOST DATA IS COMPROMISED OF CR TYPE X-RAY STUDIES PERFORMED BETWEEN (B)(6) 2010 THROUGH (B)(6) 2010 ONLY. | 1
1750829 | THE CUSTOMER REPORTED A NEAR PT INCIDENT INVOLVING PICIS' ED (EMERGENCY DEPT.) ELECTRONIC HEALTH RECORD APPLICATION RELATED TO ALLERGY DATA RECEIVED VIA INTERFACE FROM THE CUSTOMER'S ENTERPRISE INFO SYSTEM, WHICH WAS 'CODED' USING A 3RD PARTY DRUG DATABASE TERMINOLOGY NOT SUPPORTED BY PICIS. THE NUMERIC CODE REPRESENTING 'CEFTRIAXONE' IN THE ENTERPRISE SYSTEM REPRESENTED A DIFFERENT ALLERGY CONCEPT IN THE PICIS' SUPPORTED THIRD PARTY DRUG DATABASE. THE SCREEN DISPLAYED THE CORRECT ALLERGY NAME OF CEFTRIAXONE TO THE PHYSICIAN, BUT THE BACKGROUND ALLERGY CHECKING LOGIC (DETERMINED BY THE NUMERIC CODE SENT FROM THE INTERFACE) DID NOT NOTIFY THE PHYSICIAN OF A POTENTIAL ALLERGY INTERACTION, WHEN A MEDICATION ORDER FOR ROCEPHIN WAS WRITTEN. THE PT WAS GIVEN THE MEDICATION WHILE IN THE ED. THE PT DID NOT HAVE ANY ADVERSE REACTION TO THE MEDICATION. OUR INVESTIGATION FOUND THAT THE REASON FOR NO ALLERGY INTERACTION NOTIFICATION WAS A RESULT OF ALLERGIES CODED USING A NON-SUPPORTED EXTERNAL DRUG DATABASE SCHEMA BEING SENT TO ED PULSECHECK FROM THE ENTERPRISE INFO SYSTEM. PICIS IS WORKING WITH THE CUSTOMER TO REVIEW ALLERGY CODES SENT PRIOR TO JUNE 2010 TO ASCERTAIN IF THERE ARE ANY ADDITIONAL ERRONEOUS ALLERGY CODES STORED IN THE DATABASE. ADDITIONALLY, DURING THE COURSE OF OUR INVESTIGATION WE HAVE IDENTIFIED THREE SIMILAR CUSTOMER SITE CONFIGURATIONS WITH DISPARATE DRUG DATABASES. OF THESE SITES, TWO HAVE REPORTED THIS CONCERN TO PICIS, ONE OF WHICH HAS BEEN SUBMITTED IN AN EARLIER REPORT. WE ARE WORKING WITH EACH OF THESE CUSTOMERS TO PREVENT UNSUPPORTED ALLERGY CODES FROM BEING SENT TO THE PICIS APPLICATION. PICIS IS ALSO EXPLORING THE POTENTIAL FOR SIMILAR SYSTEM ADMINISTRATION ERROR IN OTHER SITES. WE WILL EVALUATE THE NEED FOR CORRECTIVE ACTION AND IMPLEMENT ACCORDINGLY. | 1
1751085 | THE CUSTOMER ((B)(6) HOSPITAL) REPORTED THAT A LABORATORY TECHNICIAN, WHILE ASSISTING IN REPLACING CLEANER SOLUTION IN A LABORATORY INSTRUMENT, WAS SPLASHED IN THE LEFT EYE WITH WASTE CONTAINER MATERIAL (BIOHAZARD AND CHEMICAL MATERIAL). THIS OCCURRED WHILE ASSISTING IN THE MAINTENANCE OF THE INSTRUMENT. THE TECHNICIAN WAS WEARING GLOVES AND A LAB COAT, BUT WAS NOT WEARING GOGGLES AS REQUESTED BY SAFETY INSTRUCTIONS. THE MAINTENANCE ON THE INSTRUMENT WAS COMPLETED ACCORDING TO THE REQUIREMENTS. THE REPORTED EVENT WAS DEEMED AN ISOLATED INCIDENT ATTRIBUTED TO FAILURE OF THE USER TO FOLLOW SAFETY INSTRUCTIONS AND THE LACK OF PROPER PROTECTIVE GEAR. | 0
1755467 | A CUSTOMER REPORTED THAT THE HEART RATE ALARM LIMITS CHANGE BACK TO FACTORY DEFAULT WHEN A TRAM MODULE IS USED WITH THE B850 MONITOR. THE DEVICE WAS REPORTEDLY NOT IN USE DURING THE INCIDENT. THE CORRECTION/REMOVAL REPORTING NUMBER IS NOT AVAILABLE AT THIS TIME. GE HEALTHCARE INVESTIGATED THE REPORTED ISSUE AND THE COMPLAINT WAS CONFIRMED. THE CUSTOMER DEFINED HEART RATE ALARM LIMITS CHANGE WHEN A TRAM MODULE IS USED WITH THE B850 MONITOR. WHEN AN INVASIVE BLOOD PRESSURE (IBP) CABLE IS PLUGGED INTO PRESSURE CHANNELS P1 OR P2, THE CUSTOMER DEFINED HEART RATE ALARM LIMITS CHANGE TO THE FACTORY DEFAULT SETTING OF 150/50. WHEN IT IS PLUGGED INTO PRESSURE CHANNEL P3, THE CUSTOMER DEFINED HEART RATE ALARM LIMITS CHANGES TO 180/30. THE ROOT CAUSE OF THE REPORTED ISSUE IS THAT WHEN SINGLE HR (HEART RATE) ALARM MODE IS SELECTED, ALL HR/PR (HEART RATE/PULSE RATE) LIMIT CHANGES FROM THE TRAM MODULE ARE INTERPRETED AS ECG HR LIMIT CHANGES TO IMPACT THE ONE COMMON LIMIT. WHEN CONNECTING AN IP CHANNEL, THE CHANNEL SENDS ITS CONFIGURATION, WHICH IS IGNORED, AND THE CONFIGURATION ON THE MONITOR IS FORCED BACK TO THE TRAM MODULE. THE ECG HR LIMITS ARE HOWEVER NOT DISTINGUISHED AS INVASIVE PRESSURE (IP) CONFIGURATION AND ARE LET THROUGH, RESULTING IN THE HR LIMITS BEING SET ACCORDING TO THE IP PR (CHANNEL DEPENDENT) DEFAULT CONFIGURATION FROM THE TRAM. THIS IS A TRAM SPECIFIC ISSUE AS ALARM LIMITS ARE ONLY ACCEPTED FROM TRAM (DUE TO POSSIBLE SIMULTANEOUS CONNECTION TO TRANSPORT PRO). GE HEALTHCARE WILL BE REPORTING THIS ISSUE TO THE FDA PER 21 CFR PART 806. CORRECTIONS TO AFFECTED PRODUCTS WILL BE DESCRIBED IN THIS REPORT. | 1
1755937 | CUSTOMER REPORTED THAT THE PROVUE READ A NEGATIVE RESULT AS POSITIVE. ISSUE OCCURRED DURING AN ANTIBODY SCREEN. CELL II WAS REPORTED AS 1+, HOWEVER, UPON REVIEW OF THE IMAGE, THE WELL WAS NEGATIVE. NO ERRONEOUS RESULTS WERE REPORTED. OCD FIELD ENGINEER WAS ON-SITE AND MADE THE NECESSARY REPLACEMENTS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
1757147 | WHEN A "?" SYMBOL IS PRESENT IN A NOTE, SUCH AS A DAILY PROGRESS NOTE, WRITTEN BY A CLINICIAN IN ECAREMANAGER AND SUBSEQUENTLY INTERFACED TO THE HOSPITAL INFORMATION SYSTEM, THE "?" IS STRIPPED FROM THE OUTBOUND HL7 MESSAGE. THIS MAY NOT ADEQUATELY CONVEY THE INTENT OF THE NOTE WHEN VIEWED IN THE HOSPITAL INFORMATION SYSTEM. THE "?" REMAINS VISIBLE IN THE NOTE VIEWED USING ECAREMANAGER. ADDITIONAL MODEL #V3.7.1. ADDITIONAL DEVICE MANUFACTURE DATE: 04/2010. | 1
1759917 | THIS MDR IS BEING RETROSPECTIVELY SUBMITTED UPON RECENT KNOWLEDGE OF A SIMILAR COMPLAINT FOR WHICH MDR WAS SUBMITTED PREVIOUSLY. PMS RECEIVED A COMPLAINT ALLEGING THAT THE BRIGHTVIEW SYSTEM WAS MANIFESTING INTERMITTENT MOTION STOPS DURING RADIUS MOTION. IT WAS IDENTIFIED THAT THE COMPLAINT ISSUE WAS SIMILAR TO A PREVIOUSLY REPORTED ISSUE WHERE THE LEAD SCREW NUT ON THE CAMERA DETECTOR CAN COME LOOSE AND THE FAILURE CAN RESULT IN THE DETECTOR DRIFTING, UNDER THE INFLUENCE OF GRAVITY, TO ITS HARD LIMIT. THIS MDR IS BEING RETROSPECTIVELY SUBMITTED UPON RECENT KNOWLEDGE OF A SIMILAR COMPLAINT FOR WHICH MDR WAS SUBMITTED PREVIOUSLY. THE SYSTEM WAS AT A BETA SITE WHEN THIS COMPLAINT WAS RECEIVED. THE LEAD SCREW WAS REPLACED. THE ISSUE HAS NOT RECURRED AFTER THE REPLACEMENT OF THE LEAD SCREW. (B)(4), MDR#2916556-2009-00004. | 0
1765650 | PROBE WAS BEING PRETESTED AND THE PROBE FROSTED TO THE HANDLE. | 0
1765811 | UPON INITIAL PRETESTING, FROSTING MIGRATED UP THE SHAFT. NO PT CONTACT. | 0
1774313 | THE ISSUE INVOLVES CERNER PROVISION DOCUMENT IMAGING UTILIZED TO CAPTURE NON-CLINICAL IMAGES. IN CERNER PROVISION DOCUMENT IMAGING, AN OLDER VERSION OF THE IMAGES MAY BE DISPLAYED. PT CARE COULD BE ADVERSELY AFFECTED, AS CLINICIANS COULD DELIVER INCORRECT CARE BASED ON AN OUT-OF-DATE NON-CLINICAL DOCUMENT SUCH AS ADVANCED DIRECTIVES OR LIVING WILLS. CERNER HAS NOT RECEIVED COMMUNICATION ON ANY ADVERSE PT EVENTS AS A RESULT OF THIS ISSUE. CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON (B)(4), 2010 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE AND A SOFTWARE PATCH IS INCLUDED TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORPORATION CONSIDERS THIS ISSUE TO BE RESOLVED AND NO FURTHER NARRATIVE IS REQUIRED FOR FOLLOW-UP. | 1
1774848 | THE CUSTOMER REPORTED THAT AFTER THE UPGRADE, IT WAS OBSERVED THAT THE "MORE" BUTTON IS MISSING WHICH ENABLES THE LVI SCREEN TO DISPLAY LONG SETUP NOTES. THERE WAS NO REPORT OF INJURY TO THE PT AND NO PT DATA WAS PROVIDED. ALTHOUGH THERE WAS NO REPORTED INJURY IN THIS CASE, THE AVAILABLE INFO SUGGESTS A MALFUNCTION IN THE DEVICE. THOUGH STILL UNDER INVESTIGATION, VARIAN HAS DETERMINED THAT A MDR IS APPROPRIATE, AS THIS MALFUNCTION, SHOULD IT RECUR, COULD POTENTIALLY CAUSE A SERIOUS INJURY. ADDITIONAL FOLLOW-UP TO THIS MDR IS EXPECTED UPON COMPLETION OF THE INVESTIGATION. | 1
1779463 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO THE DL2000 DATA MANAGEMENT UNIT FAILURE TO CLEAR THREE SAMPLE DATA RAN ON (B)(6) 2010. THE INITIAL TEST WAS STILL PENDING WHEN THE NEW ORDER WAS ENTERED ON (B)(6) 2010. THEREFORE, DL2000 COMBINED THE NEW TESTS RESULT WITH THE PREVIOUS ONE AND UPLOADED THE RESULTS POSTED FROM (B)(6) 2010. THE RESULT WAS NOT REPORTED OUT OF THE LABORATORY. PATIENT TREATMENT WAS NOT IMPACTED BY THIS EVENT. THE HOSPITAL HAD INCREASED THE NUMBER OF BARCODE DIGITS FOR SAMPLE IDS TO SIX. TO ACCOMMODATE THIS CHANGE, SETTINGS WERE CHANGED AT THE DL2000. THE LAB CONFIGURED THE DL2000 DATA MANAGEMENT TO DELETE PROGRAMMING >5 DAYS OLD. THE UNIT CLEARED ALL SAMPLES EXCEPT THREE. SERVICE OBTAINED DATA AND FORWARDED TO THE SOFTWARE VENDOR. THE ROOT CAUSE FOR THIS EVENT IS CUSTOMER ERROR AND DL2000 DATA MANAGEMENT. | 1
1783180 | DURING A PROSTATE CRYO PROCEDURE, THE WARMER CEASED FUNCTIONING AND FROZE INTO THE URETHRA. THE UNIT WAS STARTED AS NORMAL, IT WAS CHECKED AGAIN AND RUNNING PRIOR TO THE FIRST FREEZE. UPON PREPARING TO BEGIN THE SECOND FREEZE CYCLE, I NOTICED THAT THE SPINNER WAS NO LONGER MOVING. UPON FURTHER EXAMINATION, THE WARMING CATHETER WAS STUCK IN THE URETHRA AND I COULD NOT GET IT FLOWING AGAIN. I DETECTED NO LEAKS, KINKS, MACHINE FAILURE OR OTHER EXPLANTATION. WHILE TRYING TO GET THE WARMER FLOWING AGAIN, I INADVERTENTLY POPPED A HOLE IN THE WARMING POUCH. THIS WAS IMMEDIATELY REPLACED. ONCE THE WARMER HAD THAWED ENOUGH TO MOVE AND PROBE PLACEMENT WAS RECHECKED, THE WARMER RAN PERFECTLY THE ENTIRE REST OF THE TIME. THE WARMER WAS ALSO LEFT IN FOR 30 MINUTES IN RECOVERY TO ALLOW FOR THE MAXIMUM PROTECTION OF THE URETHRA AS POSSIBLE. THE HOSPITAL HAS HELD ONTO THE CATHETER AND TUBING FOR THEIR OWN INVESTIGATION. DEVICE NOT RETURNED FOR TESTING. | 0
1783211 | CRYOABLATION PERFORMED NORMALLY BUT PHYSICIAN REPORTED URETHRAL STRICTURE DURING THE RECOVERY PERIOD THAT REQUIRED URETHRAL DILATION. PHYSICIAN HAS HAD SEVERAL (5) OCCURRENCES IN THE PAST SIX WEEKS. PHYSICIAN NOTED THAT OCCURRENCES STOPPED AFTER THE REPORT. DEVICE NOT RETURNED FOR TESTING. | 0
1785154 | THE TUBING BLEW OFF THE INJECTOR END, SPRAYING CONTRAST IN THE ROOM. THIS HAPPENED TWO TIMES. INJECTOR SETTINGS: VOLUME: 35ML, FLOW: 12MLS/SEC, PRESSURE: 900 PSI. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF TWO REPORTS FOR THIS EVENT. 1721504-2010-00213. DEVICE EVALUATION. THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION/ INVESTIGATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IN COMPLETED. EVALUATION, METHOD: DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
1785184 | A CUSTOMER REPORTED A LOSS OF COMMUNICATION AT THE CENTRAL STATION. NO ADDITIONAL INFO IS AVAILABLE AT THIS TIME. GE HEALTHCARE IS UNAWARE OF ANY PT INJURIES OR DEATH ASSOCIATED WITH THIS EVENT. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. THE GE FIELD ENGINEER ATTEMPTED TO CONTACT THE CUSTOMER ON THREE SEPARATE OCCASIONS TO OBTAIN MORE INFO. NO RESPONSE HAS BEEN RECEIVED AT THIS TIME. | 1
1785308 | A SOFTWARE DEFECT IN DONOR MANAGEMENT SYSTEM (ALL VERSIONS) HAS THE POTENTIAL TO INADVERTENTLY ALLOW A DONOR WHO HAS FAILED THE SCREENING TEST TO CONTINUE WITH THE DONATION PROCESS AND DONATE A UNIT OF PLASMA. EVALUATION SUMMARY: A SOFTWARE DEFECT IN DONOR MANAGEMENT SYSTEM (ALL VERSIONS) HAS THE POTENTIAL TO INADVERTENTLY ALLOW A DONOR WHO HAS FAILED THE SCREENING TEST TO CONTINUE WITH THE DONATION PROCESS AND DONATE A UNIT OF PLASMA. EXAMPLES MIGHT INCLUDE A DONOR WHO FAILED A HEMATOCRIT, PULSE, BLOOD PRESSURE (SYSTOLIC OR DIASTOLIC), WEIGHT, TEMPERATURE OR TOTAL PROTEIN TEST. HAE CONDUCTED AN INVESTIGATION AND IT WAS DETERMINED THAT: A DEFECT DID EXIST WITHIN THE DMS APPLICATION THAT COULD BE REPRODUCED UNDER A VERY SPECIFIC SET OF CIRCUMSTANCES. THE DEFECT WAS NOT LIMITED TO JUST THE PULSE SCREENING TEST BUT TO ALL SCREENING TESTS (E.G. HEMATOCRIT, PULSE, TEMPERATURE, BLOOD PRESSURE, WEIGHT, TOTAL PROTEIN) THAT CAN BE INPUT ON THE SAME SCREEN (IE (B)(4) PRE-DONATION SCREENING). ADDITIONALLY, ALL CLIENT DATABASES WERE INVESTIGATED BETWEEN JULY 15 AND JULY 16, 2010 FOLLOWED BY AN OFFICIAL COMMUNICATION LETTER DESCRIBING THE ISSUE WAS SENT ON JULY 20, 2010. ONLY 3 OF 11 CUSTOMERS WERE AFFECTED BY THE ISSUE. A TOTAL OF 8 DONORS (FOR THOSE 3 CUSTOMERS) WERE IMPACTED BUT NO UNITS WERE INADVERTENTLY COLLECTED EVEN THOUGH THE DEFECT HAD OCCURRED. | 1
1786382 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADD'L DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICIAN ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1788528 | THE CUSTOMER REPORTED THAT THEY CREATED A WORKFLOW WHERE THEIR FACILITIES CAN ASSIGN A MEDICAL RECORD NUMBER (MRN) THAT ANOTHER FACILITY CAN USE. AN OCCURRENCE HAPPENS WHEN TWO DIFFERENT SITES USE THE SAME MRN AND SEND IT TO THEIR DATABASE, (B)(4) WHEN QUERYING BY MRN, CAN PULL ANY PATIENTS WITH THE SAME MRN. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE FILED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1803526 | CUSTOMER STATED THAT A SAMPLE SHOWING MFAGGLUTINATION IN MANUAL GEL IS NEGATIVE ON THE PROVUE. OCD FILED ENGINEER WAS ON SITE. FE PERFORMED READER CAMERA ADJUSTMENTS AND PRODUCED A NEW REFERENCE IMAGE. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
1805283 | THE CUSTOMER REPORTED THAT DURING ROUTINE CHART QA, THEY FOUND THAT 2 RA ARCS WERE TREATED WITH COUCH ROTATION VALUES OVERRIDDEN. TREAT HISTORY IN ARIA RTC INDICATES THAT 2 RA ARCS WERE DELIVERED WITH PARAMETERS OVERRIDDEN BY THE USER. SEQUENCE OF EVENTS COULD NOT BE REPRODUCED BY THE CUSTOMER - STATED THAT INTERLOCKS PREVENT TREATMENT. THE CUSTOMER ACKNOWLEDGES THE OVERRIDE BY THE THERAPIST. ONE FRACTION (2 ARCS) DELIVERED (B)(4). NO SERIOUS INJURY TO THE PT WAS REPORTED AND NO FURTHER PT DATA WAS PROVIDED. BASED ON THE INITIAL REPORT, VARIAN'S MEDICAL ADVISOR DETERMINED THAT THE EVENT DID NOT RESULT IN SERIOUS INJURY. MEDICAL PROFESSIONAL REPORT INDICATED: PT WAS BEING TREATED TO THE LUNG WITH AN ARC TECHNIQUE. ON ONE DAY PT RECEIVED AN EXTRA 100 CGY FOR AN OVER DOSE FOR THE WEEK OF 11% (1000 INSTEAD OF 900 CGY). THIS IS AN 11% OVERDOSE FOR THE WEEK AND 2.5% IN THE COURSE OF 4140. THE OVERDOSE IS WITHIN THE PLANNED TARGET VOLUME. THIS OVER DOSE IS WELL WITHIN THE ACCEPTED LEVEL OF TOLERANCE. HOWEVER, THE INITIAL INFO INDICATED AN ALLEGED DEVICE MALFUNCTION. THOUGH STILL UNDER INVESTIGATION, VARIAN HAS DETERMINED THAT A MDR IS APPROPRIATE. ADD'L F/U TO THIS MDR IS EXPECTED UPON COMPLETION OF THE INVESTIGATION. | 1
1805687 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE RESULT WHEN TESTING A PATIENT SAMPLE WITH CAPTURE-R READY-SCREEN 3 (CRRS 3) ON THE ECHO. THE PATIENT HAS A KNOWN ANTI-E. THE SAMPLE WAS RE-TESTED ON THE ECHO AND RESULTED POSITIVE (4+). THE REVIEW OF THE IMAGES: CRRS 3: THE SAMPLE WAS NON REACTIVE WITH CELLS 1-3, POSITIVE CONTROL WAS 4+. WELL IMAGES WITH CELLS 1 AND 3 APPEARED NEGATIVE WITH TIGHT BUTTONS AND NO ADHERENCE. WELL IMAGE OF CELL 2, DISPLAYED A WEAK POSITIVE REACTION WITH FUZZY BUTTON AND A SMALL OF AMOUNT ADHERENCE. CELL 2 IS E+E-. THE CUSTOMER DID NOT RETURN PRODUCT OR SAMPLE FOR FURTHER SEROLOGICAL TESTING. | 0
1808328 | DURING THE REVIEW OF WADIANA LOGS RECEIVED FROM THE CUSTOMER AS PART OF TC10-186 PRODUCT NOTIFICATION REVIEW OF PREVIOUSLY REPORTED RESULTS ON THE ORTHO PROVUE ANALYZER, THE FOLLOWING INCIDENT WAS IDENTIFIED. THE INCIDENT WAS RELATED TO FAILURE TO DISPENSE PLASMA FOLLOWING A CANCELLED MICROTUBE PER TC 10-174 CANCELLED MICROTUBES USING SOFTWARE VERSION 3.1 ON THE ORTHO PROVUE ANALYZER. THE CONCERN IS RELATED TO THE CROSSMATCH-IAT TEST. EVENT OCCURRED ON (B)(6) 2010 AT 12:24 PM TO BATCH 6591. PROVUE FAILED TO DELIVER PATIENT PLASMA INTO MICROWELL #5 TO THE MTS ANTI-IGG GEL CARD (BARCODE # 03171001042612002745) FOR THE IAT CROSSMATCH TO SAMPLE ID (B)(4) (POSITION 6 ON CAROUSEL) TO DONOR ID# (B)(6). PROVUE HAD RESULTED THE TEST WITH A NEGATIVE RESULT. INCIDENT WAS IDENTIFIED DURING A REVIEW OF THE CUSTOMER'S FILES. (B)(4). | 0
1813105 | REFERENCE: MDR 1220063-2009-00045. PER TELEPHONE COMMUNICATION WITH DRAEGER MEDICAL REGARDING THE SUBMITTAL OF ADDITIONAL MDRS FOR EACH SERIAL NUMBER LISTED ON THE INITIAL MDR SUBMITTED, DRAEGER IS SUBMITTING THIS REPORT. (B)(4). IT WAS REPORTED THAT THE DEVICE SWITCHED FROM DISCHARGE MODE SPORADICALLY AND ON THEIR OWN TO ACTIVE MODE. A DEVICE ALSO SWITCHED FROM ACTIVE MODE TO DISCHARGE MODE. THERE WAS NO PATIENT INJURY REPORTED. AN INVESTIGATION WAS INITIATED WITH THE SUPPLIER OF THE SWITCH PANEL. IT WAS DETERMINED THAT THIS CONDITION IS DUE TO A CONTAMINATION PRESENT IN THE FIXED KEY PANEL MATERIAL. A RISK ANALYSIS WAS PERFORMED. DRAEGER MEDICAL IS INITIATING A CORRECTIVE ACTION TO ADDRESS THIS PROBLEM. NO PATIENT DEATHS OR INJURIES HAVE BEEN REPORTED AS A RESULT OF THIS CONDITION. | 1
1818559 | A CUSTOMER REPORTED A LOSS OF MONITORING (ECG WAVEFORM DROPOUT) UP TO ONE HOUR DURATION FOR MULTIPLE TELEMETRY PTS. "NO TELEM" MESSAGES WERE DISPLAYED AT THE CENTRAL STATION. NO PT INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1818563 | IT WAS REPORTED THAT A NURSE WAS SITTING AT THE CIC (CENTRAL STATION) AND NOTICED THAT A PT APPEARED TO BE IN VENTRICULAR TACHYCARDIA. THE CIC REPORTEDLY DID NOT ALARM. ADDITIONAL PT INFORMATION INCLUDING THE OUTCOME OF THE PT IS UNK AT THIS TIME. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1826945 | THE CUSTOMER REPORTED THAT THE PROVUE PROBE IS DRIPPING. NO ERROR MESSAGES WERE POSTED. THE CUSTOMER INDICATED THAT THE PROVUE WAS TAKEN OUT OF USE. NO ERRONEOUS RESULTS WERE POSTED. PROBE ISSUES CAN LEAD TO THE DILUTION OF REAGENTS AND SAMPLES, CARRYOVER, AND CROSS-CONTAMINATION WHICH MAY LEAD TO ERRONEOUS TEST RESULTS. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND OBSERVED A CRACKED WASH STATION. THE FE REPLACED THE WASH STATION AND PROBE. ALL ADJUSTMENTS PERFORMED. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4) | 0
1829190 | THE CUSTOMER REPORTS THAT CANCELLED AND REJECTED IMAGES ARE BEING RETRIEVED AND DISPLAYED FROM THE ARCHIVE WHEN USING THE WADO (WEB ACCESS TO DICOM PERSISTENT OBJECTS) PROTOCOL. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1829409 | A CUSTOMER REPORTED A LOSS OF TELEMETRY MONITORING OCCURRED DUE TO A PRESUMED CIC (CENTRAL STATION) HARD DRIVE FAILURE. PATIENTS WERE REPORTEDLY RELOCATED FOR ALTERNATE MONITORING. NO DEATH OR INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1834509 | FIRST FREEZE WAS FUNCTIONING NORMALLY AT 2 MIN, AND 5:30 MIN WITH FLOW WHEEL MOVING. AT 7 MIN THE FLOW WHEEL WAS NOT MOVING BUT THE TUBING FELT WARM. A CHECK FOR KINKS IN THE TUBING WAS PERFORMED AND NONE FOUND. WE IMMEDIATELY STOPPED FREEZING AND STARTED THAWING. AT THIS POINT, THE WARMING CATHETER WAS FROZEN TO THE URETHRA. WE LET THE GLAND THAW AND IRRIGATED THE CATHETER USING WARM FLUID TO THE BLADDER THROUGH THE OPENING IN THE TOP OF THE CATHETER. AT 6 MIN OF THAWING, THE CATHETER CAME OUT OF THE PATIENT. A CYSTOSCOPY WAS PERFORMED AND NO VISIBLE DAMAGE WAS OBSERVED. A NEW WARMING CATHETER WAS INSERTED AND THE SECOND FREEZE THAW CYCLE WAS PERFORMED. THE FACILITY WOULD NOT ALLOW THE REMOVAL OF THE CONTAMINATED WARMER. DEVICE NOT RETURNED FOR TESTING. NO RESPONSE TO REQUEST FOR PATIENT UPDATE. | 0
1835897 | CUSTOMER REPORTED POSSIBLE CARRYOVER DURING ANTIBODY SCREEN TESTING. A RETROSPECTIVE REVIEW INDICATES THAT A PRENATAL PATIENT HAD ANTI-D (128) AND ANTI-C (2). CARRY OVER WAS SEEN WITH ANTI-D ON BOTH RH POSITIVE CELLS. NO ERRONEOUS RESULTS REPORTED ON ANY SAMPLES. OCD FIELD ENGINEERS PERFORMED REQUIRED REPAIRS TO RETURN THE INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
1846036 | UPON INITIAL PRETESTING, THE ENTIRE SHAFT FROZE. NO PATIENT CONTACT. | 0
1846042 | THE CUSTOMER REPORTED THAT POWERSCRIBE AND PACS ARE LOSING SYNC AND DISPLAYING DIFFERENT PATIENT INFO BETWEEN THE TWO SYSTEMS. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. INVESTIGATION RESULTS REVEALED THAT THE REPORTING SITE WAS OPERATING BELOW THE MINIMUM VERSION REQUIREMENTS SET BY BOTH GE AND NUANCE (POWERSCRIBE). IN ADDITION, AD-HOC EXAMS ARE VIEWED WHILE IN DICTATION MODE WHICH INCREASES THE POSSIBILITY OF THE TWO SYSTEMS LOSING SYNC WITH EACH OTHER. WHEN SYNC IS LOST, A WARNING MESSAGE IS PRINTED ON THE RA-1000 WORKLIST STATING SUCH. PACS 3.1.1.2 HAS MADE THE SYNC MUCH TIGHTER BETWEEN THE TWO TOOLS. (B)(4). | 1
1849256 | THE ACCOUNT STATED THAT THE ARCHITECT TROPONIN REAGENT HAS GENERATED FALSE ELEVATED RESULTS ON TWO PATIENT SAMPLES. THE INITIAL RESULT FOR PATIENT #1 WAS NEGATIVE AT <0.03, THE SAMPLE WAS THEN TESTED ON ANOTHER ARCHITECT ANALYZER (S/N (B)(4)) AND THE RESULT WAS POSITIVE AT 0.17 (THE ACCOUNT BELIEVES THIS IS A FALSE POSITIVE RESULT). UNITS OF MEASUREMENT WERE NOT PROVIDED. THE PATIENT'S CUT-OFF FOR THE TROPONIN ASSAY IS <0.03. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). THE SUSPECT MEDICAL DEVICE CHANGED FROM THE ARCHITECT STAT TROPONIN-I REAGENT (MANUFACTURING SITE (B)(4)) TO THE ARCHITECT I2000SR ANALYZER (MANUFACTURING SITE (B)(4)). ARCHITECT SYSTEM (B)(4) GENERATED DISCREPANT HIGH TROPONIN RESULTS COMPARED TO RESULTS GENERATED ON THE CUSTOMER'S OTHER ARCHITECT SYSTEM. PERFORMANCE IMPROVED AFTER FIELD SERVICE CLEANED THE PROCESS PATH AND COMPONENTS, OPTICS AND SHUTTER, BUT DISCREPANT RESULTS RECURRED. THEN FIELD SERVICE REMOVED THE WASH ZONE ASSEMBLIES DUE TO CONTINUAL LEAKAGE ISSUE AND INSTALLED THE NEW STYLE 109 WASH ZONE 1 AND 2 MANIFOLDS AND VALVES, AND ALSO STRIPPED DOWN AND CLEANED THE PROCESS PATH, AFTER WHICH NO FURTHER REPORTS OF DISCREPANT RESULTS HAVE BEEN RECEIVED. AN EVALUATION WAS PERFORMED IN RESPONSE TO THE DISCREPANT HIGH TROPONIN RESULTS GENERATED ON ARCHITECT (B)(4). THE EVALUATION INCLUDED A REVIEW OF THE INFORMATION PROVIDED BY THE CUSTOMER, A REVIEW OF COMPLAINT AND TRENDING DATA, AND A REVIEW OF THE ARCHITECT OPERATIONS MANUAL AND STAT TROPONIN-I PACKAGE INSERT. A REVIEW OF I2000SR COMPLAINT AND TRENDING DATA DID NOT IDENTIFY A PRODUCT ISSUE OR ADVERSE TREND ASSOCIATED WITH DISCREPANT TROPONIN-I RESULTS. THE REVIEW DETERMINED THE CURRENT RATE OF ERRATIC COMPLAINTS AND OCCURRENCES PER MILLION TESTS FOR THE ARCHITECT I2000SR FALLS BELOW THE INTERNAL RATE DOCUMENTED AT LAUNCH FOR THE ARCHITECT I2000. LABELING IN THE STAT TROPONIN-I PACKAGE INSERT IS ADEQUATE WITH REGARD TO TESTING, CONTROL AND CALIBRATION REQUIREMENTS AND RESULT INTERPRETATION. LABELING IN THE ARCHITECT SYSTEM OPERATIONS MANUAL IS ADEQUATE WITH REGARD TO MAINTENANCE AND COMPONENT REPLACEMENT, AND TROUBLESHOOTING. BASED ON THE AVAILABLE INFORMATION, A SPECIFIC CAUSE FOR THE DISCREPANT TROPONIN RESULTS WAS NOT IDENTIFIED. PROBABLE CAUSES INCLUDE ISSUES WITH THE PROCESS PATH AND/OR THE WASH ZONE 1 AND 2 MANIFOLDS AND VALVES, WHICH WERE BOTH ADDRESSED ON SITE BY FIELD SERVICE. A DEFICIENCY WAS NOT IDENTIFIED.><(B)(4). CONCOMITANT MEDICAL PRODUCTS: TROPONIN REAGENT LIST 2K41-38, LOT 42858UN10. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1849632 | SPACELABS RECEIVED A REPORT THAT THE WIRELESS MONITOR RESETS WHEN ACCESSING WLAN BIOMED TAB. SPACELABS WAS ABLE TO CONFIRM THE REPORTED SYMPTOMS IN THE MONITOR. THE PROBLEM WAS TRACED TO AN ERROR IN THE WIRELESS NETWORK DRIVER - A THIRD PARTY (B)(4) PRODUCT USED IN THE MONITOR. THE DRIVER MAINTAINS A LIST OF ALL OF THE WIRELESS ACCESS POINTS WITHIN A GIVEN AREA. THE ROOT CAUSE WAS DETERMINED TO BE THE DRIVER'S IMPROPER MANAGEMENT OF THE WIRELESS ACCESS POINT LIST WHEN BEACONS FROM MORE THAN 64 DISTINCT ACCESS POINT SERVICE SET IDENTIFIER (SSIDS) COULD BE RECEIVED AT THE MONITOR'S PHYSICAL LOCATION. SPACELABS HAS BEEN ABLE TO MITIGATE THE ISSUE WITH THIS DRIVER AND IS IMPLEMENTING A FIELD ACTION TO REPLACE THE SOFTWARE IN ALL AFFECTED PRODUCT. | 1
1849640 | SPACELABS RECEIVED A REPORT THAT THE WIRELESS MONITOR RESETS WHEN ACCESSING WLAN BIOMED TAB. SPACELABS WAS ABLE TO CONFIRM THE REPORTED SYMPTOMS IN THE MONITOR. THE PROBLEM WAS TRACED TO AN ERROR IN THE WIRELESS NETWORK DRIVER - A THIRD PARTY (B)(4) PRODUCT USED IN THE MONITOR. THE DRIVER MAINTAINS A LIST OF ALL OF THE WIRELESS ACCESS POINTS WITHIN A GIVEN AREA. THE ROOT CAUSE WAS DETERMINED TO BE THE DRIVER'S IMPROPER MANAGEMENT OF THE WIRELESS ACCESS POINT LIST WHEN BEACONS FROM MORE THAN 64 DISTINCT ACCESS POINT SERVICE SET IDENTIFIER (SSIDS) COULD BE RECEIVED AT THE MONITOR'S PHYSICAL LOCATION. SPACELABS HAS BEEN ABLE TO MITIGATE THE ISSUE WITH THIS DRIVER AND IS IMPLEMENTING A FIELD ACTION TO REPLACE THE SOFTWARE IN ALL AFFECTED PRODUCT. | 1
1849650 | SPACELABS RECEIVED A REPORT THAT A WIRELESS MONITOR BLACKED OUT DURING TRANSPORT. THE CUSTOMER REPORTED THAT THE MONITOR BLACKED OUT FOR ABOUT 15-30 SECONDS WHILE TAKING A PT UP TO ICU. SPACELABS WAS ABLE TO CONFIRM THE REPORTED SYMPTOMS IN THE MONITOR. THE PROBLEM WAS TRACED TO AN ERROR IN THE WIRELESS NETWORK DRIVER - A THIRD PARTY (B)(4) PRODUCT USED IN THE MONITOR. THE DRIVER MAINTAINS A LIST OF ALL OF THE WIRELESS ACCESS POINTS WITHIN A GIVEN AREA. THE ROOT CAUSE WAS DETERMINED TO BE THE DRIVER'S IMPROPER MGMT OF THE WIRELESS ACCESS POINT LIST WHEN BEACONS FROM MORE THAN 64 DISTINCT ACCESS POINT SERVICE SET IDENTIFIERS (SSIDS) COULD BE RECEIVED AT THE MONITOR'S PHYSICAL LOCATION. SPACELABS HAS BEEN ABLE TO MITIGATE THE ISSUE WITH THIS DRIVER AND IS IMPLEMENTING A FIELD ACTION TO REPLACE THE SOFTWARE IN ALL AFFECTED PRODUCTS. | 1
1849655 | SPACELABS RECEIVED A REPORT THAT THE WIRELESS MONITOR RESETS WHEN ACCESSING WLAN BIOMED TAB. SPACELABS WAS ABLE TO CONFIRM THE REPORTED SYMPTOMS IN THE MONITOR. THE PROBLEM WAS TRACED TO AN ERROR IN THE WIRELESS NETWORK DRIVER - A THIRD PARTY (B)(4) PRODUCT USED IN THE MONITOR. THE DRIVER MAINTAINS A LIST OF ALL OF THE WIRELESS ACCESS POINTS WITHIN A GIVEN AREA. THE ROOT CAUSE WAS DETERMINED TO BE THE DRIVER'S IMPROPER MGMT OF THE WIRELESS ACCESS POINT LIST WHEN BEACONS FROM MORE THAN 64 DISTINCT ACCESS POINT SERVICE SET IDENTIFIER (SSIDS) COULD BE RECEIVED AT THE MONITOR'S PHYSICAL LOCATION. SPACELABS HAS BEEN ABLE TO MITIGATE THE ISSUE WITH THIS DRIVER AND IS IMPLEMENTING A FIELD ACTION TO REPLACE THE SOFTWARE IN ALL AFFECTED PRODUCT. NOTE: THIS REPORT IS BEING SUBMITTED AS A RESULT OF A REEVALUATION OF PAST COMPLAINTS BY SPACELABS. ALTHOUGH A REPORT HAD NOT PREVIOUSLY BEEN FILED FOR THIS INCIDENT, WE HAVE CONCLUDED THAT A REPORT IS APPROPRIATE. WE ARE NOW REPORTING THIS INCIDENT IN ACCORDANCE WITH 21 CFR 803. | 1
1861240 | IT HAS BEEN DETERMINED THAT AN ERROR EXISTS IN THE RISK MITIGATION FOR THE FOLLOWING CAUSE: RECIPIENT HARMED BY UNSUITABLE UNIT (RHZM01). MITIGATION IS LOST WHEN DATA FROM THE MOBILE APPLICATION IS MERGED INTO THE PRIMARY DATABASE WHEN THE FOLLOWING CONDITION OCCURS: THE APPLICATION IS CONFIGURED TO ALLOW THE BLOOD CENTER STAFF TO MANUALLY MATCH DONORS IDENTIFIED IN ERROR AS NEW IN THE FIELD TO DONORS ALREADY EXISTING IN THE PRIMARY DATABASE. THE USER IS NOT BEING CORRECTLY NOTIFIED WHEN THE DONOR IS CURRENTLY DEFERRED. ADD'L LOT#S: 5.1M, 5.1.1M. THE RISK MITIGATION HAS BEEN REPAIRED BY: CONFIRMATION WITH ALL CLIENTS THAT THEY HAVE RECEIVED THE NOTIFICATION OF THE PRODUCT FAILURE. CONFIRMING WITH ALL CLIENTS THAT THEY HAVE RECEIVED THREE CLEARLY DEFINED CONFIGURATION OPTIONS THAT WILL REPAIR THE MITIGATION. REPAIR OPTION TITLES: ELIMINATE THE INTERACTIVE PORTION OF THE MOBILE UPLOAD PROCESS; RETAIN THE INTERACTIVE PORTION OF THE MOBILE UPLOAD PROCESS - REGISTER ALL NEW DONORS AS NEW; RETAIN THE INTERACTIVE PORTION OF THE MOBILE UPLOAD PROCESS - CONTINUE MATCHING NEW DONORS TO EXISTING DONORS. THE DETAILS AS TO HOW TO CONFIGURE THE DEVICE FOR EACH OPTION HAVE BEEN DISTRIBUTED TO EACH CLIENT. EVAL SUMMARY: DEVICE - BBCS HAS EVALUATED THE DEVICE AND DETERMINED THAT THE PROGRAM IS NOT WORKING AS DESIGNED AND REQUIRES MODIFICATIONS WHICH WILL OCCUR WITH THE NEXT RELEASE OF THE PROGRAM. ALL CLIENTS HAVE BEEN PROVIDED WORKAROUNDS THAT FULLY MITIGATE THE IDENTIFIED RISK. LABELING: THE LABELING OF THE DEVICE IS ACCURATE AND WRITTEN TO CURRENT SPECS. | 1
1864934 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1865202 | THE CUSTOMER REPORTED THAT WHILE THE (B)(4) CENTRAL STATION WAS IN USE MONITORING PATIENTS ALONG WITH AMBULATORY TELEPACKS, ALL EXCEPT ONE OF THE TELEPACKS STOPPED COMMUNICATING WITH THE CENTRAL. A "COMMUNICATION LOSS" MESSAGE WAS DISPLAYED AT THE CENTRAL. NO PT INJURY WAS REPORTED. THE COMPANY REP REPLACED A SWITCH IN THE RACK ROOM. THE SYSTEM WAS TESTED TO FACTORY SPECS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. | 1
1873574 | A CUSTOMER REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED WHEN FOUR (4) CIC (CENTRAL STATIONS) REBOOTED SIMULTANEOUSLY. NO INJURY WAS REPORTED. THIS IS THE FIRST OF FOUR REPORTS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1873589 | A CUSTOMER REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED WHEN FOUR (4) CIC (CENTRAL STATIONS) REBOOTED SIMULTANEOUSLY. NO INJURY WAS REPORTED. THIS IS THE FOURTH OF FOUR REPORTS. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. THE SERIAL NUMBER OF THIS DEVICE IS UNK AT THIS TIME. THE DEVICE MANUFACTURE DATE CANNOT BE DETERMINED AT THIS TIME, WITHOUT A SERIAL NUMBER OF THE DEVICE. | 1
1879092 | THE CUSTOMER STATED, THE ARCHITECT C8000 ANALYZER GENERATED FALSELY ELEVATED CC CO2 AND ANION GAP FOR MULTIPLE PATIENT SAMPLES. THE CUSTOMER RECALIBRATED THE CO2 ASSAY AND REPEATED THE SAMPLES, SOME OF WHICH REQUIRED CORRECTED REPORTS. ONE PATIENT GENERATED AN INITIAL CO2 RESULT OF 39 MEQ/L WITH A REPEAT RESULT OF 25 MEQ/L. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). IN RESPONSE TO THIS ISSUE AN INVESTIGATION WAS INITIATED TO FURTHER EXAMINE THE CUSTOMER'S OBSERVATION. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING, A REVIEW OF THE MOST RECENT PRODUCT IMPROVEMENT TEAM CLINICAL CHEMISTRY METRICS AND A REVIEW OF THE SERVICE HISTORY FOR THE ANALYZER IN QUESTION. A REVIEW OF THE SERVICE HISTORY FOR ARCHITECT C8000 ANALYZER, (B)(4) DID NOT IDENTIFY ANY OTHER COMPLAINTS REGARDING INCIDENTS OF INCONSISTENT, ABERRANT, OR ERRATIC PATIENT OR CONTROL RESULTS. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS PROBABLE CAUSES AND CORRECTIONS FOR ERRATIC RESULTS. A DEFINITIVE CAUSE OF THE CUSTOMER'S ISSUE COULD NOT BE DETERMINED. THE SYSTEM LOGS INDICATED NO ISSUES WITH THE INSTRUMENT OR BUBBLES IN THE REAGENT. SINCE THE ISSUE OCCURRED WITH TWO INSTRUMENTS THAT SHARE THE SAME WATER SUPPLY, THE LIKELY CAUSE MAY BE AN INTERMITTENT ISSUE WITH THE WATER SUPPLY. THE CUSTOMER HAD MAINTENANCE PERFORMED ON THE WATER SUPPLY PRIOR TO THE SERVICE CALL FOR THE ARCHITECT. THE FIELD SERVICE REPRESENTATIVE INSPECT THE INSTRUMENT AND IDENTIFIED NO ISSUES. A REVIEW OF THE COMPLAINT TREND REPORTS FOR THE PERIOD OF (B)(4) 2009 TO (B)(4) 2010 INDICATED THERE WERE NO ADVERSE TRENDS ASSOCIATED WITH THE COMPLAINT ISSUE. BASED UPON THE INVESTIGATION, IT HAS BEEN DETERMINED THAT ARCHITECT C8000, LIST NUMBER 1G06-11, (B)(4), IS PERFORMING AS INTENDED. NO PRODUCT DEFICIENCY WAS IDENTIFIED.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1880530 | THE DEVICE IS INDICATED FOR USE BY HEALTH CARE PROFESSIONALS IN THE COLLECTION AND TRANSPORT TO A CLINICAL LABORATORY OF NEONATE BLOOD AS DRIED BLOOD SPOTS. THE NEONATE BLOOD IS TESTED TO SCREEN THE INFANT FOR CONGENITAL ABNORMALITIES OF METABOLISM AND OTHER CONDITIONS. THE DEVICE IS PRESENTED AS JOINED PARTS WHICH CONSTITUTE THE TOTAL DEVICE PRESENTATION; FORMS PROVIDING THE INFANTS DEMOGRAPHIC DETAILS, AND THE (B)(4) TO WHICH CAPILLARY BLOOD SPOTS ARE APPLIED. THE DEVICE IS INDICATED AS A BLOOD SPECIMEN RECEPTACLE. THE BLOOD IS TRANSPORTED TO A LABORATORY AS DRIED BLOOD SPOTS. NO OTHER INDICATIONS OR CLAIMS ARE SUPPORTED. THE COMPLAINT DESCRIBED BY THE CUSTOMER IS OF LOW ASSAYS OF BIOTINIDASE FROM BLOOD SAMPLES APPLIED TO (B)(4). THE CUSTOMER ALLEGES THAT THE LOW ASSAY IS CAUSED BY THE (B)(4) LOT W092. THE DEVICE IS INDICATED AS A BLOOD SPECIMEN RECEPTACLE. THE BLOOD IS TRANSPORTED TO A LABORATORY AS DRIED BLOOD SPOTS. NO OTHER INDICATIONS OR CLAIMS ARE SUPPORTED. THE COMPLAINT DESCRIBED BY THE CUSTOMER IS OF LOW ASSAYS OF BIOTINIDASE FROM BLOOD SAMPLES APPLIED TO (B)(4). THE CUSTOMER ALLEGES THAT THE LOW ASSAY IS CAUSED BY THE (B)(4) LOT W092. INVESTIGATION IN PROCESS. A FOLLOW UP REPORT WILL BE SUPPLIED IN 30 DAYS, ON OR BEFORE (B)(6) 2010. | 0
1881842 | THE CONNECTOR FAILED AFTER BEING CONNECTED TO THE CONTRAST INJECTOR SPRAYING THE LAB WITH CONTRAST. NO HARM OR INJURY WAS REPORTED. DEVICE EVALUATION. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION/ INVESTIGATION. THE DEVICE RECORD WAS REVIEWED AND FOUND ONE EXCEPTION DOCUMENT FOR THIS LOT. THE EXCEPTION WAS FOR ONE DEVICE FAILING PRESSURE TEST. THE LOT WAS SAMPLED AND RE-TESTED. SAMPLE DEVICES HAD 100% PASS OF SECONDARY PRESSURE TEST. NO OTHER EXCEPTIONS FOUND IN DEVICE HISTORY RECORD. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN AN EVALUATION IS COMPLETED. EVALUATION METHOD: HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
1882397 | SPACELABS RECEIVED A REPORT THAT A PT'S NUMERIC HEART RATE VALUE PRINTED ON THE MCARE MONITOR STRIP RECORDING DOES NOT MATCH THE HEART RATE OBTAINED FROM THE WAVEFORM R-R INTERVAL MEASURED ON THE SAME STRIP. THE REPORTER INDICATED THAT THE NUMERIC HEART RATE VALUE WAS THE SAME AS THE PT'S RESPIRATION RATE. SPACELABS INVESTIGATION INTO THIS MATTER IS ON-GOING. THE EQUIPMENT IS STILL IN USE AT THE HOSPITAL. NO ONE HAS BEEN INJURED AS A RESULT OF THIS ALLEGED MALFUNCTION. WE WILL PROVIDE A SUPPLEMENTAL REPORT ONCE OUR INVESTIGATION IS COMPLETE. | 1
1883413 | A CUSTOMER ALLEGED AN APNEA MONITOR FAILED TO AUDIBLY ALARM IN A VOLUNTARY MEDWATCH (B)(4) REPORT SUBMITTED TO THE FDA IN (B)(6) 2009. THE FDA FORWARDED THIS REPORT TO US (THE MANUFACTURER) ON (B)(6) 2009. THE REPORT ALLEGED NO HARM OR INJURY ASSOCIATED WITH THE PERFORMANCE OF THE APNEA MONITOR (SN (B)(4)). A REVIEW OF THE VOLUNTARY MEDWATCH REPORT (B)(4) REVEALED THAT THE DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER WHO HAD PROVIDED THE APNEA MONITOR TO THE COMPLAINANT HAD PREVIOUSLY BEEN CONTACTED BY US (THE MANUFACTURER) REGARDING A SIMILAR COMPLAINT WITH DIFFERENT APNEA MONITOR (REF. MDR 2518422-2009-00011). THIS CONTACT OCCURRED ON (B)(6) 2009. DURING THAT CONTACT THE DME INFORMED US THAT THE APNEA MONITOR SN (B)(4) HAD BEEN PLACED INTO USE IN THE COMPLAINANT'S HOME ON (B)(6) 2006, AFTER COMPLETION OF THE CHECK OUT PROCEDURE IDENTIFIED IN PRODUCT LABELING (RESPIRONICS SMART MONITOR 2 CHECKOUT PROCEDURE MANUAL, PN 1020817 REV. 2) AND EDUCATING THE CAREGIVER REGARDING THE OPERATION OF THE MONITOR. THE DME ADDITIONALLY STATED THAT THE COMPLAINANT HAD REQUESTED REPLACEMENT OF THE SUBJECT APNEA MONITOR ON (B)(6) 2006, BUT DID NOT PROVIDE AN EXPLANATION OF WHY THE COMPLAINANT REQUESTED THE REPLACEMENT. THEY ALSO (WE BELIEVE AS A RESULT OF THEIR TESTING OF THE DEVICE) DID NOT INFORM US OF ANY PRODUCT PROBLEMS AFTER THEY REMOVED THE MONITOR FROM THE COMPLAINANT'S HOME. METHOD: (REVIEW OF COMPLAINT AND SERVICE PERFORMED). RESULTS: (NO COMPLAINTS OR SERVICE RECORDS WERE IDENTIFIED). OUR REVIEW OF THE HISTORY FOR THE SUBJECT DEVICE, THE INFORMATION FOR THIS REPORT, AND OUR COMMUNICATION WITH THE DME HAS REVEALED NO EVIDENCE OF THE DEVICE MALFUNCTIONING OR CAUSING ANY HARM/INJURY. THIS LEADS US TO BELIEVE THE DEVICE OPERATED AS DESIGNED WHEN THE EVENT THAT LED TO THE CUSTOMER'S COMPLAINT ALLEGATION (THAT THE DEVICE DID NOT ALARM AS DESIGNED) OCCURRED. TO FURTHER SUBSTANTIATE OUR CONCLUSION THAT THE SUBJECT DEVICE OPERATED AS DESIGNED DURING THE REPORTED EVENT (AND EXPLAIN WHY NO COMPLAINT WAS MADE AGAINST THE DEVICE AT THE TIME OF THE EVENT) WE COMPLETED A REVIEW OF OUR COMPLAINT DATABASE FOR COMPLAINTS AGAINST THE MONITOR SN (B)(4). NO COMPLAINTS OR SERVICE RECORDS WERE RECORDED AGAINST THE SUBJECT DEVICE ON OR AROUND (B)(6) 2006, OR AT ANY TIME SINCE THE DEVICE WAS RELEASED INTO DISTRIBUTION IN (B)(6) 2003. BASED ON THIS FINDING WE CONCLUDE THAT THE APNEA MONITOR IDENTIFIED IN THIS REPORT OPERATED AS DESIGNED RELATIVE TO THE COMPLAINANT'S ALLEGATION. WE DRAW THIS CONCLUSION BASED ON THE FACT THAT PRODUCT LABELING CLEARLY AND ADEQUATELY DETAILS OPERATIONAL CHECKS, AND CARE AND HANDLING GUIDELINES; AND BECAUSE WE FIND NO EVIDENCE THAT THE MONITOR HAD EVER REQUIRED SERVICING OF ANY TYPE AS A RESULT OF FAILING ANY TESTS AND OR CHECKS, OR PERFORMING IN A MANNER THAT INDICATING SERVICING WAS REQUIRED. THE LABELING WE FIND THAT SUPPORTS OUR CONCLUSION IS INCLUDED IN "THE PROFESSIONAL MANUAL FOR THE SMART MONITOR 2 (PN 1008895, REV. B):" WHICH STATES "THIS RESPIRONICS EQUIPMENT AND THE AUTHORIZED ACCESSORIES ARE DESIGNED TO WORK AS DESCRIBED IN THE OPERATOR'S MANUAL. THE USER(S) OF THIS EQUIPMENT SHOULD NOT USE PARTS THAT HAVE FAILED, EXHIBIT EXCESSIVE WEAR, ARE CONTAMINATED, OR OTHERWISE INEFFECTIVE. THE MONITOR AND ITS ACCESSORIES SHOULD NOT BE MODIFIED" AND "THE USER OF THIS EQUIPMENT IS RESPONSIBLE FOR READING, UNDERSTANDING, AND FOLLOWING THE WARNING AND CAUTION STATEMENTS THROUGHOUT THIS MANUAL." ADDITIONAL LABELING STATES THE CAREGIVER HAS THE RESPONSIBILITY OF ENSURING THE PROPER FUNCTION OF THE APNEA MONITOR USING THE DEVICE'S FUNCTIONAL SELF-TEST CHECKS, AND THAT THESE TESTS SHOULD BE PERFORMED AT LEAST ONCE A WEEK OR ACCORDING TO THE INSTRUCTIONS GIVEN BY YOUR HEALTH CARE PROFESSIONAL. PRODUCT LABELING ALSO STATES THAT IF THE DEVICE DOES NOT PERFORM ACCORDING TO PRODUCT LABELING (INCLUDING THE SELF-TEST CHECKS) TO CONTACT THE DEALER (DME), AND FINALLY, PRODUCT LABELING STATES "DO NOT USE YOUR MONITOR IF THE ALARM SOUNDS WEAK OR DOES NOT ACTIVATE UPON INITIAL POWER ON" (REFERENCE "PARENT'S GUIDE FOR THE SMART MONITOR 2" (PN 572-4000-00). BASED ON ALL AVAILABLE DATA WE BELIEVE THAT THE APNEA MONITOR ASSOCIATED WITH THIS REPORT OPERATED AS DESIGNED RELATIVE TO THE REPORTED ALLEGATION AND THAT NO FURTHER ACTION IS APPROPRIATE. | 1
1884274 | THE CUSTOMER STATED AN ARCHITECT I2000SR ANALYZER GENERATED A FALSE POSITIVE TROPONIN RESULTS FOR ONE PATIENT SAMPLE. THE ARCHITECT ANALYZER GENERATED AN INITIAL TROPONIN RESULT OF 0.16 NMOL/L. THE PATIENT WAS REDRAWN TWO HOURS LATER AND A TROPONIN RESULT OF <0.03 NMOL/L WAS GENERATED. THE INITIAL DRAW WAS THEN REANALYZED AND A TROPONIN RESULT OF <0.03 NMOL/L WAS GENERATED. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, REVIEW OF RETURNED SYSTEM LOGS AND SERVICE HISTORY, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. THE RETURNED RESULT LOG ONLY CONTAINED DATA FOR (B)(4) 2010 TO (B)(4) 2010 AND DID NOT AID IN THE DETERMINATION OF THE CAUSE OF THIS ISSUE. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE EVALUATION NO PRODUCT DEFICIENCY WAS IDENTIFIED.><(B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1886089 | REFERENCE : MDR 1220063-2010-00040. PER TELEPHONE COMMUNICATION WITH DRAEGER MEDICAL REGARDING THE SUBMITTAL OF ADDITIONAL MDRS FOR EACH SERIAL NUMBER LISTED ON THE INITIAL MDR SUBMITTED, DRAEGER IS SUBMITTING THIS REPORT. IT WAS REPORTED THAT THE DISPLAY OF THE INFINITY DELTA PRESENTS ALEATORY DISFUNCTIONS. IT WAS ALSO REPORTED THAT ONCE THE MONITOR IS RUNNING, THE "MENU" AND "STAND BY" WINDOW SUDDENLY APPEAR AND THERE IS NO WAY TO GET OUT OF IT BY PRESSING ANY OTHER KEY UNTIL IT DISAPPEARS BY ITSELF. THERE ARE NO ERRORS FOUND (IN THE ERROR LIST) THAT COULD BE ASSOCIATED TO THIS IRREGULARITY PRESENT BY THE EQUIPMENT. THERE WAS NO PT INJURY REPORTED. (B)(4). WHEN DRAEGER INITIALLY RECEIVED THIS REPORT, IT WAS DEEMED NOT REPORTABLE. HOWEVER, THE RESULTS OF THE INVESTIGATION DETERMINED THAT THE EVENTS INVOLVED REPORTS THAT SOME OF THE MONITOR'S FIXED KEYS BECAME INOPERATIVE OR ACTIVATE WITHOUT USER INTERACTION. THE WORST CASE SCENARIO OF THIS TEMPORARY MALFUNCTION COULD BE THAT THE MONITOR MAY DISCHARGE A PT AUTOMATICALLY. THEREFORE, IT WAS DETERMINED THAT THIS COMPLAINT IS REPORTABLE. AN INVESTIGATION WAS INITIATED WITH THE SUPPLIER OF THE SWITCH PANEL. IT WAS DETERMINED THAT THIS CONDITION IS DUE TO A CONTAMINATION PRESENT IN THE FIXED KEY PANEL MATERIAL. A RISK ANALYSIS WAS PERFORMED. DRAEGER MEDICAL IS INITIATING A CORRECTIVE ACTION TO ADDRESS THIS PROBLEM. NO PT DEATHS OR INJURIES HAVE BEEN REPORTED AS A RESULT OF THIS CONDITION. | 1
1887189 | CUSTOMER REPORTED UNEXPECTED NEGATIVE RESULT ON THE GALILEO. AFTER CLEANING THE MANIFOLD ON THE INSTRUMENT, THE CUSTOMER REPEATED THE SAMPLE USING THE SAME REAGENTS. THE EXPECTED RESULTS WERE OBTAINED. | 0
1887311 | THE CUSTOMER STATES THAT ONE PATIENT SAMPLE HAS GENERATED A FALSELY ELEVATED (POSITIVE) ARCHITECT TOTAL B-HCG ASSAY RESULT OF 484.07 MIU/ML. THE SAMPLE TESTED AS "NEGATIVE" ON A NON-ABBOTT PLATFORM. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN ABBOTT FIELD SERVICE ENGINEER (FSE) INSPECTED THE ARCHITECT I2000SR ANALYZER AND FOUND DRIED YELLOWISH RESIDUE SPREAD BY THE SAMPLE PROBE AND CONCLUDED THAT THIS MAY HAVE CAUSED THE ERRATIC RESULTS. THERE WAS A POSSIBILITY OF THE RESIDUAL MATERIAL FALLING INTO THE REACTION VESSELS BELOW THE LOAD/UNLOAD DIVERTER ASSEMBLY. THE FSE PERFORMED SYSTEM MAINTENANCE AND CLEANED THE WASH ZONE MANIFOLD AND LOAD/UNLOAD DIVERTER AREA. SINCE PERFORMING THE CLEANING PROCEDURE, THE CUSTOMER HAS NOT OBSERVED ANY ADDITIONAL INITIALLY FALSE POSITIVE B-HCG ASSAY RESULTS. A REVIEW OF THE INSTRUMENT LOG WAS PERFORMED. THE RESULT LOG CONFIRMED THE B-HCG RESULTS MENTIONED BY THE CUSTOMER. SINCE THE OTHER PERTINENT INSTRUMENT LOGS WERE NOT AVAILABLE, A DEFINITE CAUSE OF THE ERRATIC B-HCG RESULTS COULD NOT BE DETERMINED. COMPLAINT ACTIVITY WAS REVIEWED AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (LIST NUMBER 201837-108) CONTAINS INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. SINCE THE OTHER PERTINENT INSTRUMENT LOGS WERE NOT AVAILABLE, A DEFINITE CAUSE OF THE ERRATIC B-HCG RESULTS COULD NOT BE DETERMINED. THE ANALYZER IS PERFORMING AS EXPECTED FOLLOWING THE FSE MAINTENANCE ACTIVITIES. A PRODUCT MALFUNCTION WAS NOT IDENTIFIED. REVIEW OF COMPLAINT TRACKING AND TRENDING.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1890064 | THE CUSTOMER STATES THAT THE PROVUE CAMERA RESULTED A REACTION IN THE GEL CARD AS NEGATIVE WHEN VISUAL INSPECTION OF THE GEL CARD WAS POSITIVE. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND CLEANED THE LENS OF THE CAMERA AND VISOR. THE FE PERFORMED THE CARD ALIGNMENT AND RAN A REFERENCE IMAGE. THE CUSTOMER RAN QC. ALL QC WITHIN ACCEPTABLE RANGE. INSTRUMENT IS OPERATING AS EXPECTED. | 0
1890833 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. THE FOLLOWING INCIDENT DESCRIPTION WAS PROVIDED TO CARIDIANBCT QA: THE CUSTOMER/OPERATOR NOTICED BEFORE RINSEBACK THAT THE REMOVED PLASMA VOLUME WAS LESS THAN DISPLAYED. PATIENT WAS FINE (BP AND PULSE NORMAL) AND A PHYSICIAN EXAMINED THE PATIENT AND FOUND HIM FINE AS WELL. THE CUSTOMER INFORMED CARIDIANBCT THAT NO MEDICAL INTERVENTION WAS REQUIRED. FIN (B)(4). PER A SERVICE CALL, NO ISSUES WERE FOUND WITH THE MACHINE AND IT IS OPERATING WITHIN DESIGN SPECIFICATIONS. INVESTIGATION: A SIMULATED RUN WAS PERFORMED ON (B)(6) 2010 USING A TPE SET. THE SET WAS PRIMED AND AN ERROR OCCURRED STATING "PLASMA VALVE NOT OPERATING CORRECTLY", TRIED CLEARING ALARM BUT ALARM WOULD NOT CLEAR UNTIL THE PLASMA LINE THAT WAS ORIGINALLY NOT INSTALLED CORRECTLY IN THE VALVE WAS REPOSITIONED. THE ALARM THEN CLEARED AND PRIME CONTINUED WITH NO FURTHER ALARMS. CARIDIANBCT INVESTIGATORS COULD NOT CREATE THE CONDITION BY MISLOADING THE PLASMA VALVE TUBING BECAUSE THE SET WOULD NOT PASS ALARM TESTS. ROOT CAUSE: ALTHOUGH THE SET WAS RETURNED FOR INVESTIGATION, NO ROOT CAUSE COULD BE DETERMINED. IT IS LIKELY THAT THE ROOT CAUSE OF THIS INCIDENT IS RELATED TO AN OCCLUSION IN THE PLASMA COLLECT TUBING LINE THAT SELF-CORRECTED DURING THE PROCEDURE AS EVIDENCED BY SOME PLASMA WASTE BEING COLLECTED. IT IS UNCONFIRMED WHETHER THIS BLOCKAGE WAS PRODUCED BY A MFG DEFECT OF THE COBE SPECTRA DISPOSABLE OR BY THE WAY THE DISPOSABLE WAS LOADED AND HANDLED AT THE CUSTOMER SITE. SAFETY RISK ASSESSMENT: POST-PROCEDURE, THE PATIENT WAS WELL WITHOUT MEDICAL INTERVENTION TO TREAT THE HYPERVOLEMIA. | 0
1900246 | THE ROTATOR BROKE WHILE INJECTING CONTRAST. NO HARM OR INJURY REPORTED. DEVICE EVAL: THE DEVICE EVAL HAS NOT BEEN COMPLETED. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. EVAL METHOD - OTHER, DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS - OTHER, A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVAL HAS BEEN COMPLETED. | 0
1904620 | THE CUSTOMER STATED SEVERAL FALSELY ELEVATED CHLORIDE RESULTS WERE GENERATED ON THE ARCHITECT C8000 ANALYZER. THE CUSTOMER INDICATED THE ICT REFERENCE SOLUTION WAS THOUGHT TO BE THE CAUSE AS THE ISSUE OCCURRED RIGHT AFTER A NEW LOT OF SOLUTION WAS PLACED ON THE ARCHITECT. PATIENT 4 GENERATED AN INITIAL CHLORIDE OF 116 MEQ/L WITH A REPEAT OF 107 MEQ/L. THE INCORRECT RESULT WERE NOT REPORTED OUT OF THE LABORATORY. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). IN RESPONSE TO THIS ISSUE AN INVESTIGATION WAS INITIATED TO FURTHER EXAMINE THE CUSTOMER'S OBSERVATION. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND A REVIEW OF THE ARCHITECT C8000 SYSTEM LOGS. A REVIEW OF THE ARCHITECT C8000 SYSTEM LOGS CONFIRMED THE ELEVATED CHLORIDE RESULTS. ADDITIONALLY, A REVIEW OF THE SYSTEM MAINTENANCE LOGS REVEALED THE "FLUSH BULK SOLUTIONS" PROCEDURE HAD BEEN PERFORMED EARLY ON THE DAY THE DISCREPANT RESULTS WERE GENERATED AND AGAIN AFTER THE RESULTS WERE GENERATED. A POSSIBLE CAUSE FOR THE ISSUE IS THAT THE ICT MODULE MAY HAVE CONTAINED A BUBBLE AND/OR WAS CONTAMINATED AND THE BULK SOLUTION FLUSH MAY HAVE RESOLVED THE ISSUE. THE INSTRUMENT HISTORY LOG INDICATED AN "ICT REFERENCE SOLUTION INVENTORY LOW" ERROR WAS GENERATED, SUGGESTING THE ICT SOLUTION WAS INSTALLED JUST PRIOR TO THE DISCREPANT CHLORIDE RESULTS. A REVIEW OF COMPLAINT TRACKING AND TRENDING FOR THE PERIOD OF (B)(4) 2010 THROUGH (B)(4) 2010 DID NOT REVEAL ANY ISSUES OR ADVERSE TRENDS RELATED TO THE ICT REFERENCE SOLUTION LOT 37992UN10. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS PROBABLE CAUSES AND CORRECTIONS FOR ERRATIC RESULTS, AS WELL AS, GUIDELINES CONCERNING THE AUTOMATIC PROCESSING MODULE, INCLUDING PROPER REPLACEMENT OF BULK SOLUTIONS. A REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT REVEAL ANY ADDITIONAL COMPLAINTS FOR ELEVATED OR ERRATIC CHLORIDE RESULTS FOR ARCH C8000 (B)(4). BASED UPON THE INVESTIGATION AND THE INFORMATION AVAILABLE IT WAS DETERMINED THAT THE ARCHITECT C8000, LIST NUMBER 1G06-01, (B)(4) AND THE ICT REFERENCE SOLUTION, LOT NUMBER 37992UN10 ARE PERFORMING AS INTENDED. NO PRODUCT DEFICIENCY WAS IDENTIFIED.><(B)(4). CONCOMITANT MEDICAL PRODUCTS: ICT REF SOL, LIST NUMBER 1E49-20, LOT NUMBER 37992UN10. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1904657 | THE CUSTOMER STATED SEVERAL FALSELY ELEVATED CHLORIDE RESULTS WERE GENERATED ON THE ARCHITECT C8000 ANALYZER. THE CUSTOMER INDICATED THE ICT REFERENCE SOLUTION WAS THOUGHT TO BE THE CAUSE AS THE ISSUE OCCURRED RIGHT AFTER A NEW LOT OF SOLUTION WAS PLACED ON THE ARCHITECT. PATIENT 5 GENERATED AN INITIAL CHLORIDE OF 111 MEQ/L WITH A REPEAT OF 102 MEQ/L. THE INCORRECT RESULTS WERE NOT REPORTED OUT OF THE LABORATORY. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). IN RESPONSE TO THIS ISSUE AN INVESTIGATION WAS INITIATED TO FURTHER EXAMINE THE CUSTOMER'S OBSERVATION. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND A REVIEW OF THE ARCHITECT C8000 SYSTEM LOGS. A REVIEW OF THE ARCHITECT C8000 SYSTEM LOGS CONFIRMED THE ELEVATED CHLORIDE RESULTS. ADDITIONALLY, A REVIEW OF THE SYSTEM MAINTENANCE LOGS REVEALED THE "FLUSH BULK SOLUTIONS" PROCEDURE HAD BEEN PERFORMED EARLY ON THE DAY THE DISCREPANT RESULTS WERE GENERATED AND AGAIN AFTER THE RESULTS WERE GENERATED. A POSSIBLE CAUSE FOR THE ISSUE IS THAT THE ICT MODULE MAY HAVE CONTAINED A BUBBLE AND/OR WAS CONTAMINATED AND THE BULK SOLUTION FLUSH MAY HAVE RESOLVED THE ISSUE. THE INSTRUMENT HISTORY LOG INDICATED AN "ICT REFERENCE SOLUTION INVENTORY LOW" ERROR WAS GENERATED, SUGGESTING THE ICT SOLUTION WAS INSTALLED JUST PRIOR TO THE DISCREPANT CHLORIDE RESULTS. A REVIEW OF COMPLAINT TRACKING AND TRENDING FOR THE PERIOD OF (B)(4) 2010 THROUGH (B)(4) 2010 DID NOT REVEAL ANY ISSUES OR ADVERSE TRENDS RELATED TO THE ICT REFERENCE SOLUTION LOT 37992UN10. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS PROBABLE CAUSES AND CORRECTIONS FOR ERRATIC RESULTS, AS WELL AS, GUIDELINES CONCERNING THE AUTOMATIC PROCESSING MODULE, INCLUDING PROPER REPLACEMENT OF BULK SOLUTIONS. A REVIEW OF THE INSTRUMENT SERVICE HISTORY DID NOT REVEAL ANY ADDITIONAL COMPLAINTS FOR ELEVATED OR ERRATIC CHLORIDE RESULTS FOR ARCH C8000 (B)(4). BASED UPON THE INVESTIGATION AND THE INFORMATION AVAILABLE IT WAS DETERMINED THAT THE ARCHITECT C8000, LIST NUMBER 1G06-01, (B)(4) AND THE ICT REFERENCE SOLUTION, LOT NUMBER 37992UN10 ARE PERFORMING AS INTENDED. NO PRODUCT DEFICIENCY WAS IDENTIFIED.><(B)(4). CONCOMITANT MEDICAL PRODUCTS: ICT REF SOL, LIST NUMBER 1E49-20, LOT NUMBER 37992UN10. AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1904711 | THE CUSTOMER STATED SMOKE WAS OBSERVED COMING FROM VALVE 2 ON AN ARCHITECT I2000 ANALYZER. THE CUSTOMER OBSERVED A BURNT SMELL. THE ARM OF THE ANALYZER HAD STOPPED FUNCTIONING, AND A LEAK AND SMOKE AT VALVE 2 WAS OBSERVED. THE ARM WAS REPLACED TO RESOLVE THE ISSUE. THERE WAS NO HARM TO PERSONNEL, OR ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. THE AUTOMATIC RECONSTITUTION MODULE (ARM) WAS SCHEDULED TO BE REPLACED . THE ARM IS A BACK UP FOR THIS ANALYZER. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND FOR SMOKE FROM THE V2 VALVE ON THE ARM OF THE ARCHITECT I2000 ANALYZER. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND SAFETY ANALYSIS MEMO WERE FOUND TO ADEQUATELY ADDRESS ELECTRICAL HAZARDS. BASED ON THE EVALUATION, A SYSTEMIC ISSUE WITH THE ARCHITECT SYSTEM WAS NOT IDENTIFIED. | 0
1905064 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO OBTAINING A BLOOD SAMPLE TYPED AS O NEGATIVE ON THE PK7200 AUTOMATED MICROPLATE SYSTEM. TEST RESULTS WERE REPORTED BY (B)(6) TO THE BLOOD CENTER FOR WHICH THE TESTING WAS PERFORMED AND IDENTIFIED BY THAT CLIENT AS DISCREPANT WITH THE DONOR HISTORY PRIOR TO DISTRIBUTION OF THE BLOOD COMPONENTS. MANUEL TUBE TESTING BY IBC AND BCI GAVE POSITIVE ANTIHUMAN GLOBULIN (AHG) (WEAK-D TEST) RESULT. MANUAL TESTING BY BCI AND THE REAGENT MANUFACTURER SUGGESTS THAT THE RH D ANTIGEN ON THE SAMPLE RED BLOOD CELLS IS A CATEGORY VI VARIANT. THE ORIGINAL SAMPLE WAS TESTED AT BCI AS RH POSITIVE ON THE PK7200 USING THE SAME LOTS OF REAGENTS AS USED BY IBC. NO PATIENTS WERE HARMED OR INVOLVE WITH THIS INCIDENT. NO SAMPLE-RELATED ISSUES WERE IDENTIFIED. VOLUME VERIFICATION RECORDS AND REVIEW OF DATA LOGS DO NOT SUGGEST ANY EVIDENCE OF ANALYZER MALFUNCTION. BASED ON THE REAGENT MANUFACTURER INVESTIGATION RECORDS: TESTING, DATA & BATCH RECORD REVIEW THE REAGENTS WERE MANUFACTURED APPROPRIATELY AND ARE PERFORMING AS EXPECTED. ROOT CAUSE IS UNDETERMINED FOR THIS EVENT. | 0
1905866 | THIS ISSUE WAS INITIALLY RECEIVED AS AN INQUIRY. UPON FURTHER REVIEW DURING AN AUDIT, IT WAS DETERMINED ON (B)(6) 2010 THAT THIS SHOULD BE TREATED AS A COMPLAINT. CARIDIANBCT RECEIVED AN EMAIL FROM THE PERSONAL EMAIL ACCOUNT OF A STAFF NURSE AT (B)(6) HOSPITAL ASKING ABOUT HYPOTENSION AFTER A THERAPEUTIC PLASMA EXCHANGE PROCEDURE. THE NURSE STATED THAT ONE PT WENT INTO PULMONARY EDEMA AND EXHIBITED HYPOTENSION AFTER EACH OF FOUR TPE PROCEDURES. THE INITIAL COMMUNICATION MENTIONED SEPTICEMIA, BUT SUBSEQUENT COMMUNICATION INDICATED THAT THE PT ACQUIRED THE INFECTION PRIOR TO THE TPE PROCEDURES. THE NURSE REQUESTED THAT THE COMPANY NOT DISCLOSE THE CONTACT WITH THE COMPANY. FIN (B)(4). THE COMPLAINANT PROVIDED THE FOLLOWING PROCEDURAL INFO: AC USED 655ML, REPLACEMENT -3732, REMOVAL- 4495, PLASMA REMOVED - 3874MLAC IN REMOVE BAG - 598ML, AC TO PT - 57ML, PLASMA VOLUME EXCHANGED - 1.5. BASED UPON THIS INFO AND THE PT'S CLINICAL STATUS, THE CARIDIANBCT SR. PRODUCT SUPPORT AND TRAINING SPECIALIST CONCLUDED THAT PROBABLE CAUSE OF THE HYPOTENSION WAS RELATED TO THE PT'S CLINICAL STATUS AND NOT THE CARIDIANBCT DEVICE. FURTHER ATTEMPTS WILL BE MADE TO GAIN ADDITIONAL INFO AND A SUPPLEMENT WILL BE PROVIDED IF ADDITIONAL INFO IS OBTAINED. | 0
1911076 | THE CUSTOMER STATED AN ACCELERATOR APS ANALYZER MISREAD ONE BARCODE SAMPLE. THE ACCELERATOR PROPERLY READ SAMPLE ID (B)(6), BUT WHEN A DUPLICATE TUBE WITH THE SAME ID WAS PLACED ON THE TRACK THE ACCELERATOR MISREAD THE ID AS (B)(6). THE ACCELERATOR DID NOT PROCESS THIS TUBE, AND GENERATED A "NO LIS DATA" ERROR. THE (B)(6) PORTION OF THE ID REPRESENTS THE JULIAN DATE, AND THEREFORE NO ORDERS EXISTED FOR ID (B)(6). THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. ABBOTT FIELD SERVICE ADJUSTED THE BARCODE READER TO RESOLVE THE ISSUE. A FAILURE INVESTIGATION WAS OPENED TO INVESTIGATE THE ISSUE OF BAR CODE LABELS BEING MISREAD AS 1-2 DIGIT SIDS AND CHARACTER SUBSTITUTIONS. THE FREQUENCY OF SIMILAR COMPLAINTS WAS FOUNT TO BE THE PREDICTED FREQUENCY AS STATED IN THE INPECO (OEM) WORKCELL RISK ANALYSIS, AND NO ADVERSE TREND EXISTS. LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS THE COMPLAINT ISSUE. A TECHNICAL SERVICE BULLETIN WAS ISSUED TO ENSURE THE PARAMETERS FOR THE BARCODE READERS ON THE INPUT OUTPUT CENTRIFUGE AND I2000SR INSTRUMENT MODULE ARE UNIFORM AND THAT THEY ARE OPTIMIZED FOR THE SAMPLE BARCODES USED BY THE CUSTOMER.><(B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
1912391 | THE CUSTOMER STATES THAT WHEN PROCESSING PATIENT SAMPLES USING THE ARCHITECT BHCG ASSAY, RESULTS OBTAINED ARE NOT REPRODUCIBLE. ONE PATIENT SAMPLE GENERATED A FALSELY POSITIVE BHCG RESULT OF 31 MIU/ML. WHEN THE SAMPLE WAS REPEATED, A NEGATIVE RESULT OF <1.2 MIU/ML WAS OBTAINED. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. (B)(4). THE CUSTOMER INDICATED CALIBRATION AND CONTROL PERFORMANCE FOR BHCG WERE ACCEPTABLE. TESTING WAS PERFORMED AT THE CUSTOMER SITE, AND INDICATED THAT CARRY-OVER OCCURRED WHEN TESTING NEGATIVE SAMPLES SUBSEQUENTLY TO POSITIVE SAMPLES IMPACTING THE NEGATIVE SAMPLE RESULTS TO BE FALSELY ELEVATED. THE PROBE HAD BEEN ON THE SYSTEM APPROXIMATELY A MONTH. AFTER THE PROBE WAS CHANGED, AND CALIBRATED, ADDITIONAL TESTING INDICATED NO CARRYOVER OCCURRED. REVIEW OF DOCUMENTATION IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. A DEFICIENCY WAS IDENTIFIED; POSSIBLY DUE TO SAMPLE OR REAGENT CARRY-OVER. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1912995 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYS FAILED. NO PT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SVC TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1913064 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
1915321 | THE ROTATOR BROKE DURING A LEFT VENTRICULOGRAM PROCEDURE. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED THREE (3) DEFECTIVE DEVICES BUT HAS NOT PROVIDED ANY ADDITIONAL INFORMATION OR CLINICAL DETAILS FOR THE ADDITIONAL EVENTS. THE CUSTOMER RETURNED ONE DEVICE. THEREFORE, THIS SINGLE REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVALUATION: THE EVALUATION/INVESTIGATION IS NOT COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETED. EVALUATION: CONCLUSIONS - A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED. | 0
1916415 | THE CUSTOMER STATED THAT THE ARCHITECT ANALYZER GENERATED DISCREPANT CHLORIDE RESULTS FROM ONE PATIENT SAMPLE. AN INITIAL RESULT OF 124 MEQ/L WAS REPEATED AT 111 MEQ/L. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). ICT MODULE, BULK SOLUTION AND CLEANING SOLUTION. THE CUSTOMER REPLACED THE ICT REFERENCE SOLUTION, THE ICT SAMPLE DILUENT AND FLUSHED THE ICT CUP WHICH RESOLVED THE ICT CALIBRATION ISSUE BUT DID NOT RESOLVE THE PROBLEM OF POOR PRECISION. SODIUM CONTROL RESULTS WERE LOW AND CHLORIDE CONTROL RESULTS WERE HIGH. AN ABBOTT FIELD SERVICE ENGINEER INSPECTED THE INSTRUMENT AND REPLACED THE ICT MODULE, FLUSHED THE BULK SOLUTION TUBING ALONG WITH THE ICT CUP AND REPLACED THE ICT REFERENCE AND ICT CLEANING SOLUTIONS. THE FSE CONFIRMED THE INSTRUMENT IS MEETING SPECIFICATION AFTER THE SERVICE WAS COMPLETED. A COMPLAINT SEARCH WAS PERFORMED FOR THE TIME PERIOD OF (B)(4) 2010 THROUGH (B)(4) 2010 WITH NO ADDITIONAL COMPLAINTS RELATED TO THIS INSTRUMENT WERE FOUND. NO PRODUCT DEFICIENCY WAS IDENTIFIED AS A RESULT OF THIS EVALUATION, HOWEVER, THE CUSTOMER WAS REFERRED TO THE SYSTEM OPERATIONS MANUAL AND THE ICT PACKAGE INSERT WHERE SUCH AN ISSUE IS DISCUSSED ALONG WITH PREVENTIVE AND CORRECTIVE ACTIONS.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1917079 | THE ACCOUNT GENERATED A FALSELY DEPRESSED HEMOGLOBIN RESULT (5 G/DL) THAT REPEATED IN THE NORMAL RANGE WHEN PROCESSING ON THE CELL-DYN 1800 ANALYZER. THE ACCOUNT STATED ALL THE PARAMETERS ON THE PATIENT SAMPLE WERE EXTREMELY LOW WITH NO ERRORS OR FLAGS. NO SPECIFIC TESTING DATA WAS PROVIDED. THE ACCOUNT REQUESTED SERVICE FOR THE CELL-DYN 1800 ANALYZER. THE FALSELY DEPRESSED HEMOGLOBIN RESULT WAS NOT REPORTED OUTSIDE OF THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). EVALUATION - FIELD SERVICE REPLACED A SAMPLE ASPIRATION PROBE AND SAMPLE SYRINGE. IN RESPONSE TO THIS ISSUE, AN INVESTIGATION WAS INITIATED TO FURTHER EXAMINE THE CUSTOMER'S OBSERVATION. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS AND A REVIEW OF LABELING. THE CUSTOMER HAD REPORTED ERRATIC LOW PATIENT RESULTS GENERATED WITH THE CELL-DYN 1800 INSTRUMENT. A FIELD SERVICE REPRESENTATIVE (FSR) FOUND THAT THE ERRATIC RESULTS WERE DUE TO A LEAKING SAMPLE ASPIRATION PROBE AND THE SAMPLE SYRINGE; BOTH PARTS WERE REPLACED. THE FSR SUCCESSFULLY PERFORMED PRECISION ON THE CELL-DYN. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS PROBABLE CAUSES AND CORRECTIONS FOR ERRATIC RESULTS, AS WELL AS, INFORMATION ON THE MAINTENANCE SCHEDULE FOR REPLACING THE SAMPLE ASPIRATION PROBE AND SAMPLE SYRINGE. A REVIEW OF THE COMPLAINT TREND REPORTS FOR THE PERIOD OF (B)(4) 2010 THROUGH (B)(4) 2010 FOR THE CELL-DYN ANALYZER, INDICATED THERE WERE NO ADVERSE TRENDS ASSOCIATED WITH THE CELL-DYN FOR THIS ISSUE. BASED ON THE INFORMATION AVAILABLE AND THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 1800 RELATED TO THIS ISSUE.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
1919554 | IT WAS REPORTED THAT THE SUPPORT FOR THE VITAL SIGNS ENGINE IS NOT SECURE. THERE WAS NO PT INJURY REPORTED. (B)(6). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
1930116 | WE RECEIVED A REPORT INDICATING THAT SOME OF THE CISCO IP PHONES INTERMITTENTLY FAILED TO DISPLAY MESSAGES, INCLUDING ALARMS AND ALERTS. | 1
1930244 | DURING TESTING, THE TECH'S AND DOCTOR NOTICED THAT THE FROST WAS COMING ALL THE WAY UP THE PROBE SHAFT. NO PT CONTACT, PROBE NOT USED. NO PT INJURY REPORTED. | 0
1930404 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED WHEN INPUTTING DATA INTO TLINK DATABASE, THE SYSTEM DISPLAYED A YELLOW LIGHT. THE USER OBSERVED THE "RS232" MODULE LIGHT WAS GREEN, HOWEVER, IT WAS WARM TO THE TOUCH AND NOTICE AN ELECTRICAL BURNING SMELL. AN ALTERNATE DEVICE WAS EMPLOYED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY AND THERE WERE NO ADVERSE CONSEQUENCES TO A PT AS A RESULT OF THIS EVENT. | 1
1930811 | DURING AN ISCHEMIC VT ABLATION PROCEDURE, THE PHYSICIAN NOTICED 3 STEAM POPS DURING THE PROCEDURE. A TAMPONADE WAS DISCOVERED AFTER THE PROCEDURE WAS FINISHED. FURTHER INFORMATION STATED THAT AFTER THE PROCEDURE, THE PATIENT WAS SENT TO ICU. THIRTY MINUTES LATER STAFF REPORTED SIGNS OF TAMPONADE, AND SURGERY WAS PERFORMED TO REPAIR THE RUPTURED VENTRICLE . A 800ML OF BLOOD WAS DRAINED AND NO FURTHER PROBLEMS WERE NOTED AFTER SURGERY AND PATIENT RECOVERED COMPLETELY. (B)(4). IT WAS REPORTED THAT THE PHYSICIAN NOTICED 3 STEAM POPS DURING THE PROCEDURE. BWI FIELD SERVICE ENGINEER CONFIRMED THAT THE ISSUE IS NOT RELATED TO THE CARTO 3 SYSTEM. CARTO 3 SYSTEM WAS OPERATIONAL AFTER THE EVENT AND DID NOT REQUIRE SYSTEM TEST OR SERVICE. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THE SYSTEM. THE CARTO 3 SYSTEM IS A NON-INVASIVE MEDICAL DEVICE AND THEREFORE, COULD NOT CAUSE STEAM POP DIRECTLY.><INVESTIGATION IS STILL IN PROGRESS. A 3500A SUPPLEMENTAL WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED.THE EP LAB STAFF WAS CONTACTED BY THE BWI FIELD SERVICE ENGINEER. THE EP LAB STAFF STATED THAT BWI EQUIPMENT DID NOT MALFUNCTION. THE ADVERSE EVENT WAS NOT CAUSED BY ANY BWI EQUIPMENT.CONCOMMITANT PRODUCT USED DURING THE PROCEDURE:THERMOCOOL SF NAV (PRODUCT IS NOT MARKETED IN THE USA) -- PRODUCT WAS DISCARDED/NOT RETURNED FOR INVESTIGATION.MODEL # UNKNOWNLOT # UNKNOWN(B)(4) | 0
1932202 | THE ROTATOR BREAKS WHEN INJECTING CONTRAST. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE DEVICES BUT DID NOT PROVIDE INFO OR CLINICAL DETAILS FOR THE ADDITIONAL EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL. THE SUSPECT DEVICE WILL NOT BE RETURNED FOR EVAL/INVESTIGATION. THE CUSTOMER RETURNED NEW UN-OPENED DEVICES FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVAL IS COMPLETED. EVAL: METHOD: DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVAL IS COMPLETED. | 0
1932242 | ABNORMAL CELLS FOUND OUTSIDE 22 FIELDS OF VIEW (FOV). NO TRIGGER CELLS PRESENT WITHIN 22 FOV TO PROMPT A FULL SCAN OF THE SLIDE. NO DELAY IN PT DIAGNOSIS AS THE ABNORMAL CELLS WERE FOUND DURING QC. SITE WILL BE MONITORED TO DETERMINE IF EXPECTED OCCURRENCES (AS DETERMINED BY THE PIVOTAL STUDY) ARE EXCEEDED. | 0
1934310 | THE CUSTOMER STATED THAT ONE PATIENT SAMPLE GENERATED A DISCREPANT RESULT FOR THE ARCHITECT CMV-IGG AVIDITY INDEX. A RESULT OF 66.2% WAS REPEATED AT 31% USING A DIFFERENT NON-ABBOTT METHOD (VIDAS). THE PATIENT WAS RETESTED FOR THE CMV AVIDITY ON THE ARCHITECT ANALYZER WITH A RESULT OF 66.6%. THE PATIENT WAS THEN TESTED FOR THE ARCHITECT CMV-IGM WITH A RESULT OF 14.71 S/CO. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. THE EVALUATION INCLUDED A REVIEW OF THE INFORMATION PROVIDED BY THE CUSTOMER, A REVIEW OF THIS INSTRUMENT ((B)(4)) SERVICE HISTORY, A REVIEW OF THE ARCHITECT SYSTEM OPERATIONS MANUAL AND A REVIEW OF THE ARCHITECT CMV IGG AVIDITY ASSAY PACKAGE INSERT. NO SPECIFIC CAUSE WAS IDENTIFIED BEHIND THIS ISSUE. HOWEVER, REVIEWING THE QUALITY CONTROLS REPORTS REVEALED THAT THE QUALITY CONTROL LEVEL NAMED CONT 1 HAD 2 RESULTS THAT WERE FLAGGED WITH THE (B)(6). THE OPERATIONS MANUAL LISTED MULTIPLE CAUSES ALONG WITH THE CORRECTIVE ACTIONS FOR ERRATIC ASSAY RESULTS. NO ADVERSE TRENDS WERE IDENTIFIED RELATED TO THIS ISSUE. THE INSTRUMENT SERVICE HISTORY REVIEW SHOWED THAT A SUBSEQUENT VISIT BY THE FIELD SERVICE ADDRESSED THE ISSUE AND ENDED UP WITH THE FIELD SERVICE REPLACING SEVERAL INSTRUMENT HARDWARE COMPONENTS INCLUDING THE SAMPLE, STAT, REAGENT 1 AND REAGENT 2 PROBES. NO PRODUCT DEFICIENCY WAS IDENTIFIED BEHIND THIS ISSUE. | 0
1936810 | DURING PREPARATION FOR USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE ARTERIAL MONITOR DISPLAYED TEMPERATURE READINGS WITHOUT THE TEMPERATURE PROBES ATTACHED. NO OTHER DETAILS REGARDING THE NATURE OF THE EVENT WERE PROVIDED. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY AND THERE WERE NO ADVERSE CONSEQUENCES TO A PT AS A RESULT OF THIS EVENT. | 1
1938116 | THE CUSTOMER STATED THE ARCHITECT C8000 ANALYZER WAS GENERATED ABERRANT RESULTS FOR SODIUM (NA) AND CHLORIDE (CL). PATIENT P317530 HAD AN INITIAL CL RESULT OF 105.7 THAT REPEATED AT 116.2 MMOL/L. FIELD SERVICE WAS DISPATCHED TO INSPECT AND SERVICE THE ANALYZER. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, SYSTEM LOGS AND SERVICE HISTORY, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) REPLACED ALL THE 1 ML SYRINGES, CHECKED THE ICT PROBE ALIGNMENT, CHECKED THE MIXERS AND FLUSHED THE WATER LINES. THE FSR PERFORMED AN ICT CALIBRATION, RAN ICT QUALITY CONTROLS AND A PRECISION TEST. THE ICT QUALITY CONTROLS WERE ACCEPTABLE AND THE PRECISION TEST RESULTS WERE SATISFACTORY. ACCORDING TO THE FSR, THE MOST LIKELY CAUSE IS A TIGHT ASPIRATION 1 ML SYRINGE. THE RESULT LOG SHOWED THE REPEAT SODIUM RESULT OF 112 MMOL/L WAS FLAGGED LOW. THE SYSTEM LOGS REVEALED MILIVOLT FLUCTUATIONS BETWEEN THE PRE AND POST SAMPLE IREF READS, BUT THE FLUCTUATIONS DID NOT EXCEED THE LIMIT OF 10 MV, THEREFORE NO ERROR WAS GENERATED. TRACKING AND TRENDING IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE INVESTIGATION, THE TIGHT ASPIRATION 1 ML SYRINGE IS THOUGHT TO BE THE CAUSE OF THE VOLTAGE DRIFT LEADING TO DISCREPANT ICT RESULTS. | 0
1938165 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. (B)(4). REPORT RECEIVED FROM THE UNITED STATES OF PT IDENTIFIER: (B)(6), BECOMING SYMPTOMATIC WITHIN THE FIRST 15 MINUTES OF A MONONUCLEAR (MNC) STEM CELL APHERESIS PROCEDURE STARTED AT 1050 AND STOPPED AT 1543. IT WAS THE DONOR'S SECOND HETEROLOGOUS DONATION. THIS REPORT IS BEING MADE DUE TO MEDICAL INTERVENTION FOR THE DONOR REACTION. (B)(4). THIS DISPOSABLE SET WAS UNAVAILABLE FOR INVESTIGATION (LOT NUMBER WAS UNK). TRENDS SUGGEST THAT THE EVENT REPORTED IS RANDOM AND ISOLATED TO THIS CUSTOMER ON A GENERAL BASIS, AS SPECTRA PT REACTIONS HAVE BEEN REPORTED AT A RATE OF (B)(4). THE DONOR ELECTED TO HAVE A CENTRAL LINE INSERTED FOR THE PROCEDURE; THEREFORE, THE DONOR WAS NOTHING MY MOUTH (NPO) AFTER MIDNIGHT ON MORNING OF (B)(6) 2009. AS THE DONOR HAD NOT EATEN PRIOR TO THE PROCEDURE, SHE RECEIVED 500 ML NORMAL SALINE BOLUS PRIOR TO THE PROCEDURE. THE DONOR COMPLAINED OF NAUSEA, BECAME PALE AND HAD DISCOMFORT IN HER STOMACH. SHE RECEIVED CALCIUM GLUCONATE TWICE (1 GM IN 50 ML D5W) INTRAVENOUSLY (IV) AND COMPAZINE, 10 MG INTRAVENOUSLY. THE PT WOKE UP AND COMPLAINED OF NAUSEA AND VOMITING. THE DONOR REPORTEDLY VOMITED. IT WAS REPORTED THAT "ALL LABS WERE NORMAL." THE DONOR RECOVERED WITH NO SEQUELAE. THE COBE SPECTRA ESSENTIALS GUIDE STATES; "CAUTION: THE COBE SPECTRA SYSTEM HAS MANY SAFETY FEATURES. A DONOR AND/OR PT REACTION CAN OCCUR RAPIDLY, HOWEVER, IT IS IMPERATIVE THAT THE OPERATOR CONTINUOUSLY MONITOR THE COBE SPECTRA SYSTEM AND THE DONOR AND/OR PT." "ADVERSE EFFECTS - BE AWARE OF POSSIBLE DONOR OR PT REACTIONS DURING APHERESIS PROCEDURES." CONCLUSION: KNOWN POTENTIAL SIDE EFFECT OF PROCEDURE. | 0
1938391 | THE CUSTOMER STATES THAT A PATIENT HAS CONSISTENTLY GENERATED LOW SODIUM ASSAY RESULTS ON SERUM SAMPLES TESTED ON THE ARCHITECT C8000 ANALYZER. THE FOLLOWING INFORMATION WAS PROVIDED FOR THIS PATIENT: FROM (B)(6) 2010 THROUGH (B)(6) 2010, THIS DIABETIC, POST-RENAL TRANSPLANT PATIENT BEING WEANED OFF OF STEROIDS, HAS GENERATED SODIUM RESULTS RANGING FROM 123 TO 127 MMOL/L. THE LAB'S NORMAL REFERENCE RANGE IS 130 TO 149 MMOL/L. THE PATIENT'S MOST RECENT SAMPLE (LIPEMIC; CHOLESTEROL = 220 MG/DL AND TRIGLYCERIDES = 238 MG/DL) WAS SENT TO A REFERENCE LAB WHERE A SODIUM RESULT OF 143 MMOL/L WAS GENERATED (REFERENCE LAB NORMAL RANGE IS 135 TO 145 MMOL/L). THE PATIENT'S DOCTOR HAD THIS PATIENT HOSPITALIZED BECAUSE OF THE LOW SODIUM RESULTS AND INCREASED THE PATIENT'S STEROID THERAPY, TO WHICH NO RESPONSE WAS NOTED. TESTING AT THE HOSPITAL LAB GENERATED A SODIUM RESULT OF 142 MMOL/L, WHICH IS WITHIN THE NORMAL REFERENCE RANGE USED BY THE HOSPITAL. CONTROLS HAVE REMAINED WITHIN SPECIFICATIONS AT ALL TIMES AND A PRECISION RUN PERFORMED BY THE CUSTOMER MET ALL ASSAY PACKAGE INSERT SPECIFICATIONS. THERE ARE NO SIMILAR ISSUES WITH ANY OTHER PATIENTS' SAMPLES. THE CUSTOMER HAS REQUESTED A SERVICE CALL. THERE IS NO FURTHER IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). AN ABBOTT FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE AND PERFORMED PREVENTIVE MAINTENANCE PROCEDURES. HOWEVER, IT WAS NOT UNTIL THE CUSTOMER REPLACED THE INTEGRATED CHIP TECHNOLOGY (ICT) MODULE THAT SAMPLES GENERATED EXPECTED RESULTS. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE EVALUATION BEGAN WITH A REVIEW OF THE CUSTOMER'S SODIUM QUALITY CONTROL RANGES IN USE: LEVEL 1 = 98 - 114 MMOL/L AND LEVEL 2 = 124 - 151 MMOL/L. THESE RANGES ARE BASED ON THE CONTROL MANUFACTURER'S UNITY REPORT AND ARE TOO LARGE TO EFFECTIVELY MONITOR SODIUM PERFORMANCE. ABBOTT CUSTOMER SUPPORT RECOMMENDED THEY ESTABLISH CONTROL RANGES BASED ON THEIR OWN DATA. THE CUSTOMER MAY HAVE BEEN ALERTED TO SODIUM RESULTS BEING DEPRESSED IF THEIR CONTROL RANGES HAD BEEN MORE SUITABLE. THE CUSTOMER RETURNED THE SUSPECT ICT MODULE AND THREE PATIENT SAMPLES. TESTING WITH THESE RETURNED MATERIALS VERIFIED THE ISSUE SEEN AT THE CUSTOMER SITE. THE DATA SUPPORTS THAT THE SODIUM MEMBRANE OF THE RETURNED ICT MODULE WAS CONTAMINATED. QUALITY CONTROL RESULTS WERE NOT AS DEPRESSED AS THE SAMPLE RESULTS, WHICH INDICATES THESE SPECIFIC PATIENT SAMPLES ALSO CONTAIN AN INTERFERENT / CONTAMINATE CONTRIBUTING TO THE ISSUE. THE ICT MODULE BLEACHING PROCEDURE WAS PERFORMED ON THE RETURNED ICT MODULE, AFTER WHICH THE SAMPLE SODIUM RESULTS INCREASED BY 7 - 9 MMOL/L AND CONTROL RESULTS INCREASED BY 2 MMOL/L. A SINGLE BLEACHING APPEARS TO HAVE REMOVED SOME, BUT NOT ALL OF THE CONTAMINATION. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ICT SAMPLE DILUENT PACKAGE INSERT ((B)(4), (B)(4) 2009) AND THE ARCHITECT SYSTEMS OPERATIONS MANUAL ((B)(4), (B)(4) 2010) CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. BASED ON THE AVAILABLE INFORMATION, A SPECIFIC CAUSE FOR THE DISCREPANT LOW SODIUM RESULTS BEING GENERATED BY SPECIFIC PATIENT SAMPLES WAS NOT IDENTIFIED. PROBABLE CAUSES INCLUDE SODIUM MEMBRANE EXPOSURE AND/OR CONTAMINATION OVER TIME IN CONJUNCTION WITH UNKNOWN SAMPLE FACTORS. THE CUSTOMER'S ICT MODULE HAD BEEN IN USE FOR AT LEAST SIX WEEKS AND UP TO 12,000 SAMPLES, WHICH PROVIDES AMPLE OPPORTUNITY FOR CONTAMINATION. THERE IS NO INDICATION THAT THE ICT MODULE WAS BLEACHED TO REMOVE CONTAMINATION. IT APPEARS THAT SAMPLES FROM PARTICULAR PATIENTS MAY INCLUDE SPECIFIC FACTORS (INTERFERENTS/CONTAMINANTS) THAT AFFECT SOME SODIUM MEMBRANES THAT HAVE BEEN EXPOSED TO UNKNOWN CONTAMINATES OVER TIME. METHOD: REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION. | 0
1939223 | THE CUSTOMER STATED AN APS I2000SR ANALYZER GENERATED ERROR CODE 8275, "LAS SENT SID 74874 WHEN SID K40742 PRESENT AT SAMPLE POINT". THEN THE LAS SENT SID K40742 WHEN K40742 WAS PRESENT. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. INVESTIGATIONS HAVE SHOWN A DEFICIENCY OF THE ARCHITECT SYSTEM EXISTS IN THAT MESSAGES ARE NOT PROCESSED WITHIN THE EXPECTED TIME PERIOD. THE SOFTWARE SHOULD BE CHANGED TO RECALCULATE THE TIME PERIOD MEASURED. THE 8275 ERROR MESSAGE OCCURS WHEN A SAMPLE IS UNEXPECTEDLY DETECTED AT THE SAMPLING POSITION, BUT THE TESTS ARE ALLOWED TO PROCESS TO COMPLETION. INPECO WILL ADDRESS THE ISSUE BY DEVELOPING NEW FIRMWARE THAT WILL DELETE EXTRA PENDING MESSAGES AND AVOID THE 8275 ERROR. THE COMPLAINT ISSUE INVOLVES THE INTERACTION BETWEEN THE ACCELERATOR APS AND THE ARCHITECT SYSTEMS. TRACKING AND TRENDING IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THIS COMPLAINT ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. | 0
1942039 | A SITE REPORTED A LOSS OF MONITORING OCCURRED FOR APPROX 9-11 TELEMETRY PATIENTS WHEN THE CIC (CENTRAL STATION) REBOOTED. NO INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1942406 | IT WAS REPORTED THAT A NURSE SELECTED THE "OVERRIDE ALL" AT THE 4DTC WORKSTATION BUT DID NOT VERIFY THE RESULTS. THE GANTRY ANGLE, COLLIMATOR ANGLE, JAW POSITION WAS NOT COMPLETELY WHERE IT SHOULD HAVE BEEN. THE PT WAS POSITIONED ON THE TABLE AND THE TREATMENT WAS STARTED. THE PT WAS DELIVERED 42 MUS BEFORE HOSPITAL STAFF REALIZED THAT THE POSITION WAS NOT CORRECT. THE TREATMENT WAS STOPPED IMMEDIATELY. ACCORDING TO THE HOSPITAL PHYSICIST, CRITICAL ORGANS WERE NOT IRRADIATED DURING THE OCCURRENCE. AFTER INVESTIGATION BY THE HOSPITAL STAFF ON SITE, THE HOSPITAL PHYSICIST INFORMED VARIAN OF THIS INCIDENT AS A COURTESY. NO OTHER PT INFO WAS PROVIDED. NO SERIOUS DETERIORATION IN STATE OF HEALTH IS REPORTED. THE PHYSICIST REPORTED CRITICAL ORGANS WERE NOT IRRADIATED DURING THE OCCURRENCE. THOUGH STILL UNDER INVESTIGATION, VARIAN HAS DETERMINED THAT A MDR IS APPROPRIATE. ADD'L F/U TO THIS MDR IS EXPECTED UPON COMPLETION OF THE INVESTIGATION. | 1
1942872 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. CUSTOMER CALLED STATED PERFORMING RBCX ON SPECTRA SYSTEM NOTED HEMOLYSIS IN PLASMA LINE WHILE INFUSING THE 2ND TO LAST BAG. AFTER NOTICING THE HEMOLYSIS, SHE SWITCHED TO THE LAST BAG HANGING. SHE PAUSED THE PROCEDURE BEFORE COMPLETION OF THE FINAL BAG. SHE CALLED US FOR ASSIST. PROCEDURE WAS ABORTED. FINDINGS BY LAB HEMOLYSIS CHECK ON BOTH UNITS OF BLOOD POSITIVE FOR HEMOLYSIS (1+ AND 2+). BLOOD FROM ACCESS LINE/NEGATIVE FOR HEMOLYSIS AND RETURN LINE TO PT WAS 2+ FOR HEMOLYSIS. BLOOD BANK AWARE AND INVESTIGATING UNITS WITH SAME LOT NUMBERS. THE COMPLAINANT WOULD NOT SUPPLY A PT ID. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR HEMOLYSIS. FIN (B)(4). WARNINGS AND RESPONSES TO ADVERSE CONDITIONS DURING APHERESIS PROCEDURES CAN BE FOUND IN THE SPECTRA OPERATOR'S MANUAL (B)(4) PN 777093-392. THE OPERATOR IS WARNED (SECTION XXVI 16 AND SECTION XXVIII 4) TO MONITOR PTS CLOSELY DURING APHERESIS PROCEDURES FOR ANY REACTIONS. A SERVICE REP CHECKED OUT THE EQUIPMENT AND FOUND NO ISSUES. THE CUSTOMER SENT SAMPLES FROM THE RBC UNIT BAG, THE ACCESS LINE, AND THE RETURN LINE TO A LAB FOR EVAL. THE RESULTS VERIFY THAT NO HEMOLYSIS WAS FOUND IN THE ACCESS LINE, BUT HEMOLYSIS WAS FOUND IN THE RETURN LINE AND THE RBC UNITS. CLINICAL SPECIALIST, (B)(6), CONFIRMS THAT OLDER UNITS OF RBC'S CAN SOMETIMES BECOME HEMOLYZED BEFORE BEING USED. THE RESULTS FROM THE LAB WORK POINT TO THIS BEING THE CAUSE OF THE HEMOLYSIS DURING THE PROCEDURE. THE MOST LIKELY CAUSE OF THE HEMOLYSIS WAS AGE OF THE RBC UNITS USED, THE HEMOLYSIS WAS APPARENT BEFORE THE BAGS WERE USED IN THE PROCEDURE; THERE IS NO INDICATION THAT CARIDIANBCT'S EQUIPMENT CAUSED OR CONTRIBUTED TO THE FAILURE. A SERVICE CALL WAS PLACED AND THE EQUIPMENT WAS CHECKED OUT AND FOUND TO BE FUNCTIONING AS DESIGNED. CONCLUSION: HEMOLYSIS APPARENTLY CAUSED BY OLDER RBC UNITS BEING USED FOR RBCX. | 0
1944099 | I DO ON-SITE DRUG TESTING AT A MFG FACILITY IN (B)(6). I AM CURRENTLY USING "MULTI-DRUG SCREEN TEST PANEL" DISTRIBUTED BY (B)(4). LOT NUMBER DOA0030282, EXP DATE 01/2012; CATALOG NUMBER DOA-354. I HAVE HAD NUMEROUS TESTS THAT FAILED TO COMPLETELY DEVELOP. THE RESULT HAS BEEN THAT I HAVE HAD TO REPEAT THE TEST USING ANOTHER TEST CASSETTE. I CONSIDER THIS A PRODUCT FAILURE. DIAGNOSIS OR REASON FOR USE: ON SITE DRUG TESTING. | 0
1944587 | WIRELESS FOOTSWITCH BATTERY POWER LOST AFTER SHORT USE. DISCONNECT FROM POWER SOURCE AND BATTERY POWER OK. WIRELESS FOOTSWITCH WOULD NOT WORK UNTIL IT WAS PLUGGED INTO POWER CABLE.PATIENT CARE DELAYED. FLUOROSCOPY PEDAL WOULDN'T WORK DURING THE PROCEDURE. | 1
1944699 | CAS LEARNED OF THE DEATH BY WAY OF A NOTICE THAT A LAWSUIT HAS BEEN BROUGHT. THAT NOTICE ARRIVED 11/17/2010 NAMING THE DEFENDANTS AS RESPONSIBLE FOR THE BABY'S DEATH BECAUSE IT WAS CLAIMED THAT THE MONITOR DID ALARM, AWAKENING THE PARENTS, FOR SLOW HEART RATE WHEN IN FACT THE BABY HAD STOPPED BREATHING ALTOGETHER. CAS WAS NOT NAMED AS A DEFENDANT IN THE SUIT, BUT COUNSEL FOR THE PLAINTIFF DISCOVERED THAT CAS HAS MDR REPORTING RESPONSIBILITY FOR THE MONITOR DUE TO A PURCHASED AND SALE AGREEMENT FROM 2000. THE MONITOR WAS NEVER MANUFACTURED BY CAS, BUT WAS MANUFACTURED BY AEQUITRON AND OR (B)(4)/NELLCOR. WHEN CAS PURCHASED THE MONITOR BUSINESS FROM NELLCOR THIS MONITOR WAS ALREADY DISCONTINUED, HOWEVER, CAS HAD RESPONSIBILITY FOR THE REGULATORY AFFAIRS OF THIS AND OTHER MONITORS AS WELL AS SERVICING RESPONSIBILITY. THE COUNSEL FOR THE PLAINTIFFS HAS NOT RESPONDED TO A LETTER REQUEST FOR THE MONITOR SERIAL NUMBER, THEREFORE, IT CANNOT BE DETERMINED IF THIS PARTICULAR MONITOR EVER ENTERED THE FACILITY OF CAS. CAS CLEARLY DENIES ANY INVOLVEMENT IN THE PATIENT'S OUTCOME. | 1
1945131 | DURING DELIVERY OF ANESTHESIA, THE MACHINE SUDDENLY WENT OFF. PATIENT WAS AMBUED FOR APPROXIMATELY ONE MINUTE AND MACHINE CAME BACK ON. NO APPARENT INJURY TO THE PATIENT OCCURRED AS VENTILATION WAS ASSISTED AND VITAL SIGNS REMAINED STABLE. MACHINE CONTINUED TO WORK APPROPRIATELY FOR THE REMAINDER OF THE CASE. FURTHER INVESTIGATION REVEALED NOTHING WAS PLUGGED IN BACK OF MACHINE THAT WAS NOT ALREADY INDICATED TO BE PLUGGED INTO IT. NO ALARM WAS HEARD AS IT WAS TURNED DOWN. THE SCREEN WAS REPORTED TO HAVE GONE BLANK AND DEVICE WAS NOT PUMPING. ALARM LOG DID NOT GIVE INDICATION OF POWER LOSS OR BATTERY FAILURE AS EVENT LOG ONLY HAS CAPACITY FOR TEN MESSAGES. ON THE ELEVENTH EVENT, THE FIRST EVENT WILL BE DELETED. THE POWER FAILURE EVENT IS BELIEVED TO HAVE BEEN PURGED OFF THE LIST FOR THIS REASON. MANUFACTURER'S TESTING PROCEDURE WAS COMPLETED WITH FOLLOWING RESULTS: 1.ALARM LOG DID NOT GIVE ANY INDICATION OF POWER LOSS OR BATTERY FAILURE AS MACHINE WAS IMMEDIATELY PLUGGED INTO AN EMERGENCY OUTLET AND CASE CONTINUED, 2. 25 LINK CALIBRATION PERFORMED DUE TO NO MINIMUM 02 FLOW (STANDARD PM), 3. AN AMPERAGE DRAW SIGNIFICANT ENOUGH TO TRIP A 20-AMP BREAKER WOULD USUALLY RESULT IN A NON-RECOVERABLE FAILURE (I.E. POWER SUPPLY) AS AGREED BY GE TECH SUPPORTED, 4. TESTED TO MANUFACTURER'S PROCEDURE FOR THE AESTIVA 7900. | 1
1945180 | ANESTHETIC AGENT DELIVERING AT THE WRONG CONCENTRATION. THIS SAME MACHINE WAS INVOLVED IN ANOTHER EVENT. BIOMED WAS IN OR TO EVALUATE ISSUE OF ETT NOT FUNCTIONING AND IT WAS NOTED THAT ETT ISSUE WAS PLACEMENT PROBLEM NOT DEFECT, BUT HE AND MD NOTED THAT THE MACHINE WAS NOT DELIVERING ANESTHETIC AGENT AT CORRECT CONCENTRATION.======================HEALTH PROFESSIONAL'S IMPRESSION======================THE ANESTHESIA UNIT WAS REMOVED FROM SERVICE AND REPAIRED BY REPLACING THE FRESH GAS MODULE. THIS MODULE CONTROLS THE PERCENTAGE OF ANESTHETIC AGENT DELIVERY, AND HAS THE CALIBRATION CONSTANTS AND THE ATMOSPHERIC PRESSURE. AFTER REPAIR, THE UNIT WAS CALIBRATED AND RETURNED TO SERVICE. NOTE: THIS CHANGE DID NOT OCCUR WHEN THE MFG. CAME TO EVALUATE MACHINE INITIALLY WITH THE FIRST OCCURRENCE. THEY FOUND NO DEFECT. | 0
1945181 | PATIENT IN OR FOR EMERGENCY SURGERY; RESPIRATORY STATUS WAS DIMINISHING. THE ANESTHESIA MACHINE WAS ATTACHED TO THE PATIENT (THIS MACHINE HAD BEEN IN USE ALL DAY) AFTER APPROPRIATE CHECKS WERE DONE AND ACCEPTABLE. WHEN CRNA CERTIFIED NURSE ANESTHETIST TRIED TO START THE MACHINE, THE MAIN SWITCH (THE PUSH/ROTATING KNOB ON THE SCREEN WHICH ALLOWS THE MACHINE TO ENTER DIFFERENT MODES, INCLUDING "START CASE") THAT CONTROLS THE FUNCTION FAILED AND THEY COULD NOT USE THIS MACHINE; SWITCH WAS UNRESPONSIVE AND THEY WERE NOT ABLE TO TURN THE MACHINE ON TO THE "START CASE" MODE. ANOTHER ANESTHESIA MACHINE WAS IMMEDIATELY BROUGHT INTO THE ROOM AND ATTACHED TO THE PATIENT SUCCESSFULLY.======================HEALTH PROFESSIONAL'S IMPRESSION======================NO ADVERSE OUTCOME FOR THIS PATIENT AS ANOTHER MACHINE WAS READILY AVAILABLE. | 0
1946242 | TECH WAS TESTING THE IVUS SYSTEM PRIOR TO THE PROCEDURE. WHEN ATTEMPTING TO PULL BACK THE ULTRASOUND CATHETER OF IVUS SYSTEM, THE MOTOR DRIVE UNIT DID NOT PULL BACK THE CATHETER.THE MANUFACTURER TESTED THE UNIT, AND COULD NOT DUPLICATE THE PULL BACK FAILURE, BUT DID REPLACE THE UNIT. | 0
1946275 | THE PATIENT EXPERIENCED AT LEAST A FEW BRADYCARDIA EVENTS WITHIN A 30-MINUTE TIME PERIOD, WITH HEART RATE DECREASING TO 55 BPM FROM THE NORMAL 80-90 BPM. THE PHYSIOLOGICAL MONITORING SYSTEM PROPERLY IDENTIFIED AND DISPLAYED THE PATIENT'S BRADYCARDIA CONDITION, BUT NO AUDIBLE ALARM OCCURRED, EITHER IN THE PATIENT'S LOCAL MONITOR OR IN THE REMOTE MONITORING WORKSTATION. THE SYSTEM DID NOT GENERATE AN AUTOMATIC PRINT, AND THE SYSTEM'S ALARM HISTORY FILE DID NOT INCLUDE AN ENTRY FOR THE ALARM. THE STORED "FULL DISCLOSURE" ECG WAVEFORM WAS PROPERLY ANNOTATED FOR THE BRADYCARDIA CONDITION. THIS TYPE OF EVENT HAS OCCURRED SEVERAL TIMES IN RECENT MONTHS. SYSTEM LOGS FOR SOME OF THE EVENTS HAVE BEEN SENT TO THE MANUFACTURER FOR ANALYSIS. MANUFACTURER HAS CONFIRMED THE EVENTS, BUT HAS BEEN UNABLE TO PROVIDE ANY MEANS TO CORRECT THE PROBLEM OR PREVENT FUTURE SIMILAR EVENTS FROM OCCURRING.HEALTH PROFESSIONAL'S IMPRESSION: UNDER CERTAIN CONDITIONS, THE MONITORING SYSTEM RECOGNIZES AN ALARM CONDITION, BUT FAILS TO GENERATE AN AUDIO SIGNAL OR PRINTED STRIP, AND DOES NOT ENTER THE EVENT IN THE SYSTEM'S ALARM HISTORY LOG.MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, CENTRAL, CIC PRO: ANALYSIS AND REPLIES STILL PENDING FOR THIS EVENT.MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, SOLAR: REPLIES ARE PENDING RESULTS OF DETAILED INVESTIGATION. | 1
1948632 | THE MONITOR DISPLAY HAS NOISE INTERMITTENTLY IN THE TWO LEFT QUADRANTS OF THE SCREEN. THE NOISE APPEARS TO BE FLASHING WHITE HORIZONTAL LINES. | 1
1949159 | THE ROTATOR BROKE APART DURING POWER INJECTION. THERE WAS NO SPRAY AND NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE BUT HAS NOT PROVIDED ANY ADDITIONAL INFO OR CLINICAL DETAILS FOR THE SECOND EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVALUATION: THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION/ INVESTIGATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETED. EVALUATION: CONCLUSIONS: A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
1949187 | A SITE REPORTED A LOSS OF TELEMETRY MONITORING DUE TO AN ALLEGED HARD DRIVE FAILURE. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A F/U REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
1949832 | THE ROTATOR BROKE DURING USE. NO HARM OR INJURY WAS REPORTED. CUSTOMER IS RETURNING ONE UNIT OF OLD CONFIGURATION. CUSTOMER REPORTED TWO DEFECTIVE BUT HAS NOT PROVIDED ANY ADDITIONAL INFO OR CLINICAL DETAILS FOR THE SECOND EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL: ONE USED SUSPECT DEVICE WAS RETURNED FOR EVAL. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. COMPLAINT DATABASE REVIEW FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. THE DEVICE WAS EXAMINED VISUALLY AND THE COMPLAINT WAS NOT CONFIRMED. THE UNIT RETURNED WAS VERIFIED AS NOT FROM THE SAME LOT NUMBER REPORTED BY THE CUSTOMER. THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. THE ROOT CAUSES OF THE FAILURE ARE ATTRIBUTED TO THE MFG BONDING PROCESS. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. THE DEVICE RETURNED WAS BUILT PRIOR TO IMPLEMENTATION OF THE NOTED CORRECTIVE ACTIONS. COMPLAINT DATA WILL BE MONITORED TO VERIFY EFFECTIVENESS OF CORRECTIVE ACTIONS TAKEN. EVAL, METHOD: DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATA BASE WAS REVIEWED. | 0
1950154 | SOMARAD DIGITAL X-RAY RADIOGRAPHY SYSTEM WAS IN USE AT THE (B)(6). THE STAND BEGAN TO MOVE BY ITSELF AND THE OPERATOR ACTIVATED THE EMERGENCY OFF SWITCH TO STOP THE MOVEMENT. THERE WAS A PT IN THE VICINITY BUT THE MOVEMENT DID NOT IMPACT THE PT. NO INJURY TO A PT. NO ACCIDENTAL RADIATION EXPOSED. IMIX IS REPORTING THIS INCIDENT AS MANUFACTURER AND SYSTEM INTEGRATOR OF SONORAD DIGITAL RADIOGRAPHY SYSTEM WHICH INTEGRATES IMIX DIGITAL RADIOGRAPHY DETECTOR WITH THE STAND MANUFACTURED BY (B)(4). BASED ON THE PRELIMINARY INVESTIGATION, INITIALLY, IT WAS CONSIDERED AS AN ISOLATED INCIDENT AND AN ERROR BY THE OPERATOR AND THE INVESTIGATIONS DID NOT REVEAL ANY ROOT CAUSE FOR THE PROBLEM. AFTER FURTHER STUDIES AND ENGINEERING EVALUATION OF THE UNIT, IT WAS DETERMINED THAT AN INTERFERENCE INTRODUCED BY THE ELECTRICAL NOISE FROM THE EXTERNAL KEY PAD CONTROL UNIT COULD RESULT IN SUCH A MOVEMENT. ALL OTHER POSSIBLE SCENARIOS WERE RULED OUT. AS A CORRECTIVE ACTION, A FIELD NOTICE WAS ISSUED TO DISCONNECT THE EXTERNAL KEY PAD CONTROL UNIT. THIS DOES NOT HAVE AN EFFECT ON FUNCTIONALITY OR SAFETY OF THE DEVICE. AS A PREVENTIVE ACTION, IMIX IS PLANNING TO INSTALL A COLLISION SENSOR AS AN ADDED SAFETY FEATURE TO AVOID ANY SUCH UNINTENDED MOVEMENT. IMIX'S INTERNAL CORRECTIVE ACTION (B)(4) WILL CONTINUE UNTIL EFFECTIVENESS OF THE CORRECTIVE/ PREVENTIVE ACTION IS CONFIRMED. THIS REPORT IS NOT MEETING THE FDA'S REQUIRED REPORTING TIME FRAME DUE TO THE FACT THE IMIX DID NOT IDENTIFY THIS INCIDENT AS A REPORTABLE EVENT. HOWEVER, AFTER CAREFUL STUDY OF THE EVENT/S, IMIX MANAGEMENT DECIDED TO SEND THIS REPORT TO FDA. | 0
1952177 | THE CUSTOMER REPORTED THAT ON SIX OCCASIONS, NEEDLE DISPLACEMENT HAS OCCURRED. ON ALL EVENTS, THERE WAS A HIGH RETURN PRESSURE ALARM AND ON INSPECTION OF THE VENIPUNCTURE SITE, THE NEEDLE HAD MOVED OUT OF THE VEIN AND THE PROCEDURE WAS STOPPED. THIS REPORT IS BEING SUBMITTED DUE TO MEDICAL INTERVENTION. (B)(4). NO OTHER INFORMATION AVAILABLE. A SUPPLEMENTAL REPORT WILL BE FILED WHEN ADDITIONAL INFORMATION IS AVAILABLE. | 0
1953724 | HUSBAND WITH AN ANESTHESIOLOGIST AT COUNTY HOSP IN (B)(6), PLACED A CALL TO NOTIFY WIFE OF ANESTHESIA MACHINES FAILING. KEEP TURNING THEMSELVES OFF. NO PT INVOLVED. MACHINE IS PART OF A RECALL. | 0
1958266 | (B)(6) HOSPITAL LOCATED IN (B)(6) -PART OF THE (B)(6) MEDICAL CENTER- HAS HAD NOT ONE BUT TWO GE HEALTHCARE AISYS AND AVANCE ANESTHESIA MACHINES FAIL IN THE LAST TWO DAYS. THE FIRST MACHINE FAILED YESTERDAY - (B)(6) 2011- THE SECOND MACHINE FAILED TODAY- (B)(6) 2011-. GE SENT A FIELD ENGINEER BY THE NAME OF (B)(4). (B)(4) SPOKE EXTENSIVELY ABOUT THE ONE MACHINE, THE SECOND HAD YET TO BREAK WHEN HE WAS THERE BUT AT NO TIME DID HE MENTION A RECALL OF THIS PRODUCT. BOTH MACHINES TURNED OFF THE ANESTHETIC DELIVERY AGENT MULTIPLE TIMES AND FINALLY CEASED TO READ AND OR DELIVER THE ANESTHETIC AGENT MID-WAY THROUGH TWO SEPARATE SURGERIES. AFTER THE MACHINES WERE TURNED OFF IN BOTH INSTANCES, THE PT WAS AMBU-BAGGED AND THEN THE MACHINE RESTARTED AND THEY WERE ABLE TO DELIVER THE AGENT FOR THIRTY MINUTES BEFORE THE ANESTHETIC AGENT DELIVERY SYSTEM SHUTDOWN. BOTH TIMES, THE ATTENDING ANESTHESIOLOGIST WAS REQUIRED TO INTERVENE IN ORDER TO PREVENT PERMANENT DAMAGE TO THE PATIENTS. RISK MGMT AT (B)(6) HAS A WOMAN APPRISED OF THE SITUATION, I'M FAIRLY SURE SHE KNOWS NOTHING OF THIS RECALL. HER NAME IS (B)(6) AND HER PHONE NUMBER IS (B)(6). THE SERIAL NUMBERS ON THE MACHINES SEEM TO BE DIFFERENT THAN THE ONES IN YOUR RECALL BUT THE FAILURE SEEMS TO BE IDENTICAL IN NATURE TO THE RECALLED MACHINES. THE SERIAL NUMBERS ARE AS FOLLOWS: MACHINE #1 (B)(4). THE SECOND MACHINE HAS THE FOLLOWING NUMBERS: (B)(4). THE BIG PROBLEM WE ARE NOW FACING IS THAT AS A LEVEL 2 TRAUMA CENTER WITH 200+ BEDS, 7 OF OUR REMAINING NINE MACHINES ARE EXACTLY THE SAME AS THESE DEFECTIVE MACHINES. PLEASE MAKE CONTACT WITH THE APPROPRIATE HOSPITAL PERSONNEL TO APPRISE THEM OF YOUR RECALL WITH THE HOPES WE CAN RESOLVE THIS IMMEDIATELY. | 0
1958969 | THE CUSTOMER INSTALLED A NEW PRESSURE MONITOR SENSOR ON THE AXSYM ANALYZER AND POWERED ON THE ANALYZER. THE CUSTOMER THEN OBSERVES SMOKE COMING FROM THE NEW PRESSURE MONITOR SENSOR. THE CUSTOMER TURNED OFF THE POWER AND REPLACED THE PRESSURE MONITOR SENSOR WITH ANOTHER ONE. PRIOR TO OBSERVING SMOKE, THE CUSTOMER GENERATED SEVERAL ERROR CODES (3516, 3558, 3530 AND 5060-HOMING FAILURE, SAMPLING SYRINGE, SAMPLING CENTER). THE CUSTOMER DID NOT OBSERVE LIQUID IN THE PRESSURE MONITOR SENSOR AREA AND NO INJURIES WERE REPORTED. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. THE ISSUE WAS RESOLVED BY REPLACING THE PRESSURE MONITOR SENSOR BY ANOTHER ONE. (B)(4). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. A REVIEW OF THE COMPLAINT AND THE CURRENT AXSYM LABELING WAS PERFORMED. THE OPERATIONS MANUAL CONTAINED PRECAUTIONS FOR POTENTIAL HAZARDS. THE PRESSURE MONITORING ADDENDUM CONTAINS INSTRUCTIONS FOR REPLACING AND INSTALLING THE PRESSURE MONITOR SENSORS. THE ADDENDUM ALSO CONTAINS INSTRUCTIONS FOR RESOLVING ERROR CODES 3558, 3516 AND 3530 ALONG WITH MULTIPLE CAUSES AND CORRECTIVE ACTIONS. THE CAUSES INCLUDED FUNCTIONAL ISSUES WITH THE PRESSURE MONITOR SENSOR. NO SUBSEQUENT ISSUES WERE REPORTED AFTER REPLACING THE PART AND NO OTHER COMPLAINTS OF SIMILAR ISSUES WERE IDENTIFIED FOR THIS PART. NO PRODUCT DEFICIENCY WAS IDENTIFIED RELATED TO THIS ISSUE. | 0
1959090 | DURING (B)(6) 2010 SERVICE CALL TO A CUSTOMER SITE FOR THE ELITE MICROPLATE SYSTEM/AUTOMATED MICROPLATE WORKSTATION, IT WAS DETERMINED THAT THE TADM (TOTAL ASPIRATION AND DISPENSE MONITORING) CLOT DETECTION SYSTEM FOR THE INSTRUMENT HAD BEEN TURNED OFF SINCE A SERVICE CALL IN 2007. THE TADM SYSTEM IS INTENDED TO NOTIFY THE USER WHEN A CLOT PREVENTS ASPIRATION OF THE CORRECT VOLUME OF SPECIMEN. WHEN A CLOTTED SPECIMEN IS TESTED ON THE ELITE AND THE TADM IS NOT ENABLED, EITHER THE SAMPLE WILL BE REPORTED AS NEGATIVE OR THE ENTIRE PLATE WILL BE REJECTED. THE ELITE SYSTEM WAS BEING USED FOR HIV ANTIBODY TESTING WITH THE (B)(4) GS (B)(4) PLUS O EIA AT THE TESTING SITE. WITHOUT THIS CLOT DETECTION, SOME SAMPLES MAY HAVE BEEN REPORTED AS (B)(4) NEGATIVE WHEN ADEQUATE SAMPLE VOLUME HAD NOT ACTUALLY BEEN TESTED. SAMPLES THAT MAY HAVE BEEN AFFECTED ARE BEING RETESTED, AND NO FALSE NEGATIVE SAMPLES HAVE BEEN IDENTIFIED TO DATE. | 0
1959896 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED DISCREPANT GLUCOSE RESULTS. AN INITIAL RESULT OF 224 MG/DL WAS GENERATED. THE SAMPLE WAS REPEATED WITH RESULTS OF 178, 136, AND 81 MG/DL GENERATED. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR RESOLVED THE ISSUE BY PERFORMING SERVICE ON THE CUVETTE WASHER AND ALIGNING THE MIXERS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY GLUCOSE REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC/DISCREPANT RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1959945 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED DISCREPANT GLUCOSE RESULTS. AN INITIAL RESULT OF 248 MG/DL WAS GENERATED. THE SAMPLE WAS REPEATED AND A RESULT OF 107 MG/DL WAS GENERATED. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR RESOLVED THE ISSUE BY PERFORMING SERVICE ON THE CUVETTE WASHER AND ALIGNING THE MIXERS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY GLUCOSE REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC/DISCREPANT RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. | 0
1960020 | THE PHILIPS, PATIENT INFORMATION CENTER (PIC) COMPUTER WHICH DISPLAYS 16 PATIENT'S TELEMETRY WAVE FORMS FROZE FOR A FEW MINUTES, THE SYSTEM WAS THEN REBOOTED. THE REBOOTING PROCESS TOOK ABOUT 5 MINUTES. SHORTLY AFTER THE DISPLAY MONITOR CAME BACK ON, IT WAS NOTED THAT ONE PATIENT HAD BRADYCARDIA. THIS PATIENT WAS AN UNRESPONSIVE DNR PATIENT, FOR WHICH COMFORT MEASURES ONLY WERE BEING ADMINISTERED. THE PATIENT EXPIRED ABOUT ONE HOUR LATER.HEALTH PROFESSIONAL'S IMPRESSION: IT WAS DETERMINED THAT THE PATIENT WAS RECEIVING COMFORT MEASURES ONLY AND WAS AT THE END OF LIFE PRIOR TO THE INCIDENT. HOWEVER, THERE WERE 15 OTHER PATIENTS THAT WERE AT RISK DURING THIS PERIOD FOR WHICH WE WERE UNABLE TO VIEW THEIR WAVE FORMS. THERE WERE NO ADVERSE OUTCOMES TO THE OTHER 15 PATIENTS.MANUFACTURER RESPONSE FOR MONITOR, CENTRAL, PHYSIOLOGICAL, TELEMETRY: WAITING FOR RESPONSE FROM FIELD SERVICE REP. | 1
1960036 | AT APPROXIMATELY 19:41 HOURS THE PATIENT MONITORING FOR A SINGLE PATIENT WAS LOST DUE TO A SUSPECTED SIGNAL LOSS.SYSTEM LOGS WERE RAN AND DID NOT INDICATE ANY ERRORS WITHIN THE WIRELESS NETWORK. NURSING STAFF CONFIRMED THAT A "NO SIGNAL" ALARM WAS RECEIVED AT THE PHILIPS INFORMATION CENTER (PIC).THE PATIENT WAVE REVIEW SHOWS NO MONITORING DATA (ECG WAVE FORMS) FOR THIS PATIENT FOR 39 MINUTES.HEALTH PROFESSIONAL'S IMPRESSION: NO CARDIAC MONITORING FOR A SINGLE PATIENT FOR 39 MINUTES.MANUFACTURER RESPONSE FOR TRANSMITTER, PHYSIOLOGICAL, TELEMETRY:MFG. IS SCHEDULING RESOURCES TO VISIT SITE AND GATHER DATA TO FURTHER ASSIST WITH THE INVESTIGATION. | 1
1960281 | ON (B)(6) 2010, AN IMRIS CUSTOMER SUPPORT SPECIALIST REPORTED THAT THE OPERATING ROOM TABLE MOVED TO THE MAXIMUM HEIGHT OR TILT CONTINUOUSLY UPON TURNING ON THE OPERATING ROOM TABLE PENDANT (CONTROLLER). THERE WERE NO REPORTS OF INJURY TO THE PATIENT OR CLINICAL STAFF. INITIAL INSPECTION OF THE OPERATING ROOM TABLE DETERMINED THAT THE ISSUE COULD BE ATTRIBUTED TO THE OPERATING ROOM PENDANT SINCE THE REPLACEMENT OF THE PENDANT RESULTED IN A CORRECTION OF THE PROBLEM. A REPLACEMENT PENDANT HAS BEEN SHIPPED TO THE SITE PENDING THE RETURN AND EVAL OF THE ORIGINAL DEVICE COMPONENT. THE OPERATING ROOM TABLE PENDANT WILL BE RETURNED TO IMRIS FOR EVAL. A F/U REPORT WILL BE SUBMITTED AT THE CONCLUSION OF THE DEVICE COMPONENT EVAL. | 0
1961763 | THE USER FACILITY REPORTED THAT PRIOR TO THE START OF A PATIENT PROCEDURE, HOSPITAL STAFF OBSERVED A CRACK IN THE BRAKE COVER AND PART OF THE BRAKE HOSE BEHIND THE BRAKE COVER OF THE EQUIPMENT MANAGEMENT SYSTEM (EMS). THE STAFF TOOK NO ACTION TO ADDRESS THE CRACKED BRAKE COVER AND AS THE PATIENT WAS BEING PREPARED, THE BRAKE HOSE RUPTURED, CAUSING A LOUD NOISE AND RELEASING NITROGEN INTO THE OPERATING ROOM. NITROGEN IS USED TO FILL THE BLADDER BEHIND THE BRAKE COVER WHICH PREVENTS THE EMS FROM MOVING DURING A PATIENT PROCEDURE. THE PATIENT WAS TRANSFERRED TO ANOTHER OPERATING ROOM WHERE THE PROCEDURE WAS PERFORMED SUCCESSFULLY WITHOUT INCIDENT. NO INJURIES WERE REPORTED TO HOSPITAL STAFF OR PATIENT. A STERIS SERVICE TECHNICIAN INSPECTED THE EQUIPMENT MANAGEMENT SYSTEM (EMS) AND FOUND NO PROBLEMS WITH THE SYSTEM OUTSIDE OF THE BROKEN BRAKE COVER. THE TECHNICIAN WAS UNABLE TO CHECK THE PRESSURE TO THE SYSTEM AS THE NITROGEN GAS SUPPLY HAD BEEN TURNED OFF TO THE OPERATING ROOM. THE TECHNICIAN REPLACED THE BRAKE COVER AND BRAKE HOSE, AND HAD THE FACILITY TURN ON THE NITROGEN GAS SUPPLY SO THAT THE CORRECT PRESSURE SETTING COULD BE VERIFIED. THE EMS HAS BEEN PLACED BACK INTO SERVICE; NO FURTHER ISSUES HAVE BEEN REPORTED WITH THIS EQUIPMENT. UPON FURTHER INSPECTION, THE TECHNICIAN FOUND THAT THE EMS PRESSURE REGULATORS IN FOUR OTHER OPERATING ROOMS AT THE FACILITY WERE SET TOO HIGH AND THAT ALL OF THE BRAKE COVERS IN THESE OPERATING ROOMS HAD VARYING SIZES OF CRACKS DUE TO THE HIGH PRESSURE SETTINGS. THE TECHNICIAN ADJUSTED THE NITROGEN PRESSURE IN THESE FOUR OPERATING ROOMS AND THE USER FACILITY HAS REPLACED THE CRACKED BRAKE COVERS. THE MAINTENANCE MANUAL STATES ON PP 3-3 THAT AT EACH INSPECTION "1.2 CHECK OVERALL SYSTEM FOR APPEARANCE, VISUAL DAMAGE, MISSING/LOOSE PARTS OR DEFORMATION". AND "2.4 CHECK COVERS FOR SIGNS OF CRACKING/BINDING". ADDITIONALLY, THE MAINTENANCE MANUAL STATES ON PP 6-1 "EVERY YEAR, ENGAGE A PROFESSIONAL TESTING SERVICE TO CHECK ALL PRESSURE HOSES AND CABLES". THE EMS IS NOT UNDER STERIS SERVICE CONTRACT AND IS CURRENTLY MAINTAINED AND SERVICED BY THE FACILITY'S BIOMEDICAL DEPARTMENT. STERIS OFFERED TO CONDUCT IN-SERVICE TRAINING REGARDING PROPER PREVENTIVE MAINTENANCE PROCEDURES OF THIS EMS HOWEVER THE USER FACILITY DECLINED. | 0
1963552 | THE PRESSURE TUBING BREAKS UP WHILE UNDER PRESSURE. NO HARM OR INJURY WAS REPORTED. DEVICE EVAL: THE SUSPECT DEVICE HAS NOT BEEN RETURNED FOR EVAL/INVESTIGATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. EVAL METHOD: OTHER, DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. | 0
1966153 | THE MONITOR CAUGHT FIRE. THERE WAS NO PATIENT IN THE ROOM AT THE TIME OF THE EVENT. UNIT WAS OPENED AND DETERMINED THAT THE SOURCE WAS THE BATTERY. THE BATTERY HAD JUST BEEN CHANGED/REPLACED ON 7/12/2010 AND HAD A DATE CODE OF 6/29/2010.MANUFACTURER RESPONSE FOR BATTERY, PHYSIOLOGICAL MONITOR, (B) (6) AND MONITOR, PHYSIOLOGICAL, PROPAQ CS: I HAVE BEEN IN COMMUNICATION WITH A R & D BATTERIES REPRESENTATIVE. HERE IS THEIR RESPONSE "IN OUR EVALUATION WE FOUND THE POWER SUPPLY TO THE UNIT FAILED AND ONE OF OUR BATTERIES IN THE PACK WAS OVERCHARGED BY THIS FAILURE. IT IS OUR DETERMINATION THAT THERE WAS A FAILURE IN THE CHARGING CIRCUIT OF THE INSTRUMENT. ONE OF THE BATTERIES IN THE UNIT IS STILL FUNCTIONAL, SINCE THIS IS ATTACHED TO THE OTHER BATTERY THROUGH THE CIRCUIT BOARD, AND THE DAMAGED BATTERY IS ONLY CONNECTED THROUGH THE CONNECTOR THAT IS INTERFACED WITH THE DEVICE, IT IS OUR CONCLUSION THAT AN OUTSIDE INFLUENCE (OVER CHARGE) FROM THE DEVICE CAUSED THIS FAILURE."THE BATTERY HAS A FUSE (PART OF THE BATTERY PACK) ON AN INTERFACE BOARD AND IT IS HARD TO BELIEVE THAT THIS DEVICE DID NOT FAIL. | 1
1966439 | AT 5:00 P.M., A PATIENT'S MONITOR SCREEN AT THE DESK LOST COMMUNICATION DISPLAYING "NO COMMUNICATION" MESSAGE. APPROXIMATELY 10 SECONDS LATER, ANOTHER PATIENT'S SCREEN LOST COMMUNICATION. A TOTAL OF 6 PATIENTS MONITOR SCREENS (WAVEFORMS) WENT DOWN WITHIN SECONDS OF EACH OTHER. STAFF MANUALLY RE-ENTERED PATIENT INFORMATION AND TELEMETRY MONITORING RESUMED WITHIN APPROXIMATELY 7 MINUTES. BIOMED ENGINEERING PERFORMED A FULL OPERATIONAL CHECK OUT OF THE UNIT AND FOUND UNIT FUNCTIONAL; PENDING VISIT BY GE MEDICAL SYSTEMS FOR FURTHER EVALUATION. | 1
1966472 | AFTER COMING OFF BYPASS AND DURING THE ECHOCARDIOGRAM, ALL THE PATIENT'S HEMODYNAMIC MONITORING SCREENS WENT BLANK. THE SURGEONS DID NOT HAVE ACCESS TO THE PATIENT'S VITAL SIGNS FOR ABOUT ONE MINUTE. THE MONITORS CAME BACK ON AND NO FURTHER ISSUES.MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, SOLAR 9500:GE ENGINEER'S ANALYSIS OF THE SYSTEM'S REBOOT LOG FILE IS THE REBOOT OCCURRED DUE TO A TREND REQUEST FROM A CENTRAL INFORMATION CENTER (CIC). ACCORDING TO THE GE ENGINEER, THIS OCCURRENCE HAS BEEN AN ONGOING ISSUE WITH SOFTWARE VERSIONS PRIOR TO VERSION 4C. THE MONITORING SYSTEM IN QUESTION HAD VERSION 2C SOFTWARE. | 1
1966474 | WHEN A PATIENT IS DISCHARGED FROM THE MONITOR, THE MONITOR IS PUT INTO STANDBY MODE. THE PROBLEM IS THAT THE XDS BOX WILL SWITCH THE MONITOR BACK INTO MONITORING MODE AFTER ABOUT 10 MINUTES. BECAUSE NO PATIENT IS CONNECTED, THERE IS AN IMMEDIATE LEADS OFF ALARM AT THE CENTRAL STATION. WITH 53 MP2 MONITORS, THIS REPRESENTS A SIGNIFICANT QUANTITY OF NUISANCE ALARMS. OUR FEAR IS STAFF WILL BECOME CARELESS WHEN SILENCING ALARMS AND PUTTING MONITORS BACK INTO STANDBY.====================== HEALTH PROFESSIONAL'S IMPRESSION======================INAPPROPRIATE ALARMS CAUSED BY FAULTY OPERATION OF THE MONITOR. | 1
1968647 | THE CUSTOMER REPORTS INCIDENTS IN WHICH THE PT'S MEDICAL RECORD NUMBERS MATCH UP WITH EXISTING REQUEST NUMBERS. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1969996 | OUR REP IS REPORTING TO US THAT DURING AN ARTHROSCOPIC PROCEDURE, THEY COULD NOT CONTROL THE PUMP'S PRESSURE OVER 50. THIS RESULTED IN THE PT'S JOINT SPACE BEING OVERINFLATED SEVERAL TIMES. AT THIS POINT, THEY STOPPED USING THE PUMP; IN AN ATTEMPT TO AVOID COMPARTMENT SYNDROME, THE SURGEON APPLIED PRESSURE/MASSAGED THE SWOLLEN AREA TO MOVE THE FLUID OUT, AND THIS CAUSED SOME EXTRAVASATION. AN EXTRA 6 OR 7 MINS WERE ADDED ONTO THE PROCEDURE AND THE PROCEDURE WAS CONCLUDED SUCCESSFULLY WITHOUT FURTHER ISSUE OR HARM TO THE PT. MITEK IS AT THIS POINT IN TIME IN THE INFO GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A F/U REPORT. | 0
1973393 | A CUSTOMER REPORTED INCORRECT TIME/DATE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. (B)(4). | 0
1974094 | THE AGFA REPRESENTATIVE INITIATED AN UPDATE DURING THE WORK DAY RESULTING IN THE INABILITY TO PROCESS RADIOGRAPHS TAKEN OF THE PATIENT. ULTIMATELY ONE IMAGE WAS LOST DUE TO MACHINE FAILURE. FORTUNATELY, THE ONE SALVAGED IMAGE WAS SUFFICIENT AND PATIENT DID NOT HAVE TO RETURN FOR ANOTHER IMAGE.AGFA REPRESENTATIVE CAUSED THIS MALFUNCTION WHEN HE DID THE UPGRADE. IN THE FUTURE, HE WILL CALL AHEAD WHEN AN UPGRADE IS NEEDED TO MAKE SURE THAT THERE IS NO ONE IN IMAGING. | 1
1974716 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. (B)(6) (CARIDIAN BCT EMPLOYEE) WAS AT A CUSTOMER SITE TRAINING. A TPE PROCEDURE WAS BEING PERFORMED ON A PT WITH TTP. THE CUSTOMER CHOSE TO SET THE AC RATIO AT 15. THERE APPEARED TO BE CLUMPS/CLOTS IN THE FILTER BELOW THE RESERVOIR. FLUID LEVEL IN THE RESERVOIR NEVER WENT BELOW THE LOWER LEVEL SENSOR, BUT THERE APPEARED TO BE LOTS OF TINY BUBBLES WHICH APPEARED TO BE COMING FROM THE RETURN PRESSURE SENSOR AND GOING DOWN THE RETURN LINE AND INTO THE WARMER LINE THAT WAS CONNECTED TO THE RETURN LINE. NO ALARMS OCCURRED. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR AIR TO THE PT. (B)(6). THERE IS A POTENTIAL FOR SPECTRA OPTIA TO RETURN SMALL AMOUNTS O AIR TO THE PT THROUGH THE RETURN LINE WHEN THERE IS A SMALL LEVEL OF PLATELET AGGREGATION IN THE RETURN RESERVOIR. THIS AGGREGATION CAN BE CAUSED BY THE USE OF AN INSUFFICIENT LEVEL OF ANTICOAGULATION FOR A GIVEN PT. THE LEVEL OF AGGREGATION MAY BE INSUFFICIENT TO ACTIVATE THE ALARMS DESIGNED TO DETECT THIS CONDITION AT THE CURRENT ALARM LIMITS. A SERVICE CALL WAS PLACED FOR CHECKOUT OF THE EQUIPMENT. RDF ANALYSIS OF THE INCIDENT WAS PERFORMED. BASED ON THE RUN DATA FILE ANALYSIS AND PICTURES TAKEN DURING THE RUN, IT CAN BE CONCLUDED THAT A PLATELET CLUMP BLOCKED THE LOW LEVEL SENSOR WHILE BEING STUCK IN THE RESERVOIR FILTER. THIS CAUSED THE FLUID TO GO BELOW THE LOW LEVEL SENSOR AND ALLOW AIR TO BE DRAWN INTO THE RETURN LINE. BASED ON AIM IMAGE ANALYSIS, IT IS CLEAR THAT PLATELET CLUMPS WERE NOT PRESENT IN THE CHANNEL. FRESH FROZEN PLASMA WAS BEING USED AS A REPLACEMENT FLUID DURING THIS PROCEDURE AND MAY HAVE CAUSED THE ACTIVATION OF PLATELETS IN THE RESERVOIR. A CLOT IN THE FILTER BLOCKED THE LEVEL SENSOR ENOUGH TO CAUSE AIR TO BE DRAWN INTO THE RETURN LINE. A SERVICE CALL WAS PLACED FOR CHECKOUT OF THE EQUIPMENT. ALL LEVEL AND PRESSURE SENSORS CHECKED OUT WITH NO ISSUES; THE EQUIPMENT PASSED AUTOTEST AND A SALINE RUN WITH NO PROBLEMS. CONCLUSION: THE DEVICE OPERATED AS INTENDED. | 0
1975377 | DURING PREPARATION FOR USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE MONITOR WAS BLANK. AS A RESULT, AN ALTERNATE DEVICE WAS EMPLOYED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY AND THERE WERE NO ADVERSE CONSEQUENCES TO THE PT AS A RESULT OF THIS EVENT. EVALUATION IN PROGRESS, BUT NOT YET CONCLUDED. | 1
1978661 | SEVERAL OF THE (30 EACH) MOSIAQ WORKSTATIONS, INCLUDING THE ONES AT THE LINACS WERE LOCKING-UP, ONE OR TWO AT A TIME. THE PHYSICIST ON DUTY REPORTED TO IS AND MOSIAQ, WHO WAS LOGGED IN REMOTELY, TO SHUT-DOWN THE SERVER SINCE THE IX WAS THE ONLY MACHINE WORKING AT THIS TIME. REPRESENTATIVES FROM MOSIAQ REVIEWED THE SOFTWARE AND FILES AND RE-BOOTED THE SERVER. ABOUT ONE HOUR LATER ALL SYSTEMS WERE UP, EXCEPT FOR ONE IN DOSIMETERY. IT WILL NEED TO BE REPLACED WITH A NEW COMPUTER. THE FAULTY COMPUTER IN DOSIMETERY HAS SUBSEQUENTLY BEEN REPLACED. MOSIAQ APPEARS TO BE MORE STABLE SINCE. STATEMENT BY DAILY CALIBRATION TECHNICIAN: A SERVER WAS REBOOTED AND I AM NOT SURE IF ACCESS TO THE SERVER WAS REQUIRED FROM THE OUTSIDE VENDER. | 1
1981197 | FALSE POSITIVE FOR COCAINE. URINE SCREEN. DIAGNOSIS OR REASON FOR USE: OPIATE CONTRACT. | 0
1982048 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE STRIPPER PLATE SPLASHED BLOOD ONTO THE BARCODE READER ON THE COULTER LH750 ANALYZER. WHEN BLOOD GETS ON THE BARCODE READER, IT IS CLEANED WITH DISINFECTANT. ALL OPERATORS WEAR PERSONAL PROTECTIVE EQUIPMENT (PPE). THERE WAS NO EXPOSURE TO MUCOUS MEMBRANES OR OPEN LESIONS. THE OPERATOR DID NOT SEEK MEDICAL ATTENTION AND PATIENT TREATMENT WAS NOT AFFECTED IN THIS EVENT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(4) 2011 AND REPLACED TWO CHECK VALVES IN THE NEEDLE WASTE LINE. PERFORMANCE OF THE ANALYZER WAS VERIFIED. THE ROOT CAUSE WAS TWO BAD CHECK VALVES ON NEEDLE WASTE LINE. AS PER PRODUCT LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
1982208 | DEVICE TO RUN OVER 24 HOURS. MEDICATION RAN IN OVER 14 HOURS. DOSE OR AMOUNT: CEFAZOLIN 6PM, FREQUENCY: OVER 24 HRS, ROUTE: IV. DATES OF USE: (B)(6) 2011. DIAGNOSIS OR REASON FOR USE: ABDOMINAL ABCESS. | 0
1982669 | CUSTOMER CLAIMS READING 30 TO 40 POINTS HIGH WHEN COMPARED TO DOCTOR. CUSTOMER STATES UNIT WAS SHOWING IN THE 180 OR OVER 200. CUSTOMER STATES HER NORMAL PRESSURE IS 120/80 OR LESS. CUSTOMER CLAIMS SHE WAS BEING TREATED FOR EXTREME HIGH BLOOD PRESSURE BY HER DOCTOR ACCORDING TO HER READINGS ON HER UNIT. CUSTOMER STATES BECAUSE OF THE MEDICINE SHE WAS GIVEN, SHE CRASHED HER CAR AND BLACKED OUT. CUSTOMER STATES THAT WHEN SHE CRASHED HER CAR, SHE WAS BEING OVER TREATED WITH MEDICINE BY HER DOCTOR AND HER BLOOD PRESSURE BOTTOMED OUT BECAUSE OF HER UNIT. CUSTOMER WAS TREATED BY AMBULANCE AND WAS IN THE HOSPITAL FOR SIX DAYS. CUSTOMER CLAIMS ON APPROX (B)(6) 2008, HER DOCTOR PRESCRIBED AND ORDERED HER A HOME BP MONITOR. CUSTOMER CLAIMS DUE TO A HISTORY OF EXTREMELY HIGH BP, HER MD ORDERS WERE TO TAKE A READING 2-3 TIMES PER DAY AND PHONE IN THE READINGS. CUSTOMER CLAIMS MD WOULD THEN ADJUST HER MEDICATION BASED ON THE MONITOR'S BP READINGS. CUSTOMER CLAIMS FOR THE LAST SEVERAL MONTHS, THE UNIT WOULD GIVE AN ERROR MESSAGE AND NOT PROVIDE A READING. CUSTOMER STATES SHE WOULD THEN TAKE MULTIPLE BACK-TO-BACK READINGS UNTIL A VALUE WAS PROVIDED ON THE HOME MONITOR. CUSTOMER ALSO STATED SHE WOULD WRAP THE CUFF VERY TIGHTLY TO THE POINT OF PAIN TO ASSURE A TIGHT FIT IN AN ATTEMPT TO GET A READING ON THE FIRST TRY. CUSTOMER CLAIMS HER MD DID NOT HAVE HER COME INTO THE OFFICE TO TAKE A PROFESSIONAL READING ON THE OFFICE EQUIPMENT PRIOR TO ADJUSTING HER MEDICATION. CUSTOMER STATES NO COMPARISON READINGS DONE WITH HER HOME UNIT ON ROUTINE VISITS TO THE DOCTOR'S OFFICE. CUSTOMER STATES THE BP CUFF IS ORIGINAL TO THE UNIT. THE CONSUMER HAS NOT RETURNED TO UNIT FOR INSPECTION TO DATE. IF THE CUSTOMER PROVIDES THE UNIT FOR INSPECTION, A FOLLOW UP REPORT WILL BE SUBMITTED. ERROR MESSAGE INDICATORS ARE: EE - CUFF EITHER OVER OR UNDER-INFLATED. CUFF NOT APPLIED CORRECTLY. MONITOR COULD NOT DETECT PULSE WAVE. AIR PLUG IS NOT CONNECTED. CORRECTION - REMOVE THE ARM CUFF. READ "TAKING A MEASUREMENT." REAPPLY CUFF AND TAKE ANOTHER MEASUREMENT. REMAIN STILL UNTIL THE MEASUREMENT IS COMPLETE. | 0
1983113 | THE PATIENT'S FILEMAKER HOLTER REPORT WAS FAXED TO CLINIC AT RETIREMENT VILLAGE. THE FOLLOWING MORNING, UPON REVIEW OF THE ".ZPT" FILE TO VERIFY THAT IT MATCHED THE DATA IN FILEMAKER, IT WAS DISCOVERED THAT ANOTHER PATIENT'S DATA HAD MERGED INTO THIS PATIENT'S REPORT. AT THIS POINT, THE INFORMATION WAS REVIEWED WITH AN EP PHYSICIAN, WHO FELT THAT NO HARM WOULD BE DONE TO THE PATIENT FROM THE INCORRECT DATA. THE OUTSIDE CLINIC WAS CALLED TO INFORM THEM OF THE ISSUE AND THE CORRECTED REPORT WAS FAXED. THE FIRST REPORT SHOWED AVERAGE HEART RATE OF 73 WITH MINIMUM OF 51 BEATS PER MINUTE AND MAXIMUM OF 107. THE AMENDED REPORT NOTED AVERAGE HEART RATE IN THE 50'S WITH A MAXIMUN OF 90 BEATS PER MINUTE. THE PATIENT DID HAVE BRADYCARDIA AT 0200 WITH THE LOWEST HEART RATE OF 32. PATIENT HAD EPISODES OF JUNCTIONAL ESCAPE RHYTHM.PHYSICIAN FELT THERE WAS NO HARM ASSOCIATED WITH THIS EVENT. MERGING APPEARS TO HAPPEN WHEN IMPORTING DATA INTO FILEMAKER FROM TWO SATELLITE CLINICS AT THE SAME TIME. THERE IS NO MAINTENANCE ON THE PHILIPS SCANNERS, EXCEPT FOR SOFTWARE UPDATES. THE SOFTWARE HAD RECENTLY BEEN UPDATED ON THE SCANNERS PRIOR TO THIS EVENT. WHEN PROBLEMS ARE ENCOUNTERED, THE SOFTWARE IS RE-INSTALLED. MERGING APPEARS TO ONLY HAPPEN DURING FILEMAKER IMPORTATION AND IT IS DIFFICULT TO DETERMINE IF THE GLITCH IS WITH FILEMAKER OR PHILIPS. | 1
1984322 | THE CUSTOMER STATED DISCREPANT RESULTS WERE GENERATED ON THE CELL-DYN 1700 ANALYZER. A PATIENT SAMPLE WAS TESTED THREE TIMES AND GENERATED HEMOGLOBIN RESULTS OF 5.0, 10.3 AND 15.9 G/DL. WHEN THE SPECIMEN WAS TESTED A FOURTH TIME, THE VALUE WAS VERY SIMILAR TO THE PREVIOUS RESULT OF 15.9 G/DL. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><(B)(4). A CLOG AT THE PRE-MIX CUP WAS SUSPECTED. THE CUSTOMER WAS INSTRUCTED TO WASH THE PRE-MIX CUP WITH DILUTED BLEACH AND MONITOR THE ISSUE. THE CUSTOMER CONFIRMED THAT THERE WERE NO FURTHER ISSUES AFTER COMPLETING THE RECOMMENDED ACTIONS. THE ISSUE WAS LIKELY CAUSED BY DEBRIS IN THE PRE-MIX CUP. THE CELL-DYN 1700 SYSTEM OPERATOR'S MANUAL PROVIDES BASIC TROUBLESHOOTING INFORMATION FOR PROBLEM IDENTIFICATION, ISOLATION, AND CORRECTIVE ACTION. INSTRUCTIONS FOR OBTAINING TECHNICAL ASSISTANCE ARE INCLUDED AS WELL. A NON-STATISTICAL TREND (NST) COMPLAINT REVIEW WAS PERFORMED FOR THE REPORTED ISSUE FROM (B)(4) 2010 THROUGH (B)(4) 2011. NO NST WAS IDENTIFIED DURING THE SEARCHED PERIOD. BASED ON THE INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED FOR THE CELL-DYN 1700 RELATED TO THE REPORTED ISSUE. THE ISSUE WAS ISOLATED TO DEBRIS IN THE PRE-MIX CUP AREA, WHICH WAS RESOLVED BY CLEANING. | 0
1987066 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THE DEVICE DID NOT RESPOND. THE APPLICATION SLOWED DOWN. THE DEVICE WAS USED FOR THE PROCEDURE. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT AS A RESULT OF THIS EVENT. | 1
1988210 | THE MONITOR WAS BEING USED TO MONITOR A PT IN SURGERY. THE SCREEN GRAYED OUT IN AREAS, MAKING THE READINGS NOT VISIBLE. ANOTHER MONITOR WAS BROUGHT IN TO MONITOR THE PT. THE COMPANY REP WAS CONTACTED, WHO CHANGED OUT THE INTERNAL COMPONENT OF ALL OF THE MONITORS. | 1
1989991 | (B)(6) WAS PROVIDED AN ADVOCATE DUO GLUCOMETER, SERIAL NUMBER (B)(4), MODEL TD-3223, WHICH PRODUCED INACCURATE RESULTS. THE GLUCOMETER AT ISSUE WAS ACTUALLY COMPARED TO THE GLUCOMETER IN (B)(6) HOSP ON (B)(6) 2009 AND FOUND TO BE ERRONEOUS. FOR A MORE DETAILED ACCOUNT OF THE OCCURRENCE PLEASE SEE THE COMPLAINT AT LAW. DATES OF USE: (B)(6) 2009 - (B)(6) 2009. DIAGNOSIS OR REASON FOR USE: TEST GLUCOSE LEVEL FOR DIABETIC PT. | 0
1990059 | CRITICALLY ILL PATIENT WAS IN CATH LAB HAVING AN EMERGENCY C-PORT (CARDIAC CATHETERIZATION) PROCEDURE FOR TREATMENT OF AN ACUTE MI. ATTEMPTED TO PERFORM CARDIAC OUTPUT MEASUREMENT. THE SYSTEM KEPT GIVING UNSTABLE TEMPERATURE WARNINGS FOR THE CARDIAC OUTPUT FUNCTION. ALTERNATIVE METHODOLOGY HAD TO BE USED TO PERFORM MEASUREMENT. A SECOND PATIENT REQUIRED A PROCEDURE IMMEDIATELY AFTER THE MCKESSON REPRESENTATIVE MADE A CHANGE TO THE NIC CARD CONFIGURATION FOR THE RT DISPLAY COMPUTER TO CORRECT THE CARDIAC OUTPUT ISSUE; THE SYSTEM WOULD NOT REBOOT. THE PATIENT CARE WAS DELAYED AND THE PATIENT HAD TO BE TAKEN TO AN ALTERNATE ROOM FOR THEIR PROCEDURE. THE MCKESSON REPRESENTATIVE USED DIAGNOSTIC TOOLS TO RESTART THE SYSTEM. THERE HAVE NOT BEEN ANY ISSUES WITH THE CARDIAC OUTPUT FUNCTION SINCE THE CHANGE TO THE NIC CARD CONFIGURATION AND THE REBOOT.MANUFACTURER RESPONSE FOR HEMODYNAMIC MONITOR SYSTEM, MCKESSON: SERVICE CALL PLACED IMMEDIATELY AND REPRESENTATIVE SENT TO EVALUATE AND REPAIR SYSTEM.MANUFACTURER RESPONSE FOR HEMODYNAMIC MONITOR SYSTEM, MCKESSON: SERVICE CALL PLACED TO MCKESSON AND REPRESENTATIVE CAME TO EVALUATE AND REPAIR THE SYSTEMS. | 1
1991991 | DOCTOR ATTEMPTED TO USE AN ULTRASOUND MACHINE PRIOR TO PERFORMING A STAT C-SECTION ON A PATIENT TO DETERMINE IF THERE WAS A FETAL HEART RATE. THE MACHINE POWERED ON, BUT DID NOT COMPLETE THE BOOT SEQUENCE IN THE EXPECTED FASHION. THE DEVICE WAS UNPLUGGED AND ANOTHER ATTEMPT WAS MADE, BUT THE PROBLEM PERSISTED. | 1
1993905 | DURING THE (B)(6) OF 2009, OUR RADIOLOGY PRACTICE SHIFTED FROM FILM-BASED MAMMOGRAPHY TO DIGITAL MAMMOGRAPHY. WITH DIGITAL MAMMOGRAPHY, OUR RADIOLOGISTS BEGAN PERFORMING ONLINE DIAGNOSIS USING CARESTREAM RIS-DIC SYSTEM. WHEN USING THE SYSTEM, THE RADIOLOGIST WILL CHOOSE A DROP DOWN MENU TO SELECT A BREAST DENSITY TYPE AND A BIRADS CATEGORY (OVERALL DIAGNOSIS). USING THE MOUSE, THE RADIOLOGIST MAY SELECT A PARTICULAR BREAST DENSITY OR A BIRADS CATEGORY ON AN ADMINISTRATIVE MONITOR, TO SHIFT THE MOUSE TO FOCUS TO THE (B)(6)-MONITORS TO SCROLL THROUGH DIGITAL IMAGES. IN SO DOING, THE SELECTIONS PREVIOUSLY MADE ON THE ADMINISTRATIVE MONITOR FOR BREAST DENSITY AND BIRADS CATEGORY CHANGE AS THE RADIOLOGIST SCROLLS THROUGH IMAGES, EVEN THOUGH THE MOUSE IS POINTED TO A (B)(6) MONITOR. THE BIRADS CATEGORY ULTIMATELY SELECTED HAS SEVERAL CONSEQUENCES. IT IS IMPORTED AS A FINAL RESULT IN MAMMOGRAPHY REPORTS. IT IS USED TO DETERMINE THE CONTENT OF THE LAY LETTER TO THE PTS, INFORMING THEM THAT THE RESULT WAS NORMAL OR THAT ADDITIONAL IMAGING IS NEEDED. IT IS ALSO TO ESTABLISH THE INTERVAL FOR THE PTS NEXT APPOINTMENT, RANGING FROM THE IMMEDIATE NEED TO RETURN FOR ADDITIONAL IMAGING TO SIX MONTHS FOR A SHORT F/U OR TWELVE MONTHS FOR AN ANNUAL REMINDER. WHEN (B)(6) FIRST BECAME AWARE OF THE DEFICIENCIES IN (B)(6) 2009, WE NOTIFIED CARESTREAM PERSONNEL THROUGH A SERIES OF EMAILS THAT "USE OF A DROP DOWN MENU IS UNACCEPTABLE FOR MEDICAL DIAGNOSIS AND IT NEEDED TO BE MODIFIED. A SOLUTION WAS PROPOSED TO SWITCH THE DROP DOWN MENU TO DISCRETE RADIO BUTTONS. (B)(6) PERSONNEL HAVE FOLLOWED UP WITH CARESTREAM ON THIS ISSUE NUMEROUS TIMES. WE HAVE BEEN UNABLE TO GET CARESTREAM HEALTH TO DEAL WITH THIS ISSUE SERIOUSLY AND TO DEVELOP A SOLUTION. DATES OF USE: (B)(6) 2005 -- (B)(6) 2011. | 1
1994231 | THE CONNECTOR FOR A BIOSENSE WEBSTER NAVISTAR THERMOCOOL ABLATION CATHETER FIT INTO A SJM REFLEXION HD CATHETER. THIS LED TO THE REFLEXION HD BEING DIRECTLY ATTACHED TO THE STOCKERT GENERATOR. THE CONNECTOR OF BOTH THE REFLEXION HD AS WELL AS THE NAVISTAR THERMOCOOL HAS THE SAME TYPE OF "END" THAT ALLOWS IT TO BE CONNECTED TO EITHER CATHETER. THE OTHER END OF THE CONNECTOR IS VERY DIFFERENT FOR EACH CATHETER. THE REFLEXION HD HAS "TAILS" WHILE THE NAVISTAR HAS A "REDEL" CONNECTOR. UNFORTUNATELY, THE END OF THE CONNECTOR WAS NOT OBSERVED BY THE PHYSICIAN BEFORE THE CONNECTION WAS MADE SINCE IT WAS HIDDEN BY THE STERILE DRAPE. EVEN THOUGH RF ENERGY WAS NOT DELIVERED, THE STOCKERT GENERATOR, WHICH SENDS OUT 50HZ SIGNALS @ BASELINE SENT ENERGY THROUGH THE REFLEXION HD CATHETER WHICH RESULTED IN THE HEATING OF BLOOD IN THE LEFT ATRIUM WHICH APPEARED AS ECHO CONTRAST "BUBBLES" ON THE ICE CATHETER. ADDITIONALLY, THE ECG SIGNAL WAS ABNORMAL DUE TO THE INCORRECT CONNECTOR. ONCE BOTH THESE ISSUES WERE IDENTIFIED BY THE PHYSICIAN. THE PROCEDURE WAS ABORTED, ALL CATHETERS REMOVED AND THE PATIENT EVALUATED FOR SEQUELAE. THE PATIENT DEVELOPED NO APPARENT PROBLEMS DUE TO THIS INCIDENT. AFTER THE PROCEDURE, THE PHYSICIAN DID A "BENCH TEST" OF THE SAME CATHETER SET UP THAT WAS IN THE PATIENT AND WAS ABLE TO REPRODUCE THE SAME "BUBBLES" IN A WATER BATH. ====================== HEALTH PROFESSIONAL'S IMPRESSION======================BY CONNECTING THE REFLEXION HD TO THE STOCKERT GENERATOR, "ENERGY" WAS DELIVERED EVEN THOUGH THE RF DELIVERY BUTTON WAS NOT PUSHED. THIS WAS PROBABLY DUE TO THE FACT THAT THE STOCKERT GENERATOR ALWAYS SENDS OUT A 50 HZ SIGNAL WHEN THE MACHINE IS TURNED ON IN THE STANDBY MODE. THIS "ENERGY" HEATED THE BLOOD IN THE PATIENT'S LEFT ATRIUM AND CAUSED THE "BUBBLES" THE PHYSICIAN SAW ON ECHO.====================== MANUFACTURER RESPONSE FOR CONNECTORS, REFLEXION HD AND BIOSENSE NAVISTAR THERMOCOOL CATHETER AND STOCKERT GENERATOR, (BRAND NOT PROVIDED)======================THEY ASKED ABOUT DETAILS OF SITUATION AS WELL AS PATIENT CONDITION. | 0
1994313 | THE ROTATOR BREAKS WHEN INJECTING CONTRAST. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED 3 DEFECTIVE DEVICES/EVENTS BUT DID NOT PROVIDE SEPARATE INFO OR CLINICAL DETAILS FOR THE ADD'L EVENTS. THE CUSTOMER IS NOT EXPECTED TO RETURN ANY USED DEVICES FOR EVAL. THEREFORE, THIS SINGLE FORM FDA 3500A REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL: THE SUSPECT DEVICES WILL NOT BE RETURNED FOR EVAL. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL IS COMPLETED. EVAL METHOD: DEVICE HISTORY RECORD WAS REVIEWED. | 0
1994892 | THE TUBING BLEW OFF THE INJECTOR UNDER PRESSURE. THERE WAS SPRAY OF CONTRAST. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE DEVICES BUT DID NOT PROVIDE INFO OR CLINICAL DETAILS FOR THE ADDITIONAL EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A REPORT WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL. THE DEVICE HAS NOT BEEN RETURNED FOR EVAL. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. EVAL: METHOD: DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. | 0
1996256 | THE ACCOUNT STATED THAT AN ELEVATED TROPONIN RESULT WAS GENERATED. THE ACCOUNT IS USING A CUTOFF OF 0.03 NG/ML. AN INITIAL TROPONIN RESULT OF 0.056 NG/ML WAS GENERATED WITH A REPEAT RESULT OF 0.029 NG/ML. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR CHECKED THE INSTRUMENT AND FOUND THE VACUUM LINES FOR VACUUM VESSELS 1 AND 2 CONTAMINATED WITH PARTICULATE MATTER. IN ADDITION A DAMAGED VALVE WAS FOUND. THE FSR CLEANED THE VACUUM LINES AND REPLACED THE DAMAGED VALVE. THE CUSTOMER INDICATED THAT NO FURTHER PROBLEMS WITH TROPONIN RESULTS WERE SEEN. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE ARCHITECT STAT TROPONIN-I REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. | 0
1998139 | A CUSTOMER CONTACTED BAXTER'S (B)(4) AND REPORTED RECEIVING A CHECK PATIENT LINE ALARM ON THE HOMECHOICE MACHINE DURING THE INITIAL DRAIN. THE PATIENT STATED THAT THERE IS AIR IN THE PATIENT LINE. THE TECHNICAL SERVICE REPRESENTATIVE (TSR) ASSISTED THE HOME PATIENT (HP) WITH ENDING THERAPY. THE HP WILL START OVER WITH NEW SUPPLIES. (B)(4) CONTACTED THE PATIENT ON (B)(6) 2011. ACCORDING TO THE PATIENT, HE NOTIFIED HIS NURSE OF THIS EVENT. THE PATIENT STATED HE DID NOT NOTICE ANY HOLES OR LEAKS IN THE SUPPLIES, ONLY SOME FIBRIN IN THE DRAIN BAG. THE PATIENT HAS RESUMED THERAPY WITH NO FURTHER ISSUES. THERE WAS NO PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. (B)(4). THE SAMPLE IS NOT AVAILABLE FOR EVALUATION, THEREFORE, BAXTER CANNOT DETERMINE ROOT CAUSE. SINCE THE LOT NUMBER IS UNKNOWN A BATCH REVIEW CANNOT BE PERFORMED. SHOULD ANY ADDITIONAL INFORMATION BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED.><(B)(4). THIS REPORT IS FOR A CHECK PATIENT LINE ALARM AND WAS NOT CONFIRMED IN THE LAB DUE TO A LACK OF SAMPLE. THE LOT NUMBER IS UNKNOWN; THEREFORE, A BATCH REVIEW WAS NOT PERFORMED. FROM THE DATA WITHIN THE COMPLAINT INFORMATION, THE ROOT CAUSE WAS NOT IDENTIFIED. LABELING REVIEW FOUND LABELING TO BE ADEQUATE. SIMILAR REPORTS HAVE BEEN RECEIVED FOR THE REPORTED PROBLEM. BAXTER WILL CONTINUE TO MONITOR SIMILAR REPORTS TO DETERMINE IF FURTHER ACTIONS ARE REQUIRED. | 0
1998282 | ANESTHESIA UNIT WAS DELIVERING AT HIGHER CONCENTRATION OF ANESTHETIC AGENT THAN WAS SET. PATIENT RECEIVED 3X MORE ANESTHETIC. PATIENT WAS REVERSED INSTANTLY BY ANESTHESIOLOGIST AND SUFFERED NO ILL EFFECTS. THE SERVICE HISTORY OF THE UNIT WAS CLEAN. THERE IS NO ERROR MESSAGE WITH THIS TYPE OF EVENT. BIOMED HAS DETERMINED THAT THE PROBLEM IS WITH THE FRESH GAS MODULE WHICH HAS BEEN REPORTED BEFORE BY US.======================MANUFACTURER RESPONSE FOR DATEX OHMEDA ANESTHESIA DELIVERY, DATEX OHMEDA======================THIS IS SECOND EVENT OF THIS TYPE. WE ARE STILL AWAITING THEIR EVALUATION FROM FIRST EVENT. | 0
2001760 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
2002463 | THE CUSTOMER CALLED WITH A BLOOD GLUCOSE READING OF 438 MG/DL. THE CUSTOMER HAD ALREADY TREATED WITH A BOLUS. THE CUSTOMER THEN MENTIONED THAT HE HAD BEEN HOSPITALIZED FOR BLOOD GLUCOSE LEVELS OVER 1000 MG/DL SEVERAL MONTHS AGO. NOTHING FURTHER WAS REPORTED. CURRENTLY, IT IS UNK WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE, THEREFORE, CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | 0
2002605 | DURING AN INSPECTION OF THE WORKSTATION, THE FUSE OF THE WORKSTATION WAS FOUND CHARRED. THERE WAS NO USER OR PT INJURY ASSOCIATED WITH THIS EVENT. THE WORKSTATION WAS REPORTEDLY REMOVED FROM SERVICE. THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO (B)(4) FOR EVALUATION. THE USER FACILITY WAS PROVIDED WITH A REPLACEMENT DEVICE. THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE REGISTRATION NUMBER FOR THIS CORRECTION IS (B)(4). THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | 1
2003942 | DURING INTRACRANIAL ANGIOPLASTY WITH STENTING, THE PATIENT WAS UNDER GENERAL ANESTHESIA. THE ANESTHESIA MACHINE QUIT WORKING; EVERYTHING TURNED OFF AND SCREEN WENT BLANK. THE PATIENT WAS MANUALLY VENTILATED AND ANOTHER MACHINE WAS BROUGHT IN. NO HARM TO PATIENT.WE SUSPECT THAT WE HAVE SEEN THIS PROBLEM BEFORE AS WE HAVE EXPERIENCED SIMILAR ISSUES, SUCH AS THE SCREEN GOING BLANK, WITH THE ANESTHESIA MACHINES PREVIOUSLY. ALSO, WE ARE NOT AWARE OF DRAEGER NOTIFYING OUR FACILITY THAT NEW PROCESSOR BOARDS WERE BEING IMPLEMENTED DURING THE TIME OF THE IMPLEMENTATION (MAY 2007).======================MANUFACTURER RESPONSE FOR DRAGER FABIUS GS ANESTHESIA MACHINE, DRAGER FABIUS GS======================VISUAL ANALYSIS REVEALED A BURNT COMPONENT (Z10) ON PROCESSOR CONTROL BOARD. THIS PART IS A FERRITE ARRAY USED TO FILTER THE 12V SUPPLY VOLTAGE TO THE MAIN DISPLAY. IN THE EVENT OF A SHORT ON THE BOARD, THE 12V SUPPLY VOLTAGE IS SHORTED TO GROUND AND THE FERRITE BEGINS TO MELT. AS THE COMPONENT FAILS, A SLIGHT BURNING ODOR MAY BE DETECTED AND THE DISPLAY WILL BLANK OUT. ROOT CAUSE OF THE DISPLAY PROBLEM WAS A SHORT CIRCUIT ON THE PROCESSOR CONTROL BOARD RESULTING IN THE FERRITE ARRAY (Z10) FAILING DUE TO THERMAL BREAKDOWN. AS PART OF ROUTINE QUALITY IMPROVEMENT, DRAEGER PHASED OUT THE USE OF THE FERRITE ARRAY COMPONENT ON THE FABIUS PROCESSOR BOARD IN MAY 2007 AND REPLACED IT WITH A ZERO OHMS ARRAY COMPONENT. THE FABIUS GS INVOLVED IN THIS COMPLAINT WAS MANUFACTURED IN 2003. | 1
2009483 | PATIENT COMPLAINED THAT THE TELEMETRY BOX (B) (6) FELT HOT, A SITTER IN THE ROOM CHECKED THE BOX AND BOX WAS INDEED HOT, WHEN SHE OPENED THE TELEMETRY BOX (B) (6), THE BATTERIES LOOKED AS IF THEY WERE MELTING. THE BOX WAS TAKEN OFF THE PATIENT. THE PATIENT WAS NEVER HARMED, THE OUTSIDE OF TELEMETRY BOX LOOKED NORMAL, SENT TO BIO MED.><PATIENT COMPLAINED THAT THE TELEMETRY BOX NUMBER 58 FELT HOT, A SITTER IN THE ROOM CHECKED THE BOX AND BOX WAS INDEED HOT, WHEN SHE OPENED THE TELEMETRY BOX NUMBER 58, THE BATTERIES LOOKED AS IF THEY WERE MELTING. THE BOX WAS TAKEN OFF THE PATIENT. THE PATIENT WAS NEVER HARMED, THE OUTSIDE OF TELEMETRY BOX LOOKED NORMAL, SENT TO BIO MED. | 1
2010315 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADD'L DATA/FAILURE INVESTIGATION: FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION, CAUSING DRAWER FAILURE. | 1
2012318 | IT WAS REPORTED THAT A PT DIED WHILE CONNECTED TO A DELTA MONITOR. THE PT MONITOR ALARMED, ACCORDING TO THE BED "NEIGHBOR". IT WAS FURTHER REPORTED THAT THERE WAS NO CONNECTION TO A CENTRAL OR NURSE CALL SYSTEM INTENDED OR REQUESTED BY THE WARD (HEAD OF THE WARD). DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2016920 | THE USER RECEIVED QUESTIONABLE RESULTS FOR ONE PATIENT URINE SAMPLE THAT WERE ABNORMAL THEN REPEATED AS NORMAL. OF THE DATA PROVIDED, THE RESULTS FOR THE FOLLOWING ASSAYS WERE DISCREPANT. THE INITIAL LEUKOCYTE RESULT WAS 500 LEU/UL AND THE REPEAT RESULT WAS NEGATIVE. THE INITIAL PROTEIN RESULT WAS 500 MG/DL AND THE REPEAT RESULT WAS NEGATIVE. THE INITIAL GLUCOSE RESULT WAS 250 MG/DL AND THE REPEAT RESULT WAS NORMAL. THE USER WAS NOT SURE WHICH RESULTS WERE CORRECT AND DID NOT HAVE ANY INFORMATION ABOUT THE PATIENT'S CONDITION OR WHETHER PATIENT WAS AFFECTED BY ERRONEOUS RESULTS. NO ADVERSE EVENTS WERE ALLEGED REGARDING THE EVENT. THE LOT NUMBER OF THE CHEMSTRIP 10 MD URINE TEST STRIPS WAS 23063844. IT WAS UNKNOWN IF THE INITIAL REPORTER SENT REPORT TO THE FDA.><NO PRODUCT HAS BEEN RETURNED AND THEREFORE NO EVALUATION HAS OCCURRED.><A SPECIFIC ROOT CAUSE COULD NOT BE DETERMINED. NO MATERIALS WERE RETURNED FOR INVESTIGATION. RETENTION MATERIAL WAS TESTED AND WAS FOUND WITHIN SPECIFICATION. | 0
2016943 | THE PATIENT REPORTED OBTAINING FOUR TO FIVE OCCLUSION ALARMS PER DAY ON THE REPORTED INSULIN PUMP SINCE (B)(6) 2011. ON (B)(6) 2011, THE PATIENT WENT TO THE EMERGENCY ROOM WHERE HER BLOOD GLUCOSE LEVEL WAS 400 MG/DL. THE PATIENT WAS ADMITTED TO ICU WITH A DIAGNOSIS OF DKA AND WAS TREATED WITH INTRAVENOUS INSULIN. THE PATIENT REPORTED THE OCCLUSION ALARMS OCCURRED RANDOMLY, AND DURING A BOLUS INJECTION. THE PATIENT ALSO NOTED THE PUMP SOUNDED AS IF IT WAS 'STRUGGLING'. DURING A TELEPHONE CALL TO CUSTOMER SERVICE ON (B)(6) 2011, THE PATIENT NOTED SHE WAS USING AN INFUSION SET THAT EXPIRED IN 2008. DURING THE (B)(6) 2011, CUSTOMER SERVICE TELEPHONE CALL, THE PATIENT CONFIRMED THE INFUSION SET SHE WAS USING WAS CURRENT AND HAD AN EXPIRATION DATE OF 2012; NO LOT NUMBER WAS PROVIDED. TROUBLESHOOTING REVEALED THE PATIENT CHANGED THE INFUSION SET AFTER THE SECOND ALARM, HER INSERTION TECHNIQUE WAS CORRECT, AND THERE WERE NO BENDS OR KINKS IN THE TUBING. THE PUMP WAS RETURNED TO ANIMAS FOR INVESTIGATION. EVALUATION REVEALED THE PUMP WAS DELIVERING INSULIN ACCURATELY, THE SENSOR WAS DETECTING THE CORRECT FORCE, AND WHEN AN OCCLUSION ALARM WAS REPRODUCED, THE PUMP ALARMED WITH AUDIBLE AND SCREEN NOTIFICATION. | 0
2018277 | THE CUSTOMER REPORTS THAT THE ORTHO PROVUE RESULTED A FALSE NEGATIVE REACTION THAT WAS POSITIVE IN MANUAL GEL IN THE ANTIBODY SCREEN TEST. NO INCORRECT RESULTS WERE REPORTED AS A RESULT OF THIS INCIDENT. THE CUSTOMER INDICATES THEIR UNDERSTANDING OF THE POSSIBLE CAUSES OF THE DISCREPANCIES AND INDICATES THEIR CONFIDENCE THAT THE PROVUE IS OPERATING AS EXPECTED. THE CUSTOMER INDICATES THAT EACH EVENT OF THE FALSE NEGATIVE WOULD NOT HAVE PREVENTED THE IDENTIFICATION OF EACH ANTIBODY. THE CUSTOMER UNDERSTANDS THAT NO ADDITIONAL INVESTIGATION WOULD BE PERFORMED. THE CUSTOMER INDICATES THAT SERVICE TO THE PROVUE IS NOT REQUIRED. THE CUSTOMER HAS AGREED TO CONTACT CTS IF ANY ADDITIONAL DISCREPANCIES ARE OBSERVED AGAIN. (B)(4). CUSTOMER STATES PROVUE OPERATING AS EXPECTED. | 0
2019216 | IMPAC DEVICE LOST COMMUNICATION DURING THE DELIVERY OF 123 MU'S. IMPAC STATED 35 REMAINED WHEN THE LINAC COMPUTER DISPLAYED 45, WHICH WAS THE ACCURATE NUMBER. THE VERY OBSERVANT THERAPISTS NOTICED THE DISCREPANCY AND DELIVERED THE CORRECT DOSE. THE PATIENT WAS NOT HARMED.IT APPEARS THAT A SOFTWARE UPGRADE RESOLVED THIS ISSUE. | 1
2021340 | FOUR PORTABLE CHEST IMAGES OF THE PT WERE PERFORMED. DURING EACH INSTANCE, A DELAY OF IMAGES ARRIVING TO THE AGFA PACS OCCURRED. THE CLINICAL CONSEQUENCE WAS A DELAYED DIAGNOSIS ON (B)(6) WHILE THE PT WAS IN RESPIRATORY DISTRESS. THERE WERE OTHER DELAYS ON (B)(6) DURING ROUTINE RADIOGRAPHIC MORNING ROUNDS. THE PT EXPIRED OVER THE WEEKEND; HOWEVER, ACCORDING TO (B)(6) MED CTR, THE DELAY DID NOT IMPACT THE PT'S ULTIMATE OUTCOME (THE CHILD EXPIRED), AND DID NOT CAUSE A DELAY IN RECEIVING ANY SPECIFIC MEDICAL INTERVENTIONS. SEVERAL ISSUES HAVE BEEN IDENTIFIED AS THE POTENTIAL ROOT CAUSE TO THE INCIDENT AND AGFA HEALTHCARE IS ACTIVELY INVESTIGATING EACH ONE. EVAL SUMMARY: ON (B)(6) 2011, AN EMAIL ESCALATION WAS RECEIVED FROM THE (B)(6) MED CTR INDICATING THAT A PROBLEM HAD OCCURRED WITH THE IMPAX 6.3.1 SU 16 (PACS SERVER). DIAGNOSTIC IMAGES TAKEN OF A CHILD WERE DELAYED. OVERALL, A DELAY TIME BETWEEN 45-60 MINUTES FROM THE TIME SOME OF THE FILMS WERE TAKEN TO THE TIME THE PICU TEAM WAS ABLE TO VIEW THE RESULTS. ACCORDING TO (B)(6) (MED CTR) - THIS DELAY DID NOT IMPACT THE PT'S ULTIMATE OUTCOME (THE CHILD EXPIRED) AND DID NOT CAUSE A DELAY IN RECEIVING ANY SPECIFIC MEDICAL INTERVENTIONS. DETERMINATION OF CAUSE: THE DEVICE IN QUESTION (IMPAX 6.3.1 SU 16) WAS NOT THE CAUSE OF THE PRODUCT PROBLEM. AGFA HEALTHCARE IS CONTINUING TO INVESTIGATE THE INCIDENT TO DETERMINE THE ROOT CAUSE. THERE ARE SEVERAL CONTRIBUTING FACTORS THAT MAY RESULT IN IMAGES NOT ALWAYS BEING AVAILABLE IN A TIMELY FASHION FOR VIEWING ON THE IMPAX CLIENT. ALL OF THE FOLLOWING CONTINUE TO BE EXPLORED: DATA INTEGRITY ISSUES WITH DATA MIGRATED FROM LEGACY GE PACS. MIGRATION TECHNIQUES. HIGHLY CUSTOMIZED WORKFLOW TO PROVIDE "QC REQUIRED" OPERATION, AND CROSS CLUSTER OPERATION, BUT A PLAN IS BEING DEVELOPED TO RETURN TO A STANDARD CONFIGURATION. PROBLEMS WITH SOFTWARE/CONFIGURATION, PARTICULARLY WITH SCU AND SCP PROCESSES ON THE 3 WORK FLOW MANAGERS (WFM). POTENTIAL PROBLEMS WITH THE CITRIX NETSCALER LOAD BALANCER. POSSIBLE ARP POISONING. CORRECTIVE ACTION: AGFA HEALTHCARE WAS ABLE TO TEMPORALLY RESOLVE THE PROBLEM FOLLOWING DRIVER UPDATES AND MODIFICATIONS TO CONFIGURATION (DEACTIVATING WORK FLOW MANAGER 3). AGFA SUPPORT AND THE ONSITE CSE WILL CONTINUE TO MONITOR THE PERFORMANCE OF THE SYSTEM UNTIL THE INVESTIGATION AND CORRECTIVE ACTIONS ARE COMPLETE. THE DATE AGFA HEALTHCARE BECAME AWARE OF THE INCIDENT: (B)(4) 2011. | 1
2021737 | UNDERWENT STANDARD PRK WITH THE VISX STAR S4. I HAVE STARBURST IN LOW LIGHT CONDITIONS (AROUND MY HOUSE), HALOS, AND GLARE. LOSS OF CONTRAST SENSITIVITY, BLURRINESS, DISTURBING NUMBER OF FLOATERS. VERY UNHAPPY AND DISTRESSED ABOUT MY SIGHT. TOTAL REGRET. WAS NOT GIVEN THE NAME OF LASER OR CORRECTION TO BE DONE BEFORE PROCEDURE. WAS NOT GIVEN VISX STAR PAMPHLET. WAS NOT GIVEN ANY WRITTEN INFO ON PRK PRIOR TO SURGERY. I WAS ENCOURAGED TO GO UNDER TREATMENT BY THE SURGEON. WAS TOLD THAT I WOULD NOT HAVE PROBLEMS DRIVING AT NIGHT. WAS NOT ASKED WHY I WANTED PROCEDURE. I WAS NOT GIVEN A CONSENT FORM UNTIL MINS BEFORE SURGERY. CONSENT FORM WAS HIDDEN BY ANOTHER PAPER. WAS NOT TOLD THAT THE AREA TREATED WOULD BE SMALLER THAN MY PUPIL SIZE. WAS TOLD AFTERWARD THAT STANDARD AND CUSTOM WOULD NOT MATTER FOR MY TREATMENT. WAS NOT TOLD THAT PROCEDURE COULD INDUCE ASTIGMATISM. | 0
2022864 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE NEEDLE CARTRIDGE ON THE COULTER LH 750 ANALYZER WAS LEAKING BLOOD. NO INJURIES OCCURRED AND MEDICAL ATTENTION WAS NOT SOUGHT. THERE WERE NO REPORTS OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. NO CHANGE TO PATIENT TREATMENT ATTRIBUTED TO OR CONNECTED TO THIS COMPLAINT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(4) 2011 AND CLEARED THE PLUGGED WASTE LINE. THE FSE FLUSHED THE WASTE LINES AND VACUUM LINES WITH HOT WATER. AS PER PRODUCT LABELING, BCI URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
2024968 | AS THE PHYSICIAN APPLIED THE CAUTERY USING THE ERBE ICC, IT APPEARED THAT MORE CAUTERY THAN USUAL OCCURRED AND THE MACHINE DID NOT MAKE THE USUAL SOUNDS THAT IT MAKES DURING A POLYPECTOMY PROCEDURE. THE PHYSICIAN IMMEDIATELY DISCONTINUED THE CAUTERY AND THE ERBE ICC MACHINE WAS EXCHANGED AND THE PROCEDURE CONTINUED. | 0
2026022 | THE CUSTOMER DISCONTINUED USE OF THE HOMECHOICE (HC) MACHINE AND RETURNED IT TO BAXTER FOR AN UNKNOWN REASON. ADDITIONAL INFORMATION WAS OBTAINED FROM GLOBAL PHARMACOVIGILANCE INDICATING THE PATIENT PASSED AWAY ON AN UNKNOWN DATE. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED WHETHER AN AUTOPSY WAS PERFORMED. A CAUSALITY STATEMENT WAS NOT PROVIDED FOR THE FATAL EVENT. THE NURSE DECLINED TO PROVIDE ADDITIONAL INFORMATION. (B)(4). THE ASSIGNABLE CAUSE FOR THE REPORT OF THE PATIENT EXPIRED AND DISCONTINUED USE OF THE HOMECHOICE WAS UNDETERMINED. THE DEVICE WAS EVALUATED BY THE BAXTER PRODUCT ANALYSIS LAB (PAL). THE UNIT WAS RECEIVED OPERATIVE AND IN POOR CONDITION WITH A DAMAGED DOOR COVER. THE DEVICE PASSED THE HOME CHOICE RETURN INSTRUMENT TEST/EVALUATION (RITE) ELECTRICAL TEST, BUT FAILED THE RITE FUNCTIONAL TEST DUE TO RELOAD THE SET (RLS) 151 ALARM. INSPECTION REVEALED DAMAGED VOLUMETRIC STANDARD LEFT TUBING. A REVIEW OF THE RETURNED DEVICE LOGS, WHICH CONTAINED DATA FROM THE MOST RECENT THERAPIES PERFORMED BY THE CUSTOMER, REVEALED NO PREVIOUS OCCURRENCES OF A RLS 151 ALARM. CONTINUED REVIEW OF THE DEVICES LOGS REVEALED NO FAILURE, MALFUNCTION OR INCREASED INTRAPERITONEAL VOLUME EVENTS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT PASSING AWAY. A SERVICE HISTORY WAS PERFORMED AND NO ISSUES WERE IDENTIFIED THAT MAY HAVE CONTRIBUTED TO THE ISSUE OF RLS 151. THE PREVIOUS RETURN OF THE DEVICE WAS NOT FOR ANY ISSUES RELATED TO RLS 151. THE DEVICE WAS SUBSEQUENTLY ROUTED TO SERVICE. THE ROOT CAUSE INVESTIGATION FOR RLS 151 IS IN PROGRESS THROUGH RENQ-CAPA-(B)(4). RENAL QUALITY ENGINEERING, ALONG WITH PLANT SERVICING AND QUALITY PERSONNEL, WILL CONTINUE TO MONITOR PRODUCT LINES FOR TRENDS AND WILL TAKE CORRECTIVE/PREVENTIVE ACTION AS APPROPRIATE.><(B)(4). THE DEVICE HAS BEEN RECEIVED FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION. | 0
2027646 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADDITIONAL DATA / FAILURE INVESTIGATION. FIELD SVC TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
2027647 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
2027656 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
2027731 | THE REPORTER, THE PATIENT'S FATHER, REPORTED THAT ON (B)(6) 2011, THE PATIENT OBTAINED BLOOD GLUCOSE READINGS OF 500 MG/DL TO 'HI MG/DL' (GREATER THAN 600 MG/DL) WHILE USING THE INSULIN PUMP. DURING THIS TIME PERIOD, THE PATIENT EXPERIENCED NO SYMPTOMS OF HIGH OR LOW BLOOD GLUCOSE LEVELS AND TESTED NEGATIVE FOR KETONES. THE PATIENT PROGRAMMED BOLUSES FOR THE ELEVATED BLOOD GLUCOSE LEVELS, BUT REPORTED HIS BLOOD GLUCOSE LEVEL DID NOT BECOME LOWER UNTIL A MANUAL INSULIN INJECTION WAS GIVEN. AT THE TIME OF THE TROUBLESHOOTING TELEPHONE CALL, THE PATIENT'S BLOOD GLUCOSE READING WAS 130 MG/DL. TROUBLESHOOTING REVEALED THERE WERE NO SIGNS OF KINKS OR BUBBLES IN THE CANNULA, NO PROBLEMS AT THE SITE, NO LEAKS, AND THE TOTAL BOLUS/BASAL UNITS DELIVERED MATCH THOSE PROGRAMMED INTO THE PUMP. THE CUSTOMER SERVICE REPRESENTATIVE ALSO DETERMINED DURING THIS TELEPHONE CALL THAT THE PATIENT WAS USING REFRIGERATED INSULIN INSTEAD OF ROOM TEMPERATURE, WHICH MAY CAUSE BUBBLE FORMATION. THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVALUATION. IF THE DEVICE IS RETURNED, AN EVALUATION SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME.><DEVICE EVALUATION: THE PUMP HAS BEEN RETURNED AND EVALUATED BY PRODUCT ANALYSIS ON (B)(6) 2011 WITH THE FOLLOWING FINDINGS: A REVIEW OF THE PUMP'S TOTAL DAILY DOSE HISTORY INDICATED THAT INSULIN DELIVERY TOTALS CORRECTLY REFLECTED PROGRAMMED VALUES. THE PUMP HISTORY INDICATED THAT MULTIPLE OCCLUSION ALARMS OCCURRED, WHICH COULD NOT BE DUPLICATED DURING TESTING. THE PUMP WAS EXERCISED FOR 29 HOURS WITH NO INSULIN DELIVERY ISSUES OCCURRING. AN EZPRIME OPERATION WAS PERFORMED, WITH NO ISSUES OCCURRING. A NORMAL BOLUS AND AN AUDIO BOLUS WERE PERFORMED AND WERE CORRECTLY RECORDED IN THE PUMP HISTORY. THE UNITS REMAINING WERE CORRECTLY CALCULATED DURING BOTH TYPES OF BOLUSES. UNRELATED TO THE COMPLAINT, EVALUATION REVEALED THAT THE BOLUS BUTTON COVER WAS MISSING. THE BOLUS BUTTON WAS FOUND TO BE RESPONDING APPROPRIATELY WHEN PRESSED. INSULIN CAN BE DELIVERED USING THE NORMAL BOLUS FEATURE. THIS DEFECT IS GENERALLY OBVIOUS AND DETECTABLE TO THE USER AND IS NOT LIKELY TO CAUSE OR CONTRIBUTE TO AN ADVERSE EVENT. | 0
2027921 | IN OUR RADIOLOGY DEPARTMENT, A SOFTWARE UPGRADE WAS PERFORMED ON THE PHILIPS COMPUTED RADIOGRAPHY (PCR) ELEVA FROM VERSION PCR 1.1.5 TO PCR 1.2. THIS UPGRADE WAS PERFORMED, BEGINNING IN LATE FALL OF 2010. ON THE NEW 1.2 VERSION, THERE IS A SMALL SYMBOL THAT TRANSFERS OVER TO THE PACS (PICTURE ARCHIVE AND COMMUNICATIONS SYSTEMS). OUR CONCERN WITH THIS FEATURE IS THE ICON CAN BE MISTAKEN FOR THE MARKER DENOTING THE PATIENT'S RIGHT (R) SIDE. ON CERTAIN RADIOGRAPHIC VIEWS, WE UTILIZE AP-PA MIRRORING IN CONJUNCTION WITH A HANGING PROTOCOL. IN THE MIRRORING PROCESS, THE ICON COULD BE PLACED ON THE PATIENT'S LEFT SIDE, BUT YOU SEE THE R MIRROR ICON. IN OUR PRACTICE, THE TECHNOLOGISTS PLACE THE APPROPRIATE LATERALITY "L" OR "R" MARKER ON THE CORRECT SIDE, BUT ON PACS THERE ALSO IS THE MIRROR ICON THAT COULD BE ON THE LEFT SIDE OF THE PATIENT. THIS IS OF CONCERN TO US AS RADIOLOGISTS, SURGEONS, CLINICIANS MAY NOT KNOW THIS IS AN AP-PA MIRROR ICON AND MAY MISTAKE THE ICON FOR THE WRONG SIDE OF THE PATIENT'S BODY. AS OF (B)(6), 2011 THE AFFECTED FEATURES HAVE BEEN TURNED OFF AND SOME SYSTEMS REVERTED TO PRIOR SOFTWARE.DETAILED INFORMATION HAS BEEN PROVIDED TO THE MANUFACTURER AND THEY HAVE INDICATED THEY ARE WORKING TO RESOLVE THE ISSUE. | 1
2028784 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED AN INITIAL POTASSIUM RESULT OF 6.6 MMOL/L WHICH WAS QUESTIONED BY THE PHYSICIAN. THE SAMPLE WAS REPEATED AND A RESULT OF 3.7 MMOL/L WAS GENERATED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). AN ABBOTT FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE ACCOUNT. THE FSE REPLACED A LEAKY 1 ML ICT ASPIRATION SYRINGE. THE FSE ALSO REBUILT THE REAGENT 1 SYRINGE WITH NEW TIPS AND O-RINGS. THE FSE CLEANED THE ICT TO VALVE TUBING. THE FSE REPLACED THE REAGENT 1 PROBE AND THE SHAFT TUBING. THE ICT ASPIRATION TUBING WAS ADJUSTED AND TIGHTENED. THE ISSUE WAS RESOLVED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY INTEGRATE CHIP TECHNOLOGY (ICT) REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
2031659 | THE INFUSOR BAG WAS NOTED TO BE LOSING PRESSURE CAUSING ARTERIAL LINE TO PARTIALLY FAIL. I WAS ABLE TO RETURN ARTERIAL LINE TO PROPER FUNCTION AFTER RE-INFLATING INFUSOR BAG. THE INFUSOR BAG WAS REMOVED AND REPLACED. I WAS INFORMED BY THE STAFF THAT THIS HAS OCCURRED TO SEVERAL PATIENTS RECENTLY. CENTRAL SUPPLY WAS NOTIFIED OF THE FAILURE OF THE PRODUCT. THERE WAS NO HARM TO THE PATIENT.===========REPORTER'S RESPONSE============THE REPORTER INDICATED THEY HAVE ABOUT 20 PATIENTS INVOLVED IN THIS INCIDENT. ANOTHER 10-20 PATIENTS THEY SUSPECT WERE INVOLVED ON PREVIOUS INCIDENTS. THE NUMBERS REFLECT NUMBER OF PATIENTS AT THEIR INSTITUTION AND NOT AT ANY OF THEIR DIVISIONS. THE REPORTER ALSO INDICATES THEY HAD SIMILAR EVENT MAYBE 5 YEARS AGO. "DURING THIS LATEST EPISODE WE HAD AT LEAST TWO PATIENTS LOSE THEIR ARTERIAL LINES BECAUSE THEY CLOTTED OFF SECONDARY TO THE INFUSION BAG DEFLATING AND THE SUBSEQUENT LOSS OF A CONTINUOUS INFUSION OF SALINE THROUGH THE ARTERIAL CATHETER. WE HAD SEVERAL OTHER ARTERIAL LINES LOSING FUNCTION (CLOTTING OF THE CATHETER SUSPECTED) DURING THIS TIME FRAME FOR "UNKNOWN" REASONS THAT MAY HAVE BEEN RELATED TO THE LOSS OF THE SALINE INFUSION BUT THERE WAS NOT DOCUMENTATION OF THE INFUSION BAGS BEING DEFLATED." THEY BELIEVE THE LOT NUMBER LISTED ON THIS REPORT IS THE ONLY LOT NUMBER INVOLVED IN THIS EVENT. "THE LEAKS APPEARED TO BE OCCURRING INTERNALLY IN THE STOP COCK. IF WE SEALED THE LUER LOCK CONNECTION AT THE STOP COCK WITH A DEAD END CAP THE LEAKS STOPPED. IF WE LEFT THE LUER LOCK CONNECTION OPEN AND HAD THE STOP COCK TURNED TO THE OFF POSITION (TOWARD BAG) WE LOST PRESSURE OVER A SPAN OF AN HOUR OR TWO." | 0
2033632 | THE PROBLEM WAS REPORTED BY THE USER SITE DURING CLINICAL USE: THE IMAGES TRANSFERRED TO THE PACS (PICTURE ARCHIVING AND COMMUNICATIONS SYSTEMS), AND ALL MR IMAGES/STUDY WERE STORED IN (B)(6) UNDER THE SAME PT NAME. UPON INVESTIGATION IT WAS DISCOVERED THAT MULTIPLE PT RECORDS WERE INCORRECTLY MERGED IN MDC PACS DUE TO A CONFIGURATION THAT DOES NOT MATCH TO THE SITE WORKFLOW. THE REPORTING SITE DOES NOT ISSUE UNIQUE PT IDENTIFIERS IN ITS WORKFLOW, ALTHOUGH THE SYSTEM WAS CONFIGURED TO USE PT ID AS A SINGLE ENTITY ATTRIBUTED FOR IDENTIFICATION OF PT ID. THERE WAS NO REPORT OF PT OR USER HARM ASSOCIATED WITH THIS ISSUE. THE MDC PACS IS A SOFTWARE APPLICATION THAT IS USED FOR RECEIVING, MANAGING, ARCHIVING, DISTRIBUTING AND RECORDING MEDICAL IMAGES ONTO PORTABLE DIGITAL MEDIA (INCLUDING BUT NOT LIMITED TO COMPACT DISK (CD) AND DVD). PRODUCT EVAL DETERMINED THE PRODUCT OPERATED AS DESIGNED. THE REPORTED MALFUNCTION HAS BEEN ASSESSED IN A HEALTH HAZARD EVAL, AND IT WAS DETERMINED THAT THE RISK IS ACCEPTABLE BASED ON ADDITIONAL ANALYSIS THAT SUPPORTS THE ISSUE OCCURRENCE AS A RESULT OF THE ACTIONS OF A SINGLE UNTRAINED INSTALLER, WHO DID NOT COMPLETELY REVIEW INSTALLATION DOCUMENTATION AND DID NOT CONFIGURE THE PRODUCT IN ACCORDANCE WITH THE SITE WORKFLOW. NO OTHER SITES WITH THIS WORKFLOW HAVE BEEN INSTALLED BY THIS UNTRAINED INSTALLER, AND NO OTHER SITES HAVE REPORTED THIS ISSUE. IF THE SYSTEM WAS MISCONFIGURED IN THIS MANNER AT A SITE THAT DOES NOT USE UNIQUE PT IDENTIFIERS, IT IS EXPECTED THAT THE ISSUE WOULD BE DETECTED QUICKLY BY THE USERS. | 1
2035364 | THURSDAY EVENING ABOUT 9:00 PM, THE DEVICE WAS DISCONNECTED FROM ITS BATTERY CHARGER AND PLACED AT THE FOOT END OF THE TABLE OF THE SIEMENS ESPREE MRI SCANNER IN PREPARATION FOR THE ARRIVAL OF AN ICU PATIENT. THAT PATIENT WAS KEPT IN THE HALLWAY OUTSIDE THE ROOM WHILE THEIR IV LINES WERE BEING SWITCHED OVER TO TWO MR COMPATIBLE PUMPS, ONE OF WHICH ENCOUNTERED AN UNEXPECTED AIR-IN-LINE CONDITION CAUSING A DELAY. DURING THAT DELAY, THE ECG MODULE STARTED TO SMOLDER AND QUICK-THINKING STAFF REMOVED IT FROM THE EXAM ROOM AND CALLED A CODE RED. THERE WAS NO HARM TO THE PATIENT OR STAFF. THE PRINCIPLE CONCERN, HOWEVER, IS THAT THE DEVICE HAS A SHORT LEAD SET SO THAT WHEN IT IS CONNECTED TO THE ELECTRODES ON A PATIENT'S CHEST, IT IS THEN LAID ON THE CHEST AND CONNECTED VIA A FIBEROPTIC CABLE TO THE ACTUAL DISPLAY MONITOR. A TECH GRABBED THE SMOLDERING DEVICE BY THE LEAD SET AND TOOK IT DOWN A DIFFERENT HALLWAY AND PLACED IT ON ANOTHER MR TABLE THAT IS KEPT THERE. A LITTLE LATER, HE MOVED IT TO THE FLOOR. THE SMOKE FROM THE DEVICE DID NOT SET OFF THE SMOKE DETECTORS. (I FOUND A REFERENCE TO THE COMPANY NEXERGY ON THE CHARGER.) THE EVENT WAS CLEARLY THE RESULT OF SOME KIND OF BATTERY FAILURE, BUT WE DON'T KNOW THE EXACT NATURE.MANUFACTURER RESPONSE, ACCORDING TO REPORTER, FOR PHYSIOLOGICAL MONITOR, MRI, INTERFACE, VERIS: THEY CHECKED THEIR RECORDS AND TELL US THIS HAS NOT HAPPENED ANYWHERE OR AT ANYTIME BEFORE. | 0
2035795 | MEASURED AND INTERPRETED THE MIDNIGHT STRIP ON THE PATIENT; NOTICED THE UPPER AND LOWER LEADS OF THE ECG STRIP WERE NOT IN SYNCH. THERE WAS A 0.08 SECOND DELAY ON THE LOWER LEAD. CHANGED TELEMETRY BOX, WITH THE SAME RESULTS AND SAME DELAY IN LEADS WERE SHOWN. THEY CHANGED THE DASH 4000 (BEDSIDE MONITOR). THAT FIXED THE PROBLEM. THE UPPER AND LOWER LEADS WERE THEN IN SYNCH. | 1
2036050 | TITLE VOLUME MONITOR ON ANESTHESIA MACHINE WAS LOCKED IN SAFETY POSITION DURING A HERNIA REPAIR CASE. MACHINE WOULD NOT SHOW TITLE O2 VOLUME. NO ADVERSE PATIENT OUTCOME; PATIENT WAS NOT AFFECTED IN ANY WAY. CLINICAL ENGINEERING (CE) WAS NOTIFIED AND THEY WERE UNABLE TO REPAIR. CE NOTIFIED BIOMED TECHNOLOGIES WHICH IS A COMPANY WE CONTRACT WITH TO FIX ANESTHESIA MACHINES. BIOMED TECHNOLOGIES REPLACED THE PART. | 1
2036419 | THE CUSTOMER STATED RBC AND MCV QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 8/10/2010, EXPIRES 6/18/2012. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. (B)(4). | 0
2036436 | THE CUSTOMER STATED RBC, PLATELET, HEMATOCRIT, MCH AND MCHC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH A DIFFERENT LOT AND PERFORMED A DECONTAMINATION TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CELL-DYN EMERALD CLEANER MANUFACTURED ON 08/10/2010, EXPIRES ON 06/18/2012. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. (B)(4). | 0
2036483 | IT WAS REPORTED THAT THE KEYPAD BUTTONS HAD TO BE PRESSED MULTIPLE TIMES IN ORDER TO ACTIVATE THE DESIRED PUMP FUNCTIONS. THE KEYPAD IS REPORTEDLY INTACT AND THE PUMP HAS NOT BEEN EXPOSED TO MOISTURE. THERE WAS NO ADVERSE EVENT ASSOCIATED WITH THIS COMPLAINT. THE PUMP WAS RETURNED TO ANIMAS FOR EVALUATION AND EVALUATED. THE KEYPAD APPEARED TO BE INTACT; HOWEVER, THE UP/DOWN/OK/CONTRAST KEYPAD BUTTONS WERE INTERMITTENTLY RESPONDING TO USER INPUTS AND THERE WAS EVIDENCE OF CONTAMINATION UNDER THE KEY CONTACTS. | 0
2036955 | THE PATIENT UNDERWENT A PROCEDURE IN THE CATH LAB. DURING THE PROCEDURE, THE PINK BORDER REFERENCING A "DISCONNECT MANAGER" PROBLEM SHOWED UP. DURING THIS SITUATION, THE CASE IS SUPPOSED TO STORE ON THE LOCAL DATABASE UNTIL THE SYSTEM RECONNECTS AND THEN INFORMATION WILL TRANSFER. WHEN THE MONITOR TECH WENT TO PREVIEW THE CASE, THE MAJORITY OF THE DOCUMENTATION WAS MISSING. MCKESSON IS UNABLE TO RESTORE THE REPORT. THIS DOES NOT AFFECT THE IMAGES OR THE PHYSICIAN'S DICTATION. IT STRICTLY AFFECTS THE PROCEDURE REPORT. THE ANCILLARY APPLICATION TEAM WILL SUPPLY THE VENDOR'S INFORMATION. NO HARM TO THE PATIENT. | 1
2037606 | THE LANTIS ROIS FAILED ON SEVERAL OCCASIONS TO ELECTRONICALLY RECORD ALL OR PART OF PT TREATMENT. THIS IS AN INTERMITTENT ISSUE OF UNK CASE. THE SYSTEM OPERATES 6 LINEAR ACCELERATORS AT 3 CLINICAL SITES AND THE PROBLEM HAS OCCURRED ON SEVERAL OF THE ACCELERATORS. LANTIS SENDS AN ALERT MESSAGE TO THE OPERATOR WHEN THIS OCCURS SO THE OPERATOR KNOWS TO MANUALLY RECORD THE TREATMENT. THIS ULTIMATELY HAS THE POTENTIAL TO HUMAN ERROR OR PT OVERDOSE. THE (B)(6) CONTINUES TO WORK WITH SIEMENS TO RESOLVE THIS ISSUE. SIEMENS' CURRENT PROPOSED SOLUTION IS TO REBUILD THE ENTIRE DATABASE. EVENT DATE: (B)(6) 2009 - (B)(6) 2011. | 1
2038514 | THE REPORTER, THE PATIENT'S MOTHER, REPORTED THAT ON (B)(6) 2011, THE PATIENT OBTAINED A BLOOD GLUCOSE READING OF 450 MG/DL, AND A POSITIVE KETONES RESULT OF 3.7 MMOL/L. THE PATIENT DID NOT REPORT EXPERIENCING ANY SYMPTOMS OF HIGH OR LOW BLOOD GLUCOSE LEVELS. THE PATIENT'S MOTHER CHANGED THE PATIENT'S INFUSION SITE, CARTRIDGE AND TUBING, AND GAVE THE PATIENT A CORRECTING BOLUS USING THE PUMP. THE PATIENT'S SUBSEQUENT BLOOD GLUCOSE READINGS RANGED FROM 100 MG/DL TO 140 MG/DL. TROUBLESHOOTING REVEALED THERE WERE NO LEAKS FOUND FROM THE CARTRIDGE OR IN THE CARTRIDGE COMPARTMENT. THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVALUATION. IF THE DEVICE IS RETURNED, AN EVALUATION SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. | 0
2039039 | IT WAS REPORTED THAT THE UP/DOWN/OK KEYPAD BUTTONS HAD TO BE PRESSED MULTIPLE TIMES IN ORDER TO ACTIVATE THE DESIRED PUMP FUNCTIONS. THE KEYPAD IS REPORTEDLY INTACT AND THE PUMP HAS NOT BEEN EXPOSED TO MOISTURE. THERE WAS NO ADVERSE EVENT ASSOCIATED WITH THIS COMPLAINT. THE PUMP HAS BEEN RETURNED TO ANIMAS; HOWEVER, THE INVESTIGATION HAS NOT BEEN COMPLETED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. ONCE THE EVAL HAS BEEN COMPLETED A SUPPLEMENTAL REPORT WILL BE FILED. | 0
2040378 | PROBE PASSED TEST BUT STARTED FROSTING UP THE SHAFT DURING FREEZE. THE DOCTOR TOOK THE NECESSARY PRECAUTIONS TO PREVENT THE SKIN FROM BEING DAMAGED. PT WAS NOT INJURED DURING THE PROCEDURE DUE TO THIS ISSUE. PER INITIAL REPORT, PT WAS NOT INJURED. | 0
2042167 | THE REPORTER REPORTED THAT ON (B)(6) 2011 THE PT OBTAINED THE ELEVATED BLOOD GLUCOSE READING OF 360 MG/DL AND SHE TESTED POSITIVE FOR KETONES. AT THAT TIME, THE PT EXPERIENCED THE SYMPTOMS OF NAUSEA AND BLURRED VISION. AT 11:11 PM THE PT CHANGED THE CARTRIDGE AND OBSERVED INSULIN LEAKING FROM THE TUBING CONNECTION TO THE CARTRIDGE, AND THE CARTRIDGE COMPARTMENT SMELLED OF INSULIN. AT 12:00AM THE PT CHANGED THE CARTRIDGE AGAIN, AND HER BLOOD GLUCOSE LEVEL RESPONDED TO THE BOLUSES DELIVERED. TROUBLESHOOTING REVEALED THERE WAS NO PROBLEM WITH THE INFUSION SITE, THERE WAS NO BLOOD IN THE CANNULA, THE PUMP'S PROGRAMMING WAS CORRECT, AND THE TOTAL DAILY DOSES OF INSULIN DELIVERED MATCHED THE BASAL/BOLUSES PROGRAMMED. THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVALUATION. IF THE DEVICE IS RETURNED, AN EVALUATION SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. | 0
2042590 | PLACEMENT OF EPIDURAL IN PROGRESS ON PATIENT IN ACTIVE LABOR. UNABLE TO OBTAIN ADEQUATE QUALITY ELECTRONIC FETAL MONITOR TRACING VIA EXTERNAL MODE. PHANTOM SIGNAL OVERRIDING TRUE EXTERNAL FETAL HEART RATE SIGNAL. MULTIPLE CHANGES OF EXTERNAL ULTRASOUND TRANSDUCER WERE PERFORMED WITH NO IMPROVEMENT OF SIGNAL. APPLIED INTERNAL FSE. THREE DAYS AFTER THE EVENT, A CLINICAL ENGINEERING TECHNICIAN PREFORMED THE FOLLOWING TESTS ON THE MONITOR (IN THE ROOM WHERE THE EVENT OCCURRED), MECG, FECG, ULTRASOUND, FETAL MOVEMENT DETECTION AND UTERINE ACTIVITY, AS LISTED IN THE SERVICE MANUAL PAGES; PAGE 4-24 TO PAGE 4-36. ALL TESTS PASSED AND A COPY OF THE STRIP CHART RECORDING WAS LEFT ON THE FETAL MONITOR. HOWEVER, THERE WAS A NOTICEABLE INTERFERENCE NOISE OCCURRING EVERY 8 SECONDS AND LASTING ABOUT 1/2 A SECOND. THREE FETAL MONITORS WERE TESTED WITH THE SAME CONSISTENT PROBLEM OBSERVED. THE ELECTRICITY OUTLETS WERE CHECKED WITH A POLARITY TESTER, AND THEY WERE FOUND TO BE OK. THE NEXT DAY, UNIT STAFF INDICATED THAT NO INTERFERENCE WAS OBSERVED, HOWEVER THE PROBLEM HAS OCCURRED IN OTHER PATIENT ROOMS. AS THIS SEEMS TO BE AN INTERMITTENT ISSUE, CLINICAL ENGINEERING WILL BE PAGED WHEN THE PROBLEM IS OBSERVED AGAIN SO A LIVE ASSESSMENT CAN BE MADE. THE MANUFACTURER HAS BEEN INFORMED OF THE ISSUE. | 1
2042967 | A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER REPORTED AN INFANT APNEA MONITOR DID NOT ALARM DURING AN APNEIC EVENT AND THE PATIENT EXPIRED AS A RESULT. THE DME STATED THAT THE MOTHER OF THE INFANT ADMINISTERED CARDIOPULMONARY RESUSCITATION (CPR) AND THE EQUIPMENT DID NOT ALARM. THE REPORTED EVENT OCCURRED ON (B)(6), 2011 AT 10:15 AM. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION AND DID NOT CONFIRM THE CUSTOMER'S COMPLAINT OF THE UNIT NOT ALARMING. THE INFANT APNEA MONITOR WAS VISUALLY EXAMINED AND TESTED BY THE MANUFACTURER USING A SIMULATOR IN ACCORDANCE WITH THE SMARTMONITOR 2 CHECKOUT PROCEDURE MANUAL (PN 1020817). THE INFANT APNEA MONITOR DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS AND PASSED ALL REQUIRED TESTING. THE SMARTMONITOR 2 WAS SET UP WITH A 20 SECOND DELAY BEFORE ANNUNCIATING AND RECORDING AN ALARM FOR AN APNEA CONDITION. THE SMARTMONITOR 2 WAS PROGRAMMED WITH PARAMETERS THAT WERE APPROPRIATE FOR RECORDING AND ALARMING FOR RESPIRATION AND HEART RATE EVENTS DURING INFANT MONITORING. FURTHERMORE, THE INFANT APNEA MONITOR'S MEMORY DATA WAS DOWNLOADED AND ANALYZED BY TRAINED ASSOCIATES. THE DOWNLOADED MEMORY REVEALED THE APNEA MONITOR WAS NOT IN USE DURING THE TIME OF THE REPORTED EVENT. THE DEVICE WAS TURNED OFF ON (B)(6), 2011 AT 3:59:05 AM AND NOT TURNED BACK ON UNTIL (B)(6), 2011 AT 1:30:52 PM. THE SMARTMONITOR 2 DEVICE IS NOT INTENDED TO PREVENT LOSS OF BREATHING OR HEART ACTIVITY. THE SMARTMONITOR 2 PARENTS' GUIDE (PN 572-4000-00) FURTHER STATES: "THE SMARTMONITOR 2 IS A MONITORING DEVICE ONLY. IT DOES NOT PREVENT THE LOSS OF BREATHING OR HEART ACTIVITY, NOR WILL IT RESTORE BREATHING OR HEART ACTIVITY. IT WILL NOT PREVENT DEATH. ANYONE USING THE SMARTMONITOR 2 TO MONITOR AN INFANT SHOULD BE TRAINED IN CURRENT INFANT CARDIOPULMONARY RESUSCITATION (CPR), WHICH IS THE PROPER WAY TO RESTORE BREATHING AND HEART ACTIVITY." ALTHOUGH THERE WAS A DEATH REPORTED, THE MANUFACTURER CONCLUDED THAT THE DEVICE DID NOT CAUSE NOR CONTRIBUTE TO THE INFANT DEATH. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE ALLEGATION OF ALARM FAILURE DURING A PATIENT EVENT AND DETERMINED THAT THE DEVICE WAS NOT IN USE DURING THE PATIENT EVENT. THE MONITOR PASSED ALL REQUIRED TESTING AND DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS. BASED ON ALL AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES THAT THE DEVICE DOES FUNCTION TO SPECIFICATION AND THAT NO FURTHER ACTION IS APPROPRIATE. | 1
2044358 | A 1.4 GHZ TELEMETRY TRANSMITTER CONNECTS TO A PHILIPS DEDICATED TELEMON (BEDSIDE/TRANSPORT MONITOR) DURING PATIENT TRANSPORT TO DISPLAY THE VITAL SIGNS FOR THE CAREGIVER. WHEN TRANSPORT OCCURS OR WHEN VIEWING THE PATIENT'S VITAL SIGNS IN THE PATIENT ROOM THERE MUST BE TWO CABLES CONNECTED TO EACH OTHER BETWEEN THE MONITOR AND TRANSMITTER. ONE OF THE CONNECTORS ON THE MONITOR TETHER CABLE # 9898-031-4611 AND ONE OF THE CONNECTORS ON THE SECOND TRANSMITTER TETHER CABLE # 9898-031-43481 ARE FREQUENTLY BREAKING. THE HOSPITAL WAS NOTIFIED IN THE FALL OF 2010 THAT PHILIPS WAS GOING TO MAKE A CHANGE TO THESE CABLES. THE CONNECTIONS ON THESE CABLES ARE BREAKING SOONER THAN THEY SHOULD WITH REGULAR USE.======================MANUFACTURER RESPONSE FOR PHYSIOLOGICAL MONITOR, TELEMETRY, TETHER CABLES, INTELLIVUE:======================FACILITY HAS WORKED WITH THE MANUFACTURER TO ATTACH A STRAIN RELIEF ON EACH MONITOR TO KEEP THE CONNECTOR FROM GETTING BROKEN AT THE CONNECTION TO THE MONITOR. A COMPLAINT WAS FILED WITH THE PHILIPS SALES REPRESENTATIVE ABOUT THE LONGEVITY OF THE CABLES. | 1
2045573 | PT USES A ONE TOUCH GLUCOSE METER TO CHECK AND RECORD HIS BLOOD SUGAR READINGS. WHEN THE DEVICE IS CONNECTED TO THE COMPUTER TO DOWNLOAD THE CONTENTS, USING THE ONE TOUCH SOFTWARE, IF THERE IS A DISCREPANCY BETWEEN THE COMPUTER AND DEVICES TIME AND DATE THE ONE TOUCH SOFTWARE USUALLY NOTIFIES US TO CORRECT THE DIFFERENCE. IN THIS SITUATION, THE SOFTWARE DID NOT INDICATE TO US THAT THE TIME AND DATE WERE INCORRECT IN THE PT'S METER. IT CREATED THE REPORTS AS USUAL, HOWEVER, THE DATE ON THE REPORTS REFLECTED THE YEAR 2015 INSTEAD OF 2011. INITIALLY, THIS ERROR WAS NOT NOTICED AND THE DOCTOR CHANGED THE PT'S MEDICATION AND TREATMENT BASED ON HIS LATEST READINGS FROM HIS GLUCOSE METER. HOWEVER, WHEN THE PT LEFT AND THE DOCTOR WAS REVIEWING HIS NOTES, HE IDENTIFIED THE TIME AND DATE ERROR ON THE ONE TOUCH SOFTWARE REPORT. WE IMMEDIATELY CALL THE PT AND HAD HIM RETURN TO THE OFFICE TO INSPECT HIS METER AND ACTUAL READINGS. WE DETERMINED THAT HIS METER IS OPERATING PROPERLY, HOWEVER, THERE IS SOME SORT OF GLITCH IN THE ONE TOUCH COMPUTER SOFTWARE. WE REPORTED THIS INCIDENT TO OUR ONE TOUCH REPRESENTATIVE AND SHE INSTRUCTED US THAT WHEN WE SEE THIS DATE ERROR, WE NEED TO DELETE THAT PARTICULAR PT'S ARCHIVED INFORMATION IN THE ONE TOUCH COMPUTER SYSTEM TO PURGE THE ERROR FROM THE PROGRAM. THAT STILL DOES NOT CORRECT THE ISSUE THAT THE SOFTWARE IS NOT ACTING PROPERLY AND NOTIFYING US THAT THERE IS A DISCREPANCY IN THE METER AND COMPUTER TIME AND DATE, WHICH IN THIS PT'S SITUATION COULD HAVE BEEN LETHAL WITH THE PRESCRIPTION CHANGES THAT WERE MADE BASED ON THE INCORRECT DATA. | 1
2045592 | DURING ROUTINE TESTING OF THE DEVICE AT THE SERVICE CENTER, THE QUALITY TECHNICIAN REPORTED THAT THE UPPER HOUSING BOSS WAS BROKEN. SINCE THE EVENT OCCURRED DURING ROUTINE TESTING OF THE DEVICE, THERE WAS NO PATIENT INVOLVEMENT DURING THIS EVENT. EVALUATION IN PROGRESS, BUT NOT YET CONCLUDED. | 0
2045867 | I AM WRITING TO REPORT A PROBLEM IN THE DESIGN OF THE CERNER POWERCHART COMPUTER PROVIDER ORDER ENTRY (CPOE) PRODUCT. AS I WILL DESCRIBE BELOW, THESE DESIGN FLAWS ARE LARGELY RESPONSIBLE FOR APPROXIMATELY 100 MEDICATION ERRORS PER DAY IN OUR 240 BED HOSP. THE (B)(6) HOSP IN (B)(6) IS A 240 BED TEACHING HOSP. WE ARE OWNED BY THE (B)(6). WE HAVE TO BE ONE OF THE ONLY HOSPITALS IN THE COUNTRY TO FULLY IMPLEMENT CPOE WITH TWO ENTIRELY DIFFERENT ELECTRONIC MEDICAL RECORDS (EMR). IN 2007, WE COMPLETED A COMPREHENSIVE CPOE IMPLEMENTATION USING IDX 3.1 (WHICH WAS LATER PURCHASED BY GENERAL ELECTRIC). OUR SELF-REPORTED MEDICATION ERROR RATE WAS <2 ERROR REPORTS PER DAY. IN (B)(6) OF 2010, WE TRANSITIONED OUR EMR TO CERNER MILLENNIUM - POWERCHART (B)(4). THIS IMPLEMENTATION WAS DONE AS A FIRST STEP TO CONVERGE ONTO A SINGLE ELECTRONIC MEDICAL RECORD ACROSS THE (B)(6) ENTERPRISE. DESPITE EXTENSIVE NURSING INFORMATICS SUPPORT (APPROX. 7000 HRS PER MONTH FOR THE FIRST SIX MONTHS AFTER IMPLEMENTATION), OUR ERROR RATE WENT FROM APPROX. 176 MEDICATION ERRORS PER DAY ONE MONTH AFTER IMPLEMENTATION TO APPROX. 100 MEDICATION ERRORS PER DAY CURRENTLY. IT HAS BEEN STABLE AT THIS LEVEL FOR AT LEAST THE PAST 3 MONTHS. MANY OF THESE ERRORS INVOLVE HIGH RISK MEDICATIONS (E.G. HEPARIN, MORPHINE). IN ORDER TO UNDERSTAND THE ETIOLOGY OF THE ERRORS, I NEED TO EXPLAIN HOW CERNER PROCESSES CPOE ORDERS. WITH CERNER, ORDERS ARE GROUPED INTO PLANS ("POWERPLANS"). THESE PLANS CAN INCLUDE SUB-PLANS ("PHASES") AND SUB-SUB-PLANS ("SUB PHASES"). A TYPICAL ORDERS DISPLAY SCREEN APPEARS AS FOLLOWS: (B)(4). UNDER THE PLANS IS AN ORDER TAB WHICH DISPLAYS OTHER ORDERS. THIS DESIGN ALLOWS TWO DISTINCT SOURCES OF ERROR. FIRST, THERE IS NO CONSISTENT WAY TO VIEW ORDERS FOR MEDICATIONS. A MEDICATION ORDER CAN DISPLAY EITHER IN A PLAN, SUB-PLAN, OR UNDER THE ORDERS TAB. IN ORDER TO FIND A GIVEN MEDICATION, CERNER MANDATES THAT YOU CLICK THROUGH EACH AND EVERY PLAN. THIS FACILITATES DUPLICATION OF BOTH MEDICATION AND NON-MEDICATION ORDERS (THERE IS A MEDICATION CHECKER WHICH IS SO POORLY DESIGNED AND DOES LITTLE TO AID THE PHARMACISTS IN DETECTING DUPLICATE MEDICATIONS). SECOND AND MORE DANGEROUS, HIGH RISK INTRAVENOUS MEDICATIONS CAN BE RUN EITHER INSIDE OR OUTSIDE OF THE PLAN. CERNER PROGRAMMING DOES NOT KEEP HIGH RISK MEDICATIONS IN A CONSISTENT SPOT. IT IS VERY EASY TO STOP MONITORING FOR HEPARIN (WHICH IS INCLUDED IN THE HEPARIN PLAN) WHILE CONTINUING THE HEPARIN (WHICH IS OUTSIDE THE PLAN) OR VICE-VERSA. AS YOU CAN PROBABLY GUESS, WE FORMED MULTIPLE INTERDISCIPLINARY TEAMS TO ADDRESS THESE DEFICIENCIES. I WAS THE LEAD PHYSICIAN ON THE ORDERS TEAM. ALTHOUGH WE WERE UNABLE TO REDUCE THE ERROR RATE, IT IS NOT CLOSE TO AN ACCEPTABLE LEVEL. AGAIN, I BELIEVE THIS IS BECAUSE OF THE DESIGN OF THE CERNER PRODUCT WHICH IS WHY I CHOSE TO BRING THIS TO YOUR ATTENTION. WITH ISSUES THIS COMPLEX, I AM A FIRM BELIEVER THAT "SEEING IS BELIEVING." IF DESIRED, I WOULD BE HAPPY TO ARRANGE A WEB CONFERENCE TO DEMONSTRATE MY CONCERNS. | 1
2048186 | IT WAS REPORTED THAT THE CUSTOMER WAS HOSPITALIZED WITH A LOW BLOOD GLUCOSE LEVEL OF 2.8 MMOL/L. THE CUSTOMER STATED THAT SHE DID NOT BELIEVE THE ISSUE WAS WITH INSULIN PUMP AS SHE HAD BEEN SICK WITH INFLUENZA AND HAD NOT EATEN THE DAY OF THE EVENT. NOTHING FURTHER WAS REPORTED. CURRENTLY IT IS UNK WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE, THEREFORE, CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. MEDWATCH 1 OF 2 (INSULIN PUMP): 3004209178-2011-80924. | 0
2050123 | IT WAS REPORTED THAT THE LOWER LIMIT OF HEART FREQUENCY WAS NOT RECOGNIZED. THE CLINICAL BIOMED DEPARTMENT CHECKED THE MOTOR USING TWO DIFFERENT SIMULATORS AND NO FAILURE WAS IDENTIFIED. THERE WAS NO PT INJURY REPORTED. (B)(4). CORRECTION: THIS REPORT WAS INADVERTENTLY SUBMITTED UNDER 1220063-2011-2011. DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. CORRECTION: THIS REPORT WAS INITIALLY, INADVERTENTLY SUBMITTED UNDER MFR REPORT # 1220063-2011-2011. | 1
2053052 | RITONAVIR WAS APPROVED WITH THE DOSING OF 600 MG TWICE A DAY. NOWADAYS THAT IS A DOSING NOT USED BUT BECAUSE IT IS THE APPROVED DOSE THE ELECTRONIC MEDICAL RECORD THAT WE USE IN OUR INSTITUTION -EPIC- DEFAULTS THIS AS THE DEFAULT DOSING. WE HAD MORE THAN ONCE INSTANCE WHILE REFILLING NORVIR, IT WAS DEFAULTED TO THE UNACCEPTABLE DOSE FOR TODAY'S STANDARDS. WE WERE TOLD BY EPIC WHO OUTSOURCES THE FORMULARY TO AN OUTSIDE VENDOR THAT UNLESS THAT DOSE IS REMOVED BY THE FDA, IT WILL CONTINUE TO APPEAR AS SUCH. I AM WRITING TO SUGGEST THAT THE FDA MODIFIES THE APPROVED DOSING OF RITONAVIR TO THE CURRENT LOWER DOSES USED FOR BOOSTING. | 1
2053955 | IT WAS REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED FOR FOUR MINUTES FOR UP TO 16 PTS WHEN AN APEXPRO TELEMETRY SYSTEM REBOOTED ON ITS OWN DUE TO A PRESUMED HARDWARE FAILURE. THERE WERE NO PT INJURIES REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
2055317 | TELEMETRY BOX WAS ON A PATIENT WHEN THE MONITORING ROOM PERSON REPORTED A 38 SECOND PAUSE (LOSS OF SIGNAL). PATIENT WAS ASYMPTOMATIC IN THE BATHROOM. EKG NORMAL. LATER, AFTER BOX WAS TAKEN OFF OF THE PATIENT, IT READ THE SAME ASYSTOLE PATTERN WHEN IT SHOULD HAVE READ NO SIGNAL. SENT BOX FOR DIAGNOSTICS TO BIOMED.BIOMED TESTED THE DEVICE AND NO PROBLEM WAS FOUND. ALL REQUIRED TESTS WERE WITHIN NORMAL OPERATING PARAMETERS OPERATING TO THE MANUFACTURER SPECIFICATIONS AND RETURNED UNIT TO SERVICE. | 1
2055619 | THE PATIENT WAS BROUGHT TO THE CATH LAB FOR AN ATRIAL FIBRILLATION ABLATION PROCEDURE. THE PATIENT WAS INTUBATED AND SEDATED BY THE ANESTHESIOLOGIST. THE DOCTOR PLACED THE FEMORAL AND SUPERIOR VENA CAVA (SVC) CATHETERS. THE CARTO MAPPING SYSTEM WAS TURNED ON AND CALIBRATED AT THE BEGINNING OF THE CASE. THE DOCTOR WAS PREPARING TO PLACE THE TRANSSEPTAL CATHETER. INITIAL MAPPING TO BEGIN, BUT EQUIPMENT FAILED TO WORK. CLINICAL COMPANY REP PRESENT FROM BIOSENSE WEBSTER (BW) TROUBLESHOT EQUIPMENT FOR MORE THAN 1 HOUR AND PLACED A PHONE CALL TO TECHNICAL ENGINEERS AT BW. THE ATTEMPTS TO REMEDY PROBLEM WERE UNSUCCESSFUL. FINALLY THE DOCTOR ABORTED THE CASE. BIOSENSE WEBSTER SENT THE TECHNICAL ENGINEER TO TROUBLE SHOOT AND REPAIR EQUIPMENT THE NEXT DAY. A CABLE, CIRCUIT BOARD AND MAGNETIC BOARD UNDER THE TABLE WERE SENT TO BIOSENSE WEBSTER AND ALL WERE REPLACED AND MAILED TO THE EP LAB TO REPAIR THE CARTO. THE NEWLY MAILED CABLE WAS MALFUNCTIONING AND A SECOND CABLE WAS ORDERED TO BE INSTALLED ON THE FOLLOWING DAY. THE PATIENT WILL HAVE TO RETURN FOR THE PROCEDURE.====================== MANUFACTURER RESPONSE FOR CARTO, (BRAND NOT PROVIDED)======================THE COMPANY HAS BEEN VERY RESPONSIVE AND IS TAKING FULL RESPONSIBILITY. THEY DISCOVERED THE WRONG COMPUTER CARD WAS SENT AND ARE SENDING ANOTHER CARD. | 1
2057194 | IT WAS REPORTED THAT ALARM PAGES FROM THE PT NET CENTRAL STATION WERE NOT ARRIVING AT THE NURSE PAGERS. THERE WAS NO INJURY OR DEATH OR LOSS OF MONITORING ASSOCIATED WITH THIS EVENT. A MISSED ALARM PAGE COULD RESULT IN A DELAY OF TREATMENT FOR A CRITICAL PHYSIOLOGICAL EVENT. IT IS POSSIBLE THAT THE CAREGIVER MAY RELY ON THE PAGER AS A PRIMARY SOURCE OF ALARM NOTIFICATION. THE EVENT MAY BE A DEVICE MALFUNCTION. RECURRENCE OF THIS POSSIBLE MALFUNCTION COULD BE LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2057601 | THE CUSTOMER REPORTS THAT THE IMAGE ORIENTATION IS NOT DISPLAYING AS EXPECTED. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(6). | 1
2057711 | AN ERROR OCCURRED IN THE (B)(4) SOFTWARE AFTER A PT BECAME HYPOGLYCEMIC. THE SOFTWARE ERROR CAUSED AN ERROR NOTIFICATION TO BE DISPLAYED TO THE USER, AND PREVENTED THE USER FROM CONTINUING TO MANAGE THE PT ON THE SYSTEM. THE USER, FOLLOWING HER TRAINING, CONTINUED MANAGING THE PT MANUALLY (OFF THE SYSTEM). AT NO TIME DID THE SYSTEM ISSUE AN INCORRECT RECOMMENDATION. THE ERROR CONDITION WAS CORRECTED BY A COMPANY TECHNICIAN AND THE NURSE WAS THEN ABLE TO CONTINUE MANAGING THE PT ON THE SYSTEM. A CAPA WAS OPENED TO FULLY ANALYZE AND CORRECT THE CONDITION THAT LED TO THE ERROR. ALTHOUGH THE FAULT WAS CLEARLY EVIDENT TO THE USER, THIS ERROR OCCURRED WHEN THE PT BECAME HYPOGLYCEMIC AND THE SOFTWARE WAS CLASSIFIED AS A MAJOR LEVEL OF CONCERN ON THE 510(K). FOR THIS REASON, IT WAS DECIDED THAT AN MDR SHOULD BE FILED WITHIN 30 DAYS. AT APPROX 12:45 PM ON (B)(6) 2011, (B)(6), REC'D A NOTIFICATION ON HIS CELL PHONE OF A GLUCOCARE APPLICATION ERROR THAT HAD OCCURRED AT 12:17 PM AT (B)(6). PRELIMINARY INVESTIGATION OF THE ERROR INFORMATION SENT IN THE NOTIFICATION REVEALED THAT THE ERROR WAS AN ASSERTION ERROR IN THE LOGIC OF THE (B)(6) PROTOCOL; THE OCCURRENCE OF SUCH AN ERROR CAUSES THE DISPLAY OF AN ERROR SCREEN TO THE NURSE USING THE SYSTEM. (B)(6) THEN CONNECTED TO THE (B)(6) SYSTEM TO FURTHER INVESTIGATE THE ERROR CONDITION AND DETERMINE THE PT INVOLVED PRIOR TO CALLING THE UNIT. IT WAS NOTICED THAT A PT IN THE (B)(6) HAD HAD A GLUCOSE READING OF 69 MG/DL ENTERED AT 12:17 PM AND WAS ALERTING THE NURSE TO STOP THE INSULIN DRIP. (B)(6) PHONE THE UNIT AND ASKED FOR THE NURSE WHO WAS CARING FOR THE PT (B)(6). (B)(6) CONFIRMED SHE HAD GOTTEN AN ERROR, AND THAT SHE HAD STOPPED THE INSULIN DRIP BUT WAS UNABLE TO CONFIRM THE INSULIN DRIP CHANGE IN THE SYSTEM DUE TO THE ERROR. (B)(6) CONFIRMED (B)(6) WAS MANAGING THE PT MANUALLY AND ASKED IF SHE HAD TAKEN A SUBSEQUENT GLUCOSE READING; (B)(6) CONFIRMED SHE HAD OBTAINED A READING OF 96 MG/DL AT 1 PM. (B)(6) INFORMED (B)(6) THAT THE PROTOCOL SHOULD BE STOPPED AND THEN RESTARTED IN THE SYSTEM USING THE MID-PROTOCOL START FUNCTION WITH A CURRENT AND PRIOR GLUCOSE LEVEL. (B)(6) ASKED (B)(6) TO OBTAIN A CURRENT GLUCOSE READING. (B)(6) DID SO AND REPORTED THAT THE BLOOD GLUCOSE WAS NOT 115 MG/DL. (B)(6) PERFORMED A MID-PROTOCOL START AT APPROX 1:30 PM AND CONFIRMED THAT THE SYSTEM APPEARED TO BE WORKING PROPERLY. FURTHER INVESTIGATION OF THE ROOT CAUSE OF THE ERROR OVER THE NEXT FEW HOURS REVEALED A BUG RELATED TO STOPPING THE PROTOCOL AFTER A SPECIFIC SERIES OF EVENTS, WHICH WOULD BE EVIDENT IN THE PT'S HISTORY. (B)(6) CHECKED THE HISTORY OF THE PT AND CONFIRMED THAT THIS HAD INDEED OCCURRED. (B)(6) ALSO DETERMINED THAT THE DATA CORRUPTION THAT HAD CAUSED THE ERROR WAS STILL PRESENT AND WOULD CAUSE THE ERROR TO RECUR IF THE PT BECAME HYPOGLYCEMIC. IT WAS ALSO DETERMINED THAT THE SHORT-TERM FIX TO ENSURING THAT THE ERROR DID NOT RECUR WITH THIS PT WAS TO DELETE THE CORRUPT DATA IN THE DATABASE AND RESTART THE SYSTEM. A TEST WAS SUCCESSFULLY PERFORMED TO VALIDATE THE FIX, AFTER WHICH THE FIX WAS PERFORMED ON THE PRODUCTION SYSTEM AT (B)(6) WITHOUT INCIDENT. BOTH PRODUCTION SYSTEMS (B)(6) ARE BEING REGULARLY CHECKED FOR ANY SIGNS OF DATA CORRUPTION UNTIL A DEFINITIVE FIX IS INSTALLED. A SYSTEM UPDATE IS PLANNED WITHIN 30 DAYS OF THE OCCURRENCE OF THIS ERROR. | 1
2063954 | IT WAS REPORTED THAT A PATIENT EXPERIENCED A BURNING SENSATION ON THE LIPS, TONGUE, AND THROAT FOLLOWING A TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) PROCEDURE. TWO WEEKS FOLLOWING THE EXAM THE PATIENT STILL REPORTED A SWOLLEN TONGUE AND WAS REFERRED TO A SPECIALIST FOR INDEPENDENT ADVICE. PRELIMINARY EVALUATION INDICATES THAT THE TRANSDUCER PROBE MAY NOT HAVE BEEN CLEANED, DISINFECTED, OR RINSED PROPERLY RESULTING IN RESIDUAL CHEMICAL ON THE PROBE. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
2066170 | IT WAS REPORTED THAT IMAGES WERE NOT DISPLAYING PROPERLY. THE FIELD ENGINEER ARRIVED AND FOUND THE GRADIENT CABLES SWAPPED. THIS WAS THE ROOT CAUSE OF THE FLIPPED IMAGES. HE RETURNED THE CABLES TO THEIR CORRECT LOCATIONS AND THE IMAGES WERE FINE. THERE WAS NO MISDIAGNOSIS OR MISTREATMENT AS A RESULT OF THIS INCIDENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2066580 | IN JANUARY A BIOMED TECHNICIAN WAS CALLED TO THE EMERGENCY ROOM TO ADDRESS A POSSIBLE DISPLAY PROBLEM AT THE PHILIPS CENTRAL STATION. DURING THE PROCESS OF TROUBLESHOOTING THE DISPLAY (NO PROBLEM WAS FOUND) A SECONDARY ISSUE WAS NOTED. TWO OF THE THREE PHILIPS CENTRALS WERE DISPLAYING ALARM CONDITIONS BUT NO AUDIO WAS NOTED. ON FURTHER INSPECTION IT WAS FOUND THAT BOTH SETS OF ALARM SPEAKERS WERE UNPLUGGED (THE POWER BLOCK WAS UNPLUGGED). THE BIOMED TECHNICIAN REPOWERED (PLUGGED IN THE POWER SOURCE) THESE SPEAKERS AND THEN ROUNDED (INSPECTED) THE NURSING UNIT AND FOUND A THIRD CENTRAL WHERE THE AUDIO CABLE HAD BECOME DISCONNECTED FROM THE SPEAKER INPUT. ALTHOUGH THESE ARE LIKELY BUT NOT NECESSARILY USER INTRODUCED PROBLEMS IT IS THE OPINION OF THE BIOMED THAT THIS PARTICULAR SPEAKER ARRANGEMENT MAKES DEFEATING THE ALARM SYSTEM TOO EASY WHICH CAN LEAD TO NO ALARM AUDIO WHEN IT IS NEEDED. WE HAVE HAD THESE SINCE 2007 AND OVER THE YEARS I HAVE SEEN THIS SEVERAL TIMES. | 1
2067016 | THE CUSTOMER STATES THAT THE PROVUE GEL CAMERA MISREAD A 1+ REACTION AS NEGATIVE IN THE ANTIBODY SCREEN TEST. NO INCORRECT OR ERRONEOUS RESULTS WERE REPORTED AS A RESULT OF THIS INCIDENT. THERE WAS NO HARM TO ANY PATIENT. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND REPLACED THE PROBE AND WASH STATION. THE FE ALSO PERFORMED ALL PROBE ADJUSTMENTS AND RAN DIAGNOSTICS TO VERIFY ANALYZER OPERATION. THE CUSTOMER RAN AND PASSED QC. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4) | 0
2067036 | NURSE NOTED TELEMETRY PACK FELT HOT TO THE TOUCH. OPENED THE TELEMETRY PACK AND THE BATTERIES FELT VERY HOT TO TOUCH. THE OUTER COVERING OF ONE BATTERY MELTED OFF EXPOSING THE INNER SILVER PART OF THE BATTERY. | 0
2067324 | AS DESCRIBED BY THE ANESTHESIOLOGIST: PATIENT TO OR, INDUCED AND NOTED TO HAVE IMMEDIATE DROP IN O2 SATURATION, THE CRNA CERTIFIED REGISTERED NURSE ANESTHETIST STUDENT WAS UNABLE TO MOVE AIR WITH THE BAG/MASK. THE PATIENT WAS IMMEDIATELY INTUBATED AND THE ENDOTRACHEAL TUBE WAS SEEN VISUALLY TO PASS THROUGH THE VOCAL CORDS. THE MACHINE CIRCUIT WAS ATTACHED. HOWEVER, NEITHER HAND BAGGING OR MACHINE VENTILATION COULD MOVE ANY AIR INTO THE PATIENT. HIGH POSITIVE PRESSURE WAS APPLIED WITH A SUBSEQUENT RISING OF THE O2 SATS SATURATIONS INTO THE MID 80'S SUCH THAT THE ENTIRE HYPOXIC TIME WAS LESS THAN 1 MINUTE. OTHER MEDICAL INTERVENTIONS WERE MADE TO RULE OUT ANY MEDICAL CONDITIONS AND SUBSEQUENTLY THROUGH A PROCESS OF ELIMINATION; THE AIRWAY OR MACHINE WAS IMPLICATED. THE PATIENT WAS SWITCHED OVER TO A TRANSPORT CIRCUIT WHICH IS ESSENTIALLY A BAIN CIRCUIT AND PORTABLE O2 TANK AND WAS EASILY VENTILATED WITH RAPID NORMALIZATION OF O2 SAT'S SATURATIONS. THE MACHINE AND CIRCUIT WERE EXAMINED POST EVENT AND AN ABNORMALITY WAS FOUND IN THE HME WHICH CONSISTED OF A PLASTIC MEMBRANE COVERING ONE OF THE OPENINGS. | 0
2068351 | CUSTOMER REPORTS THAT THE REPORT ADDENDUM IS MISSING ON PACS. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
2069577 | A PT WAS ANESTHETIZED FOR A LONG PROCEDURE USING THE GE ADVANCE ANESTHESIA DELIVERY SYSTEM. DURING THE COURSE OF THE ANESTHETIC, IT BECAME NECESSARY TO CHANGE THE SODA-SORB IN THE MACHINE BECAUSE THE ABSORBENT WAS EXHAUSTED, EVIDENCED BY THE PT REBREATHING HIGHER LEVELS OF CO2. IN THE PROCESS OF CHANGING THE CANNISTER TO A FRESH ONE, AN O-RING BECAME DISLODGED FROM THE DEVICE. THIS WAS NOT RECOGNIZED IMMEDIATELY BECAUSE THE OPERATING ROOM WAS DARK FOR THE PROCEDURE, AND THE O-RING IS VERY SMALL. WHEN THE CANNISTER WAS REPLACED INTO THE SYSTEM, AND THE O-RING WAS NOT ON THE DEVICE, THE MACHINE LOST ITS SEAL AND WE WERE UNABLE TO VENTILATE THE PT. IT WAS COMPLICATED BY THE FACT THAT NO ONE KNEW THERE WAS AN O-RING ON THE CANNISTER THAT COULD BE DISLODGED AND ALSO THAT GE DID NOT SUPPLY ANY ADDITIONAL O-RINGS TO REPLACE ONE IF IT COMES OFF. WE HAD TO STOP THE SURGERY, VENTILATE THE PT WITH AN AMBU BAG WHILE STAFF WERE CRAWLING AROUND ON THE FLOOR TO FIND THE O-RING. WHEN IT WAS LOCATED, WE REPLACED IT (AFTER HAVING TO FIND EXACTLY WHERE IT HAD BEEN DISLODGED), AND THE PROCEDURE CONTINUED WITHOUT EVENT. THIS COULD HAVE CAUSED HARM BY NOT BEING ABLE TO VENTILATE THE PT. | 0
2069646 | THE REPORTER, THE PT'S MOTHER, REPORTED THAT ON TWO OCCASIONS, THE PUMP DID NOT RECORD A BOLUS AND THE PT SUBSEQUENTLY HAD A LOW BLOOD GLUCOSE LEVEL. ON (B)(6) 2011, THE PT TOOK AN INSULIN BOLUS AT DINNER USING THE PUMP BUT THE BOLUS WAS NOT RECORDED IN THE PUMP HISTORY. THE PT'S BLOOD GLUCOSE LEVEL DROPPED TO 33 MG/DL. THE PT'S MOTHER CONTACTED THE PHYSICIAN FOR ASSISTANCE/ADVICE. ON (B)(6) 2011, THE PT TOOK A BOLUS AT LUNCH WHICH WAS NOT RECORDED IN THE PUMP HISTORY, AND AFTERWARDS HER BLOOD GLUCOSE LEVEL WAS TESTED TO BE 61 MG/DL. THE PT ADMINISTERED SELF-TREATMENT AND DID NOT SEEK MEDICAL ATTENTION. TROUBLESHOOTING REVEALED THESE TWO BOLUSES WERE NOT RECORDED IN THE PUMP HISTORY. THE PUMP HAS BEEN RETURNED TO ANIMAS; HOWEVER, THE INVESTIGATION HAS NOT BEEN COMPLETED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. ONCE THE EVALUATION HAS BEEN COMPLETED, A SUPPLEMENTAL REPORT WILL BE FILED. | 0
2069800 | THE CUSTOMER STATES THAT THE ORTHO PROVUE REPORTED DISCREPANT RESULTS WITH THE D TYPING ON THE SAME PATIENT. CTS DISCUSSED WEAK D SAMPLE REACTIVITY IN THE GEL CARDS AS STATED IN THE LIMITATIONS OF THE IFU. THE CUSTOMER IS REFUSING SERVICE AT THIS TIME. THE CUSTOMER IS CONFIDENT THE ISSUE IS ISOLATED TO THIS ONE SAMPLE AND CUSTOMER FEELS SERVICE IS NOT WARRANTED AT THIS TIME. (B)(4). | 0
2070006 | THE ACCOUNT STATED THAT AN INITIAL ARCHITECT TOTAL B-HCG RESULT OF 69.82 MIU/ML WAS GENERATED. THE SAMPLE WAS REPEATED AND A RESULT OF < 1.2 MIU/ML WAS GENERATED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). EVALUATION: WASH ZONE MANIFOLDS, WASH ZONE 2 PROBES. A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR REPLACED THE WASH ZONE MANIFOLDS, REPLACED THE PROBES IN WASH ZONE 2 AND CLEANED THE PROCESS PATH AFTER WHICH NO FURTHER ISSUES WERE REPORTED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). THE INITIAL REPORT WAS SUBMITTED UNDER BRAND NAME ABBOTT ARCHITECT TOTAL B-HCG REAGENT, (B)(4) MANUFACTURING SITE. IT WAS DETERMINED THAT THE INVESTIGATION SHOULD BE PERFORMED ON BRAND NAME ABBOTT ARCHITECT I2000 ANALYZER, (B)(4) MANUFACTURING SITE. | 0
2070686 | THE CUSTOMER STATED AN ARCHITECT I1000SR ANALYZER GENERATED FALSELY ELEVATED B-HCG RESULTS FOR ONE PATIENT SAMPLE. DATA WAS REVIEWED AND FOUND NOT TO AFFECT THE CLINICAL INTERPRETATION OR MEDICAL DECISION MAKING. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A PRODUCT DEFICIENCY HAS BEEN IDENTIFIED. WITH A UNIQUE COMBINATION OF INSTRUMENT AND REAGENT CONDITIONS THERE IS A POTENTIAL FOR SAMPLE CARRYOVER WHICH CAN RESULT IN FALSELY ELEVATED B-HCG CONCENTRATIONS ON THE ARCHITECT I1000SR SYSTEM. NEGATIVE SAMPLES HAVE RETURNED BOTH POSITIVE AND GREY ZONE B-HCG RESULTS WHEN A HIGH CONCENTRATION B-HCG SAMPLE IS ASSESSED PRIOR TO THE NEGATIVE SAMPLE WITH ARCHITECT TOTAL B-HCG (LIST NUMBER (B)(4)) ON AN I1000SR INSTRUMENT WHICH ALSO UTILIZES THE ARCHITECT RUBELLA IGG (LIST NUMBER (B)(4)) ASSAY. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED.><THE SUSPECT MEDICAL DEVICE HAS CHANGED FROM THE ARCH I1000SR, (B)(4) TO ARCH(B)(4), MANUFACTURING SITE (B)(4). MFR # 3005094123-2011-00564 HAS BEEN SUBMITTED TO ADDRESS THIS ERROR. | 0
2070697 | THE CUSTOMER STATED AN ARCHITECT I1000SR ANALYZER GENERATED FALSELY ELEVATED B-HCG RESULTS FOR ONE PATIENT SAMPLE. DATA WAS REVIEWED AND FOUND NOT TO AFFECT THE CLINICAL INTERPRETATION OR MEDICAL DECISION MAKING. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. THE SUSPECT MEDICAL DEVICE HAS CHANGED FROM THE (B)(4). MFR # 3005094123-2011-00556 HAS BEEN SUBMITTED TO ADDRESS THIS ERROR.><(B)(4). A PRODUCT DEFICIENCY HAS BEEN IDENTIFIED. WITH A UNIQUE COMBINATION OF INSTRUMENT AND REAGENT CONDITIONS THERE IS A POTENTIAL FOR SAMPLE CARRYOVER WHICH CAN RESULT IN FALSELY ELEVATED B-HCG CONCENTRATIONS ON THE ARCHITECT I1000SR SYSTEM. NEGATIVE SAMPLES HAVE RETURNED BOTH POSITIVE AND GREY ZONE B-HCG RESULTS WHEN A HIGH CONCENTRATION B-HCG SAMPLE IS ASSESSED PRIOR TO THE NEGATIVE SAMPLE WITH ARCHITECT TOTAL B-HCG (LIST NUMBER 7K78/6C21) ON AN I1000SR INSTRUMENT WHICH ALSO UTILIZES THE ARCHITECT RUBELLA IGG (LIST NUMBER 6C17) ASSAY. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. AN INVESTIGATION IS IN PROCESS | 0
2071298 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER LH 750 ANALYZER GAVE ERRONEOUSLY HIGH RETICULOCYTE OF 50 WITH AN R FLAG. THE CUSTOMER PERFORMED A MANUAL COUNT AND OBTAINED A CORRECT RESULT OF 1.1. NO ERRONEOUS RESULTS WERE REPORTED OUT OF THE LABORATORY. THE RESULT OF THE MANUAL COUNT WAS ACCEPTED AS THE CORRECT RESULT AND WAS REPORTED OUT. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. THE INSTRUMENT IS CURRENTLY WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS (ACCURACY) AND REPEATABILITY (PRECISION). CONTROLS WERE RUN BEFORE THE CF EVENT; ALL RECOVERED WITHIN ASSAY LIMITS, CONTROLS WERE NOT RAN AFTER THE EVENT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND REPLACED THE RED STRIP TUBING FOR THE STAIN AND RETIC MIXING CHAMBERS. THE FSE ALSO REPLACED THE DIFF AND RETIC SAMPLE LINES TO THE FLOW CELL. THE CUSTOMER WAS INSTRUCTED TO "PRIME" BEFORE RUNNING BABY SAMPLES TO HELP AVOID EXCESSIVE R FLAGS. THE ROOT CAUSE IS INSTRUMENT RELATED AS THE FSE HAD TO REPLACE HARDWARE. | 0
2071779 | THE REPORTER REPORTED THAT ON (B)(6) 2011, THE PT OBTAINED ELEVATED BLOOD GLUCOSE LEVELS: AT 7:00 AM, THE PT'S READING WAS 501 MG/DL, AT 7:30 AM, 528 MG/DL AND AT 8:30 AM, 528 MG/DL. AT THAT TIME, THE PT EXPERIENCED THE SYMPTOM OF NAUSEA. THE PT ADMINISTERED SELF-TREATMENT WITH A DOSE OF INSULIN TAKEN VIA A SYRINGE INJECTION. AT 9:00 AM, THE PT TOOK HERSELF TO THE EMERGENCY ROOM, WHERE SHE RECEIVED TREATMENT WITH INTRAVENOUS FLUIDS; NO ADD'L INSULIN WAS GIVEN. THE PT'S SUBSEQUENT BLOOD GLUCOSE READING WAS 309 MG/DL. TROUBLESHOOTING REVEALED THE PT HAD NOT CHANGED THE INFUSION SITE AFTER OBTAINING ELEVATED BLOOD GLUCOSE READINGS AS RECOMMENDED, THE INSULIN WAS REFRIGERATED AND THERE WERE BUBBLES IN THE TUBING. THE PT USED REFRIGERATED INSULIN WHICH CAN LEAD TO BUBBLES IN THE CARTRIDGE AND CANNULA. HOWEVER, AS THE PT OBTAINED BLOOD GLUCOSE LEVELS SUGGESTING SEVERE HYPERGLYCEMIA WHILE USING THE PUMP, AND RECEIVED EMERGENCY MEDICAL ATTENTION AND TREATMENT, THIS COMPLAINT IS BEING REPORTED. THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVAL. IF THE DEVICE IS RETURNED, AN EVAL SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. | 0
2071953 | WHEN THE UNIT IS TURNED "ON" OR "OFF" WHILE THE NETWORK CABLE IS PLUGGED IN, THE TIME CHANGES BY 4 TO 6 HOURS. THE INCORRECT TIME IS THEN REFLECTED IN THE TIME THAT THE EKG WAS TAKEN. THIS IS A KNOWN ISSUE BY GE AND WILL BE CORRECTED WITH THE NEXT SOFTWARE UPDATE.MANUFACTURER RESPONSE FOR ELECTROCARDIOGRAPH, MAC 1600: WILL BE CORRECTED WITH THE NEXT SOFTWARE UPDATE (DATE UNKNOWN). | 1
2073249 | DURING ROUTINE PROCESS REVIEW, IT WAS DISCOVERED THAT APNEA MONITOR SENSOR WITH BELTS DID NOT CARRY A MANUFACTER'S LABEL ON PACKAGE. WHILE NURSES KNEW AND RECOGNIZED THE PRODUCT, ABSENCE OF MANUFACTURER'S LABEL INTRODUCES POTENTIAL FOR CONFUSION.THIS PRODUCT AS IT IS DISTRIBUTED DOES CARRY A HOSPITAL-PRODUCED LABEL WITH ITEM NUMBER, BAR CODE AND SHORT DESCRIPTION. HOWEVER, THIS IS A DISTRIBUTION LABEL. ======================MANUFACTURER RESPONSE FOR APNEA MONITOR ACCESSORY KIT, START-UP ACCESSORY KIT======================VENDOR ASSURED US OF CALL BACK; HAVE YET TO HEAR BACK FROM MANUFACTURER. | 0
2073603 | A DISTRIBUTOR OF THE (B)(4) RECEIVED A REPORT FROM THE MOTHER OF A BABY STATING THAT THE DEVICE HAD BURNED HER BABY'S FOOT. PT INFO HAS BEEN REQUESTED, BUT NO RESPONSE HAS BEEN RECEIVED TO DATE. THE EXACT ACCIDENT DATE IS UNK, BUT IT IS BELIEVED THAT THE EVENT OCCURRED ON OR BEFORE (B)(6) 2011. NO ADDITIONAL INFO WAS RECEIVED ABOUT THE SENSOR TYPE OR THE ALLEGED INJURY. SEVERAL REQUESTS WERE MADE FOR ADDITIONAL INFO AND FOR RETURN OF THE UNIT FOR INVESTIGATION. NO ADDITIONAL INFO HAS BEEN RECEIVED NOR HAVE ANY UNITS BEEN RETURNED. THERE ARE THREE TYPES OF SENSORS THAT CAN BE USED WITH THE (B)(4). ALL OF THESE ARE WRAP SENSORS (REUSABLE AND DISPOSABLE MODELS) AND THE CUSTOMER COULD HAVE USED ANY OF THESE SENSORS. THERE HAVE BEEN NO COMPLAINTS INVOLVING ALLEGED BURNS IN WHICH THE MEASURED TEMP OF THE SENSOR EXCEEDED THE ALLOWED TEMPERATURE OF 41 DEGREES CELSIUS (B)(4). IN MOST CASES, IT WAS DETERMINED THAT THE CAUSE OF THE BURNS WAS DUE TO THE USER NOT ROTATING THE SENSOR AS FREQUENTLY AS INSTRUCTED IN THE OPERATOR'S MANUAL. SINCE NO SAMPLES WERE RETURNED FOR INVESTIGATION THE ROOT CAUSE FOR THIS COMPLAINT COULD NOT BE DETERMINED. | 0
2073706 | A CUSTOMER REPORTED EXPERIENCING INACCURATE BIOMETRY MEASUREMENTS WITH EQUIPMENT. THE DIAGNOSTIC PROCEDURE WILL BE RESCHEDULED. THERE WAS NO HARM OR INJURY TO THE PATIENTS. A COMPANY REPRESENTATIVE EXAMINED THE SYSTEM AND PROBE, AND FOUND THEM TO MEET ALL PRODUCT SPECIFICATIONS. NO SAMPLES RETURNED FOR EVAL AND NO ADDITIONAL INFO PROVIDED RELATED TO THIS EVENT. FOR THIS REASON, STEPS COULD NOT BE TAKEN TO REPLICATE OR CONFIRM THE REPORTED EVENT. THE ROOT CAUSE IS UNK. (B)(4). | 1
2074013 | ANESTHESIA WOULD NOT DELIVER A BREATH TO THE PATIENT INTERMITTENTLY. SYSTEM DISPLAYED ERROR "UNABLE TO DRIVE BELLOWS" AND ALARMED.======================MANUFACTURER RESPONSE FOR ANESTHESIA DELIVERY SYSTEM, AISYS======================GE SERVICE ENGINEER RESPONDED AND WAS NOT ABLE TO DUPLICATE COMPLAINT. HOSPITAL ENGINEER TESTED AND WAS ABLE TO DUPLICATE COMPLAINT AT SPECIFIC O2 FLOW OF 10LPM LITERS PER MINUTE AND TIDAL VOLUME OF 400LPM. CONTACTED GE AGAIN AND GAVE THEM THE SPECIFIC OPERATING PARAMETERS WHERE SYSTEM IS FAILING. GE SERVICE ENGINEER REPLACED VENTILATOR INTERFACE BOARD AND RETURNED MACHINE TO SERVICE. HOSPITAL ENGINEER TESTED AGAIN, AND MACHINE STILL EXHIBITS SAME ERROR. GE HAS BEEN CONTACTED AGAIN. PARTS ARE ON ORDER; NOT REPAIRED YET. | 0
2074970 | THE MONITOR DID NOT ALARM FOR AN APNEA EVENT, HOWEVER, THE INDICATOR LIGHT FOR APNEA WAS ON. DATES OF USE: (B)(6) 2011. | 1
2075021 | TABLE WAS RECLINING WITH THE PATIENT STANDING ON THE FOOT BOARD WHEN THE RIGHT SIDE OF THE FOOT BOARD BROKE, WITH THE PATIENT STILL ON IT. THE PATIENT WAS NOT HARMED AND THE TEST WAS COMPLETED. | 0
2075064 | THESE MONITORS ARE USED IN OUR PACU AREA ON PATIENTS WHO ARE COMING OUT OF SURGERY. VITAL SIGNS ARE IMPORTANT WHEN PATIENTS ARE RECOVERING FROM SURGERY AND ANESTHESIA. THE CARDIOCAP 5 HAS A SOFTWARE DEFECT: WHEN NIBP IS TAKEN ON A TIMED CYCLE, THE MONITOR WILL "MISS" RECORDING SOME OF THE RESULTS AND LEAVE A TIME GAP WITH NO VALUE IN THE TRENDS. THIS IS A PATIENT SAFETY AND DOCUMENTATION CONCERN FOR THE HOSPITAL. DATES OF USE: 2009 - 2011. | 1
2075302 | A CARIDIANBCT CUSTOMER REPORTS THAT A STAFF MEMBER HAD COMPLETED A COLLECTION PROCEDURE AND WAS REMOVING THE NEEDLE FROM THE DONORS ARM. THE STAFF MEMBER WAS USING THE ANTI-NEEDLESTICK PROTECTION AND INCURRED A NEEDLESTICK. CUSTOMER IS ALLEGING A DEFICIENCY WITH THE DISPOSABLE. STAFF MEMBER WASHED THE AREAS THOROUGHLY AND WAS SEEN IN THE EMPLOYEE HEALTH CLINIC. TESTING FOR INFECTIOUS DISEASES WAS PERFORMED ON THE STAFF MEMBER AS WELL AS THE DONOR. THE STAFF MEMBER WILL HAVE F/U APPOINTMENTS FOR TESTING AND COUNSELING ACCORDING TO ARC POLICIES. THE CUSTOMER WOULD NOT PROVIDE THE WEIGHT OF THE STAFF MEMBER. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2076099 | THE RPTR, THE PT'S FATHER, REPORTED THE INSULIN PUMP DISPLAY WAS BLANK, AND THAT THE PUMP WAS SOUNDING AN ALARM. THE PT TESTED HIS BLOOD GLUCOSE LEVEL TO BE 500 MG/DL. THE PT DID NOT EXPERIENCE ANY SYMPTOMS OF HIGH OR LOW BLOOD GLUCOSE LEVELS, AND DID NOT SEEK ANY MEDICAL ATTENTION OR TREATMENT. THE PT CORRECTED HIS BLOOD GLUCOSE LEVEL WITH INSULIN VIA A SYRINGE. THE RPTR CLAIMED THE PT DID NOT NOTICE OR REACT TO THE ALARMS FOR SOME TIME. TROUBLESHOOTING REVEALED THE BATTERY CAP WAS SECURE AND INTACT AND THERE WERE NO CRACKS IN THE PUMP CASING. NO PROGRAMMING TROUBLESHOOTING WAS PERFORMED DUE TO THE BLANK DISPLAY ISSUE. THE PT DID NOT NOTICE AND ADDRESS THE PUMP ALARMS IN A TIMELY MANNER, AND AFTERWARDS EXPERIENCED ELEVATED BLOOD GLUCOSE LEVELS. HOWEVER, THIS COMPLAINT IS BEING REPORTED SINCE THE PT EXPERIENCED A HYPERGLYCEMIC EPISODE WHILE USING THE INSULIN PUMP. THE PUMP WAS RETURNED TO ANIMAS FOR EVAL. EVAL REVEALED THE PUMP WAS ACCURATELY AND CORRECTLY DELIVERING INSULIN AS PROGRAMMED; ALL DAILY INSULIN DELIVERY TOTALS CORRECTLY REFLECTED THE USER'S PROGRAMMED BASAL RATES. THE PUMP WAS TESTED AND PASSED FLOW ACCURACY TESTING. THERE IS NO EVIDENCE THE PUMP CONTRIBUTED TO THE PT'S ELEVATED BLOOD GLUCOSE LEVEL. EVAL ALSO REVEALED A CRACK IN THE DISPLAY SCREEN. THE DISPLAY SCREEN WAS REPLACED, AND THEN FUNCTIONED PROPERLY. THE DISPLAY ISSUE IS NOT THE REASON FOR SUBMISSION OF THIS REPORT, AND DID NOT CONTRIBUTE TO THE SERIOUS INJURY. | 0
2077473 | IT WAS REPORTED THAT A PT WENT INTO ASYSTOLE FOLLOWED BY A COMPLETE ATRIOVENTRICULAR BLOCK WITH A PACEMAKER FAILURE. THE ALARM WAS NOT NOTICED BY THE CARE TEAM WHO WAS BUSY IN ANOTHER ROOM WITH MANY SIMULTANEOUS ALARMS OCCURRING AND NO SPECIFIC ALARM FOR THE ASYSTOLE. ACCORDING TO INFO COLLECTED FROM THE USERS, THE REPORTED ISSUE WAS DUE TO THE FACT THAT THERE WERE TOO MANY HIGH LEVEL ALARMS ACTIVATED SIMULTANEOUSLY. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. UNDER (B)(6) LAW, PT INFO IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSPITAL. | 1
2078546 | PT HAD AN INJECTION TO BOTH KNEES ON (B)(6) 2011 WITH 0.25% MARCAINE, 1% LIDOCAINE AND DEPOMEDROL USING A STERILE TECHNIQUE. ON (B)(6) 2011, PAIN TO HIS LEFT KNEE INCREASED AND SUBSEQUENT ASPIRATION REVEALED (B)(6) INFECTION. ANOTHER PT INJECTED BY ANOTHER HEALTH CARE PROVIDER ON THE SAME DATE ((B)(6) 2011) ALSO DEVELOPED THE SAME INFECTION TO HER RIGHT KNEE. EVALUATED ALL VARIABLES AND SIMILARITIES OF THESE TWO INJECTIONS AND THE ONLY COMMONLY USED MEDICAL DEVICE WAS A 10 CC SYRINGE FROM HENRY SCHEIN, INC EXPIRATION 12/2014, LOT # 20100125. WE HAVE QUARANTINED ALL OF OUR INJECTION MATERIALS AT THE TIME OF DISCOVERY WHICH WAS (B)(6) 2011. ALREADY CONTACTED THE FDA HOTLINE AFTER HOURS. BOTH PTS HAD SURGERY TO WASH OUT KNEES ON (B)(6) 2011 AT (B)(6) HOSPITAL (B)(6). EVENT ABATED AFTER USE STOPPED OR DOSE REDUCED: YES. | 0
2081151 | A SITE REPORTED A SYSTEM WIDE LOSS OF TELEMETRY MONITORING FOR 400 BEDS FOR APPROXIMATELY 30 MINUTES DUE TO A PRESUMED SYNCHRONIZATION MASTER HARDWARE FAILURE. A MANUAL PROCESS WAS REQUIRED TO RESTORE TELEMETRY. NO ALTERNATE MONITORING WAS ESTABLISHED DURING THIS TIME. THERE WERE NO PATIENT INJURIES REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
2082648 | THE PATIENT WAS ON THE STRESS MACHINE AND THE TREADMILL WAS IN THE BRUCE PROTOCOL DURING STAGE 3. THE PATIENT SUFFERED A CARDIAC CHANGE WHERE THEY STOPPED THE TREADMILL AND TREATED THE PATIENT WITH A CODE. AFTER THEY SETTLED THE PATIENT, THEY ENDED THE STRESS TEST AND DURING THE SAVE PROCESS, THE APPLICATION STOPPED RESPONDING. THIS RESULTED IN THE STUDY NOT BEING SAVED AND THE PATIENT DATA WAS LOST. THE ONLY HARD COPIES OF THE STRESS TEST OCCURRED BEFORE THE EVENT. AS A RESULT THE PATIENT WAS TRANSPORTED TO THE CATH LAB FOR FURTHER INTERVENTION AND THEY DID NOT HAVE THE NECESSARY PAPERWORK TO DO A PROPER EVALUATION.MANUFACTURER RESPONSE FOR CARDIAC TESTING SYSTEM, CASE: GE HEALTHCARE WAS TELEPHONED DURING THE EVENT AND THEY STATED THE ONLY OPTION IS TO REBOOT THE CASE MACHINE BECAUSE OF THE LOCKUP. ONCE REBOOTED THE DATA WAS LOST AND THE APPLICATIONS SPECIALIST STATED THAT NOTHING COULD BE RECOVERED. TECH SUPPORT WAS CONTACTED AFTER THE EVENT TO FORMALLY NOTIFY THEM OF OUR CONCERNS. GE TECH SUPPORT DISPATCHED A FIELD ENGINEER (FE). HE ARRIVED ONSITE THE NEXT DAY 4/27. THE FE DOWNLOADED LOGS FROM THE CASE SYSTEM FROM THE FOLLOWING WINDOWS LOGS: EVENT LOG, APPLICATION SECURITY, AND SYSTEM LOG. ALL UNDER SYSTEM TOOLS. TESTS ARE STORED IN RAM DURING THE PROCEDURE UNTIL THE PROCEDURE IS COMPLETELY SAVED AND ONLY TO A MAX OF 512 MEGABYTES. AWAITING FOLLOW UP FROM GE. | 1
2083300 | FROM PROVIDER REPORT: CONTINUOUS BLOOD LOSS OOZED THROUGHOUT CASE. IT WAS DIFFICULT TO TELL AMOUNT OF EBL (ESTIMATED BLOOD LOSS) BY VISUAL INSPECTION OF FIELD. SURGEONS STATE EBL WAS SIGNIFICANT. SURGEONS WERE SUCTIONING BLOOD "LOOSED" INTO NEPTUNE DEVICE THROUGHOUT CASE. IN MIDDLE OF CASE THE NEPTUNE DEVICE MEASURED 450 ML. INSPECTION OF FLUID CHAMBER = 1800 ML. CLINICAL ENGINEERING EVALUATION: EBL IS ESTIMATED BLOOD LOSS. UNIT WAS TESTED AFTER THE CASE AND CONFIRMED THAT THE VOLUME DISPLAY WAS GROSSLY INACCURATE AT THE LOW END OF THE RANGE. THE NEPTUNE HAS A DIGITAL DISPLAY TO INDICATE THE VOLUME IN THE CHAMBER. THE CHAMBER ALSO HAS GRADUATIONS ON THE SIDE FOR A COARSE MEASUREMENT. ACCORDING TO THE MANUFACTURER'S SPECIFICATION, THE DIGITAL DISPLAY SHOULD BE ACCURATE TO +/- 175ML. WE OBSERVED ERRORS IN EXCESS OF 1000ML ON SEVERAL MACHINES. EBL DURING THE CASE IS A FACTOR IN DECIDING WHETHER THE PATIENT NEEDS ADDITIONAL BLOOD TO MAKE UP FOR LOSS. | 0
2084699 | TWO HOSPITALS THAT BIOMEDICAL ENGINEERING IS RESPONSIBLE FOR HAVE THE PHILIPS MEDICAL CENTRAL MONITORING SYSTEMS. THE ISSUES ARE THE MOUSE AND OR THE KEYBOARD LOCKS UP AND THEN WE HAVE TO RE-BOOT THE SYSTEM FOR EACH OCCURRENCE. IN ADDITION THE TOUCH SCREEN DISPLAY WILL LOCK UP AND WE HAVE TO RE-BOOT THE SYSTEM FOR THIS MALFUNCTION ALSO. PHILIPS INCORPORATES SOFTWARE THEY CALL KAVOOM WHICH ALLOWS 1 MOUSE TO BE USED FOR SEVERAL DISPLAYS WITHOUT HAVING TO MANUALLY SWITCH FROM DISPLAY TO DISPLAY.OUR MOUSE PROBLEMS RANGED FROM CURSER STAYING ON 1 DISPLAY TO THE MOUSE SIMPLY LOCKING UP ON 1 DISPLAY. THE PROBLEM BEGAN SHORTLY AFTER INSTALLATION IN (B)(6) 2010. AT FIRST IT WAS JUST MEDICAL CARE UNIT (MCU).PHILIPS WAS CALLED. WE WERE INITIALLY TOLD TO RESET THE RECEIVE SWITCHES, KEYBOARD, VIDEO, MONITORS (KVM'S) AND INFORMATION CENTERS (SERVERS) TO RESOLVE THIS ISSUE. AFTER SEVERAL TIMES THE PHILIPS SERVICE REPRESENTATIVES CAME IN. FIRST THE KAVOOM SOFTWARE WAS DISABLED UNTIL THE SWITCH CONFIGURATIONS COULD BE CHECKED AND RE-CORRECTED. THE KAVOOM SOFTWARE WAS ENABLED AFTER THIS WAS COMPLETED. THE PROBLEMS DIDN'T GO AWAY. THE NEXT STEP WAS TO REPLACE THE RARITAN KVM SWITCH AT THE MCU CENTRAL WITH ANOTHER RARITAN SWITCH. THE PROBLEMS KEPT HAPPENING. EVENTUALLY THE PROBLEMS STARTED HAPPENING ON INTERMEDIATE MEDICAL UNIT (IMU). AFTER MANY RESETS PHILIPS REPLACED THE RARITAN KVM SWITCHES WITH BLACK BOX SERVOSWITCH KVM'S. THE MCU'S PROBLEMS DID SLOW DOWN, BUT NOT IN IMU. EVENTUALLY THE PHILIPS SERVICE REPRESENTATIVE REMOVED THE KVM SWITCHES FROM IMU COMPLETELY. AT ONE HOSPITAL THE SOFTWARE INFORMATION IS: ALL RUN ON WINDOWS VISTA, BUSINESS SOFTWARE. ALL USE SOFTWARE VERSION L.00.18, NETWORK VERSION E.00.23, AND APPLICATION VERSION L.00.16 AND AT THE OTHER HOSPITAL OUR M3155 INFORMATION CENTERS ARE RUNNING SOFTWARE LEVEL L.00.20. | 1
2087599 | THE PATIENT'S DUAL CHAMBER MEDTRONIC PACEMAKER WAS SET TO DEMAND PACING ONLY MODE (VVI) AT RATE OF 40 BEATS PER MINUTE DURING A PROCEDURE USING STEREOTAXIS MAGNETIC NAVIGATION SYSTEM (MNS). WHEN THE MNS SYSTEM WAS MOVED INTO NAVIGATE POSITION, THE PACEMAKER WENT INTO "MAGNET MODE" (VOO) ASYNCHRONOUS PACING AT 85 BEATS PER MINUTE. TO PREVENT COMPLICATIONS FROM INAPPROPRIATE PACING, THE OUTPUT WAS DECREASED FROM 2.0 V TO 0.5 V (TO INTENTIONALLY LOSE CAPTURE AND PREVENT PACING). DURING FIRST APPLICATION OF RADIO FREQUENCY (RF) ABLATION, ECG MORPHOLOGY CHANGED TO WHAT APPEARED TO BE PACED BEATS AT ABOUT 130 BEATS PER MINUTE AT HIGHER OUTPUT. WHEN RF ABLATION WAS TERMINATED, THE DEVICE WOULD STOP PACING AND THE PATIENT'S UNDERLYING TACHYCARDIA AT 145 BEATS PER MINUTE WAS SEEN ON ECG. THE PHENOMENON WAS REPEATEDLY REPRODUCED WITH RF APPLICATION. GROUNDING PAD LOCATION WAS MOVED FROM LEFT FLANK TO LEFT THIGH, WHICH INTERMITTENTLY RESOLVED THE ISSUE. MEDTRONIC REP ON SITE WAS CONSULTED, AND MEDTRONIC ENGINEER CALLED BY TELEPHONE, BUT NO CLEAR CAUSE WAS IDENTIFIED. MEDTRONIC REP ARRIVED IN ROOM TO PERFORM DEVICE INTERROGATION ONCE THE MNS SYSTEM WAS MOVED TO STOWED POSITION. ERROR MESSAGE ON DEVICE INDICATED IT HAD GONE TO RESET WHICH RETURNED DEVICE TO THE DEFAULT FACTORY SETTINGS OF DDD (DUAL-CHAMBER PACING) 60-130, OUTPUT OF 3.5 V. PATIENT'S LOW THRESHOLD (~1.0V) ALLOWED MAX TRACKING PACING AT 130 BEATS PER MINUTE TO CAPTURE, AND HER UNDERLYING TACHYCARDIA DID NOT REQUIRE PACING WHEN OFF RF. HAD THE PATIENT BEEN PACEMAKER DEPENDENT WITH A THRESHOLD ABOVE 3.5V, NO PACING OR LACK-OF-CAPTURE AND ASYSTOLE WOULD HAVE OCCURRED. CONVERSELY, HAD THE PATIENT HAD HEART FAILURE OR POOR CARDIAC FUNCTION, THE RAPID VENTRICULAR PACING AT 130 BEATS PER MINUTE MAY HAVE CAUSED HEMODYNAMIC COLLAPSE. THEREFORE, IT WAS FELT THIS SHOULD BE REPORTED AS A SIGNIFICANT EVENT WITH POTENTIAL TO CAUSE HARM. THE PACEMAKER IS STILL IMPLANTED AND THE PATIENT HAD A GOOD OUTCOME.======================MANUFACTURER RESPONSE FOR CARTO 3 RMT ELECTROANATOMIC MAPPING SYSTEM, EO MEDELCTROPHYSIOLOGY RECORDING SYSTEM, XPER HEMODYNAMIC RECORDING SYSTEM, ZOLL M SERIES DEFIBRILLATOR, SIEMENS ARTIS ZEE FLUORO SYSTEM, STEREOTAXIS/NIOBE REMOTE MAGNETIC NAVIGATION SYSTEM., MEDTRONIC DUAL CHAMBER PACEMAKER (PER SITE REPORTER)======================MEDTRONIC REP HERE DURING PROCEDURE. CALLED TECHNOLOGY DEPARTMENT. | 0
2088167 | SUBSEQUENT TO GENERAL SURGICAL OPERATION OF THE ABDOMEN, PT WAS ORDERED PARENTERAL DILAUDID EVERY TWO HOURS. THERE WAS OLIGURIA. ON THE SECOND AND THIRD POSTOPERATIVE DAY, THE PT COMPLAINED OF DYSPNEA. LATE IN THE AFTERNOON OF THE THIRD POSTOPERATIVE DAY, THE PT WAS ORDERED A BMP TEST. THERE HAD NOT BEEN ONE DONE ON DAY TWO. THE BMP TEST INDICATED A CREATININE MORE THAN THREE TIMES THE PT'S NORMAL AND PERI-OPERATIVE BASELINE. IT WAS ELECTRONICALLY TRANSMITTED TO THE (B)(6) BY THE LAB, SILENTLY, WITHOUT ANY INDICATION IT WAS COMPLETE. NO ONE SAW THE RESULTS. THE (B)(6) ARE OF DEFECTIVE DESIGN SINCE THERE IS NOT ANY WAY FOR THE CLINICIANS OF BEING INFORMED THAT NEW RESULTS HAVE ARRIVED, UNLESS A CARE TEAM MEMBER IS FOREVER SCANNING THE PT'S (B)(6). WITHOUT HAVING SEEN THE CREATININE RESULT, THE SEVERITY OF THE POSTOPERATIVE COMPLICATIONS WAS NOT APPRECIATED BY THE CARE TEAM OR THE RADIOLOGISTS. A COMPUTERIZED TOMOGRAPHY ANGIOGRAM WAS ORDERED AND COMPLETED DESPITE EVIDENCE ON THE BMP OF ACUTE RENAL FAILURE. THE BLOOD TEST RESULT INDICATING ACUTE RENAL FAILURE WAS MISSED BY THE MEDICAL CARE TEAM AND RADIOLOGISTS. THE RADIOLOGISTS SAW A DILATED STOMACH AND ESOPHAGUS, FILLED WITH FLUID. WHILE THE PT WAS ON A GENERAL MED-SURG WARD, AN ATTEMPT AT PASSING A NASO-GASTRIC TUBE WAS NOT SUCCESSFUL. WITHIN 12 HOURS, THE PT, WHO REMAINED ON THE GENERAL MED-SURG FLOOR VOMITED AND DIED, PRESUMABLY FROM ASPHYXIATION. | 1
2088195 | ERRONEOUS PACER SPIKES ON "T" WAVE ON MONITOR DISPLAY. PACEMAKER WAS INTERROGATED BY MANUFACTURE REPRESENTATIVE OF PACEMAKER ON UNIT WAS OPERATING PROPERLY. UNNECESSARY PATIENT CHARGES FOR PACEMAKER INTERROGATION DUE TO THE PROBLEM. THE MONITOR TECH HAD FOLLOWED INSTRUCTIONS TO REMEDY THIS SITUATION FROM THE MANUFACTURER'S CLINICAL INSTRUCTORS, PROBLEM NOT RESOLVED. | 1
2088203 | (B)(4), AND (B)(4) DEFECTS: THIS PT EXPERIENCED A CASCADE OF LIFE THREATENING ADVERSE EVENTS DUE TO TREATMENT REDUNDANCIES AND DELAYS IN TREATMENT ASSOCIATED WITH (B)(4) DEVICES. THE PT WITH PAROXYSMAL ATRIAL FIBRILLATION WAS RECEIVING HEPARIN SUBCUTANEOUSLY AWAITING A THERAPEUTIC WARFARIN DOSE TO ACHIEVE AN INR BETWEEN 2 AND 3. THE INR REACHED THAT GOAL AND WAS DELIVERED TO THE (B)(4) BY THE TESTING LAB. WE BELIEVE THAT THIS (B)(4) DEVICE IS FUNDAMENTALLY FLAWED IN SO FAR AS IT DOES NOT PROVIDE ANY WARNING THAT NEW RESULTS HAVE BEEN DEPOSITED. UNLESS SOMEONE IS CONTINUOUSLY SCANNING THE (B)(4) FOR NEW RESULTS, THIS DEFECT DE-FACTO RESULTS IN DELAYS IN CARE WITH ADVERSE CONSEQUENCES OF SAID DELAY DIRECTLY PROPORTIONAL TO THE INTRINSIC SEVERITY OF THE PTS' ILLNESS AND CO-MORBIDITIES. NO ONE ON THE MEDICAL TEAM NOR THE (B)(4) DEVICE THAT ENABLED CONCOMITANT AND ANTICOAGULANTS BECAME AWARE OF THAT THERAPEUTIC INR FOR 24 HOURS AND THE HEPARIN DOSES CONTINUED FOR HOURS CONCOMITANTLY WITH WARFARIN. LIFE THREATENING BLEEDING COMPLICATIONS ENSUED RESULTING FROM AT LEAST 3 MORE SHOTS THAN WOULD HAVE BEEN GIVEN HAD THE INR BEEN KNOWN, THIS ESCAPED THE NURSES, THE PHARMACISTS, AND THE HOSPITALISTS. ONCE THE BLEEDING WAS IDENTIFIED, REVERSAL OF THE INR WAS ATTEMPTED WITH FRESH FROZEN PLASMA AND VITAMIN K. A PTT WAS ORDERED AND THE RESULT OF GREATER THAN 130 WAS SILENTLY DEPOSITED IN TO THE (B)(4). THE SURGICAL CONSULTANTS, NURSES, PHARMACISTS, AND HOSPITALISTS DID NOT SEE THAT RESULT AND NO ACTION WAS TAKEN FOR MORE THAN 12 HOURS AFTER THAT RESULT BECAME AVAILABLE. THE PT CONTINUED TO BLEED. AGAIN, THERE WAS NOT ANY NOTIFICATION THAT A NEW LIFE-THREATENING RESULT BECAME AVAILABLE. IF ANY MEMBERS OF THE MEDICAL TEAM OPENED AND LOOKED - SEVERAL DID NOT - AT THE (B)(4), THE RESULT WAS MISSED, WE SUSPECT DUE TO HUMAN FACTORS ESPECIALLY, THE USER UNFRIENDLY PRESENTATION OF RESULTS ON THIS PARTICULAR DEVICE. THE PT REMAINS CRITICALLY ILL FROM THE BLEED, AND WOULD HAVE BENEFITTED FROM EARLY TREATMENT HAD THE DELAYS DUE TO THE FUNDAMENTAL FLAW IN THE (B)(4) AND (B)(4) DEVICE(S). | 1
2092188 | A CLINICAL SUPERVISOR NOTED THERE WERE VARIATIONS IN A PT'S AXIAL LENGTH DURING A BIOMETRY EXAMINATION. ALSO, THE CUSTOMER REPORTED THAT IT TOOK ABOUT FOUR MINS TO VIEW THE REGISTRATION SCREEN AND THE EQUIPMENT IS ONLY DISPLAYING IMMEDIATE PRINTING. THE EXAMINATION WAS COMPLETED; HOWEVER, THE DOCTOR CONSIDERED THE RESULTS TO BE UNRELIABLE. THE PT WAS SENT TO ANOTHER OFFICE FOR REMEASUREMENT. THERE WAS NO HARM OR INJURY TO THE PT. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4). | 1
2092525 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THAT THE T-LINK PROGRAM FAILED TO OPEN. THE DEVICE WAS NOT CHANGED OUT. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WERE NO ADVERSE CONSEQUENCES TO A PT AS A RESULT OF THIS EVENT. | 1
2094141 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 0
2094630 | IT WAS REPORTED THAT THE TRUSAT POWER SUPPLY GOT VERY HOT, WAS MAKING POPPING SOUNDS, AND THE CASE STARTED TO MELT. THERE WAS NO DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. THE INCIDENT DATE IS UNK, BUT IT IS BELIEVED THAT THE EVENT OCCURRED ON OR BEFORE (B)(6) 2011. | 0
2095466 | THE ROTATOR BROKE DURING A LEFT VENTRICULOGRAM PROCEDURE AT 812 PSI. INJECTOR SETTINGS: AT 9 ML/SEC FOR A TOTAL OF 36 ML AT 1000 PSI. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE DEVICES BUT DID NOT PROVIDE ANY FURTHER INFO OR CLINICAL DETAILS FOR THE ADD'L EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL: TWO USED SUSPECT DEVICES WERE RETURNED FOR EVAL. UPON VISUAL INSPECTION, THE COMPLAINT WAS CONFIRMED. THE ROTATOR WAS SEPARATED AT THE BOND BETWEEN THE COLLAR AND THE HOUSING. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND ONE SIMILAR COMPLAINT FOR THIS LOT NUMBER. THE CUSTOMER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. THE ROOT CAUSES OF THE FAILURE ARE ATTRIBUTED TO THE MFG BONDING PROCESS. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. THIS LOT WAS BUILT PRIOR TO IMPLEMENTATION OF THE NOTED CORRECTIVE ACTIONS. EVAL METHOD: DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. | 0
2096611 | TRAM MODULE FAILED TO FUNCTION. SOME WAVEFORMS WERE PRESENT, BUT THE CVP AND LA LINES WOULD NOT SHOW A TRACING OR REGISTER ON THE MONITOR. | 1
2096955 | THE CUSTOMER REPORTED THAT THE EXTENSIONS OF THE TIP, WHICH ARE CONNECTED TO THE CATHETER, BROKE WHILE INJECTING CONTRAST. THE CUSTOMER REPORTED THREE DEFECTIVE DEVICES BUT DID NOT PROVIDE ANY FURTHER INFO OR CLINICAL DETAILS FOR THE ADD'L EVENTS. THEREFORE, THIS SINGLE FORM FDA 3500A WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVAL: THE SUSPECT DEVICE HAS NOT BEEN RETURNED FOR EVAL. THE CUSTOMER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. EVAL: CONCLUSIONS: A F/U REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. | 0
2098979 | WHILE IN POWER CHAIR FROM DRIVE MEDICAL VR2 DIGITAL CONTROL UNIT THE CABLE FROM THE HAND CONTROL UNIT FELL CUT. I WAS THROWN TO GROUND, STRUCK ON CONCRETE. DUE TO INCREASE IN PAIN, MORPHINE INTERNAL PUMP HAD TO BE INCREASED 50%. I FOUND DEFECT WITHIN 24 HOURS OF RECEIVING UNIT. WAITED FOR DRIVE TO OPEN ON MONDAY, (B)(6) 2010. THEY SAID THEY WOULD SEND NEW CONTROLLER (HAND). THEY NEVER DID. THREE TIMES IN THE NEXT 9-10 MONTHS, THEY NEVER SENT UNIT OR FIX PROBLEM. ON THE FOURTH TRY, THEY GOT UNIT TO ME BUT NOT AFTER CHARGING FOR IT USED 900.00. THE REPLACEMENT PART HAS SAME DEFECT. YOU CAN NOT LOCK IN. THE DRIVE MAN (B)(6) TRIED TO SAY I DID NOT INSTALL RIGHT BUT HAD NOT HAD THE PARTS INSTALLED YET. SAID WARRANTY WAS VOID. THE MAJOR SAFETY ISSUE IS PERSON BEING THROWN FROM CHAIR. SOME ARC HAS BEEN HEARD AT DISCONNECT FIRE. WHEN CHAIR IS STUCK SOMEONE COULD FREEZE TO DEATH, HEAD AND FALLING INJURY. AT MAX SPEED AND SLIT DECLINE COULD BE DEADLY. DRIVE VR2 CONTROLLER IS DEFECTIVE. (B)(6). | 0
2101447 | LEGAL COMPLAINT RECEIVED ON (B)(6) 2011 ALLEGING THAT DURING/AFTER PLACEMENT OF SCREWS AND RODS THAT NERVES WERE DAMAGED DURING SURGERY ON (B)(6) 2009. SPINE LEVELS AND SPECIFIC PRODUCTS ARE NOT DEFINED EXCEPT TO STATE THAT A NUVASIVE NERVE MONITORING DEVICE WAS USED DURING SURGERY AND DID NOT PERFORM PROPERLY OR SAFELY. (B)(4). THE SPECIFIC DEVICE INVOLVED IN THE REPORTED EVENT HAS NOT BEEN IDENTIFIED. NO FAILURE INVESTIGATION HAS BEEN POSSIBLE TO DATE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE A SUBSEQUENT REPORT WILL BE FILED. | 0
2102234 | THE CUSTOMER REPORTED THAT THE INJECTOR TUBING CAME APART ON INJECTION. THE DEVICE WAS DISCARDED AT THE HOSPITAL. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE DEVICES BUT DID NOT PROVIDE ANY FURTHER INFO OR CLINICAL DETAILS FOR THE ADDITIONAL EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVALUATION: OTHER, THE SUSPECT DEVICE WILL NOT BE RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID REVEAL EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. THE CUSTOMER REPORTED THAT THEY WERE UNSURE IF THE LOT NUMBER REPORTED WAS CORRECT. UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED FAILURE WITHOUT THE SUSPECT DEVICE. NO CONCLUSION CAN BE DRAWN. EVALUATION METHOD: DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. RESULTS: OTHER, UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED FAILURE WITHOUT THE SUSPECT DEVICE. CONCLUSIONS: NO CONCLUSION CAN BE DRAWN. | 0
2102480 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THAT SWITCHING FROM THE TABLE TO THE GRAFT MODE, THEY LOST AN ACCURATE DISPLAY OF THE ART BLOOD FLOW. THE DEVICE WAS NOT CHANGED OUT. THE CASE WAS COMPLETED SUCCESSFULLY, AND THERE WERE NO REPORTED ADVERSE CONSEQUENCES TO THE PT AS A RESULT OF THIS EVENT. | 1
2103104 | A BED WAS BEING MOVED WHEN IT CAUGHT ONE OF THE PATIENT MONITOR CABLES. THE BRACKET SEPARATED CAUSING THE MONITOR TO FALL. THE CAUSE FOR THE SEPARATION WAS EVALUATED TO BE FROM THE BOLT COMING LOOSE DUE TO NORMAL ROTATION OF THE PIVOTAL ARM AND POSSIBLE VIBRATION. THE BRACKET WAS FROM A GCX WALL MOUNT. THIS WALL MOUNT WAS INSTALLED AND SUPPLIED BY PHILIPS MEDICAL.MANUFACTURER RESPONSE FOR PHYSIOLOGICAL MONITOR, WALL MOUNT: WAS SENT A REPLACEMENT. | 0
2104388 | THE ACCOUNT GENERATED ELEVATED ARCHITECT CALCIUM RESULTS ON 3 PATIENTS WHO REPEATED IN THE LAB NORMAL RANGE OF 8.0 TO 10.4 MG/DL. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED. NO FALSELY ELEVATED ARCHITECT CALCIUM RESULTS WERE REPORTED OUT OF THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. PATIENT 3: INITIAL ARCHITECT CALCIUM = 15.5 MG/DL, REPEAT ARCHITECT CALCIUM = 10.0 MG/DL. NO CONSEQUENCES OR IMPACT TO PATIENT (B)(4); HIGH TEST RESULTS (B)(4). AN EVALUATION WAS PERFORMED BY REVIEWING THE COMPLAINT TEXT, INSTRUMENT SERVICE HISTORY, AND OTHER CUSTOMER COMPLAINTS FOR SIMILAR ISSUES. THE CUSTOMER ISSUE WAS RESOLVED BY REPLACING THE SAMPLE AND R1 PROBES. TROUBLESHOOTING AND DIAGNOSTICS IN THE ARCHITECT SYSTEM OPERATIONS MANUAL LISTS SEVERAL PROBABLE CAUSES AND CORRECTIVE ACTIONS FOR ERRATIC RESULTS, INCLUDING SAMPLE OR REAGENT PROBE IS PARTIALLY OBSTRUCTED, PROTEIN BUILDUP INSIDE REAGENT PROBE, SAMPLE PROBE OUT OF ALIGNMENT, SAMPLE OR REAGENT PROBE IS DAMAGED, SAMPLE CONTAINS FIBRIN CLOTS OR PARTICULATE MATTER, AND BUBBLES OR FOAM ON THE SURFACE OF THE SAMPLE. CORRECTIVE ACTIONS INCLUDE CLEANING, REPLACING, AND CALIBRATING SAMPLE AND REAGENT PROBES AND REMOVING FIBRIN, BUBBLES, AND FOAM OR CENTRIFUGING THE SAMPLE. BASED UPON THE DATA AVAILABLE AND THE RESULTS OF THIS EVALUATION NO PRODUCT DEFICIENCY WAS IDENTIFIED.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
2104770 | CUSTOMER REPORTED UNEXPECTED NEGATIVE RESULTS WHEN TESTING DONOR SAMPLES WITH THE ANTI-C PHENOTYPE ASSAY ON THE GALILEO. CUSTOMER TESTED 6 DONOR SEGMENT SAMPLES AND 2 TESTED NEGATIVE FOR ANTI-C. WHEN THE SAMPLES WERE TESTED MANUALLY, THEY RESULTED AS POSITIVE. CUSTOMER RE-TESTED THE SAMPLES ON THE GALILEO WITH THE SAME LOT OF REAGENTS AND THE SAMPLES TESTED POSITIVE. ALL OTHER SAMPLES RESULTED AS EXPECTED. AFTER REVIEW OF IMAGES, INFORMED THE CUSTOMER THAT THE DILUTIONS APPEAR TOO LIGHT. A SERVICE CALL WAS MADE. -THE ISSUE WAS RESOLVED BY REPLACING A DILUTER ON THE INSTRUMENT.-PERFORMED ANTI-C PHENOTYPE ASSAY ON THE TWO SAMPLES IN QUESTION. RESULTS WERE AS EXPECTED. -UNIT WAS TESTED AND OPERATING WITHIN SPECIFICATIONS. | 0
2104960 | THE CUSTOMER HAD REPORTED FALSE POSITIVE REACTIONS CAUSED BY CARRY OVER. THE CUSTOMER REPORTED A POSITIVE BIOTESTCELL 2 CELL SCREENING RESULT FOR CAP SURVEY SAMPLE JAT-2 AND THE RESULT BACK FROM CAP IS THAT THE RESULTS SHOULD HAVE BEEN NEGATIVE. CELL 1 WAS 3+ AND CELL 2 WAS 1+, PANELS WERE RUN AND AGAIN JAT-2 WAS RUN IMMEDIATELY AFTER JAT-1 AND THE FIRST 4 CELLS OF THE PANEL FOR JAT-2 WERE POSITIVE. CAP SURVEY SAMPLE JAT-1 WAS RUN IMMEDIATELY BEFORE JAT-2 AND JAT-1 WAS VERY STRONG >4000 TITER ANTI-D. THE CORRECT FUNCTION OF THE AFFECTED LOTS OF REAGENTS WERE PROVED BY TESTING THE RETENTION SAMPLES OF QUALITY CONTROL LABORATORY ON TANGO. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DIDN'T OBSERVE ANY FALSE POSITIVE REACTIONS. A REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE AFFECTED NEGATIVELY THE QUALITY OF THE COMPLAINED LOTS. IN GENERAL CARRY-OVER EVENTS CANNOT BE EXCLUDED IF APPLYING INSTRUMENTATION WITH A DESIGN LAYOUT, THAT IS AMONGST OTHERS, ALSO TRUE FOR TANGO OPTIMO. THIS MATTER IS ALSO DISPLAYED IN ITS SPECS. DESPITE COATED NEEDLES AS WELL AS RINSING STEPS, SUCH A PHENOMENON MAY OCCUR IN VERY RARE CASES, IS USUALLY BASED ON A SAMPLE-SPECIFIC REACTION AND CAN BE CAUSED BY A HIGH PROTEIN LEVEL IN PLASMA (INDUCED BY SEVERE SICKNESS OR APPLICATION OF DRUGS) OR ANTIBODIES WITH A HIGH TITRE. SINCE ANY SPECIMEN BEING CONSPICUOUS DUE TO A POSITIVE ABS RESULT SHOULD UNDERGO RETESTING AND BECAUSE ANY ABS-POSITIVE SAMPLE FROM A BLOOD DONOR SHOULD BE EXCLUDED FROM TRANSFUSION, NO SERIOUS THREAT NEITHER FOR PT NOR USER OR DEVICE MAY ARISE DUE TO A CARRY-OVER EVENT. IN ACCORDANCE TO THE RESULTS OBTAINED AT THE CUSTOMER SITE, TESTING AT BMD CONFIRMED A CARRY-OVER EVENT FROM ONE WELL INTO THE FOLLOWING ONE OCCURRING DURING SEQUENTIAL PIPETTING OF DIFFERENT BLOOD PLASMA CAUSED BY THE HIGH TITRE OF ANTIBODY PRESENT IN THE ORIGINAL SAMPLE. EACH PIPETTING STEP OF INDIVIDUAL SAMPLES IS FOLLOWED BY A RINSING STEP OF THE CORRESPONDING NEEDLE RESULTING IN A DILUTION OF BIOLOGICAL RESIDUES BY APPROX. 1:1000. IN GENERAL THIS DILUTION FACTOR IS SUFFICIENT TO PREVENT A DETECTABLE TRANSFER OF MATERIAL INTO THE FOLLOWING SPECIMEN. | 0
2108083 | PT, A PREMATURE NEONATE IN NEONATAL ICU. PT RECEIVING TOTAL PARENTERAL NUTRITION (TPN). TPN COMPOUNDED USING BAXA AUTOMATED COMPOUNDER WITH ABACUS SOFTWARE. DEXTROSE IS ORDERED AT HOSP AS DEXTROSE IN MG/KG/MINUTE TPN SOLUTION, EACH BAG, TO BE INFUSED OVER A DURATION OF 24 HOURS. ENTIRE ORDER (ALL ORDERED COMPONENTS) ENTERED INTO COMPUTER CORRECTLY. "DURATION" OF INFUSION ENTRY LINE WAS 24 HOURS PER USUAL ROUTINE. OPERATOR ERROR: LINE FOR ENTRY OF "PERIOD" (DEFINED BY BAXA AS THE TIME PERIOD WITHIN WHICH A BAG WOULD BE RUN) IN THIS CASE, IT SHOULD HAVE BEEN A BAG INFUSED OVER 24 HOUR DURATION, EVERY 24 HOURS "PERIOD". THE NUMBER IN THIS "PERIOD" LINE (WHICH IS CHANGEABLE BUT NEVER CHANGED AT THE HOSPITAL) WAS INADVERTENTLY CHANGED TO "4." THIS RESULTED IN AN AUTOMATIC RE-CALCULATION OF THE DEXTROSE COMPONENT OF THE TPN SUCH THAT THE PUMP WOULD DELIVER 6 TIMES MORE DEXTROSE THAN ENTERED. THE LABEL THAT IS PRINTED BY THE PROGRAM LISTED THE COMPONENTS AS ENTERED (INCLUDING DEXTROSE 6 MG/KG/MIN) BUT DID CHANGE THE GIR FROM 6 TO 36, CHANGED KCAL AND CHANGED OSMOLARITY. THE RPH AND RN DOUBLE CHECK THE TPN LABEL VERSUS THE ORIGINAL PHYSICIAN ORDER BY LOOKING AT THE LISTED COMPONENTS ON THE LABEL, INCLUDING THE DEXTROSE RATE (UNCHANGED AND CORRECT AT 6 MG/KG/MIN), BUT DO NOT ROUTINELY CHECK THE GIR AS IT SHOULD BE THE SAME AS THE DEXTROSE INFUSION RATE. THUS THE LABEL IS NOT CONSISTENTLY CHANGED WHEN THIS AUTOMATIC RE-CALCULATION OCCURS, LEADING TO FALSE CONFIRMATION OF CORRECT ADDED DEXTROSE. AFTER THE RN DOUBLE CHECKED THE TPN AND CONFIRMED THE DEXTROSE RATE AS 6 MG/KG/MIN (BUT NOT THE GIR), IT WAS STARTED ON THE PT. SOME 3-4 HOURS LATER A BLOOD GLUCOSE (BG) WAS MEASURED AT APPROXIMATELY 450 MG/DL, THEN RE-MEASURED AND CONFIRMED. THE PT HAD MULTIPLE BGS OVER 600 MG/DL, INSULIN WAS STARTED WITH GRADUAL NORMALIZATION OF THE BG. DURING THE ENTIRE EPISODE THE CHILD REMAINED VIGOROUS AND WITHOUT CLINICALLY OBVIOUS ILL EFFECTS. IN SUMMARY, A CHANGE IN A SELDOM USED AND POORLY DEFINED ORDER ENTRY CELL IN A TPN COMPOUNDER PROGRAM "PERIOD" RESULTED IN AUTOMATIC RE-CALCULATION OF THE TPN CONTENTS WITH SIGNIFICANT INCREASE IN DEXTROSE ADDED TO THE TPN. THIS AUTOMATIC CHANGE WAS LARGELY TRANSPARENT, NO DOSE LIMIT ALERTS OCCURRED AND WAS HARD TO DETECT WHEN COMPARING LABEL TO ORDER. THE DEXTROSE INFUSION RATE IN MG/KG/MIN STATED ON THE LABEL DID NOT CHANGE ALTHOUGH THE GIR DID (THEY SHOULD BE IDENTICAL). THIS INCONSISTENCY MADE THE LABEL "LOOK" CORRECT WHEN DOUBLE CHECKED VERSUS THE ORDER. THE MFR WAS CONTACTED AND THE FOLLOWING CHANGES BEING MADE: THE "PERIOD" ENTRY "LOCKED" TO 24; THE DURATION ENTRY DEFAULTS TO 24 AND "ALERT" OCCURS IF CHANGED; EVALUATING LABEL FORMAT TO DETERMINE HOW "GIR" CAN BE MORE APPARENT; INTERNAL HOSPITAL SYSTEM CHANGES BEING IMPLEMENTED. | 1
2109849 | PATIENT UNDERWENT A POSTOPERATIVE X-RAY FOR ALIGNMENT VERIFICATION AFTER AN OPEN REDUCTION AND INTERNAL FIXATION OF HER LEFT TIBIAL PLATEAU. WHEN THE TECHNICIAN SAVED THE EXAMINATION, THE INCORRECT CASSETTE IMAGE IDENTIFIER WAS USED. THE SOFTWARE DOES NOT ALLOW FOR CORRECTION.THE MANUFACTURER IS SENDING A REPRESENTATIVE TO REVIEW ISSUE. | 1
2109858 | PATIENT REQUIRED LEFT SHOULDER X-RAY. EXAM WAS ENTERED AS RIGHT SHOULDER. A LEFT SHOULDER FILM WAS TAKEN, HOWEVER THE RADIOLOGY SOFTWARE WOULD NOT ALLOW FOR SIDE CORRECTION IN THE FOOTER OF THE FILM.MANUFACTURER IS SENDING REPRESENTATIVE FOR REVIEW OF SOFTWARE. | 1
2110330 | THE PROBLEM WAS THAT THE UNIT HAD A MALFUNCTIONING CLINICAL INFORMATION CENTER (CIC, BEDROCK) SYSTEM TO MONITOR PATIENTS VIA TELEMETRY. FOR SOME REASON WE WERE ONLY ABLE TO MONITOR 8 OF 16 PATIENTS ON CIC 1 AND THE OTHER TEN COULD NOT BE ADMITTED TO THIS CIC. CIC2 WAS NOT ABLE TO ADMIT 14 PATIENTS DUE TO THE SAME NO COMMUNICATIONS ("NO COMMS") ERROR. MEANING YOU SELECT THE ADMIT TAB AND WHEN YOU SELECT THE ECG FROM THE TAB, THERE WERE NO BEDS OR TELEMETRY TRANSMITTER NUMBERS (TTX) TO CHOOSE FROM. I REBOOTED BOTH CIC'S AND WAS ABLE TO VIEW OTHER DEPARTMENTS SYSTEMS WHICH INDICATED THE NETWORK WAS INTACT, BUT WHEN VIEWING THE DEPARTMENT WITH THE PROBLEM I WAS ONLY ABLE TO VIEW THE SAME 8 BEDS. I CHECKED FOR LOOSE CABLES, REBOOTED BOTH RECEIVER AND SERVERS. SINCE OUR EQUIPMENT IS LOCKED DOWN TO SPECIFIC DEPARTMENTS THE OTHER DEPARTMENT EXPERIENCED DELAYS AND THE PROBLEM STILL EXISTS. PLACED A CALL TO GE TO ESCALATE THE CALL. CALLED TECHNICAL SUPPORT AND PERFORMED THE SAME THINGS AGAIN. THIS TIME REBOOTED THE MASTER TOWER VIA COM. WAS NOW ABLE TO VIEW BEDS BUT THE TTX NUMBERS DID NOT EXIST. MANUALLY ENTERED TTX NUMBERS AND I TOLD THE UNIT NOT TO DISCHARGE UNTIL THIS PROBLEM WAS RESOLVED. UNIT CLERK DISCHARGED 3 OF THE ORIGINAL EIGHT TTX'S THAT WAS NOT EFFECTED EARLIER AND GOT NO COMMUNICATIONS ERROR, MANUALLY ADDED THE TTX'S FOR THE AFFECTED THREE TRANSMITTERS. THE NEXT DAY, I HAD THE MONITOR TECH PRINT STRIPS AND I REBOOTED SYSTEM. UPON BOOT-UP THE UNIT RETAINED THE ASSIGNED BEDS AND ASSIGNED TTX NUMBERS. WENT IN DISCHARGE EVERY PATIENT AND READMITTED THEM TO MAKE SURE THE BED AND TTX NUMBERS WERE RETAINED IN THE SYSTEM. GE SERVICE DID COME TO THE HOSPITAL AND COULD NOT FIND THE PROBLEM. I HAD GE DOWNLOAD THE LOGS AND SEND TO THE GE ENGINEERS FOR RESEARCH. WE DO NOT KNOW WHAT THE CAUSE OF THE PROBLEM OR WHY IT HAPPENED. STILL WAITING FOR RESPONSE. GE FIELD REPRESENTATIVE WAS VERY HELPFUL. | 1
2110380 | THE RN OBSERVED THE DISPLAY OF THE PATIENT'S ECG SHOWING ARTIFACT AND SHE BEGAN TO TROUBLE SHOOT FOR PROBLEMS. THE RN CHECKED THE PADS (ELECTRODES) AND THE LEADS AND THE CONNECTION TO THE MONITOR. THE RN ALSO EXCHANGED THE LEADS. NONE OF THE NURSE'S INTERVENTIONS IMPROVED THE VISUALIZATION OF THE ECG TRACING. ULTIMATELY, THE PATIENT DATA MODULE (PDM) WAS CHANGED AND AN ECG TRACING WAS OBTAINED. PATIENT OUTCOME WAS DEATH. IT IS UNKNOWN IF THE DEVICE MALFUNCTION CONTRIBUTED TO THE PATIENT OUTCOME. TWO COMPONENTS ARE BEING INVESTIGATED. THERE IS POTENTIAL THAT THE DEFAULT SETTING FOR NOTIFICATION OF THE ALARMS FOR THIS MONITOR WAS SET TO TEXT INSTEAD OF AUDIBLE ALARM, ALONG WITH THE UNRECOGNIZABLE ECG TRACING. | 1
2112419 | (B)(6) WEB EXTRACTION BASKET USED AND WIRES BROKE WHEN ATTEMPTED TO CRUSH THE STONE. STAFF WAS ABLE TO GET THE BASKET OFF THE STONE IN THE (B)(4) AND REMOVED THE BASKET. | 0
2113340 | THANK YOU ONCE AGAIN FOR LOOKING AT THIS SUBMISSION WITH PT CHART NUMBER AS THE IDENTIFIER AT (B)(6) IN AN ATTEMPT TO CONTINUE SUPPORTING DATA FOR SCENARIO #2 AS PREVIOUSLY ENTERED. MR. (B)(6), DIRECTOR OF NURSING WOULD, I'M SURE, BE HAPPY TO ASSIST THE FDA WITH THE FETAL MONITORING STRIP AND NOTES FROM THE SYSTEM TO SUPPORT MY SUSPECTED CLAIMS OF POTENTIAL LIFE THREATENING FLAWS ON THE AVALON FM FETAL MONITORING SYSTEM BY PHILIPS HEALTHCARE. RECALL # Z-0552-2010. MR. (B)(6) CAN BE REACHED AT (B)(6). AT 1343, THE MONITOR'S CCV SYSTEM DETECTS THE MATERNAL HEART RATE ABOVE 100 AND ALERTS THE USER TO A POSSIBLE COINCIDENCE WITH THE FETAL HEART RATE. WHEN LOOKING AT 1343 ON THE FETAL MONITOR STRIP ITSELF, IT IS OBVIOUS THAT THIS COINCIDENCE OF MOTHER AND BABY'S HEART RATE IS TAKING PLACE. AT 1345, "COINCIDENCE STATUS ON" ALARMS AGAIN. NOTICED HOW THERE IS AT LEAST A 20 BEAT DIFFERENCE IN BASELINE THIS TIME. AT 1348, "COINCIDENCE STATUS ON" ALARMS AGAIN. IT IS SKETCHY HERE BUT THERE IS A 10-15 BEAT DIFFERENCE EXACTLY AS THE CLASS 2 RECALL OUTLINES THIS TIME. AT 1353, MIMICKING FEATURES OF THE MATERNAL BASELINE IS NOTED IN COMPARISON TO THE FETAL BASELINE. AT 1355, "COINCIDENCE STATUS ON" ALARMS AGAIN. THE BEAT DIFFERENCE BETWEEN MATERNAL AND FETAL HEART RATES IS APPROX A 10 BEAT DIFFERENCE. AT 1358, "COINCIDENCE STATUS ON" ALARMS AGAIN. THE BEAT DIFFERENCE MATERNAL AND FETAL HEART RATES IS APPROX UP TO A 20 BEAT DIFFERENCE BUT NOTICE OFF TO THE SIDE OF THE ACTUAL COMPUTER MONITOR NOTES IT READS: COINCIDENCE STATUS ON FHR 1 MODE: NO TRANSDUCERS TOCO MODE: NO TRANSDUCERS FHR 1 MODE: US FMP STATUS: ON. WHAT DOES ALL THAT MEAN? BASICALLY, THE MONITOR IS SAYING FROM THE READINGS I GATHER FROM THE ORIGINAL PHILIPS AVALON FETAL MONITOR MANUAL: I'M DETECTING SIMILAR MATERNAL AND FETAL HEART RATES. BUT I DETECT YOU ARE USING AN (US) ULTRASOUND TRANSDUCER TO PICK UP FETAL HEART RATE BUT NOW I DETECT NO TRANSDUCER -I'M ALSO SUDDENLY DETECTING THAT -TOCO- CONTRACTION EXTERNAL MONITOR TRANSDUCER IS GONE AS WELL =SO I HAVE SWITCHED ON -FMP STATUS- TO DETECT IF THE INFANT IS STILL MOVING. THIS IS CONFUSING BUT THE MONITOR IS CONTINUING TO TRACE WITH THE SAME ULTRASOUND AND CONTRACTION CABLES EVEN THOUGH THE COMPUTER IS DETECTING THAT THE CABLES ARE GONE. I SEE THIS QUITE FREQUENTLY IN UNEVENTFUL TRACINGS AND NOT THE SAME MONITOR. AT 1402, "COINCIDENCE STATUS ON" ALARMS WITH A SKETCHY ERRATIC BEAT DIFFERENCE BETWEEN THE MOTHER AND FETUS HEART RATES AND THEN DELIVERY OCCURS. AT 1405 AND 1407, NOTICE THE ULTRASOUND CABLE CONTINUES TO TRACE THE APPEARANCE OF A FETAL HEART RATE AFTER THE DELIVERY AS IN SCENARIO #1 THAT I SUBMITTED ON MONDAY (B)(6) 2011 - THE TRACING IN THE ABSENCE OF FETAL HEART TONES AS VERIFIED WITH A SONOGRAM- SO THE QUESTION IS RAISED HOW OFTEN DOES THIS COMPUTER SYSTEM READ THAT THE ULTRASOUND AND TOCO ARE GONE BUT YET IT STILL TRACES? I SEE IT QUITE OFTEN BUT HAVE NOT ACTUALLY CALCULATED % THAT IS OCCURRING OUT OF THE 200 OF THE 450 STRIPS I WAS ABLE TO REVIEW BEFORE SUBMITTING VOLUNTARY REPORTS TO THE FDA. ALSO, MY QUESTION IN WHAT CAPACITY DOES IT IMPACT IN POTENTIALLY NEGATIVE FETAL OUTCOME EVENTS? DATES OF USE: (B)(6) 2009 - (B)(6) 2010. | 1
2113344 | THANK YOU FOR LOOKING INTO THIS EVENT. "THIS IS SCENARIO #2 POTENTIALLY LIFE THREATENING EVENT DUE TO VERY QUESTION RAISED POSSIBILITY AN UNNECESSARY CESAREAN SECTION OCCURRED DUE TO SUSPECTED SEVERE FLAWS AND CONTRADICTIONS AND UNREPORTED FLAWS IN THE CLASS 2 RECALL ON THE PHILIPS AVALON FM FETAL MONITORING SYSTEM -RECALL NUMBER Z-0552-2010- CATALOG NUMBER: M2705A. PT IDENTIFIER IS PT'S CHART NUMBER AT (B)(6) IN (B)(6). I AM UNABLE TO SEND THE ACTUAL FETAL MONITOR STRIP WITH NOTES FROM THE FETAL MONITOR COMPUTER OR ANYTHING FROM THE ACTUAL PT'S CHART AS IN SCENARIO #1 BUT AGAIN (B)(6), DIRECTOR OF NURSING WOULD BE GLAD TO ASSIST THE FDA WITH THIS POTENTIAL LIFE THREATENING FLAW NOT JUST IN THIS CASE BUT IN CURRENT AND FUTURE CASES WHERE THE MONITOR IS IN USE NATIONALLY AS WELL AS INTERNATIONALLY AS I WILL ATTEMPT TO PROVE AND RAISE SERIOUS DOUBT WITH WHAT IS HAPPENING IN THE FETAL MONITORING SYSTEM IN AT LEAST ONE OF 200 STRIPS OUT OF 450. I WAS ABLE TO REVIEW WHILE THIS FETAL MONITORING SYSTEM WAS IN PLACE AT MY FACILITY THAT WAS REPORTED TO THE FDA IN FAX AND MAILED MYSELF, (B)(6), RN IN (B)(6) 2010 WITH NO REPLY THUS FAR FROM THE FDA. (B)(6) CAN BE REACHED AT (B)(6) AND ONCE AGAIN THANK YOU FOR LOOKING INTO THIS POTENTIALLY LIFE THREATENING MATTER IN THIS MEDICAL DEVICE FOR BOTH MOTHER AND BABY. THIS FETAL MONITOR STRIP FOR PURPOSES OF THIS REPORT SHALL BEGIN AT THE TIME 2010 MARK ON (B)(6) 2009. THE MATERNAL HEART RATE IS ABOVE 100, THE PULSE OXIMETER IS IN USE AS OUTLINED IN THE CLASS 2 RECALL ON THIS FETAL MONITOR AND THE "CCV" SYSTEM IS ENABLE BY THE COMPUTER GENERATING "COINCIDENCE STATUS ON" BUT INSTEAD OF THE 10-15 BEAT DIFFERENCE TO BE INVESTIGATED BY VERIFYING FETAL HEART TONES AS OUTLINED ON PAGES ONE AND TWO OF THE CLASS 2 RECALL. THE BEAT DIFFERENCE IS ACTUALLY 10-25 BEAT DIFFERENCE. HOWEVER, THERE IS A FETAL SCALP ELECTRODE PLACED AT 2055 WHICH IS ONE OF THE VERIFICATION OF FETAL HEART TONES AS OUTLINED IN THE CLASS 2 RECALL. NOTICE, HOWEVER, THE "CCV" SYSTEM CONTINUES TO BE EMPLOYED AND THE MONITOR CONTINUES TO READ AND PRINT OUT "COINCIDENCE STATUS ON" WITH THE SAME 10-25 BEAT DIFFERENCE IN MATERNAL AND FETAL HEART TONES AS WELL AS 10-30 SECOND DELAY THAT IS NOT DISCUSSED AT ALL IN THE CLASS 2 RECALL. THE DELAY IS MENTIONED IN THE ORIGINAL MANUAL AS BEING UP TO A 6 SECOND DELAY. THIS WAS A QUESTION BROUGHT UP BY MYSELF ONCE AGAIN TO LEAD APPLICATION SPECIALIST FROM PHILIPS HEALTHCARE ON (B)(6) 2010 AT (B)(6) WITH MYSELF, (B)(6) -DIRECTOR OF MATERNAL CHILD HEALTH AT THE TIME- AND (B)(6) FROM BIOMED PRESENT. (B)(6) FROM PHILIPS STATED HOWEVER THAT THIS COULD BE "UP TO A 10 SECOND DELAY" IF SO, WHY IS THAT CHANGE NOT MENTIONED IN THE CLASS 2 RECALL? ALSO, IT IS CLEAR AS IN THIS FETAL MONITOR STRIP THAT IT IS ACTUALLY A 10-30 SECOND DELAY WHICH IN THIS CASE CAUSES AN APPEARANCE OF "LATE DECELS" THAT ARE VERIFIED WITH A FETAL SCALP ELECTRODE ON THIS POTENTIALLY COMPUTER LIFE THREATENING FLAW TO THE UNBORN FETUS AS THE FETAL HEART TONES ARE VERIFIED BUT THE COMPUTER CONTINUES AS "COINCIDENCE STATUS ON" WHICH IS A STRONG IDENTIFIER THAT THE MATERNAL HEART RATE CONTINUED TO BE DUPLICATED INSTEAD OF THE MOTHER AND FETUS. SO IT IS CLEAR THAT EVEN WITH VERIFYING FETAL HEART TONES AS OUTLINED IN THE CLASS 2 RECALL THAT THE "CCV" SYSTEM CONTINUES TO READ "COINCIDENCE STATUS ON" WHEN MATERNAL HEART RATE IS ABOVE 100 AND THE COMPUTER IS RECEIVING THAT IT IS PICKING UP THE SAME HEART RATE TWICE WHICH IN THIS CASE IS PROBABLY MATERNAL. SO, QUITE HONESTLY WAS BABY EVER MONITORED AT ALL DUE TO THE FLAWS OF THIS FETAL MONITORING SYSTEM? THIS WAS A GOOD OUTCOME FOR THIS INFANT BUT WHAT HAPPENS WITH FUTURE INFANTS IN THIS SAME QUESTIONABLY FLAWED FETAL MONITORING SYSTEM? IT IS NOT MY INTENT TO PLACE JUDGEMENT AT ALL, I AM HONESTLY RAISING THE SERIOUS QUESTIONS OF DOUBT HERE AND ASKING FOR FURTHER INVESTIGATION AND EVAL BY THE FDA. BY CONTINUING TO SEE "COINCIDENCE STATUS ON" EVEN AFTER THE SCALP ELECTRODE IS PLACED, THE MONITOR ITSELF IS QUESTIONING THAT IT IS THE SAME MATERNAL HEART RATE BUT WITH A 10-25 BEAT DIFFERENCE AS WELL AS A 10-30 SECOND DELAY WITH MIMICKING FEATURES, HOW CAN ANY SKILLED HEALTH CARE PROFESSIONAL BELIEVE WHAT THEY ARE SEEING WHEN ALL THE CLASS 2 RECALL GUIDELINES ARE BEING FOLLOWED. WHAT WAS THE ACTUAL FETAL HEART RATE? I WANTED TO ADD HERE ALSO THAT NEITHER THE ORIGINAL MANUAL OR THE CLASS 2 RECALL MENTION THAT THE "CCV" FEATURE IS TO EVEN REMAIN EMPLOYED WITH A FETAL SCALP ELECTRODE BUT THIS FETAL MONITORING STRIP CLEARLY DEMONSTRATES IT AS OCCURRING IN THIS INSTANCE, AND JUST IN CASE THIS QUESTION IS RAISED BY PHILIPS, HEALTHCARE. THIS IS NOT THE ONLY TIME A FETAL SCALP ELECTRODE CONTINUED WITH THE SAME POTENTIAL "CCV" FLAWS. I HAVE OTHER EXAMPLES OUT OF THE 200 STRIPS I REVIEWED OF THE 450 WHILE THEIR FETAL MONITORING SYSTEM WAS IN PLACE. MAY I ALSO ADD IT WAS DEFINITELY NOT THE SAME MONITOR OF THE 7 DOING THIS PRESENTATION OF THIS POTENTIAL FLAW, SO PLEASE DON'T BLAME IT ON JUST ONE OF THE 7 FETAL MONITORS MALFUNCTIONING OR THE SAME FETAL SCALP ELECTRODE CABLE MALFUNCTIONING, IT WAS SEVERAL OF THEM THAT DO THIS WITH A FETAL SCALP ELECTRODE IN PLACE BUT THIS WAS THE BEST EXAMPLE AND LENGTH AND AFFECTED OUTCOME FOR THIS SCENARIO #2. THAT IS ACTUALLY ONE OF THE FEATURES OF THIS FETAL MONITORING SYSTEM THAT I LOVE. IT TELLS THE USER WHICH MONITOR THEIR ON AND WHICH CABLE IS IN USE SO AS TO BETTER BE ABLE TO TRACK DOWN A FAULTY MONITOR OR CABLE, SO THE QUESTIONS ARE RAISED. UNNECESSARY CESAREAN SECTION DUE TO A FETAL MONITOR FLAW OR FLAWS EVEN WITH ALL THE MFRS GUIDELINES AND CLASS 2 GUIDELINES FOLLOWED? IF MATERNAL HEART RATE CONTINUES TO BE DUPLICATED WITH THESE SEPARATING RATES AND DELAYED RATE FEATURES IN FUTURE DELIVERIES WILL THE OUTCOME BE AS GOOD? AND IF THE IDEA THAT UNNECESSARY SURGERY WAS PERFORMED DUE TO THIS MEDICAL DEVICE POTENTIAL MALFUNCTIONS THIS PUT THE MOTHER AT A LIFE THREATENING RISK JUST AS ANY SURGERY DOES AND PROLONGED HER HOSPITALIZATION NOT TO MENTION THE POTENTIAL OF HEMORRHAGE THAT ALWAYS EXIST, INFECTION RISK, UNNECESSARY PAIN WITH SURGERY. I ALSO WANTED TO ADD THAT ON THE CONVERSATION WITH (B)(6) ON (B)(6) 2010, I ASKED THE QUESTION AS THE COMPUTER CONTINUES TO SAY "COINCIDENCE STATUS ON". AT WHAT POINT WILL IT ACTUALLY AUDIBLY ALARM AT THE NURSES' STATION? THE REPLY WAS "IF 80% OF ANY ONE MINUTE TIME FRAME IN ACTUAL MONITORING OF MATERNAL AND FETAL HEART RATES COINCIDE, THEN THE AUDIBLE ALARM WILL ALARM." HOW IS THIS EVER POSSIBLE WITH SEPARATION OF DUPLICATED HEART RATES EVEN WHEN BOTH ARE MOTHER'S HEART RATE AND DELAYED? DATE OF USE: (B)(6) 2009 - (B)(6) 2010. | 1
2113347 | THANK YOU ONCE AGAIN FOR LOOKING AT THIS VOLUNTARY SUBMISSION WITH PT CHART NUMBER AS THE IDENTIFIER AT (B)(6) IN (B)(6) IN YET ONE MORE ATTEMPT TO CONTINUE SUPPORTING DATA FOR SCENARIO #2. MR (B)(6), DIRECTOR OF NURSING WOULD CONTINUE TO BE HAPPY TO ASSIST THE FDA WITH THE FETAL MONITORING STRIP AND NOTES FROM THE SYSTEM TO CONTINUE TO POTENTIALLY IDENTIFY LIFE THREATENING FETAL MONITOR FLAWS ON THE AVALON FM FETAL MONITORING SYSTEM BY (B)(4), HEALTHCARE RECALL # Z-0552-2010 MR (B)(6) CAN BE REACHED AT (B)(6). IN THIS EXAMPLE, IT IS RELATIVELY UNEVENTFUL LABOR WITH A QUESTION OF VARIABLE DECELS STARTING TO OCCUR APPROX AT 1138. OTHERWISE IT APPEARS TO BE AN AVERAGE FETAL MONITOR STRIP FROM 0032-1138. AT 1150, A FETAL SCALP ELECTRODE -FSE- IS PLACED AS OUTLINED IN THE (B)(4) MANUAL AND THE CLASS 2 RECALL AT THE TOP OF PAGE 2. TWO SEPARATE TRACINGS OCCUR, MEANING THERE IS A CLEAR MATERNAL AND FETAL HEART RATE TRACINGS. AT 1206, "COINCIDENCE STATUS ON" ALARMS. BEFORE THIS POINT FROM 1150-1206, THE FETAL SCALP ELECTRODE APPEARED TO BE FUNCTIONING PROPERLY BUT SUDDENLY AT 1205, IT GIVES THE APPEARANCE OF SPIRALING AROUND THE MATERNAL HEART RATE SIMULTANEOUSLY AND W/O WELL KNOWN ARTIFACTS THAT CAN OCCUR WITH A FETAL SCALP ELECTRODE. SO THEREFORE, IF A FETAL SCALP ELECTRODE IS USED TO VERIFY FETAL HEART TONES AS OUTLINED IN THE CLASS 2 RECALL ON PAGES ONE AND TWO, THEN THE QUESTION IS THAN AGAIN RAISED. WHAT IS TO BE USED WHEN THE "VERIFY TO VERIFY" NEEDS VERIFICATION TO VERIFY WHAT'S THERE? ANOTHER WORDS ... "IF HEALTHCARE PROFESSIONALS WHEN QUESTIONING THE FETAL HEART TONES ON THE ULTRASOUND TRANSDUCER ARE VERIFIED WITH A FETAL SCALP ELECTRODE AND THEN THE SCALP ELECTRODE DECIDES TO CONTINUE TRACING W/O THE USUAL ARTIFACTS AND UNINTERRUPTED AROUND THE MATERNAL HEART RATE HOW IS THE HEALTHCARE PROFESSIONAL TO KNOW WHEN TO VERIFY WHAT'S ALREADY BEING VERIFIED AS OUTLINED IN THE CLASS 2 RECALL?" THANK YOU FOR YOUR CAREFUL CONSIDERATION IN THIS SERIOUS MATTER OF POTENTIALLY LIFE THREATENING FLAWED FETAL MONITORS THAT ARE IN USE NATIONALLY AS WELL AS INTERNATIONALLY. IN NURSING SCHOOL WE ARE TAUGHT TO BE PROACTIVE FOR OUR PTS. IN THE FACILITY THAT I WORK AT, I MADE A DIFFERENCE BY CONVINCING ADMINISTRATION TO HAVE THIS FETAL MONITOR REMOVED FROM MY FACILITY IN (B)(6) 2010. MY CONSCIOUS WON'T LET ME STOP THERE. MY SUBMISSIONS TO FDA MEDWATCH ARE FOLLOW UPS OF THE COMPLAINT SUBMITTED BY (B)(6) ON (B)(6) 2010 AT WHICH TIME ALL THAT OCCURRED WAS A REINSTATEMENT OF THE PREVIOUSLY ISSUED CLASS 2 RECALL. I FIND IT INTERESTING THOUGH THAT AFTER I SPOKE WITH (B)(6), LEAD APPLICATION SPECIALIST FROM (B)(6) HEALTHCARE ON (B)(6) 2010 WITH 11 VERY INTRICATE AND DIRECT QUESTIONS ABOUT THEIR FETAL MONITORING SYSTEM. (B)(6) HEALTHCARE ISSUED A FULL REFUND OF OVER 2 MILLION DOLLARS TO MY FACILITY WITH NO QUESTIONS ASKED. BUT AS A SINGLE MOTHER OF 3 AND 20 YRS OF LABOR AND DELIVERY EXPERIENCE, I CAN'T HELP BUT ASK MYSELF AS WELL AS EVERYONE AROUND ME AND THE FDA. WHAT IF IT WAS MY BABY? YOUR BABY? OUR BABY? WHAT WOULD YOU DO? I WISH THAT I HAD BEEN INCLUDED IN THAT REPORT, I WORKED SO HARD TO GET THE DATA FROM (B)(6) 2010. MAYBE I COULD HAVE ACTUALLY MADE A DIFFERENCE. WITH SINCERE APPRECIATION AND THANK YOU FOR LISTENING AND INVESTIGATING WHAT I HAVE TRIED TO DESCRIBED AS POTENTIALLY LIFE THREATENING FLAWS ON THE AVALON FM FETAL MONITORING SYSTEM. (B)(6), RN. | 1
2113350 | ADD'L INFO RECEIVED FROM RPTR (B)(4) 2011: FOLLOW UP REPORT TO THE PHILIPS AVALON FETAL MONITORS, I AM REFERENCING AS MY ACCESS NUMBER: (B)(4) FROM A LETTER I RECEIVED FROM (B)(4). DIVISION DIRECTOR, (B)(6) FOLLOW UP FROM 5 REPORTS I SUBMITTED TO MEDWATCH (B)(4). I WOULD LIKE TO TELL DR. (B)(4) THAT I HAVE 2 EMAILS DATED (B)(6) 2010 THAT A CONVERSATION OCCURRED WITH (B)(6), LEAD APPLICATION SPECIALIST ON (B)(6) 2010 FROM PHILIPS HEALTHCARE OF WHICH PROBLEMS NOT OUTLINED WITH THE RECALL NUMBER Z-055-2010 NOT LISTED IN THE CLASS 2 RECALL WAS DISCUSSED OVER THE TELEPHONE. CASE NUMBER AT PHILIPS THAT WAS ASSIGNED WAS (B)(4). PRIOR TO PHILIPS GIVING THE FACILITY (B)(6) A FULL REFUND, NO QUESTIONS ASKED FROM PHILIPS HEALTHCARE INC FOR THE RETURN OF THE PREVIOUSLY PURCHASED 7 AVALON FM 50 FETAL MONITORS BY THE FACILITY. THERE ARE EMAILS CONTAINING 11 PERTINENT DISCLOSING QUESTIONS AND MY FOLLOW UP IN EMAIL FORM OF THAT CONVERSATION THAT OCCURRED ON (B)(6) 2010 WITH LEAD APPLICATION SPECIALIST, (B)(6). I HAVE COPIES OF THESE EMAILS AND I AM COMING TO (B)(6) ON (B)(6) 2011 IF THEY ARE OF INTEREST TO DR (B)(6). I WAS GIVEN THE ACCESS NUMBER OF (B)(4) BUT IT IS UNCLEAR IF THIS IS THE CORRECT ROUTE TO CORRESPOND WITH THIS INFO. SN# (B)(4).><THANK YOU FOR CONTINUING TO LOOK INTO THIS EVENT. THIS EVENT IS ENTERED AS TO SUPPORT SCENARIO #2 AS PREVIOUSLY SUBMITTED. PT IDENTIFIER IS PT'S CHART NUMBER AT (B)(6) IN (B)(6). AGAIN, I AM UNABLE TO SEND THE ACTUAL FETAL MONITOR COMPUTER GENERATED STRIP OR THE NOTES IN THE FETAL MONITOR COMPUTER OR ANYTHING FROM THE ACTUAL PT'S CHART. PLEASE CONTACT (B)(6), DIRECTOR OF NURSING AS I'M SURE HE WOULD BE HAPPY TO ASSIST THE FDA WITH IDENTIFYING A POTENTIALLY SEVERELY FLAWED POTENTIALLY LIFE THREATENING FETAL MONITORING SYSTEM TO THE FDA OF POTENTIAL FLAWS NOT MENTIONED IN THE CLASS 2 RECALL ISSUED BY THE FDA (RECALL # Z-0552-2010). (B)(6) CAN BE REACHED AT (B)(6). I WILL CALL THIS SCENARIO #3 TO SUPPORT THE POTENTIAL LIFE THREATENING FLAWS AS DISCUSSED IN SCENARIO #2. IN THIS EXAMPLE, A FETAL SCALP ELECTRODE IS USED AS OUTLINED IN THE CLASS 2 RECALL ON PAGE ONE AND PAGE 2 TO VERIFY FETAL HEART TONES. FROM 1645 UNTIL OFF THE MONITOR AT 1711, THIS IS WHAT APPEARS TO BE A NORMAL LOOKING FETAL MONITOR STRIP THAT IS UNEVENTFUL. BUT UPON PLACING THE PT BACK ON THE FETAL MONITOR, THERE IS ERRATIC SIGNALS FROM THE ULTRASOUND TRANSDUCER. AT 1740-1741, A FETAL SCALP ELECTRODE WAS PLACED AS OUTLINED AT THE TOP OF PAGE 2 OF THE PHILIPS CLASS 2 RECALL. AT 1745, NOTICE HOW THE SCALP ELECTRODE WAS REGISTERING A FETAL HEART RATE 140-150'S THEN SUDDENLY DROPS TO MATERNAL HEART RATE. HOWEVER, AN INTERESTING FEATURE OCCURS. NOTICED AT 1746 AND 1748 HOW THE FETAL HEART RATE TRACING MIMICKS THE MATERNAL HEART RATE WITH AT LEAST A 20 BEAT DIFFERENCE FROM THE MATERNAL HEART RATE THAT IS IN THE 60'S WHICH IS WHY THE CCV SYSTEM WAS NEVER ENABLED. -BECAUSE MOM'S PULSE WAS BELOW 100- AT 1752, IT IS OBVIOUS THE -FSE- FETAL SCALP ELECTRODE IS PICKING UP A GOOD FETAL SIGNAL AND GOES ON TO DELIVER W/O INCIDENT. SO THIS REINFORCES THE QUESTION AGAIN. DOES THE CCV SYSTEM BECOME DISABLED OR CONTINUE TO BE ENABLED WHEN A FETAL SCALP ELECTRODE IS USED BECAUSE NOWHERE IN THE MANUAL OR CLASS 2 RECALL IS THIS QUESTION ANSWERED? ALSO, WHEN IT IS CLEAR THAT THE "CCV" SYSTEM IS NOT ENABLED AS WHEN THE MATERNAL HEART RATE WAS IN THE 60'S-NOT ABOVE 100 TO EMPLOY THE CCV SYSTEM- THERE IS A CLEAR 20 BEAT MIMICKING SEPARATION WITH THE VERIFICATION OF FETAL HEART TONES WITH A FETAL SCALP ELECTRODE, SO AS TO CLEARLY UNDERSTAND THAT GREATER THAN THE CLASS 2 RECALL DISCUSSED 10-15 BEAT DIFFERENCE ON PAGE ONE OF THE CLASS 2 RECALL, CAN ACTUALLY BE GREATER WHETHER THE "CCV" SYSTEM IS ENABLED OR NOT? DATES OF USE: (B)(6) 2009 - (B)(6) 2010.><ADDITIONAL INFO RECEIVED FROM THE REPORTER ON 12/24/2013: THIS REPORT IS A FOLLOW UP REPORT TO REPORTS I MADE TO THE FDA IN 2011. I FEEL DEEPLY COMMITTED AS I HAVE STUDIED THE AVALON FETAL MONITORING SYS FOR 4 YEARS NOW AND WITH 22 YEARS OF LABOR AND DELIVERY EXPERIENCE, I HAVE BECOME AN EXPERT ON THE AVALON FETAL MONITORING SYS. THIS REPORT IS A CULMINATION OF SOME OF THE INFO I FOUND REGARDING THE PHILIPS AVALON FETAL MONITORING SYS ON THE MAUDE WEBSITE. I WAS SO THANKFUL TO SEE THAT THE ORIGINAL PROGRAM WRITER FROM (B)(4) HAD DEVELOPED AN UPGRADE 3/18/2011 AND FIELD ORDER TO UPGRADE ALL THE UNITS IN MAY OF 2011. TESTING IN AUGUST 2011 SHOWED "THAT THE MATERNAL PULSE RATE OBTAINED FROM THE TOCO MP TRANSDUCER HAS BEEN RECORDED TOGETHER WITH THE ECG DERIVED MATERNAL HEART RATE." FROM THAT REPORT IN AUGUST OF 2011, IT APPEARS THAT THE FETAL HEART RATE WAS ACCURATE NOW COMPARED TO THE MATERNAL HEART RATE AS PREVIOUSLY STATED IN THE CLASS 2 RECALLS IN SEPTEMBER OF 2009. I STILL HAVE CONCERNS ABOUT THE FETAL SCALP ELECTRODE INACCURACIES LEADING TO AN UNNECESSARY CESAREAN SECTION MAY HAVE BEEN PERFORMED DUE TO THE FETAL SPIRAL ELECTRODE AS WELL. SO MY QUESTION HERE IS: WAS THIS FETAL MONITORING UNIT ALREADY SERVICED WITH THE FIELD ORDER TO UPGRADE ALL THE UNITS IN MAY OF 2011? IF SO, THEN PLEASE REFER TO REPORT #(B)(4) AND REPORT #9610816-2011-00335 TO SEE IF THE CCV SYSTEM WAS EMPLOYED IN BOTH SYSTEMS AFTER MATERNAL HEART RATE IN BOTH REPORTS IS GREATER THAN 100 BEATS PER MINUTE AND DOES THE "CCV" DERIVED FETAL SCALP ELECTRODE MONITOR TRACINGS HAVE ALSO BEEN CORRECTED IN THAT AUGUST 2011 REPORT? IN THE CLASS 2 RECALL, DATED 09/04/2009, PHILIPS DID AN EXCELLENT JOB AT EMPHASIZING CONFIRMATION OF FETAL LIFE BY USE OF DOPPLER, FETOSCOPE ULTRASOUND AND FETAL SPIRAL ELECTRODE. IN THE NEXT SERIES OF MAUDE REPORTS THAT I WILL LIST, PLEASE NOTE THAT THEY ARE SIMILAR REPORTS IN THAT THE SAME ISSUES ARE CLAIMED WITH FETAL SPIRAL ELECTRODES IN THE CASES OF DEMISE FETUSES AND UNNECESSARY CAESAREAN SECTIONS DUE TO FETAL SPIRAL ELECTRODE INACCURACIES, HAVE THESE FETAL MONITOR TRACINGS BEEN DETERMINED THAT THE "CCV" SYSTEM IS EMPLOYED DURING THE USE OF THE FETAL SPIRAL ELECTRODE IN EACH REPORT? PLEASE NOTE THAT THESE ARE JUST SOME OF SUCH EVENTS ON THE MAUDE WEBSITE. MANY OF SUCH OCCURRENCES MAY HAVE NEVER BEEN REPORTED TO THE FDA AS THESE TYPES OF DELIVERIES ARE USUALLY RECORDED AS IUFD (INTRAUTERINE FETAL DEMISE). I BELIEVE THAT THERE ARE HUNDREDS OF NOT THOUSANDS OF THESE TYPES OF DELIVERIES SINCE THE 1970'S. THESE EVENTS ARE WHERE A SPIRAL ELECTRODE HAS REGISTERED A NONVIABLE FETAL HEART RATE THROUGH A DEAD FETUS THAT IS ACTUALLY RECORDING THE MOTHER'S HEART RATE USING ANY FETAL MONITORING SYSTEM, NOT JUST THE PHILIPS AVALON FETAL MONITORING SYSTEM. REPORT NUMBER, REPORT DATE: 9610816-2013-00131, 06/13/2013; 9610816-2009-00205, 08/12/2009; 9610816-2009-00206, 10/19/2009; 9610816-2010-00042, 01/07/2010; 9610816-2011-00335, 05/27/2011; 9610816-2011-00345, 06/08/2011; 9610816-2005-00113, 04/01/2005; 9610816-2005-00115, 04/01/2005; 9610816-2006-00115, 03/08/2006; 9610816-2006-00244, 07/05/2006. ANOTHER QUESTION I HAVE IS IF THE SPIRAL ELECTRODE COMPONENT OF THE AVALON FETAL MONITORING SYSTEM HAS THE PATENT TO THE SPIRAL ELECTRODE WITH ANOTHER COMPANY? OR PHILIPS?. AS THE PREVIOUS COMPANY (B)(4) DID NOT HAVE THE PATENT TO THE SPIRAL ELECTRODE IN 1991, MY PERSONAL VERY FIRST FETAL MONITORING CLASS. IF SO, ARE THE FLAWS IN THIS SEPARATE COMPONENT NOT OWNED BY PHILIPS? IT IS DEFINITELY CLEAR IN THE MAUDE REPORTS THAT PHILIPS IS MAKING EVERY EFFORT TO IMPROVE THE ACCURACY OF THEIR AVALON FETAL MONITORING SYS. I FEEL THAT PHILIPS IS NOT ALWAYS TO BLAME IN THOSE REPORTS FOR THE FETAL DEATHS DESCRIBED IN THOSE SAMPLE MAUDE REPORTS. BUT RATHER THAT SOME OF THOSE FETAL DEATHS ARE PERHAPS A PHENOMENA WHERE THE FETUS IS NOT ALIVE BUT UPON PLACING A FETAL SPIRAL ELECTRODE ON A DEAD FETUS, OUR TECHNOLOGY IS ADVANCED TO THE POINT IT WILL SEARCH FOR A BABY HEART RATE AND THEN REGISTER MOM'S HEART RATE AS IT IS THE ONLY HEART RATE THE FETAL SCALP ELECTRODE DETECTS. I HAD THE EXPERIENCE OF HAVING A FETAL SPIRAL ELECTRODE REGISTER THROUGH AN UNBORN DEAD BABY AND REGISTER THE MOTHER'S HEART RATE. THERE WAS NO CONFIRMATION OF FETAL HEART TONES AS THIS IS NOT COMMON PRACTICE EVEN THOSE MANY YEARS AGO. EVEN THOSE BABIES NOT YET DEAD BUT WITH A HEART RATE LESS THAN 30.. THE FETAL SPIRAL ELECTRODE WILL NOT PICK UP THE FETAL HEART RATE LESS THAN 30. BUT IF DETECTED AN EMERGENCY CESAREAN SECTION COULD SAVE A BABIES LIFE JUST BY VERIFYING THE BABY HEART RATE WITH DOPPLER, ULTRASOUND, OR FETOSCOPE. IN DEFENSE FOR PHILIPS THEY SHOULD NOT BE ACCOUNTABLE FOR THE LACK OF CONFIRMATION OF FETAL LIFE BEFORE EVEN TURNING ON THE PHILIPS AVALON FETAL MONITORING SYSTEM AMONG THE CLINICIANS WHICH ARE MANY THAT DON'T PRACTICE CONFIRMATION OF FETAL LIFE. THROUGH THE USE OF ORGANIZATIONS THAT TEACH FETAL MONITORING SUCH AS (B)(6), INFANT LIVES COULD BE SAVED AS UPON PLACING A SPIRAL ELECTRODE ON ANY FETUSES THAT IT MUST BE 100% VERIFIED WITH DOPPLER, ULTRASOUND OR FETOSCOPE. THIS IS NOT HOW CLINICIANS CURRENTLY PRACTICE. WE RELY ON MACHINES TOO MUCH TO TELL US AN INFANTS STATUS. THE EFFECTS OF SUCH AN EVENT ON THE FAMILY IS DEVASTATING TO BE IN LABOR FOR HOURS THINKING YOU HAVE A LIVE BABY ONLY TO FIND OUT YOUR BABY IS DEAD WHEN IT'S BORN. ALSO TO THE CLINICIANS THAT MAY HAVE PTSD AFTER SUCH AN EVENT. WHEN I USE THE WORD CLINICIANS THROUGHOUT THIS REPORT.. THESE NEW STANDARD OF CARE SUGGESTIONS ARE FOR NOT JUST HOSPITALS BUT HOME BIRTHS AND MIDWIFE BIRTHING CENTERS AS WELL AS ACROSS OUR NATION BUT GLOBALLY AS WELL. THROUGH THE USE OF FETAL MONITOR INSTRUCTORS, ORGANIZATIONS SUCH AS (B)(6) AND PERHAPS PHILIPS COULD AID IN EDUCATING CLINICIANS ACROSS THE NATION TO BRING UP THE STANDARD OF CARE TO PREVENT THESE SITUATIONS FROM REPEATING. (B)(6) COULD TAKE THE INITIATIVE LIKE THE SUCCESSFUL 39 WEEK INITIATIVE, BECAUSE IT'S ALL ABOUT SAVING INFANT LIVES. SIMPLY BY CONFIRMING FETAL HEART RATE BY USING 2 METHODS NOT JUST ONE TECHNOLOGY. BEFORE EVER PLACING MOM ON EXTERNAL FETAL MONITOR AND VERIFY AGAIN IF INTERNAL FETAL SCALP ELECTRODE IS PLACED. VERIFY 100% OF TIME WHEN INTERNAL FETAL SCALP ELECTRODE IS PLACED. MATERNAL PULSE OXIMETER USED 100% OF TIME WHICH IS NOT ALWAYS DONE IN LABOR. CONSIDER VERIFYING EACH TIME PT OFF MONITOR. A POSITIVE CHANGE AS WE ADVANCE IN OUR TECHNOLOGY AND COMPANIES SUCH AS PHILIPS WORK DILIGENTLY TO CORRECT FLAWS IN THEIR NEW "CCV" TECHNOLOGY. ANY CLINICIAN IF YOU ASK THEM AND TELL THEM THAT JUST ONE INFANT COULD BE SAVED THIS WAY BY VERIFYING FETAL HEART TONES WITH 2 METHODS INSTEAD OF ONE, ANY CLINICIAN WOULD IMPROVE THEIR FETAL MONITORING PRACTICES IMMEDIATELY. THE EMPHASIS JUST NEEDS TO BE HEARD. WHAT BETTER WAY THAN THROUGH THE FDA. AS CLINICIAN WE ARE ALL DOING WHAT WE LOVE TO DO. TO BRING NEW LIFE TO THIS WORLD SAFELY BECAUSE THESE BABIES IN THEIR OWN WAY TOUCH OUR HEARTS AND LIFE IS PRECIOUS AND SO ARE THEY. WAS TRYING TO LINK THIS REPORT TO ONE I MADE IN 2011. (B)(4) IS THE ACCESS NUMBER. | 1
2114123 | THE CUSTOMER STATED AN ARCHITECT I2000SR ANALYZER GENERATED A FALSE POSITIVE TOXO IGG RESULT FOR ONE PATIENT SAMPLE. THE ARCHITECT GENERATED AN INITIAL TOXO IGG RESULT OF 5.7 AND A NEGATIVE RESULT WHEN REPEATED. THE FALSE POSITIVE TOXO IGG RESULT WAS NOT REPORTED OUTSIDE THE LABORATORY. AN ABBOTT FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO CLEAN THE LOAD/UNLOAD MANIFOLD AND NOTICED THE PROBE WAS OUT OF ALIGNMENT. THE CUSTOMER COMMENTED THAT HE HAD REPLACED THE PROBE BUT HAD NOT CALIBRATED IT. THE FSE CALIBRATED THE PROBE TO RESOLVE THE ISSUE, AND INSTRUCTED THE CUSTOMER TO CALIBRATE THE PROBE UPON REPLACEMENT. (B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND FOR THE COMPLAINT ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE OBTAINED EVIDENCE A PRODUCT DEFICIENCY WAS NOT IDENTIFIED. THE ISSUE WAS THE RESULT OF OPERATOR ERROR FOR FAILING TO CALIBRATE THE PROBE AS REQUIRED BY PROCEDURE FOLLOWING REPLACEMENT. | 0
2114132 | NIBP WAS NOT WORKING; INDICATED HARDWARE FAILURE. | 1
2117752 | FLUORO FROM A PREVIOUS CATHETERIZATION WAS MISSING FROM HORIZON CARDIOLOGY. MCKESSON (VENDOR) RESEARCHED, AND FOUND NUMEROUS STUDIES WHICH WERE NOT MIGRATED BY MCKESSON FROM OLD TO NEW LONG TERM ARCHIVE IN 2008. FORWARDED VENDOR LETTER AND SPREADSHEET OF MISSING STUDIES TO RISK MANAGEMENT. | 1
2120154 | IT WAS REPORTED THAT A PATIENT IN POOR GENERAL CONDITION WAS SUBJECTED TO ROUTINE MONITORING DURING THE COURSE OF AN ERCP EXAMINATION WITH AN (B)(4), SPO2 (OXISURE) AND NIBP. DURING THE COURSE OF THE EXAMINATION, THERE WERE REPEATED DROPS IN SPO2 LEVEL AND INCORRECT VALUES DISPLAYED ON THE MONITOR. THE MONITOR GAVE AN AUDIBLE ALARM WHENEVER THE ALARM LIMITS WERE VIOLATED. AFTER THE INCIDENT, THE MONITOR IN QUESTION WAS IMMEDIATELY TAKEN OUT OF OPERATION BY THE MEDICAL TECHNOLOGY DEPARTMENT. AN INDEPENDENT EXPERT WAS COMMISSIONED BY THE HOSPITAL MANAGEMENT TO EXAMINE THE MONITOR FOR PROPER FUNCTION AND ANY INCORRECT RESPONSE. ACCORDING TO THE MEDICAL TECHNOLOGY DEPARTMENT, THE DEVICE APPEARS TO BE FUNCTIONING NORMALLY, WITH INSPECTIONS AND SAFETY TESTS HAVING BEEN CARRIED OUT PUNCTUALLY AND AT THE CORRECT INTERVALS. THE PATIENT'S CONDITION DECLINED CONTINUOUSLY DURING THE COURSE OF THE EXAMINATION, AND AT TIMES THE SPO2 LEVEL COULD NOT EVEN BE MEASURED ANY LONGER. FOLLOWING EMERGENCY INTUBATION, THE PATIENT WAS PUT ON A VENTILATOR AND TRANSFERRED TO THE INTENSIVE CARE UNIT WHERE, ACCORDING TO THE HOSPITAL, HE DIED OVER THE WEEKEND. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2121071 | IT WAS REPORTED THAT A PT COMPLAINED OF A BURNING SENSATION ON THE LEGS AFTER A MAGNETIC RESONANCE (MR) EXAM. THE TORSO COIL CABLE WAS REPORTEDLY POSITIONED BETWEEN THE PT'S LEGS. THE HOSPITAL STATED THAT THE COIL CABLE WAS NOT TOUCHING THE PT'S SKIN. THE DAY AFTER THE EXAM, THE PT CALLED AND REPORTED THAT HE HAD TWO SMALL BLISTERS ON EACH CALF AND THE SIZE OF EACH BLISTER WAS LESS THAN A DIME. THE PT HAS NOT UNDERGONE ANY MEDICAL TREATMENT FOR THE BLISTERS. THE PT REFUSED TO RETURN TO THE MR DEPARTMENT FOR THE RADIOLOGIST TO EXAMINE THE PT. THE PT TOLD THE HOSPITAL HE ONLY CALLED SO THEY WOULD BE AWARE OF THE BLISTERS THAT HAD FORMED AFTER UNDERGOING THE MR EXAM. THE PT WAS POSITIONED FEET FIRST, SUPINE WITH ARMS AT SIDE. GE RECOMMENDED PADS WERE USED DURING THE SCAN. THERE WAS NO SKIN-TO-SKIN CONTACT REPORTED BY THE TECH. THE PULSE SEQUENCES USED WERE: (B)(4). THE PT WAS SCANNED FOR EIGHT SERIES. (B)(4). IT HAS BEEN DOCUMENTED THAT THE PT WAS WARMED IN THE AREA WHERE THE COIL CABLE WAS ROUTED BETWEEN THE PT'S LEGS. THE CUSTOMER WAS USING A DAMAGED COIL: CABLE INSULATION HAD PULLED AWAY FROM THE CONNECTOR FOR THE UPPER COIL. THE CABLE WAS ALSO SLIGHTLY MODIFIED: VELCRO WAS PLACED UNDERNEATH THE CONNECTOR ON THE TOP OF THE COIL SO IT WOULD NOT COME UNPLUGGED. THE OPERATOR MANUAL STATES THE FOLLOWING: "DO NOT USE SURFACE COILS WITH EXPOSED METAL CONDUCTORS OR DAMAGED INSULATION. SKIN CONTACT WITH METAL CONDUCTORS CAN CAUSE BURNS." "THE ELECTRICAL AND MECHANICAL ASSEMBLIES AND PARTS OF THE COIL MUST BE USED WITH CARE AND SHOULD BE ROUTINELY INSPECTED. BEFORE USING THE COIL, VISUALLY CHECK IT TO ENSURE THERE IS NO EXTERNAL DAMAGE. DO NOT USE THE COIL IF THE HOUSING OR CABLE IS BROKEN." | 0
2123682 | THE GE LUNAR SYSTEM HAD TO BE RESET. WHEN IT WAS RESET, THE POPULATION COMPARISON GROUP NEEDED TO BE SELECTED. THE COMPARISON GROUP "USA" WAS SELECTED, NOT "USA/NHANES" THIS CAUSED THE WRONG CALCULATION OF BONE LOSS. STAFF REPORTED THAT WHEN THE COMPARISON GROUP IS SELECTED, THE LIST OF COMPARISON GROUPS IS SO LONG THAT ONE NEEDS TO SCROLL THROUGH MANY SELECTIONS AND THE PROPER REFERENCE DATABASE IS NOT SEEN EASILY. THE LIST OF COMPARISON GROUPS INCLUDES MANY COUNTRIES OUTSIDE OF THE US THAT WOULD NOT BE SELECTED IN THE US. IF THE LIST WAS NOT AS LONG IT WAS FELT THAT THERE WOULD NOT BE THE RISK OF SELECTING THE WRONG REFERENCE DATABASE. STAFF SUGGESTED THAT THE LIST BE MODIFIED FOR US USERS SO THAT SO MANY UNNEEDED COMPARISON GROUPS DO NOT APPEAR.THE RESULT WAS INCORRECT DATA FOR 223 PATIENT EXAMS.THE SETTINGS WERE CORRECTED. THE MANUFACTURER WAS CONTACTED AND INSTRUCTIONS WERE GIVEN ON HOW TO CORRECT AND REANALYZE THE DATA. THE EXAMS WERE REANALYZED TO CORRECT THE RESULTS. | 1
2126506 | WE ARE INFORMED BY FDA WITH AN ADVERSE EVENT VIA EMAIL. ACCORDING TO THE EMAIL FROM (B)(6) OF FDA RECEIVED ON (B)(6) 2011, FDA RECEIVED A VOLUNTARY REPORT INVOLVING A DEATH OF A PERSON USING THE ADVOCATE DUO METER. BECAUSE THIS IS A VOLUNTARY REPORT, FDA CANNOT PROVIDE US WITH THE PT'S NAME OR PERSONAL INFORMATION DUE TO FDA'S CONFIDENTIALITY RULES. HOWEVER, ACCORDING TO THE REPORT FDA RECEIVED, THE INDIVIDUAL WAS ADMITTED TO A HOSPITAL AFTER INCURRING A FALL ON OR ABOUT (B)(6) 2009. THE METER GLUCOSE RESULTS DID NOT MATCH THE HOSPITAL GLUCOSE RESULTS FROM THEIR LABORATORY INSTRUMENT. THE DUO READINGS WERE 240 MG/DL AND THE HOSPITAL OBTAINED 439 MG/DL. ON REPEAT ANALYSIS, THE DUO READ 539 MG/DL AND THE HOSPITAL OBTAINED 440 MG/DL. FDA DO NOT HAVE ANY INFORMATION REGARDING THE LOT NUMBER OF TEST STRIPS THE PERSON WAS USING, NOR THE MODEL NUMBER OF THE DEVICE. THE PT DIED IN (B)(6) 2010. WE HAVE STARTED AN INVESTIGATION ABOUT THIS INCIDENT RIGHT AWAY WHEN WE WERE NOTIFIED BY FDA. DUE TO LACK OF INFORMATION OF STRIP LOT NUMBER AND METER SERIAL NUMBER, WE WERE UNABLE TO KNOW WHICH EXACT LOT OF METER OR STRIPS WERE USED BY PT. THEREFORE, WE PULLED OUT THE RECORD OF ADVOCATE DUO METERS AND STRIPS SHIPPED TO THE US DURING 2009 (B)(6) AND BACKTRACED TO 88 LOTS OF STRIPS AND 4 LOTS OF METERS ACCORDING TO THE DATE OF INCIDENT ON (B)(6) 2009. TO INVESTIGATE THE ROOT CAUSE AND TO VERIFY IF THIS INCIDENT IS RELATED TO THE DEVICE BASED ON THE LIMITED INFORMATION WE HAD, WE DECIDED TO SINGLE OUT 5 LARGEST AMOUNTS OF STRIP LOTS FROM 88 LOTS SHIPPED TO US FOR FURTHER TESTING. WE PERFORMED PERFORMANCE TESTS INCLUDING CONTROL SOLUTION AND WHOLE BLOOD TESTS WITH THE RESERVE TEST STRIPS FROM THE 4 SAME LOTS USING RESERVE METER SAMPLE OF THE SAME MODEL. BASED ON RESULTS OF THE TESTS, ALL CONTROL SOLUTION AND WHOLE BLOOD TEST RESULTS OF ADVOCATE DUO USING THE 5 STRIP LOTS WERE WITHIN THE ACCEPTANCE CRITERIA. RESULTS DEMONSTRATE THAT THOSE 5 STRIPS LOTS OF ADVOCATE DUO WORK NORMAL. WE ARE ALSO CONCERNED THAT THERE IS A GAP OF TIME FOR AROUND THREE MONTHS BETWEEN DATES OF THE PT ADMITTED TO HOSPITAL AND THE PT'S DEATH. WE HAVE NO IDEA WHAT HAS HAPPENED TO PT DURING WHICH TIME. WE HAVE NO INFORMATION ABOUT WHAT THE CAUSE OF DEATH IS AND NO EVIDENCE THAT SHOWS THE DEATH IS RELATED TO ADVOCATE DUO. WITHOUT DETAILS OF THE DEVICE USED BY PT AND INFORMATION ABOUT THE PT, WE ARE UNABLE TO DETERMINE IF ADVOCATE DUO IS RELATED TO OR CAUSE THE DEATH OF PT. | 0
2126968 | A 3160 INVIVO PRECESS MRI PT MONITOR WAS BEING USED DURING AN MRI WHEN IT BECAME STUCK TO THE MRI SCANNER. IT APPEARS IT WAS BEING USED WITHIN SAFE OPERATING DISTANCE. HOWEVER IT WAS NOTED THE WHEELS WERE NOT LOCKED. | 0
2128174 | ASKU.><IT WAS REPORTED THAT SESSION SYNC SENT DATA WITHOUT THE ALPHA CHARACTER PREFIXES AND IT COULD NOT BE DETERMINED WHICH PATIENT'S DATA IT WAS. A DEVICE UPDATE DISK WAS SENT TO THE CENTER. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. FURTHER ANALYSIS OF THE EVENT PROMPTED A CHANGE IN THE CONCLUSION CODE.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. | 1
2128544 | WE DISCOVERED THAT OVER A THREE DAY PERIOD ALL EKGS DONE ON ONE OF OUR DEVICES IN THE EMERGENCY DEPARTMENT WERE PERFORMED WITH THE EKG ELECTRODES IN THE WRONG PLACE. THIS INVOLVED A TOTAL OF ABOUT 32-34 EKGS. UPON INVESTIGATION, IT WAS DETERMINED THAT SOMEONE REMOVED THE EKG ELECTRODES FROM THE ACQUISITION MODULE AND REPLACED THEM IN THE WRONG POST ON THE ACQUISITION MODULE CAUSING V1 AND V3 TO BE REVERSED ON THE PATIENT. | 1
2128591 | TWO SEPARATE INSTANCES: ONE TIME A NURSE WAS CLICKING ON THE "SAVED PATIENT'S NAME" AND WHEN SHE CLICKED ON THE NAME, THE SCREEN WENT BLANK AND ERASED ALL THE PATIENT HISTORY.ANOTHER TIME THE CENTRAL MONITOR SCREENS SHUT DOWN SPONTANEOUSLY AND A FATAL ERROR MESSAGE APPEARED. NO ONE HAD TOUCHED THE MODEM OR MONITORS AT THIS TIME. THE SYSTEM REBOOTED AND CARDIAC MONITORING RESUMED. ACCORDING TO BIOMED, THE FAILURE WAS DUE TO THE AGE AND SERVICE PARTS ARE NO LONGER AVAILABLE. | 1
2129578 | STAFF ORIGINALLY HAD DIFFICULTY LOGGING ON TO THE OBIX SYSTEM. BY THE TIME CLINICAL ENGINEERING STAFF ARRIVED, THE LOGIN PROBLEM HAD RESOLVED BUT IT WAS NOTICED THAT THERE WAS NO AUDIO FOR THE ALARMS. FURTHER CONTACT WITH OBIX TECH SUPPORT SUGGESTED CHECKING THE MUTE FUNCTION ON THE COMPUTER. THE AUDIO WAS MUTED, AND AFTER CLICKING THE BUTTON, THE AUDIO WAS UNMUTED AND BEGAN WORKING PROPERLY. STAFF COMMENTED THAT THIS HAS BEEN AN ONGOING PROBLEM.THE AFFECTED ALARM AUDIO IS ONLY AT THE CENTRAL STATION (NURSES' DESK). IT DOES NOT AFFECT ALARMS AT THE BEDSIDE. VISUAL ALARMS ARE STILL FUNCTIONAL.THE MUTING FUNCTION IS ON THE WINDOWS XP OPERATING SYSTEM VOLUME CONTROL, SIMILAR TO WHAT YOU WOULD HAVE ON YOUR DESKTOP COMPUTER. THE CENTRAL STATION SOFTWARE IS RUNNING ON WINDOWS. TO ACCESS IT, YOU CLICK ON THE SPEAKER ICON IN THE SYSTEM TRAY, THEN THERE IS A CHECK BOX NEXT TO THE VOLUME SLIDER FOR MUTE. ====================== MANUFACTURER RESPONSE FOR PERINATAL DATA SYSTEM, OBIX (PER SITE REPORTER)======================APPARENTLY THEY ALREADY KNEW ABOUT THIS PROBLEM. THEY ARE WORKING ON A SOLUTION. | 1
2129696 | THE ADULT PATIENT ON THE MONITOR IN A ROOM HAD A 28 BEAT V-TACH AND THE TELEMETRY MONITOR SYSTEM (TMS) DID NOT ALARM.THE TECHNICIAN MONITORING DID SEE IT REAL TIME AND COMMUNICATE TO THE RN, HOWEVER THERE WAS NEVER AN ALARM FOR V-TACH. THE SYSTEM DIDN'T RECOGNIZE V-TACH! THE SOFTWARE MUST BE ENHANCED TO PREVENT MISSING REAL CRISIS LEVEL ALARMS. PATIENT WAS NOT HARMED BECAUSE THE TECHNICIAN NOTICED V-TACH AND ALERTED THE RN. | 1
2129703 | AT 11:39 AND 20:06 THE ADULT PATIENT HAD AN EPISODE OF VENTRICULAR TACHYCARDIA AND THE TELEMETRY MONITOR SYSTEM (TMS) DID NOT ALARM. THE TECHNICIAN MONITORING IN BOTH INSTANCES CAUGHT THE ABNORMALITY AND NOTIFIED THE RN FOR THE PATIENT REAL-TIME. BY NOT ALARMING, THE SYSTEM DID NOT RETAIN THE EVENT IN THE HISTORY WHICH PREVENTED THE CARDIOLOGIST FROM VIEWING THE EVENT. THE SYSTEM DID NOT RECOGNIZE THE ARRHYTHMIA. SOFTWARE ENHANCEMENTS ARE NEEDED THAT WOULD HELP RECOGNIZE CRISIS ALARMS. PATIENT WAS NOT HARMED IN THIS CASE, BUT IT WAS A NEAR-MISS. | 1
2129733 | AT 01:09:30 THE ADULT PATIENT HAD A RUN OF VENTRICULAR TACHYCARDIA AND THE GE TELEMETRY SYSTEM DID NOT ALARM FOR V-TACH. IT DID ALARM FOR TACHY WHICH IS A HEART RATE ALARM AND VT<2 AFTER THE ACTUAL RUN OCCURRED. THIS RUN DID MEET THE DEFINED CRITERIA FOR V-TACH. THE TECHNICIAN HAPPENED TO SEE THE RUN AND NOTIFIED THE RN IMMEDIATELY. HOWEVER, THERE IS A POTENTIAL FOR MISSING THIS TYPE OF ARRHYTHMIA IF THE SYSTEM DOES NOT ALARM PROPERLY. THE PATIENT WAS NOT HARMED IN THIS CASE, BUT IT WAS A NEAR-MISS. | 1
2130193 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT HARM REPORTED. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION. FIELD SERVICE TECH INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
2131040 | IT WAS REPORTED THAT AN ERROR OCCURRED AND THE "OUTPUT" FILE WAS NOT CREATED. THIS ERROR HAS OCCURRED AFTER THE RECENT DEVICE/SOFTWARE UPDATE WAS RUN. THE PRODUCT REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.><IT WAS REPORTED THAT AN ERROR OCCURRED AND THE "OUTPUT" FILE WAS NOT CREATED. THIS ERROR HAS OCCURRED AFTER THE RECENT DEVICE/SOFTWARE UPDATE WAS RUN. THE PRODUCT REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2131052 | IT WAS REPORTED THAT THE PACEART SYSTEM DISPLAYED AN ERROR AND "LOCKED UP" WHEN A CLINICIAN WAS TRYING TO ACCESS A PATIENT RECORD. THE CLINICIAN WAS INSTRUCTED ON HOW TO END THE TASK AND WAS ABLE TO CONTINUE WORKING WITH NO FURTHER ERRORS. THE PRODUCT REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2132435 | SIXTY FIVE DAYS AFTER THE BRAND NEW GENERAL ELECTRIC (GE) ENTERPRISE ACCESS PATIENT MONITORING SYSTEM NETWORK WENT DOWN AT 0200 HOURS IN THE MORNING. GE WAS CONTACTED AND THEIR TECHNICIAN DETERMINED THAT THE SYNCHMASTER WAS OUT. THE NETWORK WAS DOWN FOR APPROXIMATELY 5 HOURS. WHEN THE SYSTEM WAS SWITCHED OVER TO THE BACKUP SYNCHMASTER IT ALSO DID NOT WORK. THE PRIMARY DATA CABLE WAS THEN SWITCHED FROM THE PRIMARY SYNCHMASTER TO THE BACKUP SYNCHMASTER, AGAIN NO SYSTEM FUNCTIONALITY. THE HOSPITAL BIOMED PERSONNEL WERE ASKED TO REBOOT ALL 88 ACCESS POINTS AND THIS ALONG WITH THE SWITCH OVER TO THE BACKUP SYNCHMASTER PUT THE TELEMETRY SYSTEM NETWORK BACK ONLINE. WE BELIEVE THERE ARE TWO DESIGN ISSUES (1) IF THE PRIMARY SYNCHMASTER FAILS THERE SHOULD BE AN AUTOMATIC SWITCH TO THE BACK UP SYNCHMASTER WITHOUT NEEDING TO TURN THE PRIMARY "OFF" AND THE BACKUP "ON", AND (2)THERE SHOULD NOT BE ANY NEED TO "RE-BOOT". THIS FAILURE AFFECTED OVER 250 OF POTENTIALLY 300 PATIENTS AS THE ENTIRE TELEMETRY, WIRELESS MONITOR NETWORK WHEN IT WENT DOWN.MANUFACTURER RESPONSE FOR NETWORK, WIRELESS, PHYSIOLOGICAL MONITOR, ENTERPRISE ACCESS: REPRESENTATIVES OF GE SWITCHED OUT THE PRIMARY SYNCHMASTER FOR A NEW SYNCHMASTER. THEY ARE ALSO REVIEWING THE ENTIRE SYSTEM. | 1
2134696 | MONITOR ROOM REQUESTED CHANGE OF BATTERIES DUE TO LACK OF READING FROM THE TELEMETRY TRANSMITTER BOX. THE TELEMETRY CASE WAS VERY HOT TO TOUCH AND THE BATTERIES WERE EXTREMELY HOT TO TOUCH. THE WRAPPING ON THE BATTERIES MELTED AROUND THE SEAMS. THERE WAS A FAINT ODOR. THE TELEMETRY TRANSMITTER WAS REMOVED FROM THE PATIENT APPROXIMATELY AND 5 MINUTES LATER, THE BATTERIES WERE STILL HOT TO THE TOUCH AND THE CASING WAS SPLIT. | 1
2136552 | A SYSMEX XT-2000I ANALYZER USED OUTSIDE THE UNITED STATES, (B)(6) REPORTED A HEMOGLOBIN (HGB) VALUE THAT WAS 62 G/L AND MCHC=223 G/L. THE PT WAS TRANSFUSED WITH 2 UNITS OF BLOOD ON (B)(6) 2008. NO ADVERSE EFFECT TO THE PT WAS REPORTED DUE TO THE TRANSFUSION. THE RESULT WAS FLAGGED BY THE ANALYZER WITH A DELTA CHECK MESSAGE AND "POSITIVE", INDICATING THE SPECIMEN RESULTS REQUIRED REVIEW. THE USER-DEFINED ANALYZER FLAG TO DETECT LOW MCHC WAS NOT USED, AND THE FLAG FOR LOW HGB VALUES WAS SET TO DETECT VALUES BELOW 40 G/L. THE SAME SPECIMEN WAS ANALYZED ON THE SAME ANALYZER 13 MINUTES LATER WITH THE SAME HGB VALUE 62 G/L. FLAGS AND MCHC=227 G/L. A SECOND SPECIMEN WAS COLLECTED, AND ANALYZED LESS THAN 2 HOURS LATER. THE HEMOGLOBIN VALUE WAS 72 G/L, AND MCHC=256 G/L WITH POSITIVE FLAGS. THE RBC AND HEMATOCRIT (HCT) VALUES DID NOT AGREE WITH THE HGB VALUES OBTAINED ON THE ORIGINAL, THE REPEAT, NOR THE SECOND SPECIMEN. THE LAB BECAME SUSPICIOUS OF SPECIMEN WITH LOW MCHC'S AND REPEATED THE ORIGINAL SPECIMEN 11 HOURS LATER: HGB=98 G/L, MCHC=345 G/L AND THE SECOND SPECIMEN REPEATED 10 HOURS LATER: HGB=100 G/L, MCHC 355 G/L. THE INCIDENT WAS REPORTED TO SYSMEX AS INTERMITTENT LOW HGB AND MCHC VALUES. REVIEW OF PT SAMPLE DATA ANALYZED BETWEEN 05:33 AND 07:27 (B)(6) 2008, SHOWED A PATTERN OF SPECIMENS WITH LOW HGB AND MCHC VALUES, NO OTHER PARAMETERS WERE SUSPECT. THE OPERATOR RETROSPECTIVELY REPEATED 17 SUSPECT SAMPLES AFTER THE TRANSFUSION EVENT. THE FACILITY REPORTED ASPIRATING A CLOTTED SPECIMEN ON THE DAY OF THE EVENT (B)(6) 2008, AND ANALYZER MAINTENANCE AND CLEANING WAS PERFORMED. THE QC AND HGB RESULTS WERE SAID TO BE OK. THE LAB INFO SYSTEM (LIS) "DECRYPTIC" ALERT WAS SET FOR HGB LESS THAN 65 G/L, BUT NO ALERT WAS USED FOR LOW MCHC VALUES. LIS DELTA CHECK FOR HGB WAS SET TO DETECT 30 G/L HGB DECREASE. AS OF (B)(6) 2008, AN ALERT WAS SET FOR MCHC LESS THAN 295 G/L. ANALYZER QC IS VERIFIED ONLY ONCE PER 24 HOURS. LACK OF RECOGNITION OF A PATTERN OF INACCURATE RESULTS AND MINIMAL USE OF USER DEFINED DETECTION SETTINGS BOTH ON THEIR SYSMEX ANALYZER AND THEIR DECRYPTIC LAB INFO SYSTEM APPEAR TO HAVE CONTRIBUTED TO THIS INCIDENT. LIS ALERTS AND ANALYZER FLAGS HELP PREVENT REPORTING OF INVALID RESULTS AND ARE CONSIDERED AN IMPORTANT PART OF GOOD LABORATORY PRACTICE. SYSMEX SERVICE PERSONNEL EVALUATED THE DEVICE AND COULD NOT PRODUCE LOW HGB VALUES. SERVICE ON THE HGB AND SAMPLE ASPIRATION MODULE WAS PERFORMED. THE PERFORMANCE OF ANALYZER WAS VERIFIED WITH PRECISION AND CALIBRATION. QC DATA AND ERROR MESSAGES THAT MAY HAVE OCCURRED AT THE TIME OF THE INCIDENT WERE NOT AVAILABLE. THE CAUSE OF THE LOW HGB VALUES COULD NOT BE DETERMINED. ADDITIONALLY, REVIEW OF THE SULFOLYSER HEMOGLOBIN REAGENT DEVICE HISTORY RECORD FOR THE LOT ON THE ANALYZER DURING THE INCIDENT DID NOT REVEAL ANY ABNORMALITIES. | 0
2136774 | LEFT VENTRICULAR ASSIST DEVICE (LVAD) PATIENTS ARE HOOKED UP TO TELEMETRY, BUT THE SYSTEM WE HAVE SHOWS INTERFERENCE ON THE EKG MONITORS, THEREFORE, THERE IS NO WAY TO KNOW WHAT THE PATIENT'S HEART IS REALLY DOING FOR THE MONITORING TECHS. THIS PROBLEM IS HAPPENING WITH ALL THE LVAD PATIENTS IN THE HOSPITAL. | 1
2140423 | AN OBSTETRIC PT PRESENTED ON (B)(6) 2011, AT A GESTATIONAL AGE OF 38+6 WEEKS. DURING ROUTINE MEASUREMENTS, THE SONOGRAPHER CALCULATED AN ESTIMATED FETAL WEIGHT (EFW) OF 7 LB 13 OZ, THE BIPARIETAL DIAMETER (BPD) SHOWING UP AT THE 18%, THE HEAD CIRCUMFERENCE (HC) BELOW 2.3%, ABDOMINAL CIRCUMFERENCE (AC) AT 95%, AND FEMUR LENGTH (FL) AT 12%. THE PT RETURNED ON (B)(6) 2011 AND AN INTRAUTERINE DEATH WAS DETECTED. THE FETUS WAS DELIVERED AND THE WEIGHT WAS APPROX 10 LB. THE HEAD CIRCUMFERENCE ALSO APPEARED LARGE AFTER DELIVERY ALTHOUGH AN ACTUAL MEASUREMENT WAS NOT AVAILABLE. THE DOCTOR INDICATED THAT IF HE HAD KNOWN ABOUT THE TRUE WEIGHT OF THE FETUS, HE WOULD HAVE TAKEN THE MOTHER'S BLOOD GLUCOSE TO MONITOR HER TREATMENT FOR DIABETES. DATE OF INTRAUTERINE DEATH IS UNDETERMINED, BUT WAS DETECTED (B)(4) 2011. THE DEVICE WAS EVALUATED BY GE HEALTHCARE ENGINEERING AND THE SYS WAS DETERMINED TO BE OPERATING PER SPECIFICATIONS. THE PT'S DATA FILE FROM (B)(4) 2011 WAS AVAILABLE AND ANALYZED. THE MEASUREMENT FUNCTIONALITY OF THE ULTRASOUND SYS WAS ANALYZED AND THE RESULTS WERE VERIFIED AGAINST THE SCALE ON THE SIDE OF THE SCREEN. NO ABNORMALITY WAS DETECTED. THE DEVICE'S CALCULATIONS AND PERCENTILES WERE VERIFIED WITH PUBLISHED GROWTH CHARTS. NO ABNORMALITY WAS DETECTED. THE ESTIMATED FETAL WEIGHT CALCULATION WAS VERIFIED WITH THE PUBLISHED CHART AND AGAIN NO ABNORMALITY WAS DETECTED. BASED ON THE EVAL OF THE SYS AND THE RE-ANALYSIS OF THE (B)(4) 2011 DATA, THE SYS PERFORMED PER SPECIFICATIONS. THE CAUSE OF THE INTRAUTERINE DEATH IS UNDETERMINED OR HAS NOT BEEN MADE AVAILABLE TO GE HEALTHCARE. NO FURTHER ACTIONS ARE PLANNED AT THIS TIME. | 1
2140502 | UPON VALIDATION OF ALLSCRIPTS PRO EHR VERSION 9.2, IT WAS FOUND THAT ONCE LAB RESULTS ARE AMENDED, THE ORIGINAL REPORTED LAB RESULT IS NO LONGER RETRIEVABLE IN THE SYSTEM/CHART. THIS WORKED PROPERLY IN EARLIER VERSIONS AND WAS PROMPTLY REPORTED TO ALLSCRIPTS CLIENT SUPPORT. ALLSCRIPTS HAS NOT CORRECTED THE PROBLEM AFTER SEVERAL ATTEMPTS TO REPORT THE PROBLEM AND OBTAIN CORRECTIVE ACTION DOCUMENTATION. NOTHING HAS BEEN PROVIDED BY ALLSCRIPTS AS A REMEDY TO THIS PROBLEM. FOR LABORATORY COMPLIANCE, ORIGINALLY REPORTED LAB RESULTS MUST REMAIN RETRIEVABLE AND IT MUST BE CLEAR THE RESULT HAS BEEN AMENDED. WE ARE UNABLE TO COMPLY WITH THIS AS LONG AS ALLSCRIPTS PRO EHR VERSION 9.2 IS IN USE. | 1
2141485 | INTELLIVUE TELEMETRY TRANSMITTER BATTERY PACK EXPLODED CAUSING THE BATTERIES TO BURN. THE PATIENT FELT THE PACK GETTING HOT. PATIENT WAS NOT BURNED. | 0
2142240 | THE ROTATOR BROKE DURING A DIAGNOSTIC PROCEDURE OF THE CORONARY ARTERIES WHILE INJECTING CONTRAST. THE CUSTOMER REPORTED THAT THIS OCCURRED FOUR TIMES. THE DEVICES WERE DISCARDED AT THE HOSP. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF FOUR REPORTS FOR THIS COMPLAINT. 1721504-2011-00173, 00176. DEVICE EVAL: THE SUSPECT DEVICE IS NOT EXPECTED TO BE RETURNED FOR EVAL. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE CUSTOMER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. EVAL METHOD: DEVICE HISTORY RECORD WAS REVIEWED. CONCLUSIONS: OTHER, A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
2143042 | FAILURE OF RUBBER MOUNTS SECURING 4" BOX FAN, USED TO COOL THE RAID ARRAY DISK DRIVE, THIS FAILURE ALLOWS THE BOX FAN TO FALL FORWARD, SHORTING THE BOX FAN'S METAL FAN GRATE AGAINST THE RAID ARRAY CONTROLLER PCB. RESULTING IN AN ERROR OF "NO OPERATING SYSTEM FOUND" ON BOOT UP. OUR FACILITY HAS 11 OF THESE SERVES LOCATED IN DATA CLOSETS ACROSS THE CAMPUS. ALL 11 WERE INSTALLED IN 2007. OUR CLINICAL ENGINEERING TEAM IS IN THE PROCESS OF VISUALLY INSPECTING AND SECURING THE RUBBER FAN MOUNTS TO PREVENT ANY FUTURE FAILURES. | 1
2143422 | .><A CALL WAS PLACED TO TECHNICAL SERVICES BY THE CLINICIAN INDICATING THE PATIENT HAD RECEIVED THERAPY AND WHEN ATTEMPTING TO VIEW THE CARELINK REPORT THE PROCESSING WAS SLOW AND AN ERROR WAS RECEIVED STATING THE REPORT CAN NOT BE GENERATED. ADDITIONAL INFORMATION RECEIVED INDICATED THAT CARELINK TRANSMISSIONS WERE SUCCESSFULLY RECEIVED BUT COULD NOT BE PROCESSED DUE TO A POWER OUTAGE EFFECTING ONE OF THE DATABASE SERVERS. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT. | 1
2143982 | ECG AND BLOOD PRESSURE WAVEFORMS DID NOT DISPLAY ON THE PHYSICIAN'S REAL-TIME MONITOR OF THE DEVICE FOR 20 SECONDS. THERE WAS NO ADVERSE EVENT OR ADVERSE CONSEQUENCE TO THE PATIENT. THE SITUATION WAS NOTICED BY THE INITIAL REPORTER. THREE CLINICAL SITES WERE DETERMINED TO HAVE THE AFFECTED SOFTWARE BUILD. SUBSEQUENT INVESTIGATION DETERMINED THAT THE PROBLEM OCCURRED AFTER EIGHT HOURS OF CONTINUOUS OPERATION OF THE REAL-TIME MONITOR. ON (B)(6) 2006, ALL AFFECTED CLINICAL SITES WERE NOTIFIED TO RESTART THE REAL-TIME MONITOR ONCE BETWEEN THE DAY'S PROCEDURES AND TO TURN OFF THE REAL-TIME MONITOR AT THE END OF THE DAY'S PROCEDURES. ON (B)(6) 2006, THE SYSTEMS AT ALL THE AFFECTED CLINICAL SITES WERE UPDATED TO SOFTWARE BUILD 3.1.1017 WHICH PREVENTS THE PROBLEM FROM OCCURRING. | 1
2146020 | IT WAS REPORTED THAT A USER SMELLED SMOKE COMING FROM THE BACK OF THE UNIT DUE TO A PRESUMED FAILURE OF THE POWER SUPPLY. THERE WERE NO REPORTS OF INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. INITIAL REPORTER'S OCCUPATION NOT KNOWN AT THIS TIME. | 1
2147429 | THE PATIENT WAS POST CARDIORESPIRATORY ARREST AND THE PHYSICIAN OPTED TO INITIATE INDUCED HYPOTHERMIA WITH AN INTERNAL COOLING, COOLGARD 3000, SYSTEM. THE PHYSICIAN PLACED THE COOLING CATHETER USING PROPER TECHNIQUE AND ULTRASOUND GUIDANCE. AN 18 GAUGE NEEDLE WAS PLACED IN THE LEFT GROIN FEMORAL VEIN. A GUIDEWIRE PASSED EASILY THROUGH THE CATHETER. THE COOLING CATHETER WAS THEN INSERTED INTO THE VESSEL OVER THE GUIDEWIRE. THE CATHETER WAS SUTURED IN PLACE AT 20 CM. ALL PORTS FLUSHED EASILY EXCEPT FOR THE BROWN-COLORED PORT WHICH WAS SOMEWHAT POSITIONAL. THE COOLGARD 3000 WAS POWERED UP AND INITIATED PER OUTLINED PROCEDURE AND APPROXIMATELY TWO HOURS AFTER THE COOLING PROCESS HAD BEGUN, THE NURSE NOTED THAT THE PATIENT'S TEMPERATURE REMAINED AT 34.6 C. THE PATIENT'S ADMISSION TEMPERATURE (FOUR HOURS PRIOR) WAS 34.7 C. THE NURSE NOTIFIED THE PHYSICIAN. DESPITE THE TUBING FROM THE COOLGARD UNIT TO THE CATHETER FEELING COLD, THE PATIENT'S TEMPERATURE DID NOT DROP. THE PATIENT WAS NOT COOLING AS IS NORMAL. THE COOLGARD UNIT APPEARED TO BE OPERATING FINE. THE PHYSICIAN REMOVED THE COOLING CATHETER. THE PLASTIC ON THE CATHETER WHICH NORMALLY IS EVIDENT TO HAVE HAD FLUID GOING THROUGH BY "WRINKLING" WAS NOT NOTICED UPON REMOVAL.THE PHYSICIAN INSERTED A NEW COOLING CATHETER AND USED THE RIGHT FEMORAL VEIN INSTEAD OF LEFT. THE CATHETER INSERTION WENT SMOOTHLY AND THERE WAS GOOD BLOOD RETURN. AGAIN, THE PATIENT'S TEMPERATURE DID NOT DROP AS EXPECTED. THE PHYSICIAN REQUESTED THAT THE PATIENT BE EXTERNALLY COOLED WITH ICE PACKS TO COOL PATIENT TO THE TARGET TEMPERATURE OF 32 C. APPROXIMATELY FIVE HOURS LATER, THE PATIENT'S TEMPERATURE WAS AT 32 C AND MAINTAINED WITH EXTERNAL ICE PACKS. IT IS THE IMPRESSION OF THE PHYSICIAN AND NURSING THAT THE COOLING CATHETERS DID NOT PERFORM AS INTENDED. | 0
2154693 | PHILIPS BEDSIDE PATIENT MONITORS IN CICU HAVE FALSE ASYSTOLE ALARMS AT A MUCH MORE FREQUENT RATE SINCE THE INSTALLATION OF WIRELESS LAPTOPS FOR EPIC. THE MONITOR WILL DISPLAY A NORMAL ECG AND THEN SUDDENLY SHOW A FLAT LINE FOR ECG AND WITHIN 30 SECONDS ALARM ASYSTOLE. THE ECG CONTINUES WITH A FLAT LINE BETWEEN 1 TO 5 MINUTES AND THEN SUDDENLY RETURNS. ALL OTHER MONITORED PARAMETERS (INVASIVE BLOOD PRESSURE, PULSE OXIMETRY, TEMPERATURE) ARE NOT AFFECTED AND CONTINUE AS NORMAL DURING THE FALSE ASYSTOLE ALARM. INFORMATION SERVICES HAS ADDED 6 WIRELESS LAPTOPS IN EACH CICU POD AND MOST PODS HAVE THE ADDITION OF A WIRELESS ACCESS POINT ON TOP OF THE MEDICATION OR SUPPLY DISPENSER. POSSIBLE RF INTERFERENCE FROM INCREASED LEVEL OF RADIO FREQUENCY ACTIVITY IN THE CICU PODS. ON (B)(6) 2011, CLINICAL ENGINEERING CONTACTED PHILIPS RELATED TO THE MORE FREQUENT APPEARANCE OF FALSE ASYSTOLE ALARMS. PHILIPS REVIEWED THEIR TECHNICAL INFORMATION AND FOUND THAT THERE IS A FIRMWARE UPGRADE THAT MAY RESOLVE THE PROBLEM. CURRENT FIRMWARE WAS D.02.02, UPGRADED FIRMWARE IS D.02.05. THE ECG FIRMWARE IS PART OF THE INTELLIVUE X2 (MODEL M3002A). CLINICAL ENGINEERING AND PHILIPS INSTALLED THE UPGRADED FIRMWARE IN ALL X2 MODULES FOUND IN CICU. THIS UPGRADE RESOLVED THE IMMEDIATE PROBLEM. NO PATIENT MONITORS HAD FALSE ASYSTOLE ALARMS DURING THE NEXT 24 HOURS. PHILIPS INDICATED THAT THE FIRMWARE UPGRADE INCREASED NOISE IMMUNITY OR SUPPRESSION THAT OTHERWISE WOULD RESULT IN A LOSS OF ECG MONITORING. INFORMATION SERVICES WAS ASKED TO PROVIDE INFORMATION ON THE LEVEL OF RADIO FREQUENCY POWER THAT WAS ADDED TO THE SPACE TO SUPPORT THE USERS USE OF EPIC. FOR THE ROOM INDICATED, A TOTAL OF 24 TO 25 DBM (250 TO 300 MILLIWATTS) OF POWER WAS USED IN THE 2.4 GHZ AND 5.4 GHZ ISM SPECTRUMS (802.11 A/B/G (B)(6)). ON (B)(6), INFORMATION SERVICES LOWERED THE TRANSMIT POWER SETTINGS OF THE LAPTOPS IN THE ROOM INDICATED TO THE LOWEST SETTINGS AND TESTED THEIR WIRELESS CONNECTIVITY. THEY TESTED BEFORE AND AFTER CHANGING THE POWER AND THEY OBSERVED A 10% INCREASE IN THE OUTBOUND (FROM LAPTOP) CONNECTION SPEED (ABOUT 2 MB PER SECOND), A WONDERFUL RESULT; LOWER POWER AND BETTER SPEED. | 1
2155728 | COMPLAINT OF NOT BEING ABLE TO CONTROL THE NIBP FUNCTION USING THE TOUCH SCREEN OR TRIM KNOB. RESEATED THE PATIENT DATA MODULE (PDM), REBOOTED CARESCAPE B850, AND REBOOTED MONITOR PROBLEM STILL EXIST. DISCONNECTED AND RECONNECT THE DIGITAL VISUAL INTERFACE (DVI) CABLE TO MONITOR PROBLEM STILL EXIST. PLACED CALL TO TECH SUPPORT UNDER REQUEST NUMBER (B) (4). WAS TOLD TO DISCONNECT THE USB TO THE MONITOR THEN SHUT DOWN, AFTER COMPLETE SHUT DOWN RECONNECT USB TO THE MONITOR, FIRST TURN ON THE MONITOR, THEN THE CARESCAPE B850. MONITOR TOUCH SCREEN AND TRIM KNOB WORKING UP TO SPECIFICATIONS AGAIN. THE USB CABLE DOES HAVE A STRAIN RELIEF BUT MAYBE A DIFFERENT KIND WILL HELP PREVENT THIS PROBLEM FROM OCCURRING AGAIN.MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, CARESCAPE (PER SITE REPORTER): THEY SAID (GE) IT WAS A COMMON PROBLEM THEY KNEW EXACTLY WHAT TO DO. THEY SAID WE SHOULD HAVE THE STRAIN RELIEF THAT ACCOMPANIED THE UNIT INSTALLED ON IT. I TOLD THEM IT WAS INSTALLED BY THE GE INSTALLER. | 1
2156227 | THE HORIZON PERINATAL CARE (HPC) PERINATAL INFORMATION SYSTEM SOFTWARE SENDS AUDIBLE ALARMS THROUGHOUT THE L&D DEPARTMENT WHEN EITHER THE ADT STATUS CHANGES (AN ALERT) OR IF THE FETAL MONITOR CONNECTED VIA THE LMS SOFTWARE HAS A FETAL ALARM. BOTH ALARMS ARE IDENTICAL, IN VOLUME, FREQUENCY, AND TONE. THE USER CANNOT DISTINGUISH WHICH IS THE HIGH PRIORITY (LMS) VS LOW (ALERT). THE ALERTS HAPPEN OFTEN BECAUSE THE SOFTWARE WAS NOT DESIGNED TO HANDLE INTERNAL ADMISSION CHANGES FROM OUTPATIENT TO INPATIENT WITHIN THE SAME DEPARTMENT. THESE CHANGES CREATE THE ADT ALERT AND HAPPEN OFTEN EACH DAY. THERE HAS BEEN AT LEAST ONE INCIDENT WHERE THERE WAS A DELAY TO RESPONDING TO A ACCELERATED FETAL HEART RATE BECAUSE OF THIS ALARM FATIGUE SCENARIO.======================MANUFACTURER RESPONSE FOR PERINATAL INFORMATION SYSTEM SOFTWARE, HORIZON PERINATAL CARE (HPC) (PER SITE REPORTER)======================MCKESSON IS REVIEWING THE WORKFLOW TO SEE WHAT SYSTEM CHANGES ARE AVAILABLE TO ADDRESS THE ALARMS. | 1
2156244 | THE FETAL MONITOR WAS CONNECTED TO THE WRONG DATA PORT ON THE WALL. PATIENT #1 WAS CONNECTED TO PATIENT #2'S DATA PORT. PATIENT #1'S PHYSIOLOGICAL DATA WAS CAPTURED IN PATIENT #2'S ELECTRONIC RECORD. PATIENT #2'S PHYSIOLOGICAL DATA WAS NOT CAPTURED AT ALL.======================HEALTH PROFESSIONAL'S IMPRESSION======================HUMAN FACTOR. THE PATCH CABLES WAS PLUGGED INTO THE WRONG DATA PORT ON THE WALL. | 1
2160255 | THE ROTATOR BROKE DURING USE. NO FURTHER INFORMATION WAS PROVIDED BY THE CUSTOMER. NO HARM OR INJURY WAS REPORTED. THE CUSTOMER REPORTED TWO DEFECTIVE DEVICES BUT DID NOT PROVIDE ANY FURTHER INFORMATION OR CLINICAL DETAILS FOR THE ADDITIONAL EVENT. THEREFORE, THIS SINGLE FORM FDA 3500A WILL BE SUBMITTED FOR THIS COMPLAINT. DEVICE EVALUATION: TWO USED SUSPECT DEVICE WERE RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. THE CUSTOMER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. THE RETURNED DEVICES WERE FROM THE FOLLOWING LOT NUMBERS: F664336, EXPIRATION DATE: 12/31/2011, DEVICE MANUFACTURE DATE: 12/2008. F636698, EXPIRATION DATE: 08/31/2011, DEVICE MANUFACTURE DATE: 08/2008. THE COMPLAINT DATABASE WAS REVIEWED FOR THE TWO LOT NUMBERS REPORTED ABOVE. NO SIMILAR COMPLAINTS WERE FOUND. THE ROOT CAUSES OF THE FAILURE ARE ATTRIBUTED TO THE MANUFACTURING BONDING PROCESS. CORRECTIVE ACTIONS HAVE BEEN IMPLEMENTED TO ADDRESS THIS TYPE OF FAILURE. BOTH LOTS ABOVE WERE BUILT PRIOR TO IMPLEMENTATION OF THE NOTED CORRECTIVE ACTIONS. EVALUATION METHOD: DEVICE HISTORY RECORD WAS REVIEWED, COMPLAINT DATABASE WAS REVIEWED. | 0
2160989 | WE HAD A PT THAT HAD A NEUROSCIENCES CAMINO COMPLETE MICRO VENTRICULAR BOLT PRESSURE MONITORING AND DRAINAGE KIT THAT WAS NOT WORKING PROPERLY AND NEEDED TO BE REMOVED. DURING THE REMOVAL, THE END PIECE OF THE BOLT BROKE OFF AND A PORTION REMAINED LODGED IN THE PT'S SKULL. THE SURGEON MADE A CONSCIOUS DECISION TO NOT REMOVE THE RETAINED PORTION OF THE BOLT BUT RATHER, PLACED A NEW KIT. THE SURGEON FELT THE BOLT THAT WAS REMOVED MAY HAVE BEEN DEFECTIVE. | 0
2162089 | IT WAS REPORTED AFTER AN ABLATION PROCEDURE UTILIZING THE ENSITE VELOCITY SYSTEM FOR MAPPING THE CARDIAC MODEL, THE PATIENT DEVELOPED A PERICARDIAL EFFUSION. THE PHYSICIAN WAS ABLATING ON THE ROOF OF THE LEFT ATRIUM AND THE MITRAL ISTHMUS LINE WHEN IT WAS NOTED THE ABLATION CATHETER APPEARED TO BE 1 CM OR MORE OUTSIDE OF THE CARDIAC MODEL. IT WAS DIFFICULT TO DETERMINE THE POSITION OF THE ABLATION CATHETER. THE PHYSICIAN CONTINUED TO ABLATE USING FLUOROSCOPY TO CONFIRM THE PLACEMENT OF CATHETERS. THE ABLATION PROCEDURE HAD BEEN FINISHED FOR APPROXIMATELY 10 MINUTES AND ALL DEVICES USED DURING THE PROCEDURE HAD BEEN REMOVED WHEN THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EFFUSION WAS NOTED. THE PHYSICIAN TREATED THE CONDITION BY DRAINING THE EFFUSION 300 CC OF THE BLOOD FROM THE PERICARDIUM. AFTER THE EFFUSION, THE (B)(4) WAS REVIEWED ON THE ENSITE VELOCITY SYSTEM WHERE IT WAS DISCOVERED, THERE WERE TWO PERIODS DURING THE PROCEDURE WHERE THE CORONARY SINUS CATHETER HAD MOVED FROM THE ORIGINAL POSITION MARKED BY A CATHETER SHADOW PLACED ON THE ENSITE VELOCITY SYSTEM DISPLAY. THE PHYSICIAN HAD REISOLATED THE PULMONARY VEINS AND PERFORMED A RIGHT ISTHMUS ABLATION DURING THE PERIODS OF TIME WHERE THERE WAS POSITIONAL REFERENCE CATHETER MOVEMENT. THE CATHETER MOVEMENT WAS NOT DETECTED BY THE SYSTEM OPERATOR DURING THE PROCEDURE, AND THE POSITIONAL REFERENCE TOOL WAS NOT USED TO DETECT THE MOVEMENT. THE PATIENT RECOVERED FROM THE EFFUSION WITH NO FURTHER CONSEQUENCES AND WAS DISCHARGED FROM THE HOSPITAL. THE INVESTIGATION MATERIALS HAVE NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED. | 1
2163545 | SINCE THE INITIATION OF CPOE MECHANIZED AUTOMATION SYSTEMS TO ORDER BLOOD TESTS, THE INCIDENCE OF "NO DRAWS", RESULTING IN DELAYS IN DIAGNOSIS AND THERAPY HAS INCREASED. RECONCILIATION MECHANISMS DO NOT EXIST TO ASSURE THAT 100% OF THE BLOOD, URINE, AND OTHER TESTS THAT WERE ORDERED WERE COMPLETED OR ADDRESSED. THE PHLEBOTOMISTS DEPEND ON COMPUTERIZED LABEL MAKERS AS THE TRIGGER AND NOTIFICATION TO DRAW BLOOD AND OBTAIN SPECIMEN. THERE ARE OTHER NOTIFICATIONS THAT BLOOD TESTS ARE NEEDED DEPENDING ON THE TIME OF DAY. THERE ARE SEVERAL OPPORTUNITIES IN THIS COMPLEX COMPUTER ORDERED AND CPOE DEPENDENT WORKFLOW FOR ERRORS TO BE FACILITATED, INCLUDING THE FAILURE OF THE INTERFACE BETWEEN CPOE AND THE LABEL MAKER. UNEXECUTED ORDERS FREQUENTLY REMAIN ON THE CPOE SCREEN FOR DAYS WITHOUT NOTIFICATION OF OPEN ORDERS. DELAYS IN DIAGNOSIS AND THERAPY RESULT FROM THIS DEFECT IN THE ORDERING SYSTEMS. CRITICAL DIAGNOSIS WAS DELAYED FOR 48 HOURS IN THIS PT WITH ABNORMAL LIVER TESTS AND MULTI SYSTEM FAILURE. WHEN IT CAME TO ATTENTION BY SERENDIPITY THAT THERE WERE NOT ANY RESULTS FOR THE TESTS THAT WERE ORDERED, AN ATTEMPT WAS MADE USING THE CPOE TO GENERATE A COMPUTERIZED LABEL AND REQUISITION. THE SYSTEMS FAILED TO ACCOMPLISH THIS. THE REQUISITION WAS WRITTEN BY LONG HAND ON PAPER. THE ANCILLARY STAFF INVOLVED WITH THIS BECOME EXASPERATED SINCE IT HAS BEEN AN ONGOING PROBLEM THAT HAS NOT BEEN FIXED. | 1
2163546 | IN OUR FACILITY, THE CPOE PALATE OF THE UNIVERSE OF ORDERS ENABLE THE ORDERING OF CONCENTRATIONS OF DRUGS IN PRE-MIXED BAGS. THUS, INSTEAD OF SIMPLY ORDERING A FIXED DOSE OF DOPAMINE AT E.G. 5 MCG/KG/MINUTE WITH THE EQUIPMENT POPULATING THE WEIGHT FIELD, MULTIPLE CLICKS ARE REQUIRED TO SELECT THE DESIRED CONCENTRATION OF THE PRE-MIXED BAG, AND THEN ENTER THE DOSING DRIP RATE AND WEIGHT -SIC-. IF THERE IS A NEED TO USE THIS DRUG IN SYNERGY WITH ANOTHER, THE DESCRIPTIVE INDIVIDUALIZATION OF DIRECTIONS IS NOT FEASIBLE. IN OUR CASE OF ACUTE COMPLEX CLASS 4 HEART FAILURE WITH HYPOTENSION, A CONCENTRATED PRE-MIX OPTION WAS SELECTED DUE TO USER (A PHYSICIAN) CONFUSION, INEXPERIENCE, OR MISS-CLICK. THE ORDER WAS NOT EXECUTED UNTIL "CLARIFICATION" (THE ORIGINAL ORDER WAS AS CLEAR AS THE CPOE TRANSMITTED IT), WHICH DELAYED THE INITIATION OF THE THERAPY BY 3.25 HOURS. THIS DELAY RENDERED THE HEART FAILURE MORE DIFFICULT TO TREAT DUE TO ADVERSE BODY CHEMISTRY CHANGES THAT PROGRESSED DURING THE DELAY, THREATENING THE PT'S LIFE. THE ORDERING PROGRAMS ON THIS EQUIPMENT IS DANGEROUS, USER-UNFRIENDLY, ERROR PRONE, AND IS AN IMPEDIMENT TO INDIVIDUALIZING MEDICAL CARE. | 1
2170198 | PATIENT RETURNED FOR MAMMOGRAM, HOWEVER, THE RIGHT CRANIOCAUDAL IMAGE FROM TWO YEARS EARLIER IS UNABLE TO BE ELECTRONICALLY RETRIEVED FOR COMPARISON.ONE OF THE PATIENTS MAMMOGRAM IMAGES FAILED TO STORE CORRECTLY IN OUR PACS SYSTEM. IT IS UNRETRIEVEABLE. | 1
2170200 | PATIENT HAD A MAMMOGRAM IN LATE 2010. IMAGES WERE BEING REVIEWED PRIOR TO HER BREAST SURGERY 3 MONTHS LATER, AND IT WAS NOTICED THAT TWO OF THE IMAGES FROM HER 2010 STUDY DID NOT SAVE CORRECTLY IN OUR PACS SYSTEM. | 1
2170201 | IT WAS BROUGHT TO OUR ATTENTION THAT FOUR OF THIS PATIENT'S MAMMOGRAM IMAGES, TAKEN IN LATE 2010 MAY HAVE FAILED TO BE SAVED IN THE PACS SYSTEM. WE HAVE CONFIRMED THAT THEY ARE NOT IN PACS AND ARE NO LONGER AVAILABLE ON THE ACTUAL MAMMOGRAM UNIT TO RE-SEND AND STORE. | 1
2170203 | WE RECEIVED A REQUEST FOR THIS PATIENT'S PRIOR MAMMOGRAMS TO BE SENT TO ANOTHER FACILITY. UPON REVIEW OF HER IMAGES TO PREPARE TO PRINT THEM, IT WAS DISCOVERED THAT ONE OF HER IMAGES FAILED TO SAVE INTO OUR ELECTRONIC STORAGE SYSTEM. IT IS UNRECOVERABLE. | 1
2170205 | TWO IMAGES FROM PATIENT'S MAMMOGRAM IN EARLY 2010 DID NOT SAVE CORRECTLY IN THE ELECTRONIC STORAGE SYSTEM (PACS). | 1
2170206 | AFTER THE PATIENT'S MAMMOGRAM THIS MORNING, IT WAS DISCOVERED THAT TWO IMAGES FROM HER EXAM ONE YEAR PREVIOUSLY DID NOT SAVE CORRECTLY IN OUR ELECTRONIC STORAGE SYSTEM (PACS) AND THEY ARE UNRECOVERABLE. | 1
2170240 | PATIENT HAD MAMMOGRAM IN THE SPRING OF 2011. UPON REVIEW OF HER PRIOR STUDIES IT WAS NOTICED THAT THE MAMMOGRAM IMAGES FROM 3 YEARS PREVIOUSLY DID NOT SAVE CORRECTLY IN OUR ELECTRONIC STORAGE SYSTEM (PACS) AND THESE IMAGES ARE UNRECOVERABLE. | 1
2170242 | PATIENT HAD MAMMOGRAM IN SPRING OF 2011. UPON REVIEW OF PRIOR IMAGES IT WAS NOTICED THAT THREE OF THE FOUR IMAGES FROM 3 YEARS PREVIOUSLY DID NOT SAVE CORRECTLY IN OUR PACS SYSTEM. | 1
2171278 | PACS APPLICATION IS ERRONEOUSLY CHANGING THE EXAM STATUS FROM PENDING TO READ W/O A RADIOLOGIST HAVING RENDERED A REPORT. THIS ACTION TAKES THE STUDY OFF OF THE ACTIVE WORKLIST W/O THE USER BEING AWARE THAT IT STILL SHOULD BE IN THE ACTIVE QUEUE. THIS RESULTS IN SIGNIFICANT DELAYS IN INTERPRETATION, WHICH IS CRITICAL FOR STAT EXAMS. | 1
2171423 | AN IMAGING DYNAMICS COMPANY (IDC) XPLORER 1600 DIGITAL RADIOGRAPHIC SYSTEM WAS IN USE AT THE DIAGNOSTIC MEDICAL GROUP OF (B)(6). AT APPROX 2:30 PM, THE SYSTEM WAS BEING RE-POSITIONED FOR A STANDING KNEE EXAM. THE RADIOLOGY TECHNICIAN HEARD A LOUD NOISE AND THE SYSTEM STOPPED WITH AN ERROR CODE (E06). THE RADIOLOGY TECHNICIAN CONTACTED THE IDC AUTHORIZED SERVICE TECHNICIAN AND BASED ON HIS INSTRUCTIONS, THE POWER WAS RESET ASSUMING IT WAS AN ERROR CODE FOR THE GENERATOR. THE E06 ERROR CLEARED WHEN THE POWER WAS RESET. THE RADIOLOGY TECHNICIAN THEN REDUCED THE SOURCE TO IMAGE DISTANCE (SID) AND THE XPLORER 1600 SYSTEM C-ARM SUFFERED AN UNCONTROLLED CRASH TO THE FLOOR. THE SYSTEM WAS PUT "OUT OF ORDER" AND HELP WAS REQUESTED FROM THE SERVICE TECHNICIANS. NO PT WAS INVOLVED IN THIS EVENT. THERE WAS NO INJURY TO ANY HEALTH PROFESSIONAL. THERE WAS NO ACCIDENTAL RADIATION EXPOSURE. (B)(6) SUBMITTED THE ATTACHED MDR TO FDA ON (B)(6) 2006. HOWEVER IMAGING DYNAMICS COMPANY COULD NOT VERIFY THEIR SUBMISSION AND DECIDED TO SEND THIS REPORT AFTER DISCUSSION WITH FDA INSPECTOR WHO INSPECTED IMAGING DYNAMICS COMPANY IN (B)(6) 2006. DUE TO THIS REASON THIS ANCILLARY 3RD PARTY REPORT IS NOT MEETING THE 30-DAY INITIAL REPORTING REQUIREMENTS. A TEMPORARY RADIOLOGY TECHNICIAN WAS IN THE X-RAY ROOM AT THE TIME OF THE PROBLEM. THE X-RAY UNIT WAS IN POSITION FOR A 72" CHEST X-RAY, AND HE WAS STARTING TO BRING IT INTO POSITION TO DO A STANDING KNEE X-RAY. THE MOTION WAS ROUGH, AND HE ENCOUNTERED AN E06 ERROR CODE. ACCORDING TO THE TECHNICIAN, THERE WAS NO OBSTRUCTION TO THE MOVEMENT OF THE C-ARM. HE CALLED ON A MORE EXPERIENCED TECHNICIAN FOR ASSISTANCE. THEY RECYCLED POWER AND THE ERROR CODE CLEARED. AS THEY DROVE THE C-ARM DOWN, THE UNIT CRASHED FROM ABOUT 30 INCHES HEIGHT TO THE FLOOR. THE TECHNICIAN STATED THAT IT SEEMED TO HIM THAT THE UNIT WAS IN "FREE FALL" RATHER THAN A RAPID DRIVE DOWNWARD. IMAGING DYNAMICS CO (IDC) IS REPORTING THIS INCIDENT AS MFR AND SYSTEM INTEGRATOR OF THE X1600 DIGITAL RADIOGRAPHIC SYSTEM (COMPROMISING OF THE IDC XPLORER 1000 DETECTOR HEAD AND SEDECAL'S X PLUS LP RADIOLOGIC STAND). BASED ON THE PRELIMINARY INVESTIGATION BY IDC AND SEDECAL, AN INTERMEDIATE TECHNICAL SERVICE BULLETIN (TSB) WAS ISSUED ON (B)(6) 2006 TO ALL USERS OF XPLORER 1600 UNITS. THE FAULTY UNIT WAS SHIPPED TO (B)(6) FACILITY IN (B)(6) FOR FURTHER INVESTIGATION. THE ROOT CAUSE FOR THE PROBLEM WAS IDENTIFIED AS LOOSE REAR BEARINGS IN THE CARRIAGE ASSEMBLY ON THE STAND. THIS CAN CAUSE LATERAL PLAY OF THE COMPLETE ASSEMBLY AND UNDER SOME CIRCUMSTANCES CAN CAUSE THE EMERGENCY BRAKE NOT TO ACTIVATE WHEN NEEDED. ON (B)(6) 2006 A SECOND TSB WAS ISSUED TO ALL END USERS, DEALERS AND SERVICE TECHNICIANS TO INSPECT ALL UNITS AND ADJUST/REPAIR AS NECESSARY. THE IDC XPLORER 1600 SERVICE MANUAL WAS REVISED TO INCLUDE THIS CHECK AS A PERIODIC MAINTENANCE ACTIVITY. AS REQUESTED BY SEDECAL, A THIRD TECHNICAL SERVICE BULLETIN WAS ISSUED ON (B)(6) 2006 TO ADDRESS THE ERROR E06 FUNCTIONALITY CHECK. IDC'S INTERNAL CORRECTIVE ACTION (B)(4) WILL CONTINUE TO REMAIN OPEN UNTIL EFFECTIVENESS OF THE CORRECTIVE ACTION IS CONFIRMED. SEDECAL SUBMITTED THE ATTACHED MDR TO FDA ON (B)(6) 2006. HOWEVER, IMAGING DYNAMICS CO. COULD NOT VERIFY THEIR SUBMISSION AND DECIDED TO SEND THIS REPORT AFTER DISCUSSION WITH FDA INSPECTOR WHO INSPECTED IMAGING DYNAMICS CO IN (B)(6) 2006. DUE TO THIS REASON THIS ANCILLARY 3RD PARTY REPORT IS NOT MEETING THE 30- DAY INITIAL REPORTING REQUIREMENTS. ADD'L INFO FROM MFR REPORT #: 9617251-2006-00002: NO PT INJURY. IDC/SEDECAL. DIAGNOSTIC X-RAY SYSTEM. (B)(6). MODEL - XPLORER 1600. (B)(6). EVAL SUMMARY: THE PRELIMINARY INVESTIGATION REVEALED EXCESS "PLAY" IN ONE OF THE SUPPORT BEARINGS. THE UNIT HAS BEEN SHIPPED BACK TO THE MFR IN (B)(6) FOR FURTHER EVAL. AN UPDATE WILL BE PROVIDED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
2171513 | WHEN THERE IS AN INTERRUPTION TO THE HORIZON PERINATAL CARE (HPC) PERINATAL INFORMATION SYSTEM'S OPERATION, AND IF IT IS LONG ENOUGH THAT THE FIRST PATIENT WAS CONNECTED TO A FETAL MONITOR AND IS DISCHARGED DURING THE INTERRUPTION, THE NEXT PATIENT'S FETAL MONITOR INFORMATION WILL BE STORED ON THE PAST PATIENT'S RECORD.======================MANUFACTURER RESPONSE FOR PERINATAL INFORMATION SYSTEM, HORIZON PERINATAL CARE (HPC) (PER SITE REPORTER) ======================FOR UNPLANNED DOWNTIME, GO TO PAPER RECORDING OF FETAL/MATERNAL INFORMATION FROM GE FETAL MONITOR. FOR PLANNED DOWNTIME, CREATE TEMPORARY PATIENT RECORDS IN ALL REMAINING VACANT BEDS TO CONTINUE SURVEILLANCE. NEW PATIENTS NEED TO HAVE FETAL STRIPS TO DOCUMENT ACTIVITY DURING THIS TIME. | 1
2174878 | THE CUSTOMER RECEIVED A FALSE NEGATIVE REACTION WITH THE D MICRO WELL OF THE ABD/REV CARD. NO INCORRECT OR ERRONEOUS RESULTS WERE REPORTED. AS PER SECOND LEVEL CTS SUPPORT, THE PROVUE IS ACTING AS EXPECTED AND NO SERVICE IS NEEDED AT THIS TIME. THIS EVENT WILL BE ENTERED FOR DOCUMENTATION PURPOSES. NO ACTION IS REQUIRED BY CTS AT THIS TIME. NO SERVICE CALLS HAVE BEEN REQUESTED BY THE CUSTOMER. (B)(4). | 0
2175160 | PT GRAPHIC WIRE WAS ADVANCED DOWN THE POSTERIOR TIBIAL ARTERY IN THE RIGHT LEG. THE DISTAL WIRE BROKE OFF AND REMAINS IN THE DISTAL POSTERIOR TIBIAL ARTERY. | 0
2175641 | A CUSTOMER CONTACTED HAEMONETICS ON (B)(6) 2011 STATING THAT THEY WERE NOTIFIED ON AN ALLEGED DONOR EVENT AFTER PLATELETPHERESIS ON THE MCS+ 9000. THE DONOR TOLD THE BLOOD CENTER THAT THE DAY AFTER HER DONATION ON (B)(6) 2011, SHE HAD A SORE ARM AND HAD A RED LINE FROM UNDER HER ARMPIT TO HER WRIST ON THE OPPOSITE ARM TO WHERE THE VENIPUNCTURE HAD BEEN. NO ABNORMALITIES WERE NOTED DURING THE PROCEDURE. INVESTIGATION IS ONGOING. FOLLOW-UP REPORT WILL BE SUBMITTED. | 0
2177020 | OUR CUSTOMER CALLED STATING THAT WHEN USING AN ICT TRANSDUCER AND BIOPSY GUIDE TO PERFORM A TRANS-RECTAL BIOPSY PROCEDURE, PATIENTS EXPERIENCED EXCESSIVE BLEEDING. THE COMPLAINT WAS NOT SPECIFIC TO ONE PATIENT. THE USER STATED IN GENERAL THE EXCESSIVE BLEEDING OCCURRED. NO INTERVENTION WAS REQUIRED. SONOSITE, INC IS THE MANUFACTURER OF THE ICT TRANSDUCER. SONOSITE DOES NOT MANUFACTURE BIOPSY GUIDES WHICH MAY BE USED WITH THE TRANSDUCER. AS THERE WAS NO MALFUNCTION, WE INVESTIGATED THROUGH INTERVIEW. THE DEVICE WAS OPERATING TO SPECIFICATION. THE CUSTOMER'S TRANSDUCER WAS REPLACED WITH A C8 TRANSDUCER WHICH IS DESIGNED SPECIFICALLY FOR TRANS RECTAL PROCEDURE. | 0
2177831 | IT WAS REPORTED THAT A SPONTANEOUS REBOOT OF THE DASH 2000 BEDSIDE MONITOR OCCURRED. THERE WAS NO SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT, NOR WAS MEDICAL INTERVENTION REQUIRED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. THE DEVICE MANUFACTURE DATE IS UNK. | 1
2181932 | THIS X-RAY SYSTEM WILL INTERMITTENTLY EXPORT FLIPPED IMAGES TO THE PICTURE ARCHIVING COMPUTER SYSTEM (PACS). THESE FLIPPED IMAGES REFERS TO IMAGE INVERSION AND CAN BE DESCRIBED AS "MIRRORED" WHERE THE LEFT AND RIGHT ARE REVERSED. INVERTED WOULD APPLY TO THE TOP AND BOTTOM BEING REVERSED. THE RANDOM CHANGES IN THE IMAGE ORIENTATION COULD HAVE BEEN ANY OF THE FOUR POSSIBLE ORIENTATION ERRORS. A REBOOT OF THE SYSTEM WILL TEMPORALLY RESTORE NORMAL OPERATION. (CONCLUSIONS): THE ROOT CAUSE OF THIS PROBLEM IS A SOFTWARE RELATED ISSUE. A FIELD CHANGE ORDER WILL BE RELEASED TO CORRECT THIS SOFTWARE ISSUE. | 1
2182857 | THE CUSTOMER CONTACTED BECKMAN COULTER, INC. (BEC) TO REPORT THAT AN OPERATOR RECEIVED AN EYE SPLASH WHILE TROUBLESHOOTING TO REPLACE TUBING THAT HAD BACKED UP WITH SEDIMENT TO THE WRIGHT STAIN ON THEIR LH SLIDE STAINER. THE OPERATOR ATTEMPTED TO TROUBLESHOOT THE INSTRUMENT WITHOUT THE AID OF A BECKMAN COULTER TECHNICAL SPECIALIST AFTER OBSERVING THAT THE BATHS WERE NOT DRAINING OR FILLING PROPERLY. THE OPERATOR WAS NOT WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) TO PROTECT HIS EYES AT THE TIME OF THE EVENT. THE OPERATOR WAS WEARING CONTACT LENSES AT THE TIME OF THE EVENT. PER SECTION SIX OF THE CLEANING AND REPLACEMENT PROCEDURE OF THE OPERATOR'S GUIDE (BECKMAN COULTER PART NUMBER 4277300DB), OPERATORS ARE ADVISED TO WEAR PPE WHILE PERFORMING TROUBLESHOOTING ACTIVITIES. THE OPERATOR FLUSHED HIS EYES WITH WATER AT AN EYE WASH STATION AND THEN SOUGHT MEDICAL ATTENTION. THE OPTOMETRIST DETERMINED THAT THERE WAS NO SERIOUS INJURY TO THE EYE AND THAT THERE SHOULD BE NO PERMANENT DAMAGE. IT IS UNKNOWN IF THE OPERATOR HAD REVIEWED THE MATERIAL SAFETY DATA SHEET (MSDS). THE FACILITY DOES HAVE AN EXPOSURE CONTROL / RISK MANAGEMENT PLAN IN PLACE. THE OPERATOR, WHILE TROUBLESHOOTING, HAD REPLACED THE WASTE FILTER, TUBING, AND PERISTALTIC PUMP TO BATH TWO. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S SITE. THE FSE VERIFIED THE REPAIRS MADE BY THE OPERATOR PER ESTABLISHED PROCEDURES. THERE WAS NO IMPACT TO PATIENT SAMPLE RESULTS OR PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. TROUBLESHOOTING ACTIVITIES PERFORMED. EYE SPLASH. PPE NOT WORN AT TIME OF EVENT. BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. THE CONTENTS OF THE REAGENT CONTAINERS AND ITS ASSOCIATED TUBING CONTAIN MATERIAL WHICH CAN STAIN SKIN AND MUST BE HANDLED WITH CARE. THE PERISTALTIC TUBING MAY CONTAIN RESIDUAL STAIN AND MUST BE HANDLED WITH CARE. AVOID SKIN CONTACT. CLEAN UP SPILLS IMMEDIATELY IN ACCORDANCE WITH LOCAL REGULATIONS AND ACCEPTABLE LABORATORY PROCEDURES. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, WEARING PROTECTIVE EYE WEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS INSTRUMENT. | 0
2183803 | A CRITICALLY ILL PT WITH MULTIPLE FAILING ORGANS AND ANEMIA WAS FOUND TO HAVE A SPURIOUSLY AND MARKEDLY INCREASED HEMOGLOBIN THAT WAS NOT RECOGNIZED BY THE PT'S HEALTHCARE PROFESSIONALS AS NOT BELONGING TO THIS PT. ADDITIONALLY, THE PLATELET COUNT AND LEUKOCYTE COUNT WERE CONSISTENTLY THOSE OF ANOTHER PT. AS A RESULT, APPROPRIATE CARE FOR THIS PT'S ANEMIA WAS DELAYED DUE TO THE ERRONEOUS RESULT FROM MISIDENTIFICATION. THE PT WHOSE RESULTS THESE MAY HAVE BEEN MISTREATED AS WELL. MISIDENTIFICATIONS HAVE BECOME COMMONPLACE AT OUR INSTITUTION WITH THE ADVENT OF (B)(6) SYSTEMS' BURDEN OF EXTRA STEPS IN THE SPECIMEN ORDERING, LABELING, AND PROCESSING, AND IN DISPLAYING RESULTS. THE FORENSIC DETAILS ARE UNK AND (B)(6) INVESTIGATION. | 1
2185510 | EXCESS FLUID ABSORPTION (RETENTION) DURING HYSTEROSCOPIC MYOMECTOMY. | 0
2188616 | TEAM WAS TRANSFERRING THE PATIENT FROM THE INTERVENTIONAL RADIOLOGY (IR) TABLE TO THE RADIANT WARMER. THE IR RN WAS MOVING THE TRANSPORT MONITOR TO THE PATIENT'S BED. THE TRAM HOUSING (TRAM CHUTE) FELL OFF THE MONITOR (TRANSPORT PRO)AND LANDED ON PATIENT'S ARM AND CHEST. WE HAVE FOUR OF THESE TYPE DEVICES (MONITORS). WE FOUND 1 WITH A MISSING SCREW (THAT LOCKS THE TRAM HOUSING TO THE MONITOR) IN APRIL. THEN WE HAD THIS REPORTED INCIDENT. UPON INSPECTION OF THE REMAINING 2, WE FOUND THAT BOTH HAD LOOSE SCREWS THAT WERE NEAR TO FALLING OUT. THE TRAM HOUSING ON OUR FOUR TRANSPORT PRO MONITORS HAVE BEEN SECURED WITH A SCREW AND LOCKING NUT UNTIL GE IS NOTIFIED AND RESPONDS WITH AN OFFICIAL MODIFICATION. | 0
2188661 | THE NIBP MACHINE FROM A NURSING UNIT WAS NOT WORKING. THE NIBP DEVICE WAS TAGGED AND SENT TO BIOMED. THE BIOMED TECHNICIAN CALLED ME TO ALERT ME TO HIS FINDINGS. HE STATED THAT THE ENTIRE NIBP TUBING WAS FILLED WITH BLOOD. WE COULD NOT IDENTIFY WHEN THIS MACHINE WAS TAKEN OUT OF SERVICE, WHICH PATIENT IT WAS LAST USED FOR AND THE STAFF MEMBER THAT USED IT PRIOR TO THE DYSFUNCTION. THERE IS CONCERN THAT THIS PRODUCT/DEVICE (NIBP HOSE ADAPTER) FITS OUR IV TUBING CONNECTIONS. IT DOES NOT HAVE A UNIQUE FITTING IN ORDER TO PREVENT SOMEONE ACCIDENTALLY ATTACHING IT TO IV TUBING.WHEN THIS EVENT WAS BROUGHT TO MY ATTENTION, I REMOVED AN IV CAP FROM AN IV TUBING INJECTION SITE AND WAS ABLE TO CONNECT THE GE/CRITIKON NIBP MALE SLIP LUER CONNECTOR FROM THE CUFF ADAPTOR TO THE INJECTION SITE.WE DO NOT KNOW IF THIS ONE WAS ACTUALLY EVER CONNECTED TO THE IV TUBING, HOWEVER WE DID DISCOVER THAT IT WAS POSSIBLE, WHICH IS A CONCERN. | 1
2188666 | WHEN THERE WAS FIRST SIGN OF FAILING KIDNEYS IN A CRITICALLY ILL PATIENT, A URINE EVALUATION FOR MICROSCOPIC AND CULTURE AND SENSITIVITY WAS ORDERED BY ELECTRONIC CLICKING OF THE APPROPRIATE ORDER SET SELECTION TO ACCOMPLISH THE TASK. THREE DAYS LATER, THERE WERE NOT ANY RESULTS. WE LEARNED THAT THE ORDER NEVER REACHED THE INTENDED HEALTH CARE TEAM MEMBER WHO WAS CHARGED WITH THE RESPONSIBILITY OF COLLECTING THE URINE FROM A FOLEY CATHETER SYSTEM. IF IT HAD REACHED THIS SECTION OF THE CARE TEAM, THE SCREEN OF THE ORDER WAS MISSED OR APPEARED MEANINGFULLY MISLEADING TO A COMPETENT AND EXPERIENCED HEALTH CARE PROFESSIONAL. THE ELECTRONIC ORDERING SYSTEM OF ALL SUCH ELECTRONIC ORDERING MEDICAL INSTRUMENTS HAS NOT ANY RECONCILIATION MECHANISM TO DETERMINE IF ANY ORDER WAS EXECUTED AND IF THERE WAS A RESULT. THE KIDNEYS CONTINUED TO FAIL, THE LUNGS FILLED WITH FLUID, MENTAL STATUS DETERIORATED TO DELIRIUM BY THE TIME THE URINE WAS COLLECTED AND SENT TO THE MICROBIOLOGY LAB. THREE DAYS AFTER THAT, AND WITH FURTHER DETERIORATION OF KIDNEYS AND OTHER ORGANS, IT BECAME KNOWN THAT THERE WAS (B)(6) COLONY COUNT OF A VIRULENT BACTERIA WITH SENSITIVITIES PUBLISHED. THE PATIENT HAS UNDERGONE KIDNEY DIALYSIS. THAT THERE WERE (B)(6) COLONIES GROWING IN THE URINE WHICH WAS NOT POSTED ON THE EMR LAB SECTION AFTER TWO DAYS INDICATES FURTHER DYSFUNCTION OF THE ELECTRONIC INSTRUMENTS DIRECTING THE WORK FLOW OF CARE OF THIS CASE. IF IT WAS POSTED, IT WAS MISSED. IF IT WAS POSTED, NO HEALTH CARE TEAM MEMBER RECEIVED ANY INDICATION THAT NEW RESULTS HAD BEEN SENT TO THE EMR. IN A BUSY UNIT WITH CRITICALLY ILL PATIENTS, THERE IS LITTLE TIME TO CONTINUALLY HAVE TO SCROLL THROUGH USER UNFRIENDLY, DISORGANIZED AND ILL DESIGNED GRIDS OF DATA TO FIND WHAT IS NEW. WE BELIEVE THAT THE SYSTEM IS DANGEROUS AND SHOULD NOT BE PERMITTED TO BE USED IN THE CARE OF PATIENTS. | 1
2188692 | THIS IS A CASE OF MISIDENTIFICATION ASSOCIATED WITH THE USER UNFRIENDLY SCREENS LISTING PATIENT'S NAMES, RESULTING IN A CONSULTANT ENTERING A PROGRESS NOTE WITH DATA, AND ORDERS BASED ON THIS DATA THAT HAD COME FROM A DIFFERENT PATIENT. THE MEDICATION, INTENDED FOR ANOTHER PATIENT, WAS ADMINISTERED. THERE WAS NOT ANY HARM. ADDITIONAL TESTS WERE NEEDED FOR SURVEILLANCE. AT LEAST ONE OTHER PHYSICIAN, AND SEVERAL NURSES, INVOLVED WITH THE CARE OF THE SAME PATIENT MISSED THIS ERROR. THE PATIENT WHO WAS TO HAVE GOTTEN THE MEDICATION DID NOT RECEIVE IT. WE ARE NOT AWARE OF ANY ADVERSITY TO THE SECOND PATIENT. | 1
2190742 | WE HAVE BEEN CONCERNED THAT THE EMR AND ITS INCLUSIVE ELECTRONIC ORDERING AND E-MAR - MEDICATION ADMINISTRATION - CREATES INNUMERABLE IMPEDIMENTS TO SAFE AND EFFICIENT CARE AND COMMUNICATION. IN THIS CASE, AN ANESTHESIOLOGIST HAD TROUBLE PLACING AN ENDOTRACHEAL TUBE DURING AN ELECTIVE INTUBATION FOR A PROCEDURE. THUS, THE PT WAS GIVEN A DIAGNOSIS OF "DIFFICULT INTUBATION." THE PT WAS EXTUBATED AFTER THE PROCEDURE, BUT DEVELOPED RESPIRATORY DISTRESS, WHICH SHOULD HAVE BEEN PREDICTABLE BASED ON CHEST RADIOGRAPH WHICH WAS EITHER NOT LINKED TO THE EMR OR WAS NOT ACCESSED AND EVALUATED BECAUSE THERE WAS NOT ANY NOTIFICATION THAT IT WAS READY FOR VIEWING, NOR WAS THERE ANY NOTIFICATION THAT IT HAD BEEN EVALUATED BY A RADIOLOGIST. WHEN THE NEED TO REINTUBATE THE PT BECAME URGENT, THE ACUTE CARE TEAM WAS NOT AWARE OF THE DIAGNOSIS "DIFFICULT INTUBATION", AND WAS NOT PREPARED FOR THE DIFFICULTY. THE EMR WAS TO HAVE PROVIDED THAT CRITICAL DIAGNOSIS AS A HIGHLIGHTED WARNING AT THE TOP OF THE SCREENS. IT FAILED TO DO SO. THERE WERE NO OTHER WARNINGS. IN THE MAYHEM THAT FOLLOWED DUE TO THE DEFECTIVE EMR AND FAILED COMMUNICATION SO ASSOCIATED, AND THE LACK OF PREPAREDNESS, THE PT SUFFERED RESPIRATORY FAILURE, RESPIRATORY AND CARDIAC ARREST, ANOXIC BRAIN INJURY, AND DEATH. | 1
2190787 | IN OUR INSTITUTION, THE CPOE ORDERING SYSTEM ALLOWS A VIEWING OF THE CURRENT MEDICATIONS AND OTHER ORDERS AND A SCREEN TO DISCONTINUE AND REPLACE THE DOSE AND FREQUENCY. THERE HAVE BEEN SEVERAL CIRCUMSTANCES IN WHICH THE ORIGINAL ORDER FOR WARFARIN AND OTHER MEDICATIONS, A POTENTIALLY DANGEROUS AND LIFE-THREATENING THERAPY, WAS REPLACED USING THIS MECHANISM OF THE CPOE. WHEN THE PROCESS IS COMPLETED, THERE SHOULD BE ONLY ONE ORDER FOR THE MEDICATION IN QUESTION ON THE ACTIVE MEDICATION ADMINISTRATION LIST. WE HAVE DETERMINED THAT THERE IS A SOFTWARE CODE DEFECT SUCH THAT THE ORIGINAL ORDER DOES NOT GET DELETED (AND MOVED TO ANOTHER SILO). WE SEE THAT THERE ARE NOW TWO ACTIVE ORDERS FOR THE SAME MEDICATION. IF THE TIME SCHEDULE IS DIFFERENT, THERE IS ABOVE AVERAGE LIKELIHOOD THAT THE PT WILL GET BOTH DOSES. MOST USERS DESCRIBE THIS POE SYSTEM AS "DANGEROUS" AND THAT THERE IS INDICATION TO REMOVE IT FROM USE. | 1
2195233 | IN THE EARLY MORNING, A RHYTHM CAME ACROSS THE MONITOR THAT LOOKED LIKE A FAST VT HEART RATE OF APPROXIMATELY 300. THE MONITOR DID NOT ALARM, BUT A NURSE WAS RIGHT IN FRONT OF THE MONITOR. SHE WENT IMMEDIATELY TO THE PATIENT'S BEDSIDE, GRABBED HER NURSE ON THE WAY, FOUND THE PATIENT ASLEEP ON THEIR STOMACH. THE PATIENT WAS AWAKENED AND CHECKED. THE PATIENT WAS FINE, JUST A LITTLE FRIGHTENED. THE NURSE THEN CALLED BIOMED. LATER IN THE MORNING, WE RECEIVED THE INFORMATION AND SHARED THE EKG WITH THE ATTENDING PHYSICIANS. THEY WERE IN A QUANDARY; HALF OF THE PHYSICIANS FELT THIS WAS REAL AND HALF FELT IT WAS ELECTRICAL. BIOMED CAME UP AND LOOKED AT THE PROBLEM. THEY FELT THERE WAS NO PROBLEM AND THAT IT WASN'T ELECTRICAL INTERFERENCE AND THAT IT WAS THE PATIENT. WE ASKED THE EP FELLOW TO GET INVOLVED AND HE TOOK IT TO OTHER EP ATTENDING PHYSICIANS AND THEY WERE SPLIT AS TO THE RECORDINGS BEING REAL. THE EP FELLOW RETURNED LATER IN THE DAY AND CHECKED THE RECORDINGS BEFORE AND AFTER THE EVENT AND FOUND A LOT OF INTERFERENCE. | 1
2196292 | COMPLAINANT ALLEGED THAT DURING A ROUTINE SHIFT CHECK BY A CLINICIAN, THE DEVICE WAS UNABLE TO POWER ON. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. THE PRODUCT HAS BEEN RECEIVED AND A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED. | 0
2201321 | IT WAS REPORTED THAT AN S/5 PM MULTI I/O ADAPTOR MODULE APPEARED TO BE BURNED FROM THE INSIDE OUT WITH MELTING OF THE EXTERNAL ENCLOSURE. IT WAS NOT REPORTED IF THE UNIT WAS IN USE AT THE TIME. THERE WAS NO REPORT OF PT INVOLVEMENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. THE S/5 DEVICE HAS BEEN REPAIRED AND RETURNED TO SERVICE. THE BURNED MODULE IS BEING RETURNED FOR EVALUATION. | 1
2202130 | A PT WITH FAILURE OF SEVERAL ORGANS WAS BECOMING DEHYDRATED. IT WAS DECIDED TO HYDRATE THE PT WITH 500 CC OF FLUID AT 60 CC PER HOUR. THE CPOE HAS AVAILABLE FOR SELECTION 500 CC AND 500 CC WAS SELECTED. HOWEVER, THE FLUID, RAN IN TO THE PT CONTINUOUSLY, RESULTING IN A FLUID OVERDOSE OF 3X IN EXCESS OF WHAT WAS ORIGINALLY ORDERED. THE PT DEVELOPED PULMONARY EDEMA OF THE LUNGS AND FURTHER COMPLICATIONS RESULTED IN DEATH. IN THIS CASE, THE ORDER ENTRY DEVICE HAS AMBIGUOUS AND MISLEADING INFORMATION THAT RESULTED IN THE FLUID OVERDOSE, AND THEN DEATH "DOR THIS INFORM." | 1
2203999 | IT WAS REPORTED THAT PACEART EXPERIENCED AN ERROR IN THE DATA EXCHANGE LOG VIEWER POTENTIALLY RESULTING IN "ONE OR MORE FIELDS NOT MATCHING UP, CAUSING PACEART TO BE UNABLE TO MATCH THE TRANSMISSION TO A PATIENT." PACEART REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.><IT WAS REPORTED THAT PACEART EXPERIENCED AN ERROR IN THE DATA EXCHANGE LOG VIEWER POTENTIALLY RESULTING IN "ONE OR MORE FIELDS NOT MATCHING UP, CAUSING PACEART TO BE UNABLE TO MATCH THE TRANSMISSION TO A PATIENT." PACEART REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT. | 1
2205321 | IT WAS REPORTED THAT THE CLINIC DID NOT RECEIVE A FAX BACK FROM A CARDIOSIGHT TRANSMISSION. THE FAX WAS RE-SENT. THE MONITOR REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT.><IT WAS REPORTED THAT THE CLINIC DID NOT RECEIVE A FAX BACK FROM A (B)(6). THE FAX WAS RE-SENT. THE MONITOR REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. IT WAS REPORTED THAT THE CLINIC DID NOT RECEIVE A FAX BACK FROM A (B)(6). THE FAX WAS RE-SENT. THE MONITOR REMAINS IN USE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT. ALSO NOTED: THE INITIAL REPORT HAS THE INCORRECT DEVICE REPRESENTED. | 1
2206518 | IT WAS REPORTED UNABLE TO RETRIEVE CARELINK TRANSMISSION INTO PACEART FOR A PATIENT. CLIENT HAS THE DEVICE ENTERED IN PACEART WITH THE WRONG NAME FORMAT, REMOTE FOLLOW UP WAS NOT ENABLED, AND IS RUNNING GCII, BUT THE DEVICE TRANSMISSION IS ONLY SUPPORTED IN GC3 AFTER DEVICE UPDATE HAS BEEN RUN. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.><IT WAS REPORTED UNABLE TO RETRIEVE CARELINK TRANSMISSION INTO PACEART FOR A PATIENT. CLIENT HAS THE DEVICE ENTERED IN PACEART WITH THE WRONG NAME FORMAT, REMOTE FOLLOW UP WAS NOT ENABLED, AND IS RUNNING GCII, BUT THE DEVICE TRANSMISSION IS ONLY SUPPORTED IN GC3 AFTER DEVICE UPDATE HAS BEEN RUN. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. | 1
2208104 | ASKU><IT WAS REPORTED THAT FOLLOWING A SOFTWARE UPDATE TO THE CLIENT'S PACEART SYSTEM, A DUPLICATE ENTRY FOR A DIAGNOSTIC CODE WAS CREATED WHICH CHANGED THE DIAGNOSIS OF PATIENTS. TROUBLESHOOTING REMOVED THE DUPLICATE ENTRY. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. FURTHER REVIEW PROMPTED A CHANGE IN THE DEVICE ANALYSIS RESULTS. THE CHANGE IS REFLECTED IN THIS REPORT.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2208644 | DURING A KNEE ARTHROSCOPY PROCEDURE, THE FMS SOLO AND DAY TUBING WERE BEING USED. THE TUBING WAS SET UP PRIOR TO THE PATIENT ENTERING THE ROOM. WHEN THE CASE WAS STARTED AND TUBING WAS HOOKED UP TO THE FIELD, THE PUMP DID NOT SENSE THE PRESSURE DIFFERENCE AND CONTINUALLY RAN, PUSHING FLUID INTO THE KNEE. THIS WAS NOTICED AND FLOW WAS STOPPED. TUBING REPLACED AND WORKED TO CORRECT PARAMETERS FOR THE CASE WITH NO INJURY TO PATIENT. THE PRIMING OF THIS PUMP TUBING REQUIRES THAT: THE PRESSURE SENSOR BE ATTACHED FINGER TIGHT TO THE PRESSURE SENSOR PORT; PRESS THE POWER ON/STANDBY BUTTON ON THE FRONT OF THE CONSOLE; PRESS THE RUN/STOP BUTTON ON THE FRONT OF THE CONSOLE; PRESS THE FILL CHAMBER BUTTON ON THE FRONT OF THE CONSOLE. THIS NEEDS TO BE PRESSED A FEW TIMES TO FILL THE CHAMBER TO A LEVEL ABOVE THE INDICATOR WINDOW. IF THE FILL CHAMBER BUTTON WAS PRESSED REPEATEDLY BEFORE THE RUN/STOP BUTTON WAS PRESSED, IT WILL CAUSE THE FLUID TO BE PUMPED TO EQUAL THE AMOUNT OF TIMES THAT THE FILL CHAMBER BUTTON WAS PRESSED WHEN THE RUN/STOP BUTTON IS FINALLY PRESSED. IF THIS WAS DONE, AND THE CHAMBER WAS FILLED TO THE DESIRED LEVEL, AND THEN THE RUN/STOP BUTTON WAS PRESSED AGAIN TO STOP THE FLOW, AND THEN PRESSED AGAIN TO START THE FLOW AT THE BEGINNING OF THE PROCEDURE, THEN THE PUMP WOULD FLOW TO CATCH UP TO THE ACTIVATED FILL CHAMBER CYCLES. IT IS BELIEVED THIS MAY HAVE HAPPENED IN THIS CASE. | 0
2209135 | IT WAS REPORTED THAT SCHEDULING INFORMATION FOR SOME PATIENTS WAS NOT TRANSFERRING FROM PACEART INTO CARELINK, AND TRANSMISSIONS SHOWED AS "MISSED." PACEART REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT.><. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION NOTE: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT. | 1
2213022 | PATIENT WAS SCHEDULED FOR KNEE ARTHROSCOPY. AT THE END OF PROCEDURE, IT WAS NOTED THAT THE THIGH WAS SWOLLEN WITH MODERATE ERYTHEMA AND SKIN TURGOR WAS TAUT. PT WAS ASSESSED. APPEARANCES SUGGEST FLUID FROM ARTHROSCOPY ENTERED THE SURROUNDING TISSUES. | 0
2213259 | THE PT HAD HAD PRIOR INTRACRANIAL NEUROSURGERY FOR BRAIN CANCER. THE PT UNDERWENT INTERVAL BRAIN MRIS TO ASSESS INTERVAL CHANGE, IF ANY. WHEN AN MRI OF THE BRAIN WAS PERFORMED FOR NEW ONSET APHASIA, THE PACS DEVICE BROUGHT UP A PRIOR MRI FOR COMPARISON WHICH WAS PERFORMED MORE THAN A YEAR PRIOR AND WAS MORE ANCIENT THAN THE PAST THREE. THUS, THE ANALYSIS BY THE RADIOLOGIST WAS FLAWED, PUTTING THE PT AT RISK OF MIS-DIAGNOSIS AND CLINICAL ERROR IN TREATMENT. | 1
2213261 | THE PT WITH A MULTIPLE ORGAN FAILURE SYNDROME HAD AN ARRHYTHMIA CONSISTING OF WIDE COMPLEX TACHYCARDIA. AN ORDER TO NOT GIVE THE PT THE PREVIOUSLY ORDERED DIGOXIN DOSE UNTIL A SERUM LEVEL WAS VERIFIED WAS ENTERED USING CPOE. THE LEVEL RETURNED AT A SATISFACTORY RANGE. ON THE PROCEEDING DAY, THE SCREEN SHOWING RECENT ORDERS INDICATED THAT THE MEDICATION HAD BEEN DISCONTINUED. CLINICAL AND PHARMACOLOGICAL DECISIONS WERE MADE BASED ON THAT. IT WAS ERRONEOUS INFO. THE MEDICATION APPEARED TO HAVE BEEN DISCONTINUED, BUT IT WAS STILL BEING ADMINISTERED. WITH A MEDICATION OF NARROW THERAPEUTIC RATIO, THE SAFE ADMINISTRATION REQUIRES PRECISE DETAIL OF ADMINISTRATION AND SCHEDULE. THIS DETAIL INVOLVED THE CPOE SYSTEM'S FALSE LISTING OF A DISCONTINUATION OF A MEDICATION OF NARROW THERAPEUTIC RATIO PUTTING THE PT AT RISK FOR TOXICITY, VENTRICULAR TACHYCARDIA, AND DEATH. THE CLINICIAN WAS LED TO BELIEVE FROM THE ELECTRONIC RECORD THAT THE MEDICATION HAD BEEN DISCONTINUED, WHEN IT WAS NOT. | 1
2213389 | DISCORDANT HIGH ADVIA CENTAUR TROPONIN AND CKMB RESULTS WERE OBTAINED ON MULTIPLE PATIENT SAMPLES. THE LABORATORY REPEATED THE SAMPLES AND CORRECTED REPORTS WERE ISSUED TO THE PHYSICIAN(S). ONE PATIENT WAS CATHETERIZED DUE TO THE DISCORDANT TROPONIN AND CKMB RESULTS. A SIEMENS FSE (FIELD SERVICE ENGINEER) WAS DISPATCHED TO THE CUSTOMER SITE TO EVALUATE THE INSTRUMENT AND INSTRUMENT DATA. AFTER ANALYSIS OF THE INSTRUMENT AND INSTRUMENT DATA, IT WAS DETERMINED THAT THE CAUSE OF THE DISCORDANT TROPONIN AND CKMB RESULTS IS UNKNOWN. THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED. | 0
2216027 | THE PT REPORTED SOMETHING HOT IN THE BED. NURSING DISCOVERED THAT THE M4841A TELEMETRY TRANSMITTER WAS HOT TO THE TOUCH AND EMITTING AN ODOR. THE NURSE REMOVED THE TRANSMITTER FROM THE PT AND THEN REMOVED THE BATTERIES. IT WAS NOTED ONE OF THE BATTERIES PROTECTIVE COATING WAS PEELED AWAY AND IT WAS HOT. THE TRANSMITTER WAS REMOVED FROM CIRCULATION AND SENT TO CLINICAL ENGINEERING. IT WAS DETERMINED THAT THE BATTERY HAD SHORT CIRCUITED DUE TO A MISSING INSULATOR PAD IN THE BATTERY COMPARTMENT THAT SERVES TO SEPARATE THE BATTERY CONNECTOR WIRES AND AS THE LOCATION FOR THE DEVICE'S NETWORK HARDWARE IDENTIFIER. | 1
2216032 | A LOADING DOSE REGIMEN OF DIGOXIN WAS CPOED INTO THE ELECTRONIC MEDICAL RECORD, FOR THE PHARMACY AND FOR THE NURSE. THIS WAS FOLLOWED BY A MAINTENANCE DOSE THAT WAS TO BEGIN THE FOLLOWING DAY. THE ACTIVE MEDICATION AND ORDER LIST DISPLAYED THE DIGOXIN REGIMEN SUCH THAT IT MISSTATED THE DOSE, PUTTING THE PT AT RISK FOR OVER DOSE ON BOTH THE LOADING DOSE AND ON THE MAINTENANCE DOSE. THE LISTING OF THE ORDERS AS INITIATED AND DISCONTINUED ON THE ORDERS SCREEN RENDERS IT DIFFICULT, IF NOT IMPOSSIBLE TO, CLINICALLY DETECT AND FOLLOW THE COURSE OF THE CASE. THIS PLACES THE PT IN ADD'L CLINICAL PERIL. | 1
2216033 | BLOOD CHEMISTRY TEST ORDERED BY CPOE IN PT WITH MULTI-SYS FAILURE. IT HAS BECOME COMMONPLACE FOR ORDERS TO NOT BE CARRIED OUT. THE INTENDED RECIPIENT DOES NOT GET THE ORDER, ALTHOUGH THERE ARE OTHER ETIOLOGIES. THERE IS NOT ANY NOTIFICATION THAT THE ORDERED TEST WAS NOT DONE. THIS SEEMS TRIVIAL ON THE SURFACE, BUT, THE EVAL AND MGMT OF A CRITICALLY ILL PT WAS DELAYED. THERE ARE NOT ANY RECORDS OF THE INCIDENCE OF THIS CPOE PROBLEM IN WHICH A TEST IS ORDERED BUT NOT DONE. IT SUBJECTS CRITICALLY ILL PTS TO RISK. | 1
2216034 | THIS (B)(4) SYS HAS A FEATURE THAT ENABLES THE DISCONTINUATION OF A MEDICATION DOSE WHILE REPLACING IT WITH ANOTHER DOSE. AFTER THE TRANSACTION, THE RECENT ORDER SCREEN SHOWS THE DISCONTINUATION OF ONE DOSE AND THE START OF ANOTHER DOSE OF THE SAME MEDICATION. THE ACTIVE MEDICATION LIST SHOWS ONLY THE NEW DOSE. IT HAS BEEN DETERMINED THAT WHEN USING THIS FUNCTION, THE PRIOR MEDICATION AND DOSE IS NOT ALWAYS DISCONTINUED AND REMAINS ON THE ACTIVE MEDICATION LIST, ALONG WITH THE SAME MEDICATION WITH THE NEW DOSE. THERE CAN BE TWO DIFFERENT DOSES OF WITH DIFFERENT TIMES FOR THE SAME MEDICATION. THIS PUTS THE PT AT RISK FOR BEING OVERDOSED. THE DEFECTIVE CODE MANIFESTS THIS PROBLEM RANDOMLY AND UNPREDICTABLY, CREATING AND REQUIRING DISTRUST FOR THE ENTIRE ELECTRONIC ORDERING SYS. THERE HAVE BEEN CASES INVOLVING DIGOXIN, A MEDICATION WITH A NARROW THERAPEUTIC RATIO; AND OTHERS INVOLVING DILTIAZEM. | 1
2216049 | THE PT REPORTED SOMETHING HOT IN THE BED. NURSING DISCOVERED THAT THE M4841A TELEMETRY TRANSMITTER WAS HOT TO THE TOUCH AND EMITTING AN ODOR. THE NURSE REMOVED THE TRANSMITTER FROM THE PT AND THEN REMOVED THE BATTERIES. IT WAS NOTED ONE OF THE BATTERIES PROTECTIVE COATING WAS PEELED AWAY AND IT WAS HOT. THE TRANSMITTER WAS REMOVED FROM CIRCULATION AND SENT TO CLINICAL ENGINEERING. IT WAS DETERMINED THAT THE BATTERY HAD SHORT CIRCUITED DUE TO A MISSING INSULATOR PAD IN THE BATTERY COMPARTMENT THAT SERVES TO SEPARATE THE BATTERY CONNECTOR WIRES AND AS THE LOCATION FOR THE DEVICE'S NETWORK HARDWARE IDENTIFIER. | 1
2216056 | A DISPARITY BETWEEN THE DRIP RATE ON THE CPOE ORDER AND THE LISTING OF THE DRIP RATE ON THE ACTIVE MEDICATION LIST WAS REPEATEDLY DETECTED. THIS DISPARITY IS GIVING RISE TO CONFUSION AND DOUBT IN THE MIND OF THE CLINICIAN AS TO WHAT DOSE THE PT IS GETTING AND WHO KNOWS AND WHO IS DECIDING. THE DRUG SUBJECT, AMIODARONE, CARRIES A BLACK BOX WARNING THAT IT MAY CAUSE FATAL TOXICITIES. THIS DEFECT IN THE CPOE DEVICE IS AN IMPEDIMENT TO SAFE CARE. THE PT WAS AT RISK FOR OVERDOSE AND UNDER DOSE OF A COMPLEX MEDICATION. | 1
2216057 | THE PT WITH ATRIAL FIBRILLATION, RESPIRATORY FAILURE, ANEMIA, AND PRIOR GASTROINTESTINAL BLEED WAS ORDERED BY CPOE A DAILY DOSE OF WARFARIN WITH A DIRECTIVE TO ADJUST THE DOSE FOR AN INR ABOVE 2.4 SO AS TO NOT OVER ANTICOAGULATE AND RISK BLEEDING. IT HAS BEEN A REPEATED OCCURRENCE THAT THE DIRECTIVE DOES NOT ARRIVE FOR THE PHARMACY AND NURSE RECIPIENT OR THAT THE DIRECTIVE, DUE TO A PROBLEM WITH USABILITY, IS NOT READILY SEEN BY PHARMACISTS AND NURSES. ONE SUCH PROBLEM IS THE EXCESSIVE AND UNNECESSARY VERBIAGE. IN THESE INSTANCES, THE PT RECEIVED THE DRUG, CAUSING UNDUE RISK FOR ADVERSE EVENTS. THE FACT THAT THIS OCCURS CAUSES COGNITIVE DISTRACTION AND DISTRUST FOR THE ENTIRE SYS OF CARE. | 1
2226840 | HAD BARD WEB PUT IN AND HAD ADVERSE REACTION REQUIRING SECOND SURGERY. | 0
2228167 | DR (B)(6), ALLEGED THAT THREE PTS UNNECESSARILY HAD DIAGNOSTIC HEART CATHETERIZATION PERFORMED DUE TO HIS INTERPRETATION OF ECG CHARTS OBTAINED THROUGH THE USE OF LIFESYNC CORP PRODUCT. INVESTIGATION OF THE INCIDENT DETERMINED THAT ONLY ONE PT WAS PRESCRIBED CATHETERIZATION THAT WAS SUBSEQUENTLY FOUND TO HAVE NO EVIDENCE OF BLOCKAGE. NO HARM TO THE PT WAS REPORTED. HOWEVER, THE MEDICAL DIRECTOR AT THE HOSP ALLEGED THAT THE PROCEDURE WAS UNNECESSARY AND INTRODUCED A HEALTH AND SAFETY RISK TO THE PT. THIS COMPLAINT WAS NOT ASSOCIATED WITH A FAILURE OF THE LIFESYNC LEADWEAR TO MEET A SPECIFICATION OR USER REQUIREMENT. IN THIS CASE, A CARDIOLOGIST AT (B)(6) HOSP ALLEGED THAT THE OUTPUT OF THE ECG DATA USING LIFESYNC LEADWEAR IN ADDITION TO THE PT'S HISTORY OF HEART DISEASE AND APPARENT SYMPTOMS (CHEST PAINS) LED TO THE DETERMINATION THAT THE PT REQUIRED DIAGNOSTIC HEART CATHETERIZATION. FOLLOWING CATHETERIZATION, THE PT WAS NOTED TO HAVE PATENT "CLEAN" ARTERIES. UPON REQUEST BY LIFESYNC; OBJECTIVE EVIDENCE COULD NOT BE PROVIDED REGARDING THE ECG DATA THAT LED TO THE DIAGNOSTIC CATHETERIZATION. HOWEVER, CONVERSATION WITH THE NURSE MGR INDICATED THAT THE PT'S ECG HAD BEEN ANALYZED USING MASON-LIKAR LEAD PLACEMENT AS WELL AS STANDARD 12 LEAD DISTAL PLACEMENT. ACCORDING TO THE NURSE MGR, THE PT WAS ORDERED TO UNDERGO CATHETERIZATION REGARDLESS OF THE LEAD PLACEMENT USED TO OBTAIN THE ECG. NO HARM OCCURRED TO THE PT AS A RESULT OF THE CATHETERIZATION. THE (B)(6) LAB POLICY IS TO USE STANDARD 12 LEAD DISTAL LEAD PLACEMENT. LIFESYNC PRODUCT IS DESIGNED TO BE USED WITH MASON-LIKAR MODIFIED CHEST LEAD POSITIONS. STANDARD 12 LEAD DISTAL LEAD PLACEMENT AND MASON-LIKAR LEAD PLACEMENT HAVE SEVERAL KEY DIFFERENCES INHERENT TO EACH PLACEMENT METHOD THAT LEAD TO VARIATION BETWEEN ECG'S OBTAINED BY EITHER METHOD. AN ECG OBTAINED WITH MASON-LIKAR PLACEMENT SHOULD NOT BE INTERPRETED THE SAME AS AN ECG OBTAINED USING 12-LEAD DISTAL PLACEMENT. PRIOR TO THIS INCIDENT, LIFESYNC CORP PROVIDED APPLICATION NOTES AND PRODUCT LABELING TO ALL CUSTOMERS (INCLUDING (B)(6)) DETAILING THE DIFFERENCES BETWEEN MASON-LIKAR PLACEMENT AND STANDARD 12 LEAD DISTAL PLACEMENT AS WELL AS APPLICATION NOTES THAT PROVIDE INSTRUCTION FOR CONTINUOUS ST SEGMENT MONITORING USING MASON-LIKAR LEAD PLACEMENT. (B)(4). | 1
2228247 | IT WAS REPORTED THAT THE VIGILANCE II MEAN ARTERIAL VALUE (MAP) DID NOT MATCH THE GE SOLAR 8000 (RAC4A) BEDSIDE MONITOR'S MAP. NO PATIENT COMPLICATIONS WERE REPORTED. A 70ANM5 SLAVE CABLE WAS USED TO TRANSMIT INFORMATION BETWEEN THE TWO MONITORS. THE ACCOUNT WAS A NEW USER TO THE SYSTEM AS OF JANUARY 2007. (B)(6). (B)(4): INCORRECT PARAMETER CALCULATIONS. FOLLOWUP WITH THE ACCOUNT INDICATED THAT THEY ARE NOT RETURNING THE VIGILANCE MONITOR. THE DFU INSTRUCTS THE USER TO VERIFY THE SLAVED MAP VALUES WITH THOSE DISPLAYED ON THE BEDSIDE MONITOR. REFERENCE VIGILANCE II DFU, SECTION 2-6. THE STATEMENT IS AS FOLLOWS: "CAUTION: THE ACCURACY OF CONTINUOUS SVR, VQI AND O2 EL DEPENDS UPON THE QUALITY AND ACCURACY OF THE MAP, CVP AND SA02 DATA TRANSMITTED FROM THE EXTERNAL MONITORS. SINCE MAP, CVP, AND SA02 ANALOG SIGNAL QUALITY FROM THE EXTERNAL MONITOR CANNOT BE VALIDATED BY THE VIGILANCE II MONITOR, ACTUAL VALUES AND THE VALUES (INCLUDING ALL DERIVED PARAMETERS) DISPLAYED BY THE VIGILANCE II MONITOR MAY NOT BE CONSISTENT. THE ACCURACY OF CONTINUOUS SVR, VQI AND O2EI MEASUREMENT, THEREFORE, CANNOT BE GUARANTEED. TO AID IN DETERMINING THE QUALITY OF THE ANALOG SIGNALS, REGULARLY COMPARE THE MAP, CVP AND SA02 VALUES DISPLAYED ON THE EXTERNAL MONITOR TO THE VALUES DISPLAYED ON THE VIGILANCE MONITOR. REFER TO EXTERNAL INPUT DEVICE OPERATOR'S MANUAL FOR DETAILED INFORMATION. REGARDING ACCURACY, CALIBRATION, AND OTHER VARIABLES WHICH MAY IMPACT THE ANALOG OUTPUT SIGNAL FROM THE EXTERNAL MONITOR." DEVICE 2: BRAND NAME: PRESSURE SLAVE CABLE; TYPE OF DEVICE: ANALOGUE SLAVE CABLE; MANUFACTURER: EDWARDS LIFESCIENCES (B)(4). MODEL#: 70ANM5, SINGLE USE DEVICE THAT WAS REPROCESSED?: NO; DEVICE AVAILABLE FOR EVAL?: NO. (B)(4): INCORRECT PARAMETER CALCULATIONS. (B)(4). | 1
2229008 | A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER REPORTED THAT AN INFANT PASSED AWAY WHILE USING A SMARTMONITOR 2 INFANT APNEA DEVICE. THE DATE AND TIME OF THE EVENT IS UNK. THE INCIDENT IS CURRENTLY BEING INVESTIGATED BY THE LOCAL POLICE DEPARTMENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVAL. A FOLLOW-UP REPORT WILL BE FILED WHEN THE DEVICE HAS BEEN RETURNED AND AN INVESTIGATION IS COMPLETED BY THE MANUFACTURER. | 1
2237171 | THE PYXIS MACHINE ARRANGEMENT HAS 3 COMPONENTS: THE MAIN UNIT AND TWO AUXILLARY UNITS. IT WAS REPORTED THAT WHEN REMOVING ITEMS CONTAINED IN ONE OF THE AUXILLARY UNITS, DRAWER 6 ON THE MAIN UNIT WOULD ALSO OPEN.THE COMPANY SENT A REPRESENTATIVE. THEY DISCOVERED THAT TWO BOTTLES OF MAGNESIUM CITRATE IN DRAWER 6 HAD SHATTERED WHEN THE DRAWER SHUT ROUGHLY. THE SOLUTION SPREAD TO THE ELECTRICAL COMPONENTS WHICH SHORTED OUT. | 0
2240680 | THE PINNACLE TREATMENT PLANNING SYSTEM HAS PROMPTS THAT ARE NOT USED AT OUR FACILITY. STAFF NEED TO CLICK THROUGH THE PROMPTS AS THEY CANNOT BE HIDDEN. STAFF "PASSED OVER" THE ONE PROMPT FOR LOCALIZATION WHICH RESULTED IN TREATMENT FIELDS BEING 9 CM OFF DESIRED TREATMENT AREA. ADDITIONAL PROMPTS HAVE BEEN ADDED TO THE SYSTEM FOR USING THE LOCALIZATION FEATURES. | 1
2241495 | DBSS FAILED BY ALLOWING AN INCORRECT BLOOD COMPONENT OTHER THAN THE PRODUCT REQUESTED TO BE ISSUED AND TRANSFUSED. IN THIS EVENT, THE PHYSICIAN ORDERED 10 UNITS OF CRYOPRECIPITATE. THE LAB TECHNICIAN PROCESSED THE TRANSFUSION REQUEST AND INITIATED THE COMPONENT CONVERSION PROCESS (THAWING) OF 2 POOLED UNITS (1 POOLED UNIT IS 5 SINGLE UNITS) BUT COMMENTED ON THE "FATNESS" OF THE UNITS, AND WAS TOLD THAT POOLS OF 5 UNITS WOULD BE "FATTER" THAN SINGLE UNITS. THE UNITS WERE ACTUALLY CRYO-REDUCED PLASMA. UNITS WERE SET UP, ISSUED AND TRANSFUSED WITHOUT THE LAB TECHNICIAN OR TRANSFUSING NURSE NOTICING THAT THE UNITS ISSUED WERE NOT THE REQUESTED PRODUCT. ACCORDING OT THE TRANSFUSION REQUEST FORM, THE PATIENT DID NOT EXPERIENCE A TRANSFUSION REACTION. | 1
2246198 | OMRON (B)(4) CONSISTENTLY PRODUCES BLOOD PRESSURES THAT ARE SUBSTANTIALLY LOWER THAN MANUAL BLOOD PRESSURES WITH A CALIBRATED TYCOS ANEROID SPHYGMOMANOMETER. THE UNIT HAS FRESH BATTERIES AND IS BEING USED ACCORDING TO INSTRUCTIONS. SBP TYPICALLY RUNS APPROX 10-15 MMHG HIGHER THAN MANUAL BP; DBP TYPICALLY RUNS APPROX 8-12 MMHG HIGHER. | 0
2249879 | MONITOR 1 - TELEMETRIES #1, #3, AND #5 SYNCED TO HAVE THE SAME HEART RATE AND TELEMETRY WAVE FORM. PT'S TELEMETRY CHANGED TO #16, #19 AND #22 ON MONITOR #2. THIS SHOWED SEPARATE HEART RATES AND WAVE FORMS. TELEMETRIES #11 AND #12 ON MONITOR 1 WERE NOT AFFECTED. AT (B)(6), TELEMETRIES #11 AND #12 SYNCED AND READ THE SAME HEART RATES, RHYTHMS AND ECTOMY. TRIED TO CHANGE TO A DIFFERENT CHANNEL. NO OTHER CHANNEL WAS ABLE TO PICK UP A SIGNAL SO WHEN SET TO ORIGINAL CHANNEL THE TWO TELE READINGS SWITCHED. PT ON #11 HAD BEEN IN SINUS RHYTHM AND PT ON #12 HAD BEEN IN A FIB/RVR. TELE #12 BEGAN READING SINUS RHYTHM AND #11 BEGAN READING A FIB/RVR. REASSIGNED NEW TELEMETRY UNITS TO BOTH PATIENTS. STOPPED USING ON (B)(6) 2011. | 1
2250776 | PATIENT WAS IN THE MIDDLE OF HAVING HIS STRESS WHEN THE TECH HAD TO PUT THE STRESS MACHINE INTO PAUSE STAGE TO KEEP IT FROM GETTING ANY FASTER. I THEN TURNED THE SPEED DOWN WHILE THE DOCTOR ENTERED THE ROOM AND TURNED THE GRADE OF THE TREADMILL DOWN AS WELL. THE DOCTOR THEN TRIED TO TURN THE SPEED DOWN AND IT WOULD NOT ALLOW HIM TO DO ANYTHING. TECH TRIED TO PUT THE PATIENT INTO RECOVERY MODE AND IT WOULD NOT DO ANYTHING. TECH STOPPED THE TREADMILL AND BROUGHT THE PATIENT BACK TO THE STRETCHER WITH NO PROBLEM. THE PATIENT WAS NOT INJURED DURING THIS TRANSITION. TECH THEN TRIED TO END THE TEST AND THE MACHINE WOULD NOT REACT TO ANYTHING. THEN TURNED THE MACHINE OFF AND RESTARTED IT. THE MACHINE WAS STILL NOT WORKING. AFTER THE PATIENT LEFT, I CALLED GE TECH SUPPORT AND WE DISCOVERED THAT THE CONTROL PANEL ON THE FRONT OF THE STRESS MACHINE WAS NOT FUNCTIONING PROPERLY. THEN PUT A WORK ORDER FOR BIOMED TO COME AND FIX IT. WE DID LOSE PEAK AND RECOVERY EKG DATA BECAUSE OF THIS INCIDENT. | 0
2253744 | CUSTOMER REPORTED THAT THE TOSOH G8 REPORTED ERRONEOUSLY LOW SA1C RESULTS ON PTS WITH HGBC VARIANTS. THE VARIANTS WERE CONFIRMED BY ANOTHER METHOD BUT TOSOH CHROMATOGRAPHY DID NOT INDICATE HV2 % FOR HGBC. RESULTS DID NOT CORRELATE WITH PREVIOUS RESULTS FOR PTS OR REFERENCE LAB RESULTS. ONE ERRONEOUS RESULT HAD BEEN REPORTED BUT NO TREATMENT WAS DONE BASED ON THE ERRONEOUS RESULTS. THE INVESTIGATION INTO THE ROOT CAUSE OF THE ERRONEOUS RESULTS INDICATED THAT ON (B)(6) 2011 AN INCORRECT VALUE HAD BEEN MANUALLY ENTERED DURING A SERVICE CALL FOR ONE OF THE G8 IDT PARAMETERS WHICH RESULTED IN THE INCORRECT RESULTS. ONCE THE IDT PARAMETER WAS CORRECTED THE TOSOH CHROMATOGRAPHY RESULTS CORRELATED WITH THE EXPECTED RESULTS AND OTHER METHODOLOGY RESULTS. | 0
2254908 | NURSING SERVICE FOUND A M4841A TELEMETRY TRANSMITTER BOX ON THE PT TO BE HOT TO THE TOUCH. EQUIPMENT WAS REMOVED FROM THE PT AND THE BATTERY WAS REMOVED. EQUIPMENT WAS SENT TO CLINICAL ENGINEERING. THE WAS NO VISIBLE EXTERNAL DAMAGE TO THE TRANSMITTER AND THE BATTERY LOAD SPRINGS WERE IN PROPER POSITION. (B)(6) TESTED THE TRANSMITTER (B)(6), REMOVING AND REPLACING THE BATTERIES X3 DAILY AS WERE NOT ABLE TO REPRODUCE "HOT/HEAD" BATTERIES. (B)(6) WILL SEND THE TRANSMITTER FOR ADDITIONAL EVALUATION. | 0
2259128 | IT WAS REPORTED THAT THE MONITOR DID NOT POST THE ASYSTOLE ALARM ON A PT WITH PACEMAKER. USING PARAMETER SIMULATOR FLUKE PS 420, THE PACEMAKER RECOGNITION WAS FUNCTIONAL AS WELL AS THE ALARMING OF ASYSTOLE. IT WAS REPORTED THAT THE PT DIED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2259202 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND REPLACED THE SYRINGE AND READER CAMERA. THE FE PERFORMED INSTRUMENT ADJUSTMENTS AND RAN DIAGNOSTICS.THE CUSTOMER SUCCESSFULLY RAN QC. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4) | 0
2259379 | DURING PREPARATION OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THAT THE MONITOR WOULD NOT COME UP. AS A RESULT, AN ALTERNATE DEVICE WAS EMPLOYED. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WERE NO ADVERSE CONSEQUENCES TO THE PT AS A RESULT OF THIS EVENT. EVAL IN PROGRESS, BUT NOT YET CONCLUDED. | 1
2263377 | THE PATIENT WENT TO THE DOCTOR'S OFFICE FOR SHORTNESS OF BREATH AND EXHAUSTION. THE PATIENT HAD HIGH TEST RESULTS ON THE HEMOCURE ANALYZER VERSUS A LABORATORY INSTRUMENT. CAPILLARY RESULTS WERE 10G/DL; VENOUS RESULTS WERE 5G/DL. NOTE: THE PATIENT WAS NOT ON ANY MEDICATIONS. THE PATIENT WAS DIAGNOSED AS ANEMIC BUT WAS SENT HOME DUE TO AN "ACCEPTABLE" RESULT ON THE HEMOCUE ANALYZER. THE NEXT DAY (WHEN THE LABORATORY RESULTS CAME BACK), THE PATIENT WAS CALLED BACK IN FOR A BLOOD TRANSFUSION AND WAS GIVEN TWO (2) UNITS. FOLLOWING TREATMENT, THE PATIENT WAS DEEMED TO BE IN GOOD HEALTH. NOTE: THIS INCIDENT WAS FOUND DURING A RETROSPECTIVE REVIEW OF COMPLAINTS. THE REVIEW WAS PERFORMED IN ORDER TO VERIFY THAT CORRECT JUDGEMENT REGARDING FILING OF MDRS WAS PERFORMED FOR COMPLAINTS WHERE A MEDICAL INTERVENTION HAD BEEN INVOLVED IN ORDER TO PRECLUDE ANY PERMANENT DAMAGE. | 0
2263438 | UNRESPONSIVE, VENTILATED PATIENT SUSTAINED A SECOND DEGREE BURN TO FOREHEAD AND NAPE OF NECK FOLLOWING USE OF NO-RINSE SHOWER CAP. RN HEATED PRODUCT ACCORDING TO MANUFACTURER'S AND PRODUCT PACKAGE INSTRUCTIONS. BURN NOTED AFTER CAP WAS REMOVED AFTER APPROXIMATELY FIVE MINUTES. PHYSICIAN MADE AWARE AND ORDERED SILVADENE TO AFFECTED AREAS AND CONSULTED WOUND CARE. | 0
2263523 | VIDEO SIGNAL LOSS ON MONITOR DURING ESU (ELECTROSURGICAL UNIT) OPERATION VIDEO CABLE SHIELD SEPARATED FROM BNC (BI-NETWORK CO-AX CABLE) CONNECTORCABLE REPLACEDFUNCTIONAL TEST PASSEDRETURNED TO SERVICE. | 1
2264023 | MCKESSON NARCSTATION DOES NOT HAVE THE PROGRAMMING TO SCAN PRODUCTS WHEN ADDING INVENTORY INTO THE CABINETS EVEN THOUGH THE TECHNOLOGY EXISTS. THE AUTOMATED DISPENSING CABINETS OR "ACUDOSE" BY MCKESSON UTILIZES BAR-CODING FOR ASSIGNING AND RESTOCKING INVENTORY. THE SAME TECHNOLOGY SHOULD BE AVAILABLE FOR THE MCKESSON NARCSTATION. ON VARIOUS OCCASIONS, THE INCORRECT DRUG WAS ORDERED THROUGH THE WHOLESALER AND ACCIDENTALLY RESTOCKED INTO THE WRONG SPOT IN THE NARCSTATION. A SIMPLE BAR-CODING FEATURE COULD HAVE PREVENTED THIS MEDICATION ERROR. THIS LACK OF CAPABILITY COULD LEAD TO PATIENT HARM. | 1
2267000 | WHILE INVESTIGATING A DOCUMENTATION ISSUE IN ALLSCRIPTS ELECTRONIC HEALTH RECORD, SOME OF THE TASKS THAT THE PROVIDERS WERE SENDING TO THE NURSING STAFF REGARDING TEST RESULT FOLLOW UP WERE ERRORING OUT AND STAFF WERE UNAWARE. APPROXIMATELY 100 PATIENTS MAY NOT HAVE RECEIVED TEST RESULTS DUE TO THIS ERROR GOING BACK AS FAR AS (B)(4) 2010. | 1
2267005 | CLINICIANS REPORTED ALARMS BETWEEN THE VENTILATOR AND MONITORING WERE NOT FUNCTIONING AFTER BIOMED UPGRADED THE GE SOLAR SOFTWARE VERSION 4E AND 4F TO GE SOLAR VERSION 5.3. INITIAL TROUBLESHOOTING WAS UNSUCCESSFUL.INVESTIGATION REVEALED THAT THE VERSION 5.3 SOFTWARE CANNOT RECEIVE SERIAL DATA FROM THE OCTANET (WHICH SERVES AS THE INTERFACE BETWEEN THE SOLAR MONITOR AND A VENTILATOR). RESOLUTION CAME AFTER REPLACING THE OCTANET WITH A UNITY ID AND THERE HAVE NOT BEEN FURTHER REPORTED ISSUES. THERE IS NO OBVIOUS DOCUMENTED INFORMATION IN GE MANUAL MENTIONING COMPATIBILITY RISK. | 1
2268169 | PATIENT HAD 8 BEATS OF VENTRICULAR TACHYCARDIA. THE MAIN CARDIAC MONITOR (CENTRAL STATION) IN UNIT DID NOT ALARM. | 1
2268897 | PHILIPS CARDIAC SYSTEM MODEL M3155 SELF-REBOOTING RESULTING IN LOSS OF PT MONITORING ON 16 TELEMETRY MONITORED PTS FOR UP TO SEVEN MINS. THE RE-BOOT RESULTED IN THE LOSS OF ECG MONITORING INCLUDING THE ABILITY TO GENERATE ANY TYPE OF CARDIAC ALARM ON THE UNITS AFFECTED. FOR PTS ON PORTABLE WIRELESS TELEMETRY, THIS ALSO IMPACTS THE PHILIPS PORTABLE MP5 MONITORS USED AT THE BEDSIDE TO SEE THE PT WAVEFORM. FOR HARDWIRE BEDS IN ICU, THE BEDSIDE MONITOR CONTINUES TO FUNCTION INDEPENDENTLY, BUT CENTRAL MONITORING IS IMPACTED. THE RE-BOOTING IS UNPREDICTABLE AND DOES NOT HAPPEN IN THE SAME UNIT, BUT RANDOMLY IN ANY ONE OF THE TEN M3155 CPU'S IN THE SYSTEM. WHEN THE REBOOT OCCURS THERE ARE NO ALARMS OR WARNING GENERATED TO WARN THAT THE EVENT IS HAPPENING AND STAFF MAY NOT BE AWARE OF THE ISSUE UNTIL A FEW MIN AFTER THE EVENT HAS BEGUN. | 1
2268923 | ON TWO PTS, WHEN TELEMETRY PATCHES TAKEN OFF, UNDERNEATH THE PATCHES THE PT'S SKIN WAS BLISTERY AND PUS DRAINAGE COMING FROM OPEN AREAS UNDERNEATH THE PATCHES. | 0
2269744 | LAMP WILL NOT TURN ON, EVEN WITH A KNOWN GOOD ONE. TROUBLESHOOT IT TO THE BALLAST ASSEMBLY; WILL NEED TO REPLACE IT AND WILL ORDER PART. THE PART ARRIVED; WILL INSTALL AND TROUBLESHOOT BY THE END OF THE WEEK. TALKING TO TECHNICAL SUPPORT TO ISOLATE PROBLEM. GE SUPPORT WILL CALL BACK WITH MORE INFORMATION. A FEW WEEKS LATER TROUBLESHOOTING THE PROBLEM AND PERFORMD A PM (PREVENTIVE MAINTANENCE) TEST. THE SPOT LIGHT IS ABLE TO TURN ON. PERFORMED THE OPERATIONAL TEST. UNIT IS BACK IN SERVICE. | 0
2270107 | OLYMPUS WAS INFORMED THAT DURING OBSERVATION FOLLOWING AN ENDOSCOPIC MUCOSAL RESECTION (EMR) PROCEDURE, X-RAY REVEALED FREE AIR IN THE PT'S CHEST CAVITY. OLYMPUS FOLLOWED-UP WITH THE USER FACILITY TO OBTAIN ADD'L INFO REGARDING THIS REPORT. THE USER FACILITY REPORTED THAT THE PHYSICIAN HAD DIFFICULTY INTUBATING INTO THE ESOPHAGUS AND NOTED SLIGHT BLEEDING. REMOVAL OF THE LESION WAS SAID TO HAVE PERFORMED SUCCESSFULLY WITHOUT FURTHER COMPLICATION. THE PT WAS KEPT FOR OBSERVATION FOLLOWING THE PROCEDURE AND AN X-RAY WAS SAID TO HAVE PERFORMED WHICH REPORTEDLY DISCOVERED FREE AIR IN THE PT'S CHEST CAVITY LEADING THE USERS TO BELIEVE THERE WAS A PERFORATION ON THE ESOPHAGUS. THE PT WAS REPORTEDLY ADMITTED AND PROVIDED WITH ANTIBIOTICS. THE PT REPORTEDLY REMAINED AT THE HOSP FOR 24-48 HRS AND WAS SENT HOME. THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO OLYMPUS FOR EVAL. THE CAUSE OF THE REPORTED PHENOMENON COULD NOT BE CONCLUSIVELY DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | 0
2270706 | ACIST DEVICE MALFUNCTIONED PRIOR TO CORONARY ANGIOGRAPHY BEING PERFORMED. ANGIOGRAPHY PERFORMED WITH MANUAL INJECTIONS. | 1
2272152 | THIS REPORT WAS RECEIVED FROM GLOBAL PHARMACOVIGILANCE (GPV) AND IS A SPONTANEOUS REPORT BY A CONSUMER WITH SUPPLEMENTAL INFORMATION BY A NURSE FROM (B)(6) OF PERITONITIS IN A PATIENT COINCIDENT WITH DIANEAL PD2 AMBUFLEX THERAPY. DURING A CALL WITH BAXTER CUSTOMER SERVICE, THE FOLLOWING INFORMATION WAS REPORTED. ON (B)(6) 2010, THE PATIENT EXPERIENCED PERITONITIS. TREATMENT WAS NOT REPORTED. THE PATIENT WAS NOT HOSPITALIZED. ON (B)(6) 2010, THE PATIENT HAD RECOVERED FROM THE EVENT OF PERITONITIS. DIANEAL THERAPY WAS ONGOING. THE NURSE STATED THAT THE EVENT OF PERITONITIS WITH (B)(6) WAS UNRELATED TO DIANEAL THERAPY. (B)(4). AS THE DATE OF ONSET OF THIS PERITONITIS EPISODE IS UNKNOWN AND PATIENTS DISCARD SUPPLIES AFTER EACH THERAPY, THE SAMPLE WAS NOT REQUEST. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. THIS IS REPORT 1 OF 3 INVOLVED IN THIS PERITONITIS EVENT.><(B)(4). A BATCH REVIEW WAS CONDUCTED ON POTENTIALLY ASSOCIATED LOT NUMBERS: H10H03047, AND H10G17049 WITH NO DEFECTS NOTED. THE CAUSE OF THE PERITONITIS WAS UNDETERMINED. BAXTER HAS RECEIVED SIMILAR REPORTS FOR THE REPORTED PROBLEM. THE ROOT CAUSE INVESTIGATION IS IN PROGRESS. | 0
2279217 | CUSTOMER REPORTS THE SERIES ARE SPLIT INCORRECTLY. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. PHONE NUMBER UNKNOWN. INVESTIGATION REVEALED THE REPORTED EVENT IS DUE TO THREAD INTERFERENCE, AS A THREAD IS SPAWNED FOR BOTH THE EXAM AND THE HISTORICAL TO OPEN THEM AT THE SAME TIME. BOTH THREADS ARE TRYING TO ACCESS AND MODIFY A STATIC VARIABLE WHICH SPECIFIES HOW EACH IMAGE IN EACH SERIES SHOULD BE GROUPED (I.E. BY SERIES). THIS IS NOW MORE EVIDENT BECAUSE OF MULTIPLE PROCESSORS AND FASTER PC'S. THE PHYSICIANS HAVE BEEN MADE AWARE OF THE ISSUE AND KNOW HOW TO RESOLVE IT. IN ADDITION, A WORKAROUND IS BEING IMPLEMENTED AT THE SITE UNTIL THE NEXT RELEASE OF PACS WHICH WILL ADDRESS THIS ISSUE. (B)(4). | 1
2280266 | IT WAS REPORTED THAT A LARGE, PARALYZED PT WAS SCANNED ON A 1.5T HDE WRAPPED IN A SHEET ONLY WITH NO PROTECTIVE PADDING. SHE RECEIVED A 2 CM BY 3 CM BLISTER ON EACH ARM WHERE IT CONTACTED THE SYSTEM BORE. IT IS UNK AT THIS TIME WHETHER THE PT REQUIRED TREATMENT FOR HER BURNS. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
2280290 | OEC 9900 ELITE CARDIAC CONFIGURATION; PURCHASE NEW; INSTALLED (B)(6) 2009. COMMUNICATION FAILURE WARNING FOLLOWED BY CONTROL PANEL ERROR. FAILURE OF FAST STOP SWITCH - UNABLE TO STOP GENERATION OF X RAYS. ONLY WAY TO STOP X RAY GENERATION WAS TO TURN OFF THE MACHINE. PROCEDURE DISCONTINUED; PATIENT AWAKENED FROM ANESTHESIA. REASON FOR USE: INTERVENTIONAL STENTING PROCEDURE. | 0
2281136 | THE TELEMETRY SYSTEM WAS EXPERIENCING SOME PROBLEMS WITH THE SERVER MANAGING THE DEVICE. THIS MAY POSSIBLY BE RELATED TO A POWER SURG. THIS REQUIRED REBOOTING THE SERVER AND ASSOCIATED DEVICES SEVERAL TIMES. IN THE PROCESS OF REBOOTING APPARENTLY THE PATIENT ASSIGNMENTS TO THE TELEMETRY BOXES WERE SWITCHED BETWEEN PATIENTS, RESULTING IN PATIENT LABELS ON THE SYSTEM DISPLAY BEING REFERENCED TO THE INCORRECT WAVEFORMS AND ALARMS. THIS WAS DISCOVERED AND RECTIFIED THE NEXT MORNING. NO PATIENT HARM. | 1
2291631 | ACCESSED THIS PATIENT'S MEDICAL RECORD IN CERNER EMR VIA CERNER POWERCHART INTERFACE AT (B)(6) HOSPITAL OF (B)(6) ON COMPUTER IP (B)(6) AT 19:12:57, YET LATER THAT SAME EVENING WHEN ACCESSING CERNER VIA POWERCHART INTERFACE AT SAME FACILITY FROM COMPUTER IP ADDRESS (B)(6) AT 20:29:59 CERNER EMR SHOWED NO EVIDENCE OF MY PRIOR ACCESS TO THIS PATIENT'S CHART AS EVIDENCED BY "RECENT" LIST ONLY SHOWING 1 PATIENT, WHICH WAS NOT THIS PATIENT. THIS IS MOST CONCERNING AS EVIDENCE OF DYSFUNCTIONAL OR NON-WORKING AUDIT TRAIL. | 1
2291642 | ACCESSED THIS PATIENT RECORD IN CERNER EMR VIA CERNER POWERCHART INTERFACE (B)(6) HOSPITAL OF (B)(6) AT COMPUTER IP ADDRESS (B)(6) AT 15:49:53, YET LATER SAME DAY WHEN ACCESSING CERNER EMR VIA CERNER POWERCHART INTERFACE AT SAME FACILITY AT COMPUTER IP ADDRESS (B)(6) AT 20:29:59 = > CERNER EMR SHOWED NO EVIDENCE OF MY PRIOR ACCESS TO THIS PATIENT'S CHART AS EVIDENCED BY "RECENT" LIST TAB ONLY SHOWING ONE PATIENT WHICH WAS NOT THIS PATIENT. THIS IS MOST CONCERNING FOR EVIDENCE OF DYSFUNCTIONAL OR NON-WORKING AUDIT TRAIL IN THIS EMR. THIS IS RECURRING PROBLEM FOR EXTENDED PERIOD OF TIME, WITH NO RESOLUTION BY COMPANY TO DATE. I HAVE DATE-TIME-STAMPED PICTURES OF MONITOR SCREENS TO DOCUMENT THIS PROBLEM. | 1
2292708 | IT WAS REPORTED THAT THE TRANSMISSION FOR A PATIENT AND THE ECG STRIP WERE MISSING IN THE APPLICATION. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. | 1
2298304 | PATIENT WAS CONNECTED TO MONITOR IN EMERGENCY DEPARTMENT. MONITOR WOULD NOT COMPLETE EKG ANALYSIS AND "LEAD ERROR" DISPLAYED ON MONITOR. ELECTRODES WERE CHANGED AND LEAD REPLACED. MONITOR WOULD STILL NOT ANALYZE AND "UNABLE TO PERFORM EKG ANALYSIS WITH 3 LEADS" DISPLAYED ON MONITOR.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, DASH (PER SITE REPORTER).======================THE ONSITE GE HEALTHCARE BIO MEDICAL TECHNICIANS REPLACED THE TRUNK CABLE TO THE MONITOR. | 1
2298788 | PATIENT ON TELEMETRY BOX BEDSIDE MONITOR AND CENTRAL EKG MONITOR WAS DISPLAYING ECG TRACING INCORRECTLY. ECG TRACING SEEMS TO SPEED UP AND SLOW DOWN BUT HEART RATE STAYED AT NORMAL LEVELS. WHEN TOOK PATIENT OFF TELEMETRY BOX AND PUT PATIENT ON ECG HARD WIRE MONITOR, THE INCORRECT ECG TRACING WITH SPEED UP AND SLOW DOWN EPISODES STOPPED. | 1
2300762 | PATIENT ECG MODULE WAS NOT COMMUNICATING WITH PATIENT MONITOR. THERE WAS A BROKEN E-PORT ON THE MODULE.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, MODULE, CARESCAPE, PDM (PER SITE REPORTER).======================YES, THEY HAVE MADE REPAIRS BY REPLACING THE E-PORT. | 1
2300786 | NO FAILURE. IMMEDIATE REPLACEMENT MODULE USED TO SOLVE PROBLEM.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, MODULE, CARESCAPE, PDM (PER SITE REPORTER).======================YES, THEY HAVE MADE REPAIRS BY REPLACING THE E-PORT CONNECTOR. | 1
2300806 | NO FAILURE. IMMEDIATE REPLACEMENT MODULE USED TO SOLVE PROBLEM.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, MODULE, CARESCAPE, PDM (PER SITE REPORTER).======================UNIT HAS BEEN REPAIRED, E-PORT REPLACED. | 1
2306904 | THE SURGEON REQUIRED THE USE OF THE O-ARM IMAGING DEVICE DURING SURGERY. THE DEVICE WAS PLACED AROUND THE PATIENT DURING THE SURGICAL CASE WITHOUT DIFFICULTY AND IMAGES WERE OBTAINED. AFTER THE IMAGES HAD BEEN TAKEN, STAFF ATTEMPTED TO OPEN THE DEVICE, BUT THE DEVICE HAD NO POWER AND WOULD NOT OPEN. STAFF CHECKED ALL OF THE CABLES/CORDS AND ALL WERE PLUGGED IN. THERE WERE NO OBVIOUS DEFECTS, FRAYS, OR OTHER INDICATIONS THAT THE DEVICE SHOULD NOT BE IN FULL OPERATIONAL STATUS. BIOMED AND THE MANUFACTURER'S REP WERE CONTACTED. THE MEDTRONIC SUPPORT STAFF MEMBER CAME TO THE OR AND ALSO WAS UNABLE TO TROUBLESHOOT THE DEVICE. HE USED A WRENCH TO MANUALLY OPEN THE O ARM AND REMOVE THE DEVICE FROM THE SURGICAL FIELD. THE STERILITY OF THE SURGICAL FIELD WAS MAINTAINED AND THERE WAS NO HARM TO THE PATIENT. PER THE MD AND THE OR STAFF, THE PATIENT DID NOT INCUR EXTRA ANESTHESIA TIME AND THERE WERE NO PROBLEMS WITH THE IMAGES OBTAINED BY THE O-ARM DEVICE. THIS PARTICULAR PIECE OF EQUIPMENT IS LESS THAN ONE YEAR OLD. HOWEVER, THE STAFF OPERATING THE EQUIPMENT AND THE PHYSICIAN USING IT ARE VERY FAMILIAR AND EXPERIENCED WITH THIS TYPE OF DEVICE. THE MEDTRONIC REP THAT CAME TO THE OR DURING THE CASE COMMENTED THAT THEY HAVE SEEN THIS KIND OF PROBLEM BEFORE AND THAT "YOU MAY HAVE A LEMON." | 0
2307287 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A BLOODY FLUID LEAK UNDER THE COULTER LH 750 HEMATOLOGY ANALYZER. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A COAT, GLOVES AND EYE PROTECTION AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED AND MEDICAL ATTENTION WAS NOT SOUGHT. NO PATIENT RESULTS WERE AFFECTED BY THIS EVENT. BEC SET UP A SERVICE VISIT TO CHECK THE INSTRUMENT AND A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. HOWEVER, THE CUSTOMER CALLED BEC LATER INFORMING THAT THE LEAK HAD BEEN FIXED BY ONE OF THEIR TECHNICIANS. THE CUSTOMER DID NOT PROVIDE ANY MORE INFORMATION REGARDING THE LEAK. THE LOCATION AND ROOT CAUSE OF THE LEAK ARE UNKNOWN. (B)(4). | 0
2307647 | IT WAS REPORTED THAT THE INFINITY ACUTE CARE MONITORING SYSTEM (IACS) TRIED APPROX 15 MINS TO OBTAIN A VALID NON INVASIVE BLOOD PRESSURE (NIBP) VALUE. AFTERWARDS THE MANUAL START OF THE NIBP MEASUREMENT WAS NOT POSSIBLE BECAUSE THE START BUTTON WAS GRAYED OUT ON THE DISPLAY. THE SYSTEM NEEDED A RESTART TO MEASURE THE NIBP. IT WAS REPORTED THAT DUE TO THE LACK OF MONITORING FUNCTION THE PT'S STATE OF HEALTH WAS NOT CLEARLY VISIBLE / UNDERSTANDABLE AND THAT THE PEDIATRIC PT RECEIVED A PERMANENT BRAIN DAMAGE. THE INVESTIGATION OF THE REPORTED EVENT HAS BEEN STARTED BUT IS NOT FINISHED. RESULTS WILL BE REPORTED IN A FOLLOW UP REPORT. | 1
2309001 | WRONG AMOUNT OF GADAVIST WAS ENTERED INTO THE SIEMENS MAK (MED ADMINISTRATION CHECK) SYSTEM BECAUSE BAR CODE SCANNING GUN IN THE MRI SCAN ROOM IS NOT WORKING PROPERLY. THE SCAN GUN GOES INTO SOME TYPE OF "SLEEP MODE" AND NEEDS TO BE TRIGGERED MULTIPLE TIMES, THEN IT WILL SCAN CORRECTLY. WHILE TRYING TO SCAN GADAVIST BARCODE, THE SYSTEM REGISTERED A DOSE OF 65ML. WHEN I ENTERED THE CORRECT DOSE OF 10ML, THE SYSTEM ALERTED ME OF A RECENT CONTRAST ADMINISTRATION (WITHIN 48 HOURS). NURSELINKS CONFIRMED THAT 65ML DOSE WAS RECORDED. | 1
2309249 | A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WITH CAPTURE-R READY-SCREEN 3 (CRRS 3) ON THE GALILEO ECHO INSTRUMENT. THE PATIENT HAS A HISTORY OF ANTI-E. REVIEW OF IMAGES: WHEN REFERRED TO (B)(6) HOSPITAL THE SAMPLE INITIALLY TESTED NEGATIVE IN CRRS3 LOT R172 ON ECHO M00889. WELL 2 IS VISUALLY WEAKLY POSITIVE. THE SAMPLE WAS THEN TESTED ON CRRID EXTEND II LOT DN049 ON THE ECHO. WELLS 2, 4, AND 14 ARE HETEROZYGOUS E, THE THREE WELLS ARE VISIBLY WEAKLY POSITIVE. ALL OTHER WELLS ARE E NEGATIVE AND RESULTED AS NEGATIVE, VISUALLY APPEAR NEGATIVE. POSITIVE AND NEGATIVE CONTROL WELLS REACTED AS EXPECTED. NEXT THE SAMPLE WAS TESTED ON CRRID EXTEND I LOT DP050 ON THE ECHO. WELLS 1 THROUGH 6 ARE E NEGATIVE, ALL RESULTED AS NEGATIVE AND VISUALLY APPEAR NEGATIVE. WELL 7 IS HETEROZYGOUS FOR E AND RESULTED NEGATIVE, VISUALLY APPEARS WEAK POSITIVE. WELLS 8 THROUGH 14 ARE HOMOZYGOUS E. WELLS 8 AND 11 GAVE EQUIVOCAL RESULTS, VISUALLY BOTH APPEAR POSITIVE. WELL 14 RESULTED NEGATIVE, VISUALLY APPEARS WEAKLY POSITIVE. WELLS 9, 10, 12, AN 13 RESULTED 3+, VISUALLY POSITIVE. THE SAMPLE WAS RETESTED 8 DAYS LATER WITH CRRS3 LOT R172 AND RESULTED AS NEGATIVE IN ALL 3 WELLS, WELL 2, HOMOZYGOUS E, APPEARS VISUALLY POSITIVE. CUSTOMERS WERE INSTRUCTED TO VISUALLY REVIEW ALL NEGATIVE REACTIONS ON ECHO IN TECHNICAL COMMUNICATION CC-09-042-01 ISSUED ON NOVEMBER 25, 2009. | 0
2310270 | PRODUCT: DIGITAL BLOOD PRESSURE MONITOR, (B)(4) MODEL: 271245, SERIAL #: (B)(4) DISTRIBUTED BY (B)(6) MANUFACTURED BY OMRON. THIS DEFECTIVE PRODUCT HAS BEEN GIVING ME FALSE READINGS FOR BLOOD PRESSURE FOR THE LAST SEVERAL MONTHS. MY MD FINALLY HAD ME BRING IT IN TO CALIBRATE WITH HIS EQUIPMENT, AND IT WAS FOUND TO BE FAULTY AND PROVIDING FALSELY LOW READINGS. I HAVE BEEN UNDER THE CARE OF SEVERAL PHYSICIANS OVER THE LAST 3 YEARS FOR HYPERTENSION, AND MISTAKENLY THOUGHT IT WAS UNDER CONTROL DUE TO THESE FALSE READINGS. I CONTACTED THE MFR AND THEY REFUSED TO REPLACE THE PRODUCT, ONLY SAYING THEY WOULD REPAIR IT. (B)(6) WAS CONTACTED TO NO AVAIL ALSO. BOTH PARTIES WERE NOT CONCERNED IN THE LEAST REGARDING THE DEFECTIVE PRODUCT AND DID NOT ASK ME TO SEND IT IN. OMRON AND (B)(6) BOTH SHOWED NO INTEREST IN INVESTIGATING THE MATTER IN ANY WAY, WHICH WAS QUITE SHOCKING AND CONCERNING TO ME. | 0
2311550 | OUR AFFILIATE IS REPORTING TO US THAT DURING ONE DAY IN 3 SUBSEQUENT PROCEDURES, THERE WERE A SERIES OF EVENTS WITH THE USE OF THE SAME FMS DUO PUMP. IT SEEMS THAT THERE WAS SOME EXTENSIVE FLUID SPRAY, AND POSSIBLY SOME PATIENT EXTRAVASATION OR SWOLLEN COMPARTMENT ISSUES. THE SURGEON AND/OR STAFF MASSAGED THE AREA TO DISSIPATE THE FLUID. THE PROCEDURES WERE CONCLUDED SUCCESSFULLY WITHOUT FURTHER ISSUE OR HARM TO THE PATIENT. THE PATIENT'S WERE MONITORED CLOSELY DURING NORMAL RECOVERY. ALSO SEE ASSOCIATED MDRS 1221934-2011-00360 AND 1221934-2011-00362. MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT.><THE COMPLAINT DEVICE WAS RECEIVED AND EVALUATED. UPON RECEIPT, THE DEVICE WAS GIVEN A THOROUGH VISUAL AND FUNCTIONAL EXAMINATION. THE 1ST FAZE OF THE EXAMINATION WAS VISUAL: THIS VISUAL IS PERFORMED AS A MATTER OF COURSE TO DISCERN IF THE DEVICE HAS ANY OBVIOUS PHYSICAL DAMAGE, SOMETHING THAT MAY OR COULD HAVE BEEN A FACTOR IN CONTRIBUTING TO THE REPORTED EVENT; IT IS NOTED THAT THE DEVICE WAS RECEIVED WITH SOME MINOR PHYSICAL DAMAGE (BENT GUIDE LINE AND MISSING SEAL), NOT A CONTRIBUTOR TO THE EVENT; HOWEVER, THAT IS ATTRIBUTED TO HANDLING AND MAINTENANCE; A USER ISSUE. THE 2ND PORTION OF THE EXAMINATION WAS THE FUNCTIONAL AND ELECTRICAL TESTING: A BATTERY OF TESTS WERE PERFORMED TO ASSESS THE DEVICE'S OPERATIONAL STATUS. FUNCTIONALLY AND ELECTRONICALLY, THE DEVICE PERFORMED WELL WITHIN ITS DESIGN AND MANUFACTURED PARAMETERS; NO FAULT OR PERFORMANCE ANOMALY COULD BE FOUND WITH THE FMS PUMP. THE ROOT OR UNDERLYING CAUSE FOR THE REPORTED EVENT IS NOT ATTRIBUTABLE TO THE FMS PUMP, BUT LIES ELSEWHERE. AT THIS POINT IN TIME, NO FURTHER ACTION IS WARRANTED; HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FORTHCOMING INFORMATION RECEIVED THAT IS PERTINENT AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
2312209 | THE PATIENT WENT TO THE OR STATUS POST (S/P) REPAIR OF ATRIOVENTRICULAR SEPTAL DEFECT (AVSD) WITH A TWO PATCH CLOSURE IN A TRISOMY 21 PATIENT THAT WAS EXTUBATED ON CPAP. IN THE EVENING THE PATIENT APPEARED TO BE IN ACCELERATED JUNCTIONAL RHYTHM. THE PATIENT'S HEART RATE BEGAN TO DECREASE TO THE LOW 100S AND APPEARED TO BE IN HEART BLOCK. WAVE REVIEW ON TELEMETRY WAS UNAVAILABLE TO CONFIRM THE PATIENT'S EXACT RHYTHM DUE TO THEIR CONTINUED RHYTHM INSTABILITY; THE DECISION WAS MADE TO ELECTIVELY INTUBATE THE PATIENT. THE TELEMETRY WAVEFORMS WERE NOT BEING DISPLAYED ON THE CENTRAL STATION MONITOR. THE TELEMETRY WAS NOT LINKING AND NO PAST TELEMETRY RESULTS COULD BE VIEWED BY THE CARDIOLOGIST. THEY COULD ONLY BE VIEWED IN REAL TIME. THE PATIENT WAS ELECTIVELY INTUBATED AND SUBSEQUENTLY CODED WHILE THE ET TUBE WAS BE BEING REPOSITIONED AND RE-TAPED AND HAD SUDDEN PROFOUND HYPOTENSION UNRESPONSIVE TO FLUID BOLUSES. THE PATIENT WAS SUCCESSFULLY CODED; I.E. COMPRESSIONS, DEFIBRILLATION AND MEDICATIONS ON MILRINONE IV. | 1
2312488 | REPORTER INDICATED THAT A PATIENT RETURNED TO THE OFFICE SEVERAL HOURS AFTER HAVING A VNS SYSTEMS DIAGNOSTICS TEST PERFORMED DUE TO COUGHING, PAIN IN THE NECK, AND CHOKING. THE PATIENT'S VNS WAS INTERROGATED AND THE SETTINGS WERE 1MA/20HZ/500 PULSEWIDTH/30 SEC ON/60 MIN OFF, INDICATING A FAULTED SYSTEMS DIAGNOSTICS TEST HAD OCCURRED PREVIOUSLY. THE SETTINGS WERE CHANGED TO 0.75MA/20HZ/250 PULSEWIDTH/30 SEC ON/1.8 MINUTES OFF. THE COUGHING, PAIN, AND CHOKING RESOLVED AFTER THE VNS SETTINGS WERE CHANGED. THE REPORTER WAS ADVISED TO ALWAYS PERFORM A FINAL INTERROGATION OF THE VNS PRIOR TO THE PATIENT LEAVING THE OFFICE TO VERIFY SETTINGS. . | 0
2314721 | THE USER REPORTED THAT A LARGE BREAST PADDLE CRACKED 1/2 WAY ACROSS THE ENTIRE WIDTH OF THE PANEL WHILE COMPRESSING THE PATIENT. THE PATIENT WAS EVALUATED BY A PHYSICIAN. THERE WAS NO HARM TO THE PATIENT OR TECHNOLOGISTS PERFORMING THE TEST. THE ACTUAL FAILED COMPONENT HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE BREAST PADDLE IS MADE OF PLASTIC AND MAY WEAKEN AND CRACK IF THE CLEANING METHODS AND MATERIALS USED ARE OTHER THAN THOSE APPROVED BY HOLOGIC. AT THIS TIME THE REASON FOR THE REPORTED COMPONENT FAILURE IS UNKNOWN. | 0
2315635 | CARDIAC MONITOR DROPPING OFF WITH EPISODES OF SPEEDING UP TO A RATE OF 300/MINUTE AND SLOWING TO 20/MINUTE. THIS IS A PROBLEM SINCE PATIENT HEART RATE HAS PAUSES GREATER THAN 3 SECONDS AND REQUIRES A PACEMAKER. | 1
2319970 | THE CUSTOMER REPORTED THAT DURING AN ANGIOGRAPHY PROCEDURE, THE ROTATOR BROKE DURING INJECTION AT 900 PSI. THE CUSTOMER REPORTED THREE DEFECTIVE DEVICES. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF THREE REPORTS FOR THIS COMPLAINT: 1721504-2011-00326, 00333. DEVICE EVALUATION: THE SUSPECT DEVICES ARE NOT EXPECTED TO BE RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. MERIT IS UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED FAILURE WITHOUT THE SUSPECT DEVICE. NO CONCLUSION CAN BE DRAWN. METHOD: ACTUAL DEVICE NOT EVALUATED. PROCESS EVALUATION. | 0
2319987 | A NURSE WAS INJURED WHEN SHE CAME INTO CONTACT WITH A PATIENT MONITOR. DURING AN EMERGENCY MRI EXAMINATION , A DOCTOR MISTAKENLY BROUGHT A PATIENT MONITOR (MAGNETIC) FROM THE EMERGENCY ROOM INTO THE MRI ROOM. THE PATIENT MONITOR WAS ATTRACTED TO THE MAGNET AND IT CAME INTO CONTACT WITH THE NURSE WHO WAS IN THE ROOM. THE NURSE HAS A BRUISE/HEMATOMA, AND PAIN UPON PRESSURE IN HER RIGHT THIGH. THERE IS ALSO A SUBCUTANEOUS HEMORRHAGE UNDER HER RIGHT THIGH. THERE IS A SUBCUTANEOUS HEMORRHAGE, SWELLING, A FEELING OF HEAT, SHARP PAIN, AND PAIN UPON PRESSURE IN HER RIGHT FOOT JOINT. THE NURSE WAS GIVEN A COMPRESS AND WAS TOLD TO REST. SHE RESTED FOR 2 WEEKS. | 0
2321341 | A CUSTOMER CONTACTED BAXTER GLOBAL TECHNICAL SERVICES (GTS) REGARDING ASSISTANCE WITH THE HOMECHOICE (HC) UNIT DURING INITIAL DRAIN. PER THE INITIAL REPORT, HOME PATIENT (HP) HAD A CHECK PATIENT LINE ALARM. THE CAREGIVER (CG) CONNECTED THE HP TO THE MACHINE BEFORE CHECKING THE PATIENT LINE. THE PATIENT LINE WAS FILLED WITH AIR. THE TSR ASSISTED WITH ENDING THERAPY AND SO THE HP COULD RESTART WITH ALL NEW SUPPLIES. GLOBAL PRODUCT SURVEILLANCE CONTACTED THE PERITONEAL DIALYSIS NURSE (RN) ON (B)(6) 2011 ABOUT THE REPORTED PROBLEM. THE RN STATED SHE SAW THE HOME PATIENT (HP) ON (B)(6) 2011. THE RN STATED THE HP IS COMPLETING THERAPY AND DOING FINE. THERE WAS NO PATIENT INJURY OR MEDICAL INTERVENTION REPORTED. (B)(4). PER THE CUSTOMER THE SAMPLE WAS DISCARDED AND THE LOT NUMBER WAS UNKNOWN, THEREFORE NO EVALUATION OR BATCH REVIEW COULD BE PERFORMED. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THEN A FOLLOW UP MDR WILL BE SUBMITTED.><(B)(4). THIS COMPLAINT FOR THE CHECK PATIENT LINE ALARM IS NOT CONFIRMED AND THE CAUSE WAS NOT IDENTIFIED BECAUSE THE DISPOSABLE SET WAS NOT RETURNED TO BAXTER, THE LOT NUMBER WAS NOT PROVIDED, AN EVENT LOG REVIEW WAS NOT PERFORMED, AND THE USER DID NOT DESCRIBE ANY TYPE OF USE ERROR THAT MIGHT HAVE CAUSED THE ALARM. THE LABEL REVIEW FOUND THE LABELING ADEQUATE FOR THE USE ERROR IN THIS COMPLAINT. BAXTER HAS CONDUCTED A TREND REVIEW AND FOUND THAT SIMILAR REPORTS HAVE BEEN RECEIVED FOR THE REPORTED PROBLEM. THIS IS BEING INVESTIGATED THROUGH CAPA. THE ASSIGNABLE CAUSE FOR THE REPORTED PROBLEM WAS UNDETERMINED. | 0
2321756 | CALLER REPORTED A POTENTIAL (B)(4) RESULT ON ONE PATIENT. PHYSICIAN'S OFFICE CALLED TO REPORT A PATIENT THAT TESTED (B)(4) FOR OPIATE AND TCA. THEY WANTED TO KNOW WHY THE TOX SCREEN WAS NOT (B)(4) FOR ACETAMINOPHEN WHEN SHE WAS TAKING TYLENOL 3. INVESTIGATION PENDING. | 0
2326376 | THIS PT HAD BEEN ADMITTED TO ICU FOLLOWING AN ACUTE MI AND CARDIOGENIC SHOCK WITH CARDIAC ARREST. RECTAL TUBE PLACED ON (B)(6) 2011 DUE TO BLACK TAR LIKE STOOL. HER HOSP COURSE HAS BEEN COMPLICATED BY RENAL FAILURE AND RESPIRATORY FAILURE, AS WELL AS SEVERE THROMBOCYTOPENIA WITH SUSPICION FOR HIT AND WAS RECEIVING ARGATROBAN. FRIDAY MORNING, SHE DEVELOPED MASSIVE HEMORRHAGE FROM RECTUM; SHE UNDERWENT AN ANGIOGRAPHIC EVAL WITH EMBOLIZATION OF THE RECTAL ARTERY X 2 BY IR, AND REQUIRED MULTIPLE TRANSFUSIONS. BLEEDING ULCER IN DISTAL RECTUM IS THOUGH TO HAVE BEEN CAUSED BY THE BALLOON FROM RECTAL TUBE. | 0
2326820 | THE PT HAD A PHYSICIANS ORDER FOR CONTINUOUS ELECTROCARDIOGRAM (ECG) AND PULSE OXIMETRY -(SPO2)- MONITORING. THIS MONITORING WAS VIEWED ON GE'S APEX PRO CLINICAL INFO CENTER (CIC) LOCATED ON THE ACUTE CARE UNIT AND ALSO CONTINUOUSLY MONITORED BY A MONITORING TECH IN THE TELEMETRY MONITORING UNIT (TMU). AT THE TIME OF THE OCCURRENCE, THE SPO2 READING DISAPPEARED FROM PT WINDOW ON CENTRAL INFO CENTER (CIC) FOR BOTH THE UNIT AND THE TMU. THE DISAPPEARANCE OF THE READING DOES NOT CAUSE AN ALARM TO SOUND AND IS NOT HIGHLIGHTED ON THE VIEWING SCREEN. IN THE LOCATION WHERE THE SPO2 READING HAD BEEN, THE TRANSCEIVER (TX) NUMBER IS VIEWABLE AS IF THE XPOD CABLE FOR THE SPO2 PROBE HAD BEEN DISCONNECTED. UPON VISUALIZATION OF GE'S DT-4500 FH TELEMETRY TRANSCEIVER, THE PULSE OXIMETRY PROBE LED LIGHT WAS ILLUMINATED AND THE XPOD CABLE TO THE SPO2 PORT WAS SECURELY IN PLACE. THE XPOD CABLE WAS REPLACED WITH A NEW CABLE WITHOUT SOLUTION TO THE PROBLEM. THE TRANSCEIVER WAS REPLACED WITH A DIFFERENT TRANSCEIVER AND IMMEDIATELY BOTH THE ECG AND THE SPO2 READING WERE RE-ESTABLISHED WITH THE CIC FOR BOTH UNIT AND TMU. THE TRANSCEIVER WAS SET ASIDE FOR TESTING BY GE. UPON FURTHER TROUBLESHOOTING, IF THE BATTERY IS REMOVED AND REPLACED, THE TRANSCEIVER WILL RESET ITSELF AND THE SPO2 READING WILL REAPPEAR AND NORMAL OPERATION IS RESTORED. THIS WORKAROUND DOES NOT, HOWEVER, ALERT CAREGIVERS OF WHEN THE TRANSCEIVER WILL STOP READING THE SPO2. GE HAS BEEN NOTIFIED AND HAS SUPPLIED RESOURCES NEEDED TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. NO ADVERSE OUTCOME FOR PT. DATES OF USE: 10/20/2011 - 10/21/2011. | 1
2326826 | THE PT HAD A PHYSICIAN'S ORDER FOR CONTINUOUS ELECTROCARDIOGRAM (ECG) AND PULSE OXIMETRY (SPO2) MONITORING. THIS MONITORING WAS VIEWED ON GE'S APEX PRO CLINICAL INFO CENTER (CIC) LOCATED ON THE ACUTE CARE UNIT AND ALSO CONTINUOUSLY MONITORED BY A MONITORING TECHNICIAN IN THE TELEMETRY MONITORING UNIT (TMU). AT THE TIME OF THE OCCURRENCE, THE SPO2 READING DISAPPEARED FROM PT WINDOW ON CENTRAL INFO CENTER (CIC) FOR BOTH THE UNIT AND THE TMU. THE DISAPPEARANCE OF THE READING DOES NOT CAUSE AN ALARM TO SOUND AND IS NOT HIGHLIGHTED ON THE VIEWING SCREEN. IN THE LOCATION WHERE THE SPO2 READING HAD BEEN, THE TRANSCEIVER "TX" NUMBER IS VIEWABLE AS IF THE XPOD CABLE FOR THE SPO2 PROBE HAD BEEN DISCONNECTED. UPON VISUALIZATION OF GE'S DT-4500 FH TELEMETRY TRANSCEIVER, THE PULSE OXIMETRY PROBE LED LIGHT WAS ILLUMINATED AND THE XPOD CABLE TO THE SPO2 PORT WAS SECURELY IN PLACE. THE XPOD CABLE WAS REPLACED WITH A NEW CABLE WITHOUT RESOLUTION TO THE PROBLEM. THE TRANSCEIVER WAS REPLACED WITH A DIFFERENT TRANSCEIVER AND IMMEDIATELY BOTH THE ECG AND THE SPO2 READING WERE RE-ESTABLISHED WITH THE CIC FOR BOTH UNIT AND TMU. THE TRANSCEIVER WAS SET ASIDE FOR TESTING BY GE. UPON FURTHER TROUBLESHOOTING, IF THE BATTERY IS REMOVED AND REPLACED, THE TRANSCEIVER WILL RESET ITSELF AND THE SPO2 READING WILL REAPPEAR AND NORMAL OPERATION IS RESTORED. THIS WORKAROUND DOES NOT HOWEVER ALERT CAREGIVERS OF WHEN THE TRANSCEIVER WILL STOP READING THE SPO2. GE HAS BEEN NOTIFIED AND HAS SUPPLIED RESOURCES NEEDED TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. NO ADVERSE PT OUTCOME. OF NOTE, THE GE TELEMETRY TRANSCEIVER THAT IS THE SUBJECT OF THIS REPORT, WAS ISSUED BY (B)(4). | 1
2326827 | THE PT HAD A PHYSICIANS ORDER FOR CONTINUOUS ELECTROCARDIOGRAM (ECG) AND PULSE OXIMETRY -(SPO2)- MONITORING. THIS MONITORING WAS VIEWED ON GE'S APEX PRO CLINICAL INFO CENTER (CIC) LOCATED ON THE ACUTE CARE UNIT AND ALSO CONTINUOUSLY MONITORED BY A MONITORING TECH IN THE TELEMETRY MONITORING UNIT (TMU). AT THE TIME OF THE OCCURRENCE, THE SPO2 READING DISAPPEARED FROM PT WINDOW ON CENTRAL INFO CENTER (CIC) FOR BOTH THE UNIT AND THE TMU. THE DISAPPEARANCE OF THE READING DOES NOT CAUSE AN ALARM TO SOUND AND IS NOT HIGHLIGHTED ON THE VIEWING SCREEN. IN THE LOCATION WHERE THE SPO2 READING HAD BEEN, THE TRANSCEIVER (TX) NUMBER IS VIEWABLE AS IF THE XPOD CABLE FOR THE SPO2 PROBE HAD BEEN DISCONNECTED. UPON VISUALIZATION OF GE'S DT-4500 FH TELEMETRY TRANSCEIVER, THE PULSE OXIMETRY PROBE LED LIGHT WAS ILLUMINATED AND THE XPOD CABLE TO THE SPO2 PORT WAS SECURELY IN PLACE. THE XPOD CABLE WAS REPLACED WITH A NEW CABLE WITHOUT SOLUTION TO THE PROBLEM. THE TRANSCEIVER WAS REPLACED WITH A DIFFERENT TRANSCEIVER AND IMMEDIATELY BOTH THE ECG AND THE SPO2 READING WERE RE-ESTABLISHED WITH THE CIC FOR BOTH UNIT AND TMU. THE TRANSCEIVER WAS SET ASIDE FOR TESTING BY GE. UPON FURTHER TROUBLESHOOTING, IF THE BATTERY IS REMOVED AND REPLACED, THE TRANSCEIVER WILL RESET ITSELF AND THE SPO2 READING WILL REAPPEAR AND NORMAL OPERATION IS RESTORED. THIS WORKAROUND DOES NOT, HOWEVER, ALERT CAREGIVERS OF WHEN THE TRANSCEIVER WILL STOP READING THE SPO2. GE HAS BEEN NOTIFIED AND HAS SUPPLIED RESOURCES NEEDED TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. NO ADVERSE OUTCOME TO THE PT. DATES OF USE: (B)(6) 2011. | 1
2326828 | THE PT HAD A PHYSICIANS ORDER FOR CONTINUOUS ELECTROCARDIOGRAM (ECG) AND PULSE OXIMETRY -SPO2- MONITORING. THIS MONITORING WAS VIEWED ON GE'S APEX PRO CLINICAL INFO CENTER (CIC) LOCATED ON THE ACUTE CARE UNIT AND ALSO CONTINUOUSLY MONITORED BY A MONITORING TECH IN THE TELEMETRY MONITORING UNIT (TMU). AT THE TIME OF THE OCCURRENCE, THE SPO2 READING DISAPPEARED FROM PT WINDOW ON CENTRAL INFO CENTER (CIC) FOR BOTH THE UNIT AND THE TMU. THE DISAPPEARANCE OF THE READING DOES NOT CAUSE AN ALARM TO SOUND AND IS NOT HIGHLIGHTED ON THE VIEWING SCREEN. IN THE LOCATION WHERE THE SPO2 READING HAD BEEN, THE TRANSCEIVER (TX) NUMBER IS VIEWABLE AS IF THE XPOD CABLE FOR THE SPO2 PROBE HAD BEEN DISCONNECTED. UPON VISUALIZATION OF GE'S DT-4500 FH TELEMETRY TRANSCEIVER, THE PULSE OXIMETRY PROBE LED LIGHT WAS ILLUMINATED AND THE XPOD CABLE TO THE SPO2 PORT WAS SECURELY IN PLACE. THE XPOD CABLE WAS REPLACED WITH A NEW CABLE WITHOUT SOLUTION TO THE PROBLEM. THE TRANSCEIVER WAS REPLACED WITH A DIFFERENT TRANSCEIVER AND IMMEDIATELY BOTH THE ECG AND THE SPO2 READING WERE RE-ESTABLISHED WITH THE CIC FOR BOTH UNIT AND TMU. THE TRANSCEIVER WAS SET ASIDE FOR TESTING BY GE. UPON FURTHER TROUBLESHOOTING, IF THE BATTERY IS REMOVED AND REPLACED, THE TRANSCEIVER WILL RESET ITSELF AND THE SPO2 READING WILL REAPPEAR AND NORMAL OPERATION IS RESTORED. THIS WORKAROUND DOES NOT, HOWEVER, ALERT CAREGIVERS OF WHEN THE TRANSCEIVER WILL STOP READING THE SPO2. GE HAS BEEN NOTIFIED AND HAS SUPPLIED RESOURCES NEEDED TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. THE PT SUFFERED NO ADVERSE EVENT. DATE OF USE: (B)(6) 2011. | 1
2326844 | THE PT HAD A PHYSICIAN'S ORDER FOR CONTINUOUS ELECTROCARDIOGRAM -ECG- AND PULSE OXIMETRY(SPO2) MONITORING. THIS MONITORING WAS VIEWED ON GE'S APEX PRO CLINICAL INFO CENTER (CIC) LOCATED ON THE ACUTE CARE UNIT AND ALSO CONTINUOUSLY MONITORED BY A MONITORING TECH IN THE TELEMETRY MONITORING UNIT (TMU). AT THE TIME OF THE OCCURRENCE, THE SPO2 READING DISAPPEARED FROM PT WINDOW ON CENTRAL INFO CENTER -CIC-'S FOR BOTH THE UNIT AND THE TMU. THE DISAPPEARANCE OF THE READING DOES NOT CAUSE AN ALARM TO SOUND AND IS NOT HIGHLIGHTED ON THE VIEWING SCREEN. IN THE LOCATION WHERE THE SPO2 READING HAD BEEN, THE TRANSCEIVER TX NUMBER IS VIEWABLE AS IF THE XPOD CABLE FOR THE SPO2 PROBE HAD BEEN DISCONNECTED. UPON VISUALIZATION OF GE'S DT-4500 FH TELEMETRY TRANSCEIVER, THE PULSE OXIMETRY PROBE LED LIGHT WAS ILLUMINATED AND THE XPOD CABLE TO THE SPO2 PORT WAS SECURELY IN PLACE. THE XPOD CABLE WAS REPLACED WITH A NEW CABLE WITHOUT RESOLUTION TO THE PROBLEM. THE TRANSCEIVER WAS REPLACED WITH A DIFFERENT TRANSCEIVER AND IMMEDIATELY BOTH THE ECG AND THE SPO2 READING WERE RE-ESTABLISHED WITH THE CIC FOR BOTH UNIT AND TMU. THE TRANSCEIVER WAS SET ASIDE FOR TESTING BY GE. UPON FURTHER TROUBLESHOOTING, IF THE BATTERY IS REMOVED AND REPLACED, THE TRANSCEIVER WILL RESET ITSELF AND THE SPO2 READING WILL REAPPEAR AND NORMAL OPERATION IS RESTORED. THIS WORKAROUND DOES NOT, HOWEVER, ALERT CAREGIVERS OF WHEN THE TRANSCEIVER WILL STOP READING THE SPO2. GE HAS BEEN NOTIFIED AND HAS SUPPLIED RESOURCES NEEDED TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. NO ADVERSE OUTCOME TO PT. DATE OF USE: (B)(6) 2011. | 1
2334142 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THERE WAS A LEAK AROUND THE MANUAL PROBE OF THE COULTER LH 780 HEMATOLOGY ANALYZER WHILE RUNNING A SHUTDOWN. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED AND MEDICAL ATTENTION WAS NOT SOUGHT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE (B)(4) 2011 AND OBSERVED THAT THE LEAK WAS COMING FROM A BROKEN RINSE CUP AT THE PROBE WIPE. FSE REPLACED THE PROBE WIPE ASSEMBLY AND NO FURTHER LEAKING OCCURRED. REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES AND RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. THE ROOT CAUSE FOR THIS EVENT IS ATTRIBUTED TO THE BROKEN RINSE CUP AT THE PROBE WIPE. (B)(4). | 0
2336773 | ON (B)(6) 2011, WE BECAME AWARE OF AN ORDER ENTRY ERROR MADE ON OUR ABACUS SOFTWARE. ABACUS IS OUR (B)(4) BASED ORDER ENTRY SOFTWARE USED FOR TPN CALCULATIONS AND LABEL PRINTING. THE SOFTWARE INCLUDES WARNING LIMIT CAPABILITY, ALUMINUM TRACKING, DRAG AND DROP CALCIUM PHOSPHATE SOLUBILITY CURVE UTILITY AND OTHER FEATURES TO ACCOMMODATE INDIVIDUAL FACILITY PROTOCOLS. ACCORDING TO THE FACILITY, THE TPN ORDER WAS ENTERED INTO THE SOFTWARE INCORRECTLY. SPECIFICALLY, A DATA ENTRY ERROR WAS MADE IN REGARDS TO THE VOLUME OF HEPARIN ORDERED. FOR THIS REASON, THE TPN PRODUCED CONTAINED GREATER LEVELS OF HEPARIN THAN INTENDED. WHILE INVESTIGATING THIS INCIDENT, OUR TECHNICAL SUPPORT PULLED THE ASSOCIATED DEVICE FILES FOR EVALUATION. THROUGH THIS INVESTIGATION, WE FOUND THE SOFTWARE DELIVERED IN THE VOLUMES ORDERED. FURTHERMORE, THE ORDER VOLUMES WERE PRINTED ON THE TPN LABEL FOR ADDITIONAL VERIFICATION. THE DEVICE WAS FOUND TO BE OPERATING AS INTENDED, WITHOUT ANY SIGN OF ERROR. AFTER PRODUCTION, THE ORDER WAS ADMINISTERED TO THE PATIENT DURING AN 18-HOUR INFUSION. AS THE HEPARIN LEVELS WERE CLEARLY PRINTED ON THE TPN LABEL, THE ERROR WAS NOTED BY HOSPITAL STAFF DURING THE FOLLOWING MORNING'S DRIP ROUNDS. ONCE RECOGNIZED, THE ERROR WAS CORRECTED THROUGH THE ADMINISTRATION OF PROTAMINE, THEREBY REVERSING THE EFFECTS OF THE HEPARIN. NO ADVERSE OUTCOMES HAVE BEEN REPORTED CONCERNING THE PATIENT. THE USER FACILITY HAS CONFIRMED THAT THE PATIENT'S STATUS IS FINE. AFTER LEARNING OF THIS EVENT, OUR SUPPORT STAFF ASSISTED THE CUSTOMER BY ADDING A WARNING LIMIT TO PREVENT RECURRENCE. ADDITIONALLY, WE HAVE WORKED WITH THE CUSTOMER TO INSTALL MORE EXPANSIVE WARNING LIMITS. EVALUATION SUMMARY: WHILE INVESTIGATING THIS INCIDENT, OUR TECHNICAL SUPPORT PULLED THE ASSOCIATED DEVICE FILES FOR EVALUATION. THROUGH THIS EVALUATION, WE FOUND THE SOFTWARE TO BE OPERATING ACCORDING TO SPECIFICATIONS. THE ABACUS SOFTWARE CALCULATED THE ORDER VALUES EXACTLY AS ENTERED BY THE USER. FURTHERMORE, THE USER WAS PRESENTED WITH A DETAILED REPORT DESCRIBING THE CONTENTS OF THE TPN BAG. THIS REPORT CLEARLY SHOWS THE TOTAL VOLUME OF EACH INGREDIENT CONTAINED WITHIN THE BAG. STANDARD PRACTICE INCLUDES A REVIEW OF THESE REPORTS BY THE PHARMACIST AS A FINAL QUALITY CHECK. THROUGH OUR INVESTIGATION REGARDING THE PERFORMANCE OF THIS PRODUCT, WE HAVE CONCLUDED THAT NO MALFUNCTIONS OCCURRED. THE SYSTEM PERFORMED AS DESIGNED, PROVIDING THE NECESSARY TOOLS TO MAKE AN INFORMED DECISION REGARDING THE CONTENT OF THE TPN. ABACUS IS DESIGNED TO BE USED BY TRAINED PHARMACISTS AND/OR PHARMACY TECHNICIANS; THEREFORE, WHILE THE SOFTWARE PROVIDES INCREASED LEVELS OF SAFETY, IT CANNOT REPLACE THE PROFESSIONAL JUDGMENT OF A PHARMACIST. SINCE THIS EVENT, BAXA TECHNICAL SUPPORT HAS WORKED WITH THE CUSTOMER TO INSTALL MORE EXPANSIVE WARNING LIMITS TO HELP PREVENT A RECURRENCE OF THIS INCIDENT. EVALUATION METHOD: COMPUTER SOFTWARE PERFORMANCE TEST CONDUCTED. RESULTS: DEVICE PERFORMED ACCORDING TO SPECIFICATIONS. CONCLUSION: DEVICE OPERATED ACCORDING TO SPECIFICATIONS, NO DEVICE FAILURE, USER ERROR CAUSED EVENT. | 1
2337744 | SPACELABS RECEIVED A REPORT THAT A CENTRAL STATION MONITOR LOST COMMUNICATION WITH THE BEDSIDE MONITOR AND FAILED TO ALARM FOR ASYSTOLE.><SPACELABS RECEIVED A REPORT THAT A CENTRAL STATION MONITOR LOST COMMUNICATION WITH THE BEDSIDE MONITOR AND FAILED TO ALARM FOR ASYSTOLE. SPACELABS IS WORKING WITH THE CUSTOMER TO COLLECT INFORMATION TO START OUR INVESTIGATION. AT THIS TIME THERE IS INSUFFICIENT DATA TO IDENTIFY ROOT CAUSE OR DRAW ANY CONCLUSION.NO ONE HAS BEEN INJURED AS A RESULT OF THIS ALLEGED MALFUNCTION. WE WILL PROVIDE A SUPPLEMENTAL REPORT ONCE OUR INVESTIGATION IS COMPLETE.><TESTING CARRIED OUT BY SPACELABS IDENTIFIED THE ROOT CAUSE AS THE NETWORK SWITCH USED IN THE HOSPITAL. THE INVESTIGATION REVEALED THAT WHEN A D-LINK NETWORK SWITCH, MODEL NO. DES-1024 OR DES-1024(R), WAS IN USE, COMMUNICATION BETWEEN AN ELANCE BEDSIDE MONITOR AND THE CENTRAL CANNOT BE AUTOMATICALLY RESTORED ONCE THE COMMUNICATION IS LOST. IN THIS CIRCUMSTANCE, THE CENTRAL WILL NOT ALARM WHEN THE BEDSIDE MONITOR ALARMS. THE CUSTOMER HAS BEEN ADVISED NOT TO USE THE DES-1024(X) NETWORK SWITCH TO PREVENT FUTURE REOCCURRENCES. SPACELABS FIELD STAFF HAVE BEEN NOTIFIED TO WATCH OUT FOR THIS TYPE OF SWITCH DURING INSTALLATION AND RELATED INFORMATION IS BEING ADDED TO PRODUCT DOCUMENTATION. A FIELD SERVICE ENGINEER HAS HELPED THE CUSTOMER REPLACE THE SWITCH WITH A RECOMMENDED NETWORK SWITCH. AFTER CHANGING TO THE NEW SWITCH, THERE WERE NO FAILURES REPORTED IN THE LAST 2 WEEKS. THIS SUPPLEMENTAL REPORT IS CONSIDERED FINAL AND THE ISSUE IS CLOSED. | 1
2338445 | ACTUAL ERROR: YES. THIS IS A COMPUTER DESIGN THAT CAUSES ERRORS IN PRESCRIBING WITH PHYSICIAN ORDER ENTRY IN A HOSP. THIS DESIGN CAUSES MANY PTS ADMITTED ON NON-FORMULARY AGENTS CAUSES ERRORS. THIS DESIGN CAUSES MANY PTS ADMITTED ON NON-FORMULARY DRUGS TO BE CONVERTED TO THE LOWEST STRENGTH OF THE MEDICATION AVAILABLE ON FORMULARY RATHER THAN THE APPROPRIATE DOSE BASED ON OUT-PT THERAPY. WHEN REVIEWING DISCHARGE MED RECS, THE PHARMACISTS NOTED THAT MANY PTS WERE CONVERTED TO THE LOWEST STRENGTH AT ADMISSION AND WONDERED WHY THIS WAS HAPPENING. WHEN THEY QUESTIONED THE PRESCRIBERS THEY FOUND THE PRESCRIBER WAS NOT AWARE THAT THEY HAD INADVERTENTLY SELECTED THE WRONG STRENGTH BY DEFAULT. MANY PRESCRIBERS WERE UNDER THE IMPRESSION THE SYSTEM WAS AUTOMATICALLY ADJUSTING THE ORDER OR THAT PHARMACY INTERVENED AND SELECTED AN APPROPRIATE DOSE. THIS ISSUE ALSO OCCURS WITH ARB'S AS SHOWN, BUT ALSO WITH MANY OTHER DRUG CLASSES LIKE BETA-BLOCKERS AND ANTIBIOTICS. IN OUR OLD SYSTEM WITH PHARMACIST ORDER ENTRY, MANY AUTO-SUBS COULD BE DESIGNED SUCH THAT THE SELECTING THE DRUG AND STRENGTH AUTOMATICALLY TOOK THE USER TO THE EXACT FORMULARY SUB. MEDICATION ADMINISTERED TO OR USED BY THE PT: NO. WHERE DID THE ERROR OCCUR: HOSP. PT COUNSELING PROVIDED: UNK. RELEVANT MATERIALS PROVIDED: NONE. | 1
2340166 | AN INITIAL REPORT OF A SEVERE REACTION WITH TETANY WHILE DONATING PLATELETS ON AN AMICUS MACHINE WAS RECEIVED ON (B)(6) 2011. THE USER FACILITY INDICATED THAT THE DONOR WAS SENT TO THE HOSPITAL. THE DONATION TOOK PLACE ON (B)(6) 2011. THE FACILITY INDICATED THAT THE DONOR MIGHT HAVE HAD A STROKE OR A POSSIBLE AIR EMBOLISM. ON (B)(6) 2011, ADD'L DETAILS WERE RECEIVED AND THE USER FACILITY INDICATED THAT THERE WERE NO SIGNS OR INDICATIONS FROM THE PT OF DISCOMFORT FROM THE CITRATE. THE DONOR SUDDENLY GROANED AND BEGAN TO SEIZE. THE TECHS ON THE FLOOR GRABBED HIS ARMS TO KEEP THEM STRAIGHT. 911 WAS CALLED AND PARAMEDICS INDICATED HE HAD A STROKE. HIS RIGHT SIDE WAS PARALYZED AND HE WAS UNABLE TO SPEAK UPON DEPARTURE TO THE HOSP. PT RETURNED TO THE CENTER A FEW DAYS LATER FULLY RECOVERED. THE INSTRUMENT INVOLVED WAS EVALUATED AND REMAINED IN USE BETWEEN THE TIME OF THE INCIDENT AND WHEN FENWAL WAS NOTIFIED OF THE REPORTED EVENT. THE INSTRUMENT HISTORY WAS DOWNLOADED AND REVIEWED, HOWEVER, DO THE LENGTH OF TIME BETWEEN THE EVENT AND THE REPORT THE SPECIFIC DETAILS OF THE PROCEDURE WERE NOT AVAILABLE. THE AMICUS KIT THAT WAS USED DURING THE PROCEDURE WAS DISCARDED AND NO LOT AND UNABLE TO BE EVALUATED. FENWAL IS CONTINUING TO INVESTIGATE THE EVENT TO DETERMINE A POSSIBLE CAUSE AND WILL SUBMIT F/U WITH ANY ADD'L DETAILS. (B)(4). | 0
2342436 | PATIENT WAS ABOUT TO UNDERGO STRESS TEST. EKG LEADS WERE ATTACHED TO PATIENT. TREADMILL WOULD NOT OPERATE. PATIENT WAS REMOVED TO ANOTHER ROOM, STARTED ON ANOTHER TREADMILL FOR TEST. | 0
2342447 | THE MILLENNIA MONITOR IS USED TO REMOTELY MONITOR PATIENTS IN THE MRI ROOM. TWO (2) YEARS AGO THERE WAS AN UPGRADE ON THE MRI SIGNA LX 1.5T MACHINE AND THE COMPANY DID NOT INFORM THE HOSPITAL THAT WITH THE UPGRADE THERE WILL BE INTERFERENCE WITH THE MONITOR THUS HEART RHYTHM IS NOT CLEARLY MONITORED. THIS IS FOUND TO BE UNSAFE FOR CRITICALLY ILL PATIENTS BEING MONITORED DURING THE MRI PROCEDURE. | 1
2345337 | ELECTRONIC HEALTH RECORD FOR DRUG "BUPIVACAINE" DOES NOT HAVE SELECTABLE CHOICE UNDER "ROUTE OF ADMINISTRATION" FOR CONTINUOUS INFUSION BY PAIN PUMP. VOLUME OF THESE PAIN PUMPS IS 150 ML WHICH IS LETHAL DOSE OF BUPIVACAINE BY OTHER ROUTES FOR ADMINISTRATION OTHER THAN CONTINUOUS INFUSION BY PAIN PUMP. THIS PROGRAMMING DEFECT UNNECESSARILY EXPOSES PTS TO SEIZURES, CARDIAC ARREST AND LIKELY DEATH. CONVENTIONAL CPR IS INEFFECTIVE IN BUPIVACAINE INDUCED CARDIAC ARREST DUE TO DIRECT EFFECT OF BUPIVACAINE ON MYOCARDIAL MUSCLE. | 1
2352362 | ON (B)(6) 2011 AN UPPER ENDOSCOPY WITH MANUAL REMOVAL OF A DISPOSABLE BRAVO WIRELESS CAPSULE WAS PERFORMED AT THE PT'S REQUEST. DEVICE WAS INSERTED (B)(6) 2011 AND FAILED TO DISLODGE AND PASS THROUGH HER SYSTEM. | 0
2356690 | (B)(6) BEGAN USING A BAXTER HOMECHOICE AUTOMATED PERITONEAL DIALYSIS CYCLER ON OR ABOUT (B)(6) 2009, AND THEREAFTER EXHIBITED SIGNS AND SYMPTOMS SIMILAR TO THOSE DESCRIBED IN BAXTER'S (B)(6) 2010, CLASS I, RECALL NOTICE FOR THIS DEVICE. (B)(6) DIED ON (B)(6) 2010. | 0
2357898 | IT WAS REPORTED THAT THE MONITOR DID NOT ALARM FOR ASYSTOLE. A HEART RATE OF 40 WAS DISPLAYED. IT WAS REPORTED THAT THE PT DIED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2366406 | THIS MONITOR IS USED IN THE ICU. WHILE DOING A PROCEDURE ON A PT THE MONITOR BLANKED OUT. BIOMED WAS CALLED IN STAT TO REPAIR IT. THE POWER SUPPLY, HITRON MODEL HES49-12040, SPACELABS NUMBER (B)(4), FAILED. IT WAS REPLACED WITH ANOTHER ONE FROM SPACELABS. THIS IS THE 6TH POWER SUPPLY THAT HAS FAILED ON A PT THIS YEAR IN EMERGENCY ROOM, OPERATING ROOM AND ICU. THIS EQUIPMENT IS USED IN INTENSIVE CARE AND A FAILURE COULD PUT THE PT AT RISK AND/OR CAUSE THEM TO BE MOVED TO A NEW MONITORED BED, IF ONE IS AVAILABLE. PT ROOMS HAVE BEEN CLOSED DOWN UNTIL PARTS ARE REPLACED, SOME TIMES THE NEXT DAY WHEN STAFF IS AVAILABLE TO REPLACE PARTS. | 1
2371933 | (B)(4). UPON INDUCTION OF ANESTHESIA AND INTUBATION, IT WAS NOTED THAT THIS PT WAS NOT BEING VENTILATED APPROPRIATELY BY THE ANESTHESIA MACHINE. UPON FURTHER INVESTIGATION, THE CIRCUIT TUBING FOR THE MACHINE WAS NOTED TO BE CONNECTED INCORRECTLY (BELLOWS CONNECTED TO THE PT PORT AND PT CONNECTED TO THE BELLOWS PORT). ONCE THE CONNECTIONS WERE CORRECTED, APPROPRIATE VENTILATION WAS ACHIEVED. THE ANESTHESIOLOGIST CHECKED THE MACHINE AS SHE NORMALLY DOES BY HOLDING THE BELLOWS AND PUSHING THE O2 FLUSH ON THE MACHINE TO CHECK THE PRESSURE. EVERTHING SEEMED IN ORDER. SHE DID NOTE, HOWEVER THAT SHE COULD HEAR A LEAK ONCE SHE PLACED THE PT ON THE MACHINE. | 0
2371957 | THE BIOMERIEUX BACT/ALERT MICROBIAL DETECTION SYS PRESENTED WITH A MESSAGE "INSTRUMENT 2 IS NOT RESPONDING." THE CONTROLLER SCREEN ON INSTRUMENT 2 WAS COMPLETELY BLACK AND NO POWER WAS GOING TO THE CONTROLLER. THE INCUBATORS ASSOCIATED WITH THE CONTROLLER CONTAINED CULTURE BOTTLES FROM 172 PHORESIS COLLECTIONS. LOSS OF POWER TO THE CONTROLLER FOR THE CULTURE BOTTLES. CONTRIBUTING FACTORS MAY HAVE BEEN THAT THE POWER TO THE BACT CONTROLLER WAS DISRUPTED CAUSING THE INTERNAL UPS (BATTERY MODE) TO SUPPLY POWER. THE INTERNAL CIRCUITRY FAILED TO ALLOW THE UPS TO SWITCH BACK TO AC MODE WHEN THE MAIN POWER WAS RESUMED. THIS CAUSED THE BATTERY TO BE COMPLETELY DRAINED. | 0
2372013 | THE MEDICATION DOSAGE FORM VALUE WAS NOT SYNCHRONIZED BETWEEN THE FIRST DATABANK (FDB) DATA, THE COMMON VOCABULARY ENGINE (CVE), THE INVISION SERVICE MASTER, AND THE RELEVANT INVISION PROFILES. THIS RESULTED IN DOSING CLARIFYING INFORMATION, SUCH AS CR (CONTROLLED RELEASE), ER (EXTENDED RELEASE) AND ENTERIC COATED BEING OMITTED FROM THE PRINTED PRESCRIPTION. | 1
2372870 | IT WAS REPORTED THAT A COROMETRICS 174 MONITOR WAS USED ON A PT AND AN EMERGENCY CAESAREAN SECTION WAS PERFORMED. WHEN THE MONITOR WAS USED ON THE NEXT PT, THE MIDWIVES REPORTEDLY NOTED THAT THE PAPER SPEED WAS SET AT 3 CM/MIN, RATHER THAN THE EXPECTED 1 CM/MIN NORMALLY USED BY THE HOSP. THE HOSP STATED THAT THE PREOPERATIVE CHECKOUT WAS NOT PERFORMED ON THE UNIT PRIOR TO THE EVENT. THE HOSP HAS QUARANTINED THE MONITOR UNTIL THEIR INTERNAL INVESTIGATION IS COMPLETED. NO REPORTED DEVICE MALFUNCTION. THE UNIT HAD REPORTEDLY BEEN REPAIRED RECENTLY BY A GE SVC DEPOT AND RETURNED TO THE HOSP. THE SETTINGS ON THE COROMETRICS 174 WERE NOT CHECKED PRIOR TO REPAIR BECAUSE THE UNIT COULD NOT BE TURNED ON AT THE TIME IT WAS RECEIVED BY GE SVC DEPOT (THUS THE HOSP SETTING FOR PAPER SPEED WERE NOT VERIFIED). THE MAINS BOARD WAS REPLACED AND THE DEFAULT SETTINGS OF THE NEW MAINS BOARD WERE VERIFIED (PAPER SPEED 3 CM/MIN, PAPER SCALE 30 - 240). THE UNIT WAS RETURNED TO THE HOSP WITH THE DEFAULT SETTINGS. UNDER THE (B)(4) PRIVACY ACT, PT INFO IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSP. | 1
2374111 | ATTEMPTED TO INTUBATE DURING AN EMERGENCY EVENT. PATIENT WITHOUT RESPIRATORY EFFORT OR CARDIAC FUNCTION. DISPOSABLE LARYNGOSCOPE BLADE WOULD NOT LOCK INTO THE HANDLE PROPERLY. | 0
2374755 | IT WAS REPORTED THAT THE CLINICIAN OCCASIONALLY CREATES TRANSMISSIONS FOR PATIENTS AND THE TRANSMISSIONS ARE GONE WHEN THE CLINICIAN GOES BACK INTO THE PATIENT MANAGEMENT SYSTEM. IT WAS DETERMINED THAT THE WORKSTATION WAS CONNECTED TO A DIFFERENT DATABASE, SO THIS WAS CORRECTED BY CONNECTING THE WORKSTATION TO THE CORRECT DATABASE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION: THIS COMPLAINT WAS INADVERTENTLY SUBMITTED UNDER THE MEDICAL DEVICE REPORTING REGULATIONS AND THEREFORE IS BEING REDACTED, AS THE POTENTIAL FOR INJURY IS NOT LIKELY FOR THIS TYPE OF EVENT. | 1
2376457 | THE USER EXPERIENCED AN ISSUE WITH AN INCORRECT RESULT TRANSMITTING FROM THE OMNI COBAS B 221(6) BLOOD GAS ANALYZER SERIAL NUMBER (B)(4) TO THE DATACARE SOFTWARE AND INTO THE LABORATORY INFORMATION SYSTEM (LIS). THE PCO2 RESULT FROM THE ANALYZER WAS 126.6 MMHG AND THE RESULT IN THE DATACARE SYSTEM WAS 1269.6 MMHG. THIS RESULT WAS PASSED ON TO THE LIS. THE PATIENT WAS TREATED FROM THE PRINTED OFF RESULT FROM THE ANALYZER WHICH WAS THE CORRECT RESULT AND WAS NOT TREATED FROM INFORMATION IN THE DATACARE. THE PATIENT EXPIRED, BUT THE USER STATED THERE WAS NO DELAY IN TREATMENT RELATED TO THE RESULTS. THE PATIENT WAS NOT ADVERSELY AFFECTED BY THE EVENT. THE TECHNICAL SUPPORT SPECIALIST INSTRUCTED THE USER TO RESEND THE RESULT FROM THE ANALYZER TO DATACARE AND THE RESULT WAS SENT CORRECTLY. A ROOT CAUSE COULD NOT BE IDENTIFIED. THE EVENT COULD NOT BE DUPLICATED AND HAS NOT OCCURRED SINCE THE ORIGINAL EVENT. THERE WAS NOTHING IN THE ERROR LOG PROVIDED FOR INVESTIGATION THAT DEMONSTRATES THIS EVENT. CUSTOMER STATED THAT THE ANALYZER (OMNI SN (B)(4)) GENERATED A PCO2 RESULT OF 126.6 MMHG AND THE RESULT WENT TO DATACARE WHERE DATACARE INDICATED THE SAME RESULT WAS 126.6 MMHG WITH GENERIC DEMOGRAPHICS. AT THE TIME THE SAMPLE WAS RUN, THE PATIENT IDENTIFICATION WAS UNKNOWN SO THE CUSTOMER USED A GENERIC ID NUMBER OF (B)(4). THIS CAUSED THE SAMPLE TO BE HELD IN THE DATACARE QUEUE FOR APPROXIMATELY 3 HOURS UNTIL THE CUSTOMER ENTERED THE PROPER ID (DEMOGRAPHICS). WHEN THE DEMOGRAPHICS ENTERED BY THE CUSTOMER WERE THEN CHANGED TO THE PROPER PATIENT IDENTIFICATION AT DATACARE THAT IS WHEN THE RESULT OF 1269.6 MMHG WAS PASSED ON TO THE HOSPITAL (B)(6). THE PATIENT WAS NOT TREATED WITH THE INCORRECT RESULT THAT WAS ON THE DATACARE OR CUSTOMER LIS SYSTEM, THE PATIENT WAS TREATED WITH THE ANALYZER PRINTOUT WHICH WAS CORRECT. WHEN THE CUSTOMER RESENT THE SAME RESULTS TO THE DATACARE SYSTEM AGAIN THEY CROSSED CORRECTLY WITH ALL CORRECT RESULTS AND WITH THE PCO2 RESULT OF 126.6. CUSTOMER STATES BECAUSE THERE WAS NO TREATMENT GIVEN BASED ON THE INCORRECT RESULT.><. | 1
2378488 | IT WAS REPORTED THAT AFTER CHECKING FULL DISCLOSURE DATA FOR ALL TELEMETRY PTS, IT WAS DISCOVERED THAT AN ASYSTOLE EVENT HAD OCCURRED FOR ONE PT, BUT THERE WAS NO ALARM TRIGGERED AT THE CLINICAL INFO CENTER (CIC). THERE WAS REPORTEDLY A SIGNIFICANT AMOUNT OF INTERFERENCE IN THE ECG SIGNAL DUE TO ARTIFACT. THE PT WAS MOVED TO THE CRITICAL CARE UNIT (CCU) AND RECEIVED A PACEMAKER. NO REPORTED PT INJURY. THE PT WAS DISCHARGED FROM THE HOSPITAL ON (B)(6) 2011. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. UNDER (B)(4) LAW, PT INFO IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSPITAL. | 1
2379361 | I USE A HOME CHOICE PRO AUTOMATIC PD SYSTEM PROVIDED BY BAXTER HEALTH CARE CORP. RECENTLY THE FDA REQUIRED THEM TO CHANGE THEIR SOFTWARE PROGRAM TO PREVENT BYPASSING THE INITIAL DRAIN AND ALL OTHER DRAINS. DUE TO THIS CHANGE THE PT NOW MUST SUFFER THROUGH A PAINFUL PERIOD CAUSED BY THE MACHINE TRYING TO SUCK FLUID FROM THE PERITONEAL CAVITY WHEN NONE EXISTS. ACCORDING TO BAXTER THE MACHINE SHOULD ALARM AND ALLOW FOR A BYPASS BUT THIS DOES NOT HAPPEN THUS THE PT REMAINS IN EXTREME PAIN FOR SEVERAL MINUTES DURING THE BEGINNING OF THE PD CYCLE. BAXTER OR SOMEONE NEEDS TO REEVALUATE THIS SOFTWARE CHANGE AND NOT EXPERIMENT WITH THE PT IN A HOME ATMOSPHERE. THIS CHANGE IN THE SOFTWARE ALSO CAUSES THE MACHINE TO NOT PROVIDE THE PROPER THERAPY, BY TAKING AWAY COMPLETE CYCLES OF EXCHANGING FLUIDS. I CAN UNDERSTAND THE NEED FOR PREVENTING OVER FILLING, BUT NOT IN THE WAY BAXTER HAS TRIED TO ACHIEVE THIS, THEN THERE ONLY RESPONSE IS THAT THE FDA HAS REQUIRED THE CHANGE. CHANGE IS GOOD BUT IT MUST BE PROVEN CHANGE AND NOT AN EXPERIMENTAL PROGRAM TRIED ON THE AT HOME PATIENT. DIAGNOSIS OR REASON FOR USE: ESRD. | 0
2379398 | DURING SURVEY VISIT AND CODE PINK DRILL, SURVEYOR WITH SAFE PLACE BAND AND SENSOR ON TEST STATUS WAS ABLE TO ENTER THE ELEVATORS AND TAKE THEM ALL THE WAY DOWN WITHOUT THE ELEVATORS EVER LOCKING DOWN AS THE SYSTEM IS DESIGNED TO DO. EXPIRATION DATE: SENSORS 0912. | 0
2380645 | A CUSTOMER REPORTED THAT A CELL IN THEIR BACT/ALERT INSTRUMENT FAILED QC. WHILE INSPECTING THE CELL, THE CUSTOMER FOUND THE LOADED BOTTLE HAD CRACKED AND THE CONTENTS HAD LEAKED INTO THE CELL. ERROR CODE 60 "CELL FAILED QC" ALERTED THE CUSTOMER THAT THE CELL HAD MALFUNCTIONED. THE BACT/ALERT USER MANUAL INSTRUCTS THE USER TO IDENTIFY THE FAILED CELL, INSPECT THE CELL FOR DEBRIS AND REMOVE OR CLEAN IF NECESSARY, AND THEN TO RECALIBRATE THE CELL. THE CUSTOMER WAS NOT WEARING GLOVES WHILE PERFORMING THE QC CHECK AND WAS REQUIRED TO FILL OUT AN INCIDENT REPORT FOLLOWING THE HOSPITAL/LAB PROTOCOL. THE SAFETY PRECAUTIONS AND PROCEDURES SECTION IN THE BACT/ALERT USER MANUAL CONTAINS GENERAL PRECAUTIONS THAT RECOMMEND: "THE FOLLOWING PRECAUTIONS SHOULD BE OBSERVED DURING MAINTENANCE AND REPAIR, EVEN IN SITUATIONS WHERE A SPILL IS NEITHER OBSERVED NO SUSPECTED. AT A MINIMUM, DISPOSABLE GLOVES, EYE PROTECTION, AND A LABORATORY COAT SHOULD BE WORN." THE BACT/ALERT 3D SYSTEM IS DESIGNED TO LIMIT THE EFFECTS OF LEAKS AS THE USER MANUAL STATES: "EACH BOTTLE CELL WITHIN THE BACT/ALERT 3D INCUBATION MODULE IS SEALED TO HELP CONTAIN AND MINIMIZE EFFECTS OF LIQUID SPILLAGE." THE USER MANUAL ALSO REFERS THE CUSTOMER TO FOLLOW THE CLSI GUIDELINE M-29A "PROTECTION OF LABORATORY WORKERS FROM INSTRUMENT BIOHAZARDS AND INFECTIOUS DISEASE TRANSMITTED BY BLOOD, BODY FLUIDS, AND TISSUE." THE CUSTOMER IS UNSURE WHETHER THE BOTTLE WAS DROPPED OR IF THE PNEUMATIC TUBE SYSTEM WAS USED TO TRANSPORT THIS BOTTLE IN THEIR FACILITY. THE CUSTOMER REPORTED THE CRACK WAS ON THE BOTTOM OF THE BOTTLE. ALTHOUGH, THE CUSTOMER REPORT NO INJURY, THEY DID REPORT THE HEALTHCARE WORKER WHO CHECKED THE FAILED QC CELL SOUTH MEDICAL TREATMENT, PER THEIR LABORATORY PROCEDURE, TO PRECLUDE EXPOSURE TO THE BOTTLE CONTENTS. THE HEALTHCARE WORKER HAD TESTS FOR HCAB, HBAG AND HIVSC, AND ALL RESULTS WERE NEGATIVE. THE CUSTOMER IS NOT WILLING TO DISCLOSE ANY ADDITIONAL INFO ABOUT THE HEALTHCARE WORKER'S TESTS, AND THE CUSTOMER IS NOT ABLE TO DISCLOSE WHAT MICROORGANISM THE PATIENT WAS BEING TESTED FOR BECAUSE, THE SAMPLE WAS RECEIVED FROM ANOTHER HOSPITAL AND THE INFO IS PROTECTED BY HIPAA REGULATIONS. AN INVESTIGATION HAS BEEN OPENED TO EXAMINE THE BOTTLE CRACK. THERE HAVE BEEN NO OTHER BOTTLE BREAKS FOR THIS LOT. AS THERE ARE NO SIMILAR COMPLAINTS WITH THIS LOT, THERE IS NO INDICATION THAT THIS ISSUE WILL RECUR. THIS EVENT IS REPORTABLE TO THE FDA AS AN ADVERSE EVENT BECAUSE, THE CUSTOMER HAS TO SEEK MEDICAL TREATMENT TO DETERMINE EXPOSURE TO THE INOCULATED BOTTLE MEDIA. NO REPORT OF MEDICAL INJURY HAS BEEN RECORDED AT THIS TIME. | 0
2381324 | THE PT WAS HAVING AN MRI SCAN OF THE RIGHT FEMUR UTILIZING A PHILIPS INTERA 1.5T SCANNER WITH THE SENSE BODY COIL. PROPER POSITIONING AND SAFETY TECHNIQUES WERE UTILIZED DURING THE SCAN. BARRIERS, SUCH AS SPONGES AND BLANKETS WERE USED TO PREVENT THE PT COIL FROM TOUCHING THE PT, AS WELL AS CORRECT BODY WEIGHT ENTRY INTO THE MACHINE FOR SAR LEVELS. THE CABLES TO THE COIL WERE NOT KINKED, WRAPPED, OR COILED, THEY RAN FROM THE COIL OUT OF THE MACHINE WHERE THEY PLUGGED IN. AFTER THE EXAM, THE PT COMPLAINED OF A BURNING SENSATION ON HIS ANTERIOR LOWER LEG. THE PT WAS EXAMINED BY THE RADIOLOGIST AS WELL AS AN INTERNIST/ GENERAL PHYSICIAN. HE WAS DIAGNOSED WITH A SMALL 1ST DEGREE BURN MEASURING 1 X 4 CM BY THE INTERNIST WITH THE RADIOLOGIST DESCRIBING THE AREA, AS WELL AS THE SURROUNDING ERYTHEMA, TO MEASURE IN TOTAL 15 CM. THE PT WAS TREATED WITH ANTIBACTERIAL OINTMENT. | 0
2382177 | A CUSTOMER REPORTED A SN BOTTLE BROKE AND LEAKED INTO THE INSTRUMENT. WHEN REMOVING THE BOTTLE FROM THE INSTRUMENT, A SPILL OCCURRED AND INOCULATED BLOOD CAME IN CONTACT WITH THE CUSTOMER'S HANDS AND LAB COAT. THE CUSTOMER INDICATED THEY FOLLOWED PROTOCOL INSPECTING THE BOTTLE FOR ANYTHING UNUSUALLY PRIOR TO LOADING, BUT WERE UNSURE IF THERE WERE ANY CRACKS IN THE BOTTLE AT THE TIME OF LOADING. THIS CUSTOMER IS UNSURE HOW THE BOTTLE WAS TRANSFERRED TO THE LAB; HOWEVER, THIS CUSTOMER HAD A PNEUMATIC TUBE SYSTEM TO DELIVER BOTTLES FROM FLOOR TO FLOOR AND THAT MOST BOTTLES ARE SENT VIA THIS SYSTEM. THE SPILL OCCURRED WHEN UNLOADING THE BOTTLE. THE CUSTOMER DID NOT FOLLOW GENERAL LABORATORY PROCEDURE TO WEAR GLOVES PRIOR TO HANDLING THE INOCULATED BOTTLE. PER PACKAGE INSERT INSTRUCTIONS "USE DISPOSABLE GLOVES AND HANDLE INOCULATED BOTTLES CAUTIOUSLY AS THOUGH CAPABLE OF TRANSMITTING INFECTIOUS AGENTS." THE CUSTOMER ACKNOWLEDGES THAT LAB TECH WAS AWARE THE BOTTLE WAS INOCULATED AND BROKEN AND CHOSE NOT TO WEAR GLOVES. THE CUSTOMER FOLLOWED LABORATORY AND BIOMERIEUX RECOMMENDED SPILL PROCEDURE COVERING THE SPILL WITH PAPER TOWELS AND CLEANING THE SPILL WITH 10% BLEACH SOLUTION. NO OTHER BOTTLE BREAKS FROM THIS SN BOTTLE LOT HAVE BEEN REPORTED AT THIS TIME. FOUR SETS OF BLOOD CULTURES WERE TESTED AND RESULTS WERE NEGATIVE FOR HEPATITIS AND HIV. THE CUSTOMER DOES NOT KNOW WHAT DISEASE WAS BEING TESTED FOR THE INOCULATED SN BOTTLE. AN INVESTIGATION HAS BEEN OPENED TO EXAMINE THE BOTTLE BREAK. A REQUEST TO RETURN THE PRODUCT HAS BEEN MADE AND A BIOHAZARD-SHIPPING CONTAINER HAS BEEN SENT TO THE CUSTOMER. THERE HAVE BEEN NO OTHER BOTTLE BREAK REPORTS FOR THIS LOT. AS THERE ARE NO SIMILAR COMPLAINTS WITH THIS LOT, THERE IS NO INDICATION THAT THIS ISSUE WILL REOCCUR. THIS EVENT IS REPORTABLE TO THE FDA AS AN ADVERSE EVENT BECAUSE THE CUSTOMER HAD TO SEEK MEDICAL TREATMENT TO DETERMINE EXPOSURE TO THE INOCULATED BOTTLE MEDIA. NO REPORT OF MEDICAL INJURY HAS BEEN RECORDED AT THIS TIME. AT THE TIME OF THE ORIGINAL COMPLAINT, THE BOTTLE WAS BELIEVED TO BE DISPOSED. UPON FURTHER FOLLOW-UP WITH THE MICROBIOLOGY LAB MGR, IT WAS DETERMINED THE BOTTLE WAS STILL AVAILABLE. THE BOTTLE WAS REQUESTED TO BE RETURNED TO BIOMERIEUX AND A BIOHAZARD SHIPPING CONTAINER WAS SENT TO THE HOSPITAL FOR EASY PACKAGING. AS OF (B)(6), 2011, THE BOTTLE HAS NOT YET ARRIVED TO BIOMERIEUX. UPON RECEIPT, THE DEVICE INVESTIGATION WILL BE PERFORMED AND A SUPPLEMENTAL REPORT WILL BE FILED. AN INVESTIGATION HAS BEEN PERFORMED ON RETAINED BOTTLES FROM THE AFFECTED LOT: MFG AND PACKAGING REVIEW: BAC/ALERT SN LOT 1027916, PN 259790 WAS MFG ON (B)(6) 2011. A REVIEW OF THE MFG AND PACKAGING RECORDS SHOWED THAT THIS LOT WAS MFG ACCORDING TO PROCEDURE. ACCORDING TO THE PACKAGING DEFECT LOT (MG-PK. 0434), ALL ATTRIBUTES MET SPECS. ONE NCMR'S WAS ISSUED FOR THIS LOT BUT WOULD NOT HAVE IMPACTED THIS EVENT. QUALITY CONTROL RECORD REVIEW: BACT/ALERT SN LOT 1027916 MET ALL QC RELEASE C SPEC REQUIREMENTS. QUALITY CONTROL AQL INSPECTION PASSES ALL SPEC REQUIREMENTS. NO OOS INCIDENTS WERE ISSUED FOR THIS LOT. RETAIN AND RETURN SAMPLE INSPECTION: THE BROKEN BOTTLE FROM LOT 1027916 WAS RETURNED FROM THE CUSTOMER TO BIOMERIEUX QA FOR EVAL. THREE HUNDRED RETENTION SAMPLES WERE INSPECTED FOR BOTTLES WITH CRACKS. NO BOTTLES WERE FOUND TO HAVE CRACKS. TO MITIGATE AGAINST LEAKS, THE SN BOTTLE PACKAGE INSERT STATES: "VISUALLY INSPECT BOTTLES BEFORE TESTING. DO NOT USE BOTTLES WITH EVIDENCE OF DAMAGE, LEAKAGE, OR DETERIORATION." IN ADDITION, THE BACT/ALERT 3D USER MANUAL STATES "IF AN INOCULATED BOTTLE IS FOUND TO BE LEAKING OR IS ACCIDENTALLY BROKEN DURING COLLECTION OR TRANSPORT, USE THE ESTABLISHED PROCEDURES IN YOUR FACILITY FOR DEALING WITH BIOHAZARDOUS MATERIAL." THE BACT/ALERT 3D SYSTEM IS DESIGNED TO LIMIT THE EFFECTS OF LEAKS AS THE USER MANUAL STATES: "EACH BOTTLE CELL WITHIN THE BACT/ALERT 3D INCUBATION MODULE IS SEALED TO HELP CONTAIN AND MINIMIZE EFFECTS OF LIQUID SPILLAGE." THE USER MANUAL ALSO INDICATED THAT, "A DRIP TRAY IS INCORPORATED AT THE BOTTOM OF EACH DRAWER BENEATH THE OPENING END OF THE RACKS TO MINIMIZE THE EFFECT OF ANY SPILLAGE OF LIQUID." THE CUSTOMER CONFIRMS THAT THE BROKEN BOTTLE LEAKED INTO ONLY CELL 4D17 AND ONTO THE BOTTOM PANEL (DRIP TRAY). THIS CUSTOMER STATES THAT 80 PERCENT OF THE BLOOD CULTURE BOTTLES ARE SENT THROUGH A PNEUMATIC TUBE SYSTEM AND THE OTHERS ARE HAND DELIVERED; HOWEVER, THE CUSTOMER IS UNSURE HOW THE BOTTLE WAS DELIVERED TO THE LAB. A DROP TEST STUDY WAS PERFORMED IN (B)(6) 2010 THE STUDIES CONCLUDED: PRODUCT MADE 1 YR AGO HAS THE SAME PROBABILITY OF BREAKING AS THE COMPLAIN LOTS OF BOTTLES. CURRENT BOTTLES ARE NO DIFFERENT IN THEIR RESISTANCE TO BREAKAGE THAN OLDER BOTTLES. BREAKS/CRACKS OCCURRED AT A HIGHER RATE IN THE HIGH VACUUM BOTTLE (SN) AND A LESSER RATE IN THE LOW VACUUM BOTTLE (FA). THE 78 CC PLASTIC BOTTLES ARE NOT BREAK PROOF, BUT WERE DESIGNED TO BE BREAK RESISTANT. | 0
2382425 | NURSE WAS WORKING WITH THE MONITOR AND THE POLE AND WHEN THE MONITOR TILTS FORWARD IT PINCHED THE NURSE'S FINGER. THE MODEL NUMBER OF THE MONITORS IS DASH 4000, THE AGE RANGE OF THE MONITORS IN OB IS FROM 1 TO 9 YEARS. THE STANDS WERE PURCHASED FROM GE. ON THESE STANDS, THE MOUNTING PLATE THAT HOLDS THE MONITOR ON TOP OF THE VERTICAL POLE PIVOTS ON A HINGE PIN TO ALLOW CHANGING THE VIEWING ANGLE OF THE MONITOR. THERE IS A PINCH POINT AT THE HINGE PIN BETWEEN THE TOP OF THE VERTICAL POLE AND THE BOTTOM OF THE MOUNTING PLATE. | 0
2382485 | ISMP MEDICATION ERRORS REPORTING PROGRAM. COMPUTER SYSTEM DISPLAY ISSUE. ROUTES ARE ENTERED BY PHARMACY USING ABBREVIATIONS (IV, IM, NG, IT, ETC). THE ABBREVIATION NG IS USED TO INDICATE NASAL-GASTRIC TUBE. HOWEVER, ON A SEPARATE SCREEN WITHIN THE COMPUTER SYSTEM IS A RN ORDER REVIEW PAGE. THIS PAGE SPELLS OUT THE ENTIRE ROUTE NAME. THE COLUMNS CAN BE ADJUSTED TO SHOW MORE OR FEWER CHARACTERS. THE DEFAULT COLUMN CUTS OFF "NASAL-GASTRIC TUBE" TO BE "NASAL...". THIS LED AN RN TO THINK THAT AN ORAL MEDICATION SHOULD BE GIVEN VIA A NASAL INHALATION ROUTE. MEDICATION ADMINISTERED TO OR USED BY THE PT: NO. TYPE OF STAFF MADE INITIAL ERROR: NURSE. PT COUNSELING PROVIDED: UNK. (B)(6). | 1
2382487 | ISMP MEDICATION ERRORS REPORTING PROGRAM: A PERSISTENT COMPUTER SOFTWARE ISSUE PRODUCED TWO SERIOUS (NEAR-MISS) MEDICATION ERRORS IN WHICH NURSES IMPROPERLY REQUESTED (AMBULATORY CARE) CHEMOTHERAPY DOSES ON THE INPATIENTS FOR WHOM THEY WERE PROVIDING CARE. PHARMACY HAS REQUESTED THAT IT CONTACT OUR COMPUTER SOFTWARE VENDOR FOR HELP CORRECTING THIS UNSAFE CONDITION: OUTPATIENT MEDICATION AND INFUSION ORDERS INAPPROPRIATELY APPEAR ON THE INPATIENT MEDICATION ADMINISTRATION RECORD (MAR). THE INPATIENT MEDICATION ADMINISTRATION RECORD MUST DISPLAY ONLY MEDICATIONS ORDERED FOR THAT PT ADMISSION. OUTPATIENT MEDICATION ORDERS AND AMBULATORY CARE CHEMOTHERAPY AND INFUSION ORDERS MUST NOT APPEAR ON THE INPATIENT MAR. THE MEDICAL DIRECTOR OF HEMATOLOGY AND ONCOLOGY HAS EXPRESSED EXTREME CONCERN ABOUT THIS DANGEROUS SITUATION. MEDICATION ADMINISTERED TO OR USED BY THE PT: NO. TYPE OF STAFF MADE INITIAL ERROR: NURSE. PT COUNSELING PROVIDED: UNK. RELEVANT MATERIALS PROVIDED: NONE. (B)(4). | 1
2382510 | (B)(6) MEDICATION ERRORS REPORTING PROGRAM: OUR MED/SURG WARD UTILIZES REMOTE TELEMETRY MONITORING DEVICES TO MONITOR CONTINUOUS EKG AND/OR OXIMETRY. IT IS A BOX WITH ELECTRODES ON THE PT, THAT SUBMITS INFO WIRELESSLY ON ANOTHER FLOOR OF THE HOSPITAL, WHERE IT IS MONITORED BY A TECHNICIAN. AT LEAST ONCE A SHIFT, WE RECEIVE A PRINT OUT OF THE EKG THAT HAS BEEN READ TO INFORM US OF THE RHYTHM. IF THERE IS A CHANGE IN THE ECG OR HEART RATE, THAT UNIT CALLS OUR UNIT, AND THE NURSE ON OUR WARD IS NOTIFIED BY THE NURSE'S STATION OF THE CHANGE. WHILE WE DO HAVE A SMALL MONITORING SCREEN ON OUR WARD WITH THE ECG READINGS, NO ONE IS THERE TO MONITOR IT, AND NURSES ON OUR FLOOR ARE LIMITED IN THEIR SCOPE OF PRACTICE IN TERMS OF WHICH CARDIAC MEDS WE CAN GIVE IV. IF A PT CONDITION CHANGES THAT REQUIRES INCREASED MONITORING AND/OR THAT THEY RECEIVE THESE DRUGS, A UNIT TRANSFER IS REQUIRED. A NURSE WAS ASSUMING CARE OF A PT XYZ BEING MONITORED ON ONE OF THESE REMOTE TELEMETRY BOXES. DURING THE USUAL BEGINNING OF SHIFT ASSESSMENT, THE NURSE LOOKED AT THE BOX, ENSURED ALL OF THE ELECTRODES WERE IN THE RIGHT PLACE, THE BOX WAS ON, AND HAD BATTERIES. IN RETROSPECT, THIS NURSE SHOULD HAVE TAKEN THIS ASSESSMENT ONE STEP FURTHER, AND CHECKED THE TELE BOX NUMBER TO THE MONITOR AT THE NURSES STATION. THE CHARGE NURSE, ABOUT AN HOUR LATER, WAS DOING HER USUAL CHECK OF THE TELEMETRY LOG, AND NOTICED THAT PT XYZ'S NAME WAS NOT ON THE MONITOR AT THE NURSE'S STATION. BY CHECKING THE BOX NUMBER, NOT THE PT NAME, IT WAS DISCOVERED THAT THE NAME ON THE SCREEN FOR THAT PARTICULAR BOX WAS FOR SOMEONE ELSE ON A COMPLETELY DIFFERENT FLOOR, AND NOT PT XYZ AFTER ALL. FROM WHAT I COULD BEST GATHER, ON THE PREVIOUS SHIFT, OUR UNIT WAS GIVEN AN ADMISSION THAT HAD REMOTE TELEMETRY ORDERED. THE TRANSFERRING NURSE FAILED TO REPORT THAT TELEMETRY WAS REQUIRED ON THE RECEIVING UNIT. AT THAT PARTICULAR TIME, ALL OF THE TELEMETRY BOXES IN THE ENTIRE HOSP WERE IN USE. THE NURSING SUPERVISOR WAS NOTIFIED, WHO THEN WENT SEARCHING FOR A BOX. HE OR SHE BROUGHT A BOX FROM ANOTHER FLOOR. APPARENTLY, THE BOX WAS FROM A PT 123 WHO WAS IN SURGERY, AND THE BOX WAS LEFT AT THE BEDSIDE, WAITING FOR PT 123'S RETURN. WHEN THE SUPERVISOR PICKED UP THE BOX, SHE DIDN'T CALL THE MONITORING UNIT TO INFORM THEM OF THE PT SWITCH, THAT A NEW PT WAS GOING TO BE MONITORED. THE BOX WENT BACK ON PT XYZ, BUT UNDER THE PREVIOUS PT 123'S NAME. HERE IS WHERE THE MED ERROR CAME IN: THE PT 123 THEN CAME BACK FROM SURGERY. IT IS UNCLEAR IF PT 123 HAD A TELEMETRY BOX ON OR NOT. HOWEVER, PT XYZ'S EKG RHYTHM CHANGED TO RAPID AFIB IN THE 150'S. HOWEVER, THE MONITOR TECH CALLED THE NURSE ON THE OTHER UNIT, BASED ON THE PT'S 123 NAME ON THE SCREEN. PT 123 RECEIVED DIGOXIN IVP, BASED ON PT XYZ'S RHYTHM. MEANWHILE, XYZ, WHO WAS ACTUALLY IN AFIB IN THE 150'S, DID NOT RECEIVE THIS MUCH NEEDED MEDICATION UNTIL SEVERAL HOURS LATER, AFTER THE ERROR WAS DISCOVERED, ONLY BY A CHARGE NURSE WHO WAS UPDATING HER MONITORING LOG Q SHIFT. PT 123 RECEIVED A MEDICATION NOT INTENDED FOR THEM, AND PT XYZ SUFFERED FROM A DELAY IN MEDICATION ADMINISTRATION BECAUSE OF THIS TELEMETRY BOX MIXUP. ONE COULD EXTRAPOLATE THAT FURTHER MEDICATION OR TREATMENT ERRORS HAD POTENTIAL TO OCCUR HAD THIS ERROR NOT BEEN DISCOVERED WHEN IT DID. MEDICATION ADMINISTERED TO OR USED BY THE PT: YES. WHEN AND HOW WAS ERROR DISCOVERED: MEANWHILE, XYZ, WHO WAS ACTUALLY IN AFIB IN THE 150'S DID NOT RECEIVE THIS MUCH NEEDED MEDICATION UNTIL SEVERAL HOURS LATER, AFTER THE ERROR WAS DISCOVERED, ONLY BY A CHARGE NURSE WHO WAS UPDATING HER MONITORING LOG Q SHIFT. LEVEL OF STAFF WHO DISCOVERED THE ERROR: NURSE. (B)(6). | 1
2383997 | IT WAS REPORTED THAT A PT INCIDENT OCCURRED. THE PT REPORTEDLY CODED. NO ADD'L INFO WAS PROVIDED BY THE HOSPITAL. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2385124 | PATIENT WAS ON AN ANTI-ARRHYTHMIC IV INFUSION FOR ATRIAL FIBRILLATION. PATIENT'S MONITOR WAS SET TO ALARM IF PATIENT'S HEART RATE WENT BELOW 90 BEATS PER MINUTE. PATIENT'S HEART RATE DROPPED TO 70 WHEN PATIENT CONVERTED TO SINUS RHYTHM BUT THE ALARM DID NOT GO OFF TO ALERT THE NURSE THAT HEART RATE DROPPED. BIOMED WAS NOTIFIED OF THE INCIDENT. TELEMETRY BOX WAS SENT TO MANUFACTURER FOR INVESTIGATION. | 1
2385125 | PATIENT HAD SEVEN (7) RUNS OF VENTRICULAR BEATS AND THE MONITOR WAS SUPPOSED TO ALARM BUT DID NOT. EARLIER THE MONITOR SHOWED NINE (9) RUNS OF VENTRICULAR BEATS AND THE ALARM WENT OFF. | 1
2386598 | PERFORMED AN INTERNAL CHART REVIEW BECAUSE PT'S DEATH WAS NOT EXPECTED AND THE SUDDEN DETERIORATION IN THIS PT'S CONDITION REVEALED AN ERROR THAT AN INCORRECT DOSING OF ELECTROLYTES RESULTING IN ELECTROLYTE ABNORMALITIES SEVERE ENOUGH TO CAUSE DEATH. INTERNAL INVESTIGATION REVEALED DATA WAS INCORRECTLY ENTERED INTO THE COMPOUNDING MACHINE. THE RATE FIELD CONSISTS OF THE VALUE, NUMERATOR AND DENOMINATOR. THE ERROR WAS MADE IN THE DENOMINATOR FIELD, WHERE THE DROP DOWN SELECTION WAS CHANGED FROM THE DEFAULT OF /HOUR TO /KG/DAY. | 1
2388582 | PATIENT REQUIRED VENTILATOR SUPPORT. WHEN ATTEMPTING TO SET UP THE EQUIPMENT INCLUDING THE BREATHING CIRCUIT IT WAS NOTED THAT THE PACKAGE LABEL WAS NOT CORRECT. PACKAGE STATED "PEDIATRIC BREATHING CIRCUIT" BUT WAS ACTUALLY "ADULT" CIRCUIT. INACCURATE LABELING OF CRITICAL PRODUCT COULD LEAD TO DELAY IN PATIENT MANAGEMENT. | 0
2388627 | THIS PT WAS UNDERGOING RADIATION THERAPY FOR ADENOCARCINOMA OF THE RECTUM. HIS TREATMENT PLAN WAS A SEVEN FIELD IMRT TO THE PELVIS AND RECTUM. THE TOTAL WAS 25 FRACTIONS FOR A TOTAL PRESCRIPTION DOSE OF 4500 CGY. THE PT BEGAN HIS TREATMENT ON 8/1 AND HAD RECEIVED 10 FRACTIONS WITHOUT INCIDENT. ON 8/12, THE RADIATION THERAPIST NOTED THAT THE TOTAL NUMBER OF MLC POINTS HAD CHANGED. AN INVESTIGATION REVEALED A COMPUTER FILE CORRUPTION WITH THE MOSAIQ SOFTWARE DATA BASE. THIS CAUSED THE INFO PARAMETERS IN FIELD ONE TO BE REPLACED WITH THE INFO PARAMETERS IN FIELD TWO. THE INFO IN FIELD TWO AND THE REMAINING FIVE TREATMENT FIELDS WERE NOT AFFECTED BY THE FILE CORRUPTION. THE CORRUPTED FILE WAS DISCOVERED ON 8/12 AND AFFECTED ONLY THE TREATMENT RENDERED ON 8/11 ONLY. THIS WAS REPORTED TO THE MFR AT THE TIME IT WAS DISCOVERED. THERE WAS NO ADVERSE OUTCOME TO THE PT. | 1
2389164 | IT WAS REPORTED THAT A TECHNICIAN WAS ASSISTING IMAGING A PT IN A WHEELCHAIR. TUBE ARM WAS ROTATED TO 80 DEGREES. ASSISTANT WAS PRESENT IN THE ROOM TO HELP DURING THE EXAM. THE ASSISTANT WAS USING THE SWIVEL-ARM OPERATION KEY THAT WAS FACING UPWARD. THE TECHNICIAN WAS BELOW THE SWIVEL ARM POSITIONING THE PT WHEN THE CONTROL PANEL ON THE TUBE ARM CAME IN CONTACT WITH HER SHOULDER/UPPER ARM AREA. THE BUTTON TO LOWER THE TUBE ARM GOT ACTIVATED BY THIS ACCIDENTAL CONTACT CAUSING THE TUBE ARM TO COMPRESS THE TECHNICIAN'S ARM. THE ASSISTANT HAD TO PRESS THE EMERGENCY STOP BUTTON AND THE TECHNICIAN WAS RELEASED. THE EXAM OF THE PT WAS COMPLETED WITHOUT FURTHER INCIDENTS. SIEMENS LOCAL SERVICE ENGINEER, (B)(6), CONFIRMED THAT THE REPORTED EVENT CAUSED BRUISING AND MUSCLE STRAIN OF THE RIGHT SHOULDER AND UPPER ARM OF THE TECHNICIAN. THE TECHNICIAN WAS TREATED WITH PAIN MEDICATION AND PHYSICAL THERAPY. THE REPORTED SYSTEM IS WORKING WITHIN SPECS. THE RISK OF COLLISION AND CRUSHING IS DESCRIBED IN THE OPERATOR MANUAL FOR MAMMOMAT INSPIRATION (B)(4) IN CHAPTER "PERSONAL SAFETY". POSSIBLE DANGER ZONES ARE MARKED. ALSO, A WARNING FOR DANGER OF CRUSHING AND JAMMING OF BODY PARTS BY MOTOR-DRIVEN SYSTEM MOVEMENTS IS INCLUDED IN THE OPERATOR MANUAL. PMA P030010/S6. | 0
2390232 | MODULES WERE FOUND TO HAVE HOLES WITH EVIDENCE OF BURNING IN THE MORNING HOURS WHEN STAFF ARRIVED. THE HOLES ARE BELIEVED TO HAVE BURNED OVERNIGHT AND NONE OCCURRED DURING A PROCEDURE. | 0
2390246 | AT THE BEGINNING WE STARTED EXPERIENCING LONG PERIODS OF INTERFERENCE ON OUR PHILIPS TELEMETRY SYSTEM. WE WERE UNABLE TO READ CARDIAC RHYTHMS OF OUR PATIENTS ON TELEMETRY BECAUSE WAVEFORMS PRODUCED FROM TELEMETRY PACKS WERE NOT RELIABLE; THEY KEPT DROPPING OFF THE SCREEN. WE STOPPED USING THE MAJORITY OF THE SYSTEM AND WERE REQUIRED TO USE LOANER TELEMETRY MONITORS. THIS OCCURRED OVER THE COURSE OF THREE DAYS. WE DECLARED AN INTERNAL HOSPITAL EMERGENCY AND OUR ED WAS PLACED ON DIVERT. PHILIPS REPRESENTATIVES WERE ON SITE AND COULD NOT DETERMINE THE CAUSE OF INTERFERENCE. AS WE EXPERIENCED DECREASED PERIODS OF INTERFERENCE WE STARTED PUTTING PATIENTS BACK ON THE SYSTEM. THE ENTIRE SPECTRUM WAS AFFECTED WHICH INCLUDED 144 PATIENTS.======================MANUFACTURER RESPONSE FOR MONITOR, TELEMETRY SYSTEM (BRAND NOT PROVIDED) (PER SITE REPORTER).======================THEY WERE ON SITE FOR FIVE DAYS TESTING THEIR SYSTEM. THEY STATED THEY FOUND NO PROBLEMS WITH THE SYSTEM. | 1
2392193 | MULTIPLE TIMES ON (B)(6) 2007, A PATIENT HAD EPISODES OF VENTRICULAR TACHYCARDIA TERMINATED BY AN AICD DEFIBRILLATING THE PATIENT INTO A PACED RHYTHM. THE WELCH ALLYN ACUITY MONITORING SYSTEM FAILED TO ALARM CORRECTLY FOR 4 OCCASIONS. INSTEAD OF ALARMING FOR V-TACH. IT READ IT AS EITHER; TACHYCARDIA, NOISE, MISSED BEAT, OR COUPLET. | 1
2392209 | PATIENT ON CARDIAC TELEMETRY FOR SYNCOPE ON (B)(6) 2007. PATIENT HAD AN EPISODE OF VENTRICULAR STANDSTILL FOR APPROXIMATELY 12 SECONDS. WELCH ALLYN ACUITY MONITORING SYSTEM DID NOT ALARM FOR THIS EVENT, INSTEAD IT ONLY READ IT AS A MISSED BEAT WHICH IS NOT A LIFT THREATENING EVENT. THIS ARRHYTHMIA WAS NOT NOTICED UNTIL PATIENT WAS READMITTED APPROXIMATELY 3 HOURS AFTER DISCHARGE ON (B)(6) 2007 FOR ANOTHER EPISODE OF SYNCOPE. | 1

## 2012
1660997 | IT WAS REPORTED THAT A PT WITH METASTATIC BREAST CANCER HAD A CT SCAN OF HER CHEST, ABDOMEN AND PELVIS. SHE WAS RECEIVING DOXIL (LIPOSOMAL DOXORUBICIN A RADIOSENSITIZER). SHE DEVELOPED WHAT WAS DESCRIBED BY THE PT AS A SEVERE "SUNBURN" WITH PEELING FROM THE NECK DOWN TO THE PELVIS. THE PHYSICIAN SAW THE PT APPROX ONE WEEK LATER AND THE PT PRESENTED WITH A "SUNTAN." GE HEALTHCARE'S INVESTIGATION SHOWED THE CT DOSE PROVIDED TO THE PT (CTDIVOL OF 25.82 MGY) WAS NOT EXCESSIVE AND NOT EXPECTED TO GENERATE SKIN EFFECTS SOLELY FROM THE RADIOLOGICAL EXAM. THE INVESTIGATION CONCLUDED THAT THE PT'S SKIN CONDITION TO BE RADIATION RECALL, A KNOWN SKIN REACTION THAT CAN OCCUR AFTER RADIATION (IONIZING AS WELL AS ULTRAVIOLET). DOXIL, THE DRUG THAT WAS GIVEN TO THE PT PRIOR TO THE EXAM, IS WELL KNOWN IN THIS CONTEXT AND IS DOCUMENTED IN THE DRUG LABELLING. THE SYS WAS EVALUATED AND IS OPERATING WITHIN SPECS. | 0
1730589 | THIS MAY BE A POSSIBLE FALSE NEGATIVE. FOV 3,6,7,9 AND 19 SHOWED POSSIBLE REACTIVE CHANGES AND ABNORMAL CELLS. NO AUTOSCAN WAS PROMPTED. HOWEVER, THERE WAS NO DELAY IN PATIENT DIAGNOSIS BECAUSE, THIS WAS A HIGH RISK PATIENT AND THIS WAS CAUGHT DURING A QUALITY CHECK. CUSTOMER IS NOT SENDING THE SLIDE FOR HOLOGIC REVIEW. | 0
1730601 | THIS MAY BE A POSSIBLE FALSE NEGATIVE. FOV 20 SHOWED POSSIBLE REACTIVE CHANGES AND ABNORMAL CELLS. NO AUTOSCAN (FULL SCAN OF THE SLIDE) WAS PROMPTED. HOWEVER, THERE WAS NO DELAY IN PATIENT DIAGNOSIS AS THE DIAGNOSTIC CELLS WERE FOUND DURING QC TESTING. CUSTOMER IS NOT SENDING SLIDE TO HOLOGIC FOR REVIEW, | 0
1731410 | ABNORMAL CELLS FOUND OUTSIDE OF THE 22 FIELDS OF VIEW (FOV). NO TRIGGER CELLS WERE PRESENT TO PROMPT A FULL SCAN OF THE SLIDE. THERE WAS NO DELAY IN PATIENT DIAGNOSIS AS THE DIAGNOSTIC CELLS WERE FOUND DURING QC. CUSTOMER IS NOT SENDING SLIDE TO HOLOGIC FOR REVIEW. | 0
1812771 | ON (B)(6) 2010, A CUSTOMER REPORTED THAT MEASUREMENTS FOR A LESION WERE DIFFERENT IN MAGNIFIED VIEW VERSUS REGULAR VIEW WHEN MEASURED IN ISITE PACS USING FUJI CR MAGNIFICATION. REGULAR VIEW SHOWS THE CORRECT MEASUREMENT. A LESION MEASURED IN ISITE PACS FROM SIEMENS MAMMOGRAPHY DR UNIT EQUALED 10MM. THE SAME LESION MEASURED ON STANDARD VIEW FROM FUJI CR EQUALED 10MM AND ON MAGNIFICATION VIEW FROM FUJI CR EQUALED 16MM. ISITE FOLLOWS THE DICOM STANDARD, SECTION 10.7 - BASIC PIXEL SPACING CALIBRATION MACRO WHEN CALCULATING MEASUREMENTS. FUJI IS NOT SENDING ANY DICOM INFORMATION FOR MAGNIFICATION OR CORRECTION OF PIXEL SPACING VALUES TO ACCOUNT FOR MAGNIFICATION IN THE MAGNIFIED VIEW IMAGES. THERE IS A MAGNIFICATION VALUE (0018, 1114) INCLUDED IN THE NON-MAGNIFIED VIEW OF THE FUJI STUDY. ISITE PACS IS A SOFTWARE PRODUCT. WE ARE ABLE TO EVALUATE THE PRODUCT AT THE CUSTOMER SITE BY REMOTE ACCESS, AS WELL AS EVALUATE THE SAME PRODUCT VERSION INHOUSE. BASED ON THE AVAILABLE INFORMATION AT THE TIME OF THIS REPORT, WE CANNOT CONFIRM THAT THE DEVICE WAS A FACTOR IN THE INCIDENT. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION REGARDING THIS ISSUE AND THE COMPLAINT IS STILL UNDER INVESTIGATION. | 1
1820013 | A NEOVENTA MEDICAL (B)(6) EMPLOYEE HEARD ABOUT THIS CASE OF A POOR NEWBORN OUTCOME THAT HAS BEEN REPORTED TO THE MEDICAL BOARD IN (B)(6). THE INFORMATION WAS INTERNALLY ENTERED IN THE CLINICAL FEEDBACK LOG. THE CLINIC HAS NOT PROVIDED DETAILED INFORMATION REGARDING THE CASE OF DEVICE. THE STAN REMAINS IN THE HOSPITAL. IT IS OUR UNDERSTANDING THAT THEY ARE TREATING THIS AS A USER ERROR. WE BELIEVE THERE ARE NO FAILURES WITH THE DEVICE, AND THAT THE DEVICE IS STILL IN USE. NEOVENTA BELIEVES THAT THE DEVICE IS NOT CONSIDERED THE ROOT CAUSE FOR THE INCIDENT. ALTHOUGH THE INITIAL REPORTABILITY DECISION INDICATED THAT THIS WAS NOT A REPORTABLE EVENT IN THE US, IN RESPONSE TO THE FDA WARNING LETTER OF AUGUST 4, 2010, AND CAPA PROJECT (B)(4), THIS EVENT WILL BE REPORTED IN THE US ACCORDING TO THE TABLE BELOW. WE ARE AWARE THAT REPORTING TIMELINES CANNOT BE MET WITH THE RETROSPECTIVE REPORTING. | 0
1883121 | THREE PROBES PLACED INTO TUMOR, THE FREEZE CYCLE WAS BEGUN AND ALL THREE PROBES FROM THE 3PK FROSTED UP THE SHAFT THROUGH THE SKIN. THREE SMALL LESIONS 1/2CM IN DIAMETER WERE NOTED ON THE SKIN POST PROCEDURE. TESTING INDICATED A VACUUM INSULATION FAILURE IN THE CRYOPROBE. FROSTING OF THE SHAFT WAS CONFIRMED. | 0
1883770 | CUSTOMER REPORTED AN UNEXPECTED (B)(6) RESULT WHEN TESTING A PATIENT SAMPLE ON THE GALILEO. CUSTOMER STATED THAT NO ADVERSE AFFECTS HAD RESULTED FROM THE UNEXPECTED RESULT. REVIEW OF INSTRUMENT RESULTS: E03- 8 (0), (E-,E+) WELL FILL COLOR AS EXPECTED. ROI'S ACCEPTABLE. VISUALLY NEGATIVE. F03- 15 (0), (E+,E-) WELL FILL COLOR AS EXPECTED. ROI'S ACCEPTABLE. SLIGHT FUZZY BUTTON, VERY FAINT PINK BACKGROUND. VISUALLY NEGATIVE. G03- 8(0), (E-,E+) WELL FILL COLOR AS EXPECTED. ROI'S ACCEPTABLE. VISUALLY NEGATIVE. H03- 10(0), (E-,E+) WELL FILL COLOR AS EXPECTED. ROI'S ACCEPTABLE. VISUALLY NEGATIVE. A SERVICE CALL WAS MADE. PERFORMED UNEXPECTED REACTIONS CHECKLIST TO VERIFY THAT THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. ALL SETTINGS AND READINGS WERE WITHIN SPECIFICATIONS. PERFORMED QC ASSAY TO VERIFY INSTRUMENT OPERATION. INSTRUMENT IS OPERATING AS EXPECTED. THE UNEXPECTED REACTION APPEARS TO BE RELATED TO THE NATURE OF THE SAMPLE. | 0
1887024 | THE CUSTOMER'S CELL-DYN SAPPHIRE ASPIRATION BOTTOM SENSOR ASSEMBLY WAS INSPECTED AND REPLACED BY (B)(4). NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(6). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL (B)(4) WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH (B)(4) INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1889442 | REPORTER INDICATED A VNS DELL X50 HANDHELD COMPUTER SERIAL CABLE WAS SUSPECTED OF BEING FAULTY, AND THE SERIAL CABLE HAS BEEN RETURNED FOR ANALYSIS. ANALYSIS OF THE SERIAL CABLE IS PENDING. CONCLUSIONS: DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1890824 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. DURING THE RUN, THE CUSTOMER REPORTED THAT SHORTLY INTO THE TPE PROCEDURE, THEY EXPERIENCED A NUMBER OF LOW AND HIGH LEVEL RESERVOIR SENSOR DETECTED EXCESS FLUID ALARMS AND THEN SHE HAD NOTICED AIR IN THE BLOOD WARMER (BW) LINE. (B)(4). THE OPERATOR STOPPED THE PROCEDURE AND CLAMPED THE LINE TO THE PT, SO NO AIR WAS RETURNED TO THE PATIENT. THE PT IS FINE. THE DISPOSABLE SET WAS RETURNED ALONG WITH AN ATTACHED ASTOTUBE DISPOSABLE BLOOD WARMER ACCESSORY AND THE SALINE BAG ATTACHED TO THE GREEN SOLUTION SPIKE. A MFG MISASSEMBLY WAS IMMEDIATELY APPARENT. THE PLASMA REMOVAL LINE LEADING TO THE WASTE BAG WAS BONDED TO THE CASSETTE IN THE PLASMA RECIRCULATION LINE BOND SOCKET LOCATION AND THE RECIRCULATION LINE WAS BONDED TO THE REMOVE LINE BOND SOCKET LOCATION. THIS MFG ERROR IS MUTUALLY EXCLUSIVE OF THE PRESENCE OF AIR IN THE RETURN LINE. THE DEVICE OPERATOR WAS ASKED DURING SUBSEQUENT FOLLOW-UP, IF THEY HAD A SALINE DRIP GOING DURING THE TIME THEY HAD ALL OF THE ALARMS AND THEY REPLIED AFFIRMATIVELY THAT THEY DID. THE OPERATOR WAS ASKED WHETHER THEY CLAMPED THE SALINE LINE WHEN THEY WERE INSTRUCTED TO PARTIALLY EMPTY THE SALINE LINE WHEN THEY WERE INSTRUCTED TO PARTIALLY EMPTY THE SALINE DRIP CHAMBER AFTER PRIME AND THE CUSTOMER RESPONDED THAT THEY HAD NOT. IF THE OPERATOR DOES NOT CLOSE THE ROLLER CLAMP ON THE SALINE LINE PRIOR TO INVERTING THE SALINE BAG TO SQUEEZE THE DRIP CHAMBER (CLEARING IT OF SALINE) AND THEY RETURN THE BAG TO THE HANGING POSITION, AIR CAN BE TRAPPED IN THE LINE AND WILL TRAVEL TO THE BW TUBING. MFG RETRAINING WAS REQUESTED FOR PROPER ASSEMBLY OF THE PLASMA RECIRC / PLASMA REMOVE LINES. ALL OPERATORS HAD COMPLETED THIS TRAINING BY (B)(4) 2009. CARIDIANBCT WILL APPEND OPTIA PROCEDURAL TRAINING MATERIALS. SPECIFICALLY ADDING A SLIDE TO THE OPTIA OPERATOR ESSENTIALS TRAINING SLIDESHOW USED BY CARIDIANBCT INSTRUCTORS. | 0
1890834 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. AT THE END OF A TPE PROCEDURE, THE OPERATOR NOTED THAT THE REMOVED VOLUME IN THE WASTE BAG WAS SIGNIFICANTLY LESS THAN IT SHOULD BE. THE PATIENT DID TOLERATE THE PROCEDURE WELL. PRE BP 139/74, PULSE 75 - POST BP 139/66 PULSE 93. PER THE CUSTOMER, THE PATIENT DID NOT HAVE ANY RESPIRATORY DISTRESS OR OTHER SYMPTOMS OF FLUID OVERLOAD. THE DISPLAYED REMOVE VOLUME WAS 3516ML (INCLUDING 374ML OF AC). THE ACTUAL PROCEDURE VOLUMES ARE AS FOLLOWS: REMOVED VOL = 1300ML, REPLACED = 3084ML 5% ALBUMIN PLUS 195ML RINSEBACK SALINE. THE CUSTOMER STATED THE PATIENT IS STABLE AND DID NOT REQUIRE ANY MEDICAL INTERVENTION AS A RESULT OF THIS FLUID IMBALANCE WITH THE PROCEDURE. (B)(4). A SIMULATED RUN WAS PERFORMED ON (B)(6) 2010 USING A TPE SET. CARIDIANBCT INVESTIGATORS COULD NOT CREATE THE CONDITION BY MISLOADING THE PLASMA VALVE TUBING BECAUSE THE SET WOULD NOT PASS ALARM TESTS. IF THE PLASMA COLLECT LINE ON A COBE SPECTRA DISPOSABLE IS BLOCKED DURING A TPE PROCEDURE, THIS CAN RESULT IN HYPERVOLEMIA IN THE PATIENT. IN THE PAST, CARIDIANBCT COMPLAINT INVESTIGATORS RAN THE SAME TYPE OF DISPOSABLE SET WITH THE COLLECT LINE BLOCKED. THE LACK OF ALARM IS LIKELY BECAUSE AN OCCLUSION AT THIS SITE ON THE PLASMA LINE DOES NOT RESULT IN INCREASED PRESSURE WITHIN THE TUBING SET BECAUSE THERE IS AN OPEN PATH VIA THE RBC RETURN LINE. IT HAS BEEN SHOWN THAT MISLOADING OF THE DISPOSABLE TUBING BELOW THE CCM CAN CAUSE AN INDENTATION SEVERE ENOUGH TO CREATE AN OCCLUSION. ROOT CAUSE: ALTHOUGH THE SET WAS RETURNED FOR INVESTIGATION, NO ROOT CAUSE COULD BE DETERMINED. IT IS LIKELY THAT THE ROOT CAUSE OF THIS INCIDENT IS RELATED TO AN OCCLUSION IN THE PLASMA COLLECT TUBING LINE THAT SELF-CORRECTED DURING THE PROCEDURE AS EVIDENCED BY SOME PLASMA WASTE BEING COLLECTED. IT IS UNCONFIRMED WHETHER THIS BLOCKAGE WAS PRODUCED BY A MFG DEFECT OF THE COBE SPECTRA DISPOSABLE OR BY THE WAY THE DISPOSABLE WAS LOADED AND HANDLED AT THE CUSTOMER SITE. SAFETY RISK ASSESSMENT: POST-PROCEDURE, THE PATIENT WAS WELL WITHOUT MEDICAL INTERVENTION TO TREAT THE HYPERVOLEMIA. | 0
1890837 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS THAT WERE PROMPTED BY AN FDA INSPECTION. THE CUSTOMER ENTERED THE WRONG DONOR INFO, 97 POUNDS HEAVIER THAN THE ACTUAL WEIGHT, BY ENTERING THE DONOR'S PLATELET PRE-COUNT VALUE AS THE DONOR'S WEIGHT, RISKING A POSSIBLE OVER-COLLECTION FROM THE DONOR. (B)(4). THIS KIT WAS NOT AVAILABLE FOR A SPECIFIC ROOT CAUSE ANALYSIS, CORRECTIVE AND PREVENTIVE ACTIONS BY QUALITY ASSURANCE. A DHR REVIEW WAS CONDUCTED. THE REVIEW SHOWED NO MFG ANOMALIES OCCURRED DURING PRODUCTION AND NOTHING OF THIS TYPE OF INCIDENT WAS REPORTED FROM THIS LOT. STERILIZATION RECORD DID NOT SHOW ANY ANOMALIES IN THE PROCESS. A TREND ANALYSIS ON THIS LOT ALSO SHOWED THIS WAS AN ISOLATED INCIDENT, INDICATING NOT DISPOSABLE RELATED ISSUE. THE EQUIPMENT STATUS WAS ALSO EXAMINED. THE PM HISTORY WAS CHECKED AND IT SHOWED A PM WAS PERFORMED IN SEPTEMBER OF 2008 WITHOUT ANY ISSUES. THE MACHINE WAS IN SERVICE FOR THREE MONTHS WHEN THIS INCIDENT OCCURRED. THE RDF WAS DOWNLOADED AND ANALYZED. THE RDF INVESTIGATION FOUND NOTHING OUT OF THE ORDINARY WITH THE RUN. THE PREDICTED FINAL COLLECTED PLATELET VOLUME WAS 434ML AND NO RBC AND PLASMA WAS COLLECTED. THERE WAS NO ADJUSTMENT MADE DURING THE PROCEDURE. THE CUSTOMER HAS SINCE INDICATED THAT THE MACHINE HAS BEEN OPERATING WITHOUT ISSUES AFTER THIS INCIDENT, SUGGESTING THIS IS NOT MACHINE RELATED. THE DONOR IS A REPEAT DONOR FOR DOUBLE PLATELET PRODUCTS. SINCE THE PROCEDURE WAS THE COLLECTION OF PLATELETS AND IT IS BASED ON THE REMOVED PLATELET COUNTS, THE TARGET WAS 6.7X10CELL IN THIS CASE, THE FINAL PLATELET COUNTS IN THE DONOR'S BODY SHOULD NOT HAVE BEEN COMPROMISED FROM ENTERING THE WRONG BODY WEIGHT. ROOT CAUSE: THE OPERATOR WRONGLY ENTERED THE DONOR INFO. THE OPERATOR ENTERED THE DONOR'S WEIGHT AS 312 LBS WHEN THE ACTUAL WEIGHT OF THE DONOR WAS 215 LBS, MISTAKENLY TAKING THE DONOR'S PLATELET PRE-COUNT AS THE DONOR'S WEIGHT. | 0
1894187 | THE CUSTOMER STATED A "BURNING SMELL AND POPPING NOISE" OCCURRED DURING DEMOGRAPHICS ENTRY WHEN PREPARING TO RUN A SAMPLE ON THE CELL-DYN 1800 ANALYZER. THE CUSTOMER ALSO NOTED A SMALL AMOUNT OF SMOKE FROM THE BACK OF THE ANALYZER. THERE WAS NO VISIBLE FLAME. THE ANALYZER WAS UNPLUGGED FROM THE WALL. A FIELD SERVICE REPRESENTATIVE INSPECTED AND SERVICED THE ANALYZER. AN INOPERATIVE POWER SUPPLY WAS REMOVED AND REPLACED. THERE WAS NO INJURIES INVOLVED AND NO DAMAGE TO THE SURROUNDING AREA OR INSTRUMENTS. (B)(4). IN RESPONSE TO THIS ISSUE, AN INVESTIGATION WAS INITIATED TO FURTHER EXAMINE THE CUSTOMER'S OBSERVATION. THE INVESTIGATION INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND A REVIEW OF THE RISK MANAGEMENT FILE FOR THE CELL-DYN ANALYZER. ATTEMPTS TO OBTAIN THE POWER SUPPLY ASSEMBLY WERE UNSUCCESSFUL. THE RISK MANAGEMENT FILE FOR THE CELL-DYN 1800 WAS REVIEWED AND FOUND TO INDICATE THAT SYSTEM COMPONENTS (I.E. POWER SUPPLY) ARE POTENTIAL SOURCE OF BURNS, FIRE, OR SMOKE BUT CONTROLS ARE IN PLACE TO REDUCE RISK INCLUDE OVER CURRENT PROTECTION (FUSING), COVERS, AND INSTRUCTIONS IN THE OPERATOR'S MANUAL TO NOT REMOVE COVERS DURING OPERATION, SAFETY ICONS, HAZARD WARNING LABELS, AND FUSE LABELING. A REVIEW OF THE GLOBAL SERVICE AND SUPPORT (GSS) FROM (B)(4) 2009 THROUGH (B)(4) 2010 INDICATES THAT THE POWER SUPPLY ASSEMBLY, P/N 8921029202, HAS AN AVERAGE OF (B)(4)RELIABILITY WORLDWIDE. PRODUCT LABELING WAS REVIEWED AND FOUND TO ADEQUATELY PROVIDE INFORMATION TO ASSIST IN PROBLEM IDENTIFICATION, ISOLATION, AND CORRECTIVE ACTIONS, AS WELL, AS THE PROCEDURES IN AN EMERGENCY SITUATION. A NON-STATISTICAL TREND REVIEW WAS PERFORMED IN THE COMPLAINT ANALYSIS TRENDING SYSTEM FOR THE REPORTED ISSUE FROM (B)(4) 2010 THROUGH (B)(4). NO TREND WAS IDENTIFIED DURING THE SEARCHED PERIOD. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 1800 RELATED TO THE REPORTED ISSUE. THE ROOT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED BECAUSE THE POWER SUPPLY WAS NOT AVAILABLE FOR INVESTIGATION.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1894293 | THE CUSTOMER STATED AN ARCHITECT I2000SR ANALYZER GENERATED A FALSE POSITIVE TROPONIN-I RESULT FOR ONE PATIENT SAMPLE. THE ARCHITECT I2000SR GENERATED AN INITIAL TROPONIN-I RESULT OF 54.2 NG/L AND A REPEAT TROPONIN-I RESULT OF 25.7 NG/L. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. THE CUSTOMER WAS INSTRUCTED TO REPLACE THE STAT PROBE. NO RECURRENCE OF ERRATIC RESULTS HAS BEEN IDENTIFIED. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND RELATED TO DISCREPANT TROPONIN-I RESULTS. PACKAGE INSERT AND OPERATIONS MANUAL LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON THE EVALUATION, A PRODUCT DEFICIENCY WAS NOT IDENTIFIED. | 0
1904582 | AN ABBOTT FIELD SERVICE ENGINEER (FSE) INSPECTED THE CUSTOMER'S CELL-DYN SAPPHIRE ANALYZER AND FOUND IT DEFECTIVE. THE FSE REPLACED THE BOTTOM SENSOR DETECTOR BOARD AND SENT THE PART BACK FOR THE INVESTIGATION. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1906196 | THE ACCOUNT STATED THAT THEY NOTED A LOUD BANG COMING FROM THE UPS THAT WAS CONNECTED TO THE CELL-DYN RUBY ANALYZER. THE ACCOUNT ALSO OBSERVED VISIBLE SMOKE AND FIRE COMING FROM THE UPS. THE ACCOUNT DISCONNECTED THE POWER FROM THE ANALYZER AND UPS IMMEDIATELY. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. THERE WAS NO INJURIES REPORTED. FIELD SERVICE WAS REQUESTED. (B)(4). THIS IS AN INITIAL REPORT. AN INVESTIGATION IS IN PROCESS. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). A FIELD SERVICE ENGINEER REPLACED THE DEFECTIVE UNINTERRUPTED POWER SUPPLY (UPS). THE CELL-DYN RUBY INSTRUMENT WAS NOT IMPACTED BY THIS ISSUE AND WAS OPERATING WITHIN SPECIFICATIONS. THE REPORTED ISSUE WAS ISOLATED TO THE UPS. THE UPS WAS RETURNED FOR EVALUATION ON (B)(4) 2010. VISUAL INSPECTION OF THE OUTSIDE AREA OF THE UPS SHOWED NO SIGNS OF SMOKE OR BURNING ON THE SURFACE. THERE WAS NO IEC POWER CORD PROVIDED WITH THE UPS FOR A 220 VOLTAGE APPLICATION; DAMAGE TO THE POWER CORD IS UNKNOWN. THE OUTSIDE METAL COVERING WAS REMOVED FOR INTERNAL INSPECTION. THE INVESTIGATOR FOUND THAT THE LARGE MAIN PWA RESISTORS WERE BURNT; THE RESISTORS WERE 13 IN TOTAL AND IN AN AXIAL POSITION PARALLEL TO THE PWA. THE PWA SURFACE WAS BURNT DUE TO OVERHEATING OF THE RESISTORS, CONFIRMING THE POSSIBLE SMOKE SEEN BY THE CUSTOMER. DUE TO THE BURNT RESISTORS, THE UPS COULD NOT BE TESTED WITH LIVE POWER. THE INVESTIGATOR CONCLUDED THAT IT WAS POSSIBLE THAT THE SMOKE OR BURNING ON THE UPS WAS DUE TO EXTREME FLUCTUATING POWER OR SPIKE, CAUSED BY UNKNOWN WEATHER CONDITIONS. THE UPS IS A PERIPHERAL DEVICE, WHICH IS NOT PART OF THE CELL-DYN RUBY SYSTEM ACCESSORY KIT; THIS PRODUCT IS ORDERED SEPARATELY AS A CONVENIENCE TO THE CUSTOMER WHEN PLACING AN ORDER FOR A CELL-DYN RUBY SYSTEM. THE CELL-DYN RUBY SYSTEM OPERATOR'S MANUAL PROVIDES INSTRUCTIONS FOR EMERGENCY SHUT-DOWN. A NON-STATISTICAL TREND (NST) COMPLAINT REVIEW WAS PERFORMED FOR THE REPORTED ISSUE FROM (B)(6) 2010. NO NST WAS IDENTIFIED DURING THE SEARCHED PERIOD. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN RUBY INSTRUMENT RELATED TO THE REPORTED ISSUE. THE REPORTED FAILURE WAS ISOLATED TO THE UPS. | 0
1908966 | THE CUSTOMER REPORTED TO CARIDIANBCT THAT THEY HAD A FEW DONOR REACTIONS AFTER THE DONOR GAVE A PLATELET PRODUCT AND A JUMBO PLASMA. ON (B)(6), A (B)(6) MAN PASSED OUT AND HIT A LIGHT POLE ON THE WAY HOME AFTER DONATING. THIS REPORT IS BEING SUBMITTED DUE TO THE SEVERITY OF THE DONOR REACTIONS AND MEDICAL INTERVENTION.><THE CUSTOMER REPORTED TO CARIDIANBCT THAT THEY HAD A FEW DONOR REACTIONS AFTER THE DONOR GAVE A PLATELET PRODUCT AND A JUMBO PLASMA. ON (B)(6), A (B)(6) MALE STAYED AT THE CENTER AFTER DONATING AND THEN VOMITED 6X ON THE WAY HOME. ANOTHER DONOR REPORTEDLY FEELS "WOOZY" WHEN HE DONATES JUMBO PLASMA. THIS REPORT IS BEING SUBMITTED DUE TO THE SEVERITY OF THE DONOR REACTIONS AND MEDICAL INTERVENTION. (B)(4). THE CLINICAL SUPPORT SPECIALIST DISCUSSED THE EVENTS WITH THE CUSTOMER AND DETERMINED THAT IN EACH CASE, THE FLUID VOLUMES COLLECTED WERE AS DISPLAYED AND AS EXPECTED INDICATING THAT THE DEVICE FUNCTIONED AS EXPECTED. SYNCOPE AND LIGHT HEADACHES IS AN ANTICIPATED REACTION TO THE DONATION OF BLOOD PRODUCTS AND COULD HAVE BEEN RELATED TO THE STARTING BP OF THE DONORS. CONCLUSION: ALTHOUGH NO DEFINITIVE CONCLUSION CAN BE DRAWN, IT APPEARS THAT THE TRIMA DEVICE FUNCTIONED AS INTENDED AND WITHIN SPECIFICATIONS. DONOR SYNCOPE REACTIONS ARE A KNOWN SIDE EFFECT OF THIS PROCEDURE.><(B)(4). THE CLINICAL SUPPORT SPECIALIST DISCUSSED THE EVENTS WITH THE CUSTOMER AND DETERMINED THAT IN EACH CASE, THE FLUID VOLUMES COLLECTED WERE AS DISPLAYED AND AS EXPECTED INDICATING THAT THE DEVICE FUNCTIONED AS EXPECTED. SYNCOPE AND LIGHT HEADACHES IS AN ANTICIPATED REACTION TO THE DONATION OF BLOOD PRODUCTS AND COULD HAVE BEEN RELATED TO THE STARTING BP OF THE DONORS. CONCLUSION: ALTHOUGH, NO DEFINITIVE CONCLUSION CAN BE DRAWN, IT APPEARS THAT THE TRIMA DEVICE FUNCTIONED AS INTENDED AND WITHIN SPECIFICATIONS. DONOR SYNCOPE REACTIONS ARE A KNOWN SIDE EFFECT OF THIS PROCEDURE. | 0
1913197 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO A DRIPPING REAGENT PROBE ON THE PK7300 AUTOMATED MICROPLATE SYSTEM. NO PATIENTS WERE INVOLVED WITH THIS EVENT. A BCI FIELD SERVICE ENGINEER (FSE) RELOCATED THE PROBE TO AN ALTERNATE POSITION. A CLEAR ROOT CAUSE HAS NOT BEEN IDENTIFIED. | 0
1915736 | A CLIENT SITE REPORTED A SITUATION ON (B)(6) 2010 WHICH WAS DIRECTLY CAUSED BY THE CUSTOMER USING THE SOFTWARE USING AN ORACLE CONFIGURATION UNDER WHICH THE MEDICAL DEVICE SOFTWARE WAS NOT VALIDATED. IN THE REPORTED ISSUE, A CROSSMATCH TEST WAS NOT LINKED TO A COMPONENT IN A CERTAIN SCENARIO AND THE SYSTEM DID NOT DISPLAY A WARNING. HOWEVER, THE ERROR SHOULD HAVE BEEN DETECTED BY SITE FOLLOWING THEIR STANDARD OPERATING PROCEDURE WHICH STATES THAT THE PRODUCT ID TAG SHOULD BE COMPARED TO THE INFO IN THE SYSTEM. NO PT HARM OCCURRED. DURING THE COURSE OF THE INVESTIGATION OF THE REPORTED ISSUE, THE MANUFACTURER IDENTIFIED ADDITIONAL ORACLE SETTINGS THAT ARE IMPORTANT TO THE SYSTEM'S DATABASE CONFIGURATION. DOCUMENTS THAT CONTAIN ADDITIONAL LABELING HAVE BEEN DISTRIBUTED TO ALL SITES THAT CURRENTLY USE THE PRODUCT. THE DATABASE SETTING WAS CHANGED AT THE REPORTING SITE AND THE ISSUE HAS NOT OCCURRED SINCE. OTHER CLIENT SITES WERE ALSO INVESTIGATED AND NO OCCURRENCES WERE IDENTIFIED; IN ADDITION, THE DATABASE CONFIGURATIONS OF THESE OTHER SITES WERE CHECKED AND THEY ARE NOT USING THE UNVALIDATED CONFIGURATION. EVAL SUMMARY: WHEN THE SOFTWARE IS USED WITH AN ORACLE DATABASE SETTING OF COST-BASED OPTIMIZATION MODE, WHICH IS AN UNVALIDATED CONFIGURATION, THE SYSTEM MAY NOT LINK A CROSSMATCH TEST TO THE COMPONENT. THE DEVICE WAS BEING USED FOR TREATMENT; NO PT HARM OCCURRED. NO ADVERSE EVENT OCCURRED. THE REPORTING SITE WAS USING THE SOFTWARE IN AN UNVALIDATED ORACLE DATABASE CONFIGURATION, AND WAS NOT INSTALLED ACCORDING TO MANUFACTURER'S INSTALLATION INSTRUCTIONS. DURING THE COURSE OF THE INVESTIGATION OF THE REPORTED ISSUE, THE MANUFACTURER IDENTIFIED ADDITIONAL ORACLE SETTINGS THAT ARE IMPORTANT TO THE SYSTEM'S DATABASE CONFIGURATION. THE DATABASE SETTING WAS CHANGED AT THE REPORTING SITE AND THE ISSUE HAS NOT OCCURRED SINCE. OTHER CLIENT SITES WERE ALSO INVESTIGATED AND NO OCCURRENCES WERE IDENTIFIED; IN ADDITION, THE DATABASE CONFIGURATIONS OF THESE OTHER SITES WERE CHECKED AND THEY ARE NOT USING THE UNVALIDATED CONFIGURATION. | 1
1920036 | THE CUSTOMER STATES THAT ONE PATIENT SAMPLE KNOWN TO BE NEGATIVE FOR TROPONIN-I, GENERATED ARCHITECT STAT TROPONIN-I ASSAY RESULTS IN THE RANGE OF 0.00 TO 0.16 UG/L (N=7) ON THE ARCHITECT I2000SR ANALYZER. THE CUSTOMER USES A TROPONIN-I CUTOFF VALUE OF 0.06 UG/ML. THE CUSTOMER DID NOT HAVE ANY FURTHER PATIENT INFORMATION. ABBOTT CONTROL MATERIALS HAVE BEEN WITHIN SPECIFICATIONS ON ALL RUNS. THE CUSTOMER NOTED THAT THE ANALYZER'S STAT PROBE HAD NOT BEEN REPLACED SINCE JULY 2010. ALSO THE MESSAGE HISTORY LOG OF THE ANALYZER SHOWED STAT PROBE OBSTRUCTION ERRORS. THE CUSTOMER USES LITHIUM HEPARIN IN THE COLLECTION TUBES AND DOUBLE CENTRIFUGES ALL SAMPLES. A SERVICE CALL WAS INITIATED. TO DATE, NO INFORMATION HAS BEEN PROVIDED CONCERNING PATIENT MANAGEMENT. (B)(4). THE CUSTOMER WAS CONCERNED WITH ASPIRATION AND WASH ZONE ERRORS WITH ERRATIC TROPONIN- I RESULTS BEING GENERATED ON THE ARCHITECT I2000SR ANALYZER. AN ABBOTT TECHNICAL APPLICATION SPECIALIST (TAS) REVIEWED THE MESSAGE HISTORY LOG AND FOUND MANY STAT PROBE OBSTRUCTION ERRORS AND WASH ZONE 2 ASPIRATION ERRORS. THE STAT PROBE HAD NOT BEEN REPLACED SINCE (B)(4) 2010. A REVIEW OF CUSTOMER MAINTENANCE LOGS REVEALED THAT THE LAST MAINTENANCE DOCUMENTATION WAS MADE IN (B)(4) 2010. THE TAS SUGGESTED THE CUSTOMER REPLACE THE STAT PROBE. ABBOTT FIELD SERVICE (FSR) WAS REQUESTED TO REPLACE THE PROBE, CHECK THE INSTRUMENT AND OPTIMIZE INSTRUMENT PERFORMANCE. THE FSR REPLACED THE PROBE TUBING ON PROBE #1 ON WASH ZONE 2. FURTHER INSTRUMENT EVALUATION FOUND THAT THE MOST LIKELY CAUSE OF THE CUSTOMER'S CURRENT ISSUE WAS A WEAK DRAIN VALVE ON WASH ZONE 2 VACUUM VESSEL WHICH CAUSED THE INSTRUMENT ERRORS. THE FSR REPLACED BOTH WASH ZONE 2 VACUUM VESSEL VALVES, AND THE WASH ZONE 2 PROBE TUBE SENSOR. LEAKING WASH ZONE MANIFOLDS WERE ALSO REPLACED. WASH ZONE ASPIRATION TESTS PASSED AND CONTROLS WERE WITHIN RANGE AND THE SYSTEM WAS PERFORMING WITHIN SPECIFICATIONS. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-108) JANUARY, 2010 AND THE ARCHITECT STAT TROPONIN-I PACKAGE INSERT ((B)(4)) BOTH CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1921050 | IT WAS REPORTED THAT THE HANDHELD COMPUTER WOULD NOT POWER ON. TROUBLESHOOTING WAS PERFORMED, BUT THE PROBLEM PERSISTED. PER REPORTER, IT DOESN'T SEEM THAT THERE WAS ANY ABNORMAL ENVIRONMENT WHICH MIGHT CAUSE THIS EVENT. PRODUCT WAS RETURNED TO THE MFR AND AN ANALYSIS WAS PERFORMED. UPON ANALYSIS, CONTINUITY TESTING WAS ABLE TO IDENTIFY AN INTERMITTENT NEGATIVE ELECTRICAL CONNECTION IN THE POWER SUPPLY CONNECTOR OF THE HANDHELD SERIAL CABLE. THE CAUSE FOR THE OPEN CONNECTION IS ASSOCIATED WITH THE LEAF SPRING IN THE SERIAL CABLE PLUG NOT MAKING CONTACT WITH THE AC ADAPTER BARREL CONNECTOR. A ROOT CAUSE FOR THE OBSERVED LEAF SPRING CONDITION COULD NOT BE DETERMINED DURING THE ANALYSIS. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
1923451 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO A DRIPPING REAGENT PROBE ON THE (B)(4) AUTOMATED MICROPLATE SYSTEM. NO PATIENTS WERE INVOLVED WITH THIS EVENT. ROOT CAUSE AND CORRECTIVE ACTION ARE UNDER INVESTIGATION. | 0
1927950 | THE CUSTOMER REPORTED THAT THE PT WAS TREATED WITH 8 FIELD IMRT PLAN. THE PT HAS 5 FIELDS AT 180 COUCH ROTATION AND 3 FIELDS WITH 90 COUCH ROTATION. SET-UP PT AND CLICK ACQUIRE PRIOR TO EVERY TREATMENT TO ENSURE COUCH PARAMETERS ARE THE SAME BY KV-KV IMAGE AND PERFORM 2D/2D MATCH. WHEN THEY ACQUIRED, THE THERAPIST CLICK "SELECT ALL" - AND ACQUIRED COUCH ROTATION ACCIDENTLY. ACKNOWLEDGED WARNING PLAN PARAMETER HAD CHANGED. THIS MODIFIED THE PLANNED COUCH ROTATION FOR ALL FIELDS TO 180 WHEN 3 WERE ACTUALLY PLANNED AT 90 - PERFORMED ONLINE MATCH APPLIED SHIFTS. THE FIRST FIELD THAT HAD A COUCH ROTATION OF 90 DEGREES WAS TREATED WITH COUCH ROTATION OF 0 FOR 28 MU OUT OF 95 MU FOR THAT FIELD. THERAPIST STOPPED BEAM AND ROTATED COUCH ROTATION TO 90 DELIVERED REMAINDER OF TREATMENT. REMAINDER OF FIELD TREATED CORRECTLY: RIGHTS IN TREATMENT ADMINISTRATION; PLAN EDIT PREFERENCES DO NOT ALLOW EDITING AND ACQUIRING ON APPROVED PLANS. IN THE OPINION OF THE PHYSICIAN, THE MISADMINISTRATION WAS NOT CONSIDERED A SERIOUS INJURY. PT HAD 8 FIELD TREATMENT-. ONE OF THE FIELDS WAS PARTIALLY TREATED WITH COUCH AT WRONG ROTATION FOR 28 OF 90 MU. THIS RESULTED IN 8CGY BEING DELIVERED TO ORBIT AND BRAIN STEM. THIS IS HIGHLY UNLIKELY TO CAUSE INJURY. THE ISSUE AS REPORTED INVOLVES THE PLAN EDIT PREFERENCE SETTINGS DURING KV, KV IMAGING, AND THE USER EXPECTS THE PLAN EDIT PREFERENCE TO NOT BE OVERRIDDEN BY INDIVIDUAL USER RIGHTS. ALTHOUGH, THE MEDICAL PROFESSIONAL REPORT INDICATED NO SERIOUS INJURY IN THIS CASE, THE REPORTED ISSUE COULD RESULT IN SERIOUS INJURY IF THE CAPTURED SETTINGS WERE USED FOR TREATMENT ON MULTIPLE FRACTIONS. THOUGH STILL UNDER INVESTIGATION, VARIAN HAS DETERMINED THAT AN MDR IS APPROPRIATE, AS THIS MALFUNCTION, SHOULD IT RECUR, COULD POTENTIALLY CAUSE A SERIOUS INJURY. ADD'L F/U TO THIS MDR IS EXPECTED UPON COMPLETION OF THE INVESTIGATION. | 0
1928964 | THE CUSTOMER STATED THAT THE ASPIRATION PROBE OF THE CELL-DYN SAPPHIRE ANALYZER CAN NOT BE REMOVED. WHEN TRYING TO REPLACE IT, IT WAS MOVED IN A CERTAIN WAY WHERE IT WAS STUCK WITH NO WAY TO GET IT OUT. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL (B)(4) 2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH (B)(4) 2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE | 0
1933616 | THE CUSTOMER STATES THAT ONE PATIENT SAMPLE GENERATED AN INITIAL FALSE POSITIVE ARCHITECT STAT TROPONIN-I ASSAY RESULT OF 0.084 UG/L ON THE ARCHITECT I2000SR ANALYZER. THE CUSTOMER USES A TROPONIN-I CUTOFF VALUE OF 0.06 UG/ML. THE CUSTOMER DID NOT PROVIDE ANY FURTHER PATIENT INFORMATION. THE SAMPLE WAS RETESTED AND GENERATED A RESULT 0.004 UG/L. ABBOTT CONTROL MATERIALS HAVE BEEN WITHIN SPECIFICATIONS ON ALL RUNS. EARLIER, THE CUSTOMER NOTED THAT THE ANALYZER'S STAT PROBE HAD NOT BEEN REPLACED SINCE JULY 2010. ALSO THE MESSAGE HISTORY LOG OF THE ANALYZER SHOWED MANY WASH ZONE 2 ASPIRATION ERRORS. THE CUSTOMER USES LITHIUM HEPARIN IN THE COLLECTION TUBES AND DOUBLE CENTRIFUGES ALL SAMPLES. A SERVICE CALL WAS INITIATED AND THE ABBOTT FIELD SERVICE ENGINEER (FSE) REPLACED THE TUBING ON PROBE #1 OF WASH ZONE 2. THE AFOREMENTIONED FALSE POSITIVE RESULT WAS GENERATED AFTER THE FSE VISIT. A CALL FOR SERVICE TO RETURN WAS ISSUED. TO DATE, NO FURTHER INFORMATION HAS BEEN PROVIDED CONCERNING PATIENT MANAGEMENT. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). THE CUSTOMER WAS CONCERNED WITH ASPIRATION AND WASH ZONE ERRORS WITH ERRATIC TROPONIN- I RESULTS BEING GENERATED ON THE ARCHITECT I2000SR ANALYZER. AN ABBOTT TECHNICAL APPLICATION SPECIALIST (TAS) REVIEWED THE MESSAGE HISTORY LOG AND FOUND MANY STAT PROBE OBSTRUCTION ERRORS AND WASH ZONE 2 ASPIRATION ERRORS. THE STAT PROBE HAD NOT BEEN REPLACED SINCE (B)(4) 2010. A REVIEW OF CUSTOMER MAINTENANCE LOGS REVEALED THAT THE LAST MAINTENANCE DOCUMENTATION WAS MADE IN (B)(4) 2010. THE TAS SUGGESTED THE CUSTOMER REPLACE THE STAT PROBE. ABBOTT FIELD SERVICE (FSR) WAS REQUESTED TO REPLACE THE PROBE, CHECK THE INSTRUMENT AND OPTIMIZE INSTRUMENT PERFORMANCE. THE FSR REPLACED THE PROBE TUBING ON PROBE #1 ON WASH ZONE 2. FURTHER INSTRUMENT EVALUATION FOUND THAT THE MOST LIKELY CAUSE OF THE CUSTOMER'S CURRENT ISSUE WAS A WEAK DRAIN VALVE ON WASH ZONE 2 VACUUM VESSEL WHICH CAUSED THE INSTRUMENT ERRORS. THE FSR REPLACED BOTH WASH ZONE 2 VACUUM VESSEL VALVES, AND THE WASH ZONE 2 PROBE TUBE SENSOR. LEAKING WASH ZONE MANIFOLDS WERE ALSO REPLACED. WASH ZONE ASPIRATION TESTS PASSED AND CONTROLS WERE WITHIN RANGE AND THE SYSTEM WAS PERFORMING WITHIN SPECIFICATIONS. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL (PN 201837-108) JANUARY, 2010 AND THE ARCHITECT STAT TROPONIN-I PACKAGE INSERT ((B)(4)) BOTH CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION. | 0
1935575 | THE CUSTOMER STATES THAT THEY HAD NOTICED ERRATIC HEMOGLOBIN RESULTS FOR PATIENT SAMPLES PROCESSED USING A CELL-DYN 1800 ANALYZER. ONE PATIENT GENERATED THE FOLLOWING HEMOGLOBIN RESULTS FOR THREE RUNS; RUN 1: HGB=11.0 G/DL, RUN 2: HGB=6.6 G/DL, RUN 3: HGB=11.7 G/DL. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. (B)(4). THE CUSTOMER STATED THAT THE ERRATIC HEMOGLOBIN (HGB) RESULTS WERE NOT REPORTED OUTSIDE THE LAB; SAMPLES WERE SENT OUT TO THE REFERENCE LAB FOR TESTING. THE CUSTOMER REPORTED CLEANING THE APERTURE PLATES AND FLUSHING THE REAGENT LINES WITHOUT RESOLUTION. THE CUSTOMER ATTEMPTED TO RUN PRECISION ON NORMAL WHOLE BLOOD, BUT AFTER FIVE (5) RUNS THE HGB BEGAN RUNNING BETWEEN 12.9 TO 17.9 G/DL, CAUSING CV% TO BE 15.4. THE CUSTOMER TECHNICAL ADVOCATE (CTA) VERIFIED WITH THE CUSTOMER THAT QUALITY CONTROLS (QC) WERE WITHIN RANGE. THE CUSTOMER REPORTED THAT THE CORRECT LYSE WAS INSTALLED AND THE SYRINGES WERE CLEANED; NO BUBBLES OR LEAKS WERE OBSERVED. THE CUSTOMER REPORTED THAT HGB REFERENCE WAS 1789 (SPECIFICATION OF 1800 - 2200). THE CTA SCHEDULED A SERVICE VISIT FOR ISSUE RESOLUTION. A FIELD SERVICE REPRESENTATIVE (FSR) FOUND THAT THE HGB VOLTAGE WAS LOW; THE FSR MADE AN ADJUSTMENT TO THE VOLTAGE. THE CUSTOMER VERIFIED PROPER INSTRUMENT OPERATION. NO FURTHER INVESTIGATION WAS REQUIRED. THE ISSUE IS ADEQUATELY ADDRESSED IN THE PRODUCT LABELING A NON-STATISTICAL TREND (NST) REVIEW WAS PERFORMED AND NO NST WAS IDENTIFIED DURING THE SEARCHED PERIOD. BASED ON THE INVESTIGATION, NO PRODUCT ISSUE WAS IDENTIFIED FOR THE CELL-DYN 1800 RELATED TO THE REPORTED ISSUE.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1936448 | A CUSTOMER CONTACTED BECKMAN COULTER INC (BCI) REGARDING DECREASED, UN-FLAGGED RESULTS FOR WHITE BLOOD CELLS (WBC), RED BLOOD CELLS (RBC), AND HEMOGLOBIN (HGB) RESULTS GENERATED BY THE COULTER LH500 ANALYZER. THE PATIENT WAS REDRAWN AND RUN ON THE LH500 INSTRUMENT AND RECOVERED HIGHER WBC, RBC, AND HGB RESULTS. THE ORIGINAL SPECIMEN WAS RERUN THREE ADDITIONAL TIMES WERE THE RBC AND HGB RESULTS WERE SIMILAR TO THE REDRAW SPECIMEN (CORRECT) RESULTS; HOWEVER, THE WBC RESULTS WERE STILL LOW. RESULTS WERE REPORTED OUT OF THE LABORATORY. THERE WAS NO CHANGE TO PATIENT TREATMENT, NO DEATH OR SERIOUS INJURY IS ASSOCIATED WITH THIS EVENT. SAMPLES WERE RUN IN PRIMARY MODE USING STANDARD 4ML BD SAMPLE TUBE. QC WAS RUN BEFORE THE EVENT AND RECOVERED WITHIN SPECIFICATIONS. A FIELD SERVICE ENGINEER (FSE) INSTRUCTED THE CUSTOMER TO PERFORM BLOOD SAMPLING VALVE (BSV) CLEANING, VERIFY REPRODUCIBILITY AND QC. ALL RESULTS WERE ACCEPTABLE PORT BSV CLEANING. ROOT CAUSE COULD NOT BE DETERMINED BASED ON AVAILABLE INFORMATION. | 0
1936590 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) TO REPORT A CONTROL RECOVERY PROBLEM (OUT OF LIMITS, HIGH) FOR THE PLT PARAMETER ON THE NORMAL, 4 C-PLUS CELL CONTROL. DURING THE CALL THE CUSTOMER STATED THE NORMAL CONTROL WAS REMOVED OUT OF THE BOX AND THE CAP POPPED OFF THE VIAL SPILLING HALF OF THE CONTENT. THE SPILL WAS CLEANED WITH BLEACH AND ABSORBENT PAD. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT. NO EXPOSURE TO OPEN LESION OR MUCUS MEMBRANES WAS REPORTED AND MEDICAL TREATMENT/ATTENTION WAS NOT SOUGHT. SERVICE HISTORY REVIEW OF (B)(4) 2010 DID NOT SHOW ANY INCIDENCES OF CALLS MADE IN REGARDS TO CONTROL CAP POPPING OFF AND/OR CONTROL LEAK. THE CUSTOMER DECLINED REPLACEMENT CONTROL AND DID NOT HAVE ANOTHER SET OF CONTROLS TO VERIFY THE NORMAL CONTROLS. A CLEAR ROOT CAUSE FOR THE LEAK IS UNKNOWN. | 0
1939095 | THE CUSTOMER STATED THAT THE CELL-DYN SAPPHIRE ANALYZER IS GENERATING ERROR CODE 1050, ASPIRATION NEEDLE DID NOT RETURN TO THE HOME POSITION. THE INSTRUMENT WAS INSPECTED AND THE ASPIRATION BOTTOM SENSOR ASSEMBLY WAS REPLACED ALONG WITH A LEAKING RESERVOIR. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). EVALUATION: ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL (B)(4) 2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH (B)(4) 2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. WITH V4, THE CD SAPPHIRE WILL EXECUTE SENSOR CHECKS TO DETERMINE THE STATUS OF THE SENSOR AND WILL GENERATE AN ERROR MESSAGE IF AN ISSUE IS DETECTED. | 0
1940953 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. (B)(6) IS REVIEWING A DONOR RECORD FROM (B)(6) 2009. FROM HER REVIEW OF THE DONOR RECORD, IT APPEARS THAT THE VENIPUNCTURE WAS PERFORMED ON THE DONOR AND THE SAMPLES WERE COLLECTED FROM THE SAMPLE POUCH BEFORE AC PRIME HAD BEEN STARTED. SHE IS WANTING TO KNOW IF THERE IS A POSSIBILITY THAT THE AIR COULD HAVE BEEN GIVEN TO THE DONOR. CUSTOMER IS NOT ALLEGING A DEFICIENCY WITH EITHER THE MACHINE OR DISPOSABLE. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR AIR TO DONOR. (B)(4). RISK ANALYSIS: IN THIS PARTICULAR CASE, ACCORDING TO WHAT THE OPERATOR TOLD THE MGR WHO RELAYED THE INFO TO CARIDIANBCT, THE VENIPUNCTURE WAS PERFORMED AFTER THE CASSETTE WAS LOADED BUT BEFORE AC PRIME WAS COMPLETED. WHEN THE CASSETTE IS LOWERED IT PERFORMS A SERIES OF AIR EVACUATIONS AND SET TESTS. AFTER THIS POINT THE TRIMA PROMPTS THE OPERATOR TO CLOSE THE DONOR LINE CLAMPS AND THE SAMPLE BAG LINE CLAMPS BEFORE MOVING ON TO THE AC PRIME. THIS IS WHEN THE OPERATOR STATES THEY ATTACHED THE DONOR. ACCORDING TO SUPPORT/CLINICAL SPECIALIST FOR AUTOMATED COLLECTIONS, (B)(6), AT THIS POINT THERE IS NO DANGER OF THE OPERATOR RECEIVING AN AIR EMBOLISM SINCE THE AIR EVACUATION HAS BEEN COMPLETED, HOWEVER BY ATTACHING THE DONOR EARLY THE OPERATOR HAS RELEASED PRESSURE FROM THE SET AND IT WILL NOT PASS AC PRIME, WHICH IS WHAT OCCURRED IN THIS PROCEDURE. ANY AIR LEFT IN THE SET DURING AC PRIME (IN THE AC LINE) IS PUSHED BY THE AC PUMP INTO THE SET ITSELF AND THROUGH THE LEVEL SENSOR, WHICH ALARMS WITH AN AC PRIME FAILURE BEFORE ANY AIR CAN REACH THE DONOR. RDF ANALYSIS CONFIRMS THAT THE EQUIPMENT HAD COMPLETED THE PROCESS TO THE AC PRIME SCREEN (WHEN AC IS ATTACHED) WITH NO ISSUES, BUT RECEIVED AN AC NOT DETECTED ALARM AND THEN ON SECOND ATTEMPT THE RUN FAILS DUE TO THE ACCESS PRESSURE SENSOR NOT REACHING TARGET PRESSURE. CONCLUSION: OPERATOR ERROR. TRIMA V6 SOFTWARE ADDED WARNING SCREENS REGARDING HAVING THE PT CONNECTED DURING AIR EVACUATION AND TO REDUCE THE AMOUNT OF AIR PUSHED OUT DURING PUMP LOAD/UNLOAD. | 0
1940959 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVAL PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. THE FOLLOWING INCIDENT DESCRIPTION WAS PROVIDED TO CARIDIANBCT QUALITY ASSURANCE: DURING AUTOLOGOUS PBSC COLLECTION THERE WAS A LEAK AT THE BLOOD WARMER LUER CONNECTION. POST-PROCEDURE THE PT SUFFERED A LOW GRADE FEVER. THIS REPORT IS BEING FILED DUE TO THE COLLECTION OF BLOOD CULTURES FROM THE DONOR. (B)(4) - LEAK/WITH FEVER POST COLLECTION. CLINICAL SPECIALIST, CALLED THE CUSTOMER FOR F/U ON (B)(6) 2009 AND THE CUSTOMER REPORTED THAT THE PT IS DOING WELL. THE CUSTOMER SITE LEFT A VOICE MAIL ON (B)(6) 2009 FOR (B)(6) THAT DETAILED THE BLOOD CULTURE RESULTS; PRODUCT=NEGATIVE / NO GROWTH, PT'S BLOOD CULTURES = NEGATIVE / NO GROWTH, PT'S CENTRAL LINE TIP = NEGATIVE / NO GROWTH, AND BLOOD CULTURE DRAWN VIA THE CENTRAL LINE = NEGATIVE / NO GROWTH. CARIDIANBCT DISPOSABLE QUALITY RECEIVED A 10" SEGMENT OF ASTO TUBE WITH MALE LUER CONNECTION ATTACHED. THE CONNECTOR WAS EXAMINED UNDER MAGNIFICATION FOR DEFECTS, NONE WERE IDENTIFIED. THE INVESTIGATOR ATTACHED A FEMALE CAP ON THE MALE END, AND ATTEMPTED TO REPEAT THE LEAK USING WATER UNDER PRESSURE. THE INVESTIGATOR WAS NOT ABLE TO DUPLICATE THE COMPLAINT. A POSSIBLE CAUSE FOR THE LEAK INCLUDES A SLIGHT MFG INTERFERENCE DEFECT BETWEEN THE MALE LUER CONNECTION ON THE ASTO TUBE AND THE FEMALE LUER CONNECTION USED (NOT PROVIDED FROM CUSTOMER). ANOTHER POSSIBLE CAUSE FOR THE LEAK COULD BE OPERATOR ERROR IN ESTABLISHING THE CONNECTION. PER THE INFO PROVIDED FROM THE CUSTOMER, THE PT WAS SUFFERING FROM MULTIPLE MYELOMA AND THE FEVER COULD HAVE BEEN CAUSED BY THE G-CSF THAT WAS ADMINISTERED. CONCLUSION: A ROOT CAUSE FOR THE LEAK AND PT FEVER CANNOT BE IDENTIFIED AT THIS TIME DUE TO MULTIPLE POSSIBLE CAUSES AND INSUFFICIENT INFO TO IDENTIFY ONE MAIN ROOT CAUSE. | 0
1943593 | IT WAS REPORTED THAT A PT MONITORED ON A DASH 2000 EXPIRED. THE DEVICE DID NOT PROVIDE AN AUDIBLE ALARM BECAUSE, THE VOLUME WAS SET TO ZERO. ADDITIONAL PT INFO WAS PROVIDED. PT INFO IS CONSIDERED CONFIDENTIAL IN (B)(6) AND WILL NOT BE RELEASED BY THE HOSPITAL. NARRATIVE: GE HEALTHCARE PERFORMED AN ANALYSIS OF THE CIC LOGS FOR THE DATE AND TIME OF THE REPORTED EVENT. THE LOGS INDICATE THAT THE DASH MONITOR IN QUESTION PROVIDED ADVISORY AND CRISIS LEVEL ALARMS AT 40% VOLUME AT THE CIC DURING THE TIME OF THE REPORTED PT INCIDENT. ALARM GRAPH STRIPS FROM THE DASH MONITOR WERE NOT PROVIDED TO CONFIRM THE ALARM VOLUME SETTING AT THE DASH DURING THE TIME OF THE ALARM. THE SITE INDICATED THE DASH VOLUME WAS SET TO ZERO. BASED ON THE INFO AVAILABLE, THE DASH 2000 PERFORMED TO SPECIFICATIONS. | 1
1944824 | THE CUSTOMER STATED THAT THE WASTE CONTAINER OF THE CELL-DYN RUBY ANALYZER OVERFLOWS WHEN IT IS FULL. THE WASTE OUTLET TUBING ASSEMBLY WAS REPLACED IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. | 0
1951202 | DISCREPANT CBC RESULTS WERE GENERATED BY THE CELLDYN 1800 ANALYZER. THE INITIAL RESULTS GENERATED WERE: WBC = 5,200/UL, RBC = 4.34X10E6/UL, HGB = 11.9 G/DL, HCT = 41.1%, PLATELET = 216,000/UL. THE SAMPLE WAS REPEATED 2 TIMES WITH THE FOLLOWING RESULTS: WBC = 2,900/UL AND 4,000/UL, RBC = 2.65X10E6/UL AND 3.45X10E6/UL, HGB = 7.0 G/DL AND 9.3 G/DL, HCT = 24.9% AND 32.4%, PLATELET = 121,000/UL AND 159,000/UL. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT AND NOTICED THAT THERE WAS NO BUBBLE MIX. THE FSR FOUND A BAD SAMPLE SYRINGE AND FLOW PANEL TUBINGS/FITTINGS; ALL WERE REPLACED BY THE FSR. THE FSR PROCEEDED TO RUN REPLICATES AND QUALITY CONTROLS, WHICH PASSED. THE FSR VERIFIED PROPER INSTRUMENT OPERATION, NO FURTHER INVESTIGATION WAS REQUIRED. THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE PREVENTIVE MAINTENANCE SCHEDULE STATES THAT REPLACEMENT OF THE SAMPLE SYRINGE IS AN AS-REQUIRED PROCEDURE. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1952147 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE HAVE CHANGED OUR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. THE CUSTOMER STATED THE PLASMA WAS VERY LIPEMIC. THE DONOR WAS ABOUT 20 MINUTES INTO THE PROCEDURE WHEN THE CUSTOMER NOTICED POSSIBLE HEMOLYSIS IN THE TUBING SET. THE CUSTOMER NOTICED A DARK PINK CLOUD IN THE PLATELET PRODUCT WHICH THEY HAD NOT SEEN BEFORE. THE CUSTOMER STATED ALL THE PLASMA IN THE TUBING SET LOOKED SLIGHTLY RED IN COLOR, AND WAS TRANSLUCENT. THE CUSTOMER ENDED THE PROCEDURE IMMEDIATELY. THE DONOR HAD NO REACTION AT ALL AND THE MEDICAL DIRECTOR WAS NOTIFIED. PER THE CUSTOMER, THE TUBING SET WAS NOT AVAILABLE FOR RETURN. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR HEMOLYSIS. (B)(4). ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HEMOLYSIS REPORTED DURING THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND TRIMA ACCEL OPERATED AS INTENDED. ON (B)(6) 2010, THE CARIDIANBCT CLINICAL SUPPORT STAFF RECEIVED THE TUBING SET LOT NUMBER FROM THE CUSTOMER AND DISCUSSED WITH THE CUSTOMER THAT OUR RDF ANALYSIS DID NOT IDENTIFY ANY ISSUES WITH EQUIPMENT. SERVICE OF THE EQUIPMENT HAD BEEN PERFORMED AS WELL AND THE CUSTOMER WAS SATISFIED. THE CUSTOMER IS CURRENTLY USING THE TRIMA. PER THE CUSTOMER, THE DONOR WAS REPORTEDLY FINE ON FOLLOW-UP FROM THEIR SITE. THE CUSTOMER WAS NOT ABLE TO DETERMINE IF HEMOLYSIS OCCURRED. A REVIEW OF THE MANUFACTURING RECORDS WAS CONDUCTED BY THE COMPLAINT INVESTIGATOR IN RELATION TO THIS FAILURE MODE. THERE WAS ONE CSP FAILURE FOR AN UNRELATED MANUFACTURING ERROR THAT WAS RESOLVED PER STANDARD PROCEDURES. ADDITIONALLY, FOR THIS LOT THERE WERE NO PDFS OR PERMIT ECO, NO SUT OBSERVATIONS, AND THE PRODUCT MET ALL ACCEPTANCE CRITERIA FOR RELEASE. CONCLUSION: NO ROOT CAUSE COULD BE DETERMINED AS FROM THE INFO AVAILABLE OUR PRODUCT MET SPECIFICATIONS, ALTHOUGH THE DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. PER INFO FROM THE CUSTOMER, THEY BELIEVE THAT THIS WAS AN RBC SPILLOVER DUE TO THE LIPEMIA. ADDITIONALLY AS THE RDF DEMONSTRATED THAT THE TRIMA OPERATED WITHIN SPECIFICATIONS, A PROCEDURAL ERROR OUTSIDE OF TRIMA MAY HAVE OCCURRED. | 0
1952891 | THE CUSTOMER REQUESTED THE LETTER OF FA27SEP2010 AND AN ABBOTT FIELD REPRESENTATIVE INSPECTED THE CUSTOMER'S CELL-DYN ANALYZER. THE FSR IDENTIFIED AN INCORRECT POWER SUPPLY FUSE (6 AMPERE) INSTALLED IN THE ANALYZER AND REPLACED IT WITH THE CORRECT FUSE (5 AMPERE). NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN EXPANDED INVESTIGATION THROUGH CORRECTION AND REMOVAL FA27SEP2010 WAS CONDUCTED TO EVALUATE THIS ISSUE. THE ROOT CAUSE IDENTIFIED FOR THIS ISSUE IS INSTALLATION OF THE INCORRECT FUSE BY EITHER THE CUSTOMER OR THE FIELD SERVICE REPRESENTATIVE. A PRODUCT CORRECTION LETTER WAS SENT TO THE CUSTOMERS ALONG WITH THE FUSE LABEL TO BE AFFIXED TO THE BACK OF THE INSTRUMENT INSTRUCTING THE CUSTOMER TO CHECK IF THE FUSE MATCHES THE VOLTAGE OF OPERATION AND TO INSTALL THE CORRECT FUSE (PROVIDED IN THE ACCESSORY KIT SHIPPED WITH THE ANALYZER) FOLLOWING REPLACEMENT INSTRUCTIONS PROVIDED IN THE LETTER. THE CUSTOMERS ARE ALSO INSTRUCTED TO ORDER THE CORRECT FUSE IF IT IS NOT PROVIDED IN THE ACCESSORY KIT. | 0
1958035 | A CUSTOMER CONTACTED BECKMAN COULTER INC., (BCI) AND REPORTED THAT HEMOGLOBIN (HGB) FAILED AT START UP AND THE COULTER ACT DIFF 2 ANALYZER GENERATED LOW HEMATOCRIT RESULTS FOR A PATIENT SAMPLE. PRINTOUTS WERE REQUESTED MULTIPLE TIMES, BUT WERE NOT PROVIDED TO DETERMINE WHICH PARAMETERS WERE AFFECTED. THE RESULTS WERE REPORTED TO THE DOCTOR AND THE PATIENT WAS SENT TO THE HOSPITAL FOR A BLOOD TRANSFUSION BASED ON THE ACT DIFF 2 RESULTS. THE PATIENT WAS RE-DRAWN AT THE HOSPITAL, WHERE THE COMPLETE BLOOD COUNT (CBC) RESULTS DID NOT SHOW THE NEED FOR A TRANSFUSION. QC WAS RUN BEFORE AND AFTER THE EVENT AND RESULTS RECOVERED WITHIN RANGE. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED: THE FSE REPLACED THE HGB LAMP AND ADJUSTED THE VOLTAGES TO SPECS AND VERIFIED THE UNIT PER ESTABLISHED PROCEDURES. THE INSTRUMENT MEETS PERFORMANCE SPECIFICATIONS. A CLEAR ROOT CAUSE FOR THIS EVENT IS UNKNOWN. | 0
1958050 | A CUSTOMER CONTACTED BECKMAN COULTER INC., (BCI) REGARDING PATIENT RESULTS GENERATED BY THE COULTER ACT DIFF 2 ANALYZER THAT DID NOT MATCH RESULTS OBTAINED AT THE HOSPITAL. IT IS UNKNOWN WHAT PARAMETERS WERE AFFECTED OR THE EXACT DATE THE ERRONEOUS RESULTS WERE GENERATED. ERRONEOUS RESULTS WERE REPORTED OUTSIDE THE LAB. THE PATIENT SPECIMEN WAS SENT TO A REFERENCE LAB AND THERE WAS NO CHANGE TO PATIENT TREATMENT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED: THE FSE RAN QC REPRODUCIBILITY AND CARRY OVER AND OBTAINED GOOD RESULTS. THE FSE VERIFIED INSTRUMENT OPERATIONS; RESULTS MEET PUBLISHED PERFORMANCE SPECIFICATIONS. A CLEAR ROOT CAUSE FOR THIS EVENT HAS NOT BEEN DETERMINED TO DATE. | 0
1959871 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED DISCREPANT GLUCOSE RESULTS. AN INITIAL RESULT OF 182 MG/DL WAS GENERATED. THE SAMPLE WAS REPEATED WITH RESULTS OF 129, 95, AND 111 MG/DL GENERATED. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR RESOLVED THE ISSUE BY PERFORMING SERVICE ON THE CUVETTE WASHER AND ALIGNING THE MIXERS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY GLUCOSE REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC/DISCREPANT RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. | 0
1959943 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED DISCREPANT GLUCOSE RESULTS. AN INITIAL RESULT OF 314 MG/DL WAS GENERATED. THE SAMPLE WAS REPEATED AND A RESULT OF 156 MG/DL WAS GENERATED. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR RESOLVED THE ISSUE BY PERFORMING SERVICE ON THE CUVETTE WASHER AND ALIGNING THE MIXERS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY GLUCOSE REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC/DISCREPANT RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1960002 | THE ACCOUNT STATED THAT THE ARCHITECT C8000 ANALYZER GENERATED DISCREPANT GLUCOSE RESULTS. AN INITIAL RESULT OF 156 MG/DL WAS GENERATED. THE SAMPLE WAS REPEATED WITH RESULTS OF 90, 78, AND 122 MG/DL GENERATED. THE RESULTS WERE NOT REPORTED AND THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). A FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR RESOLVED THE ISSUE BY PERFORMING SERVICE ON THE CUVETTE WASHER AND ALIGNING THE MIXERS. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE CLINICAL CHEMISTRY GLUCOSE REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC/DISCREPANT RESULTS. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1960039 | THE CUSTOMER REPORTED THAT THE CENTRICITY CARDIOLOGY CA1000 MEASUREMENT CALCULATIONS ARE NOT CORRECT. UPON FURTHER REVIEW, IT WAS DETERMINED THAT THE CUSTOMER IS PERFORMING THEIR MEASUREMENTS INCORRECTLY RESULTING IN INCORRECT INFO BEING ENTERED INTO THE CA1000 TO PERFORM THE CALCULATION. THE STUDY WAS SENT FROM THE CATH LAB TO AN ENTERPRISE ARCHIVE (EA) FOR STORAGE AND THEN PULLED FROM THE EA TO A CA1000 WORKSTATION FOR REVIEW. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
1960241 | IT WAS REPORTED THAT A FIELD ENGINEER EXPERIENCED LOWER BACK PAIN DURING PLANNED MAINTENANCE OF THE VOLUSON SYSTEM AFTER REMOVING AND REINSTALLING A CARD RACK WEIGHING APPROX 35 KG (77 LBS). REMOVAL WAS NECESSARY TO ACCESS AND SERVICE THE SYSTEM, BUT SERVICE INSTRUCTIONS SPECIFY THIS IS A TWO PERSON OPERATION. THE FIELD ENGINEER COMPLETED THE SERVICE AND DROVE HOME WITH BACK PAIN. THE FOLLOWING DAY, THE FIELD ENGINEER FURTHER EXPERIENCED SHARP BACK PAIN WHILE LIFTING, CARRYING, AND LOADING A 15 KG (33 LB) BOX INTO HIS CAR. THE COMBINED INJURY RESULTED IN LOST WORK TIME OF 6 DAYS. THIS INCIDENT IS CONSIDERED A SERVICE USE ERROR. THE FIELD ENGINEER REPORTED LIFTING THE CARD RACK WEIGHING APPROX 35 KG ALONE WHICH CONTRIBUTED TO A BACK INJURY. AS PART OF THE INVESTIGATION OF THIS INCIDENT, THE SERVICE INSTRUCTION WAS REVIEWED. THE SERVICE INSTRUCTION CLEARLY STATES: "TWO PEOPLE ARE REQUIRED WHEN MOVING ON INCLINES, OR LIFTING MORE THAN 16 KG". NO FURTHER ACTIONS ARE PLANNED AT THIS TIME. | 0
1960268 | REPORTED OBSERVATION: PATIENT DEVELOPED CANALICULITIS AFTER INSERTION OF AN OASIS MEDICAL FORM FIT HYDROGEL CANALICULAR PLUG. PRODUCT REFERENCE: (B)(4), PRODUCT LOT NUMBER: UNK, PRODUCT INSERTED ON: UNKNOWN. DATE OF COMPLICATION: REPORTED TO OASIS MEDICAL, INC. ON (B)(6) 2010. ATTENDING DOCTOR SUCCESSFULLY IRRIGATED THE PLUG FROM THE CANALICULUS. PT TREATED WITH AN ANTIBIOTIC TREATMENT. THE DEVICE WAS NOT RETURNED FOR EVALUATION. SPECIFIC PRODUCT LOT INFO WAS NOT PROVIDED WITH THE INITIAL REPORT OF THE COMPLICATIONS. THE DATE OF PLUG INSERTION AND REMOVAL WAS NOT REPORTED WITH THE INITIAL REPORT OF COMPLICATIONS. THE ATTENDING DOCTOR WAS ABLE TO IRRIGATE THE PLUG FROM THE CANALICULUS. A REQUEST FOR CASE SPECIFIC INFO HAS BEEN REQUESTED FROM THE REPORTING DOCTOR. SUPPLEMENTAL INFO WILL BE SUBMITTED WHEN RECEIVED. | 0
1960294 | ON (B)(6) 2010, THE CUSTOMER REPORTED THAT THE OPERATING ROOM TABLE WAS MOVING UP AND DOWN AND TILTING ON ITS OWN, WITHOUT THE MANIPULATION OF THE OPERATING ROOM TABLE PENDANT (CONTROLLER) BY THE OPERATOR. THERE WAS NO REPORT OF INJURY TO EITHER A PT OR CLINICAL STAFF. AN IMRIS CUSTOMER SUPPORT SPECIALIST INVESTIGATED THE COMPLAINT ON-SITE, BUT WAS UNABLE TO REPRODUCE THE INCIDENT. THE OPERATING ROOM TABLE PENDANT WAS LATER RETURNED TO IMRIS FOR INVESTIGATION, AND THE INCIDENT AGAIN COULD NOT BE DUPLICATED. AS THIS WAS THE FIRST OCCURRENCE, WITH NO REPORT OF INJURY AND THE INCIDENT COULD NOT BE REPRODUCED, IMRIS DETERMINED THAT THE RESULTS OF THE INVESTIGATION DID NOT WARRANT AN MDR. HOWEVER, UPON DISCOVERY OF A SECOND INCIDENT AT ANOTHER CUSTOMER SITE, IMRIS DETERMINED THAT THIS CASE SHOULD BE RE-OPENED PENDING THE CONCLUSION OF THE INVESTIGATION INTO THE ABOVE-MENTIONED CASE, AND THAT AN MDR BE SUBMITTED. . | 0
1962133 | THE CUSTOMER STATED AN ARCHITECT I1000SR ANALYZER GENERATED A FALSE POSITIVE B-HCG RESULT FOR ONE PATIENT SAMPLE. THE ARCHITECT GENERATED AN INITIAL B-HCG RESULT OF 67.03 MIU/ML, AND REPEAT B-HCG RESULTS OF 7.94 MIU/ML AND <1.20 MIU/ML. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><(B)(4). REVIEW OF DOCUMENTATION IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. A DEFICIENCY WAS IDENTIFIED; POSSIBLY DUE TO SAMPLE OR REAGENT CARRY-OVER. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. A FINAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
1979588 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE WHITE BLOOD COUNT (WBC) PARAMETER REPORTED HALF OF THE EXPECTED VALUE ON THE COULTER HMX AUTOLOADER ANALYZER ON THIRTEEN (13) DIFFERENT PATIENTS. RESULTS WERE CONFIRMED TO BE ERRONEOUS WHEN COMPARED TO THE MANUAL SLIDE AND WHEN RAN ON AN ALTERNATE INSTRUMENT THAT CORRELATED WELL WITH THE EXPECTED RESULTS. ALL OF THE RESULTS WERE FLAGGED WITH SYSTEM GENERATED MESSAGES. THE ERRONEOUS RESULTS WERE NOT REPORTED OUT OF THE LABORATORY. ONE EXAMPLE OF SUCH RESULTS, PROVIDED BY THE CUSTOMER. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT IN THIS EVENT. SAMPLES WERE COLLECTED VIA VENIPUNCTURE IN AN EDTA 4ML SAMPLE TUBE. SAMPLES WERE NOT RERUN TO CONFIRM ACCURACY TO THE LAST ACCEPTABLE CONTROL RUN. UNIT WAS NOT PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS AND REPRODUCIBILITY AT THE TIME OF THE INCIDENT. CONTROLS RUN AT THE TIME OF THE INCIDENT RECOVERED HALF OF THE EXPECTED VALUES. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(6) 2011 FOR THIS EVENT. THE FSE INSPECTED THE OPERATION OF THE WBC BATH AND PERFORMED TROUBLESHOOTING. THE FSE REPLACED THE WBC BATH, RERAN SAMPLES, AND COMPARED RESULTS. THE FSE CONFIRMED VIA EMAIL THAT THE WBC BATH RESOLVED THE ERRONEOUS WBC ISSUE. THUS, THE ROOT CAUSE WAS DETERMINED TO BE A DEFECTIVE WBC BATH. PER BCI PRODUCT LABELING, HMX HEMATOLOGY ANALYZER WITH AUTOLOADER PROVIDES THE CUSTOMER WITH VARIOUS DATA, FLAGS, AND GRAPHICAL REPRESENTATIONS THAT ARE DESIGNED FOR USE AS AN INTEGRATED REPORT. TO ASSURE THAT THE SYSTEM PROVIDES THE MOST MEANINGFUL AND ACCURATE RESULTS, SET UP ALL DEFINITIVE LIMITS ACCORDING TO THE LABORATORY'S ESTABLISHED REFERENCE RANGES AND USE ALL REPORT OUTPUTS FOR DECISION MAKING PURPOSES. | 0
1979631 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE CUSTOMER GOT WASTE FLUID SPLASHED ON HER FACE WHEN SHE ATTEMPTED TO SECURE THE LINE ON THE COULTER HMX AUTOLOADER ANALYZER THAT CAME OFF THE WASTE CUBE TO PREVENT OTHER EMPLOYEES FROM BEING SPRAYED. THE CUSTOMER WAS WEARING PROPER PERSONAL PROTECTIVE EQUIPMENT (PPE). THE FLUID DID NOT MAKE CONTACT WITH MOUTH, NOSE, EYES, OR OPEN WOUNDS. THE CUSTOMER DID NOT SEEK OUT MEDICAL ATTENTION. THE DATE BEFORE THE INCIDENT, A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND REPLACED THE WASTE SENSOR FOR THE WASTE CUBE. THE ROOT CAUSE OF THIS EVENT IS UNKNOWN. AS PER PRODUCT LABELING, BCI URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
1981801 | THE CUSTOMER NOTES THAT ERRATIC RESULTS FOR THE ICT MODULE ASSAYS HAVE BEEN OCCURRING SINCE A NEW ICT MODULE WAS INSTALLED ON (B)(4) 2011 ON THE ARCHITECT C8000 ANALYZER. THE CUSTOMER GAVE THE FOLLOWING EXAMPLE: (B)(6), SODIUM: 116, POTASSIUM: 3.9, CHLORIDE (MMOL/L): 94; RETEST: 137, 4.6, 105. NO SUSPECT RESULTS HAVE BEEN REPORTED FROM THE LAB. THE CUSTOMER REMOVED THE ICT MODULE AND REMOVED, CLEANED AND REINSTALLED THE O-RINGS. THE CUSTOMER NOTICED A BUILD-UP OF DRIED BUFFER SALTS ON THE FITTING TO THE TOP OF THE MODULE, INDICATIVE OF A LOOSE CONNECTION. THE CUSTOMER CLEANED THE FITTING AND CONNECTED THE TUBING CORRECTLY (NOT OVERTIGHT). THE CUSTOMER THEN RAN A PRECISION RUN WITH "EXCELLENT" RESULTS. THE CUSTOMER VERIFIED THAT THEY HAD NO FURTHER ERRATIC ICT MODULE ASSAY RESULTS SINCE CORRECTING THE TUBING CONNECTION. HOWEVER, THIS DID NOT RESOLVE THE ISSUE AND A SERVICE CALL WAS INITIATED. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) VISITED THE CUSTOMER SITE AND PERFORMED ROUTINE INSTRUMENT TROUBLESHOOTING AND COMPONENT REPLACEMENT, RETURNING THE SYSTEM TO WITHIN SPECIFICATIONS. THE ISSUE WAS RESOLVED BY REPLACING THE WORN 1.0 ML SYRINGES FOR THE INTEGRATED CHIP TECHNOLOGY (ICT) ASPIRATION PUMP DUE TO SALT BUILDUP UNDER THE PLUNGER. THE ICT PROBE WAS ALSO CLEANED AND THE CUVETTE DRYER TIPS WERE REALIGNED. A PRECISION RUN WAS PERFORMED ON THE ICT MODULE AND RESULTS WERE WITHIN SPECIFICATIONS. THE ASSAYS WERE CALIBRATED AND ALL CONTROLS WERE RUN WITH ACCEPTABLE RESULTS GENERATED. THE ARCHITECT SYSTEM OPERATIONS MANUAL (LIST NUMBER 201837-108) CONTAINS INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
1986786 | A CUSTOMER CONTACTED BECKMAN COULTER INC., (BCI) STATED THAT THE LABORATORY TECHNICIAN WAS SPRAYED IN THE EYES WITH RETIC CLEARING SOLUTION (REAGENT B) FROM THE COULTER LH 750 ANALYZER. THE DAY BEFORE THIS EVENT OCCURRED, THE INSTRUMENT OPERATOR WAS TROUBLESHOOTING (TS) A RETIC CONTROL RECOVERY ISSUE AND LEFT THE SIDE COVER OFF AND RESET THE INSTRUMENT. THE FOLLOWING DAY, A LABORATORY TECHNICIAN (LAB TECH), JUST COMING ON SHIFT, WALKED BY THE ANALYZER AND HEARD A BEEPING NOISE. WHEN SHE LOOKED OVER AT THE INSTRUMENT SHE GOT SPRAYED WITH REAGENT B. THERE WAS CONTACT TO MUCOUS MEMBRANES (EYES) WITH ONLY REAGENT B. THE LAB TECH WAS WEARING GLOVES AND GOWN, BUT NO GOGGLES. THE LAB TECH WENT TO THE ER AND WAS TREATED WITH GENERAL CARE AND WAS ABLE TO RETURN TO WORK. PER FOLLOW UP INFORMATION, THE CUSTOMER INDICATED THAT THE LAB TECH IS DOING FINE. THERE WAS NO DEATH ASSOCIATED WITH THIS EVENT BUT IT RESULTED IN AN INJURY THAT REQUIRED MEDICAL ATTENTION. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED FOR THIS EVENT. THE FSE REPLACED THE RETIC SHEAR VALVE AND THE PINCH VALVE TUBING FOR THE CLEARING SOLUTION PUMP. INSTRUMENT IS CURRENTLY OPERATIONAL AND ALL CONTROLS ARE IN RANGE. THE ROOT CAUSE FOR THE SPRAYED REAGENT IS THE RETIC SHEAR VALVE NOT ALIGNING WITH THE REAGENT PORTS PROPERLY WHICH CAUSED EXCESSIVE BACK PRESSURE RESULTING IN THE FLUID LEAK AT THE "Y" FITTING. THE ROOT CAUSE OF THE CHEMICAL EXPOSURE IS THAT LACK OF EYE PERSONAL PROTECTIVE EQUIPMENT (PPE) WORN BY THE LAB TECH AND INSTRUMENT OPERATION BEING CONDUCTED WITH THE INSTRUMENT DOOR OPEN. PER LABELING, BECKMAN COULTER URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, WEARING PROTECTIVE EYE WEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS INSTRUMENT OR ANY OTHER AUTOMATED LABORATORY ANALYZER. PER LABELING, BECKMAN COULTER URGES ITS CUSTOMERS TO KEEP DOORS, COVERS AND PANELS CLOSED AND SECURED IN PLACE WHILE THE INSTRUMENT IS IN USE. | 0
1990008 | IT WAS REPORTED THAT THERE WERE PROBLEMS UPGRADING A DELL X50 HANDHELD TO VERSION 8.0 VNS SOFTWARE. AT THE INSTALLATION STEP TO SELECT CF CARD, THE MANUFACTURER REP PRESENT INDICATED THE CF CARD WAS NOT AN OPTION TO SELECT. RESEATING THE FLASHCARD DID NOT RESOLVE THE ISSUE. THE VERSION 7.1 FLASHCARD WAS INSERTED INTO THE HANDHELD, AND A HARD RESET WAS PERFORMED. THE SCREEN WAS FROZEN AT "CLEAR ALL DATA IN MEMORY" SCREEN, AND NONE OF THE BUTTONS WOULD WORK. TROUBLESHOOTING DID NOT RESOLVE THE ISSUE. THE HANDHELD AND SOFTWARE FLASHCARD WERE RETURNED TO MANUFACTURER FOR ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED VERSION 8.0 FLASHCARD AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE EVENT WAS DETERMINED TO BE ASSOCIATED WITH A DEFECTIVE FLASHCARD. A VISUAL ANALYSIS OF THE PRINTED CIRCUIT BOARD (PCB) IDENTIFIED THAT PIN 27 WAS NOT SOLDERED ONTO THE PCB CAUSING THE HANDHELD TO NOT RECOGNIZE THE FLASHCARD WHEN IT WAS INSERTED. ONCE THE PIN WAS SOLDERED ONTO THE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE HANDHELD REVEALED NO ANOMALIES AND PERFORMED ACCORDING TO SPECIFICATIONS. ANALYSIS OF THE VERSION 7.1 FLASHCARD REVEALED NO ANOMALIES AND PERFORMED ACCORDING TO SPECIFICATIONS. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY. | 0
1993351 | A (B)(6) SUPPLIER REPORTED AN INFANT APNEA MONITOR DID NOT ALARM DURING A PT EVENT. IT WAS REPORTED THAT THE INFANT ASPIRATED AND HAD A CYANOTIC EVENT. THE EXACT DATE AND TIME OF THE ALLEGED EVENT IS UNK, HOWEVER; THE DME STATED THAT IT RECEIVED A CALL REGARDING THE EVENT ON (B)(6) 2010. THE (B)(6) REPORTED THAT THE INFANT WAS HOSPITALIZED AND LATER RELEASED WITHOUT PT HARM. THE MFR RECEIVED THE DEVICE FOR EVAL AND DID NOT CONFIRM THE CUSTOMER'S COMPLAINT OF THE UNIT NOT ALARMING. THE APNEA MONITOR WAS VISUALLY EXAMINED AND TESTED BY THE MFR USING A SIMULATOR IN ACCORDANCE WITH THE SMARTMONITOR 2 CHECKOUT PROCEDURE MANUAL (PN (B)(4)). THE APNEA MONITOR DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS AND PASSED ALL REQUIRED TESTING. THE SMARTMONITOR 2 WAS SET UP WITH A 16 SECOND DELAY BEFORE RECORDING APNEAS AND A 20 SECOND DELAY BEFORE ANNUNCIATING AND RECORDING ALARM FOR AN APNEA CONDITION. THE SMARTMONITOR 2 WAS PROGRAMMED WITH PARAMETERS THAT WERE APPROPRIATE FOR RECORDING AND ALARMING FOR RESPIRATION AND HEART RATE EVENTS DURING INFANT MONITORING. THE APNEA MONITOR'S MEMORY DATA WAS DOWNLOADED AND ANALYZED BY TRAINED ASSOCIATES. THE DOWNLOADED MEMORY REVEALED THE APNEA MONITOR WAS IN USE FROM (B)(6) 2010. DURING THAT TIME, THERE WERE 56 RECORDED PT EVENTS. ALL PT EVENTS THAT MET THE PARAMETERS FOR AN ALARM CONDITION WERE ASSOCIATED WITH AUDIBLE AND VISUAL ALARMS. THE SMARTMONITOR 2 DEVICE IS NOT INTENDED TO BE USED TO MONITOR PTS FOR CYANOSIS AND HAS NO CAPABILITY TO DO SO. THE SMARTMONITOR 2 PARENTS' GUIDE (PN (B)(4)) STATES IN THE INDICATIONS FOR USE: "THE SMARTMONITOR 2 IS INTENDED FOR USE IN CONTINUOUS MONITORING OF HEART RATE AND RESPIRATION OF INFANT PTS IN A HOME, HOSP OR PORTABLE ENVIRONMENT. ITS PRIMARY FUNCTION IS DETECTION OF CENTRAL APNEA. ITS SECONDARY FUNCTION IS MEASUREMENT OF HEART RATE." THE SMART MONITOR 2 DEVICE IS NOT INTENDED TO PREVENT LOSS OF BREATHING OR HEART ACTIVITY. THE SMARTMONITOR 2 PARENTS' GUIDE (B)(4) FURTHER STATES: "THE SMARTMONITOR 2 IS A MONITORING DEVICE ONLY. IT DOES NOT PREVENT THE LOSS OF BREATHING OR HEART ACTIVITY, NOR WILL IT RESTORE BREATHING OR HEART ACTIVITY. IT WILL NOT PREVENT DEATH. ANYONE USING THE SMARTMONITOR 2 TO MONITOR AN INFANT SHOULD BE TRAINED IN CURRENT INFANT CARDIOPULMONARY RESUSCITATION (CPR), WHICH IS A PROPER WAY TO RESTORE BREATHING AND HEART ACTIVITY." THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY, AND THE MFR WAS NOT ABLE TO CONFIRM THE ALLEGATION OF ALARM FAILURE DURING A PT EVENT. THE MONITOR PASSED ALL REQUIRED TESTING AND DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS. BASED ON ALL AVAILABLE INFO, THE MFR CONCLUDES THAT THE DEVICE DOES FUNCTION TO SPECIFICATION AND THAT NO FURTHER ACTION IS APPROPRIATE. | 1
1994034 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH THE CAPTURE-R READY SCREEN (3) ASSAY ON THE GALILEO ECHO INSTRUMENT. UPON REVIEW OF THE BATCH FILE IMAGES FOR THE INITIAL RUN, IT APPEARED THAT NO SERUM/PLASMA WAS ADDED TO THE TEST WELLS. AN IMMUCOR SERVICE ENGINEER IDENTIFIED A LEAK IN THE PROBE RINSE TUBING AND AROUND THE SYRINGE MANIFOLD CAP AND SEAL. BOTH WERE REPLACED, INCLUDING THE MANIFOLD CHECK VALVES. THE UNEXPECTED REACTION CHECKLIST AND QC PERFORMED AS EXPECTED. THE EXPECTED RESULTS WERE OBTAINED WITH SAMPLE RUNS. INSTRUMENT IS OPERATING ACCORDING TO SPECIFICATIONS. | 0
1994170 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT ONE HEMOGARD VACUTAINER PLASTIC TUBE CAP CAME OFF WHILE THE SAMPLE WAS BEING RUN ON THE COULTER LH 750 ANALYZER, CONSEQUENTLY SPILLING BLOOD INSIDE THE INSTRUMENT. THE TUBE WAS IN A CASSETTE ON THE ROCKER BED WHEN THE INCIDENT OCCURRED. NO INJURIES OCCURRED AND MEDICAL ATTENTION WAS NOT SOUGHT. THERE WERE NO REPORTS OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. NO CHANGE TO PATIENT TREATMENT ATTRIBUTED TO OR CONNECTED TO THIS COMPLAINT. THE CUSTOMER CLEANED THE SPILL ACCORDING TO THE DEFINED LABORATORY PROTOCOL. AFTER CLEANING THE CUSTOMER WAS GETTING "AUTO STOP" ALARMS AND STRIPPER PLATE ERRORS. THE CUSTOMER INDICATED THAT THE TUBE WAS NOT OPENED PRIOR TO RUNNING AND THE CASSETTE WAS CORRECTLY PLACED ON THE LOADING DOCK AND ROCKER BED IN THE USUAL MANNER. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND INDICATED THAT THE STRIPPER PLATE ASSEMBLY WAS DISASSEMBLED, SO THE FSE CLEANED AND REASSEMBLED IT. THE ALIGNMENT OF THE STRIPPER PLATE WITH SAMPLE TUBE WAS ALSO CHECKED WITH SATISFACTORY RESULTS. MULTIPLE PATIENT SAMPLES WERE RUN WITHOUT ANY ISSUES. THE ROOT CAUSE COULD NOT BE DETERMINED TO DATE FOR THIS EVENT. PER LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
1995433 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER ACT DIFF 2 ANALYZER WAS LEAKING AND WAS GENERATING LOW CONTROL RESULTS. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. THERE WAS NO EXPOSURE TO OPEN LESIONS OR MUCOUS MEMBRANES, AND NO ONE SOUGHT MEDICAL ATTENTION. THE CUSTOMER WAS INSTRUCTED TO TROUBLESHOOT AND FOUND THAT BOTH BATHS WERE OVERFLOWING. CUSTOMER MANUALLY DRAINED THE BATH LINES, DID A STARTUP, AND OPERATION WAS ACCEPTABLE. CUSTOMER THEN REPORTED A CLEAR LIQUID UNDER THE INSTRUMENT. CUSTOMER WAS GIVEN TROUBLESHOOTING INSTRUCTIONS AND A VACUUM FAILURE WAS OBSERVED. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND DISCOVERED THAT THE CAP CAME OFF THE FITTING ON THE VIC WHICH CAUSED THE VACUUM ERRORS. PER LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
1996539 | THE CUSTOMER REQUESTED ABBOTT SERVICE TO INSPECT THE CELL-DYN SAPPHIRE ANALYZER FOR MIX-HEAD ERRORS. THE INSTRUMENT WAS INSPECTED AND THE MIX-HEAD CABLE SENSOR WAS REPLACED ALONG WITH THE ASPIRATION BOTTOM SENSOR IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4): ASPIRATION BOTTOM SENSOR WORN-OUT FROM NORMAL USE. AN INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED AND THE ISSUE WAS RELATED TO A WORN-OUT PART DUE TO NORMAL USE. A MANDATORY TECHNICAL SERVICE BULLETIN (TSB) IS ISSUED WITH INSTRUCTIONS FOR FIELD SERVICE TO INSPECT AND REPLACE THE PART WHEN NECESSARY. THE PREVENTIVE MAINTENANCE PROCEDURES WILL BE UPDATED IN THE OPERATORS MANUAL TO INCLUDE INSPECTION OF THE ASPIRATION PROBE SENSOR EVERY SIX MONTHS AND REQUEST REPLACEMENT IF REQUIRED. A DESIGN IMPROVEMENT IS IN PROCESS TO IMPROVE THE CABLE AND SEAL THE BOARD AND SWITCH ASSEMBLIES. A CORRECTION THROUGH A FOLLOW-UP CORRECTION AND REMOVAL FA01OCT2010 WAS ISSUED TO INCLUDE INSTALLING A NEW SOFTWARE (V4). THE NEW V4 SOFTWARE RELEASED WITH FA01OCT2010 INCLUDES A DESIGN IMPROVEMENT FOR THE ASPIRATION BOTTOM SENSOR. (B)(4). | 0
2000413 | CUSTOMER REPORTED THAT THE ECHO (SOFTWARE V 1.2.0.19.2) READ A SAMPLE BARCODE INCORRECTLY. A SAMPLE WITH BARCODE (B)(4) WAS READ AS (B)(4) BY THE INSTRUMENT. CUSTOMER RELOADED SAMPLE ON THE INSTRUMENT AND THE BARCODE READ CORRECTLY. THE SAMPLE RESULTED AS O POSITIVE WITH NEGATIVE SCREEN RESULTS BOTH TIMES. REPLACED CONTROLLER PCB AND RAM CHIP ON REAGENT PCB DUE TO AN MISREAD BARCODE. PERFORMED QC TESTING AND RE-TESTED SAMPLES. SYSTEM IS OPERATING WITHIN SPECIFICATIONS WITH NO ERRORS OBSERVED. ADVISED CUSTOMER TO MONITOR. | 1
2001634 | A CUSTOMER REPORTED TO BECKMAN COULTER INC., (BCI) TO REPORT RECEIVING NEGATIVE TEST RESULTS WITH COLON CANCER PATIENTS THAT WERE TESTED WITH HEMOCCULT ICT. THE RESULT WAS NOT REPORTED OUT OF THE LABORATORY; HENCE, NO INJURY OR DEATH HAS BEEN REPORTED IN CONNECTION WITH THIS EVENT. THE TEST WAS ON FECAL SAMPLE. INVESTIGATION BY RAPIDS QC DID NOT CONFIRM THE CUSTOMER COMPLAINT. SERVICE EXPLAINED THE IMPORTANCE OF PROPER BUFFER ADDITION TECHNIQUE TO THE CUSTOMER. THE CUSTOMER NOTED THAT THEY MAY NOT HAVE PERFORMED THE TEST AS PER THE PRODUCT INSTRUCTIONS. | 0
2001973 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THE COULTER LH500 ANALYZER RECOVERED ERRONEOUS CBC RESULTS (WBC, RBC, HGB, PLT). THE ERRONEOUS HGB VALUE WAS REPORTED OUT AFTER THE SAMPLE WAS REPEATED AND RECOVERED SIMILAR RESULTS. THE CUSTOMER INDICATED THAT THIS SPECIMEN WAS A SHORT DRAW. THE PHYSICIAN QUESTIONED THE HGB VALUE BECAUSE IT WAS NOT CONSISTENT WITH THE PATIENT'S PREVIOUS HISTORY. A SAMPLE WAS REDRAWN PER THE PHYSICIAN'S REQUEST AND THE WBC, RBC AND HGB RESULTS WERE HIGHER BUT THE PLATELET VALUE WAS LOWER THAN EXPECTED. THE SAMPLE WAS REPEATED 2 ADDITIONAL TIMES AND RECOVERED VALUES THAT THE LAB CONSIDERED CORRECT. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT ATTRIBUTED TO OR CONNECTED WITH THIS EVENT. CONTROLS WERE RUN BEFORE THE INCIDENT. THE CUSTOMER REPORTED FAILED REPRODUCIBILITY FOR HGB AND PLT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND INDICATED THAT THE CUSTOMER MISINTERPRETED REPRODUCIBILITY SPECS. FSE TRAINED THE CUSTOMER. FSE REVIEWED REPRODUCIBILITY DATE AND CARRYOVER WHICH WAS ACCEPTABLE. | 0
2003975 | IT WAS REPORTED THAT WHILE CONNECTING TO THE CARTO XP SYSTEM, THE USER INADVERTENTLY CONNECTED THE ST. JUDE SPIRAL TO THE MAP PORT AND THE NAVISTAR THERMOCOOL TO THE LASSO PORT (XPRESS SYSTEM) AND AS A RESULT, BLOOD APPEARED TO BE BOILING AND THE PATIENT EXPERIENCED ELEVATED ST SEGMENTS ON THE BS ECG AS A RESULT. AIR IN THE LA WAS ALSO NOTED ON ECHO. NO ABLATION WAS PERFORMED ON THE PATIENT AND THE CASE WAS CANCELLED. THE PATIENT WAS KEPT UNDER OBSERVATION OVERNIGHT. DUE TO SOME CONFLICTING INFORMATION BETWEEN THE CUSTOMER'S COMMUNICATION AND THE "USER FACILITY REPORT (B)(4)" RECEIVED FROM THE FDA, BWI HAD TO REQUEST ADDITIONAL CLARIFICATION FROM THE FACILITY, WHICH INCLUDED A DEMONSTRATION WITHIN THE FACILITY FOR THE PRODUCT'S PERFORMANCE. (B)(4). FDA REQUEST FOR REPORT # (B)(4). (B)(4). BWI FIELD SERVICE ENGINEER AND THE CLINICAL ACCOUNT SPECIALIST ATTEMPTED TO RECREATE THE EVENT IN A WET BATH WITH BIOSENSE WEBSTER CATHETERS; HOWEVER, THEY WERE NOT ABLE TO REPRODUCE THE PROBLEM. ONLY WHEN THE ST JUDE CATHETER WAS CONNECTED TO THE CARTO XP SYSTEM, THE EVENT COULD BE REPRODUCED AND THE BUBBLES FORMED. ADDITIONALLY, IT WAS COMMUNICATED BY THE CUSTOMER AS WELL AS VERIFIED BY A FIELD SERVICE ENGINEER THAT THE SYSTEM DID NOT MALFUNCTION AND THE CUSTOMER DECLINED SERVICE FOR THE CARTO XP SYSTEM. NO FURTHER SERVICE NEEDED. ACCORDING TO IFU FOR CARTO XP SYSTEM, THE SYSTEM IS DESIGNED TO BE USED ONLY WITH SINGLE-USE CARDIAC CATHETERS EQUIPPED WITH BIOSENSE WEBSTER SENSORS. THIS WAS A CONTRAINDICATED USE OF THE EQUIPMENT AND THE ONLY MALFUNCTION OBSERVED WAS DUE TO A NON-BWI PRODUCT INTERFACE WITH THE PATIENT THAT CAUSED ELEVATED ST SEGMENT AND POSSIBLE BUBBLES THAT CAN BE AVOIDED BY FOLLOWING THE IFU DIRECTIONS IN THE EQUIPMENT MANUAL. THIS INFORMATION WAS RELAYED TO THE EP LAB AND THE PHYSICIAN. A DHR REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT.><DETAILED EVENT INFORMATION: AFTER TWO SUCCESSFUL TRANSSEPTAL PUNCTURES, IT WAS NOTICED THE MAP CATHETER WAS NOT RECOGNIZED BY CARTO XP SYSTEM. AT THE SAME TIME IT WAS NOTED THAT THERE WAS A LOT OF "CONTRAST/AIR/BUBBLES" ON THE SEQUOIA IMAGE. ADDITIONALLY, ST ELEVATIONS WERE NOTED ON THE BS ECG. UPON RECONNECTING THE HYPERTRONIC MAP CONNECTOR FROM THE MAP PORT ON THE PIU, THE CATHETER RECOGNITION ISSUE WAS NOT RESOLVED. WHEN ASKED TO RECONNECT THE CATHETER, IT WAS OBSERVED THAT THE CATHETER WAS CONNECTED TO THE SJM HD REFLEXION SPIRAL THAT WAS CONNECTED TO THE HYPERTRONIC TO THE PIU. WHEN THEY WERE REVERSED, BWI CATHETER WAS RECOGNIZED BY THE SYSTEM, AND THE ICE IMAGE WENT BACK TO NORMAL. IT WAS INDICATED TO THE PHYSICIAN THAT THIS WAS PROBABLY THE CAUSE OF THE NOISE ON ICE. THE PHYSICIAN THOUGHT THERE WAS AIR IN THE IV LINES CAUSING THE BUBBLES ON THE IMAGE, AND PROCEEDED TO INVESTIGATE. AFTER RULING OUT ANY AIR INFILTRATION, HE TRIED REVERSING THE CATHETERS + CONNECTORS AGAIN TO SEE IF IT WAS REPEATABLE. INDEED IT WAS. SINCE THE PHYSICIAN WAS UNSURE IF THIS WAS JUST ELECTRICAL NOISE OR ACTUAL BUBBLES ON THE ICE, AND WITH THE ECG CHANGES, THE CASE WAS ABORTED. THE FIELD SERVICE ENGINEER AND THE CLINICAL ACCOUNT SPECIALIST RECREATED THE EVENT IN A WET BATH WITH THESE CATHETERS. RULING OUT SYSTEM COMPONENTS ONE AT A TIME, IT TURNED OUT THAT WHEN THE ST. JUDE REFLEXION CATHETER WAS CONNECTED TO THE PIU, INDEED BUBBLES FORMED, ON SEVERAL ELECTRODES. THIS WAS REPEATABLE WITH THE STOCKERT, CFP, STIMULATION CABLE AND ICE CATHETER DISCONNECTED. ADDITIONALLY THE WET BATH WAS TESTED NEAR AND AWAY FROM THE LOCATION PAD. THIS INFORMATION WAS RELAYED TO THE PHYSICIAN. WHAT HE THOUGHT WAS A CARTO CATHETER MALFUNCTION ISSUE WAS PROVED TO BE UNTRUE. HE WAS SATISFIED THAT THERE WERE NO FUNCTIONAL ISSUES WITH BWI HARDWARE OR DISPOSABLE CATHETERS. HOWEVER, HE WAS VERY CONCERNED THAT THE REFLEXION COULD BE CONNECTED TO OUR SYSTEM WITH OUR CONNECTOR. ADDITIONALLY, IT WAS BROUGHT TO THE FSE AND CAS'S ATTENTION THAT WHILE THE PHYSICIAN WAS MANEUVERING THE SJM CATHETER OUT OF THE SRO SHEATH, THAT IT OFTEN DEPLOYED IN AN ABNORMAL CONFIGURATION. THE PATIENT WAS MALE, (B)(6) AND DIABETIC (INSULIN DEPENDENT). THE PHYSICIAN'S PROGNOSIS INDICATES THAT THE PATIENT IS IN NO IMMINENT DANGER AND THAT THE PROCEDURE WILL BE PERFORMED AT A LATER DATE AS SEEN FIT BY THE PHYSICIAN AND THE PATIENT. THE PATIENT WAS STABLE AND WAS RELEASED SUBSEQUENTLY. NO ADDITIONAL INFORMATION IS AVAILABLE REGARDING THE PATIENT OR THE PROCEDURE. CONCOMITANT PRODUCTS: NAVI-STAR THERMO-COOL ELECTROPHYSIOLOGY CATHETER, US CATALOG # NI75TCFH, LOT # 14087282. STOCKERT 70 RF GENERATOR, US CATALOG # S7001, (B)(4). COOLFLOW IRRIGATION PUMP, US CATALOG # CFP002, (B)(4). COMPETITOR PRODUCTS INVOLVED: SJM HD REFLEXION CATHETER, US CATALOG # D402864, LOT # 3210542. SJM DUO DECA CATHETER, US CATALOG # UNKNOWN, LOT # UNKNOWN. (B)(4). | 0
2005178 | THE CUSTOMER STATES THAT WHEN PROCESSING A PATIENT SAMPLE USING THE ARCHITECT C16000 ANALYZER THAT A DISCREPANTLY LOW SODIUM RESULT OF 115 MMOL/L WAS GENERATED COMPARED TO THE REPEAT RESULT OF 136 MMOL/L. NO ADVERSE OUTCOMES WERE REPORTED RELATED TO THIS ISSUE. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). AN EVALUATION WAS PERFORMED BY REVIEWING THE COMPLAINT TEXT, INSTRUMENT SERVICE HISTORY, AND OTHER CUSTOMER COMPLAINTS FOR SIMILAR ISSUES. THE ISSUE IS ADDRESSED IN THE PRODUCT LABELING WITH CORRECTIVE ACTION RECOMMENDATIONS. THE CUSTOMER ISSUE WAS RESOLVED BY REPLACING A POWER CABLE. AFTER REPLACING POWER CABLE HAD NO FURTHER ISSUES WITH LOW RESULTS. COMPLAINT DOCUMENTATION WAS REVIEWED. NO ADVERSE TRENDS WERE IDENTIFIED RELATED TO THIS ISSUE. BASED UPON THE DATA AVAILABLE AND THE RESULTS OF THIS EVALUATION NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
2012257 | THE FOLLOWING INCIDENT DESCRIPTION WAS PROVIDED TO CARIDIANBCT: THE CUSTOMER HAD A QUESTION ABOUT WHY THEY HAD TO INCREASE THE RUN TIME BECAUSE NOT ALL OF THE REPLACEMENT FLUID HAD BEEN GIVEN. THE CUSTOMER REPORTED THAT THE PT WAS FINE AND DID NOT REQUIRE MEDICAL INTERVENTION. (B)(4). THE CUSTOMER REPORTED TWO ISSUES THAT THEY COULD NOT EXPLAIN. THE FIRST ISSUE WAS THAT THEY HAD 8 BOTTLES OF 500 ML ALBUMIN FOR REPLACEMENT, BUT HAD TO INCREASE THE RUN TIME, BECAUSE NOT ALL OF THE ALBUMIN WAS GIVEN WHEN THEY REACHED 4000 ML. THE FINAL REPLACEMENT VOLUME DISPLAYED WAS 5061 ML. THE SECOND ISSUE WAS THAT THE PRIME SALINE BAG WAS EMPTY WHEN THEY WANTED TO PERFORM RINSEBACK. THE CUSTOMER WANTED TO KNOW IF THE MACHINE COULD HAVE DRAWN THE SALINE FROM THE PRIME BAG, WHICH WOULD THEN MAKE THE REPLACEMENT VOLUME HIGHER THAN INTENDED. THE CUSTOMER DID NOT NOTICE ANY OBSTRUCTIONS IN THE TUBING SET. PER THE RUN SHEET, THE LAST BOTTLE OF ALBUMIN TOOK THE LONGEST. THEY DID NOT WEIGH THE PLASMA WASTE BAG, BUT THE CUSTOMER SAID IT WAS VERY FULL AND SEEMED TO CORRESPOND WITH THE DISPLAYED REMOVE VOLUME. BASED ON THE VOLUMES PROVIDED ON THE RUN SHEET AND AN AVAILABLE REPLACEMENT VOLUME OF 4000 ML, THERE WAS A FLUID DEFICIT OF 808 ML FOR THIS PT. THE PT DID NOT HAVE ANY SIDE EFFECTS FROM THIS PROCEDURE. THE CUSTOMER'S BIOMED CHECKED OUT THE SPECTRA MACHINE AND DID NOT FIND ANYTHING WRONG. THE SAME MACHINE WAS USED FOR THREE MORE TREATMENTS ON THE SAME PT WITHOUT INCIDENT. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. BASED ON THE CUSTOMER'S RUN SHEET AND THE FACT THAT THE FIRST SEVEN BOTTLES EMPTIED NORMALLY, IT IS LIKELY THAT THE FLOW PATH BECAME PARTIALLY OCCLUDED AND CAUSED THE ALBUMIN TO EMPTY MORE SLOWLY. IT IS UNCONFIRMED WHETHER A BLOCKAGE WAS PRODUCED BY A MFG DEFECT OF THE COBE SPECTRA DISPOSABLE OR BY THE WAY THE DISPOSABLE WAS HANDLED AT THE CUSTOMER SITE. | 0
2012649 | SCOPE WAS FOUND IN STORAGE CABINET LEAKING FLUID ON 1 OLYMPUS SCOPE, BEING SETUP FOR A CASE. NO PRIOR PROCEDURES WERE DONE USING THE INDENTED SCOPE. OLYMPUS SCOPE WAS WASHED AND HIGH LEVEL DISINFECTED USING SYSTEM 83. SCOPE WAS THEN PLACED IN CABINET FOR STORAGE AND SCOPE WAS FOUND TO BE LEAKING FLUID. SERVICE TECH, WAS IMMEDIATELY DISPATCHED TO HOSPITAL FOR EVAL AND SERVICE. CUSTOM TECHNICIAN PERFORMED A COMPLETE SERVICE ON MACHINE, TECHNICIAN ALSO OBSERVED THAT THE FACILITY WAS IMPROPERLY ATTACHING SCOPE(S) TO THE SYSTEM 83 AND A TRAINING WAS RE-PERFORMED ONSITE. WITH VERIFICATION THAT SYSTEM WAS WORKING PROPERLY AND HAD STAFF MEMBERS SIGN OFF AFTER RETRAINING WAS COMPLETED. | 0
2013077 | IT WAS REPORTED THAT HIGHER THAN EXPECTED SPO2 READINGS ARE SOMETIMES DISPLAYED ON DELTA MONITORS WITH MASIMO PODS FOR PATIENTS WITH ABOVE-NORMAL BILIRUBIN LEVELS. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATION THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2013440 | AN ABBOTT FIELD SERVICE ENGINEER (FSE) WAS SPLASHED IN HIS MOUTH WITH WASTE FLUID FROM A CELL-SYN 3200 ANALYZER. THE CUSTOMER HAD CHANGED THE WASTE TANK THE PREVIOUS NIGHT FROM A FULL TANK TO AN EMPTY TANK. THE SETTING ON THE CELL-DYN 3200 ANALYZER WASTE TANK APPEARED TO BE INCORRECT AND THE WASTE TUBE BECAME DETACHED FROM THE WASTE TANK WHEN THE ANALYZER WAS STARTED UP DURING THE FSE'S VISIT. THE FSE IMMEDIATELY RINSED HIS MOUTH WITH WATER. THE NEXT DAY THE FSE SOUGHT MEDICAL ATTENTION. THERE WAS NO IMPACT TO THE CUSTOMER, OR PATIENT MANAGEMENT. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. LABELING WAS REVIEWED AND FOUND TO ADEQUATELY ADDRESS SET UP OF THE CELL-DYN 3200 WASTE LINE, AS WELL AS HAZARD INFORMATION. THE RISK MANAGEMENT FILE FOR THE CELL-DYN 3200 ALSO ADEQUATELY ADDRESSES CONTROLS IN PLACE TO REDUCE RISKS TO OPERATORS. TRACKING AND TRENDING IDENTIFIED NO SIMILAR COMPLAINTS FOR THE TIME PERIOD OF (B)(4) 2010 THROUGH (B)(4) 2011. BASED ON THE INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED FOR THE CELL-DYN 3200 RELATED TO THE REPORTED ISSUE. | 0
2015325 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REPORTING A LEAK IN THE LH 750 SLIDE STAINER ANALYZER DUE TO A FLUID THAT WAS OBSERVED OVER THE CATCH TRAY. THE LEAK WAS CONTAINED INTERNALLY WITHIN THE SLIDE STAINER AND DID NOT LEAK OUTSIDE THE INSTRUMENT. NO INJURIES OCCURRED AND MEDICAL ATTENTION WAS NOT SOUGHT. THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED FOR THIS EVENT AND CONFIRMED THAT THE LEAK WAS METHANOL. THE LEAK WAS CONSIDERED MINIMAL AND WAS CLEANED UP ACCORDING TO THE DEFINED LABORATORY PROTOCOL BY THE FSE. THE FSE REPLACED THE PERISTALTIC PUMP AND VERIFIED REPAIR PER ESTABLISHED PROCEDURES. RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. THE ROOT CAUSE FOR THE LEAK WAS ASSOCIATED WITH THE PERISTALTIC PUMP. PER PRODUCT LABELING, BCI URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
2020000 | CUSTOMER REPORTED AN RH DISCREPANCY WHEN TESTING A PATIENT SAMPLE ON THE ECHO. INITIAL TESTING RESULTED RH NEGATIVE. THE SAMPLE WAS TESTED BY MANUAL TUBE METHOD RESULTED IN 3+ REACTIONS WITH BOTH ANTI-D (MONOCLONAL BLEND) SERIES 4 AND ANTI-D (MONOCLONAL BLEND) SERIES 5. CUSTOMER RE-TESTED THE SAMPLE ON THE INSTRUMENT LATER WHICH RESULTED IN NTDS IN BOTH ANTI-D WELLS. REVIEWED THE CAMERA REPORT AND ALL CAMERA VALUES WERE WITHIN SPECIFICATION. NOTED THAT THERE WAS SLIGHT SPECKLING IN THE ANTI-D1 AND ANTI-D2 WELLS ON THE ORIGINAL BATCH AND ON REPEAT TESTING BOTH ANTI-D1/D2 WELLS APPEARED TO HAVE MORE SPECKLING AND GRADED EQUIVOCAL. A SERVICE CALL WAS MADE. ADJUSTMENTS WERE MADE TO THE INSTRUMENT AS NEEDED. UNIT WAS TESTED AND IS NOW OPERATING WITHIN SPECIFICATIONS. | 0
2020242 | IT WAS REPORTED THAT THE MONITOR WAS IN USE MONITORING A PATIENT WHEN THE MONITOR WAS SWITCHED OFF FOR A PERIOD OF TIME AND THEN SWITCHED BACK ON AGAIN. UNKNOWN TO THE USER, THE MONITOR DEFAULTED BACK TO THE AMERICAN SETTINGS FOR CHART SPEED AND THE RESULTING READINGS WERE ACTED UPON, RESULTING IN THE PATIENT HAVING AN EMERGENCY CESAREAN SECTION. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2021274 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER LH 755 ANALYZER GAVE ERRONEOUSLY HIGH RETIC RESULTS ON A SPECIFIC SPECIMEN WITHOUT INSTRUMENT GENERATED RETIC FLAG, BUT GAVE A NUCLEATED RED BLOOD COUNT (NRBC) INSTRUMENT GENERATED MESSAGE FOR THE DIFFERENTIAL. THE RESULTS WERE REPORTED OUT OF THE LABORATORY; HOWEVER, A MANUAL RETIC COUNT RECOVERED LOWER. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. CONTROLS WERE RUN BEFORE AND AFTER THE INCIDENT AND RECOVERED WITHIN ASSAY LIMITS. THE INSTRUMENT IS CURRENTLY WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS AND REPRODUCIBILITY. ACCORDING TO THE RAW DATA ANALYSIS, THE PATTERN DOES NOT DETECT ANY ABNORMAL DATA PATTERN AS IR IS NOT SEEN AS POSSESSING ANY INTERFERENCE THAT WOULD BE SUSPECTED TO AFFECT THE RETICULOCYTES PERCENTAGE. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND REPLACED THE RETIC PAK LEVEL. PER PRODUCT LABELING: BECKMAN COULTER INC. DOES NOT CLAIM TO IDENTIFY EVERY ABNORMALITY IN ALL SAMPLES. BECKMAN COULTER INC. SUGGESTS USING ALL AVAILABLE FLAGGING OPTIONS TO OPTIMIZE THE SENSITIVITY OF INSTRUMENT RESULTS BASED ON YOUR PATIENT POPULATION. ALL FLAGGING OPTIONS INCLUDE REFERENCE RANGES (H/L), ACTION AND CRITICAL LIMITS, DEFINITIVE FLAGS, SUSPECT FLAGS, PARAMETER CODES, DELTA CHECKS, DECISION RULES AND SYSTEM ALARMS. BECKMAN COULTER INC. RECOMMENDS AVOIDING THE USE OF SINGLE MESSAGES OR OUTPUTS TO SUMMARIZE SPECIMEN RESULTS OR PATIENT CONDITIONS. THE LH 700 SERIES APPLIES INSTRUMENT-GENERATED AND/OR LABORATORY-DEFINED FLAGS, CODES, AND/OR MESSAGES TO EACH SET OF PATIENT RESULTS. FLAGS, CODES, SUSPECT AND DEFINITIVE MESSAGES ARE USED TO ALERT YOU TO AN INSTRUMENT MALFUNCTION, SPECIMEN ABNORMALITY, ABNORMAL DATA PATTERN, OR ABNORMAL RESULTS. BECKMAN COULTER RECOMMENDS REVIEW, APPROPRIATE TO YOUR PATIENT POPULATION, OF ALL RESULTS DISPLAYING A FLAG, CODE OR OTHER MESSAGE. | 0
2024461 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REPORTING DILUTED BLOOD WAS FOUND UNDER THEIR COULTER LH 780 ANALYZER THAT MAY HAVE COME FROM THE NEEDLE AREA. THE CUSTOMER WAS ADVISED TO STOP USING THE INSTRUMENT UNTIL SERVICE ARRIVED. THE OPERATOR WAS WEARING PROPER PPE AND THERE WAS NO EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS AND MEDICAL ATTENTION WAS NOT SOUGHT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND IDENTIFIED THE SOURCE OF THE LEAK AS A PINHOLE IN THE TUBING CONNECTED TO THE SHEAR VALVE. THE TUBING WAS REPLACED AND REPAIR WAS VERIFIED PER ESTABLISHED PROCEDURES. THE INSTRUMENT WAS VALIDATED AND NO FURTHER LEAK WAS OBSERVED. | 0
2024502 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING WHILE ATTEMPTING TO PERFORM MAINTENANCE ON THEIR COULTER LH 750 HEMATOLOGY ANALYZER, THEY OPENED THE FRONT COVER AND NOTICED A BLOOD LEAK. THE OPERATOR DID NOT COME INTO CONTACT WITH THE LEAK. THERE WAS NO EXPOSURE TO OPEN WOUNDS OR MUCOUS MEMBRANES AND MEDICAL ATTENTION WAS NOT SOUGHT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND REPLACED THE TUBING FROM THE NEEDLE ASSEMBLY TO THE SLIDEMAKER. INSTRUMENT WAS VERIFIED AND NO FURTHER BLOOD LEAK WAS OBSERVED. | 0
2027494 | THE CUSTOMER STATED THAT THE CELL-DYN (B)(4) ANALYZER GENERATED A SHEAR VALVE POSITION ERROR WHICH WAS REPLACED IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN EXPANDED INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. A PRODUCT CORRECTION LETTER AND CUSTOMER REPLY FORM WILL BE SENT TO ALL ACTIVE (B)(4) AND (B)(6) CUSTOMERS WHO RECEIVED THE AFFECTED PART NUMBERS. THE PRODUCT CORRECTION LETTER WILL INSTRUCT CUSTOMERS TO RETURN THE CUSTOMER REPLY FORM TO ABBOTT ACKNOWLEDGING THE RECEIPT AND UNDERSTANDING THE ISSUE OR REQUEST ASSISTANCE. THE NON-CONFORMING PARTS WILL BE REPLACED IN THE FIELD THROUGH A FOUR-MONTH MANDATORY TECHNICAL SERVICE BULLETIN (TSB) ISSUED FOR EACH OF THE CELL-DYNS INVOLVED. | 0
2029052 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) REPORTING A LIQUID LEAK APPEARING TO BE DILUENT UNDERNEATH THE COULTER LH 780 ANALYZER. THE OPERATOR WAS WEARING PPE AT THE TIME OF THE EVENT AND DID NOT COME INTO CONTACT WITH THE LIQUID. THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS AND MEDICAL ATTENTION WAS NOT SOUGHT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE (B)(4) 2011, REPLACED THE LEAKING TUBING PINCH VALVE AND VERIFIED REPAIR PER ESTABLISHED PROCEDURES. RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. THE ROOT CAUSE FOR THE LEAK WAS ASSOCIATED WITH THE TUBING PINCH VALVE. | 0
2030094 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER LH 750 ANALYZER MIS-READ THREE (3) BAR CODE LABELS, WITH EACH GENERATING A NO MATCH MESSAGE. UPON RERUN THE LABEL WAS READ CORRECTLY. NO ERRONEOUS RESULTS WERE REPORTED OUTSIDE THE LABORATORY. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS EVENT. THE CUSTOMER DOES NOT USE CHECK DIGITS TO VERIFY BAR CODES. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(4) 2011 AND CLEANED AND TESTED THE BARCODE READER. THE FSE RERAN THE SAMPLES AND SOME LABELS READ CORRECTLY AND SOME INCORRECTLY OR WERE NOT READ BY THE SYSTEM (NO READ MESSAGE). PER THE FSE, THE LABELS HAD GAPS IN THE BARS WHEN VIEWED UNDER A MAGNIFYING GLASS. THE FSE SUSPECTS THIS ISSUE MAY BE DUE TO THE BAR CODE LABELS THE CUSTOMER USES. THE FSE INFORMED THE TECHNICIANS TO VERIFY EVERY SAMPLE ID ON THE TUBE WITH THE RESULT PRINTOUT SAMPLE ID AND NAMES. PER PRODUCT LABELING, RISK OF MISIDENTIFICATION. USE OF POOR QUALITY, DIRTY, IMPROPERLY PLACED OR DAMAGED BAR-CODE LABELS COULD KEEP THE INSTRUMENT FROM READING THE BAR-CODE LABELS. ENSURE THE BAR-CODE LABELS ARE UNDAMAGED. ENSURE THE BAR-CODE LABELS CONFORM TO THE SPECIFICATIONS PROVIDED IN CHAPTER 4 OF THE REFERENCE MANUAL. BECKMAN COULTER STRONGLY RECOMMENDS THE USE OF BAR-CODE CHECKSUMS TO PROVIDE AUTOMATIC CHECKS FOR READ ACCURACY. LABEL SPECIFICATIONS STATE: NO SPOTS OR VOIDS; NO INK SMEARING. | 1
2033378 | ABBOTT FIELD SERVICE INSPECTED THE CUSTOMER'S ARCHITECT C4000 ANALYZER PER MANDATORY TECHNICAL SERVICE BULLETIN 128-014 AND REPLACED THE ARCHITECT ICT PRE-AMP BOARD AS REQUIRED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). CONCLUSION: NEW WASHING PROCESS IMPACTED THE LONG TERM RELIABILITY OF THE PHOTO COUPLERS WHICH DEGRADE SLOWLY OVER TIME. THE CAUSE OF THE ARCHITECT ANALYZER PHOTO COUPLER DEGRADATION ISSUE WAS DUE TO A NEW MANUFACTURING WASHING PROCESS IMPLEMENTED BY THE SUPPLIER. ABBOTT ISSUED A PRODUCT CORRECTION LETTER TO ALL CUSTOMERS WITH AFFECTED INSTRUMENT SERIAL NUMBERS AND IMPLEMENTED A MANDATORY TECHNICAL SERVICE BULLETIN TO INSPECT ICT PRE AMP BOARDS AND REPLACE IF NEEDED. | 0
2033422 | ABBOTT FIELD SERVICE INSPECTED THE CUSTOMER'S ARCHITECT C4000 ANALYZER PER MANDATORY TECHNICAL SERVICE BULLETIN (B)(4) AND REPLACED THE ARCHITECT ICT PRE-AMP BOARD AS REQUIRED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4) NEW (B)(4) PROCESS IMPACTED THE LONG TERM RELIABILITY OF THE PHOTO COUPLERS WHICH DEGRADE SLOWLY OVER TIME. THE CAUSE OF THE ARCHITECT ANALYZER PHOTO COUPLER DEGRADATION ISSUE WAS DUE TO A NEW MANUFACTURING (B)(4) PROCESS IMPLEMENTED BY THE SUPPLIER. ABBOTT ISSUED A PRODUCT CORRECTION LETTER TO ALL CUSTOMERS WITH AFFECTED INSTRUMENT SERIAL NUMBERS AND IMPLEMENTED A MANDATORY TECHNICAL SERVICE BULLETIN TO INSPECT ICT PRE AMP BOARDS AND REPLACE IF NEEDED. | 0
2034528 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT A BLOODY LEAK WAS NOTED UNDERNEATH THE BLOOD SAMPLING VALVE, SOURCE UNKNOWN, OF THE COULTER HMX AUTOLOADER ANALYZER. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE), INCLUDING LAB COAT, GLOVES AND EYE PROTECTION. THE OPERATOR DID NOT SEEK MEDICAL TREATMENT AND THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. PATIENT TREATMENT WAS NOT AFFECTED IN ASSOCIATION WITH THIS EVENT. A BCI FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(4) 2011 AND ORDERED THE NEEDLE CARTRIDGE FOR THE CUSTOMER TO INSTALL. THE FOLLOWING DAY THE PROMARY MODE SAMPLE NEEDLE WAS REMOVED AND REPLACED. THE CONTROLS AND SYSTEM ARE NOW OPERATING WITHIN SPECIFICATIONS. AS PER PRODUCT LABELING, BCI URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
2035141 | IT WAS REPORTED THAT AFTER A HARD RESET WAS PERFORMED WITH THE HANDHELD, THE SCREEN WOULD NOT ALIGN. A SECOND HARD RESET AND TROUBLESHOOTING WAS PERFORMED, BUT THE PROBLEM PERSISTED. PRODUCT HAS BEEN REQUESTED, BUT HAS NOT BEEN RETURNED TO THE MANUFACTURER TO DATE. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2036299 | THE CUSTOMER STATED RBC, MCV AND PLATELET QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER WAS ADVISED TO USE A BLEACH DILUTION TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 8/10/10, EXPIRES 6/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. (B)(4) | 0
2036348 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 08/10/10, EXPIRES 06/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. THE ROOT CAUSE IS THE VENDOR (CDS, THE OEM SUPPLIER). | 0
2036390 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER: MANUFACTURED 08/10/10, EXPIRES 06/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. THE ROOT CAUSE IS THE VENDOR (CDS, THE OEM SUPPLIER). | 0
2036462 | IT WAS REPORTED THAT THE UP/DOWN ARROW KEYPAD BUTTONS HAD TO BE PRESSED MULTIPLE TIMES IN ORDER TO ACTIVATE THE DESIRED PUMP FUNCTIONS. THE KEYPAD IS REPORTEDLY INTACT AND THE PUMP HAS NOT BEEN EXPOSED TO MOISTURE. THERE WAS NO ADVERSE EVENT ASSOCIATED WITH THIS COMPLAINT. THE PUMP HAS BEEN RETURNED TO ANIMAS; HOWEVER, THE INVESTIGATION HAS NOT BEEN COMPLETED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. ONCE THE EVALUATION HAS BEEN COMPLETED, A SUPPLEMENTAL REPORT WILL BE FILED. | 0
2036469 | THE CUSTOMER STATED (B)(4) QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED (B)(6) 2010, EXPIRES 6/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. THE ROOT CAUSE IS THE VENDOR (B)(4). | 0
2036495 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. THE ROOT CAUSE IS THE VENDOR (CDS, THE OEM SUPPLIER).><(B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 08/10/10, EXPIRES 06/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
2036535 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER REPLACED THE CLEANER WITH THAT OF A DIFFERENT LOT AND DECONTAMINATED TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 8/10/10, EXPIRES 6/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. (B)(4) | 0
2042128 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO ERRONEOUSLY HIGH WBC RESULT WITH NO INSTRUMENT GENERATED FLAGS WHEN COMPARED TO THE REPEAT RUN ON THE COUTLER LH 755 INSTRUMENT. THE REPEAT RUN WAS CONSISTENT WITH THE LOWER WBC OBTAINED BY MANUAL COUNT WHICH THE LAB CONSIDERS CORRECT. NO ERRONEOUS RESULTS WERE REPORTED OUT OF THE LABORATORY. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. SAMPLE COLLECTION AND STORAGE INFORMATION WERE NOT PROVIDED FOR THIS INVESTIGATION. CONTROLS WERE RUN BEFORE AND AFTER THE EVENT. THE UNIT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS. RAW DATA ANALYSIS REVEALED THAT THE RAW COUNT RATE FOR RUN 1 IS SIGNIFICANTLY HIGHER THAN RUN 2. THE DIFFERENCE IN THE RAW COUNTS IS DIRECTLY RELATED TO THE DIFFERENCES IN THE WBC RESULTS. MOREOVER, THE RAW COUNT RATE TRENDS DOWN WITH COLLECTION TIME FOR RUN 1 WHILE IT REMAINED FLAT FOR RUN 2. A BCI FIELD SERVICE ENGINEER (FSE) REPLACED THE FOLLOWING COMPONENTS: VL3NO TUBE, SOL21, CHECK VALVE, AND VL10. FSE STATED UP VERIFICATION PASSES. RAN PATIENT SPECIMENS FOR SYSTEM VERIFICATION PURPOSES. ROOT CAUSE IS MOST LIKELY ATTRIBUTED TO THE PARTS THAT WERE REPLACED FOR THE SYSTEM DILUTER DURING THE INSTRUMENT SERVICING PERFORMED SUBSEQUENT TO THIS CF EVENT. | 0
2042146 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) IN REGARDS TO ERRONEOUSLY HIGH WBC RESULT WITH NO INSTRUMENT GENERATED FLAGS WHEN COMPARED TO THE REPEAT RUN ON THE COUTLER LH 755 INSTRUMENT. THE REPEAT RUN WAS CONSISTENT WITH THE LOWER WBC OBTAINED BY MANUAL COUNT WHICH THE LAB CONSIDERS CORRECT. NO ERRONEOUS RESULTS WERE REPORTED OUT OF THE LABORATORY. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. SAMPLE COLLECTION AND STORAGE INFORMATION WERE NOT PROVIDED FOR THIS INVESTIGATION. CONTROLS WERE RUN BEFORE AND AFTER THE EVENT. THE UNIT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS. RAW DATA ANALYSIS REVEALED THAT THE RAW COUNT RATE FOR RUN 1 IS SIGNIFICANTLY HIGHER THAN RUN 2. THE DIFFERENCE IN THE RAW COUNTS IS DIRECTLY RELATED TO THE DIFFERENCES IN THE WBC RESULTS. MOREOVER, THE RAW COUNT RATE TRENDS DOWN WITH COLLECTION TIME FOR RUN 1 WHILE IT REMAINED FLAT FOR RUN 2. A BCI FIELD SERVICE ENGINEER (FSE) REPLACED THE FOLLOWING COMPONENTS: VL3NO TUBE, SOL21, CHECK VALVE, AND VL10. FSE STATED UP VERIFICATION PASSES. RAN PATIENT SPECIMENS FOR SYSTEM VERIFICATION PURPOSES. ROOT CAUSE IS MOST LIKELY ATTRIBUTED TO THE PARTS THAT WERE REPLACED FOR THE SYSTEM DILUTER DURING THE INSTRUMENT SERVICING PERFORMED SUBSEQUENT TO THIS CF EVENT. | 0
2042657 | THE CUSTOMER STATED THAT FALSELY ELEVATED ARCHITECT MAGNESIUM RESULTS WERE GENERATED FOR PATIENT SAMPLES. AN INITIAL ELEVATED RESULT OF 2.63 MMOL/L RETESTED AT 0.74 MMOL/L. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). REVIEW OF THE ARCHITECT C8000 SYSTEM LOGS DID NOT IDENTIFY ANY ISSUES. CONTROLS WERE WITHIN RANGE FOLLOWING TROUBLESHOOTING, BUT PATIENT RESULTS WERE STILL ERRATIC. FIELD SERVICE DISCOVERED THAT THE AMOUNT OF WATER USED FOR CLEANING THE REAGENT PROBE WAS TOO SMALL AND ADJUSTED THE FLOW. PROPER OPERATION OF THE SYSTEM WAS VERIFIED AFTER ADJUSTING THE WATER FLOW TO THE REAGENT PROBE. QUALITY METRICS AND COMPLAINT TRACKING AND TRENDING WAS REVIEWED. NO ADVERSE TRENDS WERE IDENTIFIED RELATED TO DISCREPANT RESULTS. THE CURRENT COMBINED ERRATIC RESULT RATE FOR AEROSET AND ARCHITECT INSTRUMENTS FALLS BELOW THE ESTABLISHED INTERNAL ABERRANT RESULT RATE FOR ARCHITECT C8000 ESTABLISHED AT PRODUCT LAUNCH. BASED ON THE AVAILABLE INFORMATION, A DEFICIENCY OF THE ARCHITECT C8000 SYSTEM WAS NOT IDENTIFIED.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
2043979 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THE CUSTOMER BLEACHED THE INSTRUMENT EVERY 2-3 DAYS TO RESOLVE THE ISSUE. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. THE ROOT CAUSE IS THE VENDOR (B)(4).><(B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 8/10/10, EXPIRES 6/18/12. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS. | 0
2047557 | THE CUSTOMER'S CELL-DYN WASTE SENSOR FAILED TO DETECT A WASTE FULL. THE WASTE OUTLET TUBE ASSEMBLY WAS REPLACED IN ORDER TO RESOLVE THE ISSUE. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, FA30NOV2009, WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. | 0
2052228 | ABBOTT FIELD SERVICE INSPECTED THE CUSTOMER'S ARCHITECT C4000 ANALYZER PER MANDATORY TECHNICAL SERVICE BULLETIN (B)(4) AND REPLACED THE ARCHITECT ICT PRE-AMP BOARD AS REQUIRED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4) CONCOMITANT MEDICAL PRODUCTS: ARCHITECT ICT PRE-AMP BOARD PHOTO COUPLER, LOTS 0834 AND 0840. (B)(4) NEW (B)(4) PROCESS IMPACTED THE LONG TERM RELIABILITY OF THE PHOTO COUPLERS WHICH DEGRADE SLOWLY OVER TIME. THE CAUSE OF THE ARCHITECT ICT PRE-AMP BOARD PHOTO COUPLER DEGRADATION ISSUE WAS DUE TO A NEW MANUFACTURING (B)(4) PROCESS IMPLEMENTED BY THE SUPPLIER. ABBOTT ISSUED A PRODUCT CORRECTION LETTER TO ALL CUSTOMERS WITH AFFECTED INSTRUMENT SERIAL NUMBERS AND IMPLEMENTED A MANDATORY TECHNICAL SERVICE BULLETIN TO INSPECT THE ARCHITECT ICT PRE AMP BOARDS AND REPLACE IF NEEDED. | 0
2059702 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WITH THE ANTIBODY SCREENING ASSAY ON THE GALILEO ECHO INSTRUMENT. THE RESULT IMAGE FILES WERE REVIEWED BY AN IMMUCOR TECHNICAL SPECIALIST. THE RESULTS VISUALLY APPEARED POSITIVE WHILE THE INSTRUMENTS INTERPRETATION OF THE RESULTS WAS NEGATIVE. CUSTOMER WAS INFORMED OF THE NEED TO VISUALLY VERIFY ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS. AN IMMUCOR SERVICE ENGINEER ALSO EVALUATED THE INSTRUMENT. AS A PRECAUTIONARY MEASURE, THE CAMERA MODULE WAS REPLACED; PERFORMED SERVICE ADJUSTMENTS AND CHECKS; PERFORMED QC; AND OBSERVED SEVERAL TESTS WITHOUT ERROR. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2060786 | IT WAS REPORTED THAT THE EVALUATION SYSTEM FAILED WHEN IT WAS POWERED UP AT A CUSTOMER SITE AND SMOKE WAS NOTED. THE SYSTEM'S BACK PANEL WAS REMOVED FOR VISUAL INSPECTION AND A BURNT SPOT WAS NOTICED ON THE SYSTEM'S PASSIVE BACKPLANE BOARD. ALSO, THE POWER SUPPLY HARNESS EXHIBITED DAMAGE. THE SYSTEM WAS REMOVED FROM SERVICE. NO INJURIES WERE REPORTED. (B)(4). THE COMPLAINT DEVICE HAS NOT BEEN REC'D BY THE MANUFACTURER SO A DEVICE EVALUATION HAS NOT BEEN PERFORMED YET. THE REPORTED COMPLAINT OF BURNT SPOT AND DAMAGED POWER SUPPLY HARNESS IS CONFIRMED BASED ON PICTORIALS PROVIDED FROM THE REPORTER. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETED. | 1
2061107 | ABBOTT FIELD SERVICE INSPECTED THE ARCHITECT C4000 ANALYZER PER MANDATORY TECHNICAL SERVICE BULLETIN (B)(4) AND REPLACED THE ARCHITECT ICT PRE-AMP BOARD AS REQUIRED. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4) CONCOMITANT MEDICAL PRODUCTS: ARCHITECT ICT PRE-AMP BOARD PHOTO COUPLER, LOTS 0834 AND 0840. (B)(4) NEW (B)(4) PROCESS IMPACTED THE LONG TERM RELIABILITY OF THE PHOTO COUPLERS WHICH DEGRADE SLOWLY OVER TIME. THE CAUSE OF THE ARCHITECT ICT PRE-AMP BOARD PHOTO COUPLER DEGRADATION ISSUE WAS DUE TO A NEW MANUFACTURING (B)(4) PROCESS IMPLEMENTED BY THE SUPPLIER. ABBOTT ISSUED A PRODUCT CORRECTION LETTER TO ALL CUSTOMERS WITH AFFECTED INSTRUMENT SERIAL NUMBERS AND IMPLEMENTED A MANDATORY TECHNICAL SERVICE BULLETIN TO INSPECT THE ARCHITECT ICT PRE AMP BOARDS AND REPLACE IF NEEDED. | 0
2063247 | A STAFF MEMBER WAS USING A SEALSAFE TO SEAL A LINE THAT CONTAINED PLASMA. IT DID NOT SEAL CORRECTLY AND PLASMA SPRAYED ONTO THE FACE OF THIS STAFF MEMBER. THE STAFF MEMBER IS STABLE. THE SITE WILL FOLLOW UP IN 6 MONTHS WITH THE EMPLOYEE HEALTH CLINIC FOR POTENTIAL BLOOD BORNE PATHOGEN EXPOSURE. WHEN REQUESTED THE SITE WOULD NOT PROVIDE THE STAFF MEMBERS NAME OR WEIGHT. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE AND PREVENTATIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2063300 | THE CUSTOMER STATES THAT THE CELL-DYN SAPPHIRE ANALYZER GENERATED AN INITIAL LOW PATIENT HEMOGLOBIN RESULT OF 7.59 G/DL THAT RETESTED AT 12.40 G/DL. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><AN ABBOTT FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE AND REPLACED THE ANALYZER'S HEMOGLOBIN REAGENT SYRINGE. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE CELL-DYN SAPPHIRE ANALYZER OPERATION MANUAL PROVIDES INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. A PRODUCT DEFICIENCY WAS NOT IDENTIFIED. REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION. | 0
2065353 | AN ABBOTT FIELD SERVICE ENGINEER (FSE) REPORTS THAT A NEW DISPENSE VERIFY TOOL, USED WITH THE PRISM ANALYZER, STARTED SMOKING IMMEDIATELY WHEN THE POWER CORD WAS PLUGGED INTO AN ELECTRICAL OUTLET. THE TOOL HAD NOT BEEN CALIBRATED AND NO LIQUID HAD BEEN DISPENSED INTO THE CUPS. THE FSE IMMEDIATELY UNPLUGGED THE TOOL FROM THE OUTLET. THERE ARE NO REPORTS OF INJURY OR DAMAGE. (B)(4). THIS INVESTIGATION INDICATES THAT THE ABBOTT PRISM ANALYZER PERFORMS AS INTENDED WITH RESPECT TO THIS COMPLAINT ISSUE. NO PRODUCT DEFICIENCY WAS IDENTIFIED. NO ADDITIONAL INFORMATION OR ACTIONS HAVE BEEN IDENTIFIED TO ASSIST THIS COMPLAINANT OR THE ENTIRE CUSTOMER BASE OTHER THAN TO CONTACT THEIR AREA REPRESENTATIVE(S). THIS INVESTIGATION DOCUMENTS THE ISSUE OF SMOKE WITH THE ABBOTT PRISM DISPENSE VOLUME TOOL ONLY. THE RETURNED DISPENSE VOLUME TOOL (L/N 1G40-26, S/N (B)(4); DEVICE MANUFACTURE DATE 03/11/2010) WAS VISUALLY INSPECTED AND THE FOLLOWING WAS OBSERVED: ONE CAPACITOR (LABELED C9) ON THE PRINTED CIRCUIT BOARD INSIDE THE TOOL HAD FAILED; THIS WAS THE MOST LIKELY CAUSE OF THE SMOKE THE FSR OBSERVED. AN EVALUATION OF THE CAPACITOR DETERMINED IT WAS THE CORRECT TYPE AND IN THE CORRECT ORIENTATION. FLAMMABILITY: ENCAPSULATED MATERIALS MEET THE UNDERWRITER LABORATORIES FLAME RATING UL94 V0. THIS RATING STATES THAT THE CAPACITOR WILL SELF-EXTINGUISH AFTER THE IGNITION SOURCE IS REMOVED. COMPLAINT ACTIVITY WAS REVIEWED AND IDENTIFIED NO ADVERSE TRENDS IN ASSOCIATION WITH THE ISSUE CURRENTLY UNDER EVALUATION. THE ABBOTT PRISM OPERATIONS MANUAL (L/N 1G40-47) CONTAINS INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. THIS IS AN ISOLATED EVENT INVOLVING A PRISM DISPENSE VOLUME TOOL. NO PRODUCT DEFICIENCY WAS IDENTIFIED. METHOD: REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION. | 0
2066005 | ON (B)(6) 2011, WE RECEIVED AN E-MAIL FROM A PHYSICIAN AT CHILDREN'S HOSP OF (B)(6) THAT SOME REPORTS ARE NOT SHOWING AS BEING MARKED AS READ AFTER BEING VERIFIED, AND SOME OF THE REPORTS ARE ALSO STILL SHOWING THE "PRELIMINARY" WATERMARK EVEN THOUGH THEY SHOW AS BEING MARKED AS READ. THERE ARE NO REPORTS OF AN ADVERSE EVENT. UPON INVESTIGATION, WE FOUND THAT IF A PHYSICIAN TRIES TO SIGN THE REPORT AFTER DELETING A GRAPH THAT RESIDES IN A TABLE, THE REPORT FAILS TO SAVE. THE AFFECTED REPORT ALSO REMAINS AS PRELIMINARY INSTEAD OF VERIFIED AS EXPECTED. CHANGES MADE TO THE REPORT SINCE LAST SAVE ARE LOST. THIS BEHAVIOR DOES NOT OCCUR IF GRAPHS ARE NOT IN A TABLE. THE GRAPHS WERE PLACED IN A TABLE FOR AESTHETICS. THIS ISSUE HAS THE POTENTIAL OF CREATING A SITUATION WHERE REPORT INFO FOR TREATMENT PLANNING IS NOT AVAILABLE IF THE PHYSICIAN DOES NOT REALIZE THE SAVE FAILED. A CUSTOMER SAFETY ADVISORY NOTICE WAS CREATED AND WILL BE DISTRIBUTED TO ALL AFFECTED CUSTOMERS. A SOFTWARE PATCH WILL BE DEVELOPED AND MADE AVAILABLE. ADD'L CATALOG #: 10091604. | 1
2068341 | DURING INVESTIGATION OF THE EVENT ASSOCIATED WITH MDR 2183553-2011-00005, IT WAS DISCOVERED INTERNALLY THAT THE HFO5 PRODUCT IS ALSO AFFECTED BY THE REPORTED ISSUE. THERE WAS NO PT INVOLVEMENT FOR THE INTERNALLY DISCOVERED ISSUE. SCREEN SAVED IMAGES FROM THE MAGNETIC RESONANCE (MR) SYSTEM WILL ATTACH TO THE INCORRECT PT ON THE PICTURE ARCHIVE COMMUNICATION SYSTEM (PACS). WHEN THE USERS SEND A SCREEN SAVE IMAGE FROM THE MR SYSTEM TO THE PACS, THE SCREEN SAVE IS THEN ATTACHED TO THE CORRECT EXAM ON THE MR SYSTEM; HOWEVER, ON THE PACS, THE SCREEN SAVE MAY ATTACH TO A DIFFERENT PT/EXAM. INITIAL EVALUATION INDICATES THAT THE ISSUE IS CAUSED WHEN THE MR SYSTEM REPLICATES THE SERIES INSTANCE UID (0X0020, 0X000E) BETWEEN THE TWO SCREEN SAVE IMAGES. HOWEVER, THE TWO SCREEN SAVE IMAGES BOTH ARE USING THE SAME SERIES INSTANCE UID (0020, 000E). GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS STILL ONGOING. A FOLLOW-UP REPORT WILL BE ISSUED WHEN THE INVESTIGATION HAS BEEN COMPLETED. THIS ISSUE WAS DISCOVERED DURING INTERNAL INVESTIGATION OF THE PREVIOUS EVENT AND DOES NOT APPLY TO A SINGLE DEVICE. THEREFORE, THE SERIAL NUMBER, OPERATOR OF THE DEVICE AND DEVICE MANUFACTURE DATE ARE NOT PROVIDED. | 1
2070674 | THE CUSTOMER STATED AN ARCHITECT I1000SR ANALYZER GENERATED FALSELY ELEVATED (B)(6) RESULTS FOR ONE PATIENT SAMPLE. DATA WAS REVIEWED AND FOUND NOT TO AFFECT THE CLINICAL INTERPRETATION OR MEDICAL DECISION MAKING. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A PRODUCT DEFICIENCY HAS BEEN IDENTIFIED. WITH A UNIQUE COMBINATION OF INSTRUMENT AND REAGENT CONDITIONS THERE IS A POTENTIAL FOR SAMPLE CARRYOVER WHICH CAN RESULT IN FALSELY ELEVATED (B)(6) CONCENTRATIONS ON THE ARCHITECT I1000SR SYSTEM. NEGATIVE SAMPLES HAVE RETURNED (B)(6) AND GREY ZONE (B)(6) RESULTS WHEN A HIGH CONCENTRATION (B)(6) SAMPLE IS ASSESSED PRIOR TO THE NEGATIVE SAMPLE WITH ARCHITECT (B)(6) (LIST NUMBER 7K78/6C21) ON AN I1000SR INSTRUMENT WHICH ALSO UTILIZES THE ARCHITECT (B)(6) (LIST NUMBER 6C17) ASSAY. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. AN INVESTIGATION IS IN PROCESS.><THE SUSPECT MEDICAL DEVICE HAS CHANGED FROM THE (B)(4). MFR # 3005094123-2011-00547 HAS BEEN SUBMITTED TO ADDRESS THIS ERROR. | 0
2070770 | THE REPORTER REPORTED THAT THE PT NOTED THE INSULIN PUMP CHANGED THE DATE/TIME FOUR TIMES OVER THE PAST TWO MONTHS. ON AN UNSPECIFIED DATE, THE PT OBTAINED THE BLOOD GLUCOSE READING OF 11.0 MMOL/L (198 MG/DL) AND HE EXPERIENCED THE SYMPTOMS OF NAUSEA AND THIRST; HE WAS NEGATIVE FOR URINARY KETONES. THE PT CORRECTED HIS BLOOD GLUCOSE LEVEL USING A BOLUS INSULIN DOSE VIA THE PUMP; HE DID NOT SEEK ANY MEDICAL ATTENTION. TROUBLESHOOTING REVEALED THE PUMP PROGRAMMING HAD NOT BEEN CHANGED, AND DUE TO THE INCORRECT TIME, THE PT'S BASAL RATE WAS INCORRECT. THERE WAS NO ADVERSE EVENT ASSOCIATED WITH THIS ISSUE. THE PUMP HAS NOT BEEN RETURNED TO ANIMAS FOR EVALUATION. IF THE DEVICE IS RETURNED, AN EVALUATION SHALL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE FILED. NO CONCLUSIONS CAN BE DRAWN AT THIS TIME. | 0
2071912 | THE CUSTOMER STATED AN ARCHITECT I1000SR ANALYZER GENERATED FALSELY ELEVATED B-HCG RESULTS FOR SIX PATIENT SAMPLES. DATA WAS REVIEWED AND FOUND NOT TO AFFECT THE CLINICAL INTERPRETATION OR MEDICAL DECISION MAKING. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). A PRODUCT DEFICIENCY HAS BEEN IDENTIFIED. WITH A UNIQUE COMBINATION OF INSTRUMENT AND REAGENT CONDITIONS THERE IS A POTENTIAL FOR SAMPLE CARRYOVER WHICH CAN RESULT IN FALSELY ELEVATED B-HCG CONCENTRATIONS ON THE ARCHITECT I1000SR SYSTEM. NEGATIVE SAMPLES HAVE RETURNED BOTH POSITIVE AND GREY ZONE B-HCG RESULTS WHEN A HIGH CONCENTRATION B-HCG SAMPLE IS ASSESSED PRIOR TO THE NEGATIVE SAMPLE WITH ARCHITECT TOTAL B-HCG (LIST NUMBER 7K78/6C21) ON AN I1000SR INSTRUMENT WHICH ALSO UTILIZES THE ARCHITECT RUBELLA IGG (LIST NUMBER 6C17) ASSAY. FURTHER INVESTIGATION OF THIS QUALITY ISSUE WILL BE CONDUCTED. AN INVESTIGATION IS IN PROCESS. | 0
2072359 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER LH 750 ANALYZER GAVE ERRONEOUS DIFFERENTIAL RESULTS ON A SPECIMEN WITHOUT INSTRUMENT GENERATED MESSAGES FROM ONE PATIENT. MANUAL REVIEW SHOWED ONE (1) BLAST. ERRONEOUS RESULTS WERE NOT REPORTED OUT OF THE LABORATORY. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS INCIDENT. THERE WAS NO CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS INCIDENT. SPECIMEN WAS COLLECTED IN A 5 CC VACUTAINER TUBE AND SAMPLED <2 HOURS AFTER COLLECTION. CONTROLS WERE RUN BEFORE AND AFTER THE INCIDENT AND RECOVERED WITHIN ASSAY RANGES. THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS (ACCURACY AND PRECISION). SERVICE WAS NOT DISPATCHED FOR THIS EVENT. THE ROOT CAUSE, BASED ON RAW DATA ANALYSIS, FOR THE ERRONEOUS DIFFERENTIAL RESULTS AND THE FALSE NEGATIVE BLAST FLAGGING IS THAT THE ALGORITHM COULD NOT SET SUSPECT FLAGS BECAUSE THERE ARE NO SIGNIFICANTLY ABNORMAL PATTERNS ON THE HISTOGRAMS. PER PRODUCT LABELING, BECKMAN COULTER INC. DOES NOT CLAIM TO IDENTIFY EVERY ABNORMALITY IN ALL SAMPLES. | 0
2074997 | THE ACCOUNT STATED THE ARCHITECT C16000 ANALYZER WAS GENERATING INCONSISTENT SODIUM PATIENT RESULTS. ONE EXAMPLE PROVIDED SHOWED AN ARCHITECT C16000 SODIUM = 156.2 MMOL/L ON A SPECIMEN THAT REPEATED AT 107.1, 106.7, 111.4, 106.8, 106.7, 108.7, 107, 107, 106.9, 108, 104.9 MMOL/L. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED. THE INCORRECT ARCHITECT C16000 SODIUM RESULT WAS NOT REPORTED OUTSIDE OF THE LABORATORY. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP MDR WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><(B)(4). TO INVESTIGATE THE CUSTOMER ISSUE, THE INVESTIGATION TEAM REVIEWED THE COMPLAINT TEXT, THE RETURNED ATTACHMENTS, THE INSTRUMENT HISTORY, AND ARCHITECT SYSTEM LABELING. THE FIELD SERVICE REPRESENTATIVE RESOLVED THE CUSTOMER ISSUE BY REPLACING THE 1 ML ASPIRATION SYRINGES. THE SERVICE HISTORY REVEALED THAT ANOTHER TICKET FROM THE CUSTOMER WAS OPENED ON (B)(6) 2011 FOR DEPRESSED CONCENTRATED CALCIUM PATIENT AND QUALITY CONTROL RESULTS. THE ISSUE WAS RESOLVED BY REPLACING ALL 1 ML ASPIRATION SYRINGES ON (B)(4) 2011 AND THE TICKET WAS CLOSED ON (B)(4) 2011. THE COMPLAINT TEXT FOR THIS OTHER TICKET STATES THAT THE 1 ML ASPIRATION SYRINGES AND CHECK VALVES WERE REPLACED DUE TO DAMAGE ON (B)(4) 2011. ADDITIONALLY, THE DAMAGE TO THE 1 ML SYRINGES COULD HAVE OCCURRED WHEN THE SYRINGES WERE REPLACED 5 DAYS PRIOR. THE SYSTEM LOGS WERE REVIEWED A WEEK LATER AND NO FURTHER ISSUE WITH SODIUM FLIERS AND/OR ICT DISCREPANT RESULTS WERE OBSERVED. THE ARCHITECT OPERATIONS MANUAL SECTION 10 FOR OBSERVED PROBLEMS ERRATIC RESULTS, POOR PRECISION- ICT RESULTS PROVIDES CORRECTIVE ACTIONS THAT INCLUDES BUT NOT LIMITED TO: 1 ML SYRINGES IN THE ICT ASPIRATION PUMP OR IN THE ICT REFERENCE SOLUTION PUMP ARE NOT SEATED CORRECTLY, 1 ML SYRINGES IN THE ICT ASPIRATION PUMP OR IN THE ICT REFERENCE SOLUTION PUMP ARE LEAKING, AND ICT CHECK VALVES ARE NOT CONNECTED CORRECTLY OR NOT FUNCTIONING. NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
2075823 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COUTLER LH 750 ANALYZER RECOVERED HIGH WHITE BLOOD COUNT (WBC) RESULTS WHEN COMPARED TO AN ALTERNATE INSTRUMENT, WITH NO INSTRUMENT GENERATED FLAGS. PER THE CUSTOMER, THE HIGH WBC VALUES OCCURRED ON SOME SAMPLES. DATA WAS PROVIDED FOR SIX (6) PATIENT SAMPLES TO DEMONSTRATE THIS ISSUE. ADDITIONALLY, AN ERRONEOUS HIGH HEMOGLOBIN (HGB) RESULT WAS OBSERVED ON ONE (1) PATIENT SAMPLE. THE RESULTS PROVIDED BY THE CUSTOMER ARE SHOWN IN THE ATTACHMENT SECTION OF THIS REPORT. NO ERRONEOUS RESULTS WERE REPORTED OUT OF THE LABORATORY. THE RESULTS OBTAINED FROM THE ALTERNATE INSTRUMENT ARE CONSIDERED CORRECT AND WERE REPORTED OUT OF THE LABORATORY. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. SAMPLE COLLECTION AND STORAGE INFORMATION WERE NOT PROVIDED FOR THIS INVESTIGATION. CONTROLS ARE RUN DAILY AND WERE RUN PRIOR TO THE INCIDENT WITHIN ASSAY RANGE. HOWEVER, THE WBC PARAMETER WAS HIGH FOR THE ABNORMAL I LEVEL CONTROL AFTER THE INCIDENT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND ADJUSTED THE WBC BATH DRAIN TUBING ON THE WBC BATH. THE FSE ALSO RAN TWENTY FOUR (24) BLOOD SAMPLES ON THE LH750, ALL OF WHICH MATCHED RESULTS FROM THE ALTERNATE INSTRUMENT TO VERIFY SYSTEM OPERATION. THE ROOT CAUSE IS RELATED TO THE ADJUSTMENTS MADE TO THE WBC BATH DURING THE INSTRUMENT SERVICING SUBSEQUENT TO THIS EVENT. | 0
2079626 | GLOBUS MEDICAL, INC WAS CONTACTED VIA TELEPHONE COMMUNICATION FROM A SALES REPRESENTATIVE THAT ON (B)(6) 2010 A MALE SCRUB TECHNICIAN AT THE HOSPITAL FACILITY SUSTAINED AN INJURY TO HIS FOOT WHEN A HANDLE BROKE OFF THE REVOLVE GRAPHIC CASE HE WAS CARRYING, AND THE CASE CONTAINING REVOLVE SURGICAL INSTRUMENTS LANDED ON HIS FOOT AFTER WHICH HE EXPERIENCED PAIN AND SWELLING. THE SCRUB TECHNICIAN WENT TO THE HOSPITAL EMERGENCY ROOM FOR TREATMENT AND WAS DIAGNOSED WITH A CONTUSION AND SPRAIN ON THE TOP OF HIS RIGHT FOOT. THE SCRUB TECHNICIAN WAS OUT OF WORK ON WORKER'S COMPENSATION, AND HAS SINCE THEN RETURNED TO WORK. A COMPREHENSIVE INVESTIGATION WAS IMMEDIATELY INITIATED ON RECEIPT OF THE COMPLAINT. THE PRODUCT WAS RETURNED FOR EVALUATION AND A REVIEW WAS CONDUCTED OF ALL APPLICABLE MATERIAL RECORDS, MANUFACTURING RECORDS, STORAGE RECORDS, AND DISTRIBUTION RECORDS ACCORDING TO THE DESCRIPTIONS OF THE PRODUCT USED WITH THE CONCOMITANT DEVICE. ALL RECORDS REVEALED ALL PRODUCT LOTS WERE MANUFACTURED WITHIN SPECIFICATIONS, MAINTAINED, AND DISTRIBUTED IN ACCORDANCE WITH ALL FEDERAL, STATE AND OPERATING PROCEDURES. BASED ON RECORD REVIEW OF ALL AVAILABLE INFORMATION AND EVALUATION OF THE PRODUCT IN QUESTION, IT IS DETERMINED THE REVOLVE INSTRUMENT GRAPHIC CASE DESIGN CONFORMS TO ALL APPLICABLE STANDARDS AND THERE WAS NO DEVIATION IN THE MANUFACTURE PROCESS. THERE IS NO INDICATION THAT THE ISSUE WAS A RESULT OF PRODUCT DEFECT. | 0
2080449 | THE CUSTOMER STATED THE CELL-DYN SAPPHIRE CIRCUIT BREAKER WENT OFF AND SMOKE WAS SEEN COMING FROM THE INSTRUMENT. THE CUSTOMER UNPLUGGED THE CELL-DYN AND OBSERVED SMOKE AND A BURNT ODOR APPEARING FROM THE POWER DISTRIBUTION BOX INSTALLED UNDER THE INSTRUMENT BENCH. A LEAK FROM THE REAGENT RESERVOIR AREA, ONTO THE INSTRUMENT, ONTO THE BENCH, THEN ONTO THE POWER DISTRIBUTION BOX WAS FOUND TO BE THE CAUSE. THE CELL-DYN SAPPHIRE WBC PRE-HEATER HAD BEEN DAMAGED BY THE DRIP AND WAS REPLACED TO RESTORE THE ANALYZER TO SPECIFICATION. THE DILUENT REAGENT RESERVOIR CHECK VALVE CLAMP WAS RESHAPED TO STOP THE LEAK OF FLUID. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. THERE WAS NO INJURY TO PERSONNEL REPORTED. (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. NO MATERIAL WAS AVAILABLE FOR RETURN. THE REAGENT LEAK WAS ADDRESSED TO RESOLVE THE ISSUE. NO ADVERSE TREND WAS IDENTIFIED. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
2083322 | IT WAS REPORTED THAT THE CARTO 3 AND THE BARD EP RECORDING SYSTEM WAS RANDOMLY LOSING BOTH INTRACARDIAC AND BODY SURFACE EKGS. THESE LOSSES ALSO EFFECTED THE BARD RECORDING SYSTEM. THE CASE WAS STARTED WITH DIRECT CONNECT CABLES. CHANGING TO IC OUT CABLES DID NOT RESOLVE THE ISSUE. THE CARTO 3 AND BARD WERE REBOOTED BUT THE BEHAVIOR CONTINUED. THE CARTO WAS BYPASSED AND THE EKGS WERE CLEAN ON THE BARD SYSTEM. (B)(4). IN RESPONSE TO THIS COMPLAINT, THE BWI REPRESENTATIVE REMOVED THE BARD QUICK CONNECT CABLE AND THE ISSUE CLEARED UP. THE BWI REPRESENTATIVE ALSO CHECKED THE CONNECTIVITY BETWEEN THE BARD AND CARTO 3 SYSTEM, HOWEVER ISSUE COULD NOT BE DUPLICATED. THE CARTO 3 SYSTEM HAS BEEN USED TWICE SINCE ISSUE WAS CALLED IN WITH NO PROBLEMS BETWEEN THE BARD AND CARTO 3 SYSTEM. CARTO 3 SYSTEM WAS READY FOR USE.THE DHR ASSOCIATED WITH CARTO (B)(4) WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE.><(B)(4). INVESTIGATION IS IN PROGRESS AND A SUPPLEMENTAL WILL BE SUBMITTED. | 1
2085076 | AFTER A DONATION, WHEN THE DONOR GOT HOME SHE WAS LIGHTHEADED AND DIZZY AND STARTED VOMITING. THE DONOR IS AN EMPLOYEE AT THE CENTER, AND REPORTED FEELING FINE THE NEXT DAY. CUSTOMER IS ASKING: "DOES THE NEW SOFTWARE GIVE MORE ACDA OR ANYTHING THAT MIGHT BE MAKING PEOPLE WHO DID NOT HAVE REACTIONS PREVIOUSLY HAVE THEM NOW?". (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2087645 | THE ACCOUNT STATED THAT WHEN THEY TURNED THE AXSYM ANALYZER ON THE UPS MADE A LOUD NOISE AND SMOKE CAME OUT OF THE UNIT. NO INJURY WAS REPORTED. (B)(4). THE CUSTOMER COMPLAINT TEXT WAS REVIEWED AND INDICATES THAT THE UPS ISSUE WAS RELATED DIRECTLY TO THE UPS AND WAS NOT CAUSED BY THE AXSYM ANALYZER. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE AXSYM SYSTEM OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO PROVIDE ADEQUATE INFORMATION REGARDING POTENTIAL ELECTRICAL HAZARDS. A REVIEW OF THE SAFETY HAZARDS MEMO INDICATES THE AXSYM ANALYZER IS CERTIFIED TO THE APPROPRIATE SAFETY STANDARDS, AND ADEQUATE PROTECTION IS PROVIDED FOR THE OPERATOR AGAINST THE SPREAD OF FIRE FROM THE EQUIPMENT. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY. | 0
2087765 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. THE CUSTOMER WAS NOT WILLING TO PROVIDE THE PATIENT'S AGE OR DATE OF BIRTH. THE CUSTOMER CALLED TO REPORT AN OPERATOR ENTERED AN INCORRECT PLATELET COUNT OF 476 FOR A DONOR INSTEAD OF THEIR ACTUAL COUNT OF 239. THE CUSTOMER WANTED TO KNOW WHAT IMPACT THERE MIGHT BE DUE TO THIS MISTAKE. THE MEDICAL DIRECTOR CALLED THE DONOR AND ASKED THAT THEY RETURN FOR A POST PLATELET COUNT BUT HE REFUSED. (B)(4). INVESTIGATION: THE DISPOSABLE SET WAS NOT RETURNED FOR INVESTIGATION AS THERE WAS NO PROBLEM W/THE SET. A SERVICE CALL WAS NOT PLACED BECAUSE THE TRIMA SYSTEM OPERATED AS INTENDED WITH THE INFORMATION THE OPERATOR ENTERED. THE DONOR DID NOT EXPERIENCE ANY ADVERSE SYMPTOMS NOR WAS ANY MEDICAL INTERVENTION REQUIRED IN THIS INCIDENT. DHR WAS REVIEWED AND NO PROBLEMS WERE NOTED. ROOT CAUSE: OPERATOR ERROR. CORRECTIVE/PREVENTIVE ACTION: THE SUPPORT SPECIALIST WHO HANDLED THE CALL GAVE THE CUSTOMER SOME RE-TRAINING AS TO THE DANGER OF RUNNING A DONOR WITH THE WRONG PARAMETERS. | 0
2087887 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIME FRAME. KINKED TUBING WAS FOUND ON THE EXCHANGE LINE DURING SETUP. NO INJURY OCCURRED, AS NO DONOR WAS ATTACHED TO THE DISPOSABLE. PATIENT INFORMATION WAS NOT RELEVANT SINCE THE ISSUE OCCURRED DURING SETUP. (B)(4). INVESTIGATION: A KINK/FOLD WAS NOTED IN THE REPLACEMENT FLUID LINE. ROOT CAUSE: THE COMMON CAUSES OF TUBING KINKS, INDENTATIONS AND COMPRESSIONS ARE TYPICALLY RELATED TO SLIGHT VARIATIONS DURING TRAY PACKING. CONTENTS SHIFTING MAY ALSO OCCUR DURING PACKAGING AND TRANSPORT TO STERILIZATION. FOLLOWING EXPOSURE TO THE HEAT AND HUMIDITY OF THE STERILIZATION CYCLE, THE TUBING MAY ASSUME A DEFORMED SHAPE. CORRECTIVE/PREVENTIVE ACTION: NO CORRECTIVE AND PREVENTIVE ACTIONS BY CARIDIANBCT QUALITY ASSURANCE WILL OCCUR AT THIS TIME. | 0
2087892 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIME FRAME. PLATELET CLUMPING OR FIBRIN CLOT IN RESERVOIR OCCLUDED FILTER AND RESULTED IN AIR IN RETURN LINE. MEDICAL DIRECTOR REFUSED TO PROVIDE A PATIENT IDENTIFIER. (B)(4). INVESTIGATION: THIS SITE WAS HAVING MULTIPLE ISSUES WITH CLUMPING. THE RDF (RUN DATA FILES) ANALYSIS FOUND IN THE HEAT RECORD, FOR THIS ISSUE SHOWS THAT THERE WERE SEVERAL INLET PRESSURE TOO LOW ALARMS AND A CENTRIFUGE POWER COULD NOT BE ENABLED ALARM. DISCUSSION WITH THE CUSTOMER REVEALED THEY ARE STARTING WITH AN AC RATIO (15) THAT MAY BE TOO LOW. ESPECIALLY WHEN ADDING CALCIUM TO THE REPLACEMENT FLUID. CONCLUSION: THIS PROBLEM IS LIKELY DUE TO THE LOW AC RATIO BEING USED IN THE PROCEDURE. | 0
2087894 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. CUSTOMER REPORTED THAT TWO COLLECTIONS WERE DONE ON THIS MACHINE ((B)(6) 2010). BOTH TIMES, THE PROCEDURE STARTED FINE WITH NO ALARMS, THEN THE RETURN PRESSURE BAR GRAPH WENT TO THE TOP, AND THE DONORS HAD INFILTRATIONS AS A RESULT. NO MEDICAL INTERVENTION WAS REQUESTED OR PERFORMED. BASIC FIRST AID WAS USED TO TREAT THE DONOR INFILTRATIONS. ICE WAS APPLIED AND THE DONORS WERE GIVEN INSTRUCTIONS FOR HOME CARE. CUSTOMER WAS NOT WILLING TO GIVE THE DOBS.><THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. CUSTOMER REPORTED THAT TWO COLLECTIONS WERE DONE ON THIS MACHINE ((B)(6) 2010). BOTH TIMES, THE PROCEDURE STARTED FINE WITH NO ALARMS, THEN THE RETURN PRESSURE BAR GRAPH WENT TO THE TOP, AND THE DONORS HAD INFILTRATIONS AS A RESULT. NO MEDICAL INTERVENTION WAS REQUESTED OR PERFORMED. BASIC FIRST AID WAS USED TO TREAT THE DONOR INFILTRATIONS. ICE WAS APPLIED AND THE DONORS WERE GIVEN INSTRUCTIONS FOR HOME CARE. CUSTOMER WAS NOT WILLING TO GIVE THE DOBS. (B)(4). FIELD DIAGNOSTIC/CORRECTIONS: THE EQUIPMENT ALARMED WITH RETURN PRESSURE TOO HIGH. THE RUN DATA FILES WERE ANALYZED FOR THE PROCEDURE AND SHOW THAT THE TRIMA OPERATED AS INTENDED. EVALUATION SUMMARY: A SERVICE CALL WAS PLACED FOR THIS INCIDENT. THE SERVICE REP VERIFIED THE APS/CPS SENSORS, PUMP SPEEDS, AND OCCLUSIONS WITH NO ISSUES FOUND. AUTO TEST AND SALINE RUNS WAS PERFORMED WITH NO PROBLEMS. THE DEVICE HISTORY RECORD FOR THIS EQUIPMENT WAS REVIEWED AND NO ISSUES WERE FOUND. ROOT CAUSE: DONOR ACCESS ISSUES. NO FAILURE OF THE DEVICE WAS FOUND.><(B)(4). FIELD DIAGNOSTIC/CORRECTIONS: THE EQUIPMENT ALARMED WITH RETURN PRESSURE TOO HIGH. THE RUN DATA FILES WERE ANALYZED FOR THE PROCEDURE AND SHOW THAT THE TRIMA OPERATED AS INTENDED. EVALUATION SUMMARY: A SERVICE CALL WAS PLACED FOR THIS INCIDENT. THE SERVICE REP VERIFIED THE APS/CPS SENSORS, PUMP SPEEDS, AND OCCLUSIONS WITH NO ISSUES FOUND. AUTO TEST AND SALINE RUNS WAS PERFORMED WITH NO PROBLEMS. THE DEVICE HISTORY RECORD FOR THIS EQUIPMENT WAS REVIEWED AND NO ISSUES WERE FOUND. ROOT CAUSE: DONOR ACCESS ISSUES. NO FAILURE OF THE DEVICE WAS FOUND. | 0
2087895 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. CUSTOMER REPORTED THAT ABOUT 10 MINUTES INTO A TRIMA RUN, A MALE DONOR HAD COMPLAINED OF ARM DISCOMFORT. A SMALL BRUISE, ABOUT THE SIZE OF A QUARTER WAS OBSERVED AND THE PROCEDURE WAS DISCONTINUED. THERE WAS NO MEDICAL INTERVENTION BEYOND THE APPLICATION OF COLD COMPRESS. NO OTHER ADVERSE EFFECTS TO THE DONOR FROM THIS INCIDENT. (B)(4). THE RDF'S (RUN DATA FILES) WERE ANALYZED FOR THE PROCEDURE AND SHOW THAT THE TRIMA OPERATED AS INTENDED. THE CUSTOMER'S BIOMED CHECKED THE MACHINE. CARIDIANBCT HAS NOT BEEN MADE AWARE OF THOSE RESULTS. THE DHR (DEVICE HISTORY RECORD) WAS REVIEWED AND NO ISSUES NOTED. IF THERE IS A VERY SLOW LEAK FROM THE VENIPUNCTURE SITE DURING A COLLECTION PROCEDURE, THE TRIMA ACCEL SYSTEM MAY NOT DETECT THIS AND THEREFORE WOULD NOT TRIGGER THE RETURN PRESSURE ALARM. THIS SITUATION COULD BE CREATED OR EXACERBATED IF AN ADJUSTMENT OF THE NEEDLE IS MADE AFTER RECEIVING A RETURN PRESSURE HIGH ALARM. THE RETURN PRESSURE HIGH ALERT OCCURS WHEN THE RETURN PRESSURE REACHES +310 MM HG. THE OPERATOR DOES NEED TO MONITOR THE VENIPUNCTURE SITE AND THE DONOR DURING COLLECTION PROCEDURES TO BE SURE THEY ARE NOT HAVING ANY DISCOMFORT. ROOT CAUSE: DONOR ACCESS ISSUES. CONCLUSION: THE DEVICE OPERATED AS INTENDED. | 0
2087900 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. DURING THE PROCEDURE, AGGREGATES WERE VISIBLE IN THE COLLECT LINE TO RESERVOIR AND AT THE FILTER IN THE RESERVOIR, AND WERE VISIBLE AT THE INTERFACE IN THE CENTRIFUGE IN THE FORM OF "WAVES". THE PROCEDURE, SPECTRA OPTIA LYMPHOCYTE COLLECTION, WAS 3X TBV (15L), AND WAS STOPPED PRIOR TO THE END OF THE PROCEDURE BECAUSE OF CRAMPS/TETANY OF HANDS AND FEET (THE PATIENT HAD CRAMPS/TETANY AT THE BEGINNING OF THIS PROCEDURE DUE TO A PREVIOUS PROCEDURE). DURING THE LYMPHOCYTE COLLECTION, TINGLING OCCURRED AROUND NOSE, MOUTH AND HANDS. CALCIUM GLUCONATE (3 X 2 ML AT 15 MINUTE INTERVALS) WAS ADMINISTERED WITH NO IMPROVEMENT. ON PRESCRIPTION OF APHAERESIS PHYSICIAN: ADMINISTRATION OF MAGNITOP AND CACIT ORAL, BECAUSE OF NO IMPROVEMENT WITH CALCIUM GLUCONATE. THE AGE OF THE PATIENT WAS NOT PROVIDED BY THE CUSTOMER. (B)(4). INVESTIGATION: A PREVIOUS STEM CELL COLLECTION PROCEDURE WAS PERFORMED IN WHICH THE PATIENT BEGAN EXHIBITING CRAMPS/TETANY OF THE HANDS AND FEET FROM THE START OF THE PROCEDURE. THIS PATIENT REACTION INCREASED DURING THE SECOND PROCEDURE (LYMPHOCYTE COLLECTION CONDUCTED ON SPECTRA OPTIA) FORCING PROCEDURE CESSATION. AGGREGATES WERE VISIBLE IN THE COLLECT LINE DURING THE SECOND PROCEDURE. RUN DATA FILE ANALYSIS OF THE SPECTRA OPTIA LYMPHOCYTE COLLECTION PROCEDURE REVEALED NO ISSUES WITH THE EQUIPMENT. MANUFACTURING RECORD REVIEW: A REVIEW OF THE MANUFACTURING RECORDS WAS CONDUCTED BY THE COMPLAINT INVESTIGATOR IN RELATION TO THIS FAILURE MODE. THERE WERE NO ISSUES WITH THIS LOT. ROOT CAUSE: THE DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. PLATELET AGGREGATION AS DESCRIBED RESULTING IN VISIBLE AGGREGATES IS INDICATIVE OF UNDER-ANTICOAGULATED SPECTRA OPTIA SYSTEM. THE RECOMMENDED RESPONSE TO OBSERVANCE OF THIS BEHAVIOR IS TO AGGRESSIVELY DECREASE THE INLET: AC RADIO TO REDUCE THE SIZE AND OCCURRENCE RATE OF AGGREGATE FORMATION. THE PATIENT REACTION OF CRAMPS/TETANY OF THE HANDS AND FEET WAS EXPECTED TO SOME DEGREE DUE TO THE PRESCRIBED PRE-MEDICATION AND SIMILAR REACTION ON PREVIOUS STEM CELL COLLECTION PROCEDURE. CONCLUSION: UNDER-COAGULATION. | 0
2087911 | THIS REPORT IS BEING FILED AS A RESULT OF CHANGES TO OUR MDR EVALUATION PROCESS. WE CONTINUE TO REFINE OUR MDR PROCESS TO BETTER ALIGN WITH CURRENT AGENCY POLICY. WE REGRET THAT THIS CHANGE HAS CAUSED THIS MDR TO BE FILED OUTSIDE THE REQUIRED TIMEFRAME. THE CUSTOMER REPORTED A CHANNEL LEAK TO THE LRS CHAMBER. CUSTOMER IGNORED REPEATED ATTEMPTS TO GATHER PATIENT IDENTIFIER INFORMATION. EVEN THOUGH THIS DEVICE IS NOT CURRENTLY CLEARED BY FDA, THIS ISSUE COULD OCCUR ON A CLEARED DEVICE. (B)(4). INVESTIGATION: THE CHANNEL WAS RETURNED FOR EVALUATION. A SMALL LEAK NEAR THE SNOUT WELD OF THE CHANNEL CONNECTOR TO THE CHANNEL VINYL WAS CONFIRMED. NO OTHER ANOMALIES WERE FOUND. CONSULTED RF ENGINEERING AS SUBJECT MATTER EXPERTS TO DETERMINE IF THIS FAILURE IS A RESULT OF THE MANUFACTURING PROCESS CREATING THE CHANNEL ASSEMBLY. LEAK WAS DETECTED BY CENTRIFUGE BASIN LEAK DETECTOR AND MACHINE STOPPED PUMPS AND PAUSED THE PROCEDURE. MANUFACTURING RECORD REVIEW: FOR THIS LOT, THERE WERE NO IN-PROCESS INSPECTION FAILURES OR OBSERVATIONS, NO SIMULATED TESTING OBSERVATIONS OR FAILURES, AND THE PRODUCT MET ALL ACCEPTANCE CRITERIA FOR RELEASE. ROOT CAUSE: RF ENGINEERING CONFIRMED THAT THIS FAILURE MODE CAN BE DUE TO THE MANUFACTURING PROCESS WHICH CREATES THE CHANNEL ASSEMBLY. | 0
2089002 | THE ACCOUNT STATED THAT THE ARCHITECT I2000SR ANALYZER GENERATED A POSITIVE TOXO IGG RESULT OF 19.85 IU/ML AND A NEGATIVE TOXO IGM RESULT. THE PATIENT WAS KNOW TO BE NEGATIVE. THE SAMPLE WAS RECENTRIFUGED AND REPEATED WITH A NEGATIVE RESULT. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. (B)(4). REVIEW OF THE MESSAGE HISTORY FOR THE ARCHITECT I2000SR FOUND MANY ERROR CODE 3350, ASPIRATION ERROR, SAMPLE PIPETTOR. FURTHER REVIEW FOUND ERROR CODE 5000, SAMPLE PIPETTOR MOVEMENT RESTRICTED WAS GENERATED INDICATING A PROBABLE PROBE CRASH. THE CUSTOMER WAS DIRECTED TO REPLACE AND CALIBRATE THE PROBE. FEWER ERROR 3350 WERE NOTED AFTER THE PROBE WAS CHANGED. THE PROBE DAMAGED AS A RESULT OF A PROBE CRASH WAS DETERMINED TO BE THE PROBABLE CAUSE FOR THE ISSUE. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE OF ERRATIC RESULTS AND PROBE REPLACEMENT. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
2092530 | CUSTOMER REPORTS THAT THE PACS EXTEND FOR POWERSCRIBE 5.0 IS OPENING THE WRONG RECORD FOR REPORTING OF THE EXAM. WHEN THE RADIOLOGIST WAS DICTATING ON PT 1, AND WAS INTERRUPTED TO LOOK AT PT B (WHICH HAD ALREADY BEEN DICTATED), THEN CLOSED PT B, PT B CONTINUED TO DISPLAY. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. GE REFERENCE #: (B)(4). | 1
2092888 | DURING DEBRIDEMENT SURGERY, A PT HAD AN AIR EMBOLI WHILE RECEIVING AN RBC TRANSFUSION DURING SURGERY. THE HOSPITAL NOTED AIR IN THE PRODUCT BEFORE TRANSFUSION. ADMINISTERED OXYGEN AND DID CENTRAL LINE ASPIRATION. VITAL SIGNS RETURNED TO NORMAL 10 MINUTES AFTER TREATMENT. THE SET WILL NOT BE RETURNED AS IT HAS BEEN THROWN AWAY. PT INFO IS NOT AVAILABLE AT THIS TIME. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2092909 | THE CUSTOMER COMPLAINED ABOUT THE NEEDLE GUARD ON THE TRIMA DISPOSABLE SET. THE CUSTOMER STATED THAT THE BUTTERFLY WINGS ARE NOT ALWAYS ALIGNED WITH THE SLEEVES ON THE PROTECTOR, THEREFORE, THEY SOMETIMES HAVE TO REPOSITION THEM, WHICH CREATES A HAZARD. TWO NURSES HAVE STUCK THEMSELVES DOING SO. CARIDIANBCT IS AWAITING FURTHER INFO FROM THE CUSTOMER. (B)(4). INVESTIGATION, EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2093992 | A CUSTOMER REPORTED WHILE SETTING UP A DOUBLE NEEDLE AMICUS PROCEDURE THE RETURN LINE WAS NOT PRIMED BY THE OPERATOR PRIOR TO VENIPUNCTURE AND APPROXIMATELY 1.2 ML OF AIR WAS INFUSED. DONOR DID NOT EXPERIENCE ANY SYMPTOMS OF AIR EMBOLISM. THE PROCEDURE WAS DISCONTINUED IMMEDIATELY. THE FACILITIES MEDICAL DIRECTOR WAS INFORMED OF THE SITUATION AND DONOR FOLLOW UP WAS RECOMMENDED. DONOR CONDITION WAS STABLE AND LEFT ON OWN ACCORD. THE CENTER FOLLOWED UP WITH THE DONOR AT A LATER DATE AND THERE WERE NO ADVERSE EFFECTS REPORTED. PLEASE NOTE THAT THE OPERATORS MANUAL INDICATES A WARNING AND A CAUTION REGARDING THE PRIMING OF THE KIT. WARNING: OPENING THE RETURN LINE ROLLER CLAMP AFTER VENIPUNCTURE AND BEFORE THE BLOOD SAMPLING POUCH IS FILLED MAY RESULT IN AN AIR EMBOLISM. CAUTION: THE KIT SHOULD BE ALMOST COMPLETELY FULL OF FLUID (SALINE AND ACD) AT THE END OF PRIME. THE CONTAINERS ON THE FRONT SCALE HOOKS SHOULD HAVE LITTLE OR NO FLUID IN THEM, EXCEPT FOR THE CONTAINER ON THE FRONT LEFT SCALE HOOK ON A DOUBLE NEEDLE OR MNC KIT, OR THE CONTAINER ON THE FRONT RIGHT SCALE HOOK ON A SINGLE NEEDLE KIT. SOME AIR SHOULD BE EXPECTED NEAR THE FISTULA(S). IF THERE ARE LARGE AMOUNTS OF AIR IN THE KIT IMMEDIATELY FOLLOWING PRIME, DO NOT BEGIN COLLECTION. REMOVE THE KIT AND TRY INSTALLING AND PRIMING ANOTHER KIT. (B)(4). | 0
2094440 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE COULTER LH 500 INSTRUMENT GENERATED PRINTOUTS DID NOT MATCH THE LABORATORY INFORMATION SYSTEM (LIS) PRINTOUTS FOR BASOPHIL PARAMETER COUNT (BA#) FOR ONE PATIENT'S INITIAL AND RERUN RESULTS. THE RESULTS WERE REPORTED OUT OF THE LABORATORY, BUT WERE IMMEDIATELY CAUGHT AND A CORRECTED REPORT WAS ISSUED. NO TREATMENT WAS GIVEN OR WITHHELD DUE TO THE ERRONEOUS RESULTS. THERE WAS NO DEATH, SERIOUS INJURY, OR CHANGE TO PATIENT TREATMENT IN THIS EVENT. SAMPLE WAS COLLECTED IN A PLASTIC BD TUBE. THE INSTRUMENT IS PERFORMING WITHIN QC SPECIFICATIONS. CONTROLS WERE RUN BEFORE AND AFTER THE INCIDENT AND RECOVERED WITHIN SPECIFICATIONS. SERVICE WAS NOT DISPATCHED FOR THIS EVENT. THE ROOT CAUSE OF THIS ISSUE IS UNKNOWN AT THIS TIME. | 1
2095947 | WHILE A RADIOLOGIST IS DICTATING HIS WORKLIST (ACTIVE WORKLIST SETTING). THE IMAGE AREA (IA) DOESN'T REFRESH PROPERLY LEAVING THE STUDY THAT IS ALREADY DICTATED ON THE SCREEN, WHILE THE TEXT AREA AND PATIENT SLIDER HAVE INFORMATION FROM THE NEXT STUDY. THE POTENTIAL ISSUE IS THAT TWO DIFFERENT PATIENTS CAN POTENTIALLY DISPLAY IN THE IA SIMULTANEOUSLY. EVAL SUMMARY: DESCRIPTION OF EVENT: THIS ISSUE INTERMITTENTLY MANIFESTS AFTER DICTATION OF A STUDY HAS BEEN COMPLETED. WHAT CAN POTENTIALLY OCCUR IS THAT THE TEXT AREA AND THE IMAGE AREA PATIENT INFORMATION HAS UPDATED TO THE NEXT STUDY; HOWEVER, THE ACTUAL IMAGES DISPLAYED REMAIN ON THE PREVIOUSLY DICTATED STUDY. NOTE: THAT ALL DEMOGRAPHICS ARE CORRECT AND EACH PATIENT IS IDENTIFIED CORRECTLY. EVALUATION: AGFA HEALTHCARE EVALUATED THE IMPAX 6.5 CLIENT SOFTWARE AND LOCATED THE ERROR IN THE SOFTWARE RESPONSIBLE FOR THIS PROBLEM. A SOFTWARE FIX (IMPAX 6.5 SU1) WAS CREATED TO ADDRESS AND FIX THE ISSUE. (B)(4). | 1
2098215 | THE CUSTOMER STATED RBC QUALITY CONTROLS WERE OUT OF RANGE LOW ON A CELL-DYN EMERALD ANALYZER USING CELL-DYN EMERALD CLEANER LOT 4349. THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. (B)(4). CONCOMITANT MEDICAL PRODUCTS: CELL-DYN EMERALD CLEANER MANUFACTURED 8/10/10, EXPIRES 6/18/12 A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. AN INVESTIGATION IS IN PROCESS.><A PRODUCT DEFICIENCY WAS IDENTIFIED. THE INVESTIGATION DEMONSTRATED CONCLUSIVELY THAT ONLY CELL-DYN EMERALD CLEANER LOT 4349 WAS AFFECTED. (B)(4). | 0
2099346 | WHILE REPLACING THE CELL-DYN WASTE CONTAINER, THE CUSTOMER IDENTIFIED A BROKEN CABLE OF THE WASTE SENSOR. THE WASTE SENSOR WAS THEN REPLACED IN ORDER TO RESOLVE THIS ISSUE. NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN EXPANDED INVESTIGATION HAS BEEN CONDUCTED TO EVALUATE THIS ISSUE. THE INVESTIGATION OF THE RETURNED PARTS FROM THE FIELD SHOWED THAT THE PART HAS BEEN IN USE GREATER THAN TWO YEARS. ALTHOUGH, THE WASTE OUTLET TUBE ASSEMBLY MEETS THE RELIABILITY REQUIREMENT, THIS PART WILL WEAR OUT AND THERE EXISTS A POTENTIAL FOR THE WASTE CONTAINER TO OVERFLOW WITH LIQUID WASTE WHEN THE WASTE SENSOR FAILS AT THE END OF ITS LIFE CYCLE. THEREFORE; THE PART WILL NEED PERIODIC REPLACEMENT AS IT IS IN CONTACT WITH BIOHAZARD MATERIAL AND CUSTOMER USAGE/HANDLING. THE PART CAN BE USED ON MULTIPLE CELL-DYN INSTRUMENT MODELS. A REVIEW OF THE CELL-DYN SYSTEM OPERATORS MANUALS SHOWED ADEQUATE COVERAGE OF BIOHAZARD EXPOSURE AND WASTE OVERFLOW AS WELL AS TROUBLESHOOTING INFORMATION FOR POTENTIAL CAUSES. ABBOTT RECOMMENDED A REPLACEMENT SCHEDULE FOR THE WASTE LINE ASSEMBLY FOR THE CELL-DYN SYSTEMS. THE WASTE BOTTLE CABLE IS A SUBCOMPONENT OF THE WASTE LINE ASSEMBLY, CHANGING THE WASTE LINE ASSEMBLY WILL RESULT IN CHANGING THE WASTE BOTTLE CABLE. AS PART OF THE CORRECTIVE ACTION, A PRODUCT CORRECTION LETTER, (B)(4), WAS ISSUED TO ALL AFFECTED CUSTOMERS. IN THIS COMMUNICATION, ABBOTT RECOMMENDED REPLACING THE PART EVERY SIX MONTHS. AN UPDATE TO THE PRODUCT LABELING WILL BE ADDED AS WELL WITH REGARDS TO THIS RECOMMENDATION. A TAG THAT CAN BE AFFIXED TO THESE ITEMS WAS SENT WITH THE CUSTOMER LETTER. THIS TAG INCLUDES FIELDS TO RECORD INSTALLATION AND REPLACEMENT DATES. THE TAGS WILL ALSO BE INTRODUCED INTO REPLACEMENT ASSEMBLIES. | 0
2101214 | A CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTION ON THE ECHO. THE SAMPLE HAD AN ANTI-C AND ANTI-E. A SERVICE CALL WAS MADE. THE REAGENT PROBE WAS REPLACED. THE SAMPLE WAS RE-TESTED AND RESULTED AS EXPECTED. THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. | 0
2102246 | THE CUSTOMER REPORTS THAT THEIR THIRD PARTY (B)(4) SYSTEM IS CONFUSING THE ASCENSION NUMBER AND PT ID NUMBER. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. GE REFERENCE #: (B)(4). | 1
2103364 | IT WAS REPORTED THAT DURING PT USE, THE CARESCAPE B850 PT MONITOR EMITTED AN UNUSUAL CONSTANT BEEP-BEEP TONE THAT DID NOT AUDIBLY CORRESPOND TO AN EXPECTED ALARM TONE OR OTHER RECOGNIZABLE TONE, AND DID NOT CORRELATE TO ANY VISUAL PT CONDITION. AS LONG AS THIS TONE IS SOUNDING NO OTHER DISCERNIBLE AUDIBLE ALARMS WILL BE ANNOUNCED. VISUAL ALARMS WERE NOT AFFECTED. THERE WAS NO DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT, NOR AN INDICATION THAT INTERVENTION WAS PROVIDED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. (B)(4), PT INFO IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSP. CLARIFICATION: THE ISSUE WAS DISCOVERED DURING A RETROSPECTIVE REVIEW OF COMPLAINTS RELATED TO A KNOWN ISSUE THAT RESULTED IN THE EVENT REPORTED IN MDR 2124823-2011-00052. | 1
2104958 | THE CUSTOMER HAD REPORTED FALSE POSITIVE REACTIONS CAUSED BY CARRY OVER. THE CORRECT FUNCTION OF THE AFFECTED LOTS OF REAGENTS WERE PROVED BY TESTING THE RETENTION SAMPLES OF QC LABORATORY ON TANGO. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DIDN'T OBSERVE ANY FALSELY POSITIVE REACTIONS. THE REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE AFFECTED NEGATIVELY THE QUALITY OF THE COMPLAINED LOT. IN GENERAL CARRY-OVER EVENTS CANNOT BE EXCLUDED IF APPLYING INSTRUMENTATION WITH A DESIGN LAYOUT, THAT IS AMONGST OTHERS, ALSO TRUE FOR TANGO OPTIMO. THIS MATTER IS ALSO DISPLAYED IN ITS SPECS. DESPITE COATED NEEDLES AS WELL AS RINSING STEPS, SUCH A PHENOMENON MAY OCCUR IN VERY RARE CASES, IS USUALLY BASED ON A SAMPLE-SPECIFIC REACTION AND CAN BE CAUSED BY A HIGH PROTEIN LEVEL IN PLASMA (INDUCED BY SEVERE SICKNESS OR APPLICATION OF DRUGS) OR ANTIBODIES WITH A HIGH TITRE. SINCE ANY SPECIMEN BEING CONSPICUOUS DUE TO A POSITIVE ABS RESULT SHOULD UNDERGO RETESTING AND BECAUSE ANY ABS-POSITIVE SAMPLE FROM A BLOOD DONOR SHOULD BE EXCLUDED FROM TRANSFUSION, NO SERIOUS THREAT NEITHER FOR PT NOR USER OR DEVICE MAY ARISE DUE TO A CARRY-OVER EVENT. IN ACCORDANCE TO THE RESULTS OBTAINED AT THE CUSTOMER SITE, TESTING AT BMD CONFIRMED A CARRY-OVER EVENT FROM ONE WELL INTO THE FOLLOWING ONE OCCURRING DURING SEQUENTIAL PIPETTING OF DIFFERENT BLOOD PLASMA CAUSED BY THE HIGH TITRE OF ANTIBODY PRESENT IN THE ORIGINAL SAMPLE. EACH PIPETTING STEP OF INDIVIDUAL SAMPLES IS FOLLOWED BY A RINSING STEP OF THE CORRESPONDING NEEDLE RESULTING IN A DILUTION OF BIOLOGICAL RESIDUES BY APPROX. 1:1000. IN GENERAL THIS DILUTION FACTOR IS SUFFICIENT TO PREVENT A DETECTABLE TRANSFER OF MATERIAL INTO THE FOLLOWING SPECIMEN. | 0
2109318 | THE CUSTOMER STATED THAT SUSPECT CELL-DYN 1800 HEMOGLOBIN RESULTS OF 7.4 AND 6.7 G/DL WERE GENERATED FOR A PATIENT SAMPLE. THE SAME SAMPLE WAS TESTED AGAIN AND THE HEMOGLOBIN RESULTS WERE 15.1 G/DL TWICE. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. (B)(4). AN ABBOTT CUSTOMER TECHNICAL ADVOCATE ADVISED PERFORMING WBC SUPPLEMENTAL AND APERTURE PLATE CLEANING, AND HGB FLOW CELL CLEANING WITH BLEACH. THE CUSTOMER WAS REQUESTED TO CALL BACK WITH TROUBLESHOOTING RESULTS. THE CUSTOMER PERFORMED THE RECOMMENDED TROUBLESHOOTING, BUT THE PRECISION RUN FAILED. ABBOTT FIELD SERVICE WAS DISPATCHED AND FOUND A CLOGGED BUBBLE MIX TUBING. THE TUBING WAS REPLACED. FIELD SERVICE RESEATED THE LYSE TUBING IN THE NORMALLY CLOSED VALVE AND LUBED THE SYRINGE DRIVES. QUALITY CONTROLS WERE VERIFIED WHICH PASSED. THE INSTRUMENT WAS PERFORMING WITHIN SPECIFICATIONS. A REVIEW OF COMPLAINT TRACKING AND TRENDING WAS PERFORMED AND NO ISSUES WERE IDENTIFIED. BASED ON THE INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED FOR THE CELL-DYN 1800 ANALYZER RELATED TO THE REPORTED CUSTOMER ISSUE.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
2109319 | THE CUSTOMER STATES THAT A LOUD POPPING NOISE WAS HEARD COMING FROM THE CELL-DYN 1800 ANALYZER. THIS WAS FOLLOWED BY "LOTS" OF VISIBLE SMOKE POURING OUT OF THE ANALYZER'S VENTS. A BURNING SMELL WAS ALSO NOTICED. NO FLAMES WERE SEEN. THE CUSTOMER DISCONNECTED THE INSTRUMENT'S POWER CORD FROM THE WALL OUTLET. THERE WERE NO INJURIES OR DAMAGE TO THE SURROUNDING LAB ENVIRONMENT. THE ABBOTT CUSTOMER TECHNICAL ADVOCATE INSTRUCTED THE CUSTOMER TO INSTITUTE THEIR EMERGENCY PLAN IF SMOKE CONTINUES OR FLAMES ARE SEEN. A SERVICE CALL WAS INITIATED. (B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><FOR THE FOLLOW-UP MDR SUBMITTED ON (B)(4) 2011, MFR RECEIVED DATE SHOULD HAVE BEEN (B)(4) 2011.><AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) REPLACED THE POWER SUPPLY OF THE CELL-DYN 1800 ANALYZER AND SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. THE RETURNED CELL-DYN 1800 POWER SUPPLY ASSEMBLY FAILED TESTING DUE TO INTERMITTENTLY DROPPING VOLTAGES DURING INITIALIZATION. DURING TESTING THE DEVICE WAS VISUALLY INSPECTED FOR ANY DAMAGE; THERE WAS NO VISIBLE SMOKE DAMAGE PRIOR TO TESTING OR ANY SMOKE DURING TESTING. ANY SIGN OF SMOKE DAMAGE WOULD HAVE PREVENTED THE INVESTIGATOR FROM PERFORMING ANY FURTHER INVESTIGATION OF THE RETURNED CELL-DYN 1800 POWER SUPPLY ASSEMBLY. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE CELL-DYN 1800 SYSTEM OPERATOR'S MANUAL, LIST NUMBER 07H80-01, REVISION E, CONTAINS INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. BASED ON THE CURRENT INVESTIGATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED WITH THE CELL-DYN 1800 INSTRUMENT. THE ROOT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED. | 0
2109720 | THE ACCOUNT STATED THAT THE CELL-DYN EMERALD ANALYZER GENERATED AN INCREASED HEMOGLOBIN (HGB) RESULT OF 21.4 G/DL ON (B)(6) 2011. THE PATIENT SAMPLE WAS REPEATED THE NEXT DAY AND THE HGB WAS 15.5 G/DL. THE SAMPLE WAS ALSO SENT TO A REFERENCE LAB WHICH OBTAINED A HGB OF 15.0 G/DL. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. (B)(4). THE CUSTOMER FOUND THAT CLOG ERRORS HAD BEEN GENERATED PRIOR TO THE SUSPECT RESULT. THE CUSTOMER TECHNICAL ADVOCATE (CTA) DETERMINED THAT THE CUSTOMER DID NOT HAVE THE AUTOCLEAN CYCLE ENABLED. THE CTA HAD THE CUSTOMER ENABLE THE AUTOCLEAN CYCLES AND SET AT 15 CYCLES. THE CTA REVIEWED THE RECOMMENDED ACTIONS IN THE CELL-DYN EMERALD OPERATIONS MANUAL REGARDING CLOG ERRORS ON PATIENT RESULTS. THE CTA REVIEWED THE PROPER MIXING AND HANDLING PROCEDURE WITH THE CUSTOMER. THE CUSTOMER CONFIRMED THAT NO OTHER ERRONEOUS RESULTS HAD OCCURRED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE CELL-DYN EMERALD OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
2110104 | A CUSTOMER CONTACTED BECKMAN COULTER INC (BCI) AND REPORTED A BLOOD SPILL IN THE COULTER LH750 SLIDEMAKER FOLLOWED BY VACUUM SENSOR ERRORS. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT AND THE SPILL WAS CLEANED UP ACCORDING TO THE DEFINED LABORATORY PROTOCOL. THE FACILITY HAS AN EXPOSURE PLAN. MSDS WAS NOT REVIEWED BUT READILY AVAILABLE. THERE WAS NO REPORT OF BLOOD EXPOSURE TO MUCOUS MEMBRANES, EYES, MOUTH, OR OPEN WOUNDS. THERE WAS NO REPORT OF INJURY AND NO ONE SOUGHT MEDICAL ATTENTION. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. SERVICE WAS PROVIDED AND THE FIELD SERVICE ENGINEER (FSE) DISCOVERED A SLIDE JAMMING THE SHUTTLE AND NOT ALLOWING THE SLIDE TO BE UNDER THE PROBE CORRECTLY. THE SLIDE WAS REMOVED INTACT. IN ADDITION, THE VACUUM ERROR WAS CAUSED BY THE SENSOR GETTING WET. SUBSEQUENTLY THE SENSOR WAS REPLACED. REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES AND THE RESULTS MEET PUBLISHED PERFORMANCE SPECIFICATIONS. ROOT CAUSE FOR THE BLOOD SPILL WAS ASSOCIATED WITH ONE JAMMED SLIDE ON THE SHUTTLE CAUSING SENSOR ERROR DUE TO THE ENCLOSED SURROUNDING MOISTURE. | 0
2110140 | WHILE A FIELD SERVICE ENGINEER (FSE) FROM BECKMAN COULTER INC. (BCI) WAS PERFORMING A SERVICE OPERATION ON COULTER TQ-PREP WORKSTATION, SMOKE CAME UP FROM THE CONTROLLER BOARD. THE PROBLEM WAS DETECTED IMMEDIATELY AND POWER WAS REMOVED FROM THE UNIT. LAB COAT AND GLOVES WERE WORN WHILE WORKING ON THE UNIT. THE CUSTOMER DID NOT REPORT FLAMES, ARCHING OR SHOCK. A FIRE EXTINGUISHER WAS NOT USED NOR WAS THE FIRE DEPARTMENT CALLED. THE INSTRUMENT WAS NOT OPERATIONAL, AND THEREFORE, NO ERRONEOUS RESULTS WERE GENERATED OR REPORTED OUTSIDE OF THE LABORATORY. THERE WAS NO DEATH, INJURY, OR CHANGE TO PATIENT TREATMENT. THE CONTROLLER BOARD WAS REPLACED AND UNIT IS RUNNING ACCORDING TO THE SPECIFICATIONS. THE ROOT CAUSE FOR THE EVENT IS UNKNOWN. | 0
2110457 | IT WAS REPORTED THAT MEDICATION SCHEDULES ARE NOT VISIBLE ANYMORE AFTER AN APPEND_TRANSFER. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2111504 | A CUSTOMER WAS PERFORMING A MONONUCLEAR CELL PROCEDURE. SHE STATED THAT SHE LEFT THE SALINE OPEN AND BLOOD APPEARED IN THE RETURN ACCESS LINE, UP TO THE DRIP CHAMBER. THE CUSTOMER WAS UNABLE TO PERFORM RINSEBACK WITHIN THE EQUIPMENT DUE TO BLOOD BEING CLOTTED IN THE LINE AND IN THE MICROFILTER FOR SALINE. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR HYPOVOLEMIA. NO PT INFO IS AVAILABLE AT THIS TIME. (B)(4). INVESTIGATION, EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2113546 | THE CUSTOMER STATES THAT THE ARCHITECT I2000 SR ANALYZER HAD A DAMAGED WASTE PUMP CONNECTOR THAT WAS NOT ABLE TO CONNECT TO THE WASTE HOSE CORRECTLY (IT IS UNKNOWN HOW THE CONNECTOR WAS DAMAGED). SOME OF THE FLUID WASTE SPLASHED INTO THE FACE OF THE INSTRUMENT OPERATOR. THE OPERATOR IMMEDIATELY WASHED THEIR FACE AND DID NOT SEEK ANY MEDICAL ATTENTION. NO FURTHER HEALTH CONCERNS WERE NOTED. (B)(4). A REPLACEMENT WASTE PUMP WAS INSTALLED AT THE CUSTOMER SITE TO CORRECT THE ISSUE. THERE IS NO INDICATION THAT THE ARCHITECT I2000SR ANALYZER OR THE WASTE PUMP CONTRIBUTED TO OR CAUSED THE ISSUE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. BASED ON THE AVAILABLE INFORMATION, A CAUSE FOR THE BROKEN CONNECTOR ON THE EXTERNAL WASTE PUMP WAS NOT IDENTIFIED. HOWEVER, LABELING IN THE ARCHITECT SYSTEM OPERATION MANUAL IS ADEQUATE WITH REGARD TO PRECAUTIONS, HAZARDS, USE OF PERSONAL PROTECTIVE EQUIPMENT, AND TROUBLESHOOTING THE CUSTOMER'S ISSUE. THE EVALUATION DID NOT IDENTIFY A DEFICIENCY. | 0
2114329 | THE CUSTOMER'S CELL-DYN SHEAR VALVE DRIVE ASSEMBLY WAS REPLACED PER (B)(4). NO IMPACT TO PATIENT RESULTS, PATIENT MANAGEMENT OR USER SAFETY WAS REPORTED. (B)(4). AN EXPANDED INVESTIGATION WAS CONDUCTED TO EVALUATE THIS ISSUE. A PRODUCT CORRECTION LETTER AND CUSTOMER REPLY FORM WILL BE SENT TO ALL ACTIVE CELL-DYN SYSTEM CUSTOMERS WHO RECEIVED THE AFFECTED PART NUMBERS. THE PRODUCT CORRECTION LETTER WILL INSTRUCT CUSTOMERS TO RETURN THE CUSTOMER REPLY FORM TO ABBOTT ACKNOWLEDGING THE RECEIPT AND UNDERSTANDING THE ISSUE OR REQUEST ASSISTANCE. THE NON-CONFORMING PARTS WILL BE REPLACED IN THE FIELD THROUGH A FOUR-MONTH><(B)(4). THE AWARE DATE LISTED IN THE INITIAL MDR AS (B)(4) 2011 WAS INCORRECT. THE CORRECT AWARE DATE IS (B)(4) 2011. | 0
2114405 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BCI) STATING THAT THE WASTE FITTING ON THE BACK OF THE COULTER AC-T DIFF ANALYZER BROKE CAUSING A LEAK. THE OPERATOR WAS WEARING ALL APPROPRIATE PROPER PROTECTIVE EQUIPMENT (PPE) AND THERE WAS NO EXPOSURE TO MUCUS MEMBRANES OR OPEN WOUNDS. THERE WAS NO DEATH, INJURY, CHANGE TO PATIENT TREATMENT AND MEDICAL ATTENTION WAS NOT SOUGHT FOR THIS EVENT. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND REPLACED THE BROKEN WASTE FITTING. THE ROOT CAUSE FOR THE LEAK WAS A BROKEN WASTE FITTING. PER LABELING, RISK OF PERSONAL INJURY OR CONTAMINATION. IF YOU DO NOT PROPERLY SHIELD YOURSELF WHILE USING OR SERVICING THE INSTRUMENT, YOU MAY BECOME INJURED OR CONTAMINATED. TO PREVENT POSSIBLE INJURY OR BIOLOGICAL CONTAMINATION, YOU MUST WEAR PROPER LABORATORY ATTIRE, INCLUDING GLOVES, A LABORATORY COAT, AND EYE PROTECTION. | 0
2114557 | THE CUSTOMER STATES THAT FALSELY DECREASED RESULTS FOR ONE PATIENT WERE GENERATED ON THE AXSYM ANALYZER FOR BOTH THE TOTAL PSA AND FREE PSA ASSAYS. THE CUSTOMER PROVIDED THE FOLLOWING INFORMATION: FREE PSA: 0.00 NG/ML, TOTAL PSA: 0.00 NG/ML, DATE: (B)(6) 2011; 0.610, 1.262, (B)(6) 2011. THE INITIAL RESULTS OF 0.00 NG/ML FOR BOTH ASSAYS WERE REPORTED FROM THE LAB. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. ON THE FOLLOW-UP (#1) MEDWATCH 3500A REPORT, WAS UNINTENTIONALLY LEFT BLANK. THE DATE ENTERED SHOULD HAVE BEEN (B)(4) 2011 .><(B)(4). AN INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. CONCOMITANT PRODUCTS: AXS TOTAL PSA LN: 7K53-20 LOT: 93246LF01 EXPIRES: 26 SEP 2011. EXPIRATION DATE OF FREE PSA LOT 93233LF00: 01 OCT 2011.><AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) VISITED THE CUSTOMER SITE AND REPLACED THE FMI PUMP AND THE PTFE TUBING SETS. THE FSR VERIFIED ACCEPTABLE AXSYM ANALYZER PERFORMANCE AFTER SERVICE WAS COMPLETED. THE AXSYM SYSTEM OPERATION MANUAL (LIST NO.09A26-06) AND THE TOTAL PSA PACKAGE INSERT (49-3233/R3) CONTAIN INFORMATION TO ADDRESS THE CUSTOMER'S CURRENT ISSUE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. BASED ON THE CURRENT AVAILABLE INFORMATION AND THIS EVALUATION, NO DEFICIENCY WAS IDENTIFIED FOR THE AXSYM ANALYZER LIST NO. 07A83-01 RELATED TO THE ISSUE CURRENTLY UNDER EVALUATION. REVIEW OF COMPLAINT TRACKING AND TRENDING; FIELD SERVICE INTERVENTION. TUBING. | 0
2116813 | THE CUSTOMER REPORTED THAT AT 12 MINS INTO AN RBC DONATION RUN, THE DONOR COMPLAINED OF CHEST DISCOMFORT, NUMBNESS IN HIS LEFT ARM. THAT IS WHEN THEY NOTICED THAT THE ACD-A AND SALINE LINES WERE SWITCHED. THE DONOR RECEIVED 155ML OF ACD-A AS REPLACEMENT FLUID AND 61ML OF SALINE. NO MED INTERVENTION OCCURRED. THE DONOR RECOVERED SOON AFTER. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR INJURY FROM OVER ANTI-COAGULATION. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2116814 | A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER REPORTED AN INFANT APNEA MONITOR DID NOT ALARM DURING AN APNEIC AND CYANOTIC PATIENT EVENT. THE EXACT DATE AND TIME OF THE EVENT IS UNKNOWN; HOWEVER, THE DME STATED THAT IT WAS "BELIEVED TO HAVE OCCURRED ON (B)(6) 2011." IT WAS REPORTED THAT THE CAREGIVER OF THE INFANT PERFORMED CARDIOPULMONARY RESUSCITATION (CPR) AND THE INFANT WAS SUBSEQUENTLY TAKEN TO THE HOSPITAL FOR OBSERVATION. NO DEATH OR SERIOUS INJURY RESULTED. THE MANUFACTURER RECEIVED THE DEVICE FOR EVALUATION AND DID NOT CONFIRM THE CUSTOMER'S COMPLAINT OF THE UNIT NOT ALARMING. THE APNEA MONITOR WAS VISUALLY EXAMINED AND TESTED BY THE MANUFACTURER USING A SIMULATOR IN ACCORDANCE WITH THE SMARTMONITOR 2 CHECKOUT PROCEDURE MANUAL ((B)(4)). THE APNEA MONITOR DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS AND PASSED ALL REQUIRED TESTING. THE SMARTMONITOR 2 WAS RETURNED WITH PRESCRIBED SETTINGS OF A 20 SECOND DELAY BEFORE RECORDING APNEAS AND A 20 SECOND DELAY BEFORE ANNUNCIATING AND RECORDING AN ALARM FOR AN APNEA CONDITION. THE SMARTMONITOR 2 WAS PROGRAMMED WITH PARAMETERS THAT WERE APPROPRIATE FOR RECORDING AND ALARMING FOR RESPIRATION AND HEART RATE EVENTS DURING INFANT MONITORING. THE APNEA MONITOR'S MEMORY DATA WAS DOWNLOADED AND ANALYZED BY TRAINED ASSOCIATES. THE DOWNLOADED MEMORY REVEALED THE APNEA MONITOR WAS IN USE FROM (B)(6) 2011. DURING THAT TIME, THERE WERE 52 RECORDED PATIENT EVENTS. OF THE 52 RECORDED PATIENT EVENTS, 4 PATIENT EVENTS WERE RECORDED ON THE DATE OF THE ALLEGED INCIDENT. ALL RECORDED PATIENT EVENTS WERE ASSOCIATED WITH AUDIBLE AND VISUAL ALARMS. THE SMARTMONITOR 2 DEVICE IS NOT INTENDED TO BE USED TO MONITOR PATIENTS FOR CYANOSIS AND HAS NO CAPABILITY TO DO SO. THE SMARTMONITOR 2 PARENTS' GUIDE ((B)(4)) STATES IN THE INDICATIONS FOR USE: "THE SMARTMONITOR 2 IS INTENDED FOR USE IN CONTINUOUS MONITORING OF HEART RATE AND RESPIRATION OF INFANT PATIENTS IN A HOME, HOSPITAL OR PORTABLE ENVIRONMENT. ITS PRIMARY FUNCTION IS DETECTION OF CENTRAL APNEA. ITS SECONDARY FUNCTION IS MEASUREMENT OF HEART RATE." THE SMARTMONITOR 2 IS NOT INTENDED TO PREVENT LOSS OF BREATHING OR HEART ACTIVITY. THE SMARTMONITOR 2 PARENTS' GUIDE ((B)(4)) FURTHER STATES: "THE SMARTMONITOR 2 IS A MONITORING DEVICE ONLY. IT DOES NOT PREVENT THE LOSS OF BREATHING OR HEART ACTIVITY, NOR WILL IT RESTORE BREATHING OR HEART ACTIVITY. IT WILL NOT PREVENT DEATH. ANYONE USING THE SMARTMONITOR 2 TO MONITOR AN INFANT SHOULD BE TRAINED IN CURRENT INFANT CARDIOPULMONARY RESUSCITATION (CPR), WHICH IS A PROPER WAY TO RESTORE BREATHING AND HEART ACTIVITY." THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY, AND THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE ALLEGATION OF ALARM FAILURE DURING A PATIENT EVENT. THE MONITOR PASSED ALL REQUIRED TESTING AND DETECTED AND ALARMED APPROPRIATELY FOR SIMULATED EVENTS. BASED ON ALL AVAILABLE INFORMATION, THE MANUFACTURER CONCLUDES THAT THE DEVICE DOES FUNCTION TO SPECIFICATION AND THAT NO FURTHER ACTION IS APPROPRIATE. | 1
2116817 | THE CUSTOMER REPORTED THAT THE DONOR WAS CONNECTED TO THE MACHINE WHEN THE "DO NOT CONNECT DONOR" ICON WAS PRESENT. THE DONOR HAD BEEN CONNECTED AFTER AC PRIME WAS STARTED. THE DONOR WAS DISCONNECTED WHEN THEY NOTICED THE ICON AND THE PROCEDURE WAS ENDED. THE DONOR DID NOT RECEIVE ANY FLUIDS. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR INJURY. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2116824 | THE CUSTOMER REPORTED A NEEDLE STICK INJURY THAT OCCURRED WITH ONE OF THEIR APHERESIS OPERATORS. THE OPERATOR HAD REMOVED THE NEEDLE FROM THE DONOR'S ARM, SLID THE ANTINEEDLESTICK DEVICE OVER THE NEEDLE AND WAS REMOVING THE NEEDLE FROM THE DISPOSABLE SET TO DISCARD IT WHEN SHE INCURRED THE NEEDLE STICK. THE CUSTOMER STATED THAT THE AREA OF THE INJURY WAS CLEANED THOROUGHLY AND THEY FOLLOWED THEIR BLOOD-BORNE PATHOGEN EXPOSURE POLICY. THE EMPLOYEE WILL BE GOING THROUGH FOLLOW-UP TESTING AND COUNSELING. THE CUSTOMER WAS NOT WILLING TO PROVIDE AGE OR WEIGHT INFORMATION. THE DISPOSABLE SET IS NOT AVAILABLE FOR INVESTIGATION. (B)(4). INVESTIGATION: THE CUSTOMER STATED THAT THEIR CURRENT PRACTICE IS TO REMOVE THE NEEDLE FROM THE DONOR'S ARM, LINE UP THE WINGS ON THE BUTTERFLY NEEDLE WITH THE SAFETY DEVICE, AND THEN SLIDE THE DEVICE UP OVER THE NEEDLE. THE CUSTOMER STATED THAT THIS PROCEDURE IS USED SO THEY CAN BE SURE THAT THE NEEDLE HAS LOCKED INTO PLACE PROPERLY. THE DEVICE HISTORY RECORDS WERE REVIEWED WITH NO RELEVANT ISSUES FOUND. ROOT CAUSE: BASED UPON THE CUSTOMER'S DESCRIPTION OF THEIR PROCESS, THE PROCEDURE FOR THE USE OF THE ANTI-NEEDLE STICK PROTECTION IS DIFFERENT THAN THE TRIMA OPERATOR'S MANUAL FOR THE USE OF THE ANTI-NEEDLE STICK PROTECTION. CONCLUSION: THE DESIGN OF THE ANTI-NEEDLE STICK PROTECTION REQUIRES THE USE OF THE RECOMMENDED PROCEDURE TO ENSURE THAT THE OPERATOR'S HANDS DO NOT COME INTO CONTACT WITH THE NEEDLE. THE NEEDLE MAY BE ASKEW AS IT IS BEING RETRACTED, BUT THE NEEDLE WILL LOCK INTO PLACE IF THE COMPLETE PROCEDURE IS FOLLOWED. USE OF THE DEVICE OUTSIDE OF THIS PROCEDURE HAS NOT BEEN VALIDATED BY CARIDIANBCT AND MAY LEAD TO INJURY. | 0
2120121 | CUSTOMER REPORTED AN RH DISCREPANCY WHEN TESTING A PATIENT SAMPLE WITH THE WEAK D ASSAY ON THE ECHO. THE PATIENT IS HISTORICALLY A NEGATIVE. THE ECHO GENERATED NEGATIVE RESULTS FOR ANTI-D (MONOCLONAL BLEND) SERIES 4 AND ANTI-D (MONOCLONAL BLEND) SERIES 5. THE CUSTOMER PERFORMED WEAK D ON ECHO AND POSITIVE RESULT WAS GENERATED (4+). THE ECHO GENERATED NEGATIVE RESULTS FOR DAT. THE SAMPLE WAS TESTED BY MANUAL TUBE METHOD AND NEGATIVE RESULTS WERE GENERATED FOR DAT AND WEAK D. REVIEW OF RESULT IMAGES IDENTIFIED THAT A BRIGHT AREA SEEN IN THE UPPER RIGHT CORNER ON ALL OF THE IMAGES. A SERVICE CALL WAS MADE. INVESTIGATION OF CUSTOMER COMPLAINT REVEALED ISSUES WITH CAMERA READER ASSEMBLY. REPLACED CAMERA READER ASSEMBLY. PERFORMED ALL CAMERA ALIGNMENTS AND CALIBRATIONS. PERFORMED QC TESTING WITH ACCEPTABLE RESULTS. CUSTOMER TESTED 4 SAMPLES WITH RFXGROUPSCREEN ASSAY WITH ACCEPTABLE RESULTS. SYSTEM IS OPERATING WITHIN SPECIFICATIONS. | 0
2128797 | IT WAS REPORTED THAT ON MULTIPLE OCCASIONS, DATA ENTERED FROM NETACESS BY A USER FOR A PATIENT HAS BEEN FOUND TO BE SAVED IN THE RECORD OF A DIFFERENT PATIENT. THERE WAS NO PATIENT INJURY REPORTED. WHEN DRAEGER INITIALLY RECEIVED THIS REPORT, IT WAS DETERMINED TO BE NOT REPORTABLE BECAUSE INNOVIAN IS USED FOR DATA COLLECTION, AND REAL-TIME PATIENT MONITORING IS NOT IMPACTED. HOWEVER, UPON FURTHER INVESTIGATION, THERE IS CONCERN ABOUT MIXING PATIENT DATA. THEREFORE, IT WAS DETERMINED THAT THIS COMPLAINT IS REPORTABLE. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2132747 | THE CUSTOMER REPORTED THAT DURING A PROCEDURE, THE NEEDLE RETRACTOR WAS NOT FUNCTIONAL, AND ALLEGED A RISK OF BLOOD EXPOSURE. THE PT INFO IS NOT AVAILABLE AT THIS TIME. CARIDIANBCT IS AWAITING THE RETURN OF THE DISPOSABLE SET. THIS REPORT IS BEING FILED DUE TO AN ALLEGATION OF A SAFETY RISK. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2135859 | CUSTOMER REPORTED AN ISSUE WITH THE ABD + REV A1B REACTIONS ON THIS TANGO THAT THE CUSTOMER DESCRIBES AS "CARRYOVER." THE CUSTOMER STATED THAT THEY WERE NOT SEEING ANY ISSUES ON THE SOLIDSCREENS THAT SHE IS AWARE OF AS ALL OF THE ANTIBODY SCREENS THAT THEY HAVE RUN HAD BEEN NEGATIVE. A YEARLY PREVENTIVE MAINTENANCE WAS PERFORMED ON (B)(6) 2011. ALL QUALITY CONTROLS RUN AFTER PREVENTIVE MAINTENANCE PASSED SUCCESSFUL. SINCE NO OBVIOUS INSTRUMENT MALFUNCTION COULD BE DETECTED BUT THE QC RUN AFTER THE PREVENTIVE MAINTENANCE WAS SUCCESSFUL, THE PROBLEM WAS EITHER LINKED TO BUILD UP ON THE LEFT PIPETTOR PROBE (WHICH MAY BE DUE TO INSUFFICIENT DILIGENCE REGARDING THE DAILY MAINTENANCE) OR THE DRAIN TUBINGS, WHICH THE FIELD SVC ENGINEER ALSO PROPOSED TO BE POSSIBLY CAUSATIVE. | 0
2142981 | THE PATIENT HAD A KNOWN CONDITION WHICH DISCOLORS THE PLASMA. THE CUSTOMER WANTED TO KNOW IF HE CAN DISABLE THE RED CELL DETECTOR FROM THE ALARM SCREEN. IT WAS DETERMINED THAT THE CUSTOMER WOULD HAVE THE OPTION. THE DISPOSABLE WAS DISCARDED AND WILL NOT BE RETURNED. THE PATIENT INFORMATION IS UNAVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO HEMOLYSIS WHICH MAY BE CAUSED BY THE PATIENT'S CONDITION. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2145735 | A CUSTOMER REPORTED AN RH MISTYPE OF FIVE DONOR SAMPLES THAT WERE TESTED ON GALILEO INSTRUMENT #(B)(4). ON (B)(4) 2011, AN IMMUCOR INSTRUMENT SERVICE ENGINEER PERFORMED THE GALILEO UNEXPECTED REACTION CHECK LIST. THE REAGENT PROBE WAS REPLACED DUE TO THE TEFLON TIP BEING SLIGHTLY WORN. THE CUSTOMER PERFORMED QC AND TESTED KNOWN RH NEGATIVE AND POSITIVE SAMPLES TO VERIFY INSTRUMENT OPERATION. THE EXPECTED RESULTS WERE OBTAINED. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2149495 | THE CUSTOMER REPORTED THAT AT THE END OF A STEM CELL COLLECTION PROCEDURE, THE SAMPLE LINE WAS WET TO THE TOUCH. A LEAK WAS THEN IDENTIFIED AT THE JOINT AREA. MEDICAL STAFF WAS NOTIFIED PER THE CUSTOMER'S PROTOCOL, BUT NO MEDICAL INTERVENTION WAS NECESSARY. CUSTOMER'S INTERNAL DONOR #: (B)(6). THIS REPORT IS BEING FILED DUE TO THE CUSTOMER NOTIFYING ADD'L MEDICAL STAFF AND THE (B)(4). (B)(4). INVESTIGATION: THE FULL SET WAS RETURNED FOR INVESTIGATION. DRIED BLOOD WAS NOTED AT THE SLIPFIT LUER ON THE COLLECT LINE. THE LINE WAS HEAT-SEALED ABOVE THE LUER AND THE BLUE CLAMP WAS CLOSED BELOW THE LUER. THE LEAK AT THE SLIPFIT LUER IS POSSIBLY A RESULT OF A PRESSURE BUILD UP DURING HEAT-SEALING OF THE TUBING BECAUSE THE BLUE SIDE CLAMP WAS ALSO CLOSED BELOW THE SLIPFIT LUER, OR COULD HAVE OCCURRED AS A RESULT OF THE LINE BEING STRIPPED AT THE END OF THE COLLECTION. THE DEVICE HISTORY RECORDS WERE REVIEWED. THERE WERE NO ISSUES FOUND THAT WERE RELEVANT TO THIS EVENT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2149936 | ON (B)(6) 2011 AN ANEMIC PATIENT WITH A HEMOGLOBIN (HGB) RESULT OF 7.8G/DL WAS TREATED WITH TWO UNITS OF PACKED CELLS. PER CUSTOMER, THIS RESULT WAS DEEMED CORRECT HENCE THE REASON FOR TREATMENT. THE PATIENT SAMPLE WAS TESTED ON THE COULTER LH500 HEMATOLOGY ANALYZER. FOLLOWING THE TREATMENT, THE PATIENT'S HGB WAS MONITORED RESPECTIVELY. ON (B)(6) 2011, THE CUSTOMER CALLED BECKMAN COULTER IN., (BEC) TO QUESTION THE HGB RESULTS OBTAINED ON (B)(6) 2011 FROM THEIR LH500 ANALYZER. THE RESULTS OBTAINED FROM THE LH500 ANALYZER WERE VERIFIED BY RUNNING THE SAMPLES ON A COMPETITOR'S ANALYZER AND SIMILAR RESULTS WERE OBTAINED. PER CUSTOMER, THESE RESULTS WERE REPORTED OUT OF THE LAB; HOWEVER, THERE WAS NO CHANGE TO PATIENT TREATMENT. SEVERAL ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL DATA INCLUDING RAW DATA FILES AND PRINTOUTS, BUT THEY WERE NOT PROVIDED. THE PATIENT WAS TREATED WITH TWO UNITS OF PACKED CELLS AFTER THE INITIAL SAMPLE WAS TESTED; HOWEVER, NO FURTHER CHANGE TO PATIENT TREATMENT WAS DOCUMENTED. QC WAS RUN BEFORE AND AFTER THE EVENT; HOWEVER, DATA WAS NOT SUPPLIED. THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS WITH RESPECT TO CONTROLS AND REPRODUCIBILITY. A FIELD SERVICE ENGINEER (FSE) CHECKED THAT THE REAGENTS MEET SPECIFICATIONS, WERE NOT EXPIRED AND ONBOARD PRIOR TO THIS SAMPLE BEING ANALYZED. THE FSE VERIFIED THAT THE PROPER AMOUNT OF SAMPLE VS. DILUENT WAS DISPENSED AND REPRODUCED. THE FSE PERFORMED THREE DIFFERENT REPRODUCIBILITY STUDIES AND ALL WERE WITHIN SPECIFICATIONS. THE FSE VERIFIED ASPIRATION AND HGB LAMP VOLTAGE AND STABILITY WERE FINE AND HAD A GOOD BACKGROUND COUNT. THE INSTRUMENT WAS FOUND TO BE PERFORMING AS EXPECTED. A CLEAR ROOT CAUSE FOR THIS EVENT IS UNKNOWN. | 0
2153892 | A CUSTOMER CONTACTED BECKMAN COULTER INC., (BEC) AND REPORTED A FLUID LEAK IN COULTER LH 750 HEMATOLOGY ANALYZER THAT APPEARED TO BE ISOTON. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES, GOWN AND SAFETY GLASSES. THERE WAS NO CONTACT WITH OPEN WOUNDS OR MUCOUS MEMBRANES. NO DEATH OR INJURY OCCURRED AND NO ONE REQUIRED MEDICAL ATTENTION. A FIELD SERVICE ENGINEER (FSE) FOUND A LEAK AT PINCH VALVE (PV25) WHICH IS AROUND THE AREA OF THE BLOOD SAMPLING VALVE (BSV). THE STRIP TUBING THAT PASSES THROUGH PV25 WAS CUT. THE FSE REPLACED THE TUBING FROM CUSTOMER STOCK. THE ROOT CAUSE WAS DUE TO A CUT ON THE STRIP TUBING THAT PASSES THROUGH PV25. | 0
2159563 | OPERATORS AT THE CUSTOMER SITE HAVE REPORTED TO THEIR SUPERVISOR THAT THE RETURN PUMP "REVS" DURING THE FIRST RETURN. THEY HAVE REPORTED THAT A FEW HEMATOMAS HAVE BEEN ASSOCIATED WITH THE PUMP REVS. THE CUSTOMER DESCRIBED THE HEMATOMAS AS OF THE "NORMAL" VERY MINOR VARIETY. PATIENT INFORMATION IS NOT AVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION OCCURRED. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2160309 | GENEXPERT INFINITY XPERTISE SOFTWARE, VERSION 4.1A, ALLOWS USER TO ENTER TEST ORDERS FASTER THAN THE REQUEST FOR LAB INFO SYS (LIS) ORDERS CAN BE PROCESSED. IN COMPLAINT (B)(4), A CUSTOMER (B)(6) REPORTED VISUAL OBSERVATION OF MISMATCHED PT ID AND SAMPLE RESULTS. UPON FURTHER INVESTIGATION, CEPHEID FIELD SUPPORT DETERMINED THAT WHEN BATCH LOADING CARTRIDGES ONTO THE GENEXPERT INFINITY-48, IF INPUT DATA (PATIENT ID, SAMPLE ID, AND CARTRIDGE ID) ARE ADDED VERY QUICKLY, THE SYSTEM MAY NOT PAIR THE PT SAMPLE WITH CORRECT LIS TEST ORDER. THE INCORRECT PAIRING OF LIS ORDERED PT ID TO A SAMPLE ID CAN OCCUR IF SAMPLES ARE PROCESSED FASTER THAN LIS ORDERS ARE DOWNLOADED. FOR EXAMPLE, WHEN SAMPLE ID #1 IS SCANNED, THE LIS ORDER DOWNLOAD STARTS, BUT THE TEST CAN BE STARTED BEFORE THE DOWNLOAD HAS COMPLETED. IF SAMPLE ID #2 IS SCANNED BEFORE THE FIRST LIS ORDER DOWNLOAD COMPLETES, THE SAMPLE ID #2 MAY BE INCORRECTLY PAIRED TO THE FIRST LIS ORDERED PT ID. THERE WAS NO RISK OF INJURY TO PT DURING THE ORIGINATING EVENT AT (B)(6) (REPORTING SITE, (B)(4)) BECAUSE THE SITE COMPARES ALL TEST RESULTS TO TEST ORDERS BY PT NAME, SAMPLE ID TEST ORDER NUMBER AND CARTRIDGE NUMBER. THE TRANSPOSITION OF DATA WAS DETECTED BEFORE PT RESULTS WERE REPORTED. THE EXACT CONFIGURATION THAT ALLOWED THE EVENT TO OCCUR WAS IDENTIFIED (INFINITY 48 SYSTEMS WITH GENEXPERT 4.1A USING HOST QUERY TEST ORDER DOWNLOAD CONFIGURATION AND BATCH LOADING). NO OTHER SITES USED THIS EXACT CONFIGURATION AND THEREFORE, COULD NOT HAVE EXPERIENCED THE MALFUNCTION THAT MAY HAVE RESULTED IN A DISCREPANT PT RESULT REPORT. ALL UNITED STATES CONSIGNEES WERE CONTACTED BY PHONE 04/01/2011. ALL UNITED STATES CONSIGNEES RECEIVED A COPY OF EMAIL 04/01/2011. ALL XUS CUSTOMERS WERE NOTIFIED BY 04/08/2011. EACH CONSIGNEE WERE SENT A HARD COPY CORRECTIVE ACTION NOTIFICATION BY (B)(4) BY 04/27/2011. SUMMARY OF EVAL: EVAL OF THE (B)(6) DISCREPANT SPECIMEN RESULTS INCLUDED: REVIEW OF CUSTOMER COMPLAINTS FOR SIMILAR COMPLAINTS - NONE FOUND. REVIEW OF UNRESOLVED ANOMALY LIST FOR SIMILAR ISSUE - NONE FOUND. TRACING CHRONOLOGICAL SEQUENCE OF EVENTS OF THE DATA HAND-OFF POINTS FROM LIS TO INFINITY FOR TEST ORDER DOWNLOADS, RUNNING CARTRIDGE, AND UPLOAD OF TEST RESULTS FROM INFINITY TO LIS. WALKTHROUGH OF DATABASE RECORDS AND THEIR LINKAGES. COOL WALKTHROUGHS FOR BOTH LIS TEST ORDER DOWNLOAD PATH AND PATH FOR TEST RESULT UPLOAD TO LIS. ATTEMPTS TO REPRODUCE THE EVENTS BY RUNNING (B)(6) SPECIFIC LIS CONFIGURATION USING BATCH LOADING OF CARTRIDGES. BY EVALUATING THE DATA HAND-OFF POINTS AND SEEING HOW THE DATABASE RECORDS WERE LINKING THE DATA WITHIN THE DATABASE, THE SCENARIO WAS ABLE TO BE REPRODUCED USING THE SPECIFIC LIS CONFIGURATION. THE COMBINATION OF USING HOST QUERY LIS CONFIGURATION WHERE THE SCANNING OF THE SAMPLE ID REQUESTS INFO FROM THE LIS AND THE RAPID LOADING OF CARTRIDGES BEFORE THE LIS TEST ORDER HAD DOWNLOADED WERE DETERMINED TO BE THE ROOT CAUSE. CORRECTIVE ACTION HAS BEEN TAKEN WITH CUSTOMERS AND PREVENTIVE ACTION WILL INCLUDE NEW SOFTWARE RELEASE. NO PTS WERE MISDIAGNOSED DUE TO THIS MALFUNCTION. ONLY ONE CONSIGNEE SITE WAS CONFIGURED IN SUCH A WAY AS TO CAUSE THE MALFUNCTION. BECAUSE INFINITY RESULTS WERE ALWAYS COMPARED TO LIS TEST RESULTS MANUALLY, NO PT DATA WERE AFFECTED AND THE SITE DID NOT REPORT ANY DISCREPANT RESULTS. | 1
2160956 | A CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT A BELT ON HEMATOLOGY WAS NOT TURNING AND THAT THERE WAS A RUBBER SMELL COMING FROM THE POWER PROCESSOR. NO EFFECT TO PATIENT OR USER WAS REPORTED IN CONNECTION WITH THIS EVENT. ON (B)(4) 2011, BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED FOR THE EVENT. THE FSE OBSERVED THE INSTRUMENT FOR SIX HOURS BUT COULD NOT REPRODUCE THE PROBLEM. INSTRUMENT WAS OPERATIONAL. NO FURTHER COMPLAINT WAS RECEIVED AS OF (B)(4) 2011. | 0
2163687 | A DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER REPORTED THAT A (B)(6) APNEA MONITOR FAILED TO ALARM DURING AN APNEIC AND CYANOTIC PATIENT EVENT. THE DATE OF THE ALLEGED EVENT WAS ON (B)(6) 2011, AND THE EXACT TIME IS UNKNOWN. THE DME STATED THAT THE CAREGIVER REPORTED THAT THE "BABY WAS TURNING BLUE IN THE CRIB AFTER THE MOM TOOK THE BABY OFF OF OXYGEN." THE FIRE DEPARTMENT WAS CALLED AND ARRIVED TO TREAT THE INFANT. THE BABY DID NOT STOP BREATHING BUT IT WAS NOTICED THAT THE BABY EXPERIENCED OXYGEN DESATURATION. THE BABY WAS SUBSEQUENTLY TAKEN TO THE HOSPITAL. THE DME STATED THAT THE STATUS OF THE BABY IS NOT KNOWN AT THIS TIME. THE DEVICE AND THE PATIENT DOWNLOAD HAVE NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION AFTER REPEATED ATTEMPTS BY THE MANUFACTURER TO OBTAIN THE DEVICE FROM THE DME. THE SMARTMONITOR 2 DEVICE IS NOT INTENDED TO BE USED TO MONITOR PATIENTS FOR CYANOSIS AND HAS NO CAPABILITY TO DO SO. THE SMARTMONITOR 2 PARENTS' GUIDE (PN (B)(4)) STATES IN THE INDICATIONS FOR USE: "THE SMARTMONITOR 2 IS INTENDED FOR USE IN CONTINUOUS MONITORING OF HEART RATE AND RESPIRATION OF INFANT PATIENTS IN A HOME, HOSPITAL OR PORTABLE ENVIRONMENT. ITS PRIMARY FUNCTION IS DETECTION OF CENTRAL APNEA. ITS SECONDARY FUNCTION IS MEASUREMENT OF HEART RATE." A FOLLOW-UP REPORT WILL BE FILED WHEN THE DEVICE AND PATIENT DOWNLOAD HAS BEEN RETURNED AND AN INVESTIGATION IS COMPLETED BY THE MANUFACTURER. | 1
2163870 | DISCORDANT HIGH ADVIA CENTAUR TROPONIN RESULTS WERE OBTAINED ON MULTIPLE PATIENT SAMPLES. THE LABORATORY NOTICED THAT SOME OF THE TROPONIN RESULTS WERE RUNNING HIGH, SO THE SAMPLES WERE REPEATED. THERE ARE NO KNOWN REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO THE DISCORDANT TROPONIN RESULTS. A SIEMENS FSE (FIELD SERVICE ENGINEER) WAS DISPATCHED TO THE CUSTOMER SITE TO EVALUATE THE INSTRUMENT. AFTER ANALYSIS OF THE INSTRUMENT AND INSTRUMENT DATA, THE FSE REPLACED ASPIRATE TUBING THAT HAD BROKEN LOOSE FROM A CONNECTOR HUB. THE CUSTOMER THEN CALIBRATED AND RAN CONTROLS. ALL CONTROLS WERE WITHIN RANGE. THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. NO FURTHER EVALUATION OF THE DEVICE IS REQUIRED. | 0
2168934 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL COUNT IN THE COLLECTED BLOOD PRODUCT. THE PATIENT INFORMATION IS UNAVAILABLE AT THIS TIME. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. CARIDIANBCT IS AWAITING FURTHER INFORMATION FROM THE CUSTOMER ABOUT THIS INCIDENT. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2168943 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL COUNT IN THE COLLECTED BLOOD PRODUCT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. THE DONOR INFORMATION IS NOT AVAILABLE AT THIS TIME. CARIDIANBCT IS AWAITING FURTHER INFORMATION FROM THE CUSTOMER ABOUT THIS INCIDENT. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2178107 | CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS ON CAPTURE-R READY SCREEN (CRRS) AND CAPTURE-R READY ID (CRRID) ON THE ECHO WITH ONE PATIENT SAMPLE WITH A PREVIOUSLY KNOWN ANTI-S. REVIEW OF INSTRUMENT REACTIONS: ALL SAMPLE WELLS REACTED NEGATIVE FOR BOTH CRRS AND CRRID. THE CELL REACTIONS APPEARED NEGATIVE FOR THE CELLS IN QUESTION. POSITIVE AND NEGATIVE CONTROL WELLS ON CRRID AND CRRS REACTED AS EXPECTED. SCII: S+S- , SCIII :S+S- ID CELLS 1, 2, 11: S+S-, ID CELLS 5, 6, 14: S+S+ A SERVICE CALL WAS MADE: INSTALLED NEW READER AND CONFIGURED. INSTALLED NEW WASHER MANIFOLD AND RAN LONG SYSTEM PRIME. QC AND DAILY MAINTENANCE PERFORMED. TESTED 4 SAMPLES WITH 2 KNOW POSITIVE AND 2 NEGATIVE SCREEN SAMPLES. RESULTS WERE AS EXPECTED, HOWEVER, THERE WAS A NTD ON ONE OF THE SCREENS. LOOSENED THE ROBOT Y BELT AND RE-TESTED THE SAMPLE WITH EXPECTED RESULTS THE ISSUE WAS RESOLVED BY REPLACING THE CAMERA READER. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2178564 | THE CUSTOMER CONTACTED BECKMAN COULTER, INC. (BCI) TO REPORT THAT INCORRECT PATIENT INFORMATION FOR A SINGLE PATIENT SPECIMEN WAS UPLOADED FROM A UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM TO A MIDDLEWARE DATA MANAGER. THE INCORRECT PATIENT INFORMATION CONSISTED OF AN INCORRECT PATIENT IDENTIFICATION AND INCORRECT PATIENT NAME FOR A SINGLE PATIENT. NO MISIDENTIFICATION OCCURRED ON THE SPECIMEN IDENTIFICATION (SPECIMEN ID #(B)(4)). THE INCORRECT PATIENT INFORMATION WAS IDENTIFIED AFTER UPLOAD OF THE PATIENT RESULTS FROM THE UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM TO THE MIDDLEWARE DATA MANAGER AND WAS DETECTED DURING MANUAL REVIEW OF THE PATIENT RESULTS. NO ERRONEOUS RESULTS WERE REPORTED OUT OF THE LABORATORY. THERE WAS NO REPORT OF PATIENT INJURY REQUIRING MEDICAL INTERVENTION OR CHANGE TO PATIENT TREATMENT ATTRIBUTED OR CONNECTED TO THIS EVENT. THE SPECIMEN WAS RE-ASSAYED ON ANOTHER UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM IN THE LABORATORY. THE CORRECT PATIENT INFORMATION WAS UPLOADED. A DEFINITIVE ROOT CAUSE HAS NOT YET BEEN DETERMINED FOR THIS EVENT. METHOD: MANUAL VERIFICATION OF PATIENT INFORMATION.RESULT: ERRONEOUS PATIENT INFORMATION.CONCLUSION: A DEFINITIVE ROOT CAUSE HAS NOT YET BEEN DETERMINED FOR THIS EVENT. | 1
2184900 | THE CUSTOMER CONTACTED BECKMAN COULTER, INC. (BEC) TO REPORT THAT THE PROBE OF THEIR COULTER AC·T DIFF2 HEMATOLOGY ANALYZER HAD PIERCED THE BOTTOM OF A SPECIMEN TUBE CAUSING BLOOD TO LEAK OUT. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE - GLOVES) AT THE TIME OF THE EVENT. THERE WAS NO INJURY TO THE OPERATOR OR EXPOSURE TO OPEN WOUNDS OR MUCOUS MEMBRANES. THE OPERATOR DID NOT SEEK MEDICAL ATTENTION. THE CUSTOMER REVIEWED THE MATERIAL SAFETY DATA SHEET (MSDS). THE FACILITY DOES HAVE AN EXPOSURE CONTROL / RISK MANAGEMENT PLAN IN PLACE. THE CUSTOMER WAS INSTRUCTED ON HOW TO REMOVE THE PROBE THAT WAS STUCK IN THE SPECIMEN TUBE. AFTER THE ANALYZER PASSED ITS START-UP PROCEDURE, THE CUSTOMER RAN A BLANK SAMPLE IN CAP PIERCE MODE. THE PROBE FUNCTIONED NORMALLY AND DID NOT GET STUCK. ON-SITE SERVICE WAS NOT REQUIRED AS THE CUSTOMER CALLED BACK LATER AND REPORTED THAT THE ANALYZER WAS OPERATIONAL. THE CUSTOMER REPORTED THAT SHE BELIEVES THAT THE BOTTOM OF THE SPECIMEN TUBE WAS PIERCED BECAUSE THE SPECIMEN TUBE WAS NOT FULLY INSERTED INTO THE TUBE HOLDER. PATIENT RESULTS WERE NOT IMPACTED DUE TO THIS EVENT. THERE WAS NO IMPACT TO PATIENT TREATMENT ASSOCIATED WITH THIS EVENT. CONCLUSION: SPECIMEN TUBE SUSPECTED NOT TO BE FULLY INSERTED INTO TUBE HOLDER. BEC URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. | 0
2185594 | IT WAS REPORTED THAT IMAGES WERE NOT DISPLAYING PROPERLY. THE FIELD ENGINEER ARRIVED AND FOUND THE GRADIENT CABLES SWAPPED. THIS WAS THE ROOT CAUSE OF THE FLIPPED IMAGES. HE RETURNED THE CABLES TO THEIR CORRECT LOCATIONS AND THE IMAGES WERE FINE. THERE WAS NO MISDIAGNOSIS OR MISTREATMENT AS A RESULT OF THIS INCIDENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2190720 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL COUNT IN THE COLLECTED PLATELET PRODUCT. PT IDENTIFIER AND AGE ARE NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE TO BE RETURNED FOR EVAL. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE PLATELET COLLECTION DID NOT START WHEN THE PLATELET VALVE OPENED AT 12 MINUTES BUT CLOSER TO 20 MINUTES. THIS MAY INDICATE THAT PLATELETS WERE NOT AVAILABLE AS EXPECTED; THIS CAN BE ATTRIBUTED TO INACCURATE DONOR PLATELET PRE-COUNT ENTRY OR POSSIBLY DONOR RELATED. SIGNALS IN THE RUN DATA FILE DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE REPORTED ELEVATED WBC CONTENT IN THIS PROCEDURE. THERE ARE NO EVENTS IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. THIS REPORT WAS FILED BEYOND THE 30-DAY TIME FRAME DUE TO AN INTERNAL PROCESSING ISSUE THAT IS CURRENTLY BEING EVALUATED FOR APPROPRIATE CORRECTIVE ACTIONS. | 0
2192390 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A LEAK IN THE COULTER HMX AUTOLOADER ANALYZER INSTRUMENT'S WASH TRUCK. THE USER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF LAB COAT, GLOVES AND EYE PROTECTION AT THE TIME OF THE INCIDENT. NO INJURIES OCCURRED AND MEDICAL ATTENTION WAS NOT SOUGHT. THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. NO ONE WAS SPLASHED, SPRAYED, OR INJURED. PATIENT TREATMENT WAS NOT AFFECTED IN THIS EVENT. A BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(6) 2011 FOR THIS EVENT. THE FSE OBSERVED A LEAK CONSISTING OF DILUENT AND A SMALL AMOUNT OF BLOOD IN THE SECOND PORT VACUUM LINE OF THE RINSE BLOCK. THE FSE ALSO DISCOVERED DRIED UP CLENZ AROUND THE WHITE BLOOD COUNT (WBC) BATH APERTURE HOUSING WHICH THE FSE CLEANED. THE FSE DETERMINED THAT THE ROOT CAUSE OF THE LEAK WAS DUE TO A TRAPPED SMALL QUANTITY OF PARTICULATE MATTER IN THE RINSE BLOCK VACUUM LINE. THE FSE REPLACED THE VC6 (VACUUM CHAMBER THAT DISTRIBUTES VACUUM IN THE DILUTER) TO IMPROVE THE VACUUM DISTRIBUTION. REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES AND RESULTS MEET PUBLISHED PERFORMANCE SPECIFICATIONS. PER LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. BEC IDENTIFIER FOR THIS REPORT IS (B)(4). | 0
2192960 | IT WAS REPORTED THAT THE USER WAS CREATING A SINGLE POINT-TO-POINT MEASUREMENT, BUT WAS BEING PRESENTED WITH TWO MEASUREMENTS FOR THAT LINE. THE RESULTING MEASUREMENTS WERE DIFFERENT. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
2193597 | DURING A PROSTATECTOMY WHILE THE SURGEON WAS USING THE GYRUS G400 GENERATOR, HE HAD PROBLEMS WITH CUTTING AND COAG DURING THE PROCEDURE. HE USED ALTERNATIVE METHOD TO COMPLETE THE CASE, BUT EXTENDED THE TIME OF SURGERY BY ONE HOUR. THE PT WAS UNDER ANESTHESIA ABOUT ONE HOUR LONGER SO THE BLEEDING COULD BE STOPPED. ANALYSIS OF THE GENERATOR FOUND THAT THERE WAS AN INTERNAL FAILURE ASSOCIATED WITH POWER SUPPLY BOARD. THE BOARD WAS REPLACED, AND THE SOFTWARE WAS UPGRADED TO THE LATEST VERSION. THE GENERATOR WAS THEN INSPECTED AND TESTED AND FOUND TO OPERATE TO THE MFR'S SPECS. | 0
2193620 | THE CUSTOMER REPORTED THAT WHILE THE PANORAMA CENTRAL STATION WAS IN USE MONITORING PATIENTS ALONG WITH SPECTRUM MONITORS AND AMBULATORY TELEPACKS, THE WIRELESS TRANSCEIVER STOPPED FUNCTIONING, CAUSING THE CENTRAL STATION TO LOSE CONNECTIVITY TO THE PATIENT DEVICES. NO PATIENT INJURY WAS REPORTED. THE COMPANY SERVICE REPRESENTATIVE REPLACED THE POWER SUPPLY IN THE WIRELESS TRANSCEIVER. THE SYSTEM WAS TESTED TO FACTORY SPECIFICATIONS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. | 1
2201198 | ON (B)(6) 2011, (B)(6), USA REPORTED BEING UNABLE TO ACCESS IMAGES IN THE OPERATING ROOM VIA THE CARDIO WEB SERVER. THE SYSTEM RESUMED FUNCTIONING WITHOUT INTERVENTION FROM THE SITE OR (B)(4). IT IS ESTIMATED THAT THE INCIDENT MAY HAVE AFFECTED CV WEB USERS FOR APPROX 30-60 MINUTES. THERE WERE TWO SURGICAL PROCEDURES IN PROGRESS DURING THE INTERRUPTION. THE SITE HAS NOT PROVIDED ANY EXPLICIT DETAILS DESCRIBING THE SURGERIES BUT NO ADVERSE CLINICAL EVENT WAS REPORTED AS A RESULT OF THE INTERRUPTION. ADDITIONAL INFO PERTINENT TO EVAL (CONCLUSION): THE DEVICE IN QUESTION DID NOT FAIL. THE SITE WAS IN THE PROCESS OF MIGRATING DATA TO NEW SERVERS. AS A RESULT OF THE MIGRATION, THE MEDICAL DEVICE OPERATING REQUIREMENTS WERE COMPROMISED WHEN THE SITE RE-ROUTED THE NETWORK CONNECTION. THE SITE'S HARDWARE DID NOT INCLUDE ENOUGH PORTS TO MEET IMPAX CV REQUIREMENTS DURING THE MIGRATION. | 1
2210109 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT AFTER A DOUBLE PLATELET PRODUCT COLLECTION. DONOR UNIT #: (B)(4). NO MEDICAL INTERVENTION WAS REQUIRED. THE DISPOSABLE SET IS NOT BEING RETURNED FOR EVAL. THIS REPORT IS BEING FILED DUE TO ALLEGED DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE WAS REVIEWED. SIGNALS INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. THE MEASURED WBC COUNT OF THE PRODUCT DOES NOT QUALIFY AS A WBC FAILURE PER THE CURRENT FDA GUIDELINES OF 8 X 10^6 FOR A DOUBLE PLATELET PRODUCT COLLECTION. CONCLUSION: NO FAILURE OCCURRED. | 0
2211197 | A CUSTOMER CONTACTED BAXTER TECHNICAL SERVICES (BTS) REGARDING ASSISTANCE WITH A SYSTEM ERROR 2240 ALARM DURING DRAIN 3 OF 5 ON THE HOMECHOICE MACHINE (HC). THE HOME PATIENT (HP) STATED HE IS CONNECTED NOW BUT HAD DISCONNECTED DURING THE DWELL. THE TECHNICAL SERVICE REPRESENTATIVE (TSR) ASSISTED THE HP TO CYCLE POWER TO CLEAR THE ALARM. THE HP STATED HE WOULD END THERAPY FOR THE DAY. THE HP WAS CONTACTED ON (B)(6) 2011. THE HP STATED THAT THIS WAS AN ISOLATED EVENT. THE HP STATED HE DID NOT DEVELOP ANY ADVERSE REACTIONS OR NEED ANY MEDICAL INTERVENTION. THERE WAS PATIENT INVOLVEMENT; HOWEVER, THERE WAS NO INJURY OR MEDICAL INTERVENTION REPORTED. (B)(4). THIS COMPLAINT FOR A SYSTEM ERROR 2240 WAS CONFIRMED AND THE ASSIGNABLE CAUSE IS DETERMINED TO BE USE ERROR, THE PATIENT DISCONNECTED FROM THE SET. A LABELING REVIEW OF THE HOMECHOICE APD SYSTEMS PATIENT AT-HOME GUIDE ISSUED WAS FOUND TO BE ADEQUATE FOR THE USE ERROR(S) IN THE COMPLAINT. BAXTER HAS CONDUCTED A TREND REVIEW AND FOUND THAT SIMILAR REPORTS HAVE BEEN RECEIVED FOR THE REPORTED PROBLEM. THIS ISSUE IS BEING INVESTIGATED THROUGH CAPA.><(B)(4).THERE WAS NO ALLEGATION REPORTED AGAINST THE BAXTER PRODUCT BY THE CUSTOMER. THEREFORE THE DEVICE WAS NOT REQUESTED FOR EVALUATION AND A BATCH REVIEW WILL NOT BE CONDUCTED. A FOLLOW-UP REPORT WILL BE FILED UPON COMPLETION OF BAXTER'S INVESTIGATION, OR IF ANY ADDITIONAL INFORMATION IS RECEIVED. | 0
2213988 | IT WAS REPORTED THAT THE POWER SUPPLY ON THE DEVICE HAD SHORTED. THE USER PLUGGED THE UNIT INTO THE A/C WALL OUTLET, HEARD A LOUD POP AND SAW A SPARK. THE HOSPITAL STATED THAT THE REPORTED ISSUE WAS DUE TO USER ERROR AND NOT BECAUSE OF A PART FAILURE. THEY STATED THE USER HAD PLACED A PACK OF WET CLEANSING WIPES IN THE BASKET OF THE ROLLING STAND, AND THE LIQUID FROM THE WIPES SEEPED INTO THE POWER SUPPLY INPUT CONNECTOR, SO IT SHORTED OUT WHEN THE OPERATOR CONNECTED THE POWER CORD TO THE AC LINE. THERE WAS NO DAMAGE TO THE MONITOR AND NO USER OR PT INJURY OCCURRED. THIS REPORT IS A DUPLICATE OF MFR REPORT # 9613557-2011-00013 WHICH WAS SUBMITTED ON (B)(4) 2011. THIS MFG SITE WAS RELOCATED IN (B)(4) 2011. THE NEW LOCATION AND ASSOCIATED REGISTRATION NUMBER IS 3008729547. | 1
2217621 | A CUSTOMER REPORTED AN UNEXPECTED POSITIVE CMV RESULT WHEN TESTING A DONOR SAMPLE ON THE GALILEO. TWO SAMPLES WERE DRAWN FROM THE DONOR AND TESTED ON THE SAME PLATE. ONE SAMPLE RESULTED POSITIVE, THE OTHER NEGATIVE. THE SAMPLES WERE RE-TESTED AND BOTH RESULTED NEGATIVE. REVIEW OF INSTRUMENT IMAGES: SAMPLE (B)(6) HAD A FUZZY BUTTON (VISUALLY NEGATIVE). SAMPLE (B)(6) HAD A FUZZY BUTTON WITH A HOLE IN THE CENTER OF THE BUTTON (VISUALLY THE IMAGE LOOKED SUSPECT, POSITIVE). A SERVICE CALL WAS MADE. REAGENT PIPETTOR TUBING AND NEEDLE WERE REPLACED. UNIT WAS TESTED AND FOUND TO BE OPERATING WITHIN SPECIFICATIONS. THE 2 SAMPLES IN QUESTION WERE TESTED WITH ACCEPTABLE RESULTS. | 0
2218020 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THE COULTER LH750 HEMATOLOGY ANALYZER MISREAD A BAR CODE LABEL. THE CUSTOMER REALIZED THAT THE MISREAD OCCURRED BECAUSE THEY CHECK BAR CODE NUMBER AND PATIENT NAME. THE ERRONEOUS RESULTS WERE NOT REPORTED OUTSIDE OF THE LABORATORY. THE MISREADS WERE DISCOVERED PRIOR TO RELEASING RESULTS. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS EVENT. THE LABEL READ CORRECTLY UPON RERUN. SAMPLE LABELS WERE REQUESTED BUT NOT SENT BY THE CUSTOMER. THE CUSTOMER WAS USING CODABAR BARCODE SYMBOLOGY WITH THE CHECKSUM DISABLED. PER BEC LABELING "CERTAIN FEATURES, SUCH AS CHECKSUM DIGITS, MAXIMIZE ACCURACY IN READING CODABAR, CODE 39 AND INTERLEAVED 2-OF-5 LABELS. USE CHECKSUMS TO PROVIDE PROTECTION AGAINST OCCASIONAL MISREAD ERRORS CAUSED BY PROBLEMS SUCH AS DAMAGED OR MISAPPLIED LABELS. IF YOU MUST USE BAR CODES WITHOUT CHECKSUMS, BECKMAN COULTER RECOMMENDS THAT YOU VERIFY EACH BAR-CODE READING TO ASSURE CORRECT PATIENT IDENTIFICATION". A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE TO COLLECT DATA AND PERFORMED BAR CODE READ TEST WHICH PASSED SUCCESSFULLY. | 1
2219877 | IT WAS REPORTED THAT THERE WAS A LOSS OF MONITORING FOR 50 PATIENTS FOR AN UNDISCLOSED AMOUNT OF TIME. (B)(4) CENTRAL STATION MONITORING AND BEDSIDE MONITORING WAS LOST ON 50 EAGLE PATIENT MONITORS DUE TO NETWORK TRAFFIC APPARENTLY ORIGINATING FROM A DASH 4000 PATIENT MONITOR. THE NETWORK TRAFFIC APPEARED TO CAUSE THE EAGLE MONITORS TO ENTER A REBOOT CYCLE WHICH WOULD NOT COMPLETE. ALL MONITORED WAS RESTORED AFTER THE DASH 4000 WAS REMOVED FROM THE NETWORK. THERE WERE NO PATIENT INJURIES REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
2220647 | CUSTOMER STATED THAT THEY EXPERIENCED CARRY-OVER. FIRST SAMPLE POSITIVE FOR ANTI-D AND ANTI-C WAS 4+ IN CELL 1 AND 2 OF THE BIOTESTCELL 3. SAMPLE RUN IMMEDIATELY AFTERWARDS IN THE BATCH WAS 3+ IN CELL 1 AND 2+ IN CELL 2. THE SECOND SAMPLE'S PANEL WAS NEGATIVE AND WHEN RUN INDEPENDENTLY ON THE TANGO FOR ANTI-BODY SCREEN IS CORRECTLY NEGATIVE. ANTIBODY SCREENING WAS REPRODUCED USING THE PROVIDED SAMPLE ON TANGO OPTIMO S/N (B)(4). THE ANTIBODY CONTAINING SAMPLE WAS CHOSEN TO BE THE FIRST ONE IN ORDER, FOLLOWED BY FOUR KNOWN NEGATIVE DONOR SAMPLES. IN THIS APPROACH ALL STEPS OF THE PROCESS WERE CONDUCTED AUTOMATICALLY BY TANGO OPTIMO W/O ANY MANUAL INTERVENTION. THE ORIGINAL, ANTIBODY-CONTAINING SAMPLE RESULTED IN A STRONG POSITIVE (4+) REACTION REGARDING P1 AND P2 CELLS. P3 REACTED CLEARLY NEGATIVE. THE FIRST FOLLOWING KNOWN NEGATIVE SAMPLE RESULTED IN A POSITIVE (2+) REACTION ONLY WITH P1 CELLS. ALL OTHER KNOWN NEGATIVE SAMPLES REACTED NEGATIVE WITH ALL OF THE TESTCELLS. THE CORRECT FUNCTION OF THE AFFECTED REAGENTS WERE PROVED BY TESTING THE RETENTION SAMPLES OF QC LAB ON TANGO. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DID NOT OBSERVE ANY FALSE POSITIVE REACTIONS OF THE AFFECTED REAGENTS THEMSELVES. | 0
2221954 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL CONTENT IN THE BLOOD PRODUCT. DONOR UNIT #: (B)(6). THE PT AGE, SEX AND WEIGHT IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE IS NOT BEING RETURNED. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROGRESS. A F/U REPORT WILL BE PROVIDED. | 0
2224527 | DURING MANUFACTURER REVIEW OF A VNS PATIENT'S PROGRAMMING HISTORY, IT WAS IDENTIFIED THAT THE PATIENT'S VNS SETTINGS WERE SET TO 0MA AND 60 MINUTES OFF ON (B)(6) 2008; THIS IS INDICATIVE OF A PREVIOUS FAULTED SYSTEMS DIAGNOSTICS TEST. THIS WAS NOT CORRECTED UNTIL AT LEAST (B)(6) 2008, WHEN THE SETTINGS WERE NOTED TO BE 1.25MA AND 1.1 MINUTES OFF. THERE ARE NO FAULTED SYSTEMS DIAGNOSTICS IN THE HISTORY TO EXPLAIN THE 0MA AND 60 MINUTES OFF SETTINGS NOTED ON (B)(6) 2008; IT IS SUSPECTED A FAULTED SYSTEMS TEST OCCURRED WITH AN UNKNOWN PROGRAMMING SYSTEM. ANALYSIS OF PROGRAMMING HISTORY. | 0
2226340 | A CUSTOMER REPORTED THAT UNEXPECTED NEGATIVE RESULTS WERE OBTAINED WITH A SAMPLE ORIGINALLY TESTED ON (B)(6) 2011 ON THE GALILEO ECHO INSTRUMENT ((B)(4)) AND THAT A DELAYED TRANSFUSION REACTION HAD OCCURRED. RESULT IMAGE FILES FROM THE ECHO WERE REVIEWED BY AN IMMUCOR TECHNICAL SUPPORT SPECIALIST. SAMPLES VISUALLY APPEARED TO BE POSITIVE FOR CELL 2 OF THE 3-CELL SCREENING ASSAY. AN IMMUCOR FIELD SERVICE ENGINEER PERFORMED THE UNEXPECTED REACTION CHECKLIST WITH NO DEVIATIONS IDENTIFIED. THE INSTRUMENT IS PERFORMING AS EXPECTED. THE CUSTOMER WAS ADVISED PER CUSTOMER COMMUNICATION (B)(4), TO VISUALLY INSPECT THE TEST WELLS OF ALL NEGATIVE RESULTS OBTAINED ON THE ECHO FOR ANTIBODY SCREENING AND IDENTIFICATION ASSAYS PRIOR TO RELEASE OF RESULTS. | 0
2228717 | THE WBC COUNT WAS NOT PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC COUNT IN THE PLATELET PRODUCT. NO MEDICAL INTERVENTION WAS NECESSARY FOR THIS EVENT. THE DISPOSABLE KIT WAS DISCARDED, SO WILL NOT BE AVAILABLE FOR RETURN. THE WHITE BLOOD CELL COUNT IS NOT AVAILABLE AT THIS TIME. DONOR UNIT # (B)(4). THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION WITH THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). THE RUN DATA FILES (RDF) WERE ANALYZED. THE ANALYSIS OF THE RDF DID NOT FIND ANY CONCLUSIVE CAUSE OF THE ELEVATED WBC CONTENT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND TRIMA ACCEL OPERATED AS INTENDED. IT IS POSSIBLE THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. BASED ON THE AVAILABLE DATA, IT ALSO CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2234425 | IT WAS REPORTED THAT THE CUSTOMER HAD A PT IN A BED FOR A STEREO EXAMINATION ON THE MAMMOMAT INSPIRATION SYSTEM. THE SYSTEM WAS TILTED TO AN ANGLE BETWEEN 60 AND 90 DEGREES. WHEN STARTING THE EXAM, THE SYSTEM AUTOMATICALLY RUNS +/-15 DEGREES TO ACQUIRE THE STEREO IMAGES. THE UNIT ACQUIRED A FIRST PAIR OF STEREO IMAGES, AND THEN MOVED 15 DEGREES TO ACQUIRE A SECOND SET. WHEN THE SYSTEM MOVED DOWNWARD 15 DEGREES, IT COLLIDED WITH THE PT BED. THIS CAUSED THE UNIT TO MOVE APPROXIMATELY 20 CM SIDEWAYS ON THE FLOOR AND ROTATE APPROXIMATELY 5 DEGREES OUT OF POSITION. BECAUSE THE UNIT COMPRESSION FORCE WAS LOW AT THE TIME OF THE COLLISION, THE PT'S BREAST WAS RELEASED FROM THE COMPRESSION PADDLE. THE SYSTEM DISPLAYED MULTIPLE ERROR MESSAGES AND HAD TO BE SHUT DOWN AND RESTARTED CAUSING THE EXAM TO BE DELAYED. THERE WAS NO INJURY INVOLVED IN THE DESCRIBED EVENT. THIS EVENT HAPPENED IN (B)(6). A DETAILED INVESTIGATION OF THE SYSTEM IS NECESSARY TO DETERMINE THE ROOT CAUSE OF THE DESCRIBED ISSUE, HOWEVER, BUT OUR FACTORY EXPERTS REPORTED THAT THEY HAD PROVIDED OPERATIONAL HINTS TO PREVENT THE REPORTED ISSUE FROM REOCCURRING. BORDERS ON THE FLOOR WERE MARKED TO SHOW WHERE THE SYSTEM CAN POTENTIALLY COLLIDE WITH THE PT TABLE. CUSTOMER'S ADDRESS: (B)(6). | 0
2234451 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A HIGHER THAN EXPECTED WHITE BLOOD CELL COUNT IN THE COLLECTED BLOOD PRODUCT. PT IDENTIFIER AND AGE ARE UNAVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE TO BE RETURNED FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE RDF DOES NOT SHOW THE ROOT CAUSE OF THE ELEVATED WBC COUNT. IT IS POSSIBLE THAT THE CAUSE OF THE ELEVATED WBC COUNT IS DONOR RELATED. IT IS ALSO POSSIBLE THAT THIS FAILURE IS PART OF THIS SITE'S STATISTICAL DISTRIBUTION FOR LEUKOREDUCTION. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. THIS REPORT WAS FILED BEYOND THE 30-DAY TIMEFRAME DUE TO AN INTERNAL PROCESSING ISSUE THAT IS CURRENTLY BEING EVALUATED FOR APPROPRIATE CORRECTIVE ACTIONS. | 0
2234962 | IT WAS INITIALLY REPORTED THAT THE PHYSICIAN WAS HAVING INTERMITTENT COMMUNICATION DIFFICULTIES WITH HIS PROGRAMMING SYSTEM. TROUBLESHOOTING WAS PERFORMED AND THE ISSUE SEEMED TO THE RELATED TO THE SERIAL ADAPTOR CORDS CONNECTION WITH THE HANDHELD. THE PHYSICIAN WAS SENT A NEW HANDHELD WITH A SERIAL CABLE. (B)(4) ATTEMPTS FOR PRODUCT RETURN HAVE BEEN UNSUCCESSFUL TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION AND PRODUCT ANALYSIS HAS BEEN COMPLETED. AN ANALYSIS WAS PERFORMED ON THE HANDHELD AND THE REPORTED ALLEGATION OF INTERMITTENT COMMUNICATION COULD NOT BE VERIFIED. THE HANDHELD WAS RECEIVED WITHOUT A SERIAL CABLE. DURING THE ANALYSIS, NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE SERIAL CORD ADAPTOR, WHICH WAS THE SUSPECTED ISSUE, WAS NOT RETURNED. FOLLOW-UP WITH THE OFFICE INDICATED THAT THEY DID NOT KNOW WHERE THE CORD WAS. THEY HAVE NOT HAD ANY FURTHER ISSUE SINCE THE HANDHELD WAS REPLACED. .><DEVICE FAILURE IS SUSPECTED IN THE SERIAL CORD ADAPTOR CORD WHICH WAS NOT RETURNED TO THE MANUFACTURER, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2237297 | REPORTER INDICATED A VNS DELL X5 COMPUTER WAS FREEZING DURING USE. THE EXACT SCREEN IN WHICH THE COMPUTER WAS FREEZING IS NOT KNOWN. THE COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE PENDING ANALYSIS. ATTEMPTS FOR FURTHER INFORMATION REGARDING THE SCREEN FREEZING ARE IN PROGRESS><REPORTER INDICATED THAT THE VNS COMPUTER WAS FREEZING DURING INTERROGATIONS. PRODUCT ANALYSIS OF THE VNS COMPUTER AND FLASHCARD WAS COMPLETED. NO ANOMALIES WERE IDENTIFIED WITH THE COMPUTER, FLASHCARD, OR SOFTWARE COMPONENTS, AND ALL DEVICES PERFORMED PER SPECIFICATIONS. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2237315 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WHEN TESTING A PATIENT SAMPLE WITH CAPTURE-R READY-SCREEN 3 (CRRS 3) ON THE ECHO. THE PATIENT SAMPLE WAS RE-TESTED ON THE ECHO AND RESULTED POSITIVE. AN ANTI-FYA WAS IDENTIFIED. REVIEW OF CAMERA IMAGES: BATCH ID (B)(4) LOT R157 WITH INDICATOR CELL (B)(4) EXP 11AUG2011 (BATCH SHOWING UNEXPECTED NEGATIVE) NO ERRORS OCCURED DURING PROCESSING. CELL 1 0, CELL2 0, CELL3 0. UPON REVIEWING COLOR CHECK IMAGES OF ALL WELLS, VISUALLY NO SERUM WAS DISPENSED IN THE WELLS. A SERVICE CALL WAS MADE. REPLACED LIQUID LEVEL SENSE PRINTED CIRCUIT BOARD AND PROBE. REPLACED ALL ECHO INTERNAL FINAL ASSEMBLY TUBING. INSTRUMENT WAS TESTED WITH EXPECTED RESULTS. INVESTIGATION OF THIS EVENT IDENTIFIED THAT THE UNEXPECTED NEGATIVE RESULT WAS DUE TO THE ABSENCE OF PLASMA IN THE TEST WELL. THE MOST LIKELY ROOT CAUSE IS AIR ASPIRATION RESULTING FROM A SAMPLE CONTAINING BUBBLES OR FOAM. (B)(4), EXPECTED TO BE DISTRIBUTED ON (B)(4) 2011, WILL NOTIFY CUSTOMERS OF THIS ISSUE. IMMUCOR RECOMMENDS THAT CUSTOMERS EVALUATE THEIR PRE-ANALYTICAL SAMPLE HANDLING PROCEDURES IN ORDER TO ELIMINATE THE PRESENCE OF FOAM OR BUBBLES IN THE SAMPLE. | 0
2237817 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT A BLOODY LEAK WAS NOTED AROUND THE NEEDLE OF THE COULTER HMX AUTOLOADER ANALYZER WHEN TRYING TO RUN A SAMPLE IN PRIMARY MODE. THE CUSTOMER CONTINUED TO RUN THE INSTRUMENT IN OPEN MODE WITH NO FURTHER LEAKS. THERE WAS NO REPORT OF ANY PATIENT RESULTS BEING AFFECTED BY THIS INCIDENT. THE USER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE INCIDENT. MEDICAL ATTENTION WAS NOT SOUGHT. THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS. A BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(6) 2011 FOR THIS EVENT. THE FSE INDICATED THAT THE ACTUATOR AT THE PV21 WAS STICKING AND NOT ALLOWING THE BELLOWS TO DRAIN. PV21 OPENS THE PATH FROM THE NEEDLE BELLOWS TO THE WASTE. THE ACTUATOR WAS REPLACED AND REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES PRIOR TO RETURNING IT INTO SERVICE. THE ROOT CAUSE FOR THE LEAK WAS THAT THE ACTUATOR AT PV21 WAS STICKING; THEREFORE, NOT ALLOWING THE BELLOWS TO DRAIN CAUSING THEM TO OVERFLOW AND SPILL. AS PER PRODUCT LABELING, BCI URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. BEC IDENTIFIER FOR THIS REPORT IS (B)(4). | 0
2240521 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WHEN TESTING A PATIENT SAMPLE ON THE ECHO. CUSTOMER STATES ECHO RESULTS ARE NEGATIVE OR EQUIVOCAL BUT IMAGES APPEAR POSITIVE. THE PATIENT HAS HISTORY OF ANTI-C AND ANTI-E. REVIEW OF CAMERA IMAGES: SAMPLE (B)(6) TESTED ON (B)(6) 2011 WITH CAPTURE-R READY ID EXTEND I LOT DP049 CELL 8 (C NEG, E POS HOMOZYGOUS) WAS INTERPRETED AS NEGATIVE WITH A 2 REACTION STRENGTH, VISUALLY POSITIVE ROUGH BUTTON, PINK ADHERENCE THE EQUIVOCALS IN THIS BATCH CELLS 7 (C AND E POS HOMOZYGOUS) ,9 (C NEG, E HOMOZYGOUS,10 (C NEG, E HOMOZYGOUS) AND 14 (C NEG, E HOMOZYGOUS) ALL HAD 8 FOR A REACTION STRENGTH AND ALL APPEARED CLEARLY VISUALLY POSITIVE. A SERVICE CALL WAS MADE. INVESTIGATION REVEALED THAT THE CAMERA SHOULD BE REPLACED. REPLACED CAMERA. INSTRUMENT WAS TESTED AND FOUND TO BE OPERATING WITHIN SPECIFICATIONS. | 0
2242295 | THE CUSTOMER CONTACTED BECKMAN COULTER INC (BEC) TO REPORT THAT THE BACKWASH CUP BROK IN HALF THE DAY AFTER MOD WAS INSTALLED AND WHOLE BLOOD DILUTE MIXTURE LEAKED. THE MOD RELEASED AN IMPROVED REVISION OF THE PROBE WIPE ASSEMBLY. WHOLE BLOOD DILUENT MIXTURE WAS LEAKING. THE CUSTOMER WAS WEARING HIS PPE. NO CONTACT TO MUCOUS MEMBRANE OR BROKEN SKIN. NO ONE SOUGHT MEDICAL ATTENTION. NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT AS A RESULT OF THIS INCIDENT. IT IS UNKNOWN IF THERE IS AN EXPOSURE CONTROL OR RISK MANAGEMENT PLAN IN PLACE AT THE FACILITY. ON (B)(6) 2011, A FIELD SERVICE ENGINEER (FSE) FOUND RINSE CUP BROKEN. REPLACED RINSE CUP, RAN START UP AND SEVERAL PROBE WIPE TESTS. NO ERRORS OR DEFECTS NOTED. ROOT CAUSE IS UNKNOWN. PER LABELING, BECKMAN COULTER URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IS NOT LIMITED TO, WEARING PROTECTIVE EYE WEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS INSTRUMENT OR ANY OTHER AUTOMATED LABORATORY ANALYZER. BEC INTERNAL IDENTIFICATION NUMBER IS (B)(4). | 0
2245823 | THE CUSTOMER WOULD LIKE TO KNOW THE POSSIBLE CAUSE OF THE ELEVATED WBC CONTENT IN THE PLATELET PRODUCT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN. THE PT INFO IS UNAVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROGRESS. A F/U REPORT WILL BE PROVIDED. | 0
2247439 | THE CUSTOMER INDICATED THAT THEY HAD OBSERVED A LEAK ON A LH SLIDE STAINER ON TWO DIFFERENT DAYS. THIS IS REPORT ONE OF TWO AND REPRESENTS THE LH SLIDE STAINER LEAK OBSERVED ON (B)(6) 2011 INITIALLY, THE CUSTOMER HAD INDICATED THAT THE LEAK HAD NOT CAUSED AN OVERFLOW OF THE LH SLIDE STAINER OVERFLOW TRAYS. HOWEVER, A BECKMAN COULTER INC. FIELD SERVICE ENGINEER WHILE ON SITE TO ADDRESS THE LH SLIDE STAINER LEAK, INDICATED THAT THE CUSTOMER HAD REPORTED THAT THE LH SLIDE STAINER BATH NUMBER TWO HAD LEAKED AND HAD OVERFLOWED THE INSTRUMENT TRAYS. THE HEALTHCARE WORKERS INTERFACING WITH THE MACHINE WERE WEARING PERSONAL PROTECTIVE EQUIPMENT, WHICH INCLUDED LABORATORY COATS AND GLOVES, AT THE TIME OF OCCURRENCE THERE WAS NO BIOHAZARDOUS EXPOSURE TO HEALTHCARE WORKER UNCOVERED WOUNDS OR MUCOUS MEMBRANES AND NO INJURIES WERE REPORTED. NO PERSONNEL SOUGHT ANY MEDICAL ATTENTION IN ASSOCIATION WITH THIS EVENT. NO PATIENT RESULTS WERE AFFECTED BY THIS EVENT. NO DEATH, INJURY OR MODIFICATION TO PATIENT TREATMENT WAS ASSOCIATED WITH THIS EVENT. THERE WAS AN EXPOSURE PLAN IN PLACE AT THE FACILITY AND A MATERIAL SAFETY DATA SHEET WAS AVAILABLE. SERVICE WAS DISPATCHED TO THE SITE ON (B)(4) 2011 FOR THIS EVENT. THE FIELD SERVICE ENGINEER (FSE) REPLACED THE FILL AND DRAIN TUBING TO BATH TWO AND REPLACED THE WASTE FILTER. THE INSTRUMENT WAS RETURNED INTO SERVICE AFTER COMPLETION OF VERIFIED REPAIRS. THE ROOT CAUSE OF THE LEAKS WERE ULTIMATELY ATTRIBUTED TO A DAMAGED SENSOR MECHANISM AND A PLUG IN THE STAIN FILL TUBE. MDRS ASSOCIATED WITH THIS EVENT: 1061932-2011-01445, 1061932-2011-01444. | 0
2247552 | CUSTOMER REPORTED DRAWER ON THE PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT WAS PRESENT. (B)(4). ADDITIONAL DATA/ FAILURE INVESTIGATION. FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER TO FAIL. | 1
2253457 | NOTE: THIS EVENT WAS ORIGINALLY FILED ON (B)(4) 2011, UNDER FORM 3500. VARIAN MEDICAL SYSTEMS RECEIVED A GE (B)(4) RELOAD WITH A VARIAN A-195 INSERT FOR FAILURE ANALYSIS. THE CUSTOMER (B)(6), RETURNED THIS TUBE UNIT TO VARIAN WITH THE X-RAY TUBE SERVICE REPORT STATING THAT, "OIL HEATED UP ENOUGH TO MELT THE PLASTIC/ACRYLIC PORT. THIS ALLOWED ALL OF THE OIL TO DRAIN OUT OF THE TUBE INTO THE COLLIMATOR." UNDER THE HEADING REASON FOR RETURN, THE BOX "OIL LEAK" WAS CHECKED AND ALSO THE "OTHER" STATED, "SHATTERED INSERT." REP FROM VARIAN, (B)(6), GE, AND (B)(4) MET ON CONFERENCE CALLS ON (B)(6) 2011. FROM THESE DISCUSSIONS, IT WAS DISCOVERED THAT THE GE PROTEUS SYSTEM HAD AN OLD VERSION OF SOFTWARE (VER. 5.57) INSTALLED ON IT THAT WAS KNOWN TO CAUSE THIS TYPE OF FAILURE. GE'S RECORDS SHOW A SOFTWARE UPDATE TO VERSION 5.59 IN 2009 TO PREVENT THIS TYPE OF KNOWN FAILURE. (B)(6) IS RESPONSIBLE FOR SERVICING THEIR OWN MACHINES, AND COULD NOT ACCOUNT FOR THE V 5.57 SOFTWARE ON THE PROTEUS SYSTEM AT THE TIME OF THE FAILURE. | 0
2254837 | CUSTOMER RAN WEAK D TESTING ON AN RH POSITIVE SAMPLE IMMEDIATELY FOLLOWED BY A NEGATIVE RH SAMPLE AND THE RH NEGATIVE SAMPLE TESTED POSITIVE. THE RH POSITIVE SAMPLE WAS AB POSITIVE (WEAK D 4+) AND THE RH NEGATIVE SAMPLE WAS 0 NEGATIVE (WEAK D 3+). SHE RAN THE SAME SAMPLE ON EACH OF HER 4 TANGOS AND GOT VARYING RESULTS (SN (B)(4)). DATE SAMPLE AVAILABLE: (B)(6) 2011, ALL SAMPLES SHOWED HEMOLYSIS, PLASMA COLOURED RED. ADD'L SERIAL# (B)(4). IN GENERAL CARRY-OVER EVENTS CANNOT BE EXCLUDED IF APPLYING INSTRUMENTATION WITH A DESIGN LAYOUT, THAT IS AMONGST OTHERS, ALSO TRUE FOR TANGO OPTIMO. THIS MATTER IS ALSO DISPLAYED IN ITS SPECIFICATIONS. DESPITE COATED NEEDLES AS WELL AS RINSING STEPS, SUCH A PHENOMENON MAY OCCUR IN VERY RARE CASES, IS USUALLY BASED ON A SAMPLE-SPECIFIC REACTION AND CAN BE CAUSED BY A HIGH PROTEIN LEVEL IN PLASMA (INDUCED BY SEVERE SICKNESS OR APPLICATION OF DRUGS) OR ANTIBODIES WITH A HIGH TITRE. SINCE ANY SPECIMEN BEING CONSPICUOUS DUE TO A POSITIVE ABS RESULT SHOULD UNDERGO RETESTING AND BECAUSE ANY ABS POSITIVE SAMPLE FROM A BLOOD DONOR SHOULD BE EXCLUDED FROM TRANSFUSION, NO SERIOUS THREAT NEITHER FOR PT NOR USER OR DEVICE MAY ARISE DUE TO A CARRY-OVER EVENT. THE CORRECT FUNCTION OF THE AFFECTED REAGENTS WERE PROVED BY TESTING THE RETENTION SAMPLES OF OUR QUALITY CONTROL LAB IN TANGO. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DID NOT OBSERVE ANY FALSE POSITIVE REACTIONS. ACCORDING TO THE INTENDED USE OF THE ALSO AFFECTED REAGENT SOLIDSCREEN II ANTI-D BLEND THE TEST IS USED TO TEST DONOR SAMPLES WHICH HAVE BEEN TESTED NEGATIVE WITH IGM ANTI-D ON ERYTYPE S ON THE TANGO. THE CUSTOMER TESTED RH(D) POSITIVE SAMPLES. THEREFORE WE DA NOT EXCLUDE A CUSTOMER HANDLING ERROR. A REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE AFFECTED NEGATIVELY THE QUALITY OF THE COMPLAINED LOTS. | 0
2255575 | OLYMPUS WAS INFORMED THAT THE FRONT LEFT WHEEL OF THE CART BROKE AFTER A NURSE HAD MOVED THE UNIT TO A DIFFERENT LOCATION. ONE OF THE CASTORS ON THE DEVICE REPORTEDLY BROKE, AND THE CART TOPPLED AND A PORTION OF THE WORKSTATION STRUCK THE BACK OF THE NURSE. THE NURSE REPORTEDLY SUSTAINED A MINOR UNSPECIFIED INJURY THAT REQUIRED MEDICAL ATTENTION. THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO OLYMPUS FOR EVALUATION, HOWEVER A LOCAL REPRESENTATIVE WAS ABLE TO VISIT THE FACILITY AND EXAMINE THE CASTOR. THE CASTOR BOLT WAS NOTED TO BE SIGNIFICANTLY CORRODED. OLYMPUS WAS INFORMED THAT HOSPITAL PERSONNEL HAD REPLACED THE CASTORS ON THIS CART IN (B)(6) 2011 FOR UNKNOWN REASONS. THE EVALUATION REVEALED THAT THERE WAS A FAILURE OF THE CASTOR MOUNTING STUD. OLYMPUS KEYMED HAS CONCLUDED THAT THE MOST PROBABLE CAUSE OF THE REPORTED EVENT WAS DUE TO THE IMPROPER INSTALLATION OF THE CASTOR WHEEL, THEREBY CREATING A GAP BETWEEN THE WASHER AND THE WORKSTATION BASE. THE HIGH LOAD PUT ON THE CASTOR THREAD LIKELY RESULTED IN THE FRACTURE. THE USER FACILITY WAS PROVIDED A REPLACEMENT TROLLEY. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | 0
2259951 | ADD'L INFO WAS RECEIVED ON (B)(4) 2011, AT WHICH TIME THE TICKET WAS MADE INTO A COMPLAINT. CUSTOMER REPORTED THAT IQUERY RETURNED INCORRECT RESULTS. WHILE PERFORMING A QUERY BY PT NAME, MULTIPLE RESULTS WERE SHOWN. ONE OF THE ASSOCIATED MRNS WAS FOUND TO BE FOR TWO DIFFERENT PTS. ROOT CAUSE INVESTIGATION IS IN PROGRESS. | 1
2260235 | THE SITE REPORTS THAT WHEN INSERTING KEY IMAGES WITH PRIORS, THE DATE ON THE KEY IMAGES ONLY REFLECTS THE DATE OF THE CURRENT IMAGES AND NOT THE PRIOR DATE. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
2260264 | DURING PRETEST, ICE TRAVELED UP THE ENTIRE PROBE. INCIDENT OCCURRED DURING PRETESTING - NO PT CONTACT, INJURY OR DELAY. EVAL CONFIRMED SHAFT FROSTING. THE MOST LIKELY CAUSE IS A VACUUM FAILURE. | 0
2264239 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PATIENT WAS PRESENT. (B)(4). ADDITIONAL DATA/FAILURE INVESTIGATION. FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL OBJECT OBSTRUCTION CAUSING DRAWER TO FAIL. | 1
2264256 | CUSTOMER REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT HARM. (B)(4). ADD'L DATA / FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER TO FAIL. | 1
2264625 | THE CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) TO REPORT THE UNINTERRUPTED POWER SUPPLY (UPS) WHERE THE LH750 WAS CONNECTED TO CAUGHT ON FIRE; HOWEVER, NO POWER SURGES WERE REPORTED. SMOKE AND FIRE DAMAGE WAS LIMITED TO THE UPS ONLY. NO OTHER DAMAGE WAS DONE. NO INJURY, DEATH, CHANGE IN PATIENT TREATMENT, OR MEDICAL ATTENTION WAS SOUGHT OR REPORTED BY ANY EMPLOYEE IN ASSOCIATION WITH THIS EVENT. THE FIRE DEPARTMENT WAS CALLED TO THE SCENE. ON (B)(4) 2011, A BEC FIELD SERVICE ENGINEER (FSE) FOUND THE 250V 1/16 AMP FUSE BLOWN IN THE LH POWER SUPPLY AND REPAIRED INSTRUMENT (REPLACED FUSE) AND VERIFIED INSTRUMENT PERFORMANCE PER ESTABLISHED PROCEDURES. THE RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. SYSTEM VALIDATION DOCUMENTED IN CUSTOMER QC RECORD. THE UPS FAILURE CAUSED THE LH750 FUSE TO BLOW. THE UPS WAS NOT AVAILABLE FOR EVALUATION AS IT WAS DISCARDED. (B)(4). | 0
2265467 | CUSTOMER REPORTED HAVING A PT ON THE TRACKING BOARD THAT HAS A PENICILLIN ALLERGY, BUT WHEN ZOSYN ORDER WAS DOCUMENTED IN THE ED PULSECHECK (ELECTRONIC MEDICAL RECORD) SOFTWARE APPLICATION, IT DID NOT DISPLAY APPROPRIATE ALLERGY CHECKING NOTIFICATION, YET IT DID FOR MAXIPIME. THERE WERE NO REPORTS OF PT IMPACT OR INJURY AS A RESULT OF THIS ERROR. WE HAVE REC'D AND INVESTIGATED TWO SIMILAR CUSTOMER REPORTS OF MALFUNCTION INVOLVING PICIS EMR SOFTWARE APPLICATION. CUSTOMER CLINICIANS REPORTED OBSERVING ERRONEOUS RESULTS WITH THE DEVICE ALLERGY INTERACTIONS FEATURE USED IN CHECKING MEDICATION ORDERS AGAINST DOCUMENTED PT ALLERGY HISTORY. THERE WERE NO ASSOCIATED REPORTS OF MEDICATION ERROR OR PT INJURIES AS A RESULT. OUR INTERNAL INVESTIGATION HAS CONFIRMED AN ERROR CONDITION EXISTS WITHIN SOFTWARE CODE FOR THE DEVICE, WHEN USED IN CONJUNCTION WITH MULTUM 3RD PARTY DRUG DATABASE CONTENT. THE WORKFLOW IMPACTED IS AS FOLLOWS: PT HAS A DOCUMENTED ALLERGY IN OUR DEVICE APPLICATION TO A SPECIFIC DRUG OR CATEGORY OF DRUGS, CLINICIAN ATTEMPTS TO PLACE AN ORDER FOR A MULTI-COMPONENT MEDICATION WITH ONE OF THE COMPONENTS BEING IN THE SAME CATEGORY AS THE PT'S ALLERGY, APPLICATION DOES NOT DISPLAY THE APPROPRIATE ALLERGY INTERACTION NOTIFICATION. FOR EXAMPLE: PT HAS AN ALLERGY TO CATEGORY OF 'PENICILLINS' AND THE USER ATTEMPTS TO ORDER 'ZOSYN' (PIPERACILLIN/TAZOBACTAM) - NO ALLERGY INTERACTION NOTIFICATION (PIPERACILLIN BELONGS TO THE ALLERGY CATEGORY OF 'PENICILLINS') IS DISPLAYED. IT SHOULD BE NOTED THAT IF THE COMPONENT IS THE SAME DRUG AS THE PT ALLERGY, THE ALLERGY CHECKING WILL DISPLAY CORRECTLY. ROOT CAUSE: OUR INVESTIGATION HAS CONCLUDED THAT THE ROOT CAUSE RELATES TO A SOFTWARE CODING ERROR INTRODUCED DURING DEVELOPMENT OF SOFTWARE RELEASE V5.0 SERVICE PACK L, WHICH WAS NOT DETECTED DURING VERIFICATION AND VALIDATION TESTING. SERVICE PACK L WAS RELEASED ON (B)(4) 2011 AND WAS SUBSEQUENTLY CARRIED FORWARD INTO SERVICE PACK M RELEASE, ON (B)(4) 2011. WE HAVE CONFIRMED THAT NO EARLIER OR LATER VERSIONS OF SOFTWARE ARE AFFECTED. CORRECTIVE ACTION(S): THE FOLLOWING CORRECTIVE MEASURES HAVE BEEN DEPLOYED TO PREVENT FURTHER OPPORTUNITY FOR THIS ERROR TO BE EXPERIENCED IN THE FIELD: AFFECTED SERVICE PACK HAS BEEN REMOVED FROM AVAILABILITY ON OUR USER COMMUNITY WEBSITE. IN ACCORDANCE WITH FDA QSR 21 CFR SUBCHAPTER H PART 806, WE HAVE ISSUED A WRITTEN COMMUNICATION TO ALL POTENTIALLY AFFECTED CLIENTS, EXPLAINING THE POTENTIAL FOR THIS ERROR, ALONG WITH INSTRUCTIONS TO MONITOR DOCUMENTED PT ALLERGIES CLOSELY WHEN PRESCRIBING MEDICATION ORDERS IN OUR DEVICE APPLICATION, USING MULTUM DRUG DATABASE CONTENT. ON (B)(4), A SOFTWARE SOLUTION WAS RELEASED IN SERVICE PACK N FOR V5.0, WHICH WE ARE CURRENTLY COORDINATING IMPLEMENTATION WITH ALL POTENTIALLY AFFECTED CLIENTS. WE WILL TRACK THE PROGRESS ON THIS FIELD CORRECTION AND COORDINATE THIS INFO WITH OUR LOCAL FDA DISTRICT RECALL COORDINATOR. AN INTERNAL CAPA INVESTIGATION IS IN PROGRESS TO REVIEW INTERNAL PROCESSES FOR OPPORTUNITIES WHERE IMPROVEMENT IS NEEDED TO PREVENT THIS TYPE OF ERROR IN THE FUTURE. (B)(4) | 1
2265600 | CUSTOMER CALLED REPORTING UNEXPECTED NEGATIVE REACTIONS ON THE ECHO WHEN TESTING PATIENT SAMPLES WITH CAPTURE-R READY SCREEN 3 (CRRS 3). THE RESULTS WERE VISUALLY POSITIVE BUT WERE CALLED NEGATIVE BY THE INSTRUMENT. AN ANTI-E WAS DETECTED WITH CAPTURE-R READY ID (CRRID) ON THE ECHO. REVIEW OF INSTRUMENT IMAGES: CELL 1 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL2 REACTED NEGATIVE HOMOZYGOUS FOR E; VISUAL APPEARANCE WEAK POSITIVE (THIS IS BEING INVESTIGATED UNDER (B)(4)). CELL 3 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. POSITIVE CONTROL WELL REACTED 4+ POSITIVE; VISUAL APPEARANCE POSITIVE. CRRID LOT ID146, (B)(4): CELL 1 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL2 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 3 REACTED EQUIVOCAL HOMOZYGOUS FOR E; VISUAL APPEARANCE WEAK POSITIVE. CELL 4 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 5 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 6 REACTED NEGATIVE HOMOZYGOUS FOR E; VISUAL APPEARANCE WEAK POSITIVE. CELL 7 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 8 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 9 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 10 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 11 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 12 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE CELL 13 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. CELL 14 REACTED NEGATIVE, NEGATIVE FOR E; VISUAL APPEARANCE NEGATIVE. POSITIVE CONTROL WELL REACTED 4+ POSITIVE; VISUAL APPEARANCE POSITIVE. NEGATIVE CONTROL WELL REACTED NEGATIVE; VISUAL APPEARANCE NEGATIVE. CELL 2 ON THE RS3 AND CELL AND 6 ON THE CRRID WERE VISUALLY POSITIVE BUT WERE CALLED NEGATIVE BY THE INSTRUMENT. THIS ISSUE WAS COMMUNICATED TO CUSTOMERS IN TECHNICAL COMMUNICATION (B)(4) ON NOVEMBER 25, 2009. CUSTOMERS WERE ADVISED TO VISUALLY INSPECT ALL NEGATIVE REACTIONS. A SERVICE CALL WAS MADE. UNEXPECTED REACTIONS CHECKLIST WAS PERFORMED. SOME MODULES CONTAINED EXCESSIVE SALT DEPOSITS. -CLEANED PROBE AND PROBE HOUSING ASSEMBLY. -CLEANED RINSE STATION AND REPLACED FILTER. -REPLACED PROBE SUPPLY TUBING FROM FLUIDICS TO MAIN. -CLEANED AND RINSED WASH MANIFOLD. INSTRUMENT WAS TESTED AND IS OPERATING WITHIN SPECIFICATIONS. | 0
2265835 | IT WAS REPORTED THAT A CIC PRO CLINICAL INFORMATION CENTER FAILED TO PRODUCE AN AUDIBLE ALARM FOR A PATIENT'S VENTRICULAR TACHYCARDIA EVENT. THE SYSTEM DISPLAYED A VISUAL ALARM AND PRINTED AN ALARM STRIP. THE CLINICIAN NOTICED THE VISUAL ALARM. THE PATIENT WAS CHECKED, DID NOT REQUIRE ANY TREATMENT, AND NO INJURY WAS REPORTED. DURING DISCUSSION WITH A GE HEALTHCARE SERVICE ENGINEER THE SITE TECHNICIAN TESTED THE SYSTEM'S AUDIO ALARMS AND NONE WOULD SOUND, HOWEVER ALL VISUAL INDICATORS FUNCTIONED AS EXPECTED. THE SITE TECHNICIAN WAS INSTRUCTED TO INITIATE A CONTROLLED REBOOT OF THE CIC. A CONTROLLED REBOOT INCLUDES NOTIFYING STAFF AND MINIMIZING THE LOSS OF CENTRAL MONITORING OF PATIENT EVENTS. THE CONTROLLED REBOOT RESTORED THE AUDIO ALARM FUNCTION. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
2265869 | THE CUSTOMER COMPLAINED ABOUT UNUSUAL SOLIDSCREEN REACTIONS CALLED "?" THROUGH 3+ BY TANGO OPTIMO WHEN THE SAMPLES SHOULD HAVE YIELDED NEGATIVE SOLIDSCREEN REACTIONS. A SERVICE VISIT WAS PERFORMED DURING WHICH THE ENGINEER NOTICED THAT THE COLOUR OF THE WELLS REMAINED RED AFTER SPINNING WHICH IS AN INDICATION OF HEMOLYSIS. AFTER SETTING UP FRESH WASH BUFFER, ALL THE QUALITY CONTROLS PASSED SUCCESSFULLY WITH NO FURTHER PROBLEMS. THE FINDINGS STRONGLY INDICATE THAT THE WASH BUFFER IN USE WAS OUT OF SPECS AND CAUSED HEMOLYSIS OF THE REAGENT RED CELLS. A HOMOGENEOUS DISTRIBUTION OF PIXELS WITH SIMILAR GREYSCALE VALUES ARE ALWAYS INTERPRETED AS A POSITIVE ANTIBODY SCREEN RESULT BY TANGO OPTIMO. A DIFFERENTIATION BETWEEN A SOLIDPHASE AND A COLOURED BACKGROUND DUE TO HEMOLYSED CELLS CANNOT BE MADE BY THE SYSTEM. THIS COMPLAINT WAS THUS TREATED AS A USER ERROR. THE ITI ON HAND IS PART OF A SERIES OF COMPLAINTS WITH DELAYED REPORTING TO THE AUTHORITIES. ALL OF THESE CASES WILL INDIVIDUALLY BE INVESTIGATED AND PROCESSED BUT WILL COLLECTIVELY INDUCE A CAPA. | 0
2267028 | DURING A REVIEW OF THE PATIENT'S PROGRAMMING HISTORY, A FAULTED DIAGNOSTIC TEST WAS OBSERVED ON (B)(6) 2008. THIS RESULTED IN A CHANGE TO THE PATIENT'S SETTINGS. A FINAL INTERROGATION WAS NOT PERFORMED AND THE PATIENT LEFT AT UN-INTENDED SETTINGS. THE PATIENT WAS THEN SEEN ON (B)(6) 2009 AND THE DEVICE WAS INTERROGATED AT 0MA. THE PATIENT'S GENERATOR WAS NOT PROGRAMMED BACK ON UNTIL (B)(6) 2009. ANALYSIS OF PROGRAMMING HISTORY | 0
2269085 | RACK MOUNT STARTED SMOKING/STEAMING ON FIRE INSIDE UNIT, WHEN THEY TURN SERVER ON. SERVER IS A RACKMOUNT AND WAS INSTALLED IN AN IT ROOM WITH OTHER SERVERS. THEY HAD TO UNPLUG THE SERVER AND TAKE OUTSIDE. WE CALLED HP AND THEY WENT ON SITE TO SAY THE RACK MOUNT HAS COMPLETELY MELTED AND CANNOT BE REPAIRED. THE SUBJECT DEVICE WILL BE SHIPPED TO HP FOR EVAL. | 1
2269105 | ON (B)(6) 2011, LEICA MICROSYSTEMS REC'D A COMPLAINT FROM (B)(6) REGARDING FAILURE OF A NUMBER OF SLIDES TO STAIN IN THE RUN STARTED AT 07:57AM ON (B)(6) 2011, USING BOND-MAX SERIAL NUMBER (B)(4). ON (B)(6) 2011, LEICA MICROSYSTEMS WAS ADVISED THAT MANUAL STAINING OF TWO SLIDES FOR CALRETININ AND CA125 PREPARED FROM ONE PT LIVER CORE BIOPSY SAMPLE WHICH HAD NOT BEEN STAINED IN THE RUN STARTED AT 07:57AM ON (B)(6) 2011 WAS UNSUCCESSFUL; AND THAT NO TISSUE WAS AVAILABLE TO PREPARE FURTHER SLIDES. INFO REGARDING THE IMPACT OF THIS STAINING FAILURE IN RELATION TO ESTABLISHING A DIAGNOSIS FOR THE PT SAMPLE AND WHETHER RE-BIOPSY OF THE PT IS REQUIRED AND/OR HAS BEEN PERFORMED IS BEING SOUGHT BY LEICA MICROSYSTEMS. | 0
2269393 | IT WAS REPORTED THAT THE CLINICAL INFO CENTER (CIC) WENT DOWN AND NO BACK UP CIC WAS AVAILABLE ON THE UNIT. THE HOSPITAL STATED THIS WAS NOTICED IMMEDIATELY, HOWEVER THE FAILURE WAS UNEXPECTED. MONITOR FOR SIX PTS WAS LOST FOR 90 MINUTES. NO PT INJURY REPORTED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2272122 | ALL ATTEMPTS TO THE REPORTER FOR RETURN OF THE VNS COMPUTER AND FLASHCARD HAVE BEEN UNSUCCESSFUL TO DATE.><THE VNS COMPUTER AND FLASHCARD WERE RETURNED FOR PRODUCT ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD COMPUTER WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED SCREEN FREEZING ALLEGATION WAS NOT VERIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO SPECIFICATIONS. THE SCREEN FREEZING EVENT WAS UNABLE TO BE RECREATED.><REPORTER INDICATED THAT A VNS (B)(6) COMPUTER WAS CONSISTENTLY FREEZING. RESETTING THE COMPUTER RESOLVED THE SCREEN FREEZING, BUT THE FREEZING CONTINUES TO OCCUR. ATTEMPTS FOR RETURN OF THE COMPUTER AND FLASHCARD ARE IN PROGRESS. . | 0
2273059 | WHILE REVIEWING THE PATIENT'S PROGRAMMING HISTORY ON (B)(6) 2011 IT WAS DISCOVERED THAT ON (B)(6) 2009 A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED THAT CHANGED THE PATIENT'S SETTINGS TO SYSTEM DIAGNOSTICS TEST SETTINGS OF OUTPUT=1MA/FREQUENCY=20HZ/PULSE WIDTH=500USEC/ON TIME=30SEC/OFF TIME=60MIN. THE PATIENT LEFT THE CLINICAL VISIT AT THESE INCORRECT SETTINGS AND IT WASN'T UNTIL THE PATIENT'S NEXT VISIT ON (B)(6) 2009 WHEN THE SETTINGS WERE CORRECTED. TRAINING WILL BE PROVIDED TO THE PHYSICIAN TO REMIND HIM OF THE NEED TO PERFORM A FINAL INTERROGATION TO MAKE SURE THE PATIENT IS AT THE CORRECT SETTINGS BEFORE THEY LEAVE THE CLINICAL VISIT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2275803 | CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN TESTING A SAMPLE ON THE GALILEO. THE SAMPLE WAS TESTED BY MANUAL SCREEN AND RESULTED POSITIVE. THE SAMPLE WAS RE-TESTED ON A DIFFERENT GALILEO AND RESULTED POSITIVE. THE SAME LOT OF REAGENTS WERE USED ON BOTH INSTRUMENTS. THE ANTIBODY WAS LATER IDENTIFIED AS AN ANTI-D. CUSTOMER DOES NOT HAVE SAMPLE REMAINING FOR FURTHER TESTING. REVIEW OF INSTRUMENT IMAGES: SAMPLE DISPLAYED SLIGHT RED CELL ADHERENCE WITH A TIGHT BUTTON. THAT COULD BE INTERPRETED AS A NEGATIVE REACTION OR WEAKLY REACTIVE. ROI DID NOT APPEAR TO NEED ADJUSTMENT. AUTOMATIC CAMERA ADJUSTMENT PERFORMED. CUSTOMER HAS NOT REPORTED ANY FURTHER ISSUES OF THIS NATURE. | 0
2276194 | IT WAS REPORTED THAT WHILE ACTIVELY MONITORING A PT, THIS M300 APPEARED TO DISCHARGE THE PT WITH NO USER INTERVENTION. PT DATA STOPPED BEING DISPLAYED AT THE ICS AND A 'BED DISCONNECTED' MESSAGE APPEARED AT THE ICS FOR THIS M300. THE PT WAS READMITTED TO THE ICS AND THERE WAS NO PT INJURY REPORTED. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATION THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2279233 | CUSTOMER REPORTS IMAGES THAT WERE DELETED FROM STUDIES WERE REAPPEARING IN IMAGECAST PACS. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. (B)(4). | 1
2291270 | IT WAS REPORTED WHEN SEARCHING FOR PATIENTS IN MMP (MANAGE MY PATIENT) PAGE, IT WAS VERY SLOW AND SOMETIMES NEVER RETURNED ANY RESULTS. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2291990 | ADDITIONAL INFORMATION WAS RECEIVED THAT INDICATED THAT PRODUCT ANALYSIS HAS BEEN COMPLETED ON THE HANDHELD AND FLASHCARD. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE REPORTED ALLEGATION OF A SCREEN RESPONSE ISSUE WAS VERIFIED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE TOUCHSCREEN DISPLAY WAS DEFECTIVE. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH A RESISTANCE VALUE BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. SINCE THE TOUCHSCREEN DISPLAY USES RESISTANCE VALUES TO DETERMINE THE COORDINATES OF A SCREEN TAP, THE ADDITIONAL RESISTANCE ASSOCIATED WITH PIN 1 CAUSED THE DISPLAY TO BE UNRESPONSIVE. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS INITIALLY REPORTED THAT THE PHYSICIAN HANDHELD WAS NOT RESPONDING TO TAPS ON THE SCREEN. A HARD RESET WAS ATTEMPTED, BUT WENT TO A BLANK SCREEN AND WOULD NOT PROCEED. ANOTHER HARD RESET WAS ATTEMPTED, BUT STILL DID NOT COMPLETE. ON THE THIRD ATTEMPT THE HARD RESET WENT THROUGH, BUT STILL WOULD NOT RESPOND TO TAPS ON THE SCREEN. IT WAS CONFIRMED THAT THE LOCK BUTTON WAS NOT ENGAGED AND HERE WAS NO DEBRIS ON THE SCREEN THAT COULD BE INTERFERING WITH DETECTION FOR THE SCREEN. THE PHYSICIAN WAS PROVIDED ANOTHER HANDHELD AND THE SUSPECT HANDHELD WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. PRODUCT ANALYSIS IS PLANNED BUT HAS NOT BEEN COMPLETED. .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2293487 | IT WAS REPORTED THAT THE CLIENT WAS UNABLE TO LOCATE PATIENT INFORMATION IN THE APPLICATION. THE PATIENT'S REMOTE MONITOR HAD SUCCESSFULLY SENT TRANSMISSIONS PREVIOUSLY, AS RECENTLY AS SEVERAL MONTHS AGO. FURTHER INVESTIGATION INDICATED THAT THE PATIENT HAD BEEN DISCONTINUED AND THERE IS NO WAY TO DETERMINE WHO HAD DISCONTINUED THE PATIENT OR WHEN THIS HAD HAPPENED. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. | 1
2293499 | IT WAS REPORTED THAT A REMOTE INTERROGATION OF THE PATIENT'S SINGLE CHAMBER DEVICE REPORTED AN INTERLOCK WITH MODE SWITCH, VENTRICULAR CAPTURE MANAGEMENT, AND "SLEEP." ALSO REPORTED, THIS "MESSAGE SHOULD NOT BE SEEN" AS THIS IS A SINGLE CHAMBER DEVICE. THE DEVICE REMAINS IN USE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2297235 | THE CUSTOMER CONTACTED BECKMAN COULTER INC (BEC) TO REPORT A LEAK OF BLOOD AND DILUENT INSIDE THE DXH800 INSTRUMENT. NO REPORT OF ANY PATIENT RESULTS OR TREATMENT BEING AFFECTED BY THIS INCIDENT. THE USER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES, A LAB COAT AND GLASSES AT THE TIME OF THE INCIDENT. NO INJURIES OCCURRED, NO EXPOSURE (SPLASHED OR SPAYED) TO MUCOUS MEMBRANE OR OPEN WOUNDS WAS REPORTED, AND MEDICAL ATTENTION WAS NOT SOUGHT. ON THE DAY OF THE EVENT, A BEC FIELD SERVICE ENGINEER (FSE) FOUND THAT THE PROBE WASH COLLAR WAS LEAKING BLOOD AND DILUENT INSIDE THE INSTRUMENT. THE FSE REMOVED AND CLEANED THE COLLAR WITH 50% BLEACH, PERFORMED PROBE WASTE CHAMBER FLUSH WITH 50% BLEACH. THE REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES. ROOT CAUSE FOR THE LEAK IS ATTRIBUTED TO THE PROBE WASH COLLAR WAS CLOGGED WITH RESIDUAL BLOOD. (B)(4). | 0
2297236 | .><THE CUSTOMER CONTACTED BECKMAN COULTER INC(BEC) TO REPORT A CLENZ LEAK (2ML) IN THE DRIP TRAY UNDERNEATH THE ABC/RBC BATHS OF THE DXH 800 INSTRUMENT. THE LEAK DID NOT IMPACT PATIENT RESULTS OR TREATMENT. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A GOWN, GLOVES AND GOGGLES AT THE TIME OF THE INCIDENT. NO INJURIES OCCURRED, NO EXPOSURE (SPLASHED OR SPAYED) TO MUCOUS MEMBRANE OR OPEN WOUNDS WAS REPORTED, AND MEDICAL ATTENTION WAS NOT SOUGHT. ON THE DAY OF THE EVENT, A BEC FIELD SERVICE ENGINEER (FSE) NOTED THAT THE CUSTOMER HAD UNPLUGGED THE NRBC CHAMBER BOTTOM PORT. THE FSE PLUGGED THE CHAMBER PORT. VERIFIED REPAIR PER ESTABLISHED PROCEDURES. ROOT CAUSE WAS ATTRIBUTED TO THE CHAMBER PORT BEING UNPLUGGED. (B)(4).><(B)(4). | 0
2298359 | UNDER CERTAIN CIRCUMSTANCES, WHEN USING NICOLETONE WITH THE NICOLET CORTICAL STIMULATOR, IT IS POSSIBLE THAT EVEN THOUGH STIMULATION WAS DELIVERED THROUGH THE SELECTED ELECTRODES, THE ANNOTATION LABEL WILL INCORRECTLY INDICATED THAT THE STIMULATION WAS DELIVERED VIA A DIFFERENT ELECTRODE. A VOLUNTARY FIELD CORRECTION AND REMOVAL, 806.10 WAS INITIATED (B)(4) 2011 TO THE (B)(6) OFFICE, REPORT NUMBER: (B)(4). | 0
2300693 | THE CUSTOMER STATES THAT THE ORTHO PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND ADJUSTED THE READER CAMERA. THE FE REPLACED THE GRIPPER AND SOLENOID. THE FE ALSO PERFORMED A PM AND VERIFIED MOD #(B)(4). THE CUSTOMER RAN QC AND ALL WERE WITHIN SET LAB RANGE. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
2301968 | CUSTOMER REPORTED FALSE POSITIVE REACTION CAUSED BY CARRYOVER WHEN RUNNING CAP SURVEY SAMPLES IN THE FOLLOWING ORDER: JAT15, JAT16, JAT17, JAT18, JAT19 ON TANGO OPTIMO DEVICE. ACCORDING TO CUSTOMER JAT15 REACTED 4+. JAT16 TESTED +/- IN THE RUN. IF RUN INDEPENDENTLY, JAT16 TESTED NEGATIVE. A SERVICE VISIT WAS CONDUCTED TO CHECK CAMERA VALUES. ADDITIONALLY A QC RUN WAS PERFORMED. CAMERA VALUES WERE IN RANGE, BUT HAVE BEEN SLIGHTLY ADJUSTED . QUALITY CONTROLS RAN SUCCESSFULLY WITH ALL RESULTS IN RANGE. NO INDICATION FOR ANY INSTRUMENT MALFUNCTION WAS OBSERVED BASED ON THE INFORMATION OBTAINED DURING THE SERVICE VISIT WE DO NOT SEE A CONTRIBUTION OF THE DEVICE ITSELF TO THE REPORTED PROBLEM. ALTHOUGH THE FINAL TITER OF JAT15 IS UNKNOWN, THE SCENARIO OF A CARRYOVER FROM A STRONG POSITIVE (4+) SAMPLE TO A NEGATIVE ONE IS PLAUSIBLE AND CANNOT BE PREVENTED IF THE FINAL ANTIBODY TITER EXCEEDS A CERTAIN AMOUNT. THE CORRECT FUNCTION OF THE ALLEGEDLY DEFECTIVE REAGENTS WERE PROVED BY TESTING THE RETAINED SAMPLES OF QUALITY CONTROL LABORATORY ON TANGO OPTIMO DEVICE. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DIDN´T OBSERVE ANY FALSE POSITIVE REACTIONS. THE REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE NEGATIVELY AFFECTED THE QUALITY OF THE COMPLAINED LOT. AS THE INVESTIGATION IS CURRENTLY ONGOING FURTHER RESULTS WILL BE REPORTED WHEN AVAILABLE.><THE CUSTOMER REPORTED A FALSE POSITIVE REACTION CAUSED BY CARRYOVER WHEN RUNNING CAP SURVEY SAMPLES IN THE FOLLOWING ORDER: JAT15, JAT16, JAT17, JAT18, JAT19 ON TANGO OPTIMO DEVICE. ACCORDING TO CUSTOMER JAT15 REACTED 4+. JAT16 TESTED +/- IN THE RUN. IF RUN INDEPENDENTLY, JAT16 TESTED NEGATIVE. A SERVICE VISIT WAS CONDUCTED TO CHECK CAMERA VALUES. ADDITIONALLY A QC RUN WAS PERFORMED. CAMERA VALUES WERE IN RANGE, BUT HAVE BEEN SLIGHTLY ADJUSTED . QUALITY CONTROLS RAN SUCCESSFULLY WITH ALL RESULTS IN RANGE. NO INDICATION FOR ANY INSTRUMENT MALFUNCTION WAS OBSERVED BASED ON THE INFORMATION OBTAINED DURING THE SERVICE VISIT WE DO NOT SEE A CONTRIBUTION OF THE DEVICE ITSELF TO THE REPORTED PROBLEM. ALTHOUGH THE FINAL TITER OF JAT15 IS UNKNOWN, THE SCENARIO OF A CARRYOVER FROM A STRONG POSITIVE (4+) SAMPLE TO A NEGATIVE ONE IS PLAUSIBLE AND CANNOT BE PREVENTED IF THE FINAL ANTIBODY TITER EXCEEDS A CERTAIN AMOUNT. THE CORRECT FUNCTION OF THE ALLEGEDLY DEFECTIVE REAGENTS WERE PROVED BY TESTING THE RETAINED SAMPLES OF QUALITY CONTROL LABORATORY ON TANGO OPTIMO DEVICE. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DIDN´T OBSERVE ANY FALSE POSITIVE REACTIONS. THE REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE NEGATIVELY AFFECTED THE QUALITY OF THE COMPLAINED LOT. THIS IS OUR FINAL REPORT FOR THIS INCIDENT. | 0
2303569 | IT WAS REPORTED THAT DURING AN A-FIB PROCEDURE, ERROR 1006 WAS NOTED ON THIS CARTO 3 SYSTEM, "BACK PATCH SENSOR IS DISCONNECTED". THE ON-CALL FIELD SERVICE ENGINEER WAS CONSULTED FOR FURTHER ASSISTANCE. HOWEVER, THE ERROR PERSISTED. THIS ISSUE COULD NOT BE RESOLVED OVER THE PHONE. THE CASE WAS COMPLETED USING NON-BWI DEVICE/ SYSTEM (MANUFACTURED BY ST JUDE MEDICAL). THE MALFUNCTION ITSELF WAS NOT INDICATIVE OF A REPORTABLE EVENT. THE PATIENT WAS INTUBATED AND WAS UNDER GENERAL ANESTHESIA; AND THE PROCEDURE WAS PROLONGED BY APPROXIMATELY 1.5-2 HOURS. ATTEMPTS HAVE BEEN MADE TO OBTAIN THE PATIENT'S BASELINE CONDITION (SUCH AS AGE, GENDER, AND CHRONIC HEALTH CONDITION OR MAJOR MEDICAL HISTORY), HOWEVER, NO DETAILED INFORMATION WAS PROVIDED. (B)(4). THE COMPLAINT REPORTED WAS REGARDING ERROR 1006 NOTED ON THE CARTO 3 SYSTEM: "BACK PATCH SENSOR IS DISCONNECTED" THAT RESULTED IN PROCEDURE PROLONGATION OF 1.5-2 HRS. THE ISSUE WAS NOT RESOLVED DURING THE CALL SUPPORT. THE CARTO 3 SYSTEM ASSOCIATED WITH THE REPORTED COMPLAINT WAS THOROUGHLY INSPECTED BY BWI SERVICE ENGINEER AND CONFIRMED THE REPORTED ISSUE. BACK PATCH #2 WAS FOUND FAULTY AND REPLACED WITH A NEW ONE. THE SYSTEM WAS TESTED AND FOUND READY FOR USE. THE DEFECTIVE BACK PATCH SENSOR CABLE CAUSED THE ISSUE. THE SYSTEM CORRECTLY DISPLAYED THE RELATED ERROR MESSAGE ABOUT THE PROBLEM. CARTO 3 USERS ARE INFORMED ABOUT THE POTENTIAL FAILURE OF CARTO 3 IN THE INSTRUCTION FOR USE. IN ADDITION, THE DEVICE HISTORY RECORD ASSOCIATED WITH CARTO 3 SYSTEM SERIAL # (B)(4) WAS REVIEWED AND THERE WAS NO DISCREPANCY NOTED. AS FOR THE COMPLAINT ISSUE ITSELF, A CORRECTIVE ACTION HAS BEEN OPENED TO ADDRESS AND RESOLVE THIS ISSUE.><INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. (B)(4). | 1
2304709 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2011, WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE HANDHELD AND FLASHCARD. PRODUCT ANALYSIS ON THE HANDHELD REVEALED NO ANOMALIES ASSOCIATED WITH THE HANDHELD DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE ALLEGED SCREEN FREEZE COMPLAINT IS A KNOWN ISSUE THAT HAS BEEN EVALUATED AND ADDRESSED. THIS EVENT ONLY OCCURS ON DELL X5 MODELS OF THE HANDHELD COMPUTERS EXECUTING THE POCKET PC 2002 OPERATING SYSTEM. THIS EVENT CAN BE RESOLVED FOLLOWING A SOFTWARE RESET OF THE HANDHELD OR RESEATING THE FLASHCARD AND MAY TYPICALLY OCCUR AFTER SEVERAL DAYS OF INACTIVITY OR AFTER A DAYLIGHT SAVINGS TIME CHANGE. NO OTHER ISSUES WERE OBSERVED WITH EITHER THE FLASHCARD OR SOFTWARE AND OTHER THAN THE SCREEN FREEZE OBSERVATIONS, THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(6) 2011, A VNS TREATING PHYSICIAN REPORTED THAT HIS HANDHELD'S SCREEN FROZE AT THE INTERROGATION SUCCESSFUL SCREEN AND THAT IT HAS HAPPENED SEVERAL TIMES IN THE PAST. THE PHYSICIAN DID A HARD RESET HIMSELF WHICH RESOLVED THE PROBLEM AND HE WAS ABLE TO INTERROGATE THE PATIENTS AFTERWARDS WITHOUT HAVING THE SCREEN FREEZE AGAIN. THE PHYSICIAN STILL WANTED A NEW HANDHELD THOUGH. THE PHYSICIAN'S HANDHELD AND FLASHCARD WERE RECEIVED FOR PRODUCT ANALYSIS ON (B)(6) 2011 THAT HAS NOT YET BEEN COMPLETED. . | 0
2313663 | IT WAS REPORTED THAT WHILE ACTIVELY MONITORING A PT, THE M300 APPARENTLY STOPPED TRANSMITTING PT DATA TO THE ASSOCIATED ICS AT APPROX 14:09 HOURS. THE USER REPORTEDLY DID NOT NOTICE IF THE ICS DISPLAYED A 'BEDSIDE OFFLINE' MESSAGE AT THE TIME. IT WAS REPORTED THAT AT APPROX 15:12 HRS, A USER NOTICED THAT THE ICS WAS DISPLAYING A 'BED DISCONNECTED' MESSAGE FOR THIS M300. THE M300 WAS REPORTEDLY READMITTED TO THE ICS AT APPROX 15:25 HRS. THERE WAS NO PT INJURY REPORTED. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2315430 | DURING THE STANDARD USE OF THE UNIT, WHILE IN THE PARAFFIN-WAX STEP, PARAFFIN HEATED TO A TEMPERATURE THAT MELTED THE FORTY-FIVE CASSETTES. AS A CONSEQUENCE, THE BIOPSY TISSUE INSIDE THE CASSETTES WERE DAMAGED AND A SATISFACTORY DIAGNOSIS WAS NOT POSSIBLE. THE UNIT WAS OPERATING ON A POWER SUPPLY SOURCE AND GROUND CONDITION NOT ADEQUATE BASED ON THE PROVIDED POWER REQUIREMENTS DESIGNATED ON THE PRE-INSTALLATION GUIDE AND USER MANUAL PROVIDED TO THE USER. | 0
2318302 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THE COULTER LH 750 HEMATOLOGY ANALYZER WAS GIVING NO RETIC RESULTS (RETIC VOTEOUT/INCOMPLETE COMPUTATION) AND ++++ FOR RETIC BACKGROUNDS (VALUES THAT ARE ABOVE THE OPERATING RANGE ARE INHIBITED, AND THE VALUE IS REPLACED BY PLUSES +++++). THE CUSTOMER ALSO FOUND A LEAK OF ABOUT 5ML CLEAR FLUID NEAR THE MANUAL SAMPLING PROBE. THE CUSTOMER COULD NOT LOCATE THE SOURCE OF THE LEAK. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED AND MEDICAL ATTENTION WAS NOT SOUGHT. NO PATIENT RESULTS WERE AFFECTED BY THIS EVENT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE (B)(4) 2011 AND FOUND THE INSTRUMENT GIVING RETIC VOTEOUTS AND FOUND THE PINCH VALVE PV53B BORKEN THEREBY PREVENTING THE RETIC CHAMBER FROM VENTING WHEN IT'S RUNNING THE SAMPLE. FSE REPLACED THE PINCH VALVE AND PERFORMED A WORKSTATION SOFTWARE UPGRADE. THE ROOT CAUSE FOR THE LEAK IS ATTRIBUTED TO THE BROKEN PINCH VALVE PV53B.(B)(4). | 0
2320052 | THE CUSTOMER REPORTED THAT DURING AN ANGIOGRAPHIC PROCEDURE THE ROTATOR BROKE DURING INJECTION AT 900 PSI. THE CUSTOMER REPORTED THAT THIS OCCURRED FOUR TIMES. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF FOUR REPORTS FOR THIS COMPLAINT: 1721504-2011-00312, 00313, 00315. DEVICE EVALUATION: THE SUSPECT DEVICE IS NOT EXPECTED TO BE RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. MERIT IS UNABLE TO DETERMINE THE ROOT CAUSE OF THE REPORTED FAILURE WITHOUT THE SUSPECT DEVICE. NO CONCLUSION CAN BE DRAWN. NO TESTING METHODS PERFORMED, PROCESS EVALUATION. | 0
2328394 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS NOTED THE PATIENT'S VNS WAS INTERROGATED ON (B)(6) 2005 AND FOUND TO BE AT SYSTEMS DIAGNOSTICS DEFAULT SETTINGS OF 1MA/20HZ/500 PULSEWIDTH/30 SEC ON/60 MIN OFF. THE SETTINGS WERE CORRECTED ON THIS DAY. THERE IS NO PROGRAMMING HISTORY PRIOR TO (B)(6) 2005. THE PATIENT WAS IMPLANTED WITH A NEW VNS GENERATOR ON (B)(6) 2005. IT IS SUSPECTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED AT THE IMPLANT PROCEDURE WITH AN UNKNOWN VNS PROGRAMMING SYSTEM ON (B)(6) 2005 AND THE SETTINGS WERE NOT CORRECTED UNTIL (B)(6) 2005. ANALYSIS OF PROGRAMMING HISTORY. | 0
2331604 | A CUSTOMER CONTACTED BAXTER GLOBAL TECHNICAL SERVICES (GTS) REGARDING ASSISTANCE WITH THE HOMECHOICE (HC) UNIT DURING INITIAL DRAIN. PER THE INITIAL REPORT, HOME PATIENT (HP) HAD A SYSTEM ERROR 2240 (AIR IN THE CASSETTE). THE TECHNICAL SERVICE REPRESENTATIVE (TSR) EXPLAINED THE ALARMS AND THE HP STATED SHE PRESSED GO AND WAS NOT CONNECTED. THE HP THEN CONNECTED TRYING TO RESTART THERAPY AFTER RECEIVING THE SYSTEM ERROR 2240 ALARM. THE TSR EXPLAINED TO DISCARD ALL THE SUPPLIES THE HP WOULD FINISH THE NIGHT'S THERAPY MANUALLY. GLOBAL PRODUCT SURVEILLANCE CONTACTED HP ON (B)(6) 2011. THE HP WAS ABLE TO COMPLETE THERAPY WITH NEW SUPPLIES. THE HP DID NOT NOTICE ANY DEFECTS WITH THE ORIGINAL SUPPLIES. THE HP STATED THAT IT WAS HER FAULT THAT SHE PRESSED GO BEFORE CONNECTING TO THE HC. THERE WAS NO PATIENT INJURY OR MEDICAL INTERVENTION REPORTED. (B)(4). THIS COMPLAINT FOR A SYSTEM ERROR (SE) 2240 (AIR IN SET) OCCURRED DURING INITIAL DRAIN IS CONFIRMED BECAUSE THE PATIENT REPORTED PRESSING GO PRIOR TO CONNECTING HIMSELF, WHICH IS A USE ERROR THAT CAN CAUSE SYSTEM ERROR 2240 ALARMS. THE HOME PATIENT (HP) STATED THAT IT WAS HER FAULT THAT SHE PRESSED GO BEFORE CONNECTING TO THE HC. THE LABEL REVIEW FOUND THE PATIENT AT HOME GUIDE TO BE ADEQUATE FOR THE USE ERROR IN THIS INCIDENT. BAXTER HAS CONDUCTED A TREND REVIEW AND FOUND THAT SIMILAR REPORTS HAVE BEEN RECEIVED FOR THE REPORTED PROBLEM. THIS IS BEING INVESTIGATED THROUGH CAPA. THE ASSIGNABLE CAUSE FOR THE REPORTED PROBLEM WAS USE ERROR.><(B)(4). PER THE CUSTOMER THE SAMPLE WAS DISCARDED AND THE LOT NUMBER WAS UNKNOWN, THEREFORE NO EVALUATION OR BATCH REVIEW COULD BE PERFORMED. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THEN A FOLLOW UP MDR WILL BE SUBMITTED. | 0
2333316 | A CUSTOMER CONTACTED BAXTER GLOBAL TECHNICAL SERVICES (GTS) REGARDING ASSISTANCE WITH THE HOMECHOICE (HC) UNIT DURING DRAIN 2 OF 10. PER THE INITIAL REPORT, HOME PATIENT (HP) HAD A CHECK DRAIN LINE ALARM WERE THE DRAIN LINE WAS SUBMERGED IN THE DRAIN FLUID. THE TECHNICAL SERVICE REPRESENTATIVE (TSR) DETERMINED THAT THE DRAIN LINE SUBMERGED IN FLUID. THE TSR HAD THE HP DISCONNECT THE EXTENSION AND MOVE THE BUCKET IN THE ROOM WITH HER. GLOBAL PRODUCT SURVEILLANCE CONTACTED THE PERITONEAL DIALYSIS NURSE (RN) ON (B)(6) 2011. THE RN STATED THAT THE HP WAS DOING FINE AND COMPLETING THERAPY. THERE WAS NO PATIENT INJURY OR MEDICAL INTERVENTION REPORTED. (B)(4). PER THE CUSTOMER THE SAMPLE WAS DISCARDED AND THE LOT NUMBER WAS UNKNOWN, THEREFORE NO EVALUATION OR BATCH REVIEW COULD BE PERFORMED. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THEN A FOLLOW UP MDR WILL BE SUBMITTED.><(B)(4). THIS COMPLAINT WAS FOR A REPORT OF A CHECK DRAIN LINE ALARM. THE ALARM RESOLVED WHEN THE DRAIN LINE EXTENSIONS WERE REMOVED BUT IT IS UNKNOWN WHAT CAUSED THE ALARM. DRAIN LINE SUBMERGED IN FLUID IS NOT KNOWN TO CAUSE AN ALARM. THE COMPLAINT WAS NOT CONFIRMED AND THE CAUSE WAS NOT IDENTIFIED BECAUSE THE SAMPLE WAS NOT RETURNED FOR EVALUATION AND THE LOT NUMBER WAS NOT REPORTED. THE ASSIGNABLE CAUSE FOR THE REPORTED PROBLEM WAS UNDETERMINED. | 0
2333748 | ON (B)(6) 2011, THE VNS PATIENT'S PROGRAMMING HISTORY WAS REVIEWED AND IT WAS DISCOVERED THAT A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED ON (B)(6) 2008 WHICH INADVERTENTLY CHANGED THE PATIENT'S SETTINGS FROM OUTPUT=2MA/FREQUENCY=30HZ/PULSE WIDTH=250USEC/ON TIME=30SEC/OFF TIME=1.8MIN/MAGNET OUTPUT=2.25MA/MAGNET PULSE WIDTH=250USEC/MAGNET ON TIME=60SEC TO INCORRECT SETTINGS OF OUTPUT=1MA/FREQUENCY=20HZ/PULSE WIDTH=500USEC/ON TIME=30SEC/OFF TIME=60MIN/MAGNET OUTPUT=1MA/MAGNET PULSE WIDTH=500USEC/MAGNET ON TIME=30SEC. THE PATIENT'S SETTINGS WERE NOT FOUND AND CORRECTED UNTIL THEIR NEXT VISIT ON (B)(6) 2009. ANALYSIS OF PROGRAMMING HISTORY. | 0
2333788 | WHILE REVIEWING THE PROGRAMMING HISTORY FOR THE PATIENT IT WAS NOTICED THAT THE PATIENT CAME INTO AN APPOINTMENT ON (B)(6) 2010 SET TO AN OFFICE TIME OF 60 MINUTES. IT IS UNKNOWN IF THE PATIENT WAS INTENTIONALLY PROGRAMMED TO THOSE SETTINGS OR OF IT WAS A RESULT OF AN INCOMPLETE DIAGNOSTICS THAT RESULTED IN A SETTING CHANGE THAT WAS ONLY PARTIALLY CORRECTED. THE OFF TIME WAS CORRECTED WHEN IT WAS SEEN ON INTERROGATION. DATE OF THIS REPORT, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY LISTED THE WRONG YEAR FOR THE DATE THAT THE REPORT WAS RECEIVED BY THE MANUFACTURER. DATE RECEIVED BY MANUFACTURER (MO/DAY/YR), CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY LISTED THE WRONG YEAR FOR THE DATE THAT THE REPORT WAS RECEIVED BY THE MANUFACTURER.><ANALYSIS OF PROGRAMMING HISTORY. | 0
2334957 | IT WAS REPORTED THAT AN INTERMITTENT LOSS OF TELEMETRY MONITORING OCCURRED FOR UP TO 30 ROOMS SPREAD ACROSS 5 FLOORS FOR UP TO AN HOUR AT A TIME. LIMITED ALTERNATIVE PHYSIOLOGICAL MONITORING IS AVAILABLE. THERE WERE NO PATIENT INJURIES REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 1
2339247 | IT WAS REPORTED THAT THE DEVICE UNEXPECTEDLY SHUT DOWN WHILE IN USE. REBOOTING THE DEVICE DID NOT RESOLVE THE REPORTED PROBLEM. THERE WAS NO PT INJURY REPORTED. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2341607 | ANALYSIS WAS COMPLETED ON THE RETURNED HANDHELD AND SOFTWARE. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE (B)(4) DATABASE WAS CORRUPT CAUSING THE HANDHELD TO DISPLAY SQL ERRORS WHEN AN INTERROGATION WAS PERFORMED. A CAUSE FOR THE DATABASE CORRUPTION COULD NOT BE DETERMINED BASED ON THE AVAILABLE INFORMATION. IT WAS ALSO IDENTIFIED THAT SEVERAL ARCHIVE DATABASES WERE ALSO CORRUPT. BASED ON THE DATABASE SIZE, THE ARCHIVE DATABASES ARE LIKELY ARCHIVE COPIES OF THE (B)(4) DATABASE. NO FURTHER ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS.><IT WAS REPORTED THAT THE PHYSICIAN COULD NOT ACCESS THE DATABASE IN THE HANDHELD COMPUTER ANYMORE. UPON ATTEMPTING TO VIEW THE DATABASE, THE COMPUTER RETURNS THE MESSAGE "ERROR EXECUTING THE SQL STATEMENT". A HARD AND SOFT RESET WERE ATTEMPTED, BUT THE PROBLEM PERSISTED. THE HANDHELD WAS RETURNED TO THE MANUFACTURER, BUT ANALYSIS IS PENDING. . | 0
2344349 | IT WAS INITIALLY REPORTED THAT THERE WAS A PACING SPIKE DURING RF ABLATION (IN AN AVNRT PROCEDURE) ON THE CARTO 3 SYSTEM. THERE WAS ALSO UNKNOWN INTERFERENCE ON THE SYSTEM DURING THIS TIME. THE CALLER DECLINED TROUBLESHOOTING. THE PACING SPIKE MIGHT BE A REAL STIMULATION. IT APPEARED THAT THE STIMULUS WAS DELIVERED FROM THE DISTAL ABLATION CATHETER WHILE ABLATING. THE PACING SPIKE ONLY HAPPENED ONCE AND THERE WAS NO PATIENT CONSEQUENCE. THE CASE WAS COMPLETED SUCCESSFULLY. UPON FOLLOW UP, BWI RECEIVED THE INFORMATION ON (B)(4), 2011 (IT CHANGED THE ALERT DATE) WHICH SHOWED IT APPEARED AN UNINTENDED PACING CAPTURED BEAT DURING ABLATION FROM THE ABLATION CATHETER, WHICH COULD CAUSE POTENTIAL RISK OF INDUCING VF TO THE PATIENT. (B)(4).><(B)(4). BIOSENSE FIELD SERVICE ENGINEERS PERFORMED THE RELEVANT TESTS. THE TESTS ARE PASSED. THE SYSTEM WAS FOUND OPERATIONAL. THE REPORTED ISSUE WAS NOT CONFIRMED. AN ADDITIONAL ORIGINAL EXTERNAL MANUFACTURER (OEM) MANUFACTURER ACTION (DHR REVIEW OR INVESTIGATION) WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. | 1
2348781 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2011, WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE FLASHCARD AND HANDHELD. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE PROBLEM THAT THE PHYSICIAN REPORTED REGARDING THE HANDHELD SHUTTING OFF WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS AND THE REMAINING BATTERY CHARGE LEVEL AFTER OVER 1 HOUR OF TESTING WAS 67% OF FULL CAPACITY.><ON (B)(6) 2011, A VNS TREATING PHYSICIAN REPORTED THAT HE WAS HAVING INTERMITTENT PROBLEMS WITH HIS HANDHELD AND WANTED A NEW ONE RIGHT AWAY. HE STATED THAT THE HANDHELD KEEPS SHUTTING OFF BUT HE REFUSED TO DO ANY TROUBLESHOOTING. THE PHYSICIAN'S HANDHELD AND FLASHCARD WERE RECEIVED BY THE MANUFACTURER FOR PRODUCT ANALYSIS ON (B)(6) 2011 THAT HAS NOT YET BEEN COMPLETED. . | 0
2351638 | A CUSTOMER CONTACTED BAXTER TECHNICAL SERVICES(BTS) REGARDING ASSISTANCE WITH A CHECK LINES AND BAGS ALARM DURING THE PRIME ON THE HOMECHOICE MACHINE(HC). THE TECHNICAL SERVICE REPRESENTATIVE(TSR) ASKED THE HOME PATIENT(HP) IF THEY STARTED OVER WITH ALL NEW SUPPLIES FROM EARLIER PROBLEMS AND THE HP REPLIED HE ONLY STARTED OVER WITH A NEW CASSETTE. THE TSR ADVISED THE HP THAT HE SHOULD HAVE STARTED OVER WITH A NEW CASSETTE AND BAGS. THE CARE GIVER (CG) WAS CONTACTED ON (B)(6) 2011 REGARDING THE REPORTED PROBLEM. THE CG STATED THAT THE HP'S THERAPY IS GOING WELL WITHOUT ANY OTHER PROBLEMS. THEY STARTED OVER WITH NEW SUPPLIES AND EVERYTHING WENT WELL. NO FURTHER INFORMATION WAS PROVIDED REGARDING THE REPORTED PROBLEM. THERE WAS PATIENT INVOLVEMENT BUT NO PATIENT INJURY OR MEDICAL INTERVENTION INDICATED AT THE TIME OF THE INITIAL REPORT. (B)(4). THIS COMPLAINT FOR A REPORT OF USE ERROR, REUSE OF SINGLE-USE SUPPLIES WAS CONFIRMED. THE ROOT CAUSE WAS UNDETERMINED. THE LABEL REVIEW FOUND THE LABELING ADEQUATE FOR THE USE ERROR IN THIS COMPLAINT. BAXTER HAS CONDUCTED A TREND REVIEW AND FOUND THAT SIMILAR REPORTS HAVE BEEN RECEIVED FOR THE REPORTED PROBLEM. BAXTER WILL CONTINUE TO MONITOR SIMILAR REPORTS TO DETERMINE IF FURTHER ACTIONS ARE REQUIRED. A FOLLOW-UP MDR WILL BE SUBMITTED IF ANY ADDITIONAL INFORMATION IS RECEIVED.><(B)(4).THERE WAS NO ALLEGATION REPORTED AGAINST THE BAXTER PRODUCT BY THE CUSTOMER. THEREFORE THE DEVICE WAS NOT REQUESTED FOR EVALUATION AND A BATCH REVIEW WILL NOT BE CONDUCTED. A FOLLOW-UP REPORT WILL BE FILED UPON COMPLETION OF BAXTER'S INVESTIGATION, OR IF ANY ADDITIONAL INFORMATION IS RECEIVED. | 0
2352167 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT FLUID WAS LEAKING FROM THE CBC TRANSDUCER PANEL ON THE UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM. THE USER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES, A LAB COAT AND GLASSES AT THE TIME OF THE INCIDENT. THERE WAS NO INJURY OR EXPOSURE REPORTED AND MEDICAL ATTENTION WAS NOT SOUGHT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THAT THE LEAK WAS COMING FROM THE DIFFERENTIAL MIXING CHAMBER IN THE AMTC (AIR MIX AND TEMPERATURE CONTROL) MODULE AND NOT FROM THE CBC TRANSDUCER AS ORIGINALLY REPORTED BY THE CUSTOMER. FSE REMOVED THE DEBRIS FROM THE DIFFERENTIAL MIXING CHAMBER DRAIN FITTING WHICH RESOLVED THE ISSUE. THE ROOT CAUSE FOR THE LEAK IS ATTRIBUTED TO THE DEBRIS IN THE DIFFERENTIAL MIXING CHAMBER DRAIN FITTING. (B)(4). | 0
2354834 | THE PT HAD A PLATELET DEPLETION DURING A THERAPEUTIC PLASMA EXCHANGE (TPE). THE PT INFO IS UNAVAILABLE AT THIS TIME. THE PT OUTCOME IS UNAVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFO AT THIS TIME TO DETERMINE IF A MALFUNCTION OCCURRED.><THE PATIENT'S IDENTIFICATION AND AGE WERE NOT PROVIDED BY THE CUSTOMER DUE TO PRIVACY. THE CUSTOMER STATED THAT THE PATIENT WAS AN ADULT. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. INVESTIGATION: THE CUSTOMER HAS ALLEGED A PLATELET LOSS OF 55%. CLOTTING IN THE DISPOSABLE SET WOULD BE TOO SEVERE TO ALLOW THE PROCEDURE TO COMPLETE DUE TO CELLULAR OBSTRUCTIONS WITHIN THE FLUID PATHWAY TO ACCOUNT FOR THIS MAGNITUDE OF PLATELET LOSS. A DROP THIS SEVERE WOULD REQUIRE THE LOSS OF SUBSTANTIAL QUANTITY OF PLATELETS TO THE WASTE BAG, WHICH WOULD BE CLEARLY VISIBLE AS CLOUDY SWIRL IN THE PLASMA. ALL BLOOD COMPONENTS NOT BEING REMOVED TO THE PLASMA WASTE BAG ARE RETURNED TO THE PATIENT. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. DISPOSABLE LOTS 05T3331 AND 08T3222 COMPRISE 92% OF THE OPTIA EXCHANGE KITS DELIVERED TO THIS CUSTOMER SITE IN THE 7 WEEKS PRIOR TO THE EVENT. THE DEVICE HISTORY RECORDS (DHRS) WERE REVIEWED FOR THESE LOTS. NOTHING WAS FOUND IN THE DHRS THAT WOULD HAVE CONTRIBUTED TO THE EVENT EXPERIENCED BY THE CUSTOMER. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. ANALYSIS OF THE RUN DATA LOG FOR THIS OCCURRENCE SHOWS NO INDICATIONS THAT PLATELET LOSS COULD HAVE TAKEN PLACE. CORRECTIVE/PREVENTIVE ACTION: THE CUSTOMER BEGAN AN INTERNAL INVESTIGATION INTO THE MANNER IN WHICH THEY PERFORM PLATELET PRE COUNTS AND THE TIMING RELATED TO THE PROCEDURE.><(B)(4). INVESTIGATION: THERE WERE NO PLATELETS IN THE WASTE BAG WHICH IS TYPICAL IN OPTIA TPE PROCEDURES. THERE WERE NO CLUMPS EVIDENT WITHIN THE DISPOSABLE FOLLOWING RINSEBACK, WHICH WAS COMPLETED NORMALLY. THE RUN DATA LOG SHOWS NO SPILLOVER EVENTS. THE AIM SYSTEM ALERTED THE OPERATOR THAT THE INTERFACE WAS NOT COMING UP AND THE OPERATOR CHOSE NOT TO ADJUST THE PT'S HEMATOCRIT, BUT RATHER CONTINUED THROUGH THE ALARM. THIS RESULTED IN THE INTERFACE RUNNING QUITE LOW, KEEPING THE PLATELETS WELL AWAY FROM THE WASTE EXTRACTION LINE. PER AN ON-SITE INVESTIGATION, THERE IS NO POSSIBLE WAY THAT THE SPECTRA OPTIC MACHINE CONTRIBUTED TO THE REDUCTION IN PERIPHERAL PLATELET COUNT. THERE ARE NO PLATELETS ANYWHERE IN THE KIT AND THE RUN DATA LOG CONFIRMS THE INTERFACE'S BEHAVIOR AS SUCH. CORRECTIVE ACTIONS: THE CUSTOMER WAS PROVIDED WITH INSTRUCTIONS THAT WILL MAKE IT VERY UNLIKELY FOR THE OPTIA TO CONTRIBUTE TO REDUCTIONS IN PT'S PERIPHERAL PLATELET COUNTS. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2356414 | A CUSTOMER REPORTED AN DONOR UNIT LABELED A NEG RESULTED AS AB NEG WHEN TESTED ON THE GALILEO. A REVIEW OF INSTRUMENT RESULTS: (B)(6) RAN AT 1816 INTERPRETED AS AB NEG. A10 (ANTI-A) - 79 LARGE AGGLUTINATE PRESENT/ TIGHT BUTTON, VISUALLY POSITIVE, B10 (ANTI-B) - 70 LARGE AGGLUTINATE PRESENT V/ FUZZY, VISUALLY POSITIVE, C10 (ANTI-D 4) - 6 NO AGGUTINATION, VISUALLY NEGATIVE, D10 (RH CTRL) - 6 NO AGGLUTINATION, VISUALLY NEGATIVE, E10 (SAMPLE MIX) - X EMPTY OR TOO DENSE WELL DETECTED. ROWS E3, E4, AND E10 ALL PRESENTED WITH EMPTY OR TOO DENSE WELLS DETECTED ON THE PLATE. THE AFFECTED POSITIONS WERE NOT TESTING WELLS. ADVISED CUSTOMER THAT THE COMPLAINT WILL BE REFERRED TO SERVICE TO INVESTIGATE THE PIPETTORS AND SYRINGES DUE TO EMPTY OR TOO DENSE WELLS DETECTED. REVIEW OF CUSTOMER ACCOUNT INDICATES THAT THE PIPTEST FAILED YESTERDAY (PROCEEDING TO TEST SAMPLES IN THESE CONDITIONS IS CONTRARY TO LABELING) AND THE PLUNGER OF THE SYRINGE FOR PROBE #2 WAS NOT PROPERLY SEATED. ADDITIONALLY, PROBE #2 WAS REPLACED LAST WEEK. A SERVICE CALL WAS MADE. INVESTIGATION OF CUSTOMER COMPLAINT REVEALS FAINT FUZZY REACTION ON ANTI-B FOR SAMPLE IN QUESTION. PERFORMED RINSE PUMP VERIFICATION, PIPETTOR VERIFICATION, AND DAILY MAINTENANCE WITH EXPECTED RESULTS. PERFORMED INITIALIZATION AND QC TESTING WITH EXPECTED RESULTS. THE SYSTEM WAS TESTED AND IS OPERATING WITHIN SPECIFICATIONS. | 0
2362235 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) COUNT IN THE PLATELET PRODUCT OF A TRIPLE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION BECAUSE THE CUSTOMER DISCARDED OF IT. THIS REPORT IS BEING FILED DUE TO AN ALLEGED DEVICE MALFUNCTION. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. THE MEASURED WBC COUNT OF THE PRODUCT DOES NOT QUALIFY AS A WBC FAILURE PER THE CURRENT FDA GUIDELINES OF 12 X 10^6 FOR A TRIPLE PLATELET PRODUCT COLLECTION. CONCLUSION: NO FAILURE OCCURRED. | 0
2363614 | IT WAS REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED FOR UP TO 16 PTS FOR APPROX 10 MINUTES DUE TO A PRESUMED HARDWARE FAILURE. THE PTS WERE MOVED TO ALTERNATE TELEMETRY MONITORING UNTIL THE HARDWARE COULD BE REPAIRED. THERE WERE NO PT INJURIES REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. EXACT DATE OF THE EVENT IS UNK BUT IS BELIEVED TO HAVE OCCURRED IN (B)(6). THERE WERE 3 INCIDENTS ON 3 DIFFERENT UNITS WITH EACH HAVING A SEPARATE FORM 3500A. | 1
2364991 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE HIGHER THAN EXPECTED RESIDUAL WHITE BLOOD CELL (RWBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RWBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(4). THE DISPOSABLE SET WAS DISCARDED BY THE CUSTOMER, THEREFORE IT IS NOT AVAILABLE FOR RETURN FOR EVAL. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS INVESTIGATED FOR THIS EVENT. SIGNALS IN THE RDF DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE GREATER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED IN THE RDF AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PT PRODUCT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2368136 | IT WAS REPORTED THAT THE PATEINT'S SCHEDULED TRANSMISSION DID NOT "COME THROUGH." THE MONITOR HAD PREVIOUSLY SENT A SUCCESSFUL WIRELESS TRANSMISSION. ADDITIONAL INFORMATION RECEIVED INDICATED THAT THERE WAS A POSSIBLE SERVICE BROKER ISSUE WHICH MAY HAVE PREVENTED THE NEW SCHEDULED INFORMATION FROM BEING SENT. THE MONITOR HAS SUBSEQUENTLY SENT A SUCCESSFUL WIRELESS TRANSMISSION. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION: FURTHER REVIEW PROMPTED A CHANGE IN THE DEVICE ANALYSIS RESULTS. THE CHANGE IS REFLECTED IN THIS REPORT.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2368351 | IT WAS REPORTED THAT THE USER INDICATED THAT THERE WAS A SYSTEM REPORT OBSERVATION OF A CAPTURE MANAGEMENT WARNING, WHICH RESULTED IN THE PATIENT BEING BROUGHT INTO THE CLINIC UNNECESSARILY FOR A DEVICE INTERROGATION. THE USER THOUGHT THERE WAS A LEAD PROBLEM DUE TO THE REPORT ON THE WEBSITE. FURTHER INVESTIGATION INDICATED THAT THERE IS AN ACUTE PHASE WHERE THE VOLTAGES WILL CHANGE UNTIL THE TARGET AREA FOR LEAD OPTIMIZATION IS REACHED. WHEN THE PATIENT WAS BROUGHT IN, THE LEAD THRESHOLDS WERE "EXCELLENT." THE PATIENT'S LEADS AND DEVICE ARE STILL IN USE PER COMPANY PATIENT REGISTRATION DATABASE. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. CORRECTION: FURTHER REVIEW PROMPTED A CHANGE IN THE DEVICE ANALYSIS RESULTS. THE CHANGE IS REFLECTED IN THIS REPORT.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2369561 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT ON PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE ELEVATED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF CONTINUOUS ESCAPE OF WBCS FROM THE LRS CHAMBER THROUGHOUT THE MAJORITY OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ECT.) IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2370401 | THE ABBOTT VP 2000 PROCESSOR IS A DEVICE DESIGNED TO AUTOMATE AND STANDARDIZE SLIDE SPECIMEN PROCESSING AND ROUTINE SLIDE STAINING FOR THE LABORATORY. A CUSTOMER AND A FIELD SERVICE REPRESENTATIVE IN (B)(4) ALLEGED THAT THE METAL EDGES OF AN ABBOTT VP 2000 PROCESSOR (HEATER BLOCK, COVERS) UNIT ARE VERY SHARP. THE PARTS WITH EDGES THAT ARE ALLEGEDLY SHARP INCLUDE THE COVER FOR THE AMBIENT TEMPERATURE BATHS AND WHEN THE HEATED BASINS ARE REMOVED, THE OPENINGS OF THE HEATER REAGENT BASINS. DRYING STATION AND WATER BATH HAVE ALLEGEDLY SHARP EDGES ALSO. NO INJURY WAS REPORTED. ABBOTT MOLECULAR REQUESTED THE RETURN OF THE INSTRUMENT WITH THE ALLEGED SHARP EDGES FOR COMPLAINT INVESTIGATION. THE COMPLAINT IS CURRENTLY UNDER INVESTIGATION. ABBOTT MOLECULAR COMPLETED A SHARPNESS TEST ON A VP2000 INSTRUMENT CURRENTLY IN USE AT ABBOT MOLECULAR. THIS INSTRUMENT PASSED ON ALL USER ACCESSIBLE EDGES (PER UL STANDARD 1439). | 0
2372163 | A CUSTOMER REPORTED THAT 2 DONOR SAMPLES RECIEVED FROM THE BLOOD CENTER LABELED A POS RESULTED AS AB POS ON THE GALILEO FWD ABO ASSAY. CUSTOMER BELIEVES THAT THERE WAS SPLASHING IN THE B WELL ON BOTH; THE IMAGES OF THE WELLS LOOKED GREEN AND NOT YELLOW IN THE B WELL. A REVIEW OF INSTRUMENT RESULTS: ON PLATE ID (B)(4) READ AT 1:14PM, THE REACTIONS STRENGTH IN THE ANTI- B WELL FOR SAMPLE ID (B)(4) IN THE B09 WELL WAS 57 AND THE REACTION STRENGTH FOR (B)(4) IN THE B10 WELL WAS 81. TO THE LEFT OF THESE SAMPLES THERE WERE 2 SAMPLES THAT WERE STRONGLY POSITIVE FOR THE B WELL WITH REACTION STRENGTHS OF 99. UPON REVIEW OF THE IMAGE THERE DID APPEAR TO BE SOME DISCOLORATION OF THE BACKGROUND AND THERE WAS STILL A BUTTON PRESENT IN BOTH THE B09 AND B10 WELLS; HOWEVER BOTH WERE FUZZY BUTTONS. ADDITIONALLY, THE SAMPLES DID NOT DEMONSTRATE THE STRONG YELLOW COLOR OF THE OTHER B NEGATIVE SAMPLES THAT WERE FOUND ON THIS PLATE. ON PLATE PLATE ID (B)(4) READ AT 2:38PM THE DONOR SAMPLE (B)(4) IN THE B02 WELL HAD A REACTION STRENGTH OF 9 AND THE WELL WAS CLEARLY NEGATIVE WITH A UNIFORM DIFFUSION OF THE WELL. THE DONOR SAMPLE (B)(4) IN THE B01 WELL HAD A REACTION STRENGTH OF 6 AND WAS ALSO CLEARLY NEGATIVE WITH A UNIFORM DIFFUSION OF THE WELL. FOLLOW-UP WITH CUSTOMER: THE CUSTOMER INDICATED THAT HAS SHE PERFORMED NUMEROUS ABO FWD ASSAYS WITHOUT ERROR USING THE SAME REAGENT VIALS, WITHOUT MAKING ANY ADJUSTMENTS TO THE INSTRUMENT AND USING THE SAME SAMPLE SEGMENTS. NO FURTHER ISSUES REPORTED. | 0
2373645 | IT WAS REPORTED THAT THE CARESCAPE B850 PT MONITOR DID NOT APPROPRIATELY ALARM FOR BRADYCARDIA WHILE IN PT USE. THERE WAS NO REPORT OF DEATH OR SERIOUS INJURY. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. THE HOSPITAL WILL NOT RELEASE PT INFO. | 1
2375020 | ATTEMPTS FOR PRODUCT RETURN HAVE REMAINED UNSUCCESSFUL TO DATE.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WOULD NOT HOLD A CHARGE DESPITE BEING LEFT PLUGGED IN OVERNIGHT. NO OTHER DETAILS ARE KNOWN AT THIS TIME. ATTEMPTS FOR PRODUCT ARE UNDERWAY. . | 0
2375591 | UPON TRANSFER OF THE PT STEM CELLS FROM THE COLLECTION BAG TO A TRANSFER PACK, THE SAMPLE BULB DETACHED AND SEPARATED FROM THE COLLECTION BAG. ON (B)(6) 2011, CARIDIANBCT RECEIVED A VOLUNTARY MEDWATCH FORM 3500 FROM THE CUSTOMER, WHICH IS ATTACHED TO THIS REPORT. THIS EVENT WAS REPORTED TO THE FDA THROUGH THE MEDWATCH PROGRAM, REFERENCE REPORT NUMBER (B)(4). THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFO AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR DEATH OR SERIOUS INJURY OCCURRED. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2376899 | WHILE REVIEWING THE MANUFACTURER'S PROGRAMMING HISTORY DATABASE IT WAS NOTICED THAT THERE WAS AN INCOMPLETE DIAGNOSTICS THAT RESULTED IN THE PATIENT BEING SET TO UNINTENTIONAL SETTINGS. THE SETTING CHANGE WAS NOTICE, BUT ONLY PARTIALLY CORRECTED. THE PHYSICIAN LET THE PATIENT AT AN OFF TIME OF 60 MINUTES. ANALYSIS OF PROGRAMMING HISTORY. | 0
2377613 | IT WAS REPORTED THAT THE CLINIC RECEIVED A WIRELESS ALERT NOTIFICATION WITH HIGH URGENCY ELEVEN HOURS AFTER THE PATIENT WAS SHOCKED. IT WAS FURTHER REPORTED THAT THE PATIENT DIED LATER THE SAME DAY. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS IS ONE OF FIVE REPORTS ASSOCIATED WITH THIS EVENT: 6000094000-2011-02219, 2649622000-2011-16757, 2649622000-2011-16758, 2649622000-2011-16759. FURTHER REVIEW PROMPTED A CHANGE IN THE DEVICE ANALYSIS RESULTS.><THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS IS ONE OF FIVE REPORTS ASSOCIATED WITH THIS EVENT: 6000094-2011-02219, 2649622-2011-16757, 2649622-2011-16758, 2649622-2011-16759. | 1
2379214 | THE CUSTOMER STATES THAT THE SAMPLE CAMERA MISREAD THE BARCODE LABELS OF TWO SAMPLES FROM TWO DIFFERENT BATCHES. NO ERRONEOUS RESULTS WERE REPORTED. THE CUSTOMER REPRINTED THE SAMPLE BARCODE LABELS AND RETESTED. THE CUSTOMER STATED ALL BARCODE READINGS WERE ACCEPTABLE. THE CUSTOMER INDICATED THEY WILL MONITOR THE LABEL MAKER FOR SMUDGE AND SMEARING. THE CUSTOMER WILL CONTACT CTS IF THE INCIDENT OCCURS AGAIN. (B)(4). NO SERVICE NEEDED. | 1
2379357 | THE CUSTOMER REPORTED THAT THEY FOUND A BLACK PARTICLE (1 MM DIAMETER) IN THE CONCENTRATED PLATELET BAG AFTER TRANSFERRING TO THE LINKED BAG. NO MEDICAL INTERVENTION WAS NECESSARY FOR THIS PT. THE CUSTOMER REPORTED THIS EVENT TO THEIR LOCAL AUTHORITIES. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. (B)(4). INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2379432 | THE CUSTOMER REQUESTED A RUN DATA FILE (RDF) ANALYSIS FOR A PLATELET/PLASMA PROCEDURE THAT WAS SUBMITTED FOR ROUTINE QC TESTING. THE FINAL PLATELET VOLUME DID NOT MATCH UP WITH WHAT TRIMA MEASURED, THE DISCREPANCY WAS OVER 60ML. THE CUSTOMER ALSO REPORTED THAT THE DONOR HAD A SEVERE CITRATE REACTION AT THE END OF THE PROCEDURE. THE DONOR NOTED A VIBRATION IN THE ARM, AT THE COMPLETION OF THE PROCEDURE, AFTER THEY HAD BEEN DISCONNECTED FROM THE MACHINE, THE DONOR INDICATED THEY HAD A FLUTTERY FEELING IN THE CHEST AND A METALLIC TASTE IN THE MOUTH. AN AMBULANCE WAS CALLED TO THE DONOR CENTRE, THE DONOR WAS ASSESSED, BUT CLEARED AS OK AND LEFT THE CENTRE UNASSISTED. THE DOCTOR FROM THE BLOOD SERVICE SPOKE WITH THE DONOR THE FOLLOWING DAY AND DETERMINED THIS WAS LIKELY A MINOR CITRATE REACTION. PT AGE IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVAL. THIS REPORT IS BEING FILED DUE TO MEDICAL INTERVENTION. (B)(4). INVESTIGATION: THE RDF WAS ANALYZED FOR THIS EVENT. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND TRIMA ACCEL OPERATED AS INTENDED. THERE IS NO INDICATION IN THE RDF OF ANY EVENT THAT COULD RESULT IN THE LOW MEASURED VOLUME. IT IS POSSIBLE THAT THE PRODUCT BAG WAS NOT CLAMPED PRIOR TO UNLOADING THE CASSETTE ALLOWING SOME PRODUCT TO DRAIN BACK INTO THE SET AFTER THE CASSETTE WAS UNLOADED. IT IS ALSO POSSIBLE THAT THIS VOLUME DISCREPANCY WAS CAUSED BY A WEIGHTING OR CALCULATION ERROR. THERE WERE NO SIGNALS IN THE RDF TO INDICATE ANY REASON FOR THE CITRATE REACTION AT THE END OF THE PROCEDURE. ALL AC INFUSION LEVELS APPEARED NORMAL. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2382294 | DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THE USER REPORTED THAT THE T-LINK SYSTEM STARTED TO SLOW DOWN AND THEN STOPPED THE USER FROM INPUTTING DATA DURING THE CASE. THE USER REBOOTED THE SYSTEM, THE UNIT CAME BACK UP BUT THEN SLOWED DOWN AGAIN FINALLY SHUT DOWN. THE USER COMPLETED THE CASE WITH PAPER CHARTING. THE DEVICE WAS NOT CHANGED OUT. THE USER REPORTED THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THERE WERE NO ADVERSE CONSEQUENCES TO THE PATIENT AS A RESULT OF THIS EVENT. (DEVICE NOT RETURNED). | 1
2382720 | A NURSE REPORTED THAT THE EQUIPMENT DISPLAYED VERY SIMILAR LENS CALCULATION RESULTS AMONG PATIENTS. THE PHYSICIAN FEELS THAT THIS IS NOT NORMAL. THERE WERE NO INJURIES TO ANY PATIENTS REPORTED. THE COMPANY REPRESENTATIVE EXAMINED THE SYSTEM; TESTS AND CALIBRATIONS WERE PERFORMED. THE SYSTEM WAS THEN TESTED AND MET PRODUCT SPECIFICATIONS. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. (B)(4). | 1
2383601 | THE USER REPORTED THAT DURING A LEFT VENTRICULOGRAM PROCEDURE, THE ROTATOR BROKE TWICE IN A ROW. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF THREE REPORTS FOR THIS COMPLAINT: 1721504-2011-00397. DEVICE EVAL: THE SUSPECT DEVICE HAS NOT BEEN RETURNED FOR EVAL. THE USER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. | 0
2383602 | THE USER REPORTED THAT DURING A LEFT VENTRICULOGRAM PROCEDURE THE ROTATOR BROKE TWICE IN A ROW. NO HARM OF INJURY WAS REPORTED. THIS IS ONE OF THREE REPORTS FOR THIS COMPLAINT: 1721504-2011-00398. DEVICE EVAL: THE SUSPECT DEVICE HAS NOT BEEN RETURNED FOR EVAL. THE USER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVAL HAS BEEN COMPLETED. | 0
2384368 | ON (B)(6) 2011, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT A (B)(6) PRINTER WAS PRINTING GARBLED INFORMATION AND THE BATTERIES WERE HOT TO TOUCH. BASED ON THE INFORMATION AVAILABLE AT THE TIME, THERE WERE NO INJURIES ASSOCIATED WITH THIS EVENT. (B)(4). | 1
2384370 | ON (B)(6) 2011, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT A (B)(4) PRINTER WAS PRINTING GARBLED INFORMATION AND THE BATTERIES WERE HOT TO TOUCH. BASED ON THE INFORMATION AVAILABLE AT THE TIME, THERE WERE NO INJURIES ASSOCIATED WITH THIS EVENT. (B)(4). | 1
2385317 | CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT THE COULTER LH 750 SLIDE MAKER (LH 750 SLIDEMAKER) GENERATED FLUID DETECTOR ERRORS. CUSTOMER REPORTED THAT THEY DID NOT HAVE RED STRIPE TUBING FOR THE SAMPLE LINE. CUSTOMER REPORTED THAT THE SAMPLE LINE TUBING HAD BEEN REPLACED WITH THE WRONG TUBING. CUSTOMER REPORTED THERE WAS A SMALL LEAK (0.5 CUBIC CENTIMETER) OF GREENISH FLUID IN THE TRAY OF THE LH750 SLIDEMAKER. CUSTOMER REPORTED THAT ERRONEOUS RESULTS WERE NOT GENERATED. THERE WAS NO REPORT OF ANY ADVERSE EVENT OR INJURY REQUIRING MEDICAL INTERVENTION OR PATIENT TREATMENT. BEC FIELD SERVICE ENGINEER (FSE) ENSURED THAT THE COULTER LH 750 SLIDE MAKER SAMPLE LINE HAD THE APPROPRIATE TUBING. THE FSE ALSO CLEANED THE SHUTTLE GUIDE RAILS TO ALLOW THE SHUTTLE TO REACH THE DISPENSE POSITION PROPERLY. THERE WAS NO FURTHER EVIDENCE OF LEAKING AFTER THE REPAIRS. THE FSE VERIFIED THE REPAIRS WERE PERFORMED PER ESTABLISHED PROCEDURES. THE FSE VERIFIED THE RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. BEC IDENTIFIER FOR THIS COMPLAINT IS (B)(4). | 0
2387857 | IT WAS REPORTED THAT DURING PATIENT USE THE CARESCAPE B850 PATIENT MONITOR EMITTED AN UNUSUAL CONSTANT BEEP-BEEP TONE THAT DID NOT AUDIBLY CORRESPOND TO AN EXPECTED ALARM TONE OR OTHER RECOGNIZABLE TONE, AND DID NOT CORRELATE TO ANY VISUAL PATIENT CONDITION. AS LONG AS THIS TONE IS SOUNDING NO OTHER DISCERNIBLE AUDIBLE ALARMS WILL BE ANNOUNCED. VISUAL ALARMS WERE NOT AFFECTED. THERE WAS NO DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT, NOR AN INDICATION THAT INTERVENTION WAS PROVIDED. GE HEALTHCARE'S INVESTIGATION REVEALED THAT A SOFTWARE ISSUE EXISTS IN WHICH A SOUND RESOURCE IS RE-USED, RESULTING IN THE REPORTED "BEEP-BEEP" TONE TO BE REPLAYED. THE ISSUE WILL ONLY OCCUR WHEN ALARM TONES ARE CHANGED FROM DEFAULT IEC TONES (AS DEFINED IN (B)(4)) TO LEGACY TONES (GE DEFINED TONES) AND REMOTE ALARM TONE IS CHANGED FROM LOCAL (SAME TONE AS USED IN LOCAL MONITOR) TO REPEAT (REMOTE ALARM IS REPEATED WITH ITS OWN TONE) IN THE DEFAULT SETUP MENU. IN ADDITION, ONLY WHEN A SPECIFIC SEQUENCE OF REAL PATIENT ALARMS OCCURS SIMULTANEOUSLY, BOTH WITH THE LOCAL PATIENT OR A REMOTE PATIENT MONITORED VIA THIS MONITOR, WILL THE ISSUE OCCUR. IN THIS SCENARIO, THE REPORTED "BEEP-BEEP" TONE WILL CONTINUE WITH NO CORRESPONDENCE TO THE ACTUAL PATIENT ALARM MESSAGES ON THE DISPLAY. HOWEVER, THE REAL PATIENT ALARM MESSAGES, EXISTING AND NEW, WILL BE DISPLAYED NORMALLY ON THE MONITOR. THE ALARM LIGHT AT THE TOP OF THE DISPLAY WILL CONTINUE TO PERFORM AS DESCRIBED IN THE OPERATIONS MANUAL. THE ALARM SOUND SEQUENCES OF THE POSSIBLE NEW ALARMS WILL NOT ACTIVATE AUDIBLY AT THE MONITOR, BUT ALARMS ARE TRANSFERRED NORMALLY TO THE NETWORKED CENTRAL STATION. THE MONITOR MUST BE REBOOTED IN ORDER TO RESTORE PROPER AUDIBLE ALARM FUNCTION. GE HEALTHCARE HAS CORRECTED THIS ISSUE IN THE NEXT REVISION OF SOFTWARE, WHICH IS BEING APPLIED TO FORWARD PRODUCTION. NO PATIENT OR USER HARM HAS BEEN REPORTED FOR THIS ISSUE. NO SIGNIFICANT SAFETY RISK TO PATIENT OR USER HAS BEEN IDENTIFIED WITH THE ISSUE AS CHARACTERIZED. WARNINGS EXIST IN THE USER'S MANUAL REGARDING POSSIBLE EQUIPMENT MALFUNCTIONS. THE CARESCAPE MONITOR B850 USER'S MANUAL STATES: "EQUIPMENT MALFUNCTIONS, NETWORK OR NURSE CALL DISCONNECTION, AND ALARM VOLUME SETTINGS MAY RESULT IN MISSED ALARMS. ALWAYS KEEP THE PATIENT UNDER CLOSE SURVEILLANCE." IN ADDITION, THE MANUAL STATES, "IF AN ERROR MESSAGE APPEARS DURING OPERATION, IT IS THE LICENSED MEDICAL PRACTITIONER'S RESPONSIBILITY TO DECIDE WHETHER THE UNIT IS STILL SUITABLE FOR PATIENT MONITORING. AS A GENERAL RULE, MONITORING SHOULD ONLY CONTINUE IN EXTREMELY URGENT CASES AND UNDER THE DIRECT SUPERVISION OF A LICENSED MEDICAL PRACTITIONER. THE UNIT MUST BE REPAIRED BEFORE BEING USED AGAIN ON A PATIENT. IF AN ERROR MESSAGE APPEARS AFTER POWER-UP, THE UNIT MUST BE REPAIRED BEFORE BEING USED ON A PATIENT." | 1
2387858 | IT WAS REPORTED THAT PREVENTATIVE MAINTENANCE HAD RECENTLY BEEN DONE ON THE CARDIOCAP 5 BY HOSPITAL STAFF OR HOSPITAL BIOMED AND THE AUDIBLE ALARM WAS NOT ALARMING FOR SP02. ALL OTHER FUNCTIONS WERE REPORTEDLY NORMAL. GE TECHNICAL SUPPORT ATTEMPTED TO RESOLVE THE ISSUE BUT THE HOSPITAL REQUESTED A GE FIELD SERVICE DISPATCH. THERE WAS NO REPORT OF PATIENT INVOLVEMENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2387912 | A CUSTOMER REPORTED THAT ON (B)(6) 2011 THE DATALINK/DL2000 DATA MANAGER MODIFIED AN INSTRUMENT GENERATED C-REACTIVE PROTEIN SAMPLE RESULT OF '<150.0 MG/L' TO '<0.5 MG/DL'. THE ORIGINAL SAMPLE RESULT WAS ACCOMPANIED BY A "LT" INSTRUMENT MESSAGE STANDING FOR "OUT OF ORDAC RANGE LOW." THE DL2000 HAD AN APPLICATION CONFIGURABLE RULE TO MODIFY ANY RESULT THAT CONTAINED THE 'LT' MESSAGE TO '<0.5 MG/DL'. IT IS UNKNOWN AS WHO CONFIGURED THIS RULE IN THE DATALINK/DL2000 DATA MANAGER. THE DL2000 PROCESSED THE RESULT CORRECTLY ACCORDING TO THIS INCORRECTLY ESTABLISHED RULE. AN ERRONEOUS RESULT WAS GENERATED AND REPORTED FROM THE LABORATORY HOWEVER, THERE WAS NO DEATH, SERIOUS INJURY OR MODIFICATION TO PATIENT TREATMENT ASSOCIATED OR ATTRIBUTED TO THIS EVENT. UPON RETEST OF THE SAMPLE IN A DILUTED FORM ON THE SAME INSTRUMENT, A VALUE OF "101.0 MG/L" WAS OBTAINED, REGARDED AS VALID, AND AN AMENDED REPORT WAS ISSUED. NO SAMPLE COLLECTION/HANDLING INFORMATION, SYSTEM CALIBRATION/QUALITY CONTROL INFORMATION OR PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS EVENT. ACTUAL PATIENT DATA WAS NOT PROVIDED BY THE CUSTOMER. BECKMAN COULTER INC. ASSESSMENT OF INSTRUMENT PROGRAMMING INDICATED THAT THE DATALINK/DL2000 DATA MANAGER CORRECTLY PROCESSED THE SAMPLE USING AN INCORRECTLY CONFIGURED RULE OF "IF CRP <5.0 CRP.FLAG CONTAINS <5.0 OR CRP=LT,THEN RESULT('<5.0';VALID." THE CUSTOMER DISABLED THE INVOLVED SOFTWARE CONFIGURATION RULE SO THAT THE DATALINK/DL2000 DATA MANAGER WOULD NOT CONVERT THESE TYPES OF "LT" FLAGGED RESULTS. THE ROOT CAUSE OF THIS EVENT APPEARS TO BE AN INCORRECTLY CONFIGURED RULE. | 1
2388420 | SPACELABS RECEIVED A REPORT THAT A SPACELABS CENTRAL STATION MONITOR DID NOT ALARM ON V-FIB. THE INVESTIGATION REVEALED THAT THE BEDSIDE MONITOR DID ALARM ON V-FIB, AS DOCUMENTED IN SPACELABS' ICS SYSTEM. THERE WAS NO PRINTED RECORD GENERATED FOR THE ALARM AS THE AUTO ALARM RECORDING FEATURE WAS TURNED TO OFF. ADDITIONALLY, THE ALARM VOLUME ON THE BEDSIDE MONITOR IN THE PATIENT ROOM WAS SET TO VERY LOW, ALTHOUGH THE VOLUME ON THE CENTRAL STATION WAS SET TO AN AUDIBLE LEVEL. TESTING CARRIED OUT BY SPACELABS' FIELD SERVICE ENGINEER (FSE) AND THE HOSPITAL'S BIOMED AT THE CUSTOMER'S SITE CONFIRMED THAT THE CENTRAL STATION MONITOR PASSED ALL TESTS AND WAS OPERATING TO SPECIFICATIONS. ALARMS GENERATED BY THE BEDSIDE MONITOR WERE PRESENT AND AUDIBLE AT THE CENTRAL STATION. THE FSE AND BIOMED WERE NOT ABLE TO DUPLICATE THE REPORTED FAULT. THE CUSTOMER IS CONTINUING TO USE THE CENTRAL STATION MONITOR. NO ONE HAS BEEN INJURED AS A RESULT OF THIS ALLEGED MALFUNCTION. SPACELABS WILL OFFER ADDITIONAL TRAINING TO THE CUSTOMER REGARDING RECORD GENERATION AND ALARM VOLUMES. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE IS CLOSED. | 1
2388776 | DR. (B)(6) REPORTED THAT THE VES STANDARD UPTAKE VALUE (SUV) MEASUREMENTS WERE NOT CONSISTENT WITH OTHER POST-PROCESSING SYSTEM WHILE COMPARING SUVS BETWEEN THREE POST-PROCESSING SYSTEMS. DR. (B)(6) FOUND THIS WHILE EXPLORING DIFFERENCES BETWEEN SYSTEMS AND NOT DURING PT CARE. NO INJURIES WERE REPORTED RELATED TO THIS ISSUE. THE INITIAL INVESTIGATION OF THIS ISSUE FOUND THE SUV ERROR WAS DUE TO INVALID DICOM HEADER INFO OF IMAGES PROVIDED BY A PHILIPS HEALTHCARE GEMINI 64 TIME-OF-FLIGHT WHEN USING BQML UNITS. THE INVALID DICOM HEADER ISSUE WAS FORWARDED TO PHILIPS HEALTHCARE. VITAL IMAGES BECAME AWARE ON (B)(4), 2011 THAT THERE WAS AN ISSUE WITH THE VITREA PRODUCT. THE ISSUE WAS FOUND WHILE EVALUATING A POTENTIAL WORKAROUND FOR THE INVALID DICOM HEADER. THE ERROR IS UNIQUE TO THE ACQUISITION WORKFLOW USED TO CREATE THE EXAMPLE IMAGES PROVIDED BY THE CUSTOMER. THE VITREA ERROR WAS NOT DETECTABLE UNTIL THE EFFECT OF THE ORIGINAL IMAGE DICOM ERROR WAS REMOVED. | 1
2392855 | RN AND MD GOT FETAL HEART TONES AT 164 TWICE AND ULTRASOUND SHOWED FETAL DEMISE. PARENTS VERY UPSET. | 1
2392882 | PT HAD ERCP SCHEDULED AND DR (B)(6) ATTEMPTED STONE RETRIEVAL WITH BASKET. A SMALL STONE WAS RETRIEVED BUT WHEN HE TRIED TO RETRIEVE THE LARGE SECOND STONE, THE WIRE BROKE DURING THE MECHANICAL LITHOTRIPSY. DR (B)(6) CONSULTED DR (B)(6) AND SURGERY WAS DONE THE NEXT DAY FOR REMOVAL ALONG WITH MULTIPLE STONES. | 0
2393612 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT DURING THE LOADING/UNLOADING PROCESS OF THE AUTOMATE 2500 HS S3I (SORTER AND ALIQUOTER) SYSTEM WITH TUBE RACKS THE COVER UNEXPECTEDLY CLOSED AND STRUCK THE OPERATOR ON THE HEAD. THE OPERATOR WAS HIT ON THE FOREHEAD AND ENDURED A HEAD CONTUSION (NO CUT, NO BRUISE, SLIGHT BUMP ON THE HEAD). NO MEDICAL INTERVENTION WAS NECESSARY. THE CUSTOMER INDICATED THAT THE ANALYZER WAS FITTED WITH OLD GAS SPRINGS, WHICH MAY BECOME WORN AND DETERIORATE OVER TIME. IT IS UNKNOWN WHETHER THE GAS SPRINGS ON THE COVER WERE THE ORIGINAL ONES OR HAD ALREADY BEEN REPLACED IN THE SYSTEMS HISTORY. NO PATIENT SAMPLES WERE INVOLVED IN THIS EVENT. THE CUSTOMER INDICATED THAT THE COVER ITSELF (AND THE GAS SPRINGS) HAVE AN ADDITIONAL LOCKING BRACKET WHERE ONCE THE COVER IS OPENED, WHEN PROPERLY ACTIVATED, PREVENTS A CRASHING DOWN OF THE COVER IN THE EVENT OF INADEQUATE STRENGTH OF THE SPRINGS. THE USER IS CAUTIONED IN THE USER GUIDE TO MAKE SURE THE LOCKS ARE IN PLACE WHEN OPENING THE COVER. HOWEVER, THE COVERS SHOULD NOT BE OPENED TO REMOVE OR INSERT RACKS INTO THE SYSTEM. THE OPERATOR VIOLATED THE INTENDED USE OF THE SYSTEM. THE OPERATOR DISREGARDED THE SAFETY INSTRUCTIONS AND WARNINGS GIVEN IN THE BEC IFU (INSTRUCTIONS FOR USE) MANUAL AS WELL AS THE TWO SAFETY LABELS ON THE COVER OF THE AUTOMATE SYSTEM AND DID NOT OPERATE THE SYSTEM AS SPECIFIED: IFU TELLS NOT TO REACH INTO THE DEVICE FOR LOADING/UNLOADING TRAYS. LOADING/UNLOADING TRAYS IS ONLY TO DO WITH CLOSED COVERS ON FULLY EXTENDED DRAWERS, COVERS ARE ONLY NEEDED TO BE OPENED IN CASE OF FAILURE WHICH NEEDS USER INTERACTION AND FOR CLEANING PURPOSE, THE NORMAL STATE OF THE COVERS IS CLOSED, ADDITIONALLY THE STICKER ON THE COVER ITSELF TELLS THE USER TO MAKE SURE THE COVER STAYS IN UP POSITION BEFORE REACHING INTO THE DEVICE. A BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(6) 2011 AND PERFORMED TROUBLESHOOTING ON THE SYSTEM. THE FSE REPLACED BOTH GAS SPRINGS. PARTS WERE NOT AVAILABLE FOR ANY FURTHER INVESTIGATION. THE FSE ALSO INFORMED THE CUSTOMER TO NOT VIOLATE THE IFU AND THE SAFETY STICKERS ON THE COVER. | 0
2394053 | THE ANALYZER HAD RECEIVED MAINTENANCE FROM A RADIOMETER SVC ENGINEER WHO HAD RE-MEMBRANED THE REFERENCE ELECTRODE, AND CALIBRATED THE INSTRUMENT. FOLLOWING THIS PROCEDURE THE QC RESULTS FAILED, YIELDING THE ERROR MESSAGE "PH UNSTABLE." THE TRAFFIC LIGHT ON THE INSTRUMENT DISPLAY WAS IN YELLOW, SIGNALLING THE QC ERROR. THE SVC ENGINEER THEN LEFT THE INSTRUMENT. LATER IN THE AFTERNOON, A SERIES OF SAMPLES WERE ANALYZED ON THE INSTRUMENT, AND THE RESULTS WERE USED DIAGNOSTICALLY. TWO PREMATURE NEONATAL PTS WERE TREATED WITH INCREASED AMOUNTS OF SODIUM BICARBONATE UNTIL THE ERROR WAS NOTICED. LATER ON THIS NIGHT, THE CUSTOMER RE-MEMBRANED THE REFERENCE ELECTRODE SEVERAL TIMES WITHOUT SUCCESS. ON THE NEXT MORNING, (B)(6) 2011, A SVC TECH VISITED THE CUSTOMER AGAIN AND PERFORMED MAINTENANCE ON THE INSTRUMENT. TREATMENT WITH SODIUM BICARBONATE WAS CANCELLED IMMEDIATELY. IT IS YET UNCLEAR IF ANY OF THE PTS TOOK ANY PERMANENT INJURY FROM THE EVENT. DR (B)(6) WILL DETERMINE THIS FROM A F/U EXAMINATION IN 2 - 3 WEEKS. THE EVENT WAS CAUSED BY A USE ERROR. ALL MEASUREMENT RESULTS DISPLAYED OR PRINTED BY THE INSTRUMENT CONTAINED THE ERROR MESSAGE "0476: MESSUNG INSTABIL" FOR PH, AND SHOULD THEREFORE NOT HAVE BEEN USED FOR DIAGNOSIS. EVEN THOUGH THE INSTRUMENT WAS INSUFFICIENTLY MAINTAINED AND HAD A MALFUNCTION, THE SAFETY MEASURES WORKED AS INTENDED AND THE USER WAS WARNED ABOUT THE ERROR. FOLLOWING THE EVENT A SVC TECH CLEANED THE PH/BG MEASURING CHAMBER, REPLACED THE SEALING GASKET OF THE PH ELECTRODE, REPLACED THE REFERENCE ELECTRODE, AND RE-MEMBRANED THE NEW REFERENCE ELECTRODE. | 0
2394313 | CANON (B)(4) MEDICAL SYSTEMS DIVISION (CMS) RECEIVED A REPORT OF GRAYED OUT OR BLANK IMAGES WITH CXDI-70C WIRELESS PANEL USED IN CONJUNCTION WITH NE CONTROL SOFTWARE VERSION 1.40 AND 1.40.1. UPON INVESTIGATION, IT WAS FOUND THAT IMAGES WOULD APPEAR GRAY OR BLANK WHEN SENT TO THE PACS AND WERE STORED THAT WAY IN THE PAST STUDY LIST ON THE NE SOFTWARE. THE BLANK IMAGES WERE UNUSABLE AND NECESSITATED RETAKING IMAGES WHICH WOULD SUBJECT PTS TO UNNECESSARY ADDITIONAL RADIATION EXPOSURE. CANON (B)(4) ISSUED VERSION 1.40.2.0 NE CXDI CONTROL SOFTWARE ON (B)(4) 2011 TO ADDRESS THIS ISSUE. | 1
2395316 | SIEMENS RECEIVED NOTIFICATION ON 04/08/2011 VIA E-MAIL FROM THE ACTING RADIATION THERAPY SERVICE MANAGER OF THE RADIATION THERAPY DEPARTMENT, WITH THE FOLLOWING REQUEST AND PARTICULARS: "CAN YOU INVESTIGATE THE PROBLEM OUTLINED BELOW AS AN URGENT PRIORITY? A TREATMENT FIELD IS OPENED IN LANTIS FOR ASSESSMENT OR AMENDMENT. IT IS THEN RE-APPROVED BUT NOT CLOSED. FURTHER CHANGES ARE THEN MADE TO THE FIELD AND IT IS THEN CLOSED. THE FIELD REMAINS APPROVED WITH THE FURTHER CHANGES. THEREFORE THERE IS A SERIOUS RISK OF MAKING AMENDMENTS TO A FIELD AND NOT BEING AWARE OF IT." THE SUMMARY OF POTENTIAL COMPLAINT ENTRY INDICATED "YES" TO "POTENTIAL COMPLAINT RELATIVE TO SAFETY" AND "UNKNOWN" TO BOTH "WAS THE DEVICE BEING USED TO TREAT OR DIAGNOSE PT DURING EVENT?" AND "WAS ANYONE INJURED (EVEN SLIGHTLY)?" THE MATTER WAS REFERRED TO THE THIRD-PARTY SOFTWARE PROVIDER FOR EVAL. IT WAS DETERMINED THE SYSTEM WORKED AS SPECIFIED AND THIS ASSESSMENT WAS PROVIDED TO THE COMPLAINANT SITE DURING A REGULAR SERVICE MEETING. A NEW COMPLAINT FILE WAS OPENED WHEN THE SITE CLARIFIED ON 11/25/2011 "THAT THERE WAS AN ACTUAL INCIDENT WHERE A PT WAS TREATED WITH THE WRONG FIELD SIZE POSSIBLE DUE TO THE ISSUE OUTLINED IN THE ORIGINAL COMPLAINT." THE SITE FURNISHED AN INCIDENT REPORT WHICH INDICATED "DIAGNOSIS CA PROSTATE; PRESCRIPTION AND FRACTIONATION 37# 74GY; DATE AND TIME (B)(6) 2011 3.30 PM; NO. FIELDS INVOLVED 1; NO. FRACTIONS INVOLVED 11; RADICAL TREATMENT; ERROR TYPE: DOSE/TUMOUR SITE" AND CONTAINED ADDITIONAL DETAILS ABOUT THE SITE'S INTERNAL ASSESSMENT. THE REPORT RECEIVED ON (B)(4) 2011 INDICATED A PT WAS INJURED, WITH EXTENT OF INJURIES "POSSIBLE RADIATION DAMAGE". INCIDENT REPORTED TO IRISH MEDICINE BOARD. THE SITE HAS NOT RETURNED PARTICULAR DETAILS ABOUT THE PT REPORTED TO HAVE BEEN INVOLVED. THE MANUFACTURER'S PRELIMINARY INVESTIGATION BASED ON A REVIEW OF THE SITE'S SYSTEM LOG FILES AND CLARIFICATIONS FROM THE SITE PERSONNEL, INDICATE THE SYSTEM PERFORMED AS SPECIFIED AND THE INCIDENT IS LIKELY ATTRIBUTABLE TO USER ERROR; HOWEVER, ADDITIONAL MANUFACTURER REVIEW AND INVESTIGATION IS UNDERWAY. WE BECAME AWARE OF THIS RECENT INFO ON (B)(4) 2011. | 0
2396375 | FIRE IN THE MRI. INITIAL REPORTS ARE THAT STAFF FIRST DETECTED SMOKE FOLLOWED BY VISIBLE FLAMES. STAFF AND PATIENT EVACUATED THE MRI WITHOUT INJURY. MRI TECHNOLOGIST RETURNED TO THE ROOM TO QUENCH THE MAGNET'S CRYOGEN TO REMOVE THE MAGNETIC FIELD. FIRE DEPARTMENT RESPONDED AND EXTINGUISHED FIRE. BECAUSE OF SMOKE CONDITIONS IN THE ADJACENT AREAS, SEVEN EMERGENCY DEPARTMENT PATIENTS WERE TRANSFERRED TO ANOTHER HOSPITAL FOR CONTINUED CARE. ADDITIONALLY, ICU PATIENTS LOCATED ON THE SECOND FLOOR ABOVE THE MRI WERE MOVED LATERALLY WITHOUT INCIDENT. THE EMERGENCY DEPARTMENT WENT ON DIVERT.THE MANUFACTURER RESPONDED IMMEDIATELY TO A SERVICE CALL INITIATED AT THE TIME OF THE EVENT. THEY IDENTIFIED THE LOCATION OF THE SOURCE OF THE FIRE AT A CONNECTION ON THE HIGH CURRENT CABLE ON THE BORE. | 0
2396797 | PT COMPLAINED TELEMETRY UNIT WAS HOT, NURSE CHECKED TELEMETRY UNIT (B)(4) AND REMOVED FROM PT. TELEMETRY UNIT WAS WARM TO TOUCH, BATTERIES WERE REMOVED WHICH WERE TOO HOT TO HANDLE AND INSIDE COVERING WAS MELTING. TELEMETRY UNIT REMOVED FROM SERVICE AND EVALUATED BY BIOMED. A COMPONENT ON THE MOTHERBOARD HAD COME LOOSE AND WAS SHORTING OUT THE BATTERIES WHICH CAUSED THE UNIT TO OVERHEAT. | 0
2396812 | TELEMETRY UNIT (B)(4) WAS WARM TO TOUCH AND REMOVED FROM PT. TELEMETRY UNIT REMOVED FROM SERVICE AND EVALUATED BY BIOMED. BIOMED CONNECTED THE UNIT TO A STIMULATOR BUT COULD NOT DUPLICATE PROBLEM. AFTER 30-45 MINS OF NORMAL OPERATIONS, THE PROBLEM COULD NOT BE DUPLICATED. THE TELEMETRY UNIT WAS RETURNED TO THE NURSING UNIT (B)(4) 2011. | 0
2397528 | DBSS FAILS TO PERFORM VALIDATION CHECKS WHEN UPDATES SUCH AS BLOOD TYPE GROUP ARE MADE IN THE "EDIT UNIT" SCREEN TO DETERMINE IF THE UNIT SHOULD BE QUARANTINED. IF THIS EVENT WERE TO REOCCUR, THE EVENT MAY CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. DUE TO THE FACT THAT PLASMA PRODUCTS DO NOT REQUIRE ABO CONFIRMATORY TESTING, AN INCOMING PRODUCT WOULD BE RECEIVED INTO THE FACILITY AS AVAILABLE. SINCE PLASMA PRODUCTS DO NOT REQUIRE CROSSMATCHING, THE ABO INCOMPATIBILITY WOULD NOT BE IDENTIFIED PRIOR TO THE UNIT BEING ISSUED AND SUBSEQUENTLY TRANSFUSED. THIS IN TURN COULD RESULT IN THE DEATH OR SERIOUS INJURY TO A PATIENT. | 0
2397568 | DURING INSTALLATION OF THE UNICEL DXH SLIDEMAKER STAINER AND FLOOR STAND, BECKMAN COULTER INC., (BEC) FIELD SERVICE ENGINEER (FSE) FOUND THAT THE VINYL SAFETY CAPS WERE MISSING FROM THE DXH SMS FLOOR STAND DRAWER SLIDES. IT IS UNKNOWN IF AN EXPOSURE CONTROL/RISK MANAGEMENT PLAN IS IN PLACE. NO INJURIES OCCURRED AND MEDICAL ATTENTION WAS NOT SOUGHT. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT ATTRIBUTES OR CONNECTED TO THIS COMPLAINT. THERE WAS NO EFFECT TO PATIENT RESULTS. THE FLOOR STAND CONTAINS FOUR DRAWERS FOR REAGENT AND WASTE STORAGE WITH TWO SLIDES PER DRAWER. THE VINYL SAFETY CAPS SHOULD HAVE BEEN INSTALLED ON THE EIGHT METAL EDGES OF THE DRAWER SLIDES IN MANUFACTURING TO PROTECT THE USER FROM INJURY WHILE CHANGING REAGENT OR WASTE CONTAINERS. PER LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. THE ROOT CAUSE FOR THIS ISSUE IS AN INCOMPLETE IMPLEMENTATION IN BEC MANUFACTURING. (B)(4). | 0
2398125 | A CUSTOMER REPORTED TO BECKMAN COULTER INC (BEC) THAT A CLEAR LIQUID LEAK WAS FOUND UNDERNEATH COULTER (B)(4) SLIDEMAKER. THE USER WAS WEARING PPE CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE INCIDENT. THERE WAS NO REPORT OF EXPOSURE TO MUCOUS MEMBRANES OR OPEN LESIONS. MEDICAL ATTENTION WAS NOT SOUGHT. THE MATERIAL SAFETY DATA SHEET (MSDS) WAS NOT REVIEWED, BUT AN EXPOSURE CONTROL OR RISK MANAGEMENT PLAN IS IN PLACE. PATIENT SAMPLES WERE NOT AFFECTED. THE FIELD SERVICE ENGINEER (FSE) FOUND TWO LEAKS INSIDE THE INSTRUMENT: ONE AT THE TUBING THROUGH VL14 WHERE THE TUBING HAD A HOLE AND THE OTHER AT THE SAMPLE ACCESS MODULE. THE DURO TUBING AT THE SAMPLE ACCESS MODULE WAS WORN AND ALSO HAD A HOLE IN IT. THE FSE REPLACED THE TUBING AT BOTH LOCATIONS AND THE REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES. ROOT CAUSE FOR THE LEAK IS ATTRIBUTED TO HOLES ON TUBING. BEC INTERNAL IDENTIFIER FOR THIS COMPLAINT IS (B)(4). | 0
2400485 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A SYRINGE PROBE ERROR AND A SMALL LEAK COMING FROM ONE OF THE COMPONENTS ON THE RIGHT SIDE COMPARTMENT OF THE COULTER ACT 5DIFF AUTOLOADER HEMATOLOGY ANALYZER. THE CUSTOMER WAS UNABLE TO LOCATE THE SOURCE OF THE LEAK. THE CUSTOMER WAS WEARING GLOVES AT THE TIME OF THE INCIDENT AND NO INJURY OR EXPOSURE WAS REPORTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE THE DAY OF THE EVENT AND DID NOT OBSERVE ANY LEAK OR RESIDUE FROM A LEAK. WHILE ON-SITE, FSE FOUND THE SAMPLE SYRINGE WAS NOT REACHING HOME STATE. FSE REPLACED THE SAMPLE SYRINGE HOME SWITCH, VERIFIED REPAIR AND VALIDATED SYSTEM. THE ROOT CAUSE FOR THE LEAK IS UNKNOWN. THE BEC INTERNAL IDENTIFIER FOR THIS EVENT IS (B)(4). | 0
2401275 | THE GE TELEMETRY SYSTEM (CENTRALIZED MONITORING) DID NOT ALARM FOR AN 11 BEAT RUN OF VENTRICULAR TACHYCARDIA. THE SYSTEM SOMETIMES DOES NOT RECOGNIZE V-TACH, WHICH IS A SERIOUS CONCERN TO US. THE PERSON MONITORING THE PATIENT RECOGNIZED THE EVENT, AND REPORTED IT TO THE NURSE ON THE FLOOR CARING FOR THE PATIENT. SO, THE PATIENT WAS NOT HARMED. | 1
2401782 | THE CUSTOMER REPORTED THAT DURING THE PROCEDURE, JUST AFTER THE BEGINNING, THE PLASMA TUBE OF THE COLLECTION CHAMBER OF THE CHANNEL CAME OUT OF THE CHAMBER ITSELF. THE CUSTOMER SAID THAT THIS CAUSED A BIG BLOOD LEAK INTO THE CENTRIFUGE CHAMBER. PATIENT INFORMATION IS UNAVAILABLE AT THIS TIME. THIS DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. CARIDIANBCT INTERNAL REFERENCE: (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2404698 | PATIENT WAS IN THE OR AND HAVING EXTENSIVE GASTRIC SURGERY. THE PATIENT WAS CONNECTED TO A GE SOLAR 8000I MONITOR THAT INCLUDED A PERIOPERATIVE INFORMATION MANAGEMENT (PIM) SOFTWARE INTERFACE. THE PIM SOFTWARE IS IN THE DATA PROCESSING UNIT (DPU). THE DPU CONNECTS THE SOLAR DISPLAY, CARESCAPE MODULE, NETWORK AND OTHER VARIOUS DEVICES TOGETHER. THIS IS FAIRLY NEW TECHNOLOGY IN THE OPERATING ROOM AND WHILE THIS DEVICE HAS BEEN USED THROUGHOUT THE HOUSE (HOSPITAL) WITHOUT DIFFICULTY, THE OPERATING ROOM SOLAR MONITORS HAVE HAD MULTIPLE OCCASIONS WHEN THE NON-INVASIVE BLOOD PRESSURE (NIBP) FUNCTION FAILED TO TAKE A BLOOD PRESSURE (THE NIBP WILL EXIT THE AUTO INTERVAL MODE WITHOUT AN ALARM), DESPITE THE PATIENT HAVING A PALPABLE PULSE. IN THIS CASE, THE ANESTHESIOLOGIST WAS UNABLE TO OBTAIN A BP FOR SEVERAL MINUTES AND HAD TO PLACE AN ARTERIAL LINE TO CONTINUE WITH THE SURGERY. WHEN THIS WAS FIRST PRESENTED TO THE BIOMEDICAL STAFF THEY HAD MULTIPLE CONVERSATIONS WITH GE, AND APPARENTLY THIS PROBLEM OCCURS ONLY IN THE OPERATING ROOM MODE WITH THE PIM AND WAS EXACERBATED WITH AN UPGRADE FROM SOFTWARE VERSION 1.2, TO VERSION 2.0, THE PROBLEM PERSISTS. AFTER SPEAKING WITH THE BIOMEDICAL ENGINEERS WHO HAVE BEEN IN CONTACT WITH THE GE REPRESENTATIVES, THE UPCOMING SOFTWARE UPGRADE IS GOING TO FIX THIS PROBLEM. | 1
2404705 | THE DATEX OHMEDA ANESTHESIA DELIVERY UNIT (ADU) CHECKED OUT AND WORKED FOR THE FIRST 2 CASES. FOR THE 3RD CASE WHEN THE VENTILATOR TURNED ON, VENTILATION COMMUNICATION ERROR FLASHED AND OXYGEN SUPPLY ALARM WAS NOT ACTIVATED PER MACHINE. WHEN THE VENTILATOR DID NOT FUNCTION CORRECTLY, THE PATIENT WAS RETURNED TO HAND VENTILATION AND THE REPLACEMENT MACHINE WAS BROUGHT INTO THE ROOM. THE PATIENT STABLE THROUGHOUT, OXYGEN SATURATION'S 96 TO 100. BIOMED TRACED THE PROBLEM TO A BAD SPIROMETRY CONNECTOR ON THE GAS MODULE. REPLACED GAS MODULE AND CHECKED MACHINE AND IT IS WORKING WELL NOW. RETURNED TO SERVICE. | 0
2404722 | BELOW IS AN EMAIL COMMUNICATION THAT DESCRIBES YET ANOTHER NON-INVASIVE BLOOD PRESSURE MACHINE FAILURE IN THE DATA PROCESSING UNIT (DPU) OF THE SOLAR 8000I. THE FAILURE IS FROM THE PERIOPERATIVE INFORMATION MANAGEMENT (PIM) SOFTWARE THAT IS CURRENTLY BEING USED IN THE OR. APPARENTLY, THE NIBP CUFF SIMPLY STOPS RECORDING THE PATIENTS' BLOOD PRESSURE. IN THIS CASE, IT WAS ABOUT 15 MINUTES. THIS HAS HAPPENED A LOT AT THE HOSPITAL IN THE OR'S AND WHILE THE MANUFACTURER HAS ASSURED US A FIX IS IN THE PIPELINE, IT HAS BEEN DECIDED TO SUBMIT A REPORT ON EVERY OCCASION OF NIBP FAILURE.E-MAIL: I HAD ANOTHER OCCASION OF A FAILED SOLAR 8000I NIBP MONITOR. THAT IS, IT WENT FROM AUTO NIBP INTERVAL TO OFF WITH NO BUTTONS PRESSED. THE BP STAYED ON THE BIG SCREEN SO IT WAS NOT OBVIOUS.I DID NOTICE IT AND OBTAINED A BLOOD PRESSURE ANOTHER WAY WHILE I MADE SURE THE MONITOR DID NOT ALARM. I THEN RESTARTED THE AUTO NIBP INTERVAL MODE. WE HAVE HAD MANY OF THESE FAILURES AND I SIMPLY WANTED TO DOCUMENT ONE MORE FROM THE ANESTHESIOLOGIST. | 1
2410764 | THE MCKESSON HEMODYNAMIC APPLICATION EXPERIENCED ISSUES THAT PREVENTED THE USE OF THE SYSTEMS. MAIN THREE LABS INOPERABLE NECESSITATED DIVERTING POTENTIAL EMERGENCIES (ST SEGMENT ELEVATION MYOCARDIAL INFARCTION - STEMI) TO ANOTHER FACILITY THEREBY POTENTIALLY DELAYING CARE. FORTUNATELY, NO PATIENTS HAD TO BE DIVERTED DURING THE 3 HOURS SOFTWARE WAS INOPERABLE. THIS HOSPITAL IS A LEVEL 1 TRAUMA AND CARDIAC FACILITY.======================MANUFACTURER RESPONSE FOR HORIZON CARDIOLOGY HEMO, IMAGING APP & HEMO APP, HORIZON CARDIOLOGY HEMO, IMAGING APP & HEMO APP (PER SITE REPORTER).======================MCKESSON PERFORMING ROOT CAUSE ANALYSIS OF EACH HEMO STATION. LAB 1 HAD A CORRUPT USER PROFILE THAT PREVENTED THE NORMAL ACCOUNT FROM LOGGING IN. LABS 2&3 HAD CORRUPT SQL INSTANCES ON THEIR LOCAL DATABASES. FINAL DETERMINATION OF WHAT CAUSED BOTH CORRUPTIONS IS NOT COMPLETE, BUT IT WAS DETERMINED THAT THE SYSTEMS HAD BEEN "HARD BOOTED" (PRESSING POWER BUTTON TO RESTART) MANY TIMES. THIS IS A KNOWN WAY TO CAUSE SQL CORRUPTIONS. PREVENTATIVE MEASURES ARE TO CONFIGURE ONE PC TO BE A HOT SPARE. MCKESSON COMING ON SITE TO PHYSICALLY REBUILD TO A BASELINE "PRISTINE" BUILD. THIS WILL REMOVE ANY RESIDUAL SETTINGS FROM THE DEVICES. | 1
2413313 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE GREATER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED IN THE RUN DATA FILE AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2414746 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATION TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NO A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE ELEVATED RWBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER NEAR THE END OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2414857 | THE BIOMEDICAL ENGINEERING TECHNICIAN DISCOVERED THAT THE GE ANESTHESIA MACHINE MODEL: AVANCE S5 WITH SOFTWARE VERSION: 6.20 THAT WAS INSTALLED (B)(6) 2010, HAS A PROBLEM WITH PRESSURE VENTILATION MODE'S NOT DEFAULTING PROPERLY. THE THREE MODES AFFECTED ARE PRESSURE SUPPORT WITH APNEA BACKUP MODE (PSVPRO), SYNCHRONIZED INTERMITTENT MANDATORY VENTILATION (SIMV-PC), AND PRESSURE CONTROL VENTILATION (PCV). WHEN USING THESE MODES THE PEAK INSPIRATORY PRESSURE, MAX (PINSP) DOES NOT DEFAULT BACK TO INITIAL DEFAULT SETTINGS WHEN THE CASE HAS ENDED. WHEN SELECTING A "NEW CASE" ON THE MACHINE THE "PINSP" DEFAULTS BACK TO THE LAST CASE SETTING AND NOT THE ORIGINAL DEFAULT FOR THIS MODE. THE ONLY WAY FOR THE "PINSP" TO DEFAULT BACK TO THE DEFAULT SETTING IS FOR THE CLINICIAN TO SHUT THE MACHINE COMPLETELY OFF AND REBOOT THE MACHINE. THE BIOMEDICAL ENGINEERING TECHNICIAN NOTIFIED THE GE TECHNICAL SUPPORT LINE AND MADE THEM AWARE OF THE ISSUE. GE CLINICAL SUPPORT INSTRUCTED BIOMEDICAL ENGINEERING TO CALL IN THE GE TECHNICAL SUPPORT ENGINEER TO EVALUATE THE PROBLEM. THE PROBLEM WAS REPLICATED BY GE TECHNICAL SUPPORT BUT GE HEALTHCARE DOES NOT HAVE A FIX FOR THE PROBLEM AT THIS TIME. BIOMEDICAL ENGINEERING INFORMED THE HEAD OF ANESTHESIA ABOUT THE PROBLEM. THIS IS ONE OF SIXTY ANESTHESIA MACHINES AT THIS FACILITY.======================MANUFACTURER RESPONSE FOR GAS MACHINE, ANESTHESIA, AVANCE S/5 (PER SITE REPORTER).======================BIOMEDICAL ENGINEERING NOTIFIED THE GE TECHNICAL SUPPORT LINE AND MADE THEM AWARE OF THE ISSUE. GE CLINICAL SUPPORT INSTRUCTED BIOMEDICAL ENGINEERING TO CALL IN THE GE TECHNICAL SUPPORT ENGINEER TO EVALUATE THE PROBLEM. THE PROBLEM WAS REPLICATED BY GE TECHNICAL SUPPORT, BUT GE HEALTHCARE DOES NOT HAVE A FIX FOR THE PROBLEM AT THIS TIME. | 1
2414950 | DURING ENDOSCOPY WITH ATTEMPTED PLACEMENT OF A BRAVO PH PROBE, THERE WAS A FAILURE TO DEPLOY. PROCEDURE WAS ABORTED. CAPSULE AND PACKAGING DISPOSED OF. NO IDENTIFIERS AVAILABLE. | 0
2414952 | DURING ENDOSCOPY WITH ATTEMPTED PLACEMENT OF A BRAVO PH PROBE, THERE WAS A FAILURE TO DEPLOY. TWO MORE CAPSULES ATTEMPTED - ALL WITH FAILURE TO DEPLOY (TOTAL ATTEMPTS X 3). PROCEDURE ABORTED. EACH CAPSULE TO BE REPORTED SEPARATELY. REPORT 2 OF 3. | 0
2414966 | DURING ENDOSCOPY WITH ATTEMPTED PLACEMENT OF A BRAVO PH PROBE, THERE WAS A FAILURE TO DEPLOY. TWO MORE CAPSULES ATTEMPTED - ALL WITH FAILURE TO DEPLOY (TOTAL ATTEMPTS X 3). PROCEDURE ABORTED. EACH CAPSULE TO BE REPORTED SEPARATELY. REPORT 1 OF 3. | 0
2414968 | DURING ENDOSCOPY WITH ATTEMPTED PLACEMENT OF A BRAVO PH PROBE, THERE WAS A FAILURE TO DEPLOY. TWO MORE CAPSULES ATTEMPTED - ALL WITH FAILURE TO DEPLOY (TOTAL ATTEMPTS X 3). PROCEDURE ABORTED. EACH CAPSULE TO BE REPORTED SEPARATELY. REPORT 3 OF 3. | 0
2414969 | DURING ENDOSCOPY WITH ATTEMPTED PLACEMENT OF A BRAVO PH PROBE, THERE WAS A FAILURE TO DEPLOY. SECOND CAPSULE ATTEMPTED AND SUCCESSFUL. | 0
2415620 | A CUSTOMER REPORTED INCONSISTENT RESULTS ON THE GALILEO NEO INSTRUMENT WHEN TESTING PATIENT SAMPLES WITH THE CMV (CYTOMEGALOVIRUS) ASSAY. A REVIEW OF THE IMAGE FILES SHOWED CELL BUTTONS WITH HOLES WHICH MAY BE INDICATIVE OF INCONSISTENT PLATE WASHING. HOWEVER, THE INCONSISTENT RESULTS COULD NOT BE DUPLICATED UPON REPEAT TESTING. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2416169 | IT WAS REPORTED THAT DURING PAROXYSMAL ATRIAL FIBRILLATION - PVI, NO INTRACARDIAC ELECTROGRAM (IC) SIGNAL AND THE SURFACE ECG SIGNALS APPEARED ON BOTH CARTO AND EP RECORDING SYSTEM AND CAUSED THE PROCEDURE TO BE EXTENDED SLIGHTLY. THE PATIENT WAS UNDER GENERAL ANESTHESIA FOR 4 HOURS. THE SYSTEM WAS REBOOTED, THE CATHETERS WERE DISCONNECTED FROM THE CONNECTORS, AND THEN RECONNECTED AND ALSO THE ECG CABLES AND ELECTRODES WERE CHANGED, BUT THE PROBLEMS CONTINUED. THEN, THE CARTO ECG WAS DISCONNECTED, AND THE EP RECORDING SYSTEM WAS CONNECTED DIRECTLY TO THE PATIENT. WE WERE ABLE TO CREATE AN ANATOMY OF THE LEFT ATRIUM WHILE WE WERE TRYING TO RESOLVE THE ISSUES WITH THE ELECTROGRAMS, AND ONCE WE HAD COLLECTED THE ANATOMY, THE OPERATOR WANTED TO START ABLATING, AND NEEDED SIGNALS FROM THE LASSO CATHETER. AT THAT TIME, WE REALLY WEREN'T SURE OF THE PIU BECAUSE WE STILL HAD NO SIGNALS AT ALL (PARTLY BECAUSE WE HAD BY THIS POINT DISCONNECTED THE BODY SURFACE ECG), SO AN OPTIMA CIRCULAR CATHETER WAS PLUGGED DIRECTLY INTO THE EP RECORDING SYSTEM, AND JUMPED ACROSS TO A PIN BOX ON THE PIU IN ORDER TO VISUALIZE THE CATHETER. WE THEN NOTICED THAT WE HAD SIGNALS ON CARTO, AND THE CASE WAS COMPLETED WITHOUT FURTHER PROBLEM. (B)(4). THE CABLE WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS; VISUAL INSPECTION AND ELECTRICAL TEST WERE PERFORMED. VISUAL INSPECTION PASSED, THE CATHETER WAS FOUND IN PHYSICAL NORMAL CONDITION. ELECTRICAL TEST FAILED, SHORTING WAS FOUND BETWEEN PINS 21 AND 32 AT ONE OF THE CONNECTORS. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE FOUND. IN ADDITION, DHR REVIEW VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. THE COMPLAINT CONDITION WAS CONFIRMED; DURING TESTING, THE CABLE PRESENTED A SHORTING MALFUNCTION THAT MAY CONTRIBUTE TO THE LOSS OF SIGNALS.><THE CONCOMITANT PRODUCTS: SMART TOUCH: MODEL# D-1336-02-S, LOT # UNKNOWN. CARTO 3: MODEL # M-4800-01, SERIAL # (B)(4). (B)(4). | 1
2417215 | THERMADRAPE THERMAL BLANKET EMITTED A STRONG FOUL SMELL. THE ODOR INCREASED AS PATIENTS USED THEM. WE ATTEMPTED TO PLACE THEM ON APPROXIMATELY 20 PATIENTS. ALL BLANKETS WERE REMOVED FROM STOCK. | 0
2417463 | IT WAS REPORTED THAT NO AUDIBLE ALARM SOUND WAS HEARD ON THE DINAMAP PROCARE 1000 DUE TO A LOOSE CONNECTION BETWEEN THE SPEAKER CABLE AND THE MAIN PC BOARD. THE FIELD SERVICE ENGINEER RECONNECTED THE SPEAKER CABLE, TESTED THE UNIT AND THE AUDIBLE ALARM FUNCTION WAS RESTORED. THERE WAS NO DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2418291 | THE CUSTOMER REPORTED THAT THEY OBSERVED A BLACK PARTICLE (1-2 MM) IN THE PLATELET CONCENTRATE WHEN CHECKING THE PLATELET SWIRL. PATIENT INFORMATION IS NOT AVAILABLE AT THIS TIME. CARIDIANBCT IS AWAITING THE RETURN OF THE DISPOSABLE SET FOR EVALUATION. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2418881 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THE HEMOGLOBIN (HGB) AND HEMATOCRIT (HCT) RESULTS ON THEIR COULTER ACT DIFF 2 ANALYZER WERE NOT MATCHING. WHEN THE SAMPLES WERE RE-RUN, THE HGB AND HCT MATCHED; HOWEVER, THE RBC CHANGED SIGNIFICANTLY ON SOME OF THE SAMPLES. THE RESULTS PROVIDED BY THE CUSTOMER ARE SHOWN IN THE ATTACHMENT. REVIEW OF THE DATA SHOWED RBC VALUES SIGNIFICANTLY LOWER ON RE-RUN FOR SIX OF THE SEVEN REPRESENTATIVE PATIENT SAMPLES PROVIDED BY THE CUSTOMER; RBC WAS HIGHER ON RE-RUN FOR SAMPLE 4. AN INSTRUMENT GENERATED FLAG ( * ) WAS PRESENT FOR PLT AND MPV ON THE INITIAL AND RE-RUN RESULTS FOR SAMPLES 2 AND 6. RDW WAS ALSO FLAGGED ON THE RE-RUN FOR SAMPLE 2 AS WELL AS WBC ON RE-RUN FOR SAMPLE 6. SINCE RBC IS A FACTOR IN CALCULATING MCH, MCHC AND HCT, THESE PARAMETERS WERE ALSO IMPACTED; THIS RESULTED IN THE H/H CHECK FAILS OBSERVED BY THE CUSTOMER. ERRONEOUS RESULTS WERE NOT REPORTED OUTSIDE THE LABORATORY. NO DEATH OR INJURY IS ATTRIBUTED TO THIS EVENT AND THERE WAS NO CHANGE TO PATIENT TREATMENT. PER CUSTOMER, THEY OBTAINED CORRECT RESULTS ON (B)(6) 2011 FOLLOWING THE FSE REPAIRS BY RE-RUNNING THE SAMPLES AND PERFORMING A DELTA CHECK AGAINST A PREVIOUS ANALYSIS OF THE SAME PATIENT'S BLOOD. SAMPLES ARE COLLECTED IN PLASTIC VACUTAINER TUBES AND TESTED IN AN AVERAGE OF 5 MINUTES. CONTROLS ARE RUN DAILY AND WERE RUN BEFORE THE INCIDENT AND RECOVERED WITHIN RANGE. THE UNIT IS CURRENTLY PERFORMING WITHIN PUBLISHED SPECIFICATIONS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE THE DAY OF THE EVENT AND DETERMINED THAT THE HGB AND HCT WERE ERRATIC DUE TO INTERMITTENTLY HIGH RBC. THE FSE OBSERVED AIR IN ONE OF THE RBC FILTERS AND CONSEQUENTLY REPLACED BOTH RBC BATH FILTERS AND VERIFIED THE REPAIR PER ESTABLISHED PROCEDURES. RESULTS MET PUBLISHED PERFORMANCE SPECIFICATIONS. THE ROOT CAUSE FOR ERRATIC RESULTS IS ATTRIBUTED TO AIR IN ONE OF THE RBC FILTERS CAUSING INTERMITTENTLY HIGH RBC. PER BEC LABELING, "BECKMAN COULTER DOES NOT CLAIM TO IDENTIFY EVERY ABNORMALITY IN SAMPLES AND SUGGESTS USING ALL AVAILABLE FLAGGING OPTIONS TO OPTIMIZE THE SENSITIVITY OF THE INSTRUMENT RESULTS. ALL FLAGGING OPTIONS INCLUDE REFERENCE RANGES (H/L), CODES, AND FLAGS. BECKMAN COULTER RECOMMENDS AVOIDING THE USE OF SINGLE MESSAGES OR OUTPUTS TO SUMMARIZE SPECIMEN RESULTS OR PATIENT CONDITIONS. IF ANY FLAG APPEARS, REVIEW THE RESULTS ACCORDING TO YOUR LABORATORY'S PROTOCOL." | 0
2419055 | IT WAS REPORTED THAT THE B650 MONITOR DID NOT ACTIVATE AN AUDIBLE ALARM FOR A PATIENT IN THE ICU (INTENSIVE CARE UNIT) WHEN THE PATIENT'S SPO2 DECREASED UNDER 70% AND HEART RATE DROPPED UNDER 40 (LOW HEART RATE LIMIT WAS SET TO 40). HEALTH CARE PROVIDERS WERE PRESENT IN THE ROOM AT THE TIME OF THE EVENT AND THERE WAS NO REPORTED DELAY IN TREATMENT. THE PATIENT WAS TREATED WITHIN 2 - 3 MINUTES USING O2 BAG VENTILATION AND ADRENALIN. THE PATIENT REPORTEDLY DID NOT GO INTO CARDIAC ARREST AND RECOVERED WITHOUT LONG TERM ADVERSE CONSEQUENCES. INITIAL LOG REVIEW COMPLETED BY GE HEALTHCARE ENGINEERING INDICATES THAT THE SYSTEM DID PROVIDE SPO2 ALARMS DURING THE TIME PERIOD IN QUESTION. THE ALARMS WERE SILENCED BY THE USER. HEART RATE ALARMS WERE NOT PRESENT IN THE LOG FILES PROVIDED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. | 1
2422853 | ONCENTRA ALLOWS TO EXPORT UNAPPROVED PLANS WITHOUT DOSE FOR USER WORKFLOW REASONS. THE UNAPPROVED PLAN WAS USED FOR BEAM SETUP AND CONSEQUENTLY FOR TREATMENT. THIS RESULTED UNINTENDED IN 2 BEAMS WITH 100MU EACH. THIS COULD ALMOST COMPLETELY BE COMPENSATED DURING THE FOLLOWING TREATMENTS, AND THE DEVIATION FROM THE PRESCRIBED DOSE WAS VERY SMALL. ONCENTRA HAS AN APPROVAL MECHANISM AND IT'S CLEARLY STATED IN THE DOCUMENTATION THAT ONLY APPROVED PLANS SHOULD BE USED FOR TREATMENT. THE USER HAS NOT FOLLOWED THESE INSTRUCTIONS. ONCENTRA DOES NOT ALLOW TO APPROVE PLANS WITHOUT DOSE. AS A RESULT OF THE FDA FORM 483 REPORT FEI NUMBER (B)(4) DATED (B)(4) 2011, NUCLETRON BV IS NOW SUBMITTING THIS MDR IN COMPLIANCE WITH THE CORRECTIVE ACTIONS OF THE FDA FORM 483. | 1
2423586 | USE OF NEW COMPUTERIZED RECORD INVOLVES SCANNING IN LAB DATA. THYROID LEVEL WAS ENTERED INTO WRONG PATIENT'S CHART. WHEN SENT FOR REVIEW, I REQUESTED IT TO BE ERRORED AS ON WRONG CHART. IT WAS BUT THERE IS NO DISTINCT MARKING. I RETURNED TO THE CHART A FEW WEEKS LATER TO PREPARE A REFERRAL LETTER AND IN REVIEWING LAB SAW THE ABNORMAL THYROID LEVEL AND SENT NOTE TO MY NURSE TO CALL FAMILY AND START ON THYROID WITH FOLLOW UP LEVEL. I STARTED A VERY LOW DOSE DESPITE A VERY HIGH TSH BECAUSE HE IS OVER (B)(6). MY NURSE WAS SURPRISED AT THE MISSED LAB AND IN REVIEWING COULD NOT FIND THE RESULT. ERRORED SCANS ARE NOT VISIBLE IF USING ONLY ACTIVE DOCUMENT BUT IF VIEWING ALL DOCUMENTS THEY SHOW WITHOUT VISIBLE ERRATA MARKING. I HAVE REQUESTED CERNER MAKE CHANGES BUT THEY INFORM ME TO REQUEST AN ENHANCEMENT. I'VE REQUESTED THE PATIENT RETURN FOR REPEAT LAB BUT HE HAS NOT. I SUSPECT HE WANTS TO FIND A DOCTOR WHO IS NOT USING A COMPUTER AT THIS POINT. IT TOOK SO LONG TO DO HIS VISIT WITH THE NEW SYSTEM AND NOW THIS ERROR, MAKES ME WONDER WHY I AGREED TO START USING THEM. | 1
2423983 | A CUSTOMER REPORTED A TECHNICIAN STUCK HIMSELF ON THE NEEDLE INSIDE THE BACT/ALERT ADAPTER CAP (PN 279012) WHILE TRYING TO GET THE INSERT OUT OF THE ADAPTER TO DRAW BLOOD CULTURES. THE TECHNICIAN WAS IN THE PROCESS OF INSERTING A VACUUM CONTAINER TUBE WHEN THEIR HAND SLIPPED INTO THE ADAPTER CAP CAUSING THEIR FINGER TO BE STUCK BY THE SHEATHED NEEDLE. THE NEEDLE IS USED TO PIERCE THE SEPTUM IN BACT/ALERT BLOOD CULTURE BOTTLES. THERE IS A THIN RUBBER SHEATH/COVER THAT RETRACTS WHEN THE ADAPTER CAP IS PLACED ON THE CULTURE BOTTLE. THE BACT/ALERT ADAPTER CAP IS USED TO FACILITATE THE FILLING OF BLOOD INTO A BACT/ALERT BLOOD CULTURE BOTTLE. THE BACT/ALERT BLOOD COLLECTION ADAPTER INSERT IS USED FOR POSITIONING A TRANSFER NEEDLE TO FACILITATE BLOOD COLLECTION INTO BLOOD COLLECTION TUBES. THE ADAPTER CAP PACKAGE INSERT RECOMMENDS THE TECHNICIAN TO FOLLOW THE FACILITY'S APPROVED PROCEDURES FOR USE FOR THE BACT/ALERT BLOOD COLLECTION ADAPTER CAP AND INSERT. THE TEST PROCEDURE STATES TO PLACE THE ADAPTER CAP ON BACT/ALERT BLOOD CULTURE BOTTLES FIRST, AND, IF ADDITIONAL BLOOD IS REQUIRED FOR OTHER TESTS, TO INSERT THE ADAPTER INSERT, LOCK INTO PLACE AND NOW THE APPARATUS IS COMPATIBLE WITH VACUUM COLLECTION TUBES. AN INVESTIGATION HAS BEEN OPENED FOR THIS INCIDENCE. THE ORIGINAL KIT WAS DISPOSED AND NOT AVAILABLE FOR INVESTIGATION, BUT THE CUSTOMER DID SEND SIMILAR KITS TO BE INSPECTED. THE CUSTOMER DOES NOT KNOW THE LOT NUMBER AS THE ADAPTER CAPS ARE ADDED TO A COMMUNAL BIN. THE EVENT IS REPORTABLE TO THE FDA AS AN ADVERSE EVENT BECAUSE THE CUSTOMER HAD TO SEEK MEDICAL TREATMENT TO DETERMINE EXPOSURE TO THE ADAPTER CAP NEEDLE AND THE CUSTOMER WOULD NOT PROVIDE ANY ADDITIONAL TREATMENT WAS ADMINISTERED OR THE RESULTS OF ANY TESTS. A REVIEW OF THE RETURNED ADAPTER CAPS AND ADAPTER CAP INSERTS FOUND NO FUNCTIONAL ISSUES WITH EITHER THE ADAPTER CAP OR THE INSERT. BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER COMPLAINT, THE ROOT CAUSE OF THE INCIDENT WAS DUE TO ACCIDENTAL CUSTOMER ERROR THAT RESULTED IN THE CUSTOMER'S FINGER SLIPPING INSIDE THE ADAPTER CAP AND ADAPTER CAP INSERT AND BEING STUCK BY THE BLOOD COLLECTION SET NEEDLE. THE ADAPTER CAP AND INSERT ARE BIOMERIEUX PARTS AND ARE DESIGNED TO BE COMPATIBLE WITH BLOOD COLLECTION SETS THAT ARE MADE BY OTHER MANUFACTURERS. A REVIEW OF BD'S PACKAGE INSERT INCLUDES A GENERAL PRECAUTION STATEMENT REGARDING THE USE OF SHARPS: HANDLE ALL BIOLOGIC SAMPLES AND BLOOD COLLECTION "SHARPS" (LANCETS, NEEDLES, LUER ADAPTERS, AND BLOOD COLLECTION SETS) IN ACCORDANCE WITH THE POLICIES AND PROCEDURES OF YOUR FACILITY. OBTAIN APPROPRIATE MEDICAL ATTENTION IN THE EVENT OF ANY EXPOSURE TO BIOLOGIC SAMPLES (E.G., THROUGH A PUNCTURE INJURY) SINCE SAMPLES MAY TRANSMIT VIRAL HEPATITIS, HIV (AIDS), OR OTHER INFECTIOUS DISEASES. UTILIZE ANY SAFETY ENGINEERED FEATURE IF THE BLOOD COLLECTION DEVICE PROVIDES ONE. THE BACT/ALERT BLOOD COLLECTION ADAPTER CAP IS A PLASTIC CAP THAT FITS OVER BLOOD CULTURE BOTTLES AND IS USED IN CONJUNCTION WITH A BLOOD COLLECTION SET. BIOMERIEUX DOES NOT PROVIDE OR MANUFACTURE THE BLOOD COLLECTION SET AND THIS MATERIAL IS LISTED IN THE "MATERIALS REQUIRED BUT NOT PROVIDED" SECTION. THE ADAPTER CAP PACKAGE INSERT INSTRUCTS THE USER TO "REFER TO YOUR FACILITY'S APPROVED PROCEDURES FOR USE OF THE BACT/ALERT BLOOD COLLECTION ADAPTER CAP AND INSERT." THE BIOMERIEUX ADAPTER CAP PACKAGE INSERT ALSO CONTAINS STEP BY STEP INSTRUCTIONS FOR THE TEST PROCEDURE ON HOW TO USE THE ADAPTER CAP. STEPS 1-6 PROVIDE INSTRUCTIONS ON HOW TO USE THE ADAPTER CAP WITH BACT/ALERT CULTURE BOTTLES, AND A NOTE FOLLOWS THESE STEPS THAT STATES "IF ADDITIONAL BLOOD IS REQUIRED FOR OTHER TESTS, PLACE THE ADAPTER INSERT INTO THE ADAPTER CAP AND LOCK INTO PLACE. THIS MAKES THE CAP COMPATIBLE WITH VACUUM COLLECTION TUBES." THE LAST STEP IN THE TEST PROCEDURE SECTION STATES "AFTER BLOOD COLLECTION IS COMPLETE, REMOVE THE ADAPTER CAP FROM THE CULTURE BOTTLE AND THEN REMOVE THE NEEDLE FROM THE PATIENT'S VEIN." THE STEP-BY-STEP INSTRUCTIONS PROVIDE CLEAR INSTRUCTIONS FOR THE USER TO FOLLOW. THE ADAPTER INSERT IS INTENDED TO BE USED AT THE LAST STAGES OF BLOOD COLLECTION IF ADDITIONAL BLOOD IS NEEDED TO BE DRAWN INTO VACUUM COLLECTION TUBES. TO FURTHER CLARIFY THE ADAPTER INSERT SHOULD BE USED AT THE LAST STEP, ADDITIONAL VERBIAGE WILL BE ADDED TO THE ADAPTER CAP PACKAGE INSERT TO INSTRUCT USERS NOT TO REMOVE THE INSERT FROM THE CAP AND TO DISCARD THE BLOOD COLLECTION ADAPTER CAP AND INSERT AFTER USE.><INVESTIGATION IS ONGOING. | 0
2424859 | THE FOLLOWING IS A RECAP OF THE PROGRESS TO DATE REGARDING THE TELEMETRY DROP OUT ISSUE. BACKGROUND: WE HAVE BEEN EXPERIENCING DROPOUT ON THE GE TELEMETRY PATIENT MONITORING SYSTEM FOR SOME TIME. BASED ON THE CALLS WE RECEIVED, THE DROP OUT SEEMED INTERMITTENT AND RANDOM IN TERMS OF TIMING, DURATION AND EFFECTED FLOORS/UNITS. IN MOST CASES THE DURATION WAS BRIEF (< 10 MINUTES) AND TYPICALLY HAD RESOLVED ITSELF BY THE TIME BIOMED RESPONDED TO THE CALL. THESE INSTANCES WERE REPORTED TO GE AND IN SOME CASES GE CAME ON SITE TO INVESTIGATE BUT COULD FIND NO ISSUES WITH THE SYSTEM. AFTER SEVERAL INSTANCES IN THE SECOND HALF OF 2011, THE HOSPITAL BIOMED AND GE INCREASED THEIR EFFORTS TO IDENTIFY THE CAUSE OF THE DROP OUT. IN THE CASE OF THE HOSPITAL BIOMED, THIS INVOLVED USE OF PROCURING SPECIAL EQUIPMENT, SOFTWARE AND TRAINING ON THE EQUIPMENT. USING THIS EQUIPMENT, WE WERE ABLE TO DOCUMENT THAT THERE WAS A SOURCE OF INTERFERENCE IMPACTING THE L BAND FREQUENCIES IN WHICH THE TELEMETRY FLOOR SYSTEM OPERATES. THE CHALLENGE THEN WAS TO IDENTIFY THE SOURCE OF INTERFERENCE. DUE TO THE SEEMINGLY RANDOM AND INTERMITTENT NATURE OF THE DROPOUT ISSUES THIS WAS AKIN TO THE NEEDLE IN THE HAYSTACK. THE HOSPITAL BIOMED INITIATED AN INTENSIVE SYSTEM MONITORING AND TESTING PROTOCOL INVESTIGATING VARIOUS POTENTIAL SOURCES OF INTERFERENCE (PORTABLE EQUIPMENT, TELEMETRY COMMUNICATIONS SYSTEMS, HVAC SYSTEMS, ETC). FINDINGS: VIA THE DAILY SYSTEM MONITORING, WE OBSERVED THAT THE INTERFERENCE WAS OCCURRING AT A MUCH HIGHER RATE THAN WE HAD PREVIOUSLY UNDERSTOOD. THESE INSTANCES WERE VERY BRIEF AND OFTEN DID NOT IMPACT THE UNIT FREQUENCIES (L BAND) OR RESULT IN TELEMETRY DROPOUT. WE DETERMINED A METHOD FOR LOCATING THE SOURCE AND ULTIMATELY FOUND THAT IT WAS TELEMETRY BOXES (TRANSMITTERS) THEMSELVES THAT WERE THE SOURCE OF INTERFERENCE. SPECIFICALLY, IT SEEMED THAT A TELEMETRY BOX IN A LOW BATTERY STATE WAS THE SOURCE OF INTERFERENCE. WE HAVE CONFIRMED THIS ON SEVERAL OCCASIONS. THROUGHOUT THE COURSE OF OUR INVESTIGATION WE HAVE BEEN IN CONTACT WITH GE AND RELAYING OUR FINDINGS. GE IS WORKING TO DETERMINE THE ROOT CAUSE OF THIS ISSUE. THEY STATE THAT NO REPORTS OF THIS TYPE OF ISSUE HAVE BEEN RECEIVED FROM OTHER SITES USING THE L BAND TELEMETRY SYSTEM. NEXT STEPS: 1) HOSPITAL BIOMED HAS OBTAINED LOANER TELEMETRY BOXES AND SENT TELEMETRY BOXES EXHIBITING THE ISSUE TO GE FOR ANALYSIS AND TESTING. 2) HOSPITAL BIOMED HAS IDENTIFIED THE FOLLOWING STEPS FOR CLINICAL STAFF TO MITIGATE THE TELEMETRY BOX LOW BATTERY STATE. IMPORTANT TELEMETRY BOX USE AND STORAGE INFO: DO NOT STORE THE TELEMETRY BOXES WITH THE BATTERIES INSIDE. USE NEW BATTERIES FOR EACH PATIENT. CHANGE BATTERIES AS SOON AS INDICATED BY TELEMETRY BOX LED OR CHANGE BATTERY ALARM AT THE CENTRAL/DECENTRAL STATION. 3) OUR CARDIOVASCULAR (CV) TEAM TECHNICIANS WILL BE CHECKING THE TELEMETRY BOX STORAGE DRAWERS DURING ROUNDS AND REMOVING BATTERIES FROM ANY STORED BOXES. 4) BIOMED PLANS TO REPORT THE ISSUE AND OUR FINDINGS TO THE APPROPRIATE BODIES (FDA-MAUDE, MEDSUN, ECRI). MANUFACTURER RESPONSE FOR TRANSMITTER, PHYSIOLOGICAL, CARESCAPE T14 (PER SITE REPORTER).====================== THE MANUFACTURER IS WORKING WITH US HERE AT THE HOSPITAL TO ATTEMPT TO FIGURE OUT THE ISSUE. | 1
2424884 | YESTERDAY AROUND 2:30, WHILE I WAS TRYING TO ENTER THE PATIENTS EDUCATION CHECKS OFF, A SCREEN WOULD POP UP THAT SAID AN ERROR TO THE APPLICATION OCCURRED. WHEN THIS OCCURRED I CLOSED OUT THE CARDIAC SCIENCE APPLICATION AND THEN REOPENED IT. WHEN I REOPENED IT, THREE PATIENTS WERE RESTORED, BUT I HAD TO RE-ADMIT THE OTHER THREE IN THE CLASS. I WENT BACK IN AND TRIED TO ENTER THE EDUCATION THREE ADDITIONAL TIMES, BUT THE SAME THING OCCURRED. IT RESTORED THE SAME THREE PATIENTS EACH TIME, AND I HAD TO RE-ADMIT THE SAME THREE EACH TIME. THIS IS WHEN WE CALLED YOU CLINICAL ENGINEERING AND THEY SUGGESTED HOLDING OFF ON MAKING AN ENTRY INTO THE EDUCATION DATA, AND TRYING AGAIN TO MAKE ENTRIES BETWEEN CLASSES. AT 3:45 ANOTHER CLINICIAN TOOK OVER FOR ME AND WAS ABLE TO ENTER IN THE EDUCATION FOR THE 3:45 CLASS WITHOUT ANY PROBLEMS. I CALLED BIOMED AND LET CLINICAL ENGINEERING KNOW SHE DIDN'T EXPERIENCE ANY PROBLEMS." THIS HAS NOT HAPPENED BEFORE IN THIS FASHION. THEY HAVE BEEN EXPERIENCED PROBLEMS WHERE THE SYSTEM WOULD SHUT DOWN. THIS WAS CORRECTED WITH A POWER SUPPLY REPLACEMENT, BY CLINICAL ENGINEERING. AFTER TALKING WITH CARDIAC SCIENCE WHO CONFERRED WITH THE CARDIAC SCIENCE, PRODUCT SPECIALIST THEY STATED THAT THIS IS A KNOWN "GLITCH" AND ALTHOUGH THEY ARE AWARE OF THE PROBLEM A RESOLUTION IS NOT NEAR. CARDIAC SCIENCE REPRESENTATIVE SEEMED FRUSTRATED WITH THE SOFTWARE ENGINEER THAT CREATED THE PACKAGE, STATING THAT HE DID NOT ALLOW FOR DIFFERENT WORKFLOWS WITHIN HIS DESIGN. THE ENGINEER IS NO LONGER EMPLOYED WITH CARDIAC SCIENCE. CARDIAC SCIENCE PROVIDED THE FOLLOWING RECOMMENDATION. WHILE MONITORING PATIENTS DURING THE CLASS, DO NOT GO INTO TO THE PATIENT CARE PLAN AND ENTER ANY DATA. ONLY ENTER PATIENT CARE INFORMATION BEFORE OR AFTER THE CLASS. THE OCCURRENCE OF THIS PROBLEM IS INCONSISTENT, BUT DOES APPEAR TO BE RELATED TO ENTERING INFORMATION IN THE PATIENT CARE PLAN, EDUCATION AREA. THIS SEEMS TO CAUSE THE PROBLEM DESCRIBED. CARDIAC SCIENCE DOES NOT PROVIDE A LEVEL OF CONFIDENCE THAT THEY ARE ADDRESSING THIS "KNOWN" PROBLEM. | 1
2424960 | PATIENT IN OR FOR EXPLORATORY LAPAROSCOPY WAS INDUCED AND INTUBATED AFTER ANESTHESIA MACHINE WAS CHECKED AND FOUND TO BE IN FUNCTIONING ORDER PRE-OPERATIVELY. WHEN PATIENT BEING SWITCHED TO MECHANICAL VENTILATION PRIOR TO START OF SURGERY, ANESTHESIA MACHINE BELLOWS NOT MOVING UP AND DOWN AND DEVICE FAILED TO SWITCH FROM MANUAL MODE TO VENTILATOR MODE. PATIENT CONTINUED TO BE MANUALLY VENTILATED AND FAILED ANESTHESIA MACHINE WAS IMMEDIATELY REPLACED WITH ANOTHER FROM THE ENDOSCOPY SUITE. THE FAILED ANESTHESIA MACHINE WAS REMOVED FROM SERVICE AND BIO-MEDICAL ENGINEERING WAS NOTIFIED. THERE WAS NO ADVERSE IMPACT ON THE PATIENT. THE MAINTENANCE VENDOR FOUND THE VENT SELECTOR SWITCH TO BE INOPERATIVE AND THE MALFUNCTIONED SWITCH WAS REPLACED AND THE ANESTHESIA MACHINE WAS RETURNED TO SERVICE. | 0
2425200 | WHEN PT'S SURGERY COMPLETED AND DRAPES TAKEN DOWN. THE PT'S RIGHT FOREARM WAS NOTED TO BE EDEMATOUS DISCOLORED AND BLISTERED. THE SETTING WAS ON MANUAL AND WAS NOT BEING USED AT THE TIME, THE CUFF WAS ALSO DISCONNECTED APPROX. 10 MIN BEFORE THE DRAPES CAME OFF AND THE CUFF DID NOT DEFLATE. THERE WERE IVS PRESENT IN THE RIGHT FOREARM/HAND , AN ARTERIAL LINE WHICH WAS NOT FUNCTIONING WELL. THE PT REQUIRED A FASCIOTOMY AND A VAC DRESSING WAS APPLIED. (B)(4). | 0
2426718 | PATIENT, GRAVIDA 2, PARA 1 CAME IN FOR CAESERIAN SECTION. AFTER A CAESAREAN SECTION, AS THE DRAPES WERE TAKEN OFF, A SMALL BURN WAS NOTED ON THE PATIENT'S ABDOMEN WHERE THE ULTRASOUND DEVICE WAS LOCATED. IT WAS DISCOVERED THAT THE METAL BELT CLASP MIGHT BE THE SOURCE OF THE BURN. THE BURN WAS APPROXIMATELY 1-2 MM IN DIAMETER AND SURGEON APPLIED DERMABOND TO SEAL THE BURN. | 0
2428257 | PRODUCT ANALYSIS WAS COMPLETED ON THE RETURNED HANDHELD COMPUTER AND WAND. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD COMPUTER. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. FOLLOW UP WITH THE REPORTER INDICATED NO ISSUES WERE OCCURRING WITH THE OLD VNS PROGRAMMING WAND AND NEW HANDHELD COMPUTER SYSTEM, RULING OUT ANY PROBLEMS WITH THE VNS PROGRAMMING WAND.><REPORTER INDICATED A VNS HANDHELD COMPUTER WITH 8.0 SOFTWARE WAS NOT PERFORMING AS INTENDED AND WAS DISPLAYING 'CHECK SUM' MESSAGES. THE SUSPECT COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE PENDING PRODUCT ANALYSIS. . | 0
2430296 | DURING A TRANSORAL ROBOTIC-ASSISTED SURGICAL PROCEDURE, THE ROBOTIC MONOPOLAR CAUTERY WAS BEING USED AND SUDDENLY A SMALL FLAME APPEARED AT THE SURGICAL SITE. MD IMMEDIATELY STOPPED SURGERY, PA WITHDREW INSTRUMENT, AND STERILE SALINE WAS POURED INTO THE FIELD EXTINGUISHING THE FLAME. IT APPEARED THE DISPOSABLE CAUTERY SPATULA TIP WAS WHAT HAD CAUGHT ON FIRE AS IT WAS BURNED AND A PIECE FELL OFF - IT WAS REMOVED FROM MOUTH. MOISTENED RAYTEC WERE IN PLACE SURROUNDING THE MOUTH AND RETRACTORS. CHARGE NURSE AND MDA (MD ANESTHESIOLOGIST) NOTIFIED. MOUTH AND THROAT WERE EXAMINED BY SURGEON AND MDA. NO APPARENT HARM TO PATIENT. CAUTERY TIP WAS REMOVED AND REPLACED WITH A NEW ONE. PATIENT WAS BEING VENTILATED WITH 100% OXYGEN PRIOR TO INCIDENT DUE TO LOW SPO2 WHEN N2O WAS USED AND AFTER INCIDENT WATER WAS POURED INTO THE MOUTH AND THROAT AND A SMALL CUFF LEAK WAS NOTICED. PT STABLE AND CASE CONTINUED WITHOUT INCIDENT AFTER ALL PARTIES CONCURRED IT WAS SAFE TO CONTINUE. FAMILY NOTIFIED BY MD AFTER PROCEDURE. | 0
2431021 | PT WAS SEEN IN A PHYSICIAN OFFICE. THE OFFICE USES THE SAME COMPUTER SYS AS THE HOSP SO ALLERGY INFO ENTERED AT THE OFFICE IS AVAILABLE WHEN A PT IS ADMITTED TO THE HOSP. AT HER OFFICE VISIT, THE PT REPORTED THAT SHE HAD SEASONAL ALLERGIES. THE RN ENTERED HER ALLERGY AS SEASONALE (LEVONORGESTREL-ETHYNYL ESTRADIOL). WHEN THE PT WAS LATER ADMITTED TO THE HOSP, THE MEDICATION RECONCILIATION TECH CLARIFIED WITH THE PT THAT SHE HAD NO NEED FOR ORAL CONTRACEPTIVES SINCE SHE HAD NO UTERUS BUT THAT SHE DID EXPERIENCE SEASONAL ALLERGIES. NOTHING HAPPENED TO THE PT, BUT ONE COULD IMAGINE A SCENARIO WHERE THE PT'S ORAL CONTRACEPTIVE OR HORMONE REPLACEMENT THERAPY WAS AVOIDED OR DISCONTINUED INAPPROPRIATELY BASED ON INCORRECT ALLERGY INFO. THE COMPUTER SYS IS DESIGNED SUCH THAT IF YOU TYPE SEASONAL AND PRESS ENTER TO SEARCH IT DEFAULTS TO SEASONALE. IT IS EASY TO ENTER THROUGH THE VERIFICATION SCREEN. THE PT ALSO REPORTED ALLERGY TO CAT HAIR AND THE ONLY CHOICE WAS TO ENTER THAT SHE WAS ALLERGIC TO CAT HAIR STANDARD EXTRACT. THE COMPUTER IS NOT DESIGNED SUCH THAT SEASONAL ALLERGY OR ALLERGY TO CAT HAIR CAN BE ENTERED. YOU MUST PICK SOMETHING FROM THE LIST. THE SEARCH FEATURE ON THE ALLERGY ENTRY FIELD IS DESIGNED SUCH THAT IF YOU TYPE BEES AND SEARCH, YOU WON'T FIND BEE. RECOMMENDATION: ADD COMMON ALLERGIES TO THE DATA BASE. ADD BOTH CAT AND CATS WITH SIMILAR REDUNDANCY FOR COMMON AND SERIOUS ALLERGENS (IE PEANUT, PEANUTS, NUT, NUTS) TO FACILITATE ACCURATE REPORTING. AND OF COURSE BE MORE CAREFUL AND PROVIDE MORE EDUCATION. (B)(6) MEDICATION ERRORS REPORTING PROGRAM. TYPE OF STAFF MADE INITIAL ERROR: NURSE. PT COUNSELING PROVIDED: UNK, ACCESS NUMBER: (B)(4). | 1
2431076 | ISMP MEDICATION ERRORS REPORTING PROGRAM. IN OUR HOSPITAL WE HAVE TWO COMPUTER SYSTEMS WHERE ALLERGIES ARE LISTED. ONE IS USED FOR THE INPATIENT ADMISSIONS AND ONE IS USED IN THE EMERGENCY ROOM. THERE IS AN INTERFACE ISSUE WITH THESE COMPUTER SYSTEMS AND THEY DO NOT COMMUNICATE WITH ONE ANOTHER. ALSO, WE CHANGED THE INPATIENT PHARMACY COMPUTER SYSTEM IN XXXXXXXXX 2010. WHEN A PATIENT IS ADMITTED NOW, IF THEY WERE PREVIOUSLY ADMITTED THE OLD SERIES COMPUTER SYSTEM, ALLERGY INFORMATION FLOWS INTO THE NEW (B)(6) PHARMACY COMPUTER SYSTEM AND (B)(6) CLINICAL (HOSPITAL WIDE SYSTEMS) AND ALLERGIES ARE UPDATED AND VERIFIED BY PHARMACISTS. HOWEVER, NEITHER OF THESE SYSTEMS COMMUNICATES (NOT INTERFACED) WITH THE ER SYSTEM. WE HAD A PATIENT ADMITTED TO OUR HOSPITAL ON (B)(6) 2010 WHO HAD AN ALLERGIC REACTION TO AVELOX (MOXIFLOXACIN) DURING THIS ADMISSION. AT THAT TIME THE PATIENT'S ALLERGIES WERE UPDATED IN THE SERIES COMPUTER SYSTEM AND WERE ADDED TO THE PATIENT'S FACESHEET WHICH IS PRINTED WITH EACH SUBSEQUENT ADMISSION. THE PATIENT WAS BROUGHT TO THE ER ON (B)(6) 2011 (ABOUT 6 MONTHS LATER) AND AGAIN RECEIVED AVELOX DESPITE HAVING THE ALLERGY LISTED IN THE SERIES COMPUTER SYSTEM AND ON THE FACESHEET. THE PATIENT DID NOT HAVE ALLERGIES LISTED IN THE ER COMPUTER SYSTEM AT THIS TIME. THE PATIENT WAS BROUGHT INTO THE ER AGAIN ON (B)(6) 2011 (ABOUT A MONTH AFTER) AND AGAIN RECEIVED AVELOX. SHE HAD A THIRD ALLERGIC REACTION AND HAD TO BE TREATED AT THAT TIME FOR THIS REACTION. AGAIN THE PATIENT'S ALLERGIES WERE LISTED IN THE HOSPITAL SYSTEM AND ON THE FACE SHEET BUT NOT IN THE ER SYSTEM. OF NOTE, THE PATIENT WAS SEEN IN OUR ER 13 TIMES BETWEEN THE INITIAL ALLERGIC REACTION ON (B)(6) 2010 AND THE MOST RECENT EVENT ON (B)(6) 2011. OBVIOUSLY, THERE WERE MANY OPPORTUNITIES FOR ALLERGY INFORMATION TO BE RECONCILED IN THIS PATIENT THAT WAS MISSED. WE HAVE SINCE CHANGED OUR PROCESS, ER NURSES ARE NOW REQUIRED TO CHECK ALLERGIES IN BOTH THE ER SYSTEM AND THE HOSPITAL SYSTEM FOR ANY PATIENT WHO PRESENTS. WE THOUGHT IT APPROPRIATE TO REPORT THIS BECAUSE WITH MANY HOSPITALS GOING TO MORE ELECTRONIC SYSTEMS IT IS VERY IMPORTANT TO REALIZE THAT THERE MAY BE ISSUES WITH PATIENT INFORMATION BEING COMMUNICATED BETWEEN THE SYSTEMS. IT IS ALSO VERY IMPORTANT TO EDUCATE STAFF ON THESE ISSUES AND PUT IN PLACE POINTS OF CHECK (E.G. CHECKING THE MULTIPLE SYSTEMS WHERE INFORMATION MAY BE DOCUMENTED, INTERVIEWING THE PATIENT WITH EACH ADMISSION TO OBTAIN CURRENT ACCURATE MEDICAL INFORMATION, TELLING THE PATIENT THE NAMES OF THE MEDICATIONS YOU ARE ADMINISTERING, ETC) TO PREVENT THESE ERRORS FROM OCCURRING. (B)(4). | 1
2431169 | IT WAS REPORTED THAT A LOSS OF PT MONITORING OCCURRED WHEN THE CIC (CENTRAL STATION) REBOOTED INTERMITTENTLY DUE TO A PRESUMED MOTHERBOARD FAILURE. PTS WERE WITHOUT MONITORING FOR APPROX FIVE (5) TO SIX (6) HOURS. NO INJURY OR DELAY IN TREATMENT WAS REPORTED. A GE HEALTHCARE FIELD SERVICE ENGINEER REPLACED THE MOTHERBOARD FOLLOWING THE EVENT. THE CIC THEN OPERATED AS EXPECTED, PASSING ALL TESTS. GE HEALTHCARE'S INVESTIGATION INTO THE ROOT CAUSE OF THE REPORTED OCCURRENCE IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. | 1
2431915 | IT WAS INITIALLY REPORTED THAT THE PHYSICIAN WAS HAVING TROUBLE WITH THEIR HANDHELD TURNING ON. IT WAS REPORTED THAT THE HANDHELD WAS CHARGED BUT WOULD NOT COME ON. IT WAS CONFIRMED THAT THE LOCK BUTTON WAS NOT ENGAGED. A HARD RESET WAS ATTEMPTED THREE TIMES BUT NOTHING HAPPENED. A COMPANY REPRESENTATIVE WENT TO THE SITE TO TROUBLESHOOT THE ISSUE FURTHER BUT WAS UNABLE TO RESOLVE IT. GOOD FAITH ATTEMPTS FOR PRODUCT RETURN AND ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WAS RECEIVED WITHOUT A BATTERY COVER. AS A RESULT, THE HANDHELD WOULD NOT POWER ON. ADDITIONALLY IT WAS IDENTIFIED THAT THE MAIN BATTERY WAS INSERTED BACKWARDS. THE CONDITION CAN PREVENT HANDHELD FROM OPERATING ON BATTERY POWER. NO FURTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2432861 | IT WAS REPORTED THAT THE CLINICAL INFO CENTER (CIC) WENT DOWN AND NO BACK UP CIC WAS AVAILABLE ON THE UNIT. THE HOSP STATED THIS WAS NOTICED IMMEDIATELY, HOWEVER, THE FAILURE WAS UNEXPECTED. A TOTAL OF 24 TELEMETRY PTS WERE WITHOUT TELEMETRY MONITORING FOR 30 MINUTES. THE EVENT DID NOT CONTRIBUTE TO A DEATH OR SERIOUS INJURY OR REQUIRE MEDICAL INTERVENTION. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2432926 | A CUSTOMER REPORTED THAT A SPRING THAT IS IN THE EQUIPMENT WAS BROKEN. THIS CAUSED THE RESULTS TO BE ALTERED FOR ONE PATIENT'S EXAM. THE EXAM WAS COMPLETED WITH THE SAME SYSTEM, AND THERE WAS NO IMPACT TO THE PATIENT. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. (B)(4). | 1
2432932 | A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE REACTION WHEN TESTING A PATIENT SAMPLE WITH CAPTURE-R READY SCREEN (CRRS) ON THE GALILEO ECHO INSTRUMENT. REVIEW OF THE BATCH FILES SHOWS AS FOLLOWS: SAMPLE ID (B)(4) PERFORMED IN BATCH (B)(4) WITH (B)(4) EXP 01/12/2012 RESULTED AS NEGATIVE. SCREENING CELL 1=NEG VISUALLY APPEARS NEGATIVE. SCREENING CELL 2 (FYA POS)=NEG VISUALLY APPEARS VERY WEAK POS WITH A DIFFUSED CELL BUTTON WITH SLIGHT RED CELL ADHERENCE AROUND THE BUTTON. SCREENING CELL 3 (K POS) = NEG VISUALLY APPEARS POSITIVE WITH A DIFFUSED CELL BUTTON AND RED CELL ADHERENCE AROUND THE BUTTON. OTHER SAMPLES PERFORMED IN THE SAME BATCH RESULTED AS EXPECTED. SAMPLE ID (B)(4)PERFORMED IN BATCH (B)(4) WITH (B)(4) EXP 01/12/2012 RESULTED AS NEGATIVE. SCREENING CELL 1=NEG VISUALLY APPEARS NEGATIVE. SCREENING CELL 2 (FYA POS)=NEG VISUALLY APPEARS VERY WEAK POS WITH A DIFFUSED CELL BUTTON WITH SLIGHT RED CELL ADHERENCE AROUND THE BUTTON. SCREENING CELL 3 (K POS) = NEG VISUALLY APPEARS POSITIVE WITH A DIFFUSED CELL BUTTON AND RED CELL ADHERENCE AROUND THE BUTTON. OTHER SAMPLES PERFORMED IN THE SAME BATCH RESULTED AS EXPECTED. CUSTOMERS WERE INFORMED OF THE NEED TO INSPECT ALL NEGATIVE RESULTS WITH CRRS AND CAPTURE-R READY ID IN CUSTOMER COMMUNICATION (B)(4). | 0
2433669 | A CUSTOMER CONTACTED HAEMONETICS ON (B)(6) 2009 TO REPORT THAT THE LINE SENSOR OF A MCS+ MACHINE WAS NOT WORKING AFTER AN OPERATOR TOUCHED THE LINE SENSOR AND IT SHOCKED HER. NO OPERATOR OR DONOR INJURY WAS REPORTED. A CUSTOMER CONTACTED HAEMONETICS ON (B)(6) 2009 TO REPORT THAT THE LINE SENSOR OF A MCS+ MACHINE WAS NOT WORKING AFTER AN OPERATOR TOUCHED THE LINE SENSOR AND IT SHOCKED HER. NO OPERATOR OR DONOR INJURY WAS REPORTED. FIELD SVC WAS SENT TO THE SITE TO EVALUATE THE DEVICE. UPON EVAL OF THE DEVICE LINE SENSOR WAS READING 0 AND COULD NOT BE CALIBRATED. THE LINE SENSOR WAS REPLACED AND CALIBRATED. THE MACHINE IS NOW READY FOR USE. THE PRODUCT ISSUE IDENTIFIED IN THIS REPORT WAS IDENTIFIED BY HAEMONETICS DURING A RETROSPECTIVE REVIEW OF THE COMPANY'S SVC RECORDS. THE USER ALLEGES THAT THEY WERE SHOCKED BY THE LINE SENSOR WHEN TOUCHING IT. THE AFFECTED PART WAS REPLACED. NO SERIOUS INJURY RELATED TO THE ALLEGATION WAS CONFIRMED, THEREFORE AN MDR IS BEING FILED. (B)(4). | 0
2434663 | CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT THERE WAS A LEAK OF CLEAR FLUID MIXED WITH BLOOD DRIPPING INTO THE SLIDE TRAY UNDER THE COULTER LH 750 SLIDEMAKER. CUSTOMER REPORTED THAT PATIENT RESULTS WERE NOT AFFECTED. THERE WAS NO REPORT OF ANY ADVERSE EVENT OR INJURY REQUIRING MEDICAL INTERVENTION OR PATIENT TREATMENT. BEC FIELD SERVICE ENGINEER (FSE) REPLACED CUT TUBING AT PINCH VALVE #20. BEC IDENTIFIER FOR THIS COMPLAINT IS (B)(4). | 0
2434947 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THE COLLECTION BEING PERFORMED WAS A DOUBLE PLATELET PRODUCT AND RBC COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2438425 | PATIENT'S MONITOR DID NOT ALARM; HOWEVER, IT WAS NOTED THAT THE PATIENT HAD AN ERRATIC RHYTHM, WHICH WAS OBVIOUSLY NOT THIS PATIENT. THIS HAS HAPPENED MULTIPLE TIMES WITH DIFFERENT PATIENTS ON THIS DATE. BIOMED ADVISED AND ARE WORKING ON THE SITUATION. APPARENTLY THIS IS A WIDESPREAD ISSUE WITH THE DASH MONITORS WITH SOFTWARE VERSION 5.4. ADDITIONAL FACT FINDINGS: 1- REWIND BACK TO THE ORIGINAL EVENT OF LAST SUMMER. IF YOU ANALYZE THE FREQUENCY OF THE QRS'S ON THE GRAPH PAPER YOU WOULD COME TO THE CONCLUSION THAT IT IS NOT HUMANLY POSSIBLE, NOT EVEN WITH A PEDIATRIC PATIENT. THIS ANALYSIS WAS CONFIRMED TO BE CORRECT BY GE. 2- THE SYMPTOMS FOR THESE EVENTS ARE CONSISTENT WITH THE ORIGINAL EVENT. GE HAS ACKNOWLEDGED THE TWO ISSUES TO BE THE SAME. HOWEVER WHAT APPEARS TO BE DIFFERENT ACCORDING TO GE IS THE CPU BOARD MANUFACTURING AND REVISION VERSIONS. THE ORIGINAL ISSUE WAS RELATED TO CPU BOARDS MANUFACTURED BETWEEN JANUARY 2003 AND JANUARY 2007. THESE EVENTS (SEVEN CURRENT ONES) OCCURRED WITH DASH MONITORS HAVING CPU BOARDS MANUFACTURED POST JANUARY 2007.3- PROBLEM DEFINED: ACCORDING TO GE THIS ISSUE IS DEFINITELY A DASH PROBLEM, HARDWARE RELATED. IT IS NOT PHYSIOLOGICAL/PATIENT ISSUE. IT IS CAUSED BY HIGH NETWORK TRAFFIC OVERHEATING A COMPONENT (ETHERNET CHIP) ON THE CPU BOARD, WHICH IN TURN CAUSES THE WAVEFORM TO BE CONDENSED. THE MONITOR DOES NOT ALARM WHEN THIS HAPPENS, SINCE PROCESSING HAS ALREADY TAKEN PLACE. "ARRHYTHMIA PROCESSING" WOULD HAVE SEEN THE CORRECT PATIENT RHYTHM BUT SOMEHOW WAS DISPLAYED INCORRECTLY BY THE DASH MONITOR.4- ONE OF THE CHALLENGES RELATES TO THE NEWER CPU BOARD TYPE WHICH DO NOT ACCEPT ORIGINAL SOFTWARE VERSION. THE LATEST SOFTWARE VERSION IS NOT THE SAME AS OLDER VERSION CAUSING POTENTIAL FOR USER INTERFACE ERRORS. | 1
2439175 | THE CUSTOMER WAS COLLECTING MONONUCLEAR CELLS (MNCS) AND PLASMA FROM A PROSTATE CANCER PT. AT 83 MINUTES INTO THE PROCEDURE, THERE WAS A 'LEAK IN THE CENTRIFUGE' ALARM. THE OPERATOR TERMINATED THE PROCEDURE. SHE COLLECTED 79 ML OF MNCS AND 150 ML OF PLASMA. NO RINSEBACK WAS PERFORMED, THE PT WAS DOING FINE AND WENT HOME. THE PRODUCT WAS SENT TO THE LAB TO CHECK FOR CULTURES. THE DOCTOR WILL MAKE THE DECISION ON THE TRANSFUSION OF THE PRODUCT. THE PT IDENTIFIER IS NOT AVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFO PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR DEATH OR INJURY OCCURRED IN RELATION TO THE PT TREATMENT. THE LEAK DOES NOT POSE A RISK OF CONTAMINATION AS THIS IS AN OPEN SET. (B)(4). INVESTIGATION: THE SPECTRA CHANNEL AND LOOP ASSEMBLY WERE RETURNED FOR INVESTIGATION. THE CHANNEL AND LOOP CONTAINED BLOOD. PRESSURE WAS APPLIED TO THE CHANNEL. A LEAK WAS LOCATED ON THE SINGLE SIDE OF THE CONTROL Y. A BOND VOID WAS FOUND BETWEEN THE SINGLE SIDE TUBING AND THE CONTROL Y. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2439229 | PATIENT HAS RADIATION BURN TO BACK FROM PROLONGED EXPOSURE DURING A CATH LAB PROCEDURE - FLUORO TIME WAS 139 MINUTES. ADDITIONAL INFORMATION ABOUT THE RADIATION DOSE IS NOT YET AVAILABLE - WILL SEND IN AN UPDATED REPORT ONCE WE RECEIVE THE DATA FROM GE. | 0
2442838 | THERE ARE FOUR NEW GENERIC THERAPEUTIC CLASSIFICATION (GTC) DRUG CODES ADDED TO THE MEDICATIONS DATABASE BY A THIRD-PARTY DATA SOURCE CAUSING THEIR ASSOCIATED AGENTS TO NOT DISPLAY ON THE ECAREMANAGER "MEDICATIONS VIEW" SCREEN, EVEN THOUGH THE MEDICATION DATA EXISTS IN THE DATABASE AND ARE USED APPROPRIATELY BY OTHER PARTS OF THE SYSTEM, SUCH AS THRESHOLD PROMPTS AND REPORTS. (B)(4). | 1
2443140 | THERE WAS AN OUTAGE THAT IMPACTED THE TRACEVUE DATA ACQUISITION SERVER. TRACINGS FROM FETAL MONITORS WERE NOT ARCHIVING INTO TRACEVUE FOR SEVERAL MINUTES UNTIL THE SERVERS WERE REBOOTED. FOR THIS PARTICULAR PATIENT TRACING, THE FAST UPLOAD OF DATA APPEARS TO NOT HAVE PERFORMED CORRECTLY. THERE IS NO DATA IN THE ELECTRONIC RECORD DURING THE UP LOAD TIME. TRACEVUE INDICATES THAT THE DATA WAS UPLOADED; THERE IS A GAP DURING THE TIME OF THE "UPLOAD" BUT NO DATA. I VERIFIED THE PAPER TRACING IS COMPLETE. THIS BEHAVIOR OF THE UPLOAD FUNCTION FOR THIS FETAL MONITOR-TRACEVUE CONNECTION IS NOT AS EXPECTED. | 1
2443414 | IT WAS REPORTED THAT FIVE (5) TELEMETRY PATIENTS LOST COMMUNICATION WITH THE CIC CENTRAL STATION FOR APPROXIMATELY TWO (2) HOURS. THE ISSUE WAS IMMEDIATELY NOTED BY HOSPITAL STAFF, HOWEVER NO ALTERNATIVE MONITORING WAS AVAILABLE. NO INJURY OR DELAY IN TREATMENT WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. | 1
2443480 | RECEIVED A REPORT THAT THE MACHINE ALARMED WITH MULTIPLE RETURN LINE AIR DETECTOR (RLAD) ALARMS. THE RETURN LINE AIR DETECTOR ISSUE WAS DISCOVERED DURING CHECKOUT OF THE MACHINE FOLLOWING REPLACEMENT OF THE HARD DRIVE DUE TO A REPORTED BOOT FAILURE. NO PT WAS CONNECTED AND THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE CARIDIANBCT SERVICE TECHNICIAN CHECKED OUT THE MACHINE AND INDICATED THAT IT ALARMED AT PRIME TO ONLY BE READING FLUID. THIS WILL RESULT IN A 'RETURN LINE AIR DETECTOR DETECTED FLUID TOO SOON' ALARM. TROUBLESHOOTING REVEALED THAT THE SENSOR DETECTED FLUID WHEN NONE WAS PRESENT. THE SERVICE TECHNICIAN REPLACED THE RLAD. AN OPTIA RLAD ASSEMBLY WAS RECEIVED FOR EVAL. VISUAL INSPECTION REVEALED NO ANOMALIES. THE RLAD FAILED FUNCTIONAL TEST PROCEDURE. IT DETECTED FLUID AT ALL TIMES AS REPORTED. CARIDIANBCT ENGINEERING WAS ABLE TO DUPLICATE THE REPORTED PROBLEM. ROOT CAUSE: DEFECTIVE RLAD. CORRECTIVE ACTION: THE SERVICE TECH REPLACED THE RLAD. AN INTERNAL CAPA WAS OPENED TO INVESTIGATE RLAD FAILURE ISSUES. THIS REPORT WAS FILED BEYOND THE 30-DAY TIME FRAME DUE TO AN INTERNAL PROCESS ERROR. THE PROCESS HAS BEEN CORRECTED TO PREVENT RECURRENCE. A REVIEW OF OLDER RECORDS WITH THIS FAILURE MODE WAS CONDUCTED AND NO ADD'L EVENTS THAT WERE NOT PREVIOUSLY REPORTED WERE FOUND. | 0
2443488 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION : THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS DID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED RWBC CONTENT REPORTED FOR THIS COLLECTION. SIGNALS IN THE RDF INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THIS. INVESTIGATION EVAL AND CORRECTION ACTIONS ARE IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2443703 | IT WAS REPORTED THAT THERE WAS SWELLING IN THE RIGHT FOREARM OF A PT LEADING TO A FASCIOTOMY. THE SURGEON WHO DISCOVERED THE INJURY STATED THAT THE NIBP CUFF CONTRIBUTED TO THE INJURY. THE NIBP CUFF WAS USED ON THE SAME ARM WHERE AN ARTERIAL LINE AND AN IV WERE INSERTED, AS THE LEFT ARM WAS INVOLVED WITH THE PT'S SURGICAL PROCEDURE. THE HOSPITAL FILED A MEDWATCH REPORT. GE HEALTHCARE RECEIVED A COPY OF THE MEDWATCH REPORT SUBMITTED BY THE HOSPITAL ON (B)(6) 2012. IN THE MEDWATCH REPORT IT IS DOCUMENTED THAT: "THE CUFF DID NOT DEFLATE". ON (B)(6) 2012, GE HEALTHCARE RECONFIRMED WITH UNIVERSITY OF (B)(6) MEDICAL CENTER'S RISK SPECIALIST THAT THE HOSPITAL'S MEDWATCH STATEMENT, "THE CUFF DID NOT DEFLATE", IS INCORRECT. THE HOSPITAL RISK SPECIALIST CONFIRMED TO GE HEALTHCARE THAT THE OPERATING ROOM SURGEON REPORTED TO HER THAT HE FOUND THE INVOLVED NIBP CUFF DEFLATED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
2444076 | CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT A CLEAR LIQUID WAS LEAKING UNDER THE COULTER LH 750 SLIDEMAKER. CUSTOMER REPORTED THAT PATIENT RESULTS WERE NOT AFFECTED. THERE WAS NO REPORT OF ANY ADVERSE EVENT OR INJURY REQUIRING MEDICAL INTERVENTION OR PATIENT TREATMENT. BEC FIELD SERVICE ENGINEER (FSE) REPLACED THE DISPENSE TUBING. THE FSE VERIFIED THE REPAIRS WERE PERFORMED PER ESTABLISHED PROCEDURES. (B)(4). | 0
2444284 | A COMMUNICATION WAS RECEIVED FROM THE ADVERSE INCIDENT CENTRE OF THE (B)(6), STATING THAT AN INCIDENT HAD BEEN REPORTED BY (B)(6) HOSPITAL WITH REGARD TO AN FM830 FETAL MONITOR. THE REPORTED FAULT WITH THE MONITOR WAS "CTG DISPLAYING TRACE WITH FETAL HEART RATE OF 150BPM, BUT BABY WAS DEAD." THIS REPORT IS BEING FILED UNDER EXEMPTION (B)(4) ON BEHALF OF THE MANUFACTURER HUNTLEIGH HEALTHCARE LTD. (REGISTRATION#1000589001). THE UNIT IS BEING RETURNED FROM (B)(6) HOSPITAL TO THE SERVICE DEPARTMENT AT (B)(6) FOR INVESTIGATION. ADDITIONAL INFORMATION WILL BE PROVIDED FOLLOWING THE CONCLUSION OF THE MANUFACTURER'S INVESTIGATION. | 1
2446248 | THE PHILIPS FLEXVISION SYSTEM ALLOWS SIGNALS FROM MULTIPLE VIDEO SOURCES TO BE DISPLAYED IN A VARIETY OF WAYS ON A SINGLE, LARGE SCREEN. IT IS CONTROLLED BY A LAPTOP-SIZED DEVICE THAT IS MOUNTED TO THE TABLE. THE BIG SCREEN IS MOUNTED IN A FRAME WHICH IS SUPPORTED BY A CEILING-MOUNTED TRACK SYSTEM AND TELESCOPING ACTUATOR. THIS MAKES IT FULLY LOCATABLE AND HEIGHT ADJUSTABLE WITHIN THE FIELD OF THE TABLE. AFTER ONLY THREE USES, SOMETHING IN THE ACTUATOR FAILED TO HOLD THE WEIGHT OF THE MONITOR AND FRAME AND THEY CAME DOWN HARD UPON THE FOOT END OF THE TABLE. THANKS TO THE TABLE, THE TELESCOPING ACTUATOR DID NOT FULLY EXTEND BEFORE THE MONITOR WAS STOPPED. THIS LIKELY MEANS THAT THE CEILING TRACK SYSTEM WAS SPARED THE SHOCK OF SUDDENLY STOPPING THE MONITOR'S FALL. IT DID LAND ON THE CONTROLLER DEVICE AND BREAK IT AND ITS ADJUSTABLE SUPPORT ARM BUT THOSE WERE THE ONLY VISIBLY BROKEN THINGS. THE PATIENT WAS IN THE ROOM BUT HAD NOT BEEN TRANSFERRED ONTO THE TABLE YET.PHILIPS USES SKF USA, INC. ACTUATORS FOR THEIR SUSPENDED MONITORS. THEY HAVE CONTACTED SKF ABOUT WHAT APPEARS TO BE A FAILURE OF THEIR PRODUCT.THEY WERE ON SITE VERY QUICKLY. THEY ARE TAKING THE SITUATION VERY SERIOUSLY BECAUSE A FAILURE OF THIS KIND WAS NOT EXPECTED. THERE ARE HUNDREDS OF THESE SYSTEMS THROUGHOUT THE U.S. SO THEY WANT TO UNDERSTAND WHAT HAPPENED AND PREPARE A STRATEGY FOR DEALING WITH THEM. | 0
2446282 | IT WAS REPORTED THAT THE SITE'S DELL HANDHELD COMPUTER SCREEN BECAME FROZEN AFTER INTERROGATION. THE EVENT WAS RESOLVED WITH A HARD RESET. NO ADVERSE EVENTS OCCURRED AS A RESULT OF THE ISSUE. . | 0
2446443 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: THE RDF ANALYSIS DID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED WBC CONTENT REPORTED FOR THIS COLLECTION. SIGNALS IN THE RDF INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYS EXPECTS, POSSIBLY ALLOWING SOME WBC'S TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. THE PLASMA LINE MAY BECOME OCCLUDED IF THE CENTRIFUGE COLLAR IS NOT LOADED INTO THE LATCH IN THE CORRECT ORIENTATION. UNDER THESE CONDITIONS, THE RBC LINE IS ALLOWED TO LIE ON TOP OF THE PLASMA LINE AND, UNDER CERTAIN FLOW CONDITIONS, MAY CAUSE THE PLASMA LINE TO PINCH OFF. THIS IN TURN CAUSES HIGHER FLOW THROUGH THE LRS CHAMBER, WHICH MAY LEAD TO ELEVATED LEVELS OF WBC'S IN THE PLATELET PRODUCT. CORRECTIVE/PREVENTIVE ACTION: ALGORITHMS WILL BE INCORPORATED INTO THE SOFTWARE FOR THE TRIMA ACCEL SYS. THESE ALGORITHMS WILL RECOGNIZE A POTENTIAL ELEVATED WBC COUNT DUE TO THE FAILURE MODE WITNESSED DURING THIS INVESTIGATION. THE SOFTWARE WILL FLAG THE PROCEDURE TO MEASURE OR VERIFY WBC'S IN THE PLATELET PRODUCT. THE NEW ALGORITHMS WILL BE ADDED TO THE NEXT TRIMA EQUIPMENT SOFTWARE UPGRADE, VERSION 6.0.2, WHICH IS IN THE PROCESS OF BEING VALIDATED. ADDITIONALLY, AN INTERNAL CAPA HAS BEEN OPENED TO EVALUATE REPORTS OF ELEVATED WBC'S. | 0
2448424 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING THAT WHILE TROUBLESHOOTING 'NO DIFFERENTIAL RESULTS' ON THEIR COULTER LH 750 HEMATOLOGY ANALYZER, THERE WAS A LEAK IN VALVE VL12A AND IT SPRAYED HER ON HER LAB COAT. SHE WAS WEARING A LAB COAT, GLOVES AND FACE SHIELD AT THE TIME OF THE INCIDENT. THERE WAS NO EXPOSURE TO OPEN WOUNDS OR MUCOUS MEMBRANES AND MEDICAL ATTENTION WAS NOT SOUGHT. NO PATIENT SAMPLES WERE AFFECTED BY THIS EVENT. BEC OFFERED TO SCHEDULE SERVICE; HOWEVER THE CUSTOMER REPLACED THE TUBING ON VALVE VL12 AND STATED THE UNIT IS WORKING FINE. THE ROOT CAUSE OF THE LEAK IS ATTRIBUTED TO THE LOOSE TUBING ON VL12A. | 0
2448430 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING THAT THERE WAS EXCESSIVE BLOOD (1ML) ON THE SLIDES, THE RINSE CUP, AND THE DISPENSE PROBE OF THEIR COULTER LH750 SLIDEMAKER. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES. THE LAB'S EXPOSURE CONTROL/RISK MANAGEMENT PLANS ARE IN PLACE. NO INJURY OR EXPOSURE WAS REPORTED. NO PATIENT RESULTS WERE AFFECTED BY THIS EVENT. A SERVICE VISIT WAS SET UP HOWEVER IT WAS CANCELLED PER REQUEST OF THE CUSTOMER. THE CUSTOMER CARRIED OUT THEIR ROUTINE PROBE WIPE AND RINSE CUP MAINTENANCE AND THE ISSUE WAS REMEDIATED. THE ROOT CAUSE OF THE LEAK IS UNKNOWN, BUT MAY BE ATTRIBUTED TO BUILD UP OF THE PROBE WIPE AND RINSE CUP. THE BEC INTERNAL IDENTIFIER FOR THIS EVENT IS (B)(4). | 0
2448471 | IN THE MICU, WE HAVE HAD TWO SETS OF PRESSURE TUBING BREAK WHILE THE NURSE WAS PRIMING IT. THERE WAS NO PATIENT HARM, AS THE TUBING IS PRIMED PRIOR TO PATIENT USE. | 0
2448882 | DURING A REVIEW OF PROGRAMMING HISTORY, AVAILABLE IN HOUSE, A PROGRAMMING ANOMALY WAS IDENTIFIED. ON (B)(6) 2009, THE PATIENT WAS INTERROGATED AND REPROGRAMMED. A FAULTED DIAGNOSTIC TEST THEN OCCURRED AND A FINAL INTERROGATION WAS NOT PERFORMED. AT THE PATIENT'S NEXT RECORDED APPOINTMENT, (B)(6) 2009, THE PATIENT WAS INITIALLY INTERROGATED AT 0.0MA. THE PATIENT WAS THEN REPROGRAMMED TO INTENDED SETTINGS. IT APPEARS THAT THE FAULTED SYSTEM DIAGNOSTIC TEST ON (B)(6) 2009, LIKELY RESULTED IN A CHANGE TO THE PATIENT'S SETTING WHICH WAS NOT INITIALLY CORRECTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2449098 | IT WAS REPORTED THAT THE CLIENT NOTICED THE BATTERY IMPEDANCE INCREASED THEN "ABRUPTLY" DROPS TO ZERO ON THE REPORT. TECHNICAL SERVICES EXPLAINED THE REASON THIS HAD CHANGED. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2450140 | THE TELEMETRY BOX WAS SENT DOWN WITH A BURNED BATTERY COMPARTMENT (INTERIOR ONLY) AND BURNED BATTERIES. THE BATTERY COMPARTMENT HAS A LOT OF BURN RESIDUE AROUND THE NEGATIVE TERMINAL CONNECTORS. THE BATTERIES HAD OVERHEATED TO THE POINT THAT THE PLASTIC LABEL COVER ON THE BATTERY BLISTERED AND STARTED PEELING AWAY FROM THE BATTERY CELL. THE BATTERIES APPEARED TO BE INSTALLED CORRECTLY DUE TO WHERE THE BURN MARKS WERE ON THE BATTERIES AND BATTERY COMPARTMENT. I CAN'T TELL IF THEY WERE BOTH REMOVED AND REINSTALLED IN THE SAME ORDER AS ORIGINALLY INSTALLED WHEN THE INCIDENT HAPPENED. I DID INSTALL THE ORIGINAL OVERHEATED BATTERIES IN THE TELE BOX AND "THE ON BOTH BATTERIES AND THEY TOUCHED BY THE NEGATIVE TERMINAL SIDE." A NEW TELEMETRY BOX WAS SIMULATED WITH THIS CONDITION WHILE ADMITTED AND RUNNING AT A CENTRAL. THE BATTERIES CAN BE SHORTED, WHILE STILL SHOWING A WAVEFORM ON THE CENTRAL, WITH THE BATTERY DOOR CLOSED. THE BATTERIES WILL OVERHEAT WITHIN 10 MINS TO THE POINT THEY START SMELLING AND COULDN'T BE TOUCHED. THE TEST WAS STOPPED BEFORE ANY DAMAGE TO THE BOX OCCURRED. AFTER EVALUATING THE (B)(4) TELEMETRY BOX, THERE WAS NO OBVIOUS CAUSE FOR THIS CONDITION. INSTALLING THE BATTERIES BACKWARDS IS POSSIBLE BUT THE WAY THE COMPARTMENT IS DESIGNED, IT IS IMPOSSIBLE TO CREATE A COMPLETED CIRCUIT. THERE IS NO PHYSICAL DAMAGE TO THE COMPARTMENT THAT WOULD HAVE ALLOWED THIS TO HAPPEN. THAT WOULD MEAN THAT IMPROPERLY INSTALLED BATTERIES COULD NOT HAVE CAUSED THIS TO HAPPEN. THE NEXT STEP WAS TO CLEAN THE RESIDUE AWAY TO SEE WHAT'S UNDER IT. AFTER CLEANING THE RESIDUE AWAY, THERE WAS NOTHING THAT LOOKED PHYSICALLY WRONG WITH THE BATTERY CONTACTS OR THE COMPARTMENT. COMPARING THE DEFECTIVE TRANSMITTER TO A NEW TELEMETRY BOX, THE NEW BOX HAS NO VOLTAGE ACROSS ANY BATTERY TERMINAL (WITH BATTERIES REMOVED). THE DEFECTIVE BOX MEASURED APPROX .350 VOLTS ACROSS THE 2 NEGATIVE BATTERY TERMINALS (DON'T KNOW WHY). THE BOX WAS ADMITTED TO A PHILIPS TELEMETRY TEST SETUP AND RAN FOR 45 MINS WITH NO HEAT ISSUES AT ALL. ONE POSSIBILITY FOR BATTERY OVERHEATING IS ... | 1
2450142 | THE TELEMETRY BOX WAS BROUGHT DOWN TO THE BIOMEDICAL ENGINEERING DEPARTMENT WITH AN OVERHEATED AND BURNED BATTERIES. THE BATTERY COMPARTMENT HAS A SMALL AMOUNT OF BURN RESIDUE AROUND THE OUTSIDE OF THE NEGATIVE TERMINAL OF THE OUTER BATTERY. THE BATTERIES HAD OVERHEATED TO THE POINT THAT THE PLASTIC LABEL COVER ON THE BATTERY BLISTERED AND STARTED PEELING AWAY FROM THE BATTERY CELL. THE BATTERIES APPEARED TO BE INSTALLED CORRECTLY DUE TO WHERE THE BURN MARKS WERE ON THE BATTERIES AND BATTERY COMPARTMENT. WE INSTALLED THE ORIGINAL OVERHEATED BATTERIES IN THE TELE BOX AND THE TRANSMITTER POWERED UP PROPERLY AND TRANSMITTER ECG TO THE PICC (CENTRAL STATION MONITOR). INSTALLING THE BATTERIES BACKWARDS IS POSSIBLE BUT THE WAY THE COMPARTMENT IS DESIGNED, IT IS IMPOSSIBLE TO CREATE A COMPLETED CIRCUIT. THERE IS NO PHYSICAL DAMAGE TO THE COMPARTMENT THAT WOULD'VE ALLOWED THIS TO HAPPEN. THIS WOULD MEAN THAT IMPROPERLY INSTALLED BATTERIES COULD NOT HAVE CAUSED THIS TO HAPPEN. THERE WAS NOTHING THAT LOOKED PHYSICALLY WRONG WITH THE BATTERY CONTACTS OR THE COMPARTMENT. COMPARING THE DEFECTIVE TRANSMITTER TO A NEW TELEMETRY BOX, A NEW BOX HAS NO VOLTAGE ACROSS ANY BATTERY TERMINAL (WITH BATTERIES REMOVED) AND THE AFFECTED TRANSMITTER DIDN'T EITHER. THE BOX WAS ADMITTED TO A PHILIPS TELEMETRY TEST SETUP AND RAN FOR 75 MINS WITH NO HEAT ISSUES AT ALL. ONE POSSIBILITY FOR BATTERY OVERHEATING IS IF THE BATTERY CELL CASE HAD EXPOSED METAL ON BOTH BATTERIES AND THEY TOUCHED BY THE NEGATIVE TERMINAL SIDE. A NEW TELEMETRY BOX WAS SIMULATED WITH THIS CONDITION WHILE ADMITTED AND RUNNING AT A CENTRAL. THE BATTERIES CAN BE SHORTED, WHILE STILL SHOWING A WAVEFORM ON THE CENTRAL, WITH THE BATTERY DOOR CLOSED. THE BATTERIES WILL OVERHEAT WITHIN 10 MINS TO THE POINT THEY START SMELLING AND COULDN'T BE TOUCHED. THE TEST WAS STOPPED BEFORE ANY DAMAGE TO THE BOX OCCURRED. AFTER EVALUATING THE (B)(4) TELEMETRY BOX, THERE WAS NO OBVIOUS CAUSE FOR THIS CONDITION. | 1
2450282 | ON (B)(6) 2012, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT A MARTEL PRINTER WOULD NOT ACTIVATE. THE PRINTER WAS RECHARGED FOR SEVERAL HOURS BUT TO NO AVAIL. THE CUSTOMER STATED THAT THE BATTERY WAS REMOVED AND WAS HOT TO THE TOUCH. BASED ON THE INFORMATION AVAILABLE AT THE TIME, THERE WERE NO INJURIES ASSOCIATED WITH THIS EVENT. (B)(4) | 1
2451163 | FROM INCIDENT REPORT: APPROXIMATELY 6 MONTHS AGO DURING THE EVENING SHIFT THE AUDIO ALARM FOR THE FETAL MONITORING SYSTEM (OBIX) AT THE L&D NURSES STATION WAS REPORTED TO BE NOT SOUNDING. I ASKED CALLER TO MAKE SURE THAT THE VOLUME WASN'T MUTED OR TURNED ALL THE WAY DOWN ON THE DESKTOP. SHE SAID THAT SHE HAD ALREADY CHECKED IT AND THE VOLUME WAS TURNED ALL THE WAY UP. I CAME IN TO FIND THAT THE AUDIO CABLE FOR THE OBIX ALARM HAD BEEN DISCONNECTED FROM THE DISPLAY. SINCE THE SPEAKERS ARE BUILT INTO THE DISPLAY THE AUDIO ALARM WOULD NOT HAVE SOUNDED WITH THE CABLE DISCONNECTED. I PLUGGED THE AUDIO CABLE BACK INTO THE DISPLAY AND CONFIRMED THAT THE AUDIO ALARM IS NOW SOUNDING. THIS IS THE 3RD TIME THAT THIS CABLE WAS FOUND DISCONNECTED. THE DISPLAY THAT THE AUDIO CABLE PLUGS INTO IS STATIONARY AND THE AUDIO CABLE ITSELF IS PLUGGED IN AT THE REAR OF THE DISPLAY MAKING IT VIRTUALLY IMPOSSIBLE TO HAVE BEEN ACCIDENTALLY UNPLUGGED. I SUSPECT THAT USERS ARE DISCONNECTING THE CABLE TO SILENCE THE AUDIO ALARM. THIS IS A CLASSIC CASE OF ALARM FATIGUE WHERE THE ALARMS ARE SO FREQUENT AND DISRUPTIVE THAT USERS INTENTIONALLY DEFEAT THE ALARM AUDIO. THE ALARM PARAMETERS NEED TO BE MORE CONFIGURABLE TO MINIMIZE THE NUISANCE ALARMS. ESPECIALLY BOTHERSOME IS AN ALARM FOR ECG SIGNAL DROPOUT, WHICH HAPPPENS FREQUENTLY WITH TWINS. THE PREVIOUS SYSTEM HAD A DELAY WHICH WOULD NOT ALARM UNLESS THE SIGNAL DROPPED OUT FOR LONGER THAN A PRESET TIME. OBIX ALARMS INSTANTLY WHEN ECG SIGNAL DROPS OUT.======================MANUFACTURER RESPONSE FOR PERINATAL DATA SYSTEM, OBIX (PER SITE REPORTER).======================SOME ISSUES SHOULD BE RESOLVED WITH NEXT VERSION OF SOFTWARE. | 1
2451185 | AT 0700 THE HOSPITAL'S ACTIVE DIRECTORY SERVER LOST POWER. BY 0730 ALL BEDS WENT INTO A "HOLD" STATUS. ALL MONITORING HAD STOPPED. IT CAME BACK ON LINE AT 0930 WHEN THE ACTIVE DIRECTORY(A/D) SERVER CAME BACK ON LINE. OBIX RELIES ON LIGHT WEIGHT DIRECTORY PROTOCOL (LDAP) THROUGH ACTIVE DIRECTORY FOR OPERATION. THEY ONLY HAVE THE ABILITY TO PROGRAM ONE ACTIVE DIRECTORY SERVER.THE A/D SERVER THAT OBIX CHOSE AS AN AUTHENTICATION SOURCE WAS AT A REMOTE LOCATION. WE SINCE HAD THAT CHANGED TO A LOCAL SERVER. OBIX DOES NOT CURRENTLY HAVE THE CAPABILITY FOR LDAP TO SWITCH OVER TO AN ALTERNATE A/D SERVER IN THE EVENT OF A FAILURE. WE THINK THIS IS A DESIGN FLAW AS AN A/D LOSS BRINGS THE ENTIRE SYSTEM DOWN. THIS HAS HAPPENED AT LEAST TWICE IN THE PAST FEW MONTHS.THE OBIX SYSTEM FUNCTIONS AS THE PRIMARY PATIENT MONITORING SYSTEM FOR A 27 BED L&D UNIT FOR PATIENTS WITH HIGH RISK PREGNANCIES. IF THE SYSTEM FAILS, THERE IS NO WAY TO MONITOR ALL THE PATIENTS IN THE DEPARTMENT.======================MANUFACTURER RESPONSE FOR PERINATAL DATA SYSTEM, OBIX (PER SITE REPORTER).======================THEY INDICATE THE ISSUE WILL BE RESOLVED WITH THE NEXT VERSION OF SOFTWARE. | 1
2452131 | IT WAS REPORTED THAT A LOSS OF TELEMETRY MONITORING OCCURRED FOR 16 PTS FOR APPROXIMATELY TWO HOURS. NO INJURY OR DELAY IN TREATMENT WAS REPORTED. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. | 1
2452585 | A CUSTOMER REPORTED UNEXPECTED NEGATIVE REACTIONS WERE OBTAINED WITH THE ANTIBODY SCREENING ASSAY DURING VALIDATION OF THE GALILEO ECHO INSTRUMENT ((B)(4)). A REVIEW OF THE IMAGE FILES SHOWED THAT CELL 2, WHICH CONTAINS THE K ANTIGEN, RESULTED AS NEGATIVE BUT VISUALLY APPEARED POSITIVE. CONTROLS REACTED AS EXPECTED. THE CUSTOMER WAS REFERRED TO CUSTOMER COMMUNICATION (B)(4) THAT STATES TO VISUALLY VERIFY NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION REACTIONS PRIOR TO REPORTING. AN IMMUCOR FIELD SERVICE ENGINEER PERFORMED THE UNEXPECTED REACTIONS CHECKLIST. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2453953 | DURING THE EARLY MORNING OF A RECENT NIGHT SHIFT, ALL OF THE OBIX BEDS WENT TO A "HOLD" CONDITION, RESULTING IN LOSS OF MONITORNIG FOR THE PATIENTS IN THE 27 BED UNIT. WHEN OBIX TECH SUPPORT WAS CALLED, WE WERE TOLD THAT THEIR TECHNICIAN WAS ON ANOTHER CALL AND TO CALL BACK IN 20MIN IF WE DID NOT HEAR FROM THEM. APPROXIMATELY 25 MINUTES LATER THE BEDS STARTED TO COME BACK ON LINE AND BY 30 MINUTES AFTER THE CALL, ALL BEDS AND CENTRAL MONITORING WAS BACK ON-LINE. OBIX ORIGINALLY ATTRIBUTED THIS FAILURE TO THEIR SYSTEM NOT BEING ABLE TO ACCEPT THE TIME CHANGE PUSH (CHANGE FROM DST TO STANDARD TIME). LATER WE WERE TOLD THIS IS NORMAL BUT SHOULD TAKE LESS THAN ONE MINUTE TO COMPLETE. WE THINK THIS IS ANOTHER DESIGN FLAW AND IS DEFINITELY SOMETHING THAT OBIX SHOULD HAVE CONSIDERED AT INSTALLATION.APPROXIMATELY 12 HOURS LATER ON THE SAME DATE, OBIX BEDS AGAIN WENT TO A BED HOLD STATUS. OBIX TECH SUPPORT WAS CALLED AND THEY ADVISED THAT ALL OF THE OBIX SERVICES HAD SHUT DOWN FOR UNKNOWN REASONS. OBIX STARTED THEM REMOTELY AND 30 MINUTES LATER THE SYSTEM WAS BACK ON LINE. THIS PROBLEM HAS YET TO BE EXPLAINED BY OBIX. IT HAS NOT OCCURRED SINCE.======================MANUFACTURER RESPONSE FOR PERINATAL DATA SYSTEM, OBIX (PER SITE REPORTER).======================SEE EVENT DESCRIPTION. | 1
2462509 | THE BIOMERIEUX BACT/ALERT MICROBIAL DETECTION SYSTEM PRESENTED WITH CODE 710 AND MESSAGE "INCUBATOR MODULE TEMPORARILY UNAVAILABLE WHILE POWER UP INITIALIZATION TAKES PLACE. THE CODE THEN WENT AWAY ONLY TO REPEAT AGAIN SEVERAL MORE TIMES. ERROR CODE 1 "POWER FAULT IN INCUBATION MODULE" ALSO APPEARED. INCUBATOR #5 ASSOCIATED WITH THE CONTROLLER CONTAINED CULTURE BOTTLES FROM 32 APHERESIS COLLECTIONS. CULTURE BOTTLES HAD TO BE RELOCATED TO ANOTHER INCUBATOR. PER THE BIOMERIEUX TECH, THE VOLTAGE ON THE POWER SUPPLY WAS OUT OF TOLERANCE CAUSING THE POWER INTERRUPTION. | 0
2462950 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS DID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED RWBC CONTENT REPORTED FOR THIS COLLECTION. SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THIS. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2464548 | REPORTER INDICATED A VNS HANDHELD COMPUTER WAS NOT PERFORMING AS INTENDED. MANUFACTURER TROUBLESHOOTING WITH A DIFFERENT VNS COMPUTER AND WAND ISOLATED THE COMPUTER SERIAL CABLE AS THE LIKELY SUSPECT COMPONENT. THE COMPUTER SERIAL CABLE HAS BEEN RETURNED AND IS PENDING ANALYSIS.><ANALYSIS OF THE SERIAL CABLE WAS COMPLETED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT A KNOWN GOOD HANDHELD COMPUTER WAS UNABLE TO ESTABLISH COMMUNICATION USING THE RETURNED SERIAL CABLE. VISUAL INSPECTION OF THE SERIAL CABLE IDENTIFIED A BROKEN WIRE CONNECTION ON THE TRANSCEIVER PCB. ONCE THE BLUE WIRE WAS SOLDERED ONTO THE PCB, THE SERIAL CABLE WAS ABLE TO ESTABLISH COMMUNICATION BETWEEN THE COMPUTER AND PROGRAMMING WAND. NO FURTHER ANOMALIES ON THE DEVICE PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2465297 | IT WAS REPORTED THAT THE SOLAR 8000I ALARMED VISUALLY FOR "BRADY" (BRADYCARDIA) BUT DID NOT AUDIBLY ALARM. THE NURSING STAFF REPORTEDLY DID NOT NOTICE THE VISUAL ALARM FOR TEN MINUTES. THE PATIENT SUBSEQUENTLY EXPIRED. ACCORDING TO THE NURSE MANAGER, THE NURSES TYPICALLY TURN THE AUDIBLE ALARM OFF AT THE BEDSIDE BUT NOT AT THE CENTRAL MONITORING STATION. GE HEALTHCARE'S INVESTIGATION INCLUDED A REVIEW OF ECG STRIPS AND LOG FILE ANALYSIS. ECG STRIPS FOR THE PATIENT INVOLVED IN THE REPORTED EVENT INDICATE THAT THE SOLAR 8000I ALARM VOLUME SETTING WAS SET OF OFF STARTING AT 03:14:36 WHEN A PAUSE ALARM WAS PRINTED. THE LOGS INDICATE THAT MULTIPLE ECG HEART RATE AND ARRHYTHMIA ALARMS WERE SOUNDED AT THE CIC AT 60% VOLUME AT ADVISORY, WARNING, AND CRISIS LEVELS. BASED ON THE INFORMATION AVAILABLE, THE SOLAR 8000I AND CIC ALARMED APPROPRIATELY FOR THE PATIENT CONDITIONS. THE CLINICIAN TURNED THE ALARM VOLUME OFF AT THE SOLAR 8000I BEDSIDE PRIOR THE EVENT. NO DETERMINATION CAN BE MADE AS TO WHEN THE SOLAR 8000I ALARM VOLUME WAS TURNED OFF AT THE BEDSIDE. GE HEALTHCARE CONCLUDES THE SYSTEM PERFORMED TO SPECIFICATIONS. | 1
2465886 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS WERE REVIEWED FOR THIS LOT NUMBER. THERE WERE NO ISSUES NOTED THAT WERE RELATED TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED RWBC CONTENT REPORTED FOR THIS COLLECTION. SIGNALS IN THE RDF INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THIS. IT IS ALSO POSSIBLE THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. CORRECTIVE ACTION: ALGORITHMS WILL BE INCORPORATED INTO THE SOFTWARE FOR THE TRIMA ACCEL SYSTEM. THESE ALGORITHMS WILL RECOGNIZE A POTENTIAL ELEVATED WBC COUNT DUE TO THE FAILURE MODE WITNESSED DURING THIS INVESTIGATION. THE SOFTWARE WILL FLAG THE PROCEDURE TO MEASURE OR "VERIFY WBC IN THE PLATELET PRODUCT". THE NEW ALGORITHMS WILL BE ADDED TO THE NEXT TRIM EQUIPMENT SOFTWARE UPGRADE, VERSION 6.0.3, WHICH IS IN THE PROCESS OF BEING VALIDATED. ADDITIONALLY, AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBCS. | 0
2465908 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE ELEVATED WBC CONTENT IN THIS PROCEDURE WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER. THERE ARE NO EVENTS (CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC) IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2466707 | THE CUSTOMER REPORTED THAT THE OPERATOR WAS PERFORMING A WHITE BLOOD CELL DEPLETION PROCEDURE. AT THE END OF RINSEBACK, WHEN REMOVING THE KIT, SHE STATED THAT THERE WAS A BLOOD LEAK WITH MIST GOING INTO THE AIR AND THE OPERATOR GOT BLOOD IN HER EYE. SHE USED THE EYEWASH. LABS WERE DRAWN FROM THE OPERATOR TO RULE OUT ANY CONTAMINATION. THE OPERATOR IS FINE. LABS ARE NEGATIVE, AND THE PT IS FINE. PT ID, AGE, AND GENDER (FOR THE OPERATOR) ARE NOT AVAILABLE AT THIS TIME. THE DISPOSABLE IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION WOULD LIKELY CAUSE OR CONTRIBUTE TO A DEATH OR INJURY IF THIS SAME FAILURE WERE TO RECUR.><THE CUSTOMER DECLINED TO GIVE THE OPERATOR'S INITIALS, AGE, OR WEIGHT. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. THERE WAS NO ALARM FROM THE FLUID DETECTOR AND NO BLOOD IN THE CENTRIFUGE BASIN. PREVIOUS INVESTIGATIONS HAVE SHOWN THAT INADVERTENTLY CLAMPING THE COLLECT LINE DURING THE STRIPPING PROCESS WILL BUILD UP PRESSURE ACROSS THE LUER CONNECTION AND WILL RESULT IN A BLOOD LEAK. ROOT CAUSE: BASED ON THE DESCRIPTION OF THE EXPOSURE BEING A FINE MIST, IT IS LIKELY THAT THE COLLECT WAS INADVERTENTLY CLAMPED DURING THE STRIPPING PROCESS.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2466792 | PT TAKEN TO SURGERY ON SATURDAY (B)(6) 2011. DURING INDUCTION OF ANESTHESIA, THE PT DESATURATED. THE PROBLEM WAS IDENTIFIED AS BEING WITHIN THE ADULT ANESTHESIA CIRCUIT. A NEW CIRCUIT WAS APPLIED AND THE PT'S VITAL SIGNS RETURNED TO NORMAL VALVES. UPON INSPECTION OF THE ANESTHESIA CIRCUIT, ONE OF THE INLINE FILTERS (BLUE, IVST PAST THE ELBOW) WAS PARTIALLY OCCLUDED. THE HARD PLASTIC LUMEN WAS PARTIALLY OCCLUDED WITH A MANUFACTURER'S MOLD IMPERFECTION. | 0
2466949 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVAL. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYS OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE REPORTED FAILURE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2467656 | PATIENT FOR A CABG WITH DA VINCI ROBOT. SURGEON ATTEMPTED TO USE THE DA VINCI BIPOLAR FORCEPS, BUT THEY DID NOT WORK. CIRCULATOR CHECKED THE ELECTROSURGICAL UNIT (ESU) AND NOTED THAT GROUND PAD (RETURN ELECTRODE FOR ESU) WAS NOT PLUGGED IN. GROUND PAD CONNECTED. SURGEON STATED THERE WAS NO NOISE AND THE FOOT PEDAL WAS NOT DEPRESSED, BUT AS SOON AS THE SURGEON TOUCHED THE DIAPHRAGM WITH THE BIPOLAR FORCEPS, THE MACHINE BUZZED AND THE DIAPHRAGM HAD A SMALL BURN. CIRCULATOR CHECKED THE ESU AND NOTED THAT THE FORCEPS CABLE WAS PLUGGED INTO TWO OF THE THREE MONOPOLAR 2 SOCKETS. BIPOLAR WAS PLUGGED INTO BIPOLAR SOCKETS, MACHINE FUNCTIONED NORMALLY THROUGHOUT THE REST OF THE CASE. PATIENT HAD TWO (2) GROUND PADS ON, UNSURE WHICH WAS THE ACTUAL LOT NUMBER INVOLVED IN EVENT.======================MANUFACTURER RESPONSE FOR SYSTEM,SURGICAL COMPUTER CONTROLLED INSTRUMENT, DA VINCI (PER SITE REPORTER).======================INTUITIVE REPRESENTATIVE CAME TO THE FACILITY AND TESTED THE MACHINE. NO ISSUES IDENTIFIED. | 0
2467863 | IT WAS REPORTED THAT THE MONITOR, EVEN WHEN EQUIPPED WITH AN ADDITIONAL BATTERY, WILL NOT ACHIEVE THE OPERATING TIME SPECIFIED IN THE INSTRUCTIONS FOR USE (IFU) AND THE DISPLAYED CAPACITY IS NOT FEASIBLE. THE MONITOR IS INTENSIVELY BEING USED AS A PICK AND GO DEVICE. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2469205 | IT WAS REPORTED THAT AN INFINITY CENTRAL STATION (ICS) ASSIGNS M300 BEDS TO DIFFERENT CLUSTER VIEW LOCATIONS. THERE WAS NO PT INJURY REPORTED. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL UNDER INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2469311 | THE PT WAS SEATED IN A CHAIR AND WAS BEING POSITIONED FOR A NEEDLE LOCALIZATION PROCEDURE. THE SELENIA C-ARM WAS LOWERED BY THE TECHNOLOGIST BY DEPRESSING THE FOOTSWITCH. WHEN THE FOOTSWITCH WAS RELEASED THE C-ARM CONTINUED THE DOWNWARD MOTION. THE TECHNOLOGIST DEPRESSED THE UP-COMMAND FOOTSWITCH, BUT THE C-ARM CONTINUED DOWNWARD ONTO THE PT'S LEFT LEG. THE SYSTEM WAS EVENTUALLY STOPPED VIA PHYSICAL RESISTANCE BY TWO OTHER INDIVIDUALS CALLED TO ASSIST THE TECHNOLOGIST DURING THIS EVENT. THE EMERGENCY OFF SWITCH WAS NOT USED. THE TECHNOLOGIST STATED PT'S LEG HAD A RED MARK WHERE IT CAME IN CONTACT WITH THE SELENIA C-ARM. THE PT WAS BROUGHT TO ANOTHER ROOM AND THE PROCEDURE WAS COMPLETED THAT SAME DAY. THE SELENIA SYSTEM WAS EVALUATED BY AN HOLOGIC FIELD ENGINEER ON (B)(4) 2012 AND WAS FOUND TO BE WORKING PROPERLY. THE TECHNOLOGIST STATED THE PT WAS VERY NERVOUS AND MAY HAVE INADVERTENTLY DEPRESSED THE CONTROL SWITCH AT THE BACK OF THE SYSTEM C-ARM ALLOWING FOR THE CONTINUATION OF THE DOWNWARD MOVEMENT. THE OPERATOR'S MANUAL ADVISES CAUTION IN THE PLACEMENT OF THE PT'S HANDS DURING A PROCEDURE TO AVOID THIS TYPE OF INCIDENT. IN ADDITION, THE USE OF THE EMERGENCY OFF SWITCH WOULD HAVE IMMEDIATELY SHUTDOWN THE SYSTEM, PREVENTING THE CONTINUED DOWNWARD MOTION OF THE C-ARM. THE TECHNOLOGIST DID NOT USE THE EMERGENCY OFF SWITCH DURING THIS EVENT. THE SYSTEM WAS IMMEDIATELY RETURNED TO SERVICE AFTER THE EVAL BY THE HOLOGIC FIELD ENGINEER. NO OTHER OCCURRENCES HAVE BEEN REPORTED WITH THIS SYSTEM. | 0
2469826 | THE CUSTOMER REPORTED THAT THE ROTATOR BROKE DURING USE. THE CUSTOMER ALSO REPORTED THREE ADDITIONAL EVENTS WITH SIMILAR TUBING SETS. REFERENCE REPORT NUMBERS BELOW. NO HARM OR INJURY WAS REPORTED. THIS IS ONE OF FOUR REPORTS FOR THIS COMPLAINT. 1721504-2011-00327; 1721504-2011-00328; 1721504-2011-00329; 1721504-2011-00330 ¿ THIS REPORT. DEVICE EVALUATION: ONE USED SUSPECT DEVICE WAS RETURNED FOR EVALUATION. THE CUSTOMER DID NOT SPECIFY IF THIS WAS THE INITIAL USE OF THE DEVICE. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINT FOR THIS LOT NUMBER. A VISUAL INSPECTION OF THE RETURNED DAMAGED. THE LEADING EDGE HAD BROKEN OFF THE FRONT OF THE FITTING. THE OTHER LEADING EDGE WAS PULLED-AWAY FROM THE FITTING. DAMAGE TO THE THREAD WALL IN THE FORM OF SMALL GROOVES WAS ALSO FOUND. THE LUER TAPER ALSO SHOWED HELICAL SCRATCHES. THE DAMAGE FOUND ON THE LUER INDICATES THAT A DEVICE HAD BEEN FORCED AGAINST THE LUER WITH SUFFICIENT FORCE TO CAUSE THE DAMAGE. MERIT IS UNABLE TO DETERMINE THE EXACT ROOT CAUSE OF THE REPORTED FAILURE. | 0
2471339 | THE PATIENT SAFETY CONCERN REGARDING THE DAVINCI ROBOT IS A POSSIBLE PRODUCT DESIGN ISSUE. THE PATIENT SUFFERED A COMPARTMENT SYNDROME OF THE RIGHT LOWER EXTREMITY (RLE) AFTER THE ROBOTIC-ASSISTED LAPAROSCOPY (WITH PROCEDURES). IT IS FELT THAT THE ROBOTIC ARM, WHILE DOCKED ON THE RIGHT SIDE OF THE PATIENT, EITHER BUMPED AND/OR RESTED ON THE PATIENT'S RLE. THIS WAS A 7 HOUR PROCEDURE. THE PATIENT REQUIRED EMERGENCY, 4-COMPARTMENT FASCIOTOMY FOLLOWED BY ANOTHER SURGERY FOR SKIN GRAFT (PLUS WOUND VAC) AND WOUND CLOSURE. | 0
2471509 | ON (B)(6) 2012, THE MANUFACTURER'S CONSULTANT REPORTED THAT HIS HANDHELD NEEDS TO BE REPLACED BECAUSE IT HAS A RANDOM BLACK COMPUTER GRAPHIC IN THE BACKGROUND AND THE IMAGE ON THE SCREEN LOOKS DISTORTED. HE ALSO DESCRIBED THAT THERE IS WHAT LOOKS LIKE LINE IN THE SCREEN. THE CONSULTANT FURTHER STATED HOWEVER THAT THE FUNCTIONALITY OF THE HANDHELD WAS NOT AFFECTED. A HARD RESET WAS PERFORMED, BUT THE ISSUE WAS NOT RESOLVED. ATTEMPTS HAVE BEEN MADE FOR THE RETURN OF THE HANDHELD TO THE MANUFACTURER FOR PRODUCT ANALYSIS HOWEVER IT HAS NOT YET BEEN RECEIVED.><ON (B)(6) 2012, ADDITIONAL INFORMATION WAS RECEIVED WHEN THE HANDHELD AND FLASHCARD WERE RETURNED FOR PRODUCT ANALYSIS. PRODUCT ANALYSIS WAS COMPLETED ON (B)(6) 2012. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6), 2012 WHEN IT WAS REPORTED THAT THE HANDHELD STILL FUNCTIONED PROPERLY. ATTEMPTS WERE MADE FOR THE RETURN OF THE HANDHELD FOR PRODUCT ANALYSIS BUT IT HAS NOT BEEN RECEIVED TO DATE. . | 0
2473520 | THE CUSTOMER STATES A SAMPLES TESTED NEGATIVE ON THE ORTHO PROVUE AND POSITIVE IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED AS A RESULT OF THIS INCIDENT. THERE WAS NO HARM TO ANY PATIENT. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND CHECKED THE CENTRIFUGE SPEED AND TEMPERATURE. BOTH SPEED AND TEMPERATURE WERE WITHIN SPECIFICATION. THE FE PERFORMED A PM AND RAN DIAGNOSTIC SOFTWARE WITHOUT ERRORS. THE CUSTOMER RAN CONTROLS AND ACCEPTED RESULTS. AS PER CTS-US 2ND LEVEL SUPPORT - THERE WERE NO ERRORS DURING PROCESSING AND THE READER CAMERA IS READING APPROPRIATELY. THE REVIEW SHOWS THE PROVUE IS OPERATING WITHIN SPECIFICATIONS. (B)(4). | 0
2473678 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT#: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2474313 | GE TELEMETRY SYSTEM IN PLACE WITH ASCON PHONE SYSTEM. FOR MONITORING CARDIAC ARRYTHMIAS. THE "LEADS FAIL" DID NOT SEND AN ALARM TO THE PHONES. ONE PATIENT ROOM NOT COMMUNICATING ANY ALARMS. UPON INVESTIGATION, THE SETTINGS WERE NOT COMMUNICATING WITH EACH OTHER. THIS WAS CORRECTED BY THE COMPANY. PATIENT HAD ARRHYTHMIA AND EXPIRED. | 1
2474700 | MAMMO TECH COMPRESSED BREAST OF PATIENT, AND WHEN SHE RELEASED THE COMPRESSION FOOTPEDAL, IT CONTINUED TO COMPRESS. EMERGENCY OFF SWITCHES WERE NOT UTILIZED. PATIENT PULLED HER BREAST OUT DURING COMPRESSION.GE FOUND NO PROBLEMS WHEN DOING TEST AND CALIBRATIONS. HOWEVER, FOUND THE THE CABLE TO THE FOOTSWITCH WAS TWISTED AND CAUGHT ON SWITCH, NOT ALLOWING IT TO DEACTIVATE. FOR PRECAUTIONARY PURPOSES, REP REPLACED ALL RELATED COMPONENTS AND RECALIBRATED. CABLE WAS SHORTENED AND MADE IMPOSSIBLE TO GET HUNG UNDER SWITCH. | 0
2475287 | TELEMETRY MONITOR NOT MEASURING CORRECTLY. THIS HAPPENS AFTER INCORRECT RECORDING OF EVENTS THAT ARE NOT RELATED TO THIS PATIENT. | 1
2475289 | THE PATIENT'S EKG RHYTHM APPEARS TO SPEED UP, BUT THE ALARM DOES NOT ALARM. IN FULL DISCLOSURE, THE RHYTHM IS IDENTIFIED, BUT PRECEDING THE EVENT IT IS DESIGNATED BY A BLUE LINE WITH THE DATE AND TIME. | 1
2475909 | NURSING STAFF NOTICE VITAL SIGNS MONITOR WAS NOT PLUGGED INTO A WALL OUTLET FOR CHARGING. THE NURSE PLUGGED THE VITAL SIGNS MONITOR INTO THE ELECTRICAL OUTLET. A BRIEF SPARK AND SMOKE EMITTED FROM THE CONNECTION. THE NURSE TOOK ANOTHER VITAL SIGNS MONITOR AND TRIED TO PLUG IT INTO THE SAME OUTLET AND IT TOO SPARKED. | 1
2475938 | MARQUETTE, SOLAR 8000M (GE) MONITOR STOPPED MONITORING PATIENT; SCREEN WENT BLACK; ERROR MESSAGE APPEARED ON SCREEN; BIOMED PAGED STAT; PATIENT PLACED ON TRAVEL MONITOR. IN THE DURATION OF DOWNTIME WHICH STARTED APPROXIMATELY ONE HOUR AGO AND ONGOING NOW, PATIENT WAS NOT ABLE TO BE MONITORED BY CENTRAL MONITOR AND PATIENT HAD TO BE MONITORED CONTINUOUSLY WITH DIRECT OBSERVATION. PATIENT ON VASOACTIVE INFUSION, VENTILATED, AND ARTERIAL PRESSURES MONITORED. PATIENT BEING PACED DUAL CHAMBER, RATE, RESPONSE (DDD) ON EXTERNAL PACER FOR UNDERLYING 3RD DEGREE BLOCK. BIOMED STATED THEY NEEDED TO REPLACE THE WHOLE SYSTEM.MANUFACTURER RESPONSE FOR PATIENT MONITOR, PHYSIOLOGICAL, SOLAR 8000M (PER SITE REPORTER).======================MANUFACTURER SENT QUESTIONNAIRE:WHEN COMPLETING THE QUESTIONS IN THE PROBLEM SOLUTION FIELD, DO NOT LEAVE ANY BLANK. ENTER "N/A" FOR ANY QUESTION THAT IS NOT APPLICABLE.WHAT WAS THE TYPE OF PROCEDURE? MONITORING.WHAT INTERVENTIONS WERE REQUIRED ON THE PART OF THE CLINICIAN? THEY HAD TO SWAP OUT THE SOLAR.WHAT WAS THE CONSEQUENCE TO THE PATIENT? NONE.CUSTOMER REPRESENTATIVE AND ROLE. BIOMED.IF PATIENT USE = YES & SAFETY CONCERN = POTENTIAL OR ACTUAL, THE FOLLOWING QUESTIONS NEED TO BE ASKED AND ANSWERED: WAS THIS EVENT NOTICED IMMEDIATELY? YES.DID ANY PATIENTS EXPERIENCE A MISTREATMENT OR DELAY OF TREATMENT RELATED TO THIS EVENT? NO.IF APPLICABLE, HOW LONG WAS PATIENT MONITORING LOST AND FOR HOW MANY PATIENTS? UNKNOWN.WAS ALTERNATE MONITORING AVAILABLE? NO.APPROXIMATE DATE/TIME OF FAILURE: DAY OF EVENT AROUND 10 AM.RECEIVED FEEDBACK FROM GE WITH CONFIRMATION THAT SERVICE REQUEST FOR TECHNICAL ASSISTANCE HAS BEEN CLOSED WITH NO CLEAR RESOLUTION. | 1
2476147 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT FLUID FROM THE BACKWASH LEAKED FROM A BROKEN RINSE CUP ON THEIR COULTER LH 750 HEMATOLOGY ANALYZER. APPROXIMATELY 3CC OF THE FLUID FLOWED INTO THE WHITE UNDER TRAY, BUT DID NOT OVERFLOW OUTSIDE THE INSTRUMENT. THERE IS AN EXPOSURE CONTROL/RISK MANAGEMENT PLAN IN PLACE AT THE FACILITY. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A GOWN AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED. NO PATIENT RESULTS WERE AFFECTED BY THE LEAK. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND REPLACED PROBE WIPE ASSEMBLY WHICH RESOLVED THE ISSUE. THE ROOT CAUSE FOR THE LEAK IS A BROKEN RINSE CUP. (B)(4). | 0
2476150 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT WHILE TROUBLESHOOTING AN INSTRUMENT PROBLEM, THEY FOUND A LEAK OF CLEAR FLUID (APPROXIMATELY 5 ML) FROM UNDERNEATH THE UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM AND ALSO THE NRBC QC WAS FAILING. THE CUSTOMER WAS UNABLE TO DETERMINE THE SOURCE OF THE LEAK. THERE IS AN EXPOSURE CONTROL/RISK MANAGEMENT PLAN IN PLACE AT FACILITY. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT RESULTS WITH REGARD TO THIS EVENT. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE (B)(4) 2012 AND FOUND A PINHOLE LEAK IN THE DIFFERENTIAL MIX SAMPLE LINE AT FEED THROUGH FITTING (FF)240 AND ALSO THAT THE NRBC CHAMBER DRAIN WAS PLUGGED. THE FSE REPLACED THE NRBC CHAMBER AND THE DIFFERENTIAL MIX CHAMBER SAMPLE LINE AT FF240 AND REPAIRED THE LEAK. REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES AND RESULTS MET PUBLISHED PERFORMANCE. THE ROOT CAUSE OF THE ISSUE WAS A PINHOLE LEAK IN THE DIFFERENTIAL MIX CHAMBER OF THE DXH800 INSTRUMENT. (B)(4). | 0
2476616 | REPORT RECEIVED OF A USER ERROR RESULTING IN A SUCTION SWAB DISENGAGEMENT; LABELING INSTRUCTIONS REGARDING THE USE OF BITE BLOCK WERE NOT FOLLOWED. REPORTER STATED THAT ON (B)(6) 2012, A CERTIFIED NURSES ASSISTANT (CNA) ATTEMPTED TO PROVIDE ORAL CARE TO AN UNCOOPERATIVE RESIDENT. REPORTEDLY, RESIDENT BIT DOWN ON THE SUCTION SWAB AND THE PLASTIC STRAW SEPARATED INTO TWO PIECES. RESIDENT HAD CUFFED TRACHEOSTOMY AND WAS ON VENTILATOR. REPORTER STATED CNA, PHYSICIAN ASSISTANT, NURSES, AND (B)(6) ATTEMPTED TO RETRIEVE THE SWAB HEAD WITH ATTACHED PIECE OF PLASTIC FROM BACK OF RESIDENT'S MOUTH HOWEVER, ALL WERE UNSUCCESSFUL. PARAMEDICS NOTIFIED AND THEY TOO ATTEMPTED TO REMOVE IT WITHOUT SUCCESS. RESIDENT WAS TRANSFERRED TO (B)(6) HOSPITAL'S EMERGENCY ROOM (ER). REPORTEDLY, RESIDENT WAS "AGITATED, TACHYPNEIC AND DESATURATED." WHILE IN ER, RESIDENT WAS NOTED TO HAVE "STRIDOR" AND WAS GIVEN NIMBEX AND PROPOFOL TO TEMPORARILY PARALYZE HIM. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED WHICH SHOWED A FOREIGN BODY, "LIKELY SPONGE", NOTED AT THE LEVEL OF THE EPIGLOTTIS IN THE HYPOPHARYNX. REPORTEDLY, RESIDENT WAS RELEASED FROM ER AND ADMITTED TO A NURSING UNIT IN THE HOSPITAL. RIGHT AFTER ADMISSION TO UNIT, RESIDENT REPORTEDLY "COUGHED OUT THE PIECE OF SPONGE." FOLLOW-UP LARYNGOSCOPY ON (B)(6) 2012 SHOWED NO EVIDENCE OF FOREIGN BODY. INSTRUCTIONS FOR USE STATE, "A BITE BLOCK SHOULD BE USED WHEN PERFORMING ORAL CARE ON PTS WITH ALTERED LEVELS OF CONSCIOUSNESS OR THOSE WHO CANNOT COMPREHEND COMMANDS." IT WAS REPORTED THAT ALTHOUGH THE RESIDENT WAS UNCOOPERATIVE, A BITE BLOCK WAS NOT USED AT THE TIME OF THE INCIDENT. REPORTEDLY, THE SUCTION SWAB WAS DISCARDED. LOT INFO WAS NOT AVAILABLE FROM REPORTER. THOUGH REQUESTED, NO ADD'L INFO WAS PROVIDED. NA. | 0
2477036 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) AND REPORTED A LEAK OF APPROXIMATELY 4-5 MLS OF CLEAR FLUID LEAKING FROM THE RIGHT SIDE OF THEIR COULTER HMX AUTOLOADER INSTRUMENT. THE CUSTOMER WAS UNABLE TO PINPOINT THE SOURCE, BUT THE LIQUID WAS CLEAR. THE LAB'S EXPOSURE CONTROL/RISK MANAGEMENT PLANS ARE IN PLACE. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, EYE PROTECTION, AND GLOVES. NO INJURY OR EXPOSURE WAS REPORTED AND NO PATIENT RESULTS WERE AFFECTED; THE UNIT WAS IN A SHUTDOWN AND STARTING A STARTUP. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THE LEAK AT PINCH VALVE 43(PV43). FSE REPLACED TUBING AT PV43 AND VERIFIED REPAIR PER ESTABLISHED PROCEDURES. THERE WAS NO FURTHER EVIDENCE OF LEAKING. THE ROOT CAUSE OF THE LEAK IS ATTRIBUTED TO PV43. (B)(4). | 0
2477784 | DURING THE COURSE OF ROUTINE GENERAL ANESTHESIA, THE NONINVASIVE BLOOD PRESSURE MONITOR (NIBP OR NBP) WHICH WAS SET TO AUTO MODE WITH 2.5 MINUTE INTERVAL STOPPED CYCLING WITHOUT WARNING, NOTIFICATION, OR APPARENT CAUSE. THE MONITOR WAS GE SOLAR 8000I WITH A PDM (MULTIPARAMETER MODULE) WITH VERSION 1.14 SOFTWARE. THE SOFTWARE VERSION WAS CONFIRMED WITH PHOTO OF SAME BECAUSE SOFTWARE VERSION 2.0 WAS KNOWN TO HAVE THIS PROBLEM AND WAS RECALLED AND WAS REMOVED FROM USE IN THIS HOSPITAL. TWO PHYSICIANS OBSERVED THAT THERE WAS NO BLOOD PRESSURE MEASURED FOR NINE (9) MINUTES AND USED THE CONTROL KNOBS AND TOUCH SCREEN TO TURN THE NIBP BACK TO AUTO MODE AND RESET THE INTERVAL TO 2.5 MINUTES. THE LOCATION WAS IN AN OR. WE DOWN GRADED OUR PDM SOFTWARE TO VERSION 1.4 AND ARE STILL SEEING THE SAME PROBLEM WITH THE PDM'S IN THE OR.======================MANUFACTURER RESPONSE FOR NONINVASIVE BLOOD PRESSURE MONITOR, PATIENT DATA MODULE (PDM) SOFTWARE VERSION 1.4 AND AND SOLAR 8000I (PER SITE REPORTER).======================WE HAVE NOT LET THE MANUFACTURER AWARE AS OF YET, HOWEVER, WE WILL AFTER THIS IS SUBMITTED. | 1
2478409 | IT WAS REPORTED THAT A CIC CENTRAL STATION SUDDENLY LOST POWER AND SHUT DOWN, RESULTING IN A LOSS OF TELEMETRY MONITORING FOR 7 PTS FOR APPROX 30 MINUTES. NO INJURY OR DELAY IN TREATMENT WAS REPORTED AS A RESULT OF THE LOSS OF MONITORING. A GE HEALTHCARE ON-LINE SUPPORT ENGINEER WAS CONTACTED BY THE SITE. IT WAS DETERMINED THAT THE CIC AND ITS DISPLAY WERE CONNECTED TO A UPS WHOSE POWER SWITCH WAS FOUND TO BE OFF. THE SITE BELIEVES THAT THE POWER SWITCH WAS ACCIDENTALLY BUMPED. THE USER TURNED ON THE UPS POWER SWITCH, HOWEVER, THE CIC DID NOT BOOT UP COMPLETELY (THE CIC POWER WAS ON BUT THERE WAS NO VIDEO SIGNAL AS THE CIC DISPLAY LED WAS YELLOW, INDICATING STANDBY). THE ON-SITE BIOMEDICAL TECHNICIAN WAS SUBSEQUENTLY CONTACTED AND REPORTEDLY RESOLVED THE ISSUE. THE STEPS TAKEN TO RESOLVE THE ISSUE WERE NOT PROVIDED. BASED ON THE SCENARIO AND SYMPTOMS REPORTED (SUDDEN POWER INTERRUPTION, RESULTING IN IMPROPER SHUT DOWN), IT IS BELIEVED THAT THE PROBLEM WAS RELATED TO THE VIDEO CARD NOT INITIALIZING, RESULTING IN NO VIDEO ON BOOT UP. THIS ISSUE IS GENERALLY RESOLVED BY TURNING OFF THE POWER OF THE CIC AND DISCONNECTING THE POWER CORD FOR SEVERAL MINUTES TO REMOVE ANY STANDBY VOLTAGES, THEN PLUGGING THE POWER CORD BACK IN AND TURNING ON THE SWITCH. | 1
2478511 | DURING MONITORING OF AN ANESTHETIZED PATIENT, AN ANESTHESIA MACHINE CO2 MONITOR ALARMED FOR OCCLUDED SAMPLE TUBING. THE SAMPLE TUBING WAS REPLACED. THE MONITOR CONTINUED TO ALARM AND WAS REPLACED WITH A BACKUP CO2 MONITOR. THE BACKUP MONITOR ALSO ALARMED FOR OCCLUDED SAMPLE TUBING. IT WAS EVENTUALLY DETERMINED THAT THE NEW SAMPLE TUBING WAS DEFECTIVE, AS WERE OTHER SAMPLE TUBING LINES COMING DIRECTLY OUT OF THE PACKAGING. | 0
2479824 | A CUSTOMER REPORTED PORTAL IMAGES TAKEN WITH IVIEWGT ARE NOT ALWAYS SAVED IN (B)(6) DATABASE WHILE EXPORTING IMAGES FROM (B)(6) TO VARIAN TREATMENT. THE CUSTOMER WAS ASKED TO RECREATE INDEX FILES FOR THE IVIEWGT DATABASE. AFTER RECREATING INDEX FILES, THIS ISSUE WAS REPRODUCIBLE. THE CUSTOMER CREATED A NEW (B)(6) DATABASE AND EVERYTHING STARTED WORKING FINE WITHOUT ANY ISSUE. THE INVESTIGATION INTO THIS SITUATION IS ON-GOING AT THIS TIME. ONCE THE INVESTIGATION IS COMPLETED, MORE DETAILS AND A CONCLUSION WILL BE PROVIDED. | 1
2480960 | A VENTILATOR WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR PREVENTIVE MAINTENANCE. THERE WAS NO ALLEGATION THAT THE DEVICE FAILED TO OPERATE. THE DEVICE WAS NOT IN PATIENT USE. DURING THE DEVICE EVALUATION AT THE MANUFACTURER'S SERVICE CENTER, A FAILURE RELATED TO THE ACTIVE EXHALATION CONTROL MODULE WAS OBSERVED. THE DEVICE'S ACTIVE EXHALATION CONTROL MODULE WAS REPLACED TO ADDRESS THE ISSUE. THERE WAS NO PATIENT HARM OR INJURY REPORTED. THE VENTILATOR WAS FOUND TO AUDIBLY AND VISUALLY ALARM, AS DESIGNED. A GE HEALTHCARE SERVICE TECHNICIAN PERFORMED AN ON-SITE EVALUATION AND DETERMINED THAT THE CAUSE OF THE REPORTED ISSUE WAS A FAILED SPEAKER ASSEMBLY. THE SPEAKER ASSEMBLY WAS REPLACED AND CORRECTED THE REPORTED ISSUE. THE DEVICE WAS TESTED FOLLOWING REPAIRS AND WAS FOUND TO OPERATE WITHIN SPECIFICATIONS. | 1
2481495 | A CUSTOMER CONTACTED HAEMONETICS ON (B)(6) 2010, TO REPORT THAT A MCS+ MACHINE RETURNED ALL THE SALINE DURING THE FIRST CYCLE. NO DONOR OR OPERATOR INJURY WAS REPORTED. THE DEVICE HAS NOT BEEN RETURNED THEREFORE NO FORMAL EVAL HAS BEEN PERFORMED. THE CUSTOMER STATED THAT THE "RINSING DRAW LINE" PART OF THE PROCEDURE SEEMED TO BE TAKING TOO LONG AND THEN NOTICED THAT THE SALINE BAG HAD EMPTIED. THE PROCEDURE WAS THEN STOPPED, THE PROGRAM WAS SAVED AND THE PROCEDURE WAS ENDED WITHOUT INCIDENT. ALL PROCEDURE COMPLETE NUMBERS WERE ACCURATE, THE CUSTOMER STATED THAT THERE WERE NO ALARMS DURING THE PRIME OR PROCEDURE. THE MACHINE IS NOW IN USE. HAEMONETICS HAS NOT RECEIVED ANY ADDITIONAL REPORTS OF THIS PROBLEM FROM THE CUSTOMER. THE PRODUCT ISSUE IDENTIFIED IN THIS REPORT WAS IDENTIFIED BY HAEMONETICS DURING A RETROSPECTIVE REVIEW OF THE COMPANY'S SERVICE RECORDS. NO PT OR USER HARM OR INJURY WAS REPORTED OR ALLEGEDLY ASSOCIATED WITH THE IDENTIFIED ISSUE. THE DEVICE AIR DETECTOR DID NOT ALARM TO ALERT THE OPERATOR THAT THE SALINE BAG HAD EMPTIED. THE OPERATOR USED GOOD CLINICAL JUDGMENT TO STOP THE PROCEDURE. (B)(4). | 0
2482531 | A REVIEW OF THE PATIENT'S PROGRAMMING HISTORY IN THE IN-HOUSE DATABASE WAS PERFORMED, AND IT WAS OBSERVED THAT UPON INITIAL INTERROGATION ON (B)(6) 2005, THE SETTINGS WERE AT UNINTENDED PARAMETERS THAT ARE INDICATIVE OF A FAULTED SYSTEMS DIAGNOSTIC TEST OCCURRING. THE SETTINGS WERE REPROGRAMMED ON (B)(6) 2005. IT IS NOTED THAT THE MAGNET ON TIME WAS NEVER CORRECTED BACK TO 60 SEC. IT IS UNKNOWN IF THIS WAS INTENTIONAL. ANALYSIS OF PROGRAMMING HISTORY. | 0
2488464 | IT WAS REPORTED ON (B)(6) 2012 THAT A PHYSICIAN'S HANDHELD DEVICE WAS FREEZING ON THE MAIN MENU SCREEN DESPITE MULTIPLE HARD RESETS. A REPLACEMENT HANDHELD WAS SENT TO THE PHYSICIAN OVERNIGHT AS HE HAD PATIENTS COMING IN THE NEXT DAY. ATTEMPTS AT PRODUCT RETURN OF THE HANDHELD IN QUESTION HAVE NOT BEEN SUCCESSFUL TO DATE.><THE PHYSICIAN'S DELL X5 HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2012 AND PRODUCT ANALYSIS HAS BEEN COMPLETED. THERE WERE NO ANOMALIES FOUND WITH EITHER THE HANDHELD OR FLASHCARD. BOTH THE HANDHELD AND FLASHCARD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE REPORTED "FROZEN SCREEN" ON THE MAIN MENU SCREEN WAS NOT VERIFIED. . | 0
2489076 | ANESTHESIA MACHINE IN OR 1 WOULD NOT WORK. HAD TO MOVE ESOPHAGOGASTRODUODENOSCOPY (EGD) PATIENT TO ANOTHER OR. THIS MACHINE HAD A HISTORY OF THIS SAME MALFUNCTION. | 0
2489090 | PATIENT UNDERGOING RIGHT HIP HEMIARTHROPLASTY WHICH REQUIRES PATIENT TO BE IN LATERAL POSITION ON OR TABLE. WHEN PATIENT WAS TURNED ANESTHESIA NOTE: THE PATIENT DEVELOPED SEVERE BRONCHOSPASM IN THE OR UPON CHANGE OF THE POSITION FROM LATERAL TO THE SUPINE; QUICK DESATURATION OF OXYHEMOGLOBIN, INCREASED AIRWAY PRESSURES, LOW TIDAL VOLUMES, DIFFICULT VENTILATION, SEVERELY DIMINISHED BREATH SOUNDS BILATERALY. ETT SUCTIONED OUT, NO SECRETIONS OR ASPIRATE, ALBUTEROL INHALER ADMINISTERED VIA ETT WITH GOOD RESULTS. ETT (VITAL SIGNS CO BRAND) NOTED TO BE SOFT TO PALPATION, I COULD NOT EXCLUDE KINK IN THE OROPHARYNX; PILLOW WAS REMOVED FROM UNDER THE HEAD, SUCTION CATHETER PASSED AGAIN VIA ETT AND RESISTANCE SUGGESTING KINK WAS NOTED. ABG: PH 7.14, PCO2 69, PO2 100, BICARB 23.5, BE -5.8, SAO2 94.6. OROPHARYNX SUCTIONED OUT, GLIDESCOPE BACKUP; ETT EXCHANGED TO DIFFERENT BRAND (STIFFER AND NOT SOFTENING IN THE BODY TEMPERATURE) PERFORMED OVER ESCHMAN INTRODUCER.======================MANUFACTURER RESPONSE FOR REDITUBE ENDOTRACHEAL TUBE WITH STYLET (PER SITE REPORTER).======================TO FILE MEDWATCH AND RETURN THE AFFECTED PRODUCT. | 0
2489663 | IT WAS REPORTED THAT AFTER SWITCHING ON THE MONITOR FROM STANDBY, THE MONITOR SHOWS A CURVE FOR OXYGEN MEASUREMENT, BUT NO VALUES. AFTER SWITCHING OFF/ON, THE VALUES RETURN. THIS CAN ALSO BE REACHED BY UNPLUGGING AND PLUGGING IN AGAIN OF THE MASIMO POD USB CONNECTION. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2490131 | ON (B)(6), 2010, THE CUSTOMER REPORTED THAT THE OPERATING ROOM TABLE WAS MOVING UP AND DOWN AND TILTING ON ITS OWN, WITHOUT THE MANIPULATION OF THE OPERATING ROOM TABLE PENDANT (CONTROLLER) BY THE OPERATOR. THERE WAS NO REPORT OF INJURY TO EITHER A PT OR CLINICAL STAFF. AN IMRIS CUSTOMER SUPPORT SPECIALIST INVESTIGATED THE COMPLAINT ON-SITE, BUT WAS UNABLE TO REPRODUCE THE INCIDENT. THE OPERATING ROOM TABLE PENDANT WAS LATER RETURNED TO IMRIS FOR INVESTIGATION, AND THE INCIDENT AGAIN COULD NOT BE DUPLICATED. AS THIS WAS THE FIRST OCCURRENCE, WITH NO REPORT OF INJURY AND THE INCIDENT COULD NOT BE REPRODUCED, IMRIS DETERMINED THAT THE RESULTS OF THE INVESTIGATION DID NOT WARRANT AN MDR. HOWEVER, UPON DISCOVERY OF A SECOND INCIDENT AT ANOTHER CUSTOMER SITE, (REF MDR 3003870210-2010-00006) IMRIS DETERMINED THAT THIS CASE SHOULD BE RE-OPENED PENDING THE CONCLUSION OF THE INVESTIGATION INTO THE ABOVE-MENTIONED CASE, AND THAT AN MDR BE SUBMITTED. | 0
2490207 | DURING A REVIEW OF A PATIENT'S PROGRAMMING HISTORY, IT WAS NOTICED THAT A FAULTED SYSTEMS DIAGNOSTIC TEST OCCURRED ON (B)(6) 2007 WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE NOT CORRECTED UNTIL THE SUBSEQUENT VISIT ON (B)(6) 2007. NO OTHER ANOMALIES WERE OBSERVED DURING THE REVIEW OF THE PROGRAMMING HISTORY. ANALYSIS OF PROGRAMMING HISTORY. | 0
2490580 | PATIENT'S RHYTHM WAS OBSERVED TO HAVE AN ARRHYTHMIA, BUT MONITOR DID NOT ALARM. THE RHYTHM CHANGED WAS DENOTED BY A VERTICAL BLUE LINE THAT STATED THE DATE AND TIME. THE RHYTHM CHANGED TERMINATED BACK TO PATIENT'S PRIOR RHYTHM. FULL DISCLOSURE REVEALED THE EVENT. BIOMEDICAL ENGINEERING WAS NOTIFIED. THE CHARGE NURSE IS AWARE OF THIS ISSUE TO OBSERVE AND FOLLOW UP. | 1
2490581 | PATIENT'S RHYTHM WAS OBSERVED TO CHANGE, THE CENTRAL MONITOR DID NOT ALARM FOR VENTRICULAR STANDSTILL, THE RHYTHM AUTO TERMINATED AND WENT BACK TO PATIENTS RHYTHM. THE RHYTHM CHANGED WAS DENOTED BY A BLUE VERTICAL LINE, WHICH HAD THE DATE AND TIME. | 1
2490582 | PATIENT'S RHYTHM ACCELERATED, BUT NO MONITOR ALARMS ACKNOWLEDGED THE CHANGE IN RHYTHM. THE RHYTHM IS PRECEDED BY A VERTICAL PAUSE THEN A CHANGE BACK TO THE PATIENT'S KNOWN RHYTHM. THE CHANGE IN THE RHYTHM DOES NOT TRIP THE CENTRAL ALARM. THE CHANGE IN THE RHYTHM DOES NOT TRIP THE SECONDARY ALARM (STAT VIEW PAGER). IT WOULD BE NOTICED ONLY IF DIRECT OBSERVATION OF THE EVENT OCCURRED. IT IS PRECEDED BY A BLUE VERTICAL PAUSE ON THE MONITOR, WHICH STATES THE DATE AND TIME. | 1
2490649 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE TRIPLE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE ANALYSIS OF THE RUN DATA FILE (RDF) DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CIRCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2492348 | THE O-RING PIECE OF THE VENTILATOR BECAME DISCONNECTED CAUSING ANESTHETIC GAS TO LEAK. | 0
2492884 | PER THE FAMILY THE ALARMS ARE NOT AUDIBLY SOUNDING; THE DOWNLOAD VIA THE APNEA PROGRAM SHOWS THAT THERE HAVE BEEN ALARMS. THE MONITOR ALARMS WHEN IT STARTS UP AND DOES THE SELF CHECK. | 1
2493845 | PT COMPLAINS HE HAS HAD MORE THAN 10 SURGERIES ON HIS RIGHT SHOULDER. THE FIRST AND SECOND SURGERIES HE USED ON-Q WHICH CAUSED HIS CHONDROLYSIS. THIS CAUSED HIM TO HAVE A RIGHT SHOULDER REPLACEMENT SURGERY. THIS HAS BROUGHT HIM SO MUCH PAIN, STRESS AND SUFFERING | 0
2493923 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS NOTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED ON (B)(6) 2008 WHICH CHANGED THE SETTINGS TO 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF. THE SETTINGS WERE NOT CORRECTED UNTIL (B)(6) 2009, WHEN THEY WERE CHANGED TO 1.5MA/20HZ/250PULSEWIDTH/60SEC ON/1.1MIN OFF. . | 0
2494369 | DURING MANUFACTURER REVIEW OF A VNS PATIENT'S PROGRAMMING HISTORY, IT WAS NOTED THAT ON (B)(6) 2008 THE INITIAL INTERROGATION WAS 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF. NO CHANGES WERE MADE TO THE SETTINGS. AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2007, THE SETTINGS HAD BEEN 1.5MA/30HZ/500PULSEWIDTH/30SEC ON/90MIN OFF. IT IS SUSPECTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED AT SOME POINT BETWEEN (B)(6) 2007 AND (B)(6) 2008 WITH AN UNKNOWN VNS PROGRAMMING SYSTEM. CLINIC NOTES WERE ALSO RECEIVED TO THE MANUFACTURER WHICH NOTED THE PATIENT'S OFF TIME WAS STILL SET TO 60 MINUTES. MANUFACTURER FOLLOW UP WITH THE PHYSICIAN REVEALED HE IS AWARE THAT A PROGRAMMING ANOMALY MAY HAVE OCCURRED AND HE IS AWARE THAT 60 MINUTES OFF IS NOT A THERAPEUTIC SETTING. . | 0
2494628 | A CUSTOMER FROM (B)(6) REPORTED THAT RESULTS FOR TESTS THAT HAD BEEN PREVIOUSLY PROCESSED WERE INCLUDED ON ORDERS SENT FROM THEIR LIS TO THE AMPLILINK (AL) SOFTWARE V.3.2.3. (B)(4). THE INVESTIGATION DETERMINED THAT INCORRECT TEST IDS WERE BEING SENT IN THE ASTM TRANSMISSION FROM THE CUSTOMER'S LABORATORY INFORMATION SYSTEM (LIS) TO AMPLILINK (AL) SOFTWARE VERSION 3.2.3. THIS FORMAT DOES NOT FOLLOW THE ASTM PROTOCOL DEFINITIONS PROVIDED IN THE AMPLILINK 3.2 HOST INTERFACE MANUAL. THIS CAUSED ORDERS TO BE REJECTED BY AL, AS THE ORDERS COULD NOT BE INTERPRETED. THE AMPLILINK TRACE FILES SHOWED THAT THE REJECTED ORDER WAS SENT TO THE LIS BUT THAT REJECTED ORDER WAS NOT HANDLED PROPERLY BY THE LIS. IF THE ORDER HAD BEEN SENT IN THE CORRECT FORMAT FROM THE LIS, THE ISSUE WOULD NOT HAVE OCCURRED. IN THIS PARTICULAR COMPLAINT CASE, A SOFTWARE CONCURRENCY ISSUE WAS OBSERVED. AT THE SAME MOMENT WHEN THE REJECTED ORDER WAS SENT TO LIS, A RESULT WAS ALSO SENT TO LIS. THESE TWO ASTM RECORDS WERE MIXED UP, AND THE RESULT OF SAMPLE (B)(4) WAS APPENDED TO THE REJECTED ORDER ID (B)(4). | 1
2497474 | A CUSTOMER REPORTED INCORRECT TIME/DATE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. | 1
2498173 | A CUSTOMER REPORTED THAT UNEXPECTED NEGATIVE REACTIONS WERE OBTAINED WITH THE POOLED CELL SCREEN FOR A DONOR SAMPLE TESTED ON NEO INSTRUMENT (B)(4). THE IMAGE RESULT FILES WERE REVIEWED BY AN IMMUCOR TECHNICAL SUPPORT SPECIALIST. THE IMAGES AGREED WITH THE INSTRUMENT INTERPRETATION. NO ROI MISALIGNMENT NOTED. NO PINHOLES WERE PRESENT IN THE CELL BUTTONS. POSITIVE AND NEGATIVE CONTROLS APPEARED VISUALLY POSITIVE AND NEGATIVE AS EXPECTED. UNABLE TO RULE OUT A SAMPLE-RELATED ISSUE AS THE CAUSE OF THE UNEXPECTED NEGATIVE REACTIVITY. IN ADDITION, THE INSTRUMENT INTERPRETATION WAS AT THE CUSP OF DETECTION. THE SAMPLE REACTED POSITIVE WHEN REPEATED MANUALLY WITH UNPOOLED DONORS. CUSTOMER WAS REFERRED TO THE LIMITATION SECTION OF THE PACKAGE INSERT WHICH STATES, "ANTIBODY SCREENING TESTS EMPLOYING POOLED REAGENT RED BLOOD CELLS WILL NOT BE AS SENSITIVE AS THOSE INCORPORATING THE RED BLOOD CELLS OF UNPOOLED SINGLE DONORS. POOLED REAGENT RED BLOOD CELLS SHOULD NOT BE USED WHEN THE ANTIGLOBULIN ANTIBODY SCREEN IS PERFORMED IN PLACE OF THE ANTIGLOBULIN CROSSMATCH." THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2498343 | THE CUSTOMER REPORTED THAT 7 MINUTES INTO A THERAPEUTIC PLASMA EXCHANGE (TPE) PROCEDURE USING A 5% ALBUMIN AS THE REPLACEMENT SOLUTION, A NEW TPE PATIENT BEGAN HAVING HIVES. THE PROCEDURE WAS ENDED AND IV BENADRYL AND SOLUMEDROL WERE GIVEN FOR THE ALLERGIC REACTION. WITHIN AN HOUR THE HIVES RESOLVED. THE CUSTOMER DECLINED TO PROVIDE THE PATIENT IDENTIFIER, DATE OF BIRTH OR WEIGHT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO MEDICAL INTERVENTION. (B)(4). INVESTIGATION: PER THE CUSTOMER, THE PATIENT HAD RUN OUT OF HER MEDICATIONS THE FRIDAY BEFORE THE PROCEDURE, AND HAD NOT HAD THEM REFILLED. SHE WAS NOT ON ANY ACE INHIBITORS. THE PHYSICIANS AND THE PHARMACIST AT THE CUSTOMER SITE BELIEVE THE REACTION WAS RELATED TO THE ALBUMIN THE PATIENT RECEIVED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: AFTER EVALUATING ALL THE INFORMATION, THE PATIENT'S PHYSICIAN AND PHARMACIST DETERMINED THAT THE REACTION WAS RELATED TO THE ALBUMIN THE PATIENT RECEIVED. NO ISSUES WITH THE COBE SPECTRA SYSTEM WERE REPORTED BY THE CUSTOMER. ALLERGIC REACTIONS ARE A KNOWN POSSIBLE SIDE EFFECT OF APHERESIS PROCEDURES. ACCORDING TO THERAPEUTIC APHERESIS: A PHYSICIAN'S HANDBOOK, TRANSFUSION REACTIONS MAY OCCUR IN AS MANY AS (B)(4) OF PROCEDURES. THE COBE SPECTRA OPERATOR'S MANUAL WARNS OPERATORS THAT WHEN USING BIOLOGICALLY DERIVED REPLACEMENT FLUIDS, THE PATIENT SHOULD BE CLOSELY MONITORED FOR REACTIONS. | 0
2500628 | A CUSTOMER CONTACTED BECKMAN COULTER INC (BEC) REGARDING THREE PATIENTS OF WHICH TWO GENERATED ERRONEOUS HIGH MCV AND LOW MCHC RESULTS WITHOUT INSTRUMENT GENERATED FLAGS AND ONE GENERATED LOW MCV AND HIGH MCHC RESULTS WITHOUT INSTRUMENT GENERATED FLAGS ON THREE UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEMS OVER THREE DIFFERENT DAYS. TWO PATIENTS (#1 AND #2) WERE DIAGNOSED WITH HYPONATREMIA. ERRONEOUS RESULTS WERE NOT REPORTED OUTSIDE OF THE LABORATORY AND THERE WAS NO DEATH, INJURY OR AFFECT TO PATIENT TREATMENT. THE CUSTOMER RERAN THE SAME SPECIMENS ON AN ACT DIFF INSTRUMENT WHICH RECOVERED LOWER MCV AND HIGHER MCHC RESULTS FOR PATIENTS #1 AND #3, AND RECOVERED HIGHER MCV AND LOWER MCHC FOR PATIENT # 2, WHICH THE CUSTOMER CONSIDERS CORRECT. THIS MDR REPORT DOCUMENTS THE RESULTS FOR PATIENT 2 OF 3; GENERATED BY THE SECOND DXH 800 INSTRUMENT (SERIAL NUMBER (B)(4)) ON (B)(6) 2012. ALL INSTRUMENTS WERE RUNNING WITHIN QC LIMITS. PER CUSTOMER, THEY ROUTINELY RUN CORRELATIONS BETWEEN THE INSTRUMENTS AND THEY ARE WITHIN THEIR ACCEPTABILITY LIMITS. SERVICE WAS NOT DISPATCHED FOR THIS EVENT. THE PROVIDED INSTRUMENT RAW COUNTS AND RBC HISTOGRAMS WERE REVIEWED IN DETAIL BY BEC. THERE WAS NO ABNORMALITY FOUND. THE PROVIDED DATA DOES NOT REVEAL ANY INFORMATION ABOUT THE CAUSE OF THE FAILURE. THE CAUSE FOR ERRONEOUS HIGH MCV AND LOW MCHC IS UNKNOWN. THE RESULTS GENERATED BY THE OTHER INSTRUMENTS ON DIFFERENT DAYS HAVE BEEN DOCUMENTED IN THE FOLLOWING MDR REPORTS: MDR #1061932-2012-00904 (DXH 800 SERIAL NUMBER (B)(4) ON (B)(6) 2012). MDR #1061932-2012-00953 (DXH 800 SERIAL NUMBER (B)(4) ON (B)(6) 2012). MDR #1061932-2012-00905 (DXH 800 SERIAL NUMBER (B)(4) ON (B)(6) 2012). MDR #1061932-2012-00952 (DXH 800 SERIAL NUMBER (B)(4) ON (B)(6) 2012). MDR #1061932-2012-00954 (DXH 800 SERIAL NUMBER (B)(4) ON (B)(6) 2012). | 0
2501852 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE ELEVATED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RDF INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THIS. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2502396 | IT WAS REPORTED THAT A PHYSICIAN'S HANDHELD COMPUTER WAS NOT CHARGING PROPERLY. THE COMPANY REPRESENTATIVE FOLLOWED UP WITH THE SITE AND CONFIRMED THAT THE COMPUTER WAS NOT PROPERLY CHARGING. WHEN THE SERIAL CABLE ADAPTER CONNECTOR WAS PUSHED UPWARD, IT WOULD CHARGE. HOWEVER WHEN THE ADAPTER CONNECTOR WAS RELEASED, THE BATTERY ICON WOULD STOP SHOWING THAT THE COMPUTER WAS BEING CHARGED. ATTEMPTS FOR PRODUCT RETURN HAVE BEEN UNSUCCESSFUL TO DATE.><IT WAS REPORTED THAT THE COMPANY REPRESENTATIVE WAS ABLE TO PICK UP THE SUSPECT SERIAL CABLE FOR RETURN TO THE MANUFACTURER. THE CABLE WAS RECEIVED ON (B)(4) 2012. HOWEVER, PRODUCT ANALYSIS HAS NOT BEEN COMPLETED TO DATE. THE COMPANY REPRESENTATIVE REPORTED THAT THE SERIAL CABLE ADAPTER CONNECTOR APPEARED BENT.><PRODUCT ANALYSIS FOR THE DELL X50 HANDHELD COMPUTER SERIAL DATA/POWER CABLE WAS COMPLETED ON (B)(6) 2012. IT CONFIRMED AN OPEN CONDUCTOR IN THE CABLE WHICH RESULTED IN UNRELIABLE OPERATION (POWER). THE CAUSE IS ASSOCIATED WITH A BROKEN WIRE CONNECTION IN THE SERIAL CABLE PLUG ASSEMBLY. ONCE THE WIRE CONNECTION WAS RESTORED USING A JUMPER WIRE, NO FURTHER ANOMALIES WERE IDENTIFIED AND FULL PRODUCT FUNCTIONALITY WAS RESTORED. .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2504063 | IT WAS REPORTED THAT DURING A CASE, THE GCX MOUNT HOLDING A KAPPA XLT MONITOR ON THIS APOLLO BROKE CAUSING THE MONITOR TO FALL NEARLY HITTING THE PATIENT. THERE WAS NO INJURY REPORTED. THE FAILED PARTS HAVE BEEN REQUESTED FOR INVESTIGATION BUT WERE NOT RECEIVED UNTIL NOW. THE INVESTIGATION RESULTS WILL BE PART OF A FOLLOW-UP REPORT. | 0
2504747 | THE ULTRASOUND MACHINE WHICH ALREADY HAD BEEN ACTING UP, FROZE 3 TIMES DURING THE MIDDLE OF THE EXAM, APPROXIMATELY EVERY 2-3 MINUTES, DELAYING THE EXAM LENGTH FOR THE PATIENT. THE PROBLEM IS BEING WORKED ON, HOWEVER THIS ISSUE HAS GOTTEN OUT OF HAND SINCE THE UNIT FROZE APPROXIMATELY 10 TIMES YESTERDAY. EACH TIME WE HAD TO APOLOGIZE TO THE PATIENT FOR THE DELAY. | 1
2505421 | THERE WAS 15 SAVED TELEMETRY STRIPS ALL LABELED WITHIN THE MONITOR HISTORY WHEN NURSE LEFT AT END OF SHIFT. UPON REPORTING BACK TO WORK ON NEXT SCHEDULED SHIFT, THERE WAS NO SAVED INFORMATION FOR THIS PATIENT. | 1
2505784 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE (RDF) WAS ANALYZED. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED. | 0
2506700 | FIELD SERVICE TECH FOR CAREFUSION REPORTS DRAWER ON PYXIS ANESTHESIA SYSTEM FAILED. NO PT PRESENT. (B)(4). ADD'L DATA/FAILURE INVESTIGATION: FIELD SERVICE TECHNICIAN INVESTIGATION DISCOVERED PHYSICAL ITEM OBSTRUCTION CAUSING DRAWER FAILURE. | 1
2507044 | IT WAS REPORTED THAT THE DEFICIT IS TOO HIGH AND LEAKING. ADDITIONAL INFO WILL BE PROVIDED WHEN THE INVESTIGATION HAS BEEN COMPLETED. | 0
2507313 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE DOUBLE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(4). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO AN ALLEGED DEVICE MALFUNCTION. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. THE MEASURED WBC COUNT OF THE PRODUCT DOES NOT QUALIFY AS A WBC FAILURE PER THE CURRENT FDA GUIDELINES OF 8 X 10^6 FOR A DOUBLE PLATELET PRODUCT COLLECTION. CONCLUSION: NO FAILURE OCCURRED. | 0
2508530 | IT WAS REPORTED THAT THE APEXPRO TELEMETRY SERVER (ATS) SHUT DOWN UNEXPECTEDLY DUE TO A POWER SUPPLY FAILURE, RESULTING IN A LOSS OF MONITORING AFFECTING 9 PATIENTS FOR 30 MINUTES. THERE WAS NO SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT, NOR WAS MEDICAL INTERVENTION REQUIRED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2508849 | THE CUSTOMER REQUESTED AN INVESTIGATION TO DETERMINE THE REASON THAT THEY HAVE HAD AN INCREASE IN PLATELET PRODUCTS WITH ELEVATED WHITE BLOOD CELL (WBC) CONTENT. THE CUSTOMER SUPPLIED 12 TOTAL INCIDENTS WITH THIS REQUEST. THIS REPORT IS BEING FILED FOR ONE OF THOSE INCIDENTS. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). PER A VERBAL REQUEST FROM THE FDA ON (B)(4) 2012, THE PREVIOUSLY REPORTED EVENT HAS BEEN BROKEN DOWN INTO INDIVIDUAL REPORTS FOR EACH DONOR. THE ORIGINAL REPORT WAS SUBMITTED TO THE FDA ON MEDWATCH REPORT# 1722028-2011-00153. INVESTIGATION: AS PREVIOUSLY REPORTED, THE CUSTOMER'S PRACTICES OF CHOOSING UNITS FOR QC WERE NOT COMPLETELY RANDOM. AT THE TIME, THE CURRENT CUSTOMER PROCESS PROVIDED ONLY TWO METHODS FOR THE PLATELET PROCESSING LAB TO SUBMIT A PRODUCT FOR WBC MEASURING: AS A PRODUCT WITH A "VERIFY WBCS" MESSAGE, OR AS A PRODUCT SELECTED VIA BACKGROUND SCREENING FOR WBCS. THE RUN DATA FILES (RDFS) WERE ANALYZED FOR EACH OF THE INCIDENTS SUBMITTED BY THE CUSTOMER. FOR THE INDIVIDUAL INCIDENT ON (B)(6) 2011, THE RDF SHOWS THAT THERE WAS LIKELY A PLASMA LINE OCCLUSION THAT CAUSED THE ELEVATED WBC COUNT. CONCLUSION: THE INCREASE IN ELEVATED WBC COUNTS WAS TRACED BACK TO A CHANGE IN THE PROCESS FOR DETERMINING WHICH SAMPLES WERE TO BE CHECKED FOR WBC CONTENT. THIS PROCESS WAS NO LONGER COMPLETELY RANDOM DUE TO THE USE OF A BACKGROUND SCREEN DURING PLATELET CELL COUNTING. CORRECTIVE ACTION: FOR PLASMA LINE OCCLUSION, ALGORITHMS WILL BE INCORPORATED INTO THE SOFTWARE FOR THE TRIMA ACCEL SYSTEM. THESE ALGORITHMS WILL RECOGNIZE A POTENTIAL ELEVATED WBC COUNT. THE SOFTWARE WILL FLAG THE PROCEDURE TO MEASURE OR VERIFY WBCS IN THE PLATELET PRODUCT. THE NEW ALGORITHMS WILL BE ADDED TO THE NEXT TRIMA EQUIPMENT SOFTWARE UPGRADE, VERSION 6.0.2, WHICH IS IN THE PROCESS OF BEING VALIDATED. | 0
2509610 | ONE OF FOUR MP30 WIRELESS MONITORS WAS DEFAULTED FOR MULTIPLE BED LABELS ON THE CENTRAL STATION. THIS IS A CONFIGURATION ISSUE AT MAIN NURSES STATION WHEN A MONITOR IS SELECTED. THE MONITOR/DEVICE ASSIGNED NUMBER AND THE LABEL/MONITOR # ARE NOT EASILY SELECTED AT THE CENTRAL STATION. AFTER THE PROBLEM WAS RESOLVED BY PHILIPS (CORRECTED THE CONFIGURATION FILE) THE POTENTIAL STILL EXISTS TO SELECT WRONG LABEL/LABEL/MONITOR #. | 1
2509736 | PATIENT ON TABLE FOR CARDIAC CATH LAB PROCEDURE, BUT PROCEDURE HAD NOT STARTED. TABLE INSTANTLY DROPPED, FALLING ABRUPTLY. NO ONE WAS NEAR TABLE TO ACCIDENTALLY PRESS BUTTONS OR LEVERS. ATTEMPTED TO RAISE TABLE TO PROPER HEIGHT AGAIN, AND THE SAME THING HAPPENED. NO ADVERSE OUTCOME.GE HEALTHCARE REPLACED THE TABLE LIFT ASSEMPLY TO REPAIR THE ISSUE. | 0
2511018 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR INVESTIGATION. THE WBC COUNT PER UNIT WAS NOT CLEARLY PROVIDED BY THE CUSTOMER. CARIDIANBCT IS AWAITING FURTHER INFO REGARDING THIS EVENT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT, THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. ADDITIONAL INFORMATION REGARDING SPECIFIC LOTS/MACHINES WAS NOT RECEIVED; CAUSE OF THE OBSERVED ISSUE REMAINS UNDETERMINED. | 0
2511134 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). ANALYSIS OF THE RBC DETECTOR SIGNAL OUTPUT SHOWS CONSIDERABLE FLUCTUATION SUGGESTING THE PRESENCE OF POSSIBLE AGGREGATES IN THE COLLECT LINE (PRE-PUMP INSIDE THE CHAMBER) LEADING TO DISRUPTION OF THE COLLECTION. PRESENCE OF AGGREGATES MAY DISTURB THE PLATELET BED IN THE LRS CHAMBER AND THUS LEAD TO INADVERTENT RELEASE OF WHITE CELLS INTO THE PRODUCT. A DEFINITE CAUSE FOR THE AGGREGATES COULD NOT BE DERIVED FROM THE DLOG. FURTHERMORE THERE IS A RELATIONSHIP BETWEEN HIGH LEVEL CONTAMINATION EVENTS (SATURATION) AND DONORS WITH ELEVATED BODY MASS INDEX (>30). IN THIS CASE THE BMI WAS CLOSE TO THIS LEVEL (=29.4). A POSSIBLE MISSED SATURATION EVENT CAN THEREFORE NOT COMPLETELY BE EXCLUDED EITHER. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED. | 0
2511886 | A CUSTOMER CONTACTED BECKMAN COULTER INC., (BEC) TO REPORT THAT THEY OBSERVED A DILUENT LEAK UNDERNEATH THE COULTER LH 750 SLIDEMAKER TO THE LEFT OF THE MANUAL PROBE. THE FLUID LEAK WAS NOT CONTAINED WITHIN THE INSTRUMENT. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND GLASSES AT THE TIME OF THE EVENT. THERE WAS NO EXPOSURE TO MUCOUS MEMBRANES OR OPEN WOUNDS AND MEDICAL ATTENTION WAS NOT SOUGHT. THE MATERIAL SAFETY DATA SHEET (MSDS) WAS NOT REVIEWED. THERE IS AN EXPOSURE CONTROL/RISK MANAGEMENT PLAN IN PLACE AT THE FACILITY. THERE WAS NO IMPACT TO PATIENT RESULTS. THERE WAS NO DEATH, INJURY OR CHANGE TO PATIENT TREATMENT ATTRIBUTED OR ASSOCIATED TO THIS COMPLAINT. A BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON (B)(4) 2012 FOR THIS EVENT. THE FSE FOUND TUBING THROUGH VALVE LINE (VL2) LEAKING. THE FSE REPLACED THE TUBING, PERFORMED VIP VERIFICATION, CLEANED COMPRESSOR'S AIR FILTERS, AND CHECKED VACUUMS AND PRESSURES. SERVICE WAS VERIFIED TO MEET THE SPECIFIED REQUIREMENTS PER ESTABLISHED PROCEDURES. RESULTS MEET PUBLISHED PERFORMANCE SPECIFICATIONS. PER E-MAILS FROM THE FSE, THE VL2 LINE IS ASSOCIATED WITH THE SAMPLE ACCESS MODULE (SAM)/SLIDEMAKER AND CLOSES THE LINE ON THE FIRST T-FITTING. ONLY THE TUBING WAS LEAKING; IT WAS WORN BY THE PINCH VALVE. THE LEAK CONSISTED OF A LARGE AMOUNT OF DILUENT ONLY. PER LABELING, BECKMAN COULTER, INC. URGES ITS CUSTOMERS TO COMPLY WITH ALL NATIONAL HEALTH AND SAFETY STANDARDS SUCH AS THE USE OF BARRIER PROTECTION. THIS MAY INCLUDE, BUT IT IS NOT LIMITED TO, PROTECTIVE EYEWEAR, GLOVES, AND SUITABLE LABORATORY ATTIRE WHEN OPERATING OR MAINTAINING THIS OR ANY OTHER AUTOMATED LABORATORY ANALYZER. THE CAUSE OF THE LEAK IS ATTRIBUTED TO A HOLE WORN IN THE TUBING AT VL2. (B)(4). | 0
2511906 | THE CUSTOMER REPORTED SMALL AMOUNT OF FLUID, CONSISTED OF BLOOD, LEAKED INSIDE NEAR THE RESERVOIR ROD INVOLVING LH SLIDEMAKER. BECKMAN COULTER CUSTOMER TECHNICAL SUPPORT (CTS) RECOMMENDED THE USE OF PERSONAL PROTECTIVE EQUIPMENT (PPE) PRIOR TO TROUBLESHOOTING. THE CUSTOMER HAD ON A LABORATORY COAT AND GLOVES PRIOR TO TROUBLESHOOTING. THE CUSTOMER NOTED THE FLUID LEAKED AT THE PINCH VALVE NEAR THE RESERVOIR ROD AND ONTO THE COUNTERTOP. THE CUSTOMER FOLLOWED THE EXPOSURE CONTROL/RISK MANAGEMENT PLAN AND CLEANED UP THE SPILL. NO PATIENT RESULTS WERE IMPACTED. THE CUSTOMER DID NOT COME INTO DIRECT CONTACT WITH THE FLUID. THERE WAS NO EXPOSURE TO OPEN LESIONS OR MUCUS MEMBRANES. THERE WAS NO REPORT OF INJURY OR ADVERSE EFFECT ASSOCIATED WITH THIS EVENT. THE FIELD SERVICE ENGINEER (FSE) ASSESSED THE UNIT AT THE FACILITY. THE FIELD SERVICE ENGINEER (FSE) DISCOVERED THE RED STRIPE PINCH VALVE TUBING ON DISPENSE PROBE WAS CUT AND BLOOD WAS SPRAYED INSIDE THE SLIDEMAKER. THE FSE REPLACED THE DISPENSE PROBE PINCH VALVE, PINCH VALVE MOUNT, ACTUATOR, AND PINCH VALVE TUBING. THE FSE REMOVED SEVERAL SECTIONS/MODULES AND CLEANED THE FLUID LEAK. HOUSEHOLD BLEACH WAS USED TO CLEAN AND DISINFECT SURFACES. THE UNIT CONFORMED TO THE MANUFACTURER'S PERFORMANCE SPECIFICATIONS. (B)(4). | 0
2512340 | WHILE PERFORMING AN INTERNAL REVIEW OF PROGRAMMING HISTORY IT WAS NOTED THAT A VNS DEVICE WAS INTERROGATED (B)(6) 2008 AND NOTED TO BE SET AT (1.75/15/500/30/3). A SYSTEM DIAGNOSTIC TEST WAS PERFORMED. THE SYSTEM DIAGNOSTIC RESULTED IN "FAULT" COMMUNICATION, THE DEVICE SETTINGS WERE CHANGED (1/20/500/30/60) AND NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE VISIT. THE NEXT RECORDED VISIT ((B)(6) 2008), THE DEVICE WAS INTERROGATED AND THE SETTINGS WERE 1/20/500/30/60, PRIOR TO THE PATIENT LEAVING THE DEVICE SETTINGS WERE PROGRAMMED TO INTENDED SETTINGS OF 1.25/20/500/30/5. ANALYSIS OF PROGRAMMING HISTORY CONFIRMED EVENT. | 0
2512883 | THE USER OF THE WORK AREA MANAGER (WAM) V 3.0 USED IN CONJUNCTION WITH THEIR SYSMEX HEMATOLOGY ANALYZERS REPORTED THAT RESULTS WERE AUTO VALIDATING THAT SHOULD HAVE BEEN HELD FOR ADDITIONAL REVIEW. FOUR SAMPLE RESULTS WERE IDENTIFIED HAS HAVING AUTO-VALIDATED, AND ONE REQUIRED A CORRECTED REPORT. THREE OF THE FOUR AUTO-VALIDATED RESULTS HAD CRITICAL VALUES. THE USER REPORTED SEVERAL COMMUNICATION ISSUES BEGINNING WITH THEIR LAB INFORMATION SYSTEM (LIS), MEDITECH, COMMUNICATING WITH WAM, BUT DATA WAS NOT TRANSFERRING BETWEEN THE INSTRUMENT, LINE CONTROLLER (LASC) AND WAM, LATER, THE AUTO-VALIDATION CONCERN WAS ESCALATED TO SYSMEX WAM TECHNICAL ASSISTANCE CENTER WHO DETERMINED THAT THE RULES ENGINE FILE COULD NOT BE VIEWED DUE TO AN ERROR MESSAGE "FAILED TO FIND GLOBAL PROC LIBRARY MODULE CA_VAL_FL_CHG". THE RULES FILE WAS FOUND TO BE CORRUPT, CAUSING RULES NOT TO TRIGGER TO HOLD THE RESULTS OR ALERT THE USER OF CRITICAL VALUES. THE CORRUPTED RULES TABLE WAS RECONSTRUCTED BY COPYING THE RULES IN THE TEST ENVIRONMENT TO THE PRODUCTION ENVIRONMENT TO RESOLVE THE ISSUE. THERE WAS NO ALLEGATION OF DELAY IN TREATMENT DUE TO THE AUTO-VALIDATION OF CRITICAL RESULTS, NOR ANY ADVERSE IMPACT AS A RESULT OF THE CORRECTED REPORT. THE SYSMEX WAM SOFTWARE APPLICATION IS AN INFORMATION MANAGEMENT TOOL THAT PROVIDES COMPREHENSIVE MANAGEMENT OF HEMATOLOGY INSTRUMENT TEST ORDER INFO AND RESULTS DATA. THIS SOFTWARE IS CONSIDERED A CLASS I MEDICAL DEVICE DATA SYSTEM (MDDS). THE WAM CONSOLIDATES RESULTS DATA FROM MULTIPLE INSTRUMENTS, EVALUATES RESULTS USING TEST ORDER RULES CRITERIA AND COMMUNICATES THIS INFO TO THE LIS. THE RULES ARE CRITICAL TO THE INTENDED USE OF THIS APPLICATION. RESULTS FROM THE INSTRUMENT AND WAM WERE CORRECTLY DISPLAYED, BUT THE DATABASE TABLE FOR THE RULES WAS CORRUPTED. THE ROOT CAUSE MAY BE ASSOCIATED WITH A POSSIBLE DATABASE CORRUPTION AFTER A DISK IN THE RAID ARRAY WAS IMPACTED BY A BATTERY FAILURE IN (B)(6) 2012. THE INTEGRITY OF A SECTOR ON THE HARD DISK USED BY ORACLE TO READ/WRITE DATA FROM THE RULES TABLE MAY HAVE BEEN COMPROMISED. THE FACILITY WAS AWARE OF NUMEROUS COMMUNICATION ISSUES INCLUDING SERVER MAINTENANCE, LASC RESTARTS, BATCH JOBS NOT RUNNING, ETC, WITHIN THE PAST FEW DAYS, BUT THE RULES NOT TRIGGERING TO REQUIRE FURTHER REVIEW AND VALIDATION BY THE USER PRIOR TO RELEASE OF THE RESULTS, OR TO ALERT THE USER TO REPORT CRITICAL VALUES TO THE PHYSICIAN WAS NOT IMMEDIATELY APPARENT. THE ISSUE WAS RECOGNIZED AND THE SAMPLES WERE MANAGED APPROPRIATELY BY THE LAB WITHIN 2-3 HOURS OF THE ORIGINAL SAMPLE ANALYSIS (PHYSICIAN CALLED, RESULTS VERIFIED, REPORT AMENDED). THE INTENDED USE OF THE WAM AND HEMATOLOGY ANALYZERS IS FOR IN VITRO DIAGNOSTIC USE IN SCREENING PT POPULATIONS FOUND IN CLINICAL LABORATORIES. RESULTS THAT WERE INCONSISTENT WITH THE PT'S OTHER CLINICAL SIGNS AND SYMPTOMS REQUIRE VERIFICATION PRIOR TO CLINICAL DECISION MAKING. | 1
2512885 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVAL. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE RDF DOES NOT SHOW ROOT CAUSE OF THE ELEVATED WBC COUNT MEASURED. IT IS POSSIBLE THAT THE FAILURE IS PART OF THIS SITE'S STATISTICAL DISTRIBUTION FOR LEUKOREDUCTION. IT IS POSSIBLE THAT NUMEROUS ACCESS PRESSURE ALERTS COULD HAVE DISRUPTED THE STEADY STATE OF THE SYSTEM AND COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC COUNT. THE FAILURE MAY ALSO BE DONOR RELATED. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
2513741 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #(B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE ANALYSIS OF THE RUN DATA FILE (RDF) DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED. | 0
2513775 | UPON REVIEWING A PATIENT'S PROGRAMMING HISTORY, IT WAS OBSERVED THAT A FAULTED SYSTEMS TEST OCCURRED ON (B)(6) 2006 WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED PARAMETERS. THE PATIENT'S OFF TIME WAS 1.1 MINUTES ON (B)(6) 2004, BUT THEN A FAULTED SYSTEMS TEST ON (B)(6) 2006 OCCURRED WHICH CHANGED THE OFF TIME TO 60MIN. THE OFF TIME WAS PROGRAMMED TO 180MIN ON THIS VISIT. HOWEVER, THE PATIENT'S SETTINGS WERE NOT CORRECTED TO EFFICACIOUS SETTINGS PRIOR TO THE LEAVING THE CLINIC. ANALYSIS OF PROGRAMMING HISTORY. | 0
2514386 | THE PRODUCT INFORMATION WAS PROVIDED, AND IT WAS REPORTED THAT THE DELL X50 HANDHELD COMPUTER WOULD REPEATEDLY FREEZE ON THE MAIN MENU SCREEN AND WOULD HAVE TO BE HARD RESET. THIS EVENT REPORTEDLY HAPPENED SEVERAL TIMES OVER A ONE MONTH PERIOD. THE HANDHELD COMPUTER AND RELATED FLASHCARD/SOFTWARE WERE RECEIVED BY THE MANUFACTURER ON (B)(6) 2012. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND FLASHCARD/SOFTWARE, AND THE REPORTED ALLEGATION WAS VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE, AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE HANDHELD, FLASHCARD, AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT A PHYSICIAN'S HANDHELD COMPUTER WAS LOCKING UP, AND THEREFORE, THE PHYSICIAN WAS REQUESTING A REPLACEMENT. NO ADDITIONAL INFORMATION WAS PROVIDED AND ATTEMPTS FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. ALTHOUGH PRODUCT RETURN IS EXPECTED, THE DEVICE HAS NOT BEEN RECEIVED TO DATE. DATE OF EVENT, CORRECTED DATA: THE DATE OF EVENT WAS INADVERTENTLY INCORRECT ON THE INITIAL REPORT. SERIAL #, LOT#, OTHER, CORRECTED DATA: THE PRODUCT INFORMATION WAS UNKNOWN ON THE INITIAL REPORT. DEVICE MANUFACTURER DATE, CORRECTED DATA: THE DATE WAS UNKNOWN ON THE INITIAL REPORT, AS THE PRODUCT INFORMATION WAS UNKNOWN.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2515433 | CALLER ALLEGED THAT THEIR LDX PRINTER IS SMOKING AND GETS EXTREMELY HOT. THE PRINTER BURNED A WHOLE THROUGH A LABEL WHILE PRINTING. CUSTOMER ALSO REPORTED THAT THE PRINTER DOES NOT POWER OFF AND HAS TO BE UNPLUGGED TO TURN OFF. NO REPORT OF DEATH OR SERIOUS INJURY. TECHNICAL SERVICES REPLACED PRINTER. INVESTIGATION PENDING. | 1
2515656 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) ON (B)(6) 2012 STATING THAT ON AVERAGE AT LEAST ONCE PER DAY A SPECIMEN TUBE LEAKS BLOOD FROM THE TUBE TOP WHERE THEIR UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM HAS PUNCTURED THE TUBE DURING SAMPLING AND CAUSES THE TUBE TO LEAK BLOOD DURING MIXING. PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT WAS WORN BY THE USER. THE LAB'S EXPOSURE CONTROL/RISK MANAGEMENT PLANS ARE IN PLACE. NO INJURY OR EXPOSURE WAS REPORTED. THERE HAS BEEN NO REPORT OF ANY PATIENT RESULTS BEING AFFECTED. A BEC FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND OPTIMIZED THE SAMPLE ACCESS MODULE (SAM) ADJUSTMENT AND PIERCING ADJUSTMENT. THE FAILURE MODE FOR TUBE LEAKING AFTER INSTRUMENT SAMPLING PUNCTURE IS UNKNOWN. (B)(4). | 0
2517252 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT 1 ML LEAK OF CLEAR LIQUID LEAKED FROM THE BLOOD SAMPLING VALVE (BSV) AREA CONTAINED WITHIN THEIR UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, GLOVES AND EYE PROTECTION AT THE TIME OF THE INCIDENT. THE LAB'S EXPOSURE CONTROL PLAN IS IN PLACE. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO REPORT OF ANY PATIENT RESULTS BEING AFFECTED BY THIS INCIDENT. A SERVICE VISIT WAS SET UP HOWEVER, THE CUSTOMER CANCELLED THE REQUEST. THE FIELD SERVICE ENGINEER (FSE) PERFORMED TROUBLESHOOTING OVER THE PHONE AND FOUND THAT THE ASPIRATION LINE THAT EXTENDS FROM THE ASPIRATION PROBE TO THE BSV HAD POPPED OFF ON THE BSV END. THE CUSTOMER WAS ABLE TO RECONNECT THE TUBING TO THE BSV AND THE LEAK WAS RESOLVED. THE CAUSE OF THE LEAK IS THE TUBING THAT POPPED OFF THE BSV. (B)(4). | 0
2519768 | IT WAS REPORTED THAT THE CARIOCAP 5 MONITOR WAS NOT AUDIBLY ALARMING. THE VISUAL ALARMS WERE REPORTEDLY STILL FUNCTIONAL. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2520845 | IT WAS REPORTED ON (B)(6) 2012, THAT A PHYSICIAN WAS HAVING ISSUES WITH HIS PROGRAMMING SYSTEM. THE DELL X50 HANDHELD DEVICE SCREEN HAS BLACK LINES IN THE DISPLAY HOWEVER NO PATIENT HAS BEEN AFFECTED. THERE WAS NO EVIDENCE OF ANY PHYSICAL DAMAGE TO THE HANDHELD SCREEN OR CASING. THE EDGES OF THE SCREEN AND UNDERNEATH THE PLASTIC WERE CLEANED BUT THE LINES WERE STILL PRESENT. A REPLACEMENT HANDHELD DEVICE WAS SENT TO THE PHYSICIAN AND THE DELL X50 HANDHELD IN QUESTION WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS ON (B)(6) 2012. PRODUCT ANALYSIS IS UNDERWAY AND HAS NOT BEEN COMPLETED AT THIS TIME.><PRODUCT ANALYSIS OF THE RETURNED HANDHELD DEVICE AND FLASHCARD WAS COMPLETED ON (B)(6) 2012. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DURING THE ANALYSIS OF THE HANDHELD, IT WAS IDENTIFIED THAT THE DISPLAY HAD SOME DEAD PIXELS. THE DEAD PIXELS HAD NO FUNCTIONAL IMPACT ON THE HANDHELD PERFORMANCE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2521002 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521013 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521026 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521028 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521030 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521036 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521037 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521053 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2521966 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING THAT THEIR COULTER LH 750 SLIDEMAKER WAS MAKING SLIDES WITH BLOOD ON THE SLIDE LABEL (CONTAINED TO THE LABEL ONLY, NOT INSIDE OR OUTSIDE OF THE INSTRUMENT). THE CUSTOMER TRIED FLUSHING THE DISPENSE PROBE BUT THIS DID NOT RESOLVE THE ISSUE. NO PATIENT RESULTS HAVE BEEN AFFECTED AND SMEARS ARE BEING MADE WELL. THE OPERATORS WERE WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, EYE WEAR AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE THE DAY OF THE EVENT AND FOUND SPECIMEN BLOOD MIGRATING ONTO LABEL. UPON EXAMINING THE UNIT, FSE FOUND THE PROBE TRUCK NOT EVACUATING PROPERLY. FSE VERIFIED OPERATION OF ASSOCIATED VALVES, FLUSHED VACUUM PATHWAY AND CLEANED DISPENSE PROBE AREA. THE REPAIRS PERFORMED BY THE FSE RESOLVED THE LEAK. PER FSE, WHILE PROBE WAS BACKWASHING INTO PROBE WASTE CUP, THE CUP WAS NOT DRAINING FLUID AS FAST AS IT NORMALLY WOULD DUE TO A PARTIALLY BLOCKED VACUUM PATHWAY, CAUSING THE WASTE TO STAY LONGER IN THE CUP AND THEREFORE WICK ONTO THE BOTTOM OF DISPENSE PROBE MECHANISM, WHICH WOULD THEN TRANSFER ON TO THE NEXT SLIDE LABEL. (B)(4). | 0
2522673 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2522675 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2522676 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2522677 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 1
2522680 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 1
2522696 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 0
2522703 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 1
2522704 | ABNORMAL CELLS MISSED BY THINPREP IMAGER SYSTEM. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW PRESENTED BY THE IMAGER. | 1
2522859 | CENTRAL MONITOR FOR 4TH FLOOR TELEMETRY WENT DOWN COMPLETELY ON 3RD AND 4TH FLOOR. UNABLE TO VIEW PATIENTS RHYTHMS FOR APPROXIMATELY 5 MINUTES. SYSTEM REBOOTED AND VISIBILITY (WAVEFORMS) RETURNED. CONFIRMED CORRECT PATIENTS APPEARING ON SCREEN. | 1
2523071 | ALL IMAGES SENT TO THE AMICAS VER. 6 SERVER ARE NOT AVAILABLE TO REVIEW. AT THE TIME OF THE INCIDENT, THE DATABASE WAS CONFIGURED WITH "AUTO TURNING" TURNED ON. THE DATABASE RAN OUT OF RESOURCES AND COULD NOT HANDLE THE TRANSACTIONS AT THAT TIME. THE IMPORTING DEVICE'S ERROR HANDLING ONLY REPORTED THE UNDERLYING DATABASE PROBLEM IN THE LOG FILES WITHOUT MARKING THE MISSING IMAGES AS "FAILED TO IMPORT". THUS, THE IMPORTING DEVICE MOVED TO CLOSE THE STUDY WITHOUT LEAVING THE FAILED IMAGES TO BE RE-TRIED. THE DICOM DELETION CODE SHOULD NOT HAVE DELETED THE DICOM IMAGES THAT ARE FAILED. INSTEAD, IT SHOULD LEAVE THE IMAGES SO THEY CAN BE RE-TRIED. THE ISSUE WAS IDENTIFIED AND CORRECTED IN V6, CU4 AND INSTALLED. | 1
2524814 | THE CUSTOMER REPORTED THAT INTERPRETATION STATEMENTS FROM REPORTS HAVE EEN TRUNCATED WHEN SENT AS FREE TEXT FROM XCELERA BY XCELERA CONNECT. NO PT HARM WAS REPORTED. (B)(4). THE CUSTOMER REPORTED THAT INTERPRETATION STATEMENTS REPORTS HAVE BEEN TRUNCATED WHEN SENT AS FREE TEXT FROM XCELERA BY XCELERA CONNECT. NO PT HARM WAS REPORTED. PHILIPS IS IN THE PROCESS OF OBTAINING ADD'L INFO REGARDING THIS INCIDENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETED. | 1
2528606 | THE SCREEN WOULD FREEZE OCCASIONALLY ON A SCREEN OTHER THAN THE FINAL INTERROGATION SCREEN, PER THE PHYSICIAN'S OFFICE. HOWEVER, IT WAS RESOLVED THROUGH RESET OR BY TAKING THE BATTERY OUT AND PRODUCT ANALYSIS FOUND NO ANOMALIES WITH THE SOFTWARE.><ADDITIONAL INFORMATION WAS RECEIVED REGARDING THE PRODUCT INFORMATION AND THE PROGRAMMING SYSTEM REPORTEDLY "NOT WORKING AS WELL AND NOT HOLDING A CHARGE." THE HANDHELD COMPUTER WOULD REPORTEDLY FREEZE, AND THE PHYSICIAN WOULD HAVE TO RESET THE COMPUTER OR TAKE THE BATTERY OUT. NO MISHANDLING WAS SUSPECTED TO HAVE CONTRIBUTED TO THE EVENT. THE PROGRAMMING SYSTEM WAS NOT STORED IN AN ABNORMAL LOCATION; IT STAYED PLUGGED IN THE PHYSICIAN'S OFFICE. THE COMPANY REPRESENTATIVE BELIEVES THAT THE FROZEN SCREEN WOULD OCCUR ON THE INTERROGATION SCREEN, BUT THIS HAS NOT BEEN CONFIRMED BY THE PHYSICIAN TO DATE. A REPLACEMENT HANDHELD COMPUTER WAS SENT TO THE PHYSICIAN. THE PROGRAMMING SYSTEM AND RELATED SOFTWARE WERE RECEIVED ON (B)(4) 2012, BY THE MANUFACTURER. AN ANALYSIS WAS PERFORMED ON THE HANDHELD COMPUTER AND THE REPORTED ALLEGATION WAS NOT VERIFIED. THE HANDHELD WAS RECEIVED WITH A MAIN BATTERY CHARGE OF 64%. DURING THE ANALYSIS, NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT A PHYSICIAN'S OLD HANDHELD COMPUTER WAS "NOT WORKING AS WELL" AND NOT HOLDING A CHARGE. THEREFORE, THE PHYSICIAN WAS REQUESTING A NEW COMPUTER. ATTEMPTS FOR ADDITIONAL INFORMATION AND PRODUCT RETURN HAVE BEEN UNSUCCESSFUL THUS FAR. PRODUCT RETURN IS EXPECTED, BUT THE PRODUCTS HAVE NOT BEEN RECEIVED TO DATE.><FOLLOW UP WITH THE PHYSICIAN'S OFFICE CONFIRMED THAT THE SCREEN WOULD FREEZE OCCASIONALLY ON THE INTERROGATION SCREEN, RESOLVED THROUGH TROUBLESHOOTING. THE PROGRAMMING SYSTEM IS WORKING OKAY TO DATE. NO ADDITIONAL INFORMATION WAS PROVIDED. THE SUPPLEMENTAL REPORT #2 INADVERTENTLY REPORTED THE SCREEN IN WHICH THE SCREEN WOULD FREEZE INCORRECTLY. THE SUPPLEMENTAL REPORT #1 INADVERTENTLY DID NOT INDICATE THE DEVICE EVALUATION.><SUSPECT DEVICE SERIAL #, LOT #, AND OTHER, CORRECTED DATA: THE INITIAL REPORT DID NOT REPORT THE PRODUCT INFORMATION, AS THE INFORMATION WAS UNKNOWN AT THAT TIME. DEVICE MANUFACTURER DATE, CORRECTED DATA: THE INITIAL REPORT DID NOT REPORT THE PRODUCT INFORMATION, AS THE INFORMATION WAS UNKNOWN AT THAT TIME.><. | 0
2532358 | POSSIBLE CONTAINED POTENTIAL CONTAMINATED PROVIDINE IODINE SWABS. | 0
2532751 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT # (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE (RDF) WAS ANALYZED. THE ANALYSIS OF THE RUN DATA FILE DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED. | 0
2532819 | ONE MORNING OF JANUARY 2012 ANESTHESIA MACHINES IN 7 ROOMS WOULD RESET WITH A "WATCHDOG TIMEOUT FAILURE" ALERT ON VENTILATOR MONITORS PERIODICALLY AFTER START OF CASES. DAILY MACHINE CHECKS WERE PERFORMED AS USUALLY ON MACHINES PRIOR TO THE START OF 0730 CASES WITHOUT PROBLEM. IT WAS QUICKLY DETERMINED THAT BY REMOVING SERIAL CONNECTION FROM ANESTHESIA MACHINES TO COMPUTERS RUNNING PICIS ANESTHESIA MANAGER, WATCHDOG ERROR'S MESSAGES STOPPED. AFTER FURTHER ANALYSIS OF SITUATION, IT WAS DISCOVERED THAT THE PICIS SYSTEM ADMINISTRATOR PERFORMED AN UPGRADE TO PICIS WORKSTATIONS AND DID NOT INSTALL DRIVERS IN THE PROPER LOCATION WITHIN PICIS APPLICATION. MACHINE CHECKS WERE SUCCESSFUL BECAUSE A PICIS PATIENT RECORD WAS NOT PULLED UP DURING MACHINE VERIFICATION. THERE WERE APPROXIMATELY 3 RESETS TO EACH MACHINE IN ABOUT 2 MINUTE INTERVALS BEFORE PROBLEM WAS IDENTIFIED. PATIENTS WERE BAGGED AND MONITORED DURING TEMPORARY MACHINE OUTAGES AND DID NOT SUFFER ANY ADVERSE CONSEQUENCES.======================MANUFACTURER RESPONSE FOR ANESTHESIA MACHINE, ANESTHESIA MACHINE (PER SITE REPORTER).======================WATCHDOG TIMEOUT ERROR IS CONSISTENT WITH A MACHINE COMMUNICATIONS PROBLEM. | 1
2534592 | THE REPORTER ORIGINALLY REPORTED 1 ISSUE ON (B)(6) 2012, HOWEVER, PER A FOLLOW-UP TELEPHONE CONVERSATION ON (B)(6) 2012 WITH THE COMPLAINT ORIGINATOR, THE CONTACT STATED THAT DURING AN ARTHROSCOPIC SHOULDER REPAIR WITH THE USE OF AN FMS PUMP FOR THE FLUID MANAGEMENT, THE PATIENT EXPERIENCED EXTRAVASATION DURING THE PROCEDURE; NO INTERVENTION WAS PERFORMED. THE PROCEDURE WAS CONCLUDED SUCCESSFULLY WITHOUT FURTHER ISSUES OR HARM TO THE PATIENT; HOWEVER, BECAUSE OF THE EXTRAORDINARY SWELLING DUE TO THE EXTRAVASATION WHICH EXTENDED TO ABOVE THE COLLAR BONE AND DOWN TO THE ELBOW, THE PATIENT WAS ICED DOWN AND MONITORED FOR CONTINUUM OF AIR WAY INTEGRITY. THE REPORTER STATED THAT THIS EXACT ISSUE WITH THIS SAME DEVICE HAS HAPPENED AT LEAST 4 OTHER TIMES: NONE REPORTED; NO DATES OR OTHER INFORMATION COULD BE PROVIDED FOR THESE PREVIOUS INCIDENTS. ALSO SEE ASSOCIATED MDRS 1221934-2012-00092, 1221934-2012-00093, 1221934-2012-00094 AND 1221934-2012-00095. MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT.><THE COMPLAINT DEVICE WAS RECEIVED AND EVALUATED. UPON RECEIPT, THE DEVICE WAS GIVEN A THROUGH VISUAL AND FUNCTIONAL EXAMINATION. THE 1ST FAZE OF THE EXAMINATION WAS VISUAL. THIS VISUAL IS DONE AS A MATTER OF COURSE TO DISCERN IF THE DEVICE HAS ANY OBVIOUS PHYSICAL DAMAGE, SOMETHING THAT MAY OR COULD HAVE BEEN A FACTOR IN CONTRIBUTING TO THE REPORTED EVENT. IT IS NOTED THAT THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE 2ND PORTION OF THE EXAMINATION WAS THE FUNCTIONAL TESTING. THIS PORTION OF THE EXAM IS PERFORMED TO ASSESS THE DEVICE'S OPERATIONAL STATUS. THE PERFORMANCE AND FUNCTIONAL TESTING REVEALED THAT THE DEVICE HAD A PERFORMANCE/COMPONENT ANOMALY THAT MAY OR MAY NOT HAVE BEEN CONTRIBUTORY TO THE ADVERSE EVENT; WE CANNOT TELL; TECHNIQUE MAY ALSO HAVE PLAYED A ROLE IN THE EVENT. OUTSIDE OF THESE CONSIDERATIONS, WE CANNOT DISCERN A ROOT OR UNDERLYING CAUSE FOR THE REPORTED ADVERSE EVENT. AT THIS POINT IN TIME, NO CORRECTIVE OR FURTHER ACTION IS WARRANTED, HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FUTURE INFORMATION RECEIVED THAT IS RELATIVE AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
2534700 | IT WAS REPORTED THAT THE APEXPRO TELEMETRY SERVER (ATS) SHUT DOWN UNEXPECTEDLY, RESULTING IN A LOSS OF MONITORING AFFECTING 12 TO 16 TELEMETRY PT FOR ONE HOUR. THE LOSS OF MONITORING WAS NOTICED IMMEDIATELY, AND THERE WAS NO DEATH OR SERIOUS INJURY ASSOCIATED WITH THIS EVENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2537925 | IT WAS REPORTED ON (B)(6) 2012, THAT THE PHYSICIAN'S DELL X5 HANDHELD DEVICE WAS NOT WORKING PROPERLY. THE ISSUE HAS BEEN OCCURRING FOR SEVERAL MONTHS. THE PHYSICIAN HAS BEEN UNABLE TO COMMUNICATE WITH PATIENTS' DEVICES UNLESS HE HOLDS THE HANDHELD AND SERIAL CABLE TO AVOID ANY MOVEMENT. A COMPANY REPRESENTATIVE WENT TO THE SITE TO PERFORM TROUBLESHOOTING AND VERIFIED THAT THE ISSUE WAS NOT DUE TO THE WAND. THERE WERE NO SIGNS OF DAMAGE TO THE HANDHELD. THE HANDHELD HAS BEEN REPLACED AND THE HANDHELD IN QUESTION HAS BEEN RETURNED TO THE MANUFACTURER. PRODUCT ANALYSIS IS UNDERWAY AND HAS NOT BEEN COMPLETED AT THIS TIME.><ANALYSIS OF THE RETURNED HANDHELD DEVICE AND ASSOCIATED SOFTWARE WAS COMPLETED AND APPROVED ON (B)(6) 2012. NO ANOMALIES WERE FOUND WITH EITHER THE HANDHELD DEVICE OR SOFTWARE AND BOTH PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE POWER CABLE AND SERIAL CABLE WERE ALSO RETURNED AND NO ISSUES WERE FOUND WITH THE PERFORMANCE OF EITHER CABLE. . | 0
2538971 | A PATIENT WAS SEDATED FOR AN MRI OF HIS SHOULDER. APPROPRIATE PADDING WAS USED FOR UPPER EXTREMITIES. AT CONCLUSION OF MRI AND WHEN PATIENT AWOKE, COMPLAINED OF ACHING IN ELBOWS. NURSING INTERVENTION WAS IMPLEMENTED TO TREAT JOINT PAIN - HEAT, ALTERNATED WITH COLD. HOT PACKS WERE APPLIED FOR 30 MINUTES, THEN REMOVED FOR ICE PACKS. SKIN WAS INTACT FOR BOTH APPLICATIONS. WHEN COLD PACKS REMOVED, CIRCULAR BLISTERS WERE NOTED MEASURING 2CM X 2.5CM. THERE WERE ALSO CIRCULAR REDDENED MARKS ABOUT 1/2 INCH AROUND BLISTERS MEASURING 6CM X 5CM ON RIGHT ELBOW AND 4CM X 5CM ON LEFT ELBOW. ELBOWS SLIGHTLY SWOLLEN AS WELL. PHYSICIAN EXAMINED, AND DETERMINED BURNS WERE FROM MRI. INSTRUCTED PATIENT ON CARE OF BLISTERS.MANUFACTURER'S FIELD ENGINEERS CAME ON SITE AND CONDUCTED TESTING. FOUND NOTHING WRONG WITH SCANNER, IT IS WORKING APPROPRIATELY. | 0
2539762 | THE GE MEDICAL PATIENT MONITOR, MODEL: B650 CAN FAIL TO SEND A PRINT REQUEST TO THE CIC (GE CENTRAL STATION) IF MORE THEN ONE PRINT COMMAND IS SENT AT A TIME. THIS WILL RESULT IN A LOSS OF REQUIRED ARRHYTHMIA WAVEFORM AND PATIENT VITAL SIGNS DOCUMENTATION.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, CARESCAPE (PER SITE REPORTER).======================THE PROJECT MANAGER/CLINICAL ENGINEER NOTIFIED GE MEDICAL OF THE PRINTER ISSUE. GE FIELD SERVICES DISPATCHED TECHNICAL SUPPORT TO THE HOSPITAL TO INVESTIGATE THE REPORTED COMPLAINT. THEY WERE ABLE TO REPRODUCE THE PROBLEM ON SITE AND IT IS A CONFIRMED PROBLEM WITH THIS GE MEDICAL MONITORING DEVICE. GE MEDICAL AND THE HOSPITAL HAVE SINCE REPLACED THE SIX (6) B650 PATIENT MONITORS WITH THE GE MEDICAL SOLAR 8000I MONITOR IN THE EMERGENCY DEPARTMENT. THESE MONITORS ARE WORKING CORRECTLY AND PRINTING TO THE CIC (GE CENTRAL STATION). THERE ARE NO OTHER GE B650 PATIENT MONITORS IN USE AT THE HOSPITAL AT THIS TIME. | 1
2540261 | THE CUSTOMER COMPLAINED ABOUT FALSE POSITIVE TEST RESULTS CAUSED BY THE LOW IONIC STRENGTH SOLUTION MLB 2 ON TANGO. THE CUSTOMER HAD NEITHER RETURNED THE SUPPOSEDLY DEFECTIVE PRODUCT NOR THE SAMPLES THAT HAD YIELDED ALLEGEDLY FALSE POSITIVE RESULTS. THEREFORE OUR QUALITY CONTROL LABORATORY TESTED THE RETAINED SAMPLE WITH 11 DIFFERENT CONTROLS AND SAMPLES. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DID NOT OBSERVE ANY FALSE POSITIVE REACTIONS. TESTING BY OUR QUALITY CONTROL LABORATORY CONFIRMED THE ALLEGEDLY DEFECTIVE LOT OF MLB2 FUNCTIONS CORRECTLY. A REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE NEGATIVELY AFFECTED THE QUALITY OF THE ALLEGEDLY DEFECTIVE LOT. A FIELD SERVICE ENGINEER FOUND ON THE TANGO A DEFECTIVE FLUSH PUMP RESPONSIBLE FOR CLEANING THE PROBE. THIS TAKES AN EFFECT OF FALSE POSITIVE RESULTS. THE FIELD SERVICE ENGINEER REPLACED THE DEFECTIVE FLUSH PUMP. CONTROLS (CUSTOMER QC) RUN AFTER THIS REPAIR YIELDED ONLY CORRECT RESULTS. AFTER THIS REPAIR, THE CUSTOMER HAS HAD NO MORE QUESTIONABLE RESULTS. THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT. | 0
2540531 | THE CONSUMER REPORTED THE CUFF OVERINFLATES WHICH CAUSED A BRUISE TO HIS ARM. KAZ USA HAS SENT A PRE-PAID LABEL TO RETRIEVE THE UNIT FOR EVALUATION. | 0
2540909 | ON (B)(6) 2011, A MALE PT WAS IMPLANTED WITH 65 RAPIDSTRAND RX SEEDS OF 12.58 MBQ (0.340 MCI ACTIVITY ON IMPLANT DATE) FOR BRACHYTHERAPY OF PROSTATE CANCER. PER (B)(4), THE SEEDS MAY HAVE BEEN MIXED WITH ANOTHER SEEDS ORDERS OF 16.06 MBQ (0.434 MCI) AND/OR 12.42 MBQ (0.336 MCI), RESULTING IN A POTENTIAL RADIATION OVER-DOSE OR UNDER-DOSE RESPECTIVELY (UNDER-DOSE NOT SIGNIFICANT PER MODEL). TO DATE, NO ADVERSE EXPERIENCE HAS BEEN NOTIFIED TO THE COMPANY FOR THIS PT. THE ORDERED SEED ACTIVITY AND TREATMENT PLAN FOR THIS PT WAS REVIEWED BY THE GE HEALTHCARE MEDICAL GROUP AGAINST THE SEED ACTIVITY IN THE ORDER RELEVANT TO THE POSSIBLE MIX-UP. WORST-CASE SCENARIO, 100% SWAP, WAS ASSUMED AND THE POTENTIAL OVER-DOSE OR UNDER-DOSE CALCULATED. GE HEALTHCARE HAS NOTIFIED THE PHYSICIAN WHO PERFORMED THE SEED IMPLANT IN WRITING AND BY TELEPHONE AND IS FOLLOWING UP FOR FURTHER INFO. EVAL SUMMARY: GE HEALTHCARE QUALITY ASSURANCE INVESTIGATION REPORT (B)(4)/ GE HEALTHCARE WAS MADE AWARE OF A POTENTIAL FOR INCORRECT LOADING OF RAPIDSTRAND RX NEEDLES DUE TO NON-COMPLIANCE WITH STANDARD OPERATING PROCEDURE ON THE PART OF ONE TECHNICIAN AT THE RADIOPHARMACY IN WHICH BRACHYTHERAPY IMPLANTS WERE PREPARED. FOUR INCIDENTS WERE REPORTED IN WHICH SEEDS HAD BEEN SWAPPED BETWEEN ORDERS RESULTING IN INCORRECTLY LOADED NEEDLES (MDR REPORT ID: 2915056-2011-00001, 2915056-2011-00002), OR INCORRECT CALIBRATION STRANDS BEING PROVIDED WITH THE ORDER (MDR REPORT ID: 2915056-2011-00003, 2915056-2011-00006), RESULTING IN MISLABELING OF THE DEVICE. INVESTIGATION OF THESE INCIDENTS AT THE LOADING FACILITY SHOWED THAT THE TECH HAD WORKED ON TWO ORDERS AT THE SAME TIME AND MIXED THE SEEDS BETWEEN THEM. GE HEALTHCARE REVIEWED ALL (B)(4) ORDERS THAT THE TECH COMPLETED DURING EMPLOYMENT AND IDENTIFIED (B)(4) WHERE THERE WAS THE POSSIBILITY FOR SEEDS BEING MIXED BETWEEN TWO OR MORE ORDERS. THIS WAS DETERMINED BY IT BEING POSSIBLE FOR THE TECH TO HAVE WORKED ON TWO OR MORE SAME-PRODUCT ORDERS SIMULTANEOUSLY WHERE SEEDS OF DIFFERENT ACTIVITY WERE INVOLVED. IT SHOULD BE NOTED THAT IT CANNOT BE DETERMINED THAT SEEDS WERE MIXED BETWEEN THE ORDERS, ONLY THAT THE POSSIBILITY EXISTS. MIXING SEEDS BETWEEN ORDERS IS PRESUMED TO AFFECT RADIATION EXPOSURE TO THE AREA SURROUNDING THE IMPLANT (PROSTATE GLAND, URETHRA, AND RECTUM), EITHER AN OVER-DOSE WHERE HIGHER ACTIVITY SEEDS WERE INTRODUCED INTO THE LOADED NEEDLES CARRYING INCREASED RISK OF RADIATION SIDE-EFFECTS, OR UNDER-DOSE WHERE LOWER ACTIVITY SEEDS WERE INTRODUCED INTO THE LOADED NEEDLES CARRYING RISK OF SUB-OPTIMAL TREATMENT OF THE TUMOR. THE ORDERED SEED ACTIVITY AND TREATMENT PLAN FOR EACH OF THE (B)(4) PTS WAS REVIEWED AGAINST THE SEED ACTIVITY IN THE ORDER RELEVANT TO THE POSSIBLE MIX-UP. WORST-CASE SCENARIO, 100% SWAP, WAS ASSUMED AND THE CONSEQUENT OVER-DOSE OR UNDER-DOSE CALCULATED. GE HEALTHCARE RISK ASSESSMENT CONCLUDED THAT, FROM A MEDICAL VIEWPOINT, THE RISK OF SERIOUS INJURY OR DEATH AS A RESULT OF A RECURRENCE OF SUCH MISLABELING IS REMOTE. | 0
2540912 | ON (B)(6)-2011 A MALE PT WAS IMPLANTED WITH 129 RAPIDSTRAND RX SEEDS OF 16.06 MBQ (0.434 MCI ACTIVITY ON IMPLANT DATE) FOR BRACHYTHERAPY OF PROSTATE CANCER. PER GEHC INVESTIGATION (B)(4), THE SEEDS MAY HAVE BEEN MIXED WITH ANOTHER SEEDS ORDERS, ONE OF 12.43 MBQ (0336 MCI) AND/OR ONE OF 12.58 MBQ (0.340 MCI), RESULTING IN A POTENTIAL RADIATION UNDER-DOSE. TO DATE, NO ADVERSE EXPERIENCE HAS BEEN NOTIFIED TO THE COMPANY FOR THIS PT. THE ORDERED SEED ACTIVITY AND TREATMENT PLAN FOR THIS PT WAS REVIEWED BY THE GE HEALTHCARE MEDICAL GROUP AGAINST THE SEED ACTIVITY IN THE ORDER RELEVANT TO THE POSSIBLE MIX-UP. WORST-CASE SCENARIO, 100% SWAP, WAS ASSUMED AND THE POTENTIAL OVER-DOSE OR UNDER-DOSE CALCULATED. GE HEALTHCARE HAS NOTIFIED THE PHYSICIAN WHO PERFORMED THE SEED IMPLANT IN WRITING AND BY TELEPHONE AND IS FOLLOWING UP FOR FURTHER INFO. EVAL SUMMARY: GE HEALTHCARE QUALITY ASSURANCE INVESTIGATION REPORT (B)(4)/ GE HEALTHCARE WAS MADE AWARE OF A POTENTIAL FOR INCORRECT LOADING OF RAPIDSTRAND RX NEEDLES DUE TO NON-COMPLIANCE WITH STANDARD OPERATING PROCEDURE ON THE PART OF ONE TECHNICIAN AT THE RADIOPHARMACY IN WHICH BRACHYTHERAPY IMPLANTS WERE PREPARED. FOUR INCIDENTS WERE REPORTED IN WHICH SEEDS HAD BEEN SWAPPED BETWEEN ORDERS RESULTING IN INCORRECTLY LOADED NEEDLES (MDR REPORT ID: 2915056-2011-00001, 2915056-2011-00002), OR INCORRECT CALIBRATION STRANDS BEING PROVIDED WITH THE ORDER (MDR REPORT ID: 2915056-2011-00003, 2915056-2011-00006), RESULTING IN MISLABELING OF THE DEVICE. INVESTIGATION OF THESE INCIDENTS AT THE LOADING FACILITY SHOWED THAT THE TECH HAD WORKED ON TWO ORDERS AT THE SAME TIME AND MIXED THE SEEDS BETWEEN THEM. GE HEALTHCARE REVIEWED ALL 173 ORDERS THAT THE TECH COMPLETED DURING EMPLOYMENT AND IDENTIFIED 23 WHERE THERE WAS THE POSSIBILITY FOR SEEDS BEING MIXED BETWEEN TWO OR MORE ORDERS. THIS WAS DETERMINED BY IT BEING POSSIBLE FOR THE TECH TO HAVE WORKED ON TWO OR MORE SAME-PRODUCT ORDERS SIMULTANEOUSLY WHERE SEEDS OF DIFFERENT ACTIVITY WERE INVOLVED. IT SHOULD BE NOTED THAT IT CANNOT BE DETERMINED THAT SEEDS WERE MIXED BETWEEN THE ORDERS, ONLY THAT THE POSSIBILITY EXISTS. MIXING SEEDS BETWEEN ORDERS IS PRESUMED TO AFFECT RADIATION EXPOSURE TO THE AREA SURROUNDING THE IMPLANT (PROSTATE GLAND, URETHRA, AND RECTUM), EITHER AN OVER-DOSE WHERE HIGHER ACTIVITY SEEDS WERE INTRODUCED INTO THE LOADED NEEDLES CARRYING INCREASED RISK OF RADIATION SIDE-EFFECTS, OR UNDER-DOSE WHERE LOWER ACTIVITY SEEDS WERE INTRODUCED INTO THE LOADED NEEDLES CARRYING RISK OF SUB-OPTIMAL TREATMENT OF THE TUMOR. THE ORDERED SEED ACTIVITY AND TREATMENT PLAN FOR EACH OF THE TWENTY-THREE PTS WAS REVIEWED AGAINST THE SEED ACTIVITY IN THE ORDER RELEVANT TO THE POSSIBLE MIX-UP. WORST-CASE SCENARIO, 100% SWAP, WAS ASSUMED AND THE CONSEQUENT OVER-DOSE OR UNDER-DOSE CALCULATED. GE HEALTHCARE RISK ASSESSMENT CONCLUDED THAT, FROM A MEDICAL VIEWPOINT, THE RISK OF SERIOUS INJURY OR DEATH AS A RESULT OF A RECURRENCE OF SUCH MISLABELING IS REMOTE. | 0
2541715 | MACHINE FREEZES WHILE DOING AN INTERNAL TRANSVAGINAL ULTRASOUND MAKING PATIENT FEEL UNCOMFORTABLE. HAVE TO RESTART MACHINE AND WAIT WHILE PROBE IS INSERTED IN PATIENT AND PATIENT IS IN A GYNECOLOGIC POSITION WHICH IS DIFFICULT FOR SOME PATIENTS TO MAINTAIN.CASE MANAGER ACTIVELY WORKING WITH GE AND BIOMEDICAL ENGINEERING TO RECTIFY THE PROBLEM. GE ORDERED SEVERAL NEW PARTS TO BE INSTALLED. | 1
2541857 | AS REPORTED BY THE HOME CARE DEALER, THE PARENTS ENTERED THE BABY'S ROOM AND SAW THE ALARM LIGHTS ON AND SUBSEQUENTLY INFORMED THE APNEA PROGRAM PEOPLE THAT THE MONITOR DID NOT ALARM. THERE WAS NO IMPACT TO THE BABY REPORTED. THE MONITOR WAS GIVEN TO THE HOME CARE COMPANY. AFTER THE MONITOR WAS EVALUATED BY THE HOME CARE DEALER IT WAS RETURNED TO CAS WHERE IT WAS FURTHER RUN FOR OVER 1 WEEK ALL THE WHILE CHALLENGED WITH INPUT DATA DESIGNED TO CAUSE ALARM EVENTS. THE MONITOR WORKED AS IT SHOULD. | 1
2543104 | THE CUSTOMER REPORTED THAT AFTER A DOUBLE PLATELET PRODUCT PROCEDURE AND THE DONOR HAD LEFT THE CENTER, THAT THERE WERE LARGE CLOTS IN THE PLATELET BAGS. THEY TRACED THE TUBING BACK AND NOTICED THAT THE SALINE BAG WAS HOOKED UP RATHER THAN THE ACDA BAG TO THE ANTICOAGULANT SYSTEM. THE DONOR LEFT THE CENTER FEELING FINE. UPON FOLLOW UP WITH THE DONOR, HE SAID HE FELT GOOD, JUST LIKE HE ALWAYS DOES AFTER HE DONATES. NO MEDICAL INTERVENTION WAS REQUIRED FOR THIS EVENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN AND INVESTIGATION. THIS REPORT IS BEING FILED DUE TO AN OPERATOR ERROR THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. INVESTIGATION: THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. BASED ON THE REPORT FROM THE CUSTOMER, THE OPERATOR USED NORMAL SALINE INSTEAD OF ACD-A DURING THE PROCEDURE. ROOT CAUSE: OPERATOR ERROR. | 0
2543111 | THE CUSTOMER REPORTED AIR IN THE SAMPLE POUCH. PT INFORMATION IS UNAVAILABLE AT THIS TIME. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR DEATH OR INJURY OCCURRED. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2544934 | THREE INSTANCES OF GLARE SHIELDS COMING DETACHED FROM THE C-VIEW SYSTEM DISPLAY BEZEL HAVE OCCURRED. THE FAILURES OCCURRED AT THE FOLLOWING SITES: (B)(6). (B)(6) COMPLAINED OF REPORT THAT THE SHIELD ¿FELL¿ OFF THE C-VIEW MONITOR THIS MORNING. IT ¿BRUSHED¿ THE PATIENT ON THE WAY DOWN AND SMASHED ON THE FLOOR. NO ONE WAS HURT. (B)(6) MEDICAL CENTER COMPLAINED THAT ¿SHIELD FELL OFF THE DISPLAY. IT WAS STATED THAT THE CUSTOMER WAS MOVING THE DISPLAY BY THE SHIELD VS. THE HANDLES OR THE BEZEL.¿ (B)(6) UNIVERSITY HOSPITAL COMPLAINED THAT THE ¿SHIELD FELL OFF THE DISPLAY. A LARGE CRACK WAS CREATED DOWN THE MIDDLE AND THE BINDING CORNERS ARE ALSO CRACKED.¿ INVESTIGATION BY CARROT MEDICAL CONFIRMED THAT THE METHOD OF ATTACHING THE GLARE SHIELDS TO THE BEZEL WITH ADHESIVE IS NOT SUFFICIENT. USERS HAVE INCORRECTLY BEEN PULLING ON THE GLARE SHIELDS TO POSITION THE DISPLAY RATHER THAN THE HANDLES PROVIDED ON THE DISPLAY ITSELF. MOST GLARE SHIELDS IN THE FIELD HAVE BEEN SECURED MECHANICALLY. THE SHIELDS THAT FAILED WERE NOT SECURED MECHANICALLY, BUT RATHER ADHESIVELY. ALL THOSE SHIELDS SECURED WITH ADHESIVE ARE BEING RETROFITTED WITH A MECHANICAL ATTACHMENT METHOD. ADHESIVE IS NO LONGER USED ON NEW SYSTEMS. | 0
2546827 | IT WAS REPORTED THAT THE 3SP ULTRASOUND PROBE WAS GETTING HOT DURING EXAMS WITH THE LOGIQ P6 CONSOLE. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS ISSUE. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. THIS MALFUNCTION WAS DETERMINED TO BE REPORTABLE AS THIS SAME MALFUNCTION HAS PREVIOUSLY CONTRIBUTED TO A SERIOUS INJURY WITHIN THE LAST TWO YEARS. REFERENCE MDR 9710090-2011-00001. AN URGENT MEDICAL DEVICE CORRECTION LETTER WAS SENT TO CUSTOMERS WITH AFFECTED GE LOGIQ P6 ULTRASOUND UNITS WILL BE UPDATED IN THE FIELD. (B)(4). | 0
2548640 | DURING A REVIEW OF PROGRAMMING HISTORY FOR THIS PATIENT, AN ANOMALY WAS OBSERVED. IT WAS IDENTIFIED THAT THERE WERE GENERATOR AND SYSTEM DIAGNOSTICS TESTS STORED IN A PULSE GENERATOR'S RECORD THAT OCCURRED AT THE SAME TIME. PRELIMINARY REVIEW OF THE DATABASE IDENTIFIED THAT THE DIAGNOSTIC TEST TYPE HAD CHANGED FROM A PREVIOUS DOWNLOAD OF A HANDHELD COMPUTER'S DATABASE. THE CAUSE FOR THE CORRUPTED DIAGNOSTIC DATA IN THE VNS PROGRAMMING HISTORY DATABASE IS ASSOCIATED WITH INTERRUPTION OF THE MIGRATION OF THE VNS SOFTWARE DATABASE DURING UPGRADE FROM V6.1 TO A LATER VERSION OF VNS SOFTWARE. IN ADDITION TO THE CORRUPTED DIAGNOSTIC DATA, IF THE CORRUPTION OCCURS, SQL ERRORS WILL OCCUR WHEN COMMUNICATING WITH A DEMIPULSE GENERATOR AND EMPTY FILES WILL BE GENERATED WHEN ATTEMPTING TO EXPORT THE DATABASE AS TEXT.><THE UPGRADE FROM VERSION 6.1 SOFTWARE TO VERSION 7.1 SOFTWARE OCCURRED ON (B)(4) 2007. REVIEW OF PROGRAMMING HISTORY.><THE EVENT DATE WAS INCORRECTLY REPORTED ON THE INITIALLY MDR. THE CORRECT DATE IS (B)(6) 2007. | 0
2550396 | THE CUSTOMER NOTED THE DISPENSE PROBE LEAKED CLEAR FLUID WHILE TROUBLESHOOTING SHUTTLE POSITION SENSOR ERROR MESSAGE ON THE BECKMAN COULTER LH 750 SLIDEMAKER INSTRUMENT. THE LEAK WAS CONTAINED WITHIN THE INSTRUMENT. THE DISPENSE PROBE CARRIES BLOOD, DILUENT AND CLENZ REAGENT. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, EYE PROTECTION, AND GLOVES. THE CUSTOMER DID NOT COME IN CONTACT WITH THE FLUID AND DID NOT SEEK MEDICAL ATTENTION. NO EXPOSURE (SPRAYED OR SPLASHED) TO MUCOUS MEMBRANES OR OPEN WOUNDS WERE REPORTED. NO DEATH, INJURY OR CHANGES TO PATIENT TREATMENT OR RESULTS ARE ASSOCIATED WITH THIS EVENT. THE CUSTOMER REVIEWED THE MSDS. THE FACILITY DOES HAVE A RISK MANAGEMENT / EXPOSURE CONTROL PLAN IS IN PLACE. ON THE DAY OF THE EVENT, A FIELD SERVICE ENGINEER WAS DISPATCHED. THE FSE INSTRUCTED THE CUSTOMER TO REMOVE THE BELT FROM THE SHUTTLE POSITION SENSOR. THE FSE FOUND THE PROBE LEAKING DILUENT AND THEN REROUTED THE SHUTTLE BELT, WHICH RESOLVED THE LEAK. THE FAILURE MODE OF THIS EVENT WAS ATTRIBUTED TO THE ROUTING OF THE SHUTTLE BELT. (B)(4). | 0
2550544 | THE CUSTOMER REPORTED A CLEAR FLUID LEAK WITH A HINT OF RED COLOR FROM THE LEFT SIDE OF THE BECKMAN COULTER - COULTER LH 750 SLIDEMAKER. THE LEAK WAS NOT CONTAINED WITHIN THE INSTRUMENT. THE CUSTOMER STATED THE COLLECTING TRAYS WERE FULL OF FLUID, THE APPROXIMATE VOLUME OF THE LEAK WAS UNKNOWN. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES. THE CUSTOMER DID NOT COME IN CONTACT WITH THE FLUID AND DID NOT SEEK MEDICAL ATTENTION. NO EXPOSURE (SPRAYED OR SPLASHED) TO MUCOUS MEMBRANES OR OPEN WOUNDS WERE REPORTED. NO DEATH, INJURY OR CHANGES TO PATIENT TREATMENT OR RESULTS ARE ASSOCIATED WITH THIS EVENT. THE CUSTOMER DID NOT REVIEW THE MSDS. THE FACILITY DOES HAVE A RISK MANAGEMENT / EXPOSURE CONTROL PLAN IS IN PLACE. FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO EVALUATE THE EVENT. THE FSE FOUND A LEAK AT LV20 IN THE SLIDEMAKER. VL20/LV20 CARRIES PRESSURIZED DILUENT TO RINSE THE RESERVOIRS AND DISPENSE LINES DURING NORMAL OPERATION. THE FSE REPLACED THE SILICONE TUBING AND RAN SEVERAL SAMPLES MAKING SLIDES. THE FSE DID NOT DETECT ANY ADDITIONAL LEAKS. THE CAUSE OF THE LEAK IS A HOLE IN THE TUBING AT VL20. THE LEAK WAS DILUENT LEAK ONLY AND DID NOT CONTAIN ANY HAZARDOUS FLUIDS. THE SOURCE OF THE RED COLOR IN THE FLUID IS UNKNOWN. (B)(4). | 0
2551019 | A CUSTOMER CONTACTED BECKMAN COULTER INC (BEC) STATING THAT BLOOD WAS OBSERVED ON THE INSIDE OF THE FRONT DOOR OF THEIR COULTER ACT DIFF 2 ANALYZER. THE CUSTOMER WAS RUNNING CONTROLS AND GOT WBC APERTURES ALERTS ON THE ACT DIFF 2. BEC CTS RECOMMENDED USE OF PERSONAL PROTECTIVE EQUIPMENT (PPE) BEFORE TROUBLESHOOTING AND ASKED CUSTOMER TO OPEN THE FRONT DOOR AND REMOVE FARADAY SHIELD. THE CUSTOMER THEN OBSERVED 2.5 TO 5MLS OF BLOOD ON THE INSIDE OF THE FRONT DOOR OF THE INSTRUMENT. THE LEAK WAS CONTAINED WITHIN THE INSTRUMENT. THERE IS AN EXPOSURE PLAN IN PLACE AT THE FACILITY. THE CUSTOMER WAS WEARING GLOVES AND A LAB COAT AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED AND THERE WAS NO AFFECT TO PATIENT SAMPLES. BEC CUSTOMER TECHNICAL SUPPORT (CTS) SUSPECTED THE BLOOD CAME FROM THE PROBE AND GUIDED THE CUSTOMER OVER THE PHONE IN CLEANING THE PROBE WIPE BLOCK. THE OPERATOR WAS INSTRUCTED TO CLEAN THE BLOOD WITH BLEACH ON A PAPER TOWEL AND RINSE WITH DISTILLED WATER. THE OPERATOR WAS ALSO TOLD TO REMOVE AND CLEAN THE PROBE WIPE BLOCK WITH BLEACH AND DI (DEIONIZED WATER) AND BLEACH THE WBC COUNT LINE AND VIC (VACUUM ISOLATOR CHAMBER), THEN RUN 3 RINSE MIXES AND A STARTUP. THE LEAK WAS RESOLVED; STARTUP AND CONTROLS PASSED. THE CAUSE OF THE LEAK IS LIKELY ATTRIBUTED TO THE PROBE WIPE BLOCK BEING CLOGGED DUE TO BUILD UP OF SAMPLE. THE PROBE IS THE PART OF THE INSTRUMENT THAT ASPIRATES WHOLE BLOOD, WHICH IS WHAT WAS FOUND ON THE INSIDE OF THE DOOR; THE PROBE WIPE BLOCK MAY HAVE BEEN DIRTY/POTENTIALLY CLOGGED DUE TO BUILD UP OF SAMPLE. (B)(4). | 0
2552131 | THE CUSTOMER REPORTED THAT THERE WAS NO AUDIO FROM THE LOUDSPEAKER ON THE DASH 2500 PATIENT MONITOR. THE COMPLAINT INDICATES THAT THE DEVICE FAILED THE AUDIO TEST DUE TO DC2DC BOARD FAILURE. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2553017 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY NOTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED ON (B)(6) 2007 WHICH CHANGED THE VNS SETTINGS, AS THE NEXT INTERROGATION ON (B)(6) 2007 SHOWED SETTINGS OF 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF. THE SETTINGS WERE CORRECTED AT THIS VISIT. A FINAL INTERROGATION TO VERIFY SETTINGS WAS NOT PERFORMED ON (B)(6) 2007. ANALYSIS OF PROGRAMMING HISTORY. | 0
2554748 | IT WAS REPORTED THAT A PHYSICIAN'S DELL X50 HANDHELD COMPUTER WAS FROZEN. WHEN IT WAS FIRST TURNED ON, IT SHOWED A BLANK SCREEN WHICH WAS UNRESPONSIVE. THE COMPANY REPRESENTATIVE TRIED THREE HARD RESETS, AND THE SCREEN WAS STILL AT THE "CLEAR ALL DATA" SCREEN. SHE TRIED PRESSING THE BUTTONS THAT WERE INDICATED ON THE SCREEN, BUT THE COMPUTER WAS UNRESPONSIVE. THE COMPUTER HAD PREVIOUSLY BEEN CHARGING, AND THE REPRESENTATIVE HAD JUST UNPLUGGED IT, SO THE BATTERY WAS RULED OUT AS A POTENTIAL CAUSE OF THE EVENT. A REPLACEMENT COMPUTER WAS SENT TO THE PHYSICIAN, AND THE SUSPECT COMPUTER AND SOFTWARE WERE RECEIVED BY THE MANUFACTURER ON (B)(4) 2012. HOWEVER, PRODUCT ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><PRODUCT ANALYSIS FOR THE RETURNED HANDHELD COMPUTER AND SOFTWARE WAS COMPLETED ON (B)(6) 2012. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD, AND THE REPORTED ALLEGATION OF A FROZEN SCREEN WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD, AND THE REPORTED ALLEGATION WAS ALSO NOT VERIFIED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE LOCK BUTTON WAS IN THE LOCKED POSITION. ONCE THE LOCK BUTTON WAS MOVED TO THE UNLOCKED POSITION, THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2555538 | IT WAS REPORTED THAT BEFORE AN A-FIB PROCEDURE STARTED, THE CARTO 3 SYSTEM IN USE WOULD NOT ESTABLISH COMMUNICATION WITH THE PATIENT INTERFACE UNIT. ERROR 1 APPEARED. SOME TROUBLE SHOOTING WAS ATTEMPTED. HOWEVER THE ISSUE WAS NOT RESOLVED OVER THE PHONE. THE ISSUE RESULTED IN THE CASE BEING CANCELLED. THE PATIENT WAS UNDER LOCAL ANESTHESIA. NO TRANSSEPTAL PUNCTURE HAD BEEN PERFORMED PRIOR TO THE CASE CANCELLATION. HOWEVER, SINCE THE PATIENT WAS ALREADY UNDER CONSCIOUS SEDATION; THE PATIENT REQUIRED AN EXTRA DAY STAY AT THE HOSPITAL. PATIENT'S AGE, GENDER, AND CHRONIC HEALTH CONDITION OR MAJOR MEDICAL HISTORY INFORMATION WAS NOT PROVIDED. THE DOES NOT PHYSICIAN THINK THERE WAS ANY POTENTIAL RISK TO THE PATIENT DUE TO THIS CANCELLATION. (B)(4). IT WAS REPORTED THAT BEFORE AN A-FIB PROCEDURE STARTED, THE CARTO 3 SYSTEM IN USE WOULD NOT ESTABLISH COMMUNICATION WITH THE PATIENT INTERFACE UNIT. ERROR 1 APPEARED. THE ISSUE RESULTED IN THE CASE BEING CANCELLED. THE PATIENT WAS UNDER LOCAL ANESTHESIA. NO TRANSSEPTAL PUNCTURE HAD BEEN PERFORMED PRIOR TO THE CASE CANCELLATION. HOWEVER, SINCE THE PATIENT WAS ALREADY UNDER CONSCIOUS SEDATION; THE PATIENT REQUIRED AN EXTRA DAY STAY AT THE HOSPITAL. THE CARTO 3 SYSTEM ASSOCIATED WITH THE REPORTED COMPLAINT WAS INSPECTED BY BWI FIELD SERVICE ENGINEER (FSE). THE MAIN MODULE WAS REPLACED BY THE FSE. THE ISSUE WAS RESOLVED. SYSTEM WAS FULLY FUNCTIONAL AND READY FOR USE. THE DEFECTIVE MAIN MODULE WAS SENT TO THE MANUFACTURING SITE, WHERE THEY TESTED THE MAIN MODULE/CARD WITH A WORKING SYSTEM. THE PATIENT INTERFACE UNIT (PIU) DID NOT BOOT UP AND PROBLEM WITH COMMUNICATION WAS DETECTED. DSP CARDS MASTER AND SLAVE WERE FOUND FAULTY. DSP MASTER AND SLAVE WERE REPLACED AND AS A RESULT PROBLEM WAS RESOLVED, THE PIU BOOTED UP PROPERLY. THE DEVICE HISTORY REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THIS DEVICE.><INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. (B)(4). | 1
2557733 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHILE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE ANALYSIS OF THE RUN DATA FILE (RDF) DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFO, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATIONS, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2557857 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE WAS ANALYZED. THE ANALYSIS DID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED WBC CONTENT REPORTED FOR THIS COLLECTION. SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IT THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THE ANALYSIS OF THE RDFDID NOT FIND A CONCLUSIVE CAUSE OF THE ELEVATED WBC CONTENT REPORTED FOR THIS COLLECTION. THE POSSIBLE ROOT CAUSE WAS PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2558122 | IT WAS REPORTED THAT THE TRANSESOPHAGEAL (TEE) PROBE CONTROL HANDLE ALLOWS PROPER FORWARD DEFLECTION TO THE 120 DEGREE LIMIT BUT THEN CAN EASILY BE BENT PAST THIS LIMIT MANUALLY. WITH THE PROBE IN THE PT IT WAS MANUALLY DEFLECTED PAST THE LIMIT MAKING IT DIFFICULT TO WITHDRAW. NO INJURY WAS REPORTED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
2559808 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE (RDF) WAS ANALYZED. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED. | 0
2560369 | ON (B)(6) 2010, CUSTOMER ADVISED US THAT SHE HAD BEEN REPORTING RESULTS WHEN NOT FOLLOWING TOSOH'S PROCEDURE FOR MAKING UP DILUENT FOR PREVIOUS MONTHS. THE CUSTOMER IGNORED ERROR MESSAGES FROM THE INSTRUMENT DURING THIS TIME. WE REQUESTED SHE VERIFY THE PROCEDURE FOR RECONSTITUTING DILUENT, ON (B)(6), SHE INFORMED A TOSOH BIOSCIENCE SVC ENGINEER THAT BY FOLLOWING OUR INSTRUCTIONS THE ERROR WAS CLEARED UP. SAMPLE DILUENT CONCENTRATION CAN AFFECT THE RESULTS. GENERATED BY THE AUTOMATED IMMUNOASSAY SYS. A REVIEW OF THE CUSTOMER'S RECENT HISTORY INDICATES A LACK OF FAMILIARITY BY THE END USER WITH BASIC INSTRUMENT FUNCTION, MAINTENANCE AND LAB PRACTICES. THIS IS OF CONCERN TO TOSOH BIOSCIENCE. FAX AND LETTER SENT ON (B)(4) TO CUSTOMER ADVISING OF END USER OF TRAINING WORKSHOP OPTIONS, TRAINING MATERIALS, OPERATOR USER'S GUIDES AND PERIODIC MAINTENANCE OPTIONS AVAILABLE. THIS WAS ADDRESSED TO THE (B)(6) AND THE PRACTICING PHYSICIAN. | 0
2560672 | EVENT CLARIFICATION: DURING THE SET-UP OF THE SET, THE CUSTOMER FOUND THAT THE PLATELET TUBE WAS CRUSHED. THIS ISSUE WAS FOUND DURING SETUP, SO NO DONOR WAS REASONABLY KNOWN AT THE TIME.><DURING SET-UP; CRUSHING OF PLATELET TUBE CHECK DURING THE ASSEMBLY. THIS REPORT HAS BEEN SENT TO THE ITALIAN NATIONAL MINISTRY OF HEALTH. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. INVESTIGATION: THE SET WAS RETURNED FOR EVALUATION. THE SET DID NOT APPEAR TO HAVE BEEN USED. A 1 CM FLAT AREA WAS NOTED ON THE TUBING BETWEEN THE PLATELET BAG AND THE SAMPLER. THIS WAS LOCATED APPROXIMATELY 20 CM FROM THE BAG BOND SOCKET. A SECOND KINK WAS NOTED ON THE PLASMA BAG LINE. THIS WAS APPROXIMATELY 0.5 CM LONG AND LOCATED APPROXIMATELY 25 CM FROM THE BAG BOND SOCKET. THE EFFECT OF A KINK ON POTENTIAL PERFORMANCE ISSUES WITH THE KITS HAS BEEN EVALUATED BY ENGINEERING. A TUBING KINK IS UNLIKELY TO AFFECT THE KIT PERFORMANCE AS LONG AS THE INTERIOR DIAMETER OF THE TUBING IS LARGER THAN THE INTERIOR DIAMETER OF THE NEEDLE. THE KINKS OBSERVED HAD A FLATTENED APPEARANCE, BUT WOULD OPEN READILY IF FLUID IS PASSED THROUGH THE TUBING. ROOT CAUSE: THE COMMON CAUSES OF TUBING KINKS, INDENTATIONS AND COMPRESSIONS ARE TYPICALLY RELATED TO SLIGHT VARIATIONS DURING TRAY PACKING. TUBING IS NOT AFFIXED WITHIN THE PACKAGING AND SHIFTING MAY OCCUR DURING PACKAGING AND TRANSPORT TO STERILIZATION. FOLLOWING EXPOSURE TO THE HEAT AND HUMIDITY OF THE STERILIZATION CYCLE, THE TUBING MAY ASSUME A DEFORMED SHAPE. ONLY THE MOST SEVERELY OCCLUDED KINKS IN THE TUBING SHOULD BE CAUSE FOR REJECTION TO PREVENT USAGE OF THESE DISPOSABLES IN PROCEDURES.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2561904 | IT WAS REPORTED THAT DURING THE PVI ABLATION IN THE LEFT ATRIUM, THE SOFTWARE OF THE CARTO3 WAS FROZEN AND SEVERAL REBOOTS AND USING THE DEFAULT TEMPLATE DID NOT HELP. THE PHYSICIAN DECIDED TO CANCEL AND RESCHEDULE THE CASE. SINCE THE TRANSSEPTAL PUNCTURE WAS DONE AND THE PATIENT WAS UNDER LOCAL ANESTHESIA, THE PHYSICIAN CONSIDERED THE CANCELLATION OF THE PROCEDURE MIGHT POSE A POTENTIAL RISK TO THIS PATIENT. (B)(4).><(B)(4). BIOSENSE FIELD SERVICE ENGINEER ARRIVED AT THE HOSPITAL AND REPLACED THE PROBLEMATICAL WORK STATION WITH ANOTHER ONE AND THE SYSTEM READY FOR USE SINCE THEN. THE SUSPICIOUS WS WAS TESTED BY TECHNICAL SERVICE REPAIR CENTER AND NO PROBLEM WAS FOUND. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE FOUND REGARDING THIS ISSUE. | 1
2563864 | THE CUSTOMER REPORTED A 1-2 ML OF BLOOD LEAK FROM THE SAMPLE PROBE WHICH WAS CONTAINED WITHIN THE LH750 SLIDEMAKER. THE CUSTOMER WAS WEARING GLOVES, GOGGLES AND A LAB COAT AND NO EXPOSURE OR INJURY WAS REPORTED. PATIENT RESULTS WERE NOT AFFECTED AND THERE WAS NO CHANGE TO PATIENT TREATMENT IN CONNECTION WITH THIS EVENT. BECKMAN COULTER FIELD SERVICE ENGINEER (FSE) DISCOVERED A PROBLEM WITH THE SAMPLE DELIVERY PROBE TUBING. THE FSE REPLACED THE TUBING AND VERIFIED INSTRUMENT PER ESTABLISHED PROCEDURES. (B)(4). | 0
2565106 | THE DEVICE FAILED TO ATTACH TO THE ESOPHAGUS. LATER, THE PRODUCT WAS RECALLED AFTER DATE OF EVENT. | 0
2565114 | THE DEVICE FAILED TO ATTACH TO THE ESOPHAGUS. LATER THE PRODUCT WAS RECALLED AFTER DATE OF EVENT. | 0
2566568 | IT WAS REPORTED THAT A PHYSICIAN WAS IN A PATIENT'S ROOM WHEN THE STAFF OBSERVED THE PATIENT'S SPO2 LEVEL HAD DROPPED AND THE AUDIBLE ALARM ON THE SOLAR 8000M MONITOR WAS NOT SOUNDED. THE HOSPITAL COULD NOT CONFIRM IF A VISUAL ALARM WAS PROVIDED. THE PATIENT WAS REPORTED TO BE RE-ANIMATED AND SURVIVED THE EVENT. THE PHYSICIAN WAS IN THE PATIENT'S ROOM AND THERE WAS NO DELAY IN TREATMENT. THERE WERE NO CLINICAL INFORMATION CENTER (CIC) LOG FILES OR BEDSIDE MONITOR PRINTOUTS TO CONFIRM WHETHER THE SOLAR 8000M ALARMED FOR THE EVENT. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. THE OCCUPATION OF THE INITIAL REPORTER IS UNKNOWN. | 1
2568037 | ON (B)(4) 2012, A CUSTOMER REPORTED THAT TWO RADIOLOGY RESIDENTS WERE ABLE TO SAVE A PRELIMINARY REPORT AS A FINAL REPORT ALTHOUGH THEY DID NOT HAVE AUTHORIZED PERMISSIONS TO DO SO. OVER THE COURSE OF A MONTH, INVESTIGATION OF THE PROBLEM CONTINUED BUT NO OTHER OCCURRENCES WERE IDENTIFIED. HOWEVER, MONITORING CONTINUED. ON (B)(6) 2012, THE CUSTOMER REPORTED THAT THE PROBLEM RECURRED. THERE WAS NO HARM TO PATIENTS REPORTED IN BOTH COMPLAINTS. MCKESSON MEDICAL IMAGING COMPANY HAS DETERMINED THAT THE ROOT CAUSE OF THE PROBLEM IS THAT SOME RADIOLOGY RESIDENTS AT THE JEWISH GENERAL HOSP INAPPROPRIATELY INVOKED THE "SAVE FINAL" BUTTON INSTEAD OF THE "SAVE PRELIMINARY" BUTTON. BOTH THE "SAVE FINAL" AND "SAVE PRELIMINARY" BUTTONS WERE ENABLED FOR THE RADIOLOGY RESIDENTS' ROLE DUE TO A SOFTWARE DEFECT, WHEN ONLY THE "SAVE PRELIMINARY" BUTTON SHOULD HAVE BEEN ENABLED. CONSISTENT WITH THE COMMON PRACTICE AT TEACHING HOSPITALS AND THE WORKFLOW PRESENTED IN THE DEVICE'S USER GUIDES, THE "SAVE FINAL" BUTTON IS INTENDED TO BE USED BY ACCREDITED RADIOLOGISTS WHO OVERREAD THE PRELIMINARY REPORTS CREATED BY THE RADIOLOGY RESIDENTS. IF A RADIOLOGY RESIDENT MAKES A GROSS MISDIAGNOSIS IN THEIR PRELIMINARY REPORT AND CLICKS "SAVE FINAL", IT WILL NOT BE OVERREAD BY AN ACCREDITED RADIOLOGIST AND COULD LEAD TO SIGNIFICANT HARM TO THE PT. MCKESSON MEDICAL IMAGING COMPANY CUSTOMER SUPPORT WILL WORK WITH POTENTIALLY AFFECTED CUSTOMERS TO APPLY A TECHNICAL SOLUTION TO ADDRESS THIS ISSUE. | 1
2570919 | PATIENT WAS HAVING AN MRI UNDER SEDATION AND REQUIRED ECG MONITORING. AFTER THE PROCEDURE WAS COMPLETED, MRI TECHNICIAN WAS REMOVING THE ELECTRODES FROM THE PATIENT'S CHEST. PATIENT WAS NOTED TO HAVE 2ND DEGREE BURNS AND ONE PROBABLE 3RD DEGREE BURN UNDERNEATH THE ELECTRODES WITH REDNESS IN SURROUNDING AREAS.======================MANUFACTURER RESPONSE FOR MRI SCANNER, SIGNA HDXT (PER SITE REPORTER).======================FIELD MAINTENANCE AND TROBLESHOOTING OF THE OF THE MRI. QUESTIONNAIRE COMPLETED. QUESTIONNAIRE HAS BEEN COMPLETED AND RETURNED TO THE MANUFACTURER OF THE VERIS MONITOR. FIELD REPRESENTATIVE ON SITE FOR EDUCATION.MEDI-TRACE (COVIDIEN) FIELD REPORT AND SUBMITTED SAMPLES OF LOT NUMBER PACK OF ELECTRODES.======================MANUFACTURER RESPONSE FOR ELECTRODE, ECG, MRI, MED-TRACE (PER SITE REPORTER).======================SAMPLES OF ELECTRODES FROM THE SAME LOT NUMBER WERE RETURNED TO VENDOR FOR INSPECTION/ANALYSIS. NO REPORT YET AVAILABLE. | 0
2570921 | PT WAS RECEIVING MRI OF SHOULDER. SHOULDER COIL WAS IN USE. PATIENT WAS UNABLE TO BE STILL FOR DURATION OF PROCEDURE. PROCEDURE TIME WAS LENGTHENED AS A RESULT. PT BEGAN TO COMPLAIN OF A BURNING SENSATION IN SHOULDER AREA. PROCEDURE WAS TERMINATED AT THAT TIME.ONSITE EVALUATION OF EQUIPMENT BY GE - GE REPORTS NO PROBLEMS WHEN DOING TESTING AND CALIBRATIONS. UNABLE TO REPRODUCE EFFECTS. ON SITE VISIT BY APPLICATIONS PERSON HELPFUL IN IDENTIFYING POSSIBLE RELATED CAUSES. | 0
2571571 | IT WAS REPORTED THAT THE CARESCAPE TELEMETRY SERVER (CTS) WAS EXPERIENCING RANDOM SIGNAL DROPOUT LASTING UP TO 7 MINUTES IN DURATION AFFECTING 4 PATIENTS. THE GE ANTENNA, A COMPONENT OF THE CARESCAPE TELEMETRY SYSTEM, WAS IDENTIFIED AS THE CAUSE OF THE DROPOUT. THERE WAS NO SERIOUS INJURY OR DEATH ASSOCIATED WITH THIS EVENT, NOR WAS MEDICAL INTERVENTION REQUIRED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2571789 | PATIENT HAD OUTPATIENT MRI AND WAS DISCHARGED. PATIENT RETURNED AS AN INPATIENT 7 DAYS LATER. BURNS WERE REPORTED BY THE PATIENT AT THAT TIME. BURNS WERE ON CHEST IN AREAS WHERE ELECTRODES WERE APPLIED.======================MANUFACTURER RESPONSE FOR MRI SCANNER, SIGNA HDXT (PER SITE REPORTER).======================ON SITE EVALUATION OF EQUIPMENT BY GE. GE REPORTS NO PROBLEMS WHEN DOING TEST AND CALIBRATIONS.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, MRI, VERIS (PER SITE REPORTER).======================AN INVESTIGATION QUESTIONNAIRE WAS COMPLETED. FOLLOW UP RECOMMENDATIONS WERE RECEIVED.======================MANUFACTURER RESPONSE FOR ELECTRODES, ECG, MRI (PER SITE REPORTER).======================MANUFACTURER REQUESTED SAMPLE PRODUCT FROM THE SAME LOT NUMBERS TO BE RETURNED TO THEM FOR EVALUATION. | 0
2576669 | IT WAS REPORTED THAT AFTER A SHORT TIME OF USE, THE DEVICE STARTED TO SMOKE. DIRECTLY AFTER SWITCHING ON, SMOKE WAS OBSERVED AND THE USER SWITCHED OFF THE DEVICE IMMEDIATELY AND FORWARDED IT TO REPAIR. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATION THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 0
2577234 | THE REPORTED HANDHELD AND WAND WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. ANALYSIS OF THE RETURNED HANDHELD REVEALED THAT NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE SOFTWARE INDICATED THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED BY A COMPANY REPRESENTATIVE THAT HER HANDHELD COMPUTER WOULD NOT WORK AS IT WOULD SOMETIMES FREEZE. MOREOVER, THE AREA REPRESENTATIVE INDICATED THE ISSUE WAS WITH HER HANDHELD AS THE SCREEN WOULD TURN BLACK AND DID NOT REMAIN POWERED ON AS EXPECTED. AT THE MOMENT IT IS UNKNOWN IF THERE WAS ANY TROUBLESHOOTING PERFORMED. THE REPORTED HANDHELD REMAINS IN TRANSIT TO BE RETURNED TO THE MANUFACTURER FOR ANALYSIS. . | 0
2578249 | IT WAS REPORTED THAT MULTIPLE PTS EXPERIENCED SIGNAL DROP OUT, OR LOSS OF TELEMETRY MONITORING, AT THE CLINICAL INFORMATION CENTER (CIC). ALTERNATE PT MONITORING WAS NOT AVAILABLE DURING THIS TIME. NO SERIOUS INJURY OR DEATH WAS ASSOCIATED WITH THIS REPORT, AND MEDICAL INTERVENTION WAS NOT REQUIRED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 1
2578383 | DURING ERCP, FLUORO EQUIPMENT WENT DOWN. MESSAGE ON SCREEN READ ROTOR ERROR, WE WERE UNABLE TO COMPLETE PROCEDURE. | 0
2578996 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR ID# (B)(6) - THE WBC COUNT IS UNAVAILABLE, AS PER THE CUSTOMER, IT WAS TOO NUMEROUS TO COUNT. THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE RUN DATA FILE (RDF) WAS ANALYZED. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A F/U REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE POSSIBLE ROOT CAUSE WAS PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2580006 | THESE HAPPENED OVER THE COURSE OF A MONTH:1. WHILE DOCUMENTING MEDICATION INTO WITT LOG, THE MEDICATIONS DISAPPEARED AND HAD TO BE RE-ENTERED. THIS COULD CAUSE DOUBLE DOSING.2. HEMODYNAMIC MEASUREMENTS FREEZING ON SCREEN WHILE SNAPPING, THEN SCREEN INFO MOVES VERY QUICKLY ACROSS.3. UNABLE TO DOCUMENT MEDICATION INTO WITT.4. WHILE SWITCHING FROM SCREEN TO SCREEN IN THE WITT, TABS AT TOP OF SCREEN DISAPPEARED.5. HEMODYNAMIC NUMBERS CHANGING.6. HEMODYNAMIC RECORDING NUMBERS CHANGED.7. HEMODYNAMIC DATA CHANGING IN COMPUTER.8. HEMODYNAMIC WAVEFORMS FREEZE WHEN RECORDING, MAKING IT IMPOSSIBLE TO DO ANYTHING ELSE WITH COMPUTER.9. HEMODYNAMIC NUMBERS WERE CORRECTED BY TECHNOLOGIST, BUT COMPUTER CHANGED THEM BACK TO INCORRECT VALUES.10. EKG FROZE ON MONITOR, CAUSING A DELAY IN PROCEDURE. POTENTIAL RISK OF MISSING A HARMFUL ARRHYTHMIA.11. HOST PC FROZE POST DAILY MORNING REBOOT. HAD TO HARD BOOT, CAUSING DELAY IN BEGINNING PROCEDURE.12. WITT COMPUTER CHANGED NUMBERS. SOME OF THE WAVEFORMS WERE DELETEDWHILE PERFORMING CARDIAC OUTPUTS, THE WITT MONITOR FROZE. THEY CV HAD TO EXIT OUT OF THE PROCEDURE AND GET BACK WHILE THE CASE WAS GOING ON.13. WITT CHANGED LABEL OF HEMODYNAMICS. ANOTHER PRESSURE HAD TO BE TAKEN FOR ACCURACY.14. THE WITT MONITOR FROZE WHENEVER A HEMODYNAMIC WAVEFORM WAS SAMPLED HEMODYNAMIC WAVEFORMS FROZE WHILE SAMPLING.15. THERMODILUTION CARDIAC OUTPUT SCREEN FROZE HEMODYNAMIC DATA CORRECTED BY OPERATOR, THEN REVERTED BACK BY WITT (XIMS).16. WITT FROZE WHILE DOING TD CARDIAC OUTPUTS.17. HEMODYNAMIC WAVEFORMS FROZE WHILE SAMPLING.======================MANUFACTURER RESPONSE FOR HEMODYNAMIC MONITORING SYSTEM, XPER (PER SITE REPORTER).======================SLOW RESPONSE, SOME SOFTWARE CHANGES AND UPGRADES THAT HAVE NOT SEEMED TO HELP. | 1
2580476 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS WITH THE ANTIBODY SCREENING ASSAY ON THE GALILEO ECHO ((B)(4)). A REVIEW OF THE RESULT IMAGE FILES FOR INITIAL TESTING SHOWED THE CUSTOMER RECEIVED NEGATIVE RESULTS ON ALL THREE CELLS OF THE ANTIBODY SCREENING ASSAY. THE WELLS VISUALLY APPEARED NEGATIVE AS REPORTED BY THE INSTRUMENT. THE CONTROLS ON THE BATCH WERE WEAKER THAN NORMAL WHICH RESULTED WITH A 2+ REACTION AND NOT 4+ AS USUAL. CUSTOMER PERFORMED QC TWO TIMES AFTER TESTING THE PATIENT SAMPLE AND QC FAILED EACH TIME DUE TO EQUIVOCAL REACTIONS ON VARIOUS WELLS. THE CUSTOMER RECEIVED A Q SYMBOL ON ALL WELLS INCLUDING THE CONTROL. THE CONTROL WELLS APPEARED TO HAVE A HOLE IN THE WELL ON ALL OF THE QC FAILURES. THE CUSTOMER PERFORMED QC A THIRD TIME WITH PASSING RESULTS AFTER CHANGING TO A NEW VIAL OF CAPTURE-R READY INDICATOR CELLS OF THE SAME LOT. THE CUSTOMER REPEATED THE SAMPLE WHICH RESULTED AS POSITIVE (4+) ON CELL 3 OF THE SCREENING ASSAY AS EXPECTED. ALL OTHER CELLS RESULTED AS, AND VISUALLY APPEARED, NEGATIVE AS EXPECTED. A REVIEW OF THE MAINTENANCE, CAMERA REPORTS, AND EVENT LOG SHOWS THAT ALL MAINTENANCE WAS UP TO DATE ON THE DAYS OF TESTING. THE CAMERA REPORTS AND IMAGES SUGGEST THAT THE CAMERA WAS OPERATING AS EXPECTED. THE ERROR LOG SHOWED NO ERRORS ON THE DAY OF TESTING. WE WERE UNABLE TO RULE OUT THE VIAL OF CAPTURE-R READY INDICATOR CELLS AS BEING COMPROMISED DURING INITIAL TESTING. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2581069 | KEY PAD WORN AT CODE NUMBERS. KEY PAD REPLACED, HOSPITAL CODE PROGRAMMED, OPERATION PASSED AND UNIT RETURNED TO SERVICE. | 1
2582914 | THE ISSUE INVOLVES CERNER PROVISION DOCUMENT IMAGING UTILIZED TO CAPTURE BOTH CLINICAL AND NON-CLINICAL IMAGES. IN CERNER PROVISION DOCUMENT IMAGING, THE SYS MAY DELETE IMAGES FROM THE CONTENT STORAGE SYS, APPLICATION XTENDER. THE FOLLOWING MESSAGE IS DISPLAYED WHEN YOU ATTEMPT TO RETRIEVE AFFECTED FILES FROM STORAGE: "DOCUMENT NOT FOUND." PT CARE COULD BE ADVERSELY AFFECTED, AS THE MISSING INFO FROM PTS' CHARTS COULD RESULT IN DELAY OF CARE, ORDERING OF DUPLICATE PROCEDURES OR DECISIONS FOR CARE BASED ON INCOMPLETE INFO. CERNER HAS NOT RECEIVED COMMUNICATION ON ANY ADVERSE PT EVENTS AS A RESULT OF THIS ISSUE. CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON (B)(4) 2012 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE. THE CAUSE OF THE ISSUE IS STILL UNDER INVESTIGATION. A SUGGESTED WORK AROUND AS WELL AS ADD'L ACTIONS CLIENTS CAN TAKE TO FURTHER MITIGATE THE RISK OF THE ISSUE IN THE INTERIM HAVE BEEN PROVIDED TO ALL POTENTIALLY AFFECTED CLIENTS UNTIL SUCH TIME AS A PERMANENT RESOLUTION BECOMES AVAILABLE. CERNER CORP WILL PROVIDE A F/U REPORTS WHEN THE SOFTWARE MODIFICATION IS AVAILABLE. | 1
2583461 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT# (B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE ANALYSIS OF THE RUN DATA FILE (RDF) DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR HIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN PROCESS. A FOLLOW UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2583826 | PT TO HAVE NUCLEAR MEDICINE CARDIOLITE STRESS TEST. PT WAS IN STAGE 1 OF BRUCE PROTOCOL AND AFTER 1 MINUTE 30 SECONDS TREADMILL INCREASED IN SPEED WITHOUT STAFF INJECTING. TRIED TO STOP TREADMILL WITH EMERGENCY STOP BUTTON. THIS DID NOT WORK. THEN HIT STOP TREADMILL ON SYSTEM, DID NOT WORK. PT HAD TO JUMP OFF. | 0
2583866 | THE CUSTOMER REPORTED THAT THE ORTHO PROVUE HAD A BARCODE MISREAD (B)(4). NO ERROR MESSAGES WERE POSTED. NO ERRONEOUS RESULTS WERE REPORTED. AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND OBSERVED LABELS ARE NOT FULLY ADHERING TO THE TUBE. THE FE DID A SAMPLE CAMERA ADJUSTMENT. POST ADJUSTMENT PRODUCED CONSISTENT READS. THE CUSTOMER STATED THEY WILL MONITOR THE ISSUE AND CALL BACK IF THE ISSUE RETURNS. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4) | 1
2584095 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS WITH THE 3-CELL SCREENING ASSAY ON THE GALILEO ECHO (#(B)(4)) FOR A PATIENT WITH A HISTORY OF HAVING ANTI-C. AN IMMUCOR TECHNICAL SUPPORT SPECIALIST REVIEWED THE IMAGE RESULT FILES WHICH SHOWED NEGATIVE RESULTS WERE OBTAINED FOR ALL CELLS ON THE INITIAL ANTIBODY SCREENING ASSAY. CELLS 2 AND 3 VISUALLY APPEARED TO HAVE SOME RED CELL ADHERENCE AND BOTH CELLS WERE C POSITIVE CELLS. THE POSITIVE CONTROL WELL APPEARED POSITIVE AS EXPECTED WITH A WELL SCORE OF A 4+ INTERPRETATION. REVIEW OF THE ANTIBODY IDENTIFICATION ASSAY RESULTS SHOWED THE CUSTOMER RECEIVED 3 - 4+ REACTIONS ON CELLS 2-14 EXCEPT CELL 5 WHICH RESULTED AS 1+ POSITIVE. THE POSITIVE AND NEGATIVE CONTROL WELLS APPEARED AS EXPECTED WITH WELL SCORE OF 4+ FOR POSITIVE CONTROL AND 0 FOR NEGATIVE CONTROL. THE CUSTOMER REPEATED THE SAME SAMPLE AND RECEIVED AN EQUIVOCAL (?) REACTION ON CELL 2 AND 3+ ON CELL 3. CELLS 2 AND 3 APPEARED TO HAVE SOME RED CELL ADHERENCE. THE POSITIVE CONTROL WELL APPEARED POSITIVE AS EXPECTED WITH A WELL SCORE OF A 4+ INTERPRETATION. THE CUSTOMER STATED THAT SHE DID NOT CHANGE ANY REAGENTS ON THE INSTRUMENT. ALL OTHER SAMPLES TESTED AROUND THE SAME TIME RESULTED AS EXPECTED. THE CUSTOMER WAS MADE AWARE THAT NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO ARE TO BE VISUALLY INSPECTED PRIOR TO RELEASE OF RESULTS. THIS ISSUE WAS COMMUNICATED TO CUSTOMERS IN TECHNICAL COMMUNICATION (B)(4) ON (B)(4), 2009. | 0
2585961 | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO AN INCREASE IN SEIZURES. INTERROGATION OF THE PATIENT'S GENERATOR FOUND THAT THE NORMAL AND MAGNET OUTPUT CURRENTS WERE SET TO 0MA. REVIEW OF THE ACTIONS PERFORMED BY THE PHYSICIAN AT THE PREVIOUS VISIT FOUND THAT A SYSTEM DIAGNOSTICS TEST WAS PERFORMED ON (B)(6) 2011, HOWEVER A FINAL INTERROGATION WAS NOT PERFORMED. THE SYSTEM DIAGNOSTIC TEST WAS WITHIN NORMAL LIMITS. THE PATIENT WAS NOT SEEN BY ANY PHYSICIAN FOR HER VNS BETWEEN THAT VISIT AND BEING ADMITTED TO THE HOSPITAL. IT IS LIKELY THAT THE SYSTEM DIAGNOSTIC TEST FAULTED RESULTING IN AN UNINTENDED CHANGE IN SETTINGS. THE INCREASED SEIZURE FREQUENCY WAS NOTED AS BEING BELOW THE PRE-VNS BASELINE SEIZURE FREQUENCY. NO MEDICATION CHANGES WERE BELIEVED TO HAVE CONTRIBUTED TO THE PATIENT'S INCREASE IN SEIZURES. THE PHYSICIAN ATTRIBUTES THE INCREASED SEIZURES TO THE UNINTENDED SETTINGS AS THE PATIENT HAD BEEN SEIZURE-FREE FOR 2 YEARS PRIOR TO THE SUDDEN INCREASE. A COPY OF THE PHYSICIAN'S PROGRAMMING HISTORY HAS BEEN RECEIVED BY THE MANUFACTURER THAT CONFIRMS THE SEQUENCE OF ACTIONS ON (B)(6) 2011. THE PATIENT HAS BEEN PROGRAMMED BACK ON AND IS RESUMING HER NORMAL THERAPY. AGE AT TIME OF EVENT, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY REPORTED THE INCORRECT PATIENT AGE.><ANALYSIS OF PROGRAMMING HISTORY. | 0
2586932 | IT WAS INITIALLY DISCOVERED WHEN THE GENERATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION THAT IT WOULD FOUND AT SETTINGS THAT WERE INDICATIVE OF A FAULTED DIAGNOSTIC (OUTPUT CURRENT - 0 MA, SIGNAL FREQUENCY - 20 HZ, PULSE WIDTH - 500 USEC, ON TIME - 30 SEC, OFF TIME - 60 MINUTES). IT IS UNCLEAR IF THE PATIENT HAD HAVE BEEN INTENTIONALLY PROGRAMMING TO THOSE SETTINGS OR WAS LEFT AT THOSE SETTING AFTER A FAULTED DIAGNOSTICS SINCE THEY WERE BEING EXPLANTED. . | 0
2587155 | IT WAS INITIALLY DISCOVERED WHEN REVIEWING THE MANUFACTURER'S PROGRAMMING HISTORY DATABASE THAT THE PATIENT WAS SET TO UNINTENTIONAL SETTINGS FOLLOWING AN INCOMPLETE DIAGNOSTICS. THERE WERE TWO INTERROGATIONS FOLLOWING THE SETTINGS CHANGE BUT THE SETTINGS WERE NOT CORRECTED. IT IS UNKNOWN IF THE PHYSICIAN INTENTIONALLY LEFT THE PATIENT AT THOSE SETTING OR JUST DID NOT NOTICE THE CHANGE. THE SETTINGS WERE CORRECTED AT A LATER DATE. ANALYSIS OF PROGRAMMING HISTORY. | 0
2587342 | DURING AN ARTHROSCOPIC RIGHT KNEE PARTIAL MEDIAL MENISCECTOMY THE PATIENT DEVELOPED ACUTE COMPARTMENT SYNDROME OF THE RIGHT THIGH WHICH REQUIRED A FASCIOTOMY, ADMISSION TO THE HOSPITAL AND SUSBSEQUENT SURGERY. THE FMS DUO + FLUID MANAGMENT SYSTEM WITH INTERGRATED SHAVER WAS IN USE AND MAY NOT HAVE DISPLAYED AN ACCURATE PRESSURE READING. | 0
2588546 | THIS EMR DOES NOT CARRY FORWARD ORDERS TO STOP A DRUG. E.G., PT WAS HAVING A SUSPECTED COMMON KNOWN ADVERSE EFFECT TO A NARCOTIC. ON DISCHARGE, THERE IS A SCREEN WHEREUPON I CHECK A BOX INSTRUCTING TO STOP THIS MEDICATION. THAT INSTRUCTION IS NOT CARRIED OVER TO THE PRINTED LIST OF MEDS TO BE CONTINUED AT DISCHARGE, I.E., THE STOPPED MED IS STILL LISTED. | 1
2589135 | A CUSTOMER REPORTED OBTAINING AN UNEXPECTED NEGATIVE REACTION ON THE GALILEO ECHO INSTRUMENT ((B)(4)) FOR A SAMPLE WITH A HISTORY OF HAVING ANTI-E. THE CUSTOMER WAS ADVISED TO TEST THE SAMPLE MANUALLY. TESTING PERFORMED USING NHANCE RESULTED AS NEGATIVE FOR E POSITIVE CELLS. TESTING PERFORMED WITH PEG RESULTED AS POSITIVE (2+) WITH THE E POSITIVE CELLS. AN IMMUCOR TECHNICAL SUPPORT SPECIALIST REVIEWED THE IMAGE RESULT FILES. REVIEW OF THE TEST SAMPLE RESULTS SHOWED THAT SCREENING CELLS 1 AND 3 RESULTED AS NEGATIVE AND VISUALLY APPEARED NEGATIVE. SCREENING CELL II RESULTS AS NEGATIVE AND VISUALLY APPEARS WEAK POSITIVE WITH A WELL SCORE OF 2. CONTROLS REACTED AS EXPECTED. REPEAT TESTING RESULTED THE SAME AS INITIAL TESTING. A REVIEW OF THE EVENT LOG SHOWED NO ERRORS AT THE TIME OF TESTING. THE CAMERA IMAGES APPEARED ALIGNED AND FOCUSED. THE CUSTOMER WAS ADVISED OF TECHNICAL COMMUNICATION (B)(4) WHICH INFORMS USERS THAT IN SOME INSTANCES ON THE ECHO WHERE THE INSTRUMENT MAY GENERATE A NEGATIVE WELL INTERPRETATION FOR THE ANTIBODY SCREENING AND ANTIBODY IDENTIFICATION ASSAYS AND SUBSEQUENT VISUAL INTERPRETATION OF THOSE REACTIONS ARE WEAK POSITIVE OR QUESTIONABLE (EQUIVOCAL). CUSTOMERS ARE TO PERFORM A VISUAL VERIFICATION OF NEGATIVE REACTIONS BEFORE FINAL RELEASE OF THOSE WELL RESULTS. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2589607 | SPACELABS RECEIVED A REPORT THAT A BABY WAS BORN DECEASED WHEN A SPACELABS MATERNAL OBSTETRICAL MONITOR WAS IN USE.><SPACELABS RECEIVED A REPORT THAT A BABY WAS BORN DECEASED WHEN A SPACELABS MATERNAL OBSTETRICAL MONITOR WAS IN USE. THE CUSTOMER STATED THAT THE MONITOR DISPLAYED A FETAL HEART RATE OF 180 BEATS PER MINUTE (BPM) WITH OCCASIONAL ACCELERATIONS TO 200 BPM WHEN IT WAS MONITORING THE BABY. THE CUSTOMER PLACED THE DEVICE BACK INTO SERVICE AND WAS UNABLE TO IDENTIFY WHICH MONITOR WAS IN USE DURING THE ALLEGED INCIDENT. AS A RESULT, OUR FIELD SERVICE ENGINEER WAS UNABLE TO TEST THE MONITOR. WE ARE CONTACTING THE CUSTOMER TO COLLECT MORE INFORMATION. AT THIS TIME, THERE IS INSUFFICIENT DATA TO DRAW ANY CONCLUSION. WE WILL PROVIDE A SUPPLEMENTAL REPORT ONCE ADDITIONAL INFORMATION IS OBTAINED.><THE CUSTOMER PROVIDED FETAL STRIPS DURING LABOR AND DELIVERY TO SPACELABS. A CLINICAL SPECIALIST EVALUATED THE STRIPS AND CONCLUDED THE BELOW. THE STRIP LABELED LABOR TIME DEMONSTRATES NORMAL VARIABILITY OF THE FETAL HEART RATE (FHR). THE STRIP LABELED DELIVERY TIME DEMONSTRATES: OCCASIONAL MILD AND BRIEF FETAL TACHYCARDIA EPISODES AND SEVERAL LATE FHR DECELERATIONS WHICH WOULD HAVE BEEN AN INDICATOR OF FETAL DISTRESS; SEVERAL EPISODES OF EARLY FHR DECELERATIONS SHORTLY PRIOR TO DELIVERY WHICH WOULD HAVE BEEN ASSOCIATED WITH ROUTINE HEAD COMPRESSION. BECAUSE THE CUSTOMER WAS UNABLE TO IDENTIFY THE MONITOR USED DURING THE EVENT, SPACELABS HAS TESTED ALL THE MATERNAL OBSTETRICAL MONITORS IN THE HOSPITAL. TESTING CONFIRMED THAT ALL MONITORS WORKED TO SPECIFICATIONS. NO SUPPORTING INFORMATION HAS BEEN PROVIDED BY THE HOSPITAL SUCH AS LAB WORK, MATERNAL HISTORY, AUTOPSY, ETC. AS A RESULT, WE ARE UNABLE TO IDENTIFY THE ROOT CAUSE OF THE BABY DEATH. THE CUSTOMER IS CONTINUING TO USE THE DEVICE TO MONITOR PATIENTS. WE HAVE CONCLUDED THAT THIS REPORT IS FINAL AND CONSIDER THIS ISSUE CLOSED. | 1
2591377 | WHILE PERFORMING A TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) TO EVALUATE FOR LEFT ATRIAL APPENDAGE THROMBUS, THE TEE MACHINE SUDDENLY LOST POWER AND TURNED OFF. STAFF CHECKED ALL OF THE POWER CONNECTIONS AND USED DIFFERENT OUTLETS TO TROUBLESHOOT THE PROBLEM. THIS DID NOT RESULT IN THE ABILITY TO RESTART THE TEE MACHINE.PLEASE NOTE THAT THE PROBLEM WAS WITH THE ECHO MACHINE AND NOT THE TEE PROBE. THE PROBE IS FULLY FUNCTIONAL. | 1
2592076 | THE CUSTOMER REPORTED THE APEXPRO TELEMETRY SERVER (ATS) UNEXPECTEDLY STOPPED MONITORING PATIENTS FOR FOUR HOURS. THE BIOMED AT THE CUSTOMER SITE WAS UNABLE TO PROVIDE THE NUMBER OF PATIENTS THAT WERE AFFECTED, BUT AN ATS IS ABLE TO MONITOR UP TO 16 PATIENTS AT ONE TIME. ALTERNATE MONITORING WAS NOT AVAILABLE. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2592762 | CLINIC NOTES WERE RECEIVED. ON THE OFFICE VISIT ON (B)(6) 2009, THE PATIENT WAS REPORTED TO HAVE AN SIGNIFICANT INCREASE IN SEIZURES FROM ONE EVERY THREE TO FOUR DAYS IN THE LAST ONE AND A HALF MONTHS TO HAVE AT LEAST FOUR SEIZURES PER DAY. BOTH THE NORMAL AND MAGNET MODE OUTPUT CURRENT WAS INCREASED. THE PHYSICIAN REPORTED IN THE NOTES THAT THE OFF TIME WAS 180 MINUTES WHEN IT SHOULD HAVE BEEN 1.8 MINUTES. THE FAULTED SYSTEM DIAGNOSTIC TEST ON (B)(6) 2009, WHICH WAS OBSERVED IN THE IN-HOUSE DATABASE PREVIOUSLY CAUSED THE PATIENT BEING SET AT UNINTENDED SETTINGS. THE FAULTED DIAGNOSTIC RESULTS WERE MISINTERPRETED FOR THE BATTERY BEING "BURNED OUT." THEREFORE, THE PATIENT WAS REFERRED FOR GENERATOR REPLACEMENT DUE TO SUSPECTED "END OF SERVICE" AS A RESULT WHICH OCCURRED ON (B)(6) 2009. THE GENERATOR WAS RETURNED FOR ANALYSIS. THERE WERE NO PERFORMANCE OR ANY OTHER TYPE OF ADVERSE CONDITIONS FOUND WITH THE PULSE GENERATOR. THE PULSE GENERATOR PERFORMED ACCORDING TO SPECIFICATIONS, AND THE DEVICE WAS NOT AT END OF SERVICE. FOLLOW UP WITH THE PHYSICIAN'S OFFICE ON (B)(6) 2012, REVEALED THAT THE INCREASE IN SEIZURES WAS LIKELY RELATED TO THE PATIENT'S DEVICE NOT BEING AT INTENDED SETTINGS. THE MISINTERPRETATION OF THE DIAGNOSTIC RESULTS LED THE PHYSICIAN TO REFER THE PATIENT FOR GENERATOR REPLACEMENT DUE TO END OF SERVICE.><DURING REVIEW OF PROGRAMMING HISTORY AVAILABLE IN THE IN-HOUSE DATABASE, IT WAS OBSERVED THAT A FAULTED SYSTEMS DIAGNOSTIC TEST OCCURRED ON (B)(6) 2009 WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE ALL REPROGRAMMED ON THIS CLINIC VISIT AND A FINAL INTERROGATION WAS PERFORMED, BUT THE OFF TIME WAS INADVERTENTLY NOT CORRECTED TO THE CORRECT PARAMETER. PRIOR TO THE FAULTED TEST, THE OFF TIME WAS SET AT 1.8MIN, AND IT WAS CORRECTED TO 180MIN. AS A RESULT, THE PATIENT WAS NOT RECEIVING EFFICACIOUS SETTINGS. NO OTHER ANOMALIES WERE OBSERVED. TYPE OF REPORT, CORRECTED DATA: THE INITIAL REPORT INADVERTENTLY DID NOT INCLUDE THAT THIS IS A 30 DAY REPORT.><ANALYSIS OF PROGRAMMING HISTORY. | 0
2593842 | THE ISSUE REPORTED VIA A 04/26/2012 MEDWATCH FORM WAS ABNORMAL CELLS MISSED BY THIN PREP IMAGER SYSTEM (I.E. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW (FOV) PRESENTED). THE ROOT CAUSES DERIVED FROM PAST INVESTIGATIONS OF REPORTED ABNORMAL CELLS NOT PRESENT IN THE 22 FOVS FALL MAINLY INTO TWO CATEGORIES: RARE EVENT - AN EVENT IN WHICH THE SLIDE FALLS INTO THE CATEGORY WHERE THE CELLS IN QUESTION WOULD MATCH OUR EXPERIENCE IN THE CLINICAL TRIAL AS EVIDENCED IN TABLE 8 OF THE THINPREP IMAGING SYSTEM OPERATION SUMMARY AND CLINICAL INFO. CUSTOMER MISS - DETERMINED BY A HOLOGIC CYTOLOGY APPLICATIONS SPECIALIST (CAS) WHO IN A BLINDED MANNER MICROSCOPICALLY REVIEWS THE SLIDE(S) IN QUESTION AT THE LABORATORY. THE CAS IDENTIFIES ABNORMAL CELLS (TRIGGER CELLS) IN THE ORIGINAL IDENTIFIES ABNORMAL CELLS (TRIGGER CELLS) IN THE ORIGINAL 22 FOVS WHICH SHOULD HAVE PROMPTED A FULL SCAN REVIEW. THE ROOT CAUSE IS DETERMINED TO BE THE RESULT OF DIAGNOSTIC MISINTERPRETATION AND NOT RELATED TO THE FUNCTION / PERFORMANCE OF THE THINPREP IMAGER. CUSTOMER AND INSTRUMENT INFO WERE NOT INCLUDED ON MEDWATCH FORM. THERE IS NO INFO ON PT DIAGNOSIS, NOR A DELAY IN DIAGNOSIS. IF FURTHER INFO BECOMES AVAILABLE, A SEPARATE COMPLAINT WILL BE OPENED. | 0
2593853 | THE ISSUE REPORTED VIA A 04/26/2012 MEDWATCH FORM WAS ABNORMAL CELLS MISSED BY THIN PREP IMAGER SYSTEM (I.E. ABNORMAL CELLS NOT PRESENT IN THE 22 FIELDS OF VIEW (FOV) PRESENTED). THE ROOT CAUSES DERIVED FROM PAST INVESTIGATIONS OF REPORTED ABNORMAL CELLS NOT PRESENT IN THE 22 FOVS FALL MAINLY INTO TWO CATEGORIES: RARE EVENT - AN EVENT IN WHICH THE SLIDE FALLS INTO THE CATEGORY WHERE THE CELLS IN QUESTION WOULD MATCH OUR EXPERIENCE IN THE CLINICAL TRIAL AS EVIDENCED IN TABLE 8 OF THE THINPREP IMAGING SYSTEM OPERATION SUMMARY AND CLINICAL INFO. CUSTOMER MISS - DETERMINED BY A HOLOGIC CYTOLOGY APPLICATIONS SPECIALIST (CAS) WHO IN A BLINDED MANNER MICROSCOPICALLY REVIEWS THE SLIDE(S) IN QUESTION AT THE LABORATORY. THE CAS IDENTIFIES ABNORMAL CELLS (TRIGGER CELLS) IN THE ORIGINAL IDENTIFIES ABNORMAL CELLS (TRIGGER CELLS) IN THE ORIGINAL 22 FOVS WHICH SHOULD HAVE PROMPTED A FULL SCAN REVIEW. THE ROOT CAUSE IS DETERMINED TO BE THE RESULT OF DIAGNOSTIC MISINTERPRETATION AND NOT RELATED TO THE FUNCTION / PERFORMANCE OF THE THINPREP IMAGER. CUSTOMER AND INSTRUMENT INFO WERE NOT INCLUDED ON MEDWATCH FORM. THERE IS NO INFO ON PT DIAGNOSIS, NOR A DELAY IN DIAGNOSIS. IF FURTHER INFO BECOMES AVAILABLE, A SEPARATE COMPLAINT WILL BE OPENED. | 0
2594443 | AS PART OF OUR CONTINUOUS INTERNAL QUALITY CONTROL AND IMPROVEMENT PROCESS FOR THE VISAGE 7 SOFTWARE WE HAVE FOUND THAT TWO OF THE SPECIALIZED 2D MEASUREMENT TOOLS IN SOME SITUATIONS DO NOT ALLOW FOR MANUAL PIXEL SIZE CALIBRATION AND CAN LEAD TO MISLEADING RESULTS IF USED IN COMBINATION WITH THE MANUAL SIZE CALIBRATION TOOLS. THE AFFECTED TOOLS ARE THE SPECIALIZED VERTICAL DISTANCE MEASUREMENT TOOL (USED BY SOME CUSTOMERS FOR HIP TILT MEASUREMENT), AND THE 2D FREEHAND RIO TOOL. AS PART OF OUR CONTINUOUS INTERNAL QUALITY CONTROL AND IMPROVEMENT PROCESS FOR THE VISAGE 7 SOFTWARE WE HAVE FOUND THAT TWO OF THE SPECIALIZED 2D MEASUREMENT TOOLS IN SOME SITUATIONS DO NOT ALLOW FOR MANUAL PIXEL SIZE CALIBRATION AND CAN LEAD TO MISLEADING RESULTS IF USED IN COMBINATION WITH THE MANUAL SIZE CALIBRATION TOOL. THE AFFECTED TOOLS ARE THE SPECIALIZED VERTICAL DISTANCE MEASUREMENT TOOL (USED BY SOME CUSTOMERS FOR HIP TILT MEASUREMENT) AND THE 2D FREEHAND ROI TOOL. | 1
2594612 | TEC 6 VAPORIZER MALFUNCTIONED. END TIDAL WAS 4-5 WHEN DIAL SET TO 9-10. UNIT TAKEN OUT OF SERVICE, GIVEN TO BAXTER GE REP. GE MEDICAL REP (B)(4), WHO ALSO DID MEDWATCH FORM. | 0
2595355 | THE END USER REPORTED THAT THE PT WAS BURNED DURING THE USER OF A DIGITAL TENS UNIT WITH FOUR ELECTRODES. END USER: THE END USER REPORTED THAT THE PT WAS BURNED DURING THE USE OF A DIGITAL TENS UNIT WITH FOUR ELECTRODES. DEROYAL: THE DEFECTIVE SAMPLE WAS RETURNED. THE POWER BUTTON IS STUCK AND CONTINUES TO BE UNDER REVIEW. | 0
2595356 | THE CUSTOMER REPORTED THAT SHE INADVERTENTLY LEFT THE CLAMP OPEN DURING AIR REMOVAL. THERE WAS NO DONOR INVOLVED; THE INCIDENT HAPPENED DURING SET UP AND PRIME. THE CUSTOMER SETS UP AND PRIMES THE MACHINE PRIOR TO THE DONOR'S ARRIVAL. THERE WAS NOT A DONOR OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR DEATH OR INJURY OCCURRED. (B)(4). A REVIEW OF THE DEVICE HISTORY RECORD FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE OPERATOR INADVERTENTLY LEFT THE CLAMP OPEN DURING THE AIR REMOVAL PROCESS.><(B)(4). INVESTIGATION: THE CLINICAL SUPPORT SPECIALIST INFORMED THE CUSTOMER TO UNCLAMP THE NEEDLE LINE AND SQUEEZE THE SAMPLE POUCH UNTIL ALL OF THE AIR IS EXPRESSED. THEN ONCE ALL OF THE AIR WAS REMOVED, TO RECLAMP THE NEEDLE LINE AND CONTINUE PER PROTOCOL. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2596052 | CONFUSED PT DID NOT HAVE BOWEL MOVEMENT FOR APPROXIMATELY 6 DAYS. DOCUMENTATION IN THE EMR IS OBFUSCATED WITH GIBBERISH. NURSES TEDIOUSLY CHART THE DATA, BUT NO ONE LOOKS AT THE TOTAL PICTURE. NO ONE KNEW AND NO ONE REPORTED TO DOCTORS THAT PT DID NOT HAVE BOWEL MOVEMENT. PT DEVELOPED ILEUS WITH THEN, VOMIT AND ASPIRATION, WITH RESPIRATORY FAILURE DEATH. WHY IS THE EMR CENTRAL TO THE PT DEATH? THIS IS AN EXAMPLE OF THE IMPEDIMENT TO COMMUNICATION AND SAFE AND EFFICACIOUS CARE POSED BY A POORLY DESIGNED EMR IN WHICH IMPORTANT INFORMATION IS BURIED AMIDST UNIMPORTANT INFORMATION. OF COURSE, IT IS BASIC POLICY TO BLAME THE USER, BUT THE USERS ARE IMPEDED BY THE USER UNFRIENDLY DEVICE ITSELF. THIS IS AMPLIFIED WHEN EVERY DAY, ANOTHER DIFFERENT NURSE TAKES CARE OF THE PT. | 1
2596525 | IT WAS REPORTED THAT THE INDICATION OF THE LASSO CATHETER ON THE CARTO3 AND THE IMAGE OF THE FLUOROSCOPE STARTED NOT TO MATCH DURING THE ABLATION. THE ISSUE WAS RESOLVED BY CHANGING THE CATHETER CONNECTION FROM 20POLEA TO 20B. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. UPON SEVERAL FOLLOW-UPS, BWI FINALLY RECEIVED THE INFORMATION ON (B)(4), 2012 (WHICH CHANGED THE ALERT DATE) TO CLARIFY THAT IT WAS AN ACL MAP SHIFT AND THE SHIFT ONLY OCCURRED ON THE LASSO CATHETER, BUT THE AMOUNT OF THE SHIFT REMAINED UNKNOWN. DUE TO THE POTENTIAL RISK TO PATIENT FROM THIS ACL MAP SHIFT WITH UNKNOWN AMOUNT, THIS COMPLAINT BECAME REPORTABLE. INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON THE DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED.(B)(4).><(B)(4). IT WAS REPORTED THAT THE INDICATION OF THE LASSO CATHETER ON THE CARTO3 AND THE IMAGE OF THE FLUOROSCOPE STARTED NOT TO MATCH DURING THE ABLATION. THE ISSUE WAS RESOLVED BY CHANGING THE CATHETER CONNECTION FROM 20POLEA TO 20B. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. THE SYSTEM WAS REMOVED FROM THE CUSTOMER SITE AND SENT TO CARTO MANUFACTURER (HTC) FOR INVESTIGATION. HTC REPORTED THE CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. COMPLAINT REPORTED PROBLEMS WERE NOT REPRODUCED. THE SYSTEM PASSED ACL AND MAGNETIC TESTS. THE SYSTEM WAS TESTED WITH AQUARIUM STATION. THE SYSTEM HAS BUILT CPM WITH GOOD RESOLUTION AND ALL CATHETERS WERE VISIBLE PROPERLY. DURING ABLATION THE CATHETERS LASSO B AND LASSO A STAYED WITHIN ALLOWED TOLERANCES. THE WS WORKED PROPERLY. THE SYSTEM WORKED FOR A PERIOD OF 24 HOURS WITHOUT ANY COMMUNICATION PROBLEM. CUSTOMER'S COMPLAINT WAS NOT CONFIRMED. | 1
2597432 | PT ADMITTED TO HOSP APPEARS ON ACTIVE PT LIST OF THE DOCTOR TWICE. THE DOCTOR HAS A LONG LIST. ONE OF THE TIMES THIS PT APPEARS, FIRST, IS FROM AN OLD ADMISSION AT ANOTHER HOSP ON THE EMR SYSTEM. THE CURRENT ADMISSION IS REFLECTED WITH THE PT NAME AT THE BOTTOM OF THE LIST. WHEN ENTERING ORDERS FOR THE PT ON THE CURRENT ADMISSION, THE ORDERS WERE ENTERED FOR THE LISTING THAT APPEARED FIRST ON THE LIST. MANY ORDERS FOR TESTS, PROCEDURES, AND MEDICATIONS WERE KEPT ON THAT SECTION OF THE EMR AND DID NOT GO TO THE "ACTIVE" ADMISSION AND WERE NOT CARRIED OUT. MEDICATIONS AND TESTS WERE NOT GIVEN AND DELAYED PUTTING THE PT AT RISK. | 1
2597446 | ON ANY ONE FLOOR OF THE TERTIARY CARE CENTER, EACH WITH UP TO 12 COMPUTER TERMINALS, THE SUNRISE SYSTEM WOULD NOT BOOT UP FROM THE TERMINALS. THIS CAUSED DELAYS IN CARE AND WAS HOSPITAL WIDE. IT IS UNK IF THERE WERE ANY ADVERSE EVENTS BECAUSE OF THIS, BUT, UNDOUBTEDLY, ALL PTS WERE PUT AT RISK. THIS WENT ON FOR APPROX ONE WEEK BEFORE THE VENDOR SOLVED THE PROBLEM. | 1
2597689 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT THAT WAS MEASURED IN THE PLATELET PRODUCT COLLECTION. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #(B)(6). THE DISPOSABLE KIT WILL NOT BE RETURNED AS IT HAS BEEN DISCARDED. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYS EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN PROCESS. A F/U REPORT WILL BE PROVIDED. | 0
2598416 | OLYMPUS WAS INFORMED THAT THE SUBJECT DEVICE HAD RESET DURING A PROCEDURE AND ALL OF THE EQUIPMENT ON THE CART HAD LOST POWER. OLYMPUS FOLLOWED-UP WITH THE USER FACILITY TO OBTAIN ADDITIONAL INFORMATION REGARDING THIS REPORT. THE USER FACILITY REPORTED THAT THE REPORTED PHENOMENON OCCURRED DURING A BARIATRIC PROCEDURE, IN WHICH THE USERS REPORTEDLY WERE ABLE TO REACTIVATE THE SUBJECT DEVICE AND COMPLETED THE INTENDED PROCEDURE WITH NO PT INJURY REPORTED. AN OLYMPUS FIELD SERVICE ENGINEER (FSE) VISITED THE USER FACILITY TO EVALUATE THE SUBJECT DEVICE AND IDENTIFIED THAT THE CAUSE OF THE REPORTED PHENOMENON WAS ATTRIBUTED TO THE ISOLATION TRANSFORMER, AND THE POWER SWITCHES WHICH WERE BOTH REPLACED BY THE FSE. NONE OF THE DEVICES ON THE CART WERE SAID TO HAVE BEEN DAMAGED. NO DEVICES HAVE BEEN RETURNED TO OLYMPUS FOR EVALUATION AS OF THE DEVICE OF THIS REPORT. THIS REPORT WILL BE SUPPLEMENTED IF RELEVANT AND SIGNIFICANT INFORMATION BECOMES AVAILABLE AT A LATER DATE. THIS IS ONE OF THE TWO REPORTS. ALSO REFERENCE MFR REPORT NUMBER 9611174-2012-00006. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | 1
2603151 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RDF INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVAL AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2603210 | IT WAS INITIALLY REPORTED THAT THE PROGRAMMING HANDHELD WAS NO LONGER PRODUCING SOUND. THE VOLUME SETTINGS WERE ADJUSTED, BUT IT STILL WOULD NOT BEEP WHEN INTERROGATING. A HARD RESET WAS ALSO PERFORMED BUT THE HANDHELD SOUNDS WERE STILL NOT OCCURRING. IT WAS INDICATED THAT THE HANDHELD WAS ABLE TO PERFORM INTERROGATIONS AND DIAGNOSTICS WITHOUT ISSUE; HOWEVER THE SOUND WAS JUST NOT WORKING. ATTEMPTS FOR PRODUCT RETURN AND ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
2603915 | ANALYSIS WAS COMPLETED ON THE RETURNED HANDHELD. ANALYSIS OF THE HANDHELD REVEALED THAT THE CAUSE FOR THE REPORTED COMPLAINT WAS ASSOCIATED WITH FLUID THAT WAS TRAPPED UNDER THE TOUCHSCREEN LAYER CAUSING THE DISPLAY TO BE UNRESPONSIVE. ONCE THE DISPLAY WAS REPLACED, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. MOREOVER, ANALYSIS OF THE RETURNED FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED BY A PHYSICIAN THAT HER HANDHELD COMPUTER NO LONGER FUNCTIONED PROPERLY AS SOME CLEANING AGENT PENETRATED THE DEVICE. A RESET WAS PERFORMED ON THE HANDHELD, BUT NO RESPONSE WAS RECEIVED. THE REPORTED HANDHELD WAS RETURNED TO THE MANUFACTURER AND CURRENTLY UNDERGOING ANALYSIS. .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2604814 | ON (B)(6) 2012, DURING REVIEW OF PRODUCT ANALYSIS RESULTS, IT WAS NOTED THAT THIS EXPLANTED DEVICE WAS RECEIVED WITH SETTINGS INDICATIVE OF A FAULTED DIAGNOSTIC. IN REVIEW OF THE PROGRAMMING HISTORY, A FAULTED DIAGNOSTIC WAS OBSERVED ON (B)(6) 2010, WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED SETTINGS. ALTHOUGH A FINAL INTERROGATION WAS PERFORMED, THE SETTINGS WERE NOT CORRECTED. THE DEVICE WAS EXPLANTED ON (B)(6) 2010, AND SINCE THE DEVICE WAS RECEIVED AT THESE SAME UNINTENDED SETTINGS, IT INDICATES THAT THE PATIENT'S SETTINGS WERE NOT CORRECTED PRIOR TO EXPLANT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS ISSUE. ANALYSIS OF PROGRAMMING HISTORY. | 0
2605109 | A NUMBER OF EVENTS HAVE BEEN IDENTIFIED WITH THE USE OF GE CARDIAC MONITORS, SPECIFICALLY WITH THE DASH PORT2 DOCKING STATIONS, DASH 3000, 4000, AND 5000 MONITORS. THESE EVENTS ARE CENTERED AROUND THE "NO COMM" MESSAGE THAT APPEARS ON THE MONITORS AND CIC (CENTRAL INFORMATION CENTER). THE "NO COMM" MESSAGE OCCURS NORMALLY WHEN THE MONITOR IS TURNED OFF, OR WHEN THERE IS A LOSS OF COMMUNICATION TO THE CIC. THE MESSAGE APPEARS IN RED ON THE MONITOR AND CIC.THERE ARE TWO ISSUES THAT HAVE BEEN IDENTIFIED. THE FIRST IS WHEN A "NO COMM" MESSAGE IS DISPLAYED AND NO OTHER INFORMATION IS DISPLAYED, THERE IS A THREE MINUTE DELAY IN GENERATION OF AN AUDIBLE ALARM FOR THE "NO COMM". FOR AN AUDIBLE ALARM TO BE PRODUCED, THE CIC MUST BE PROGRAMMED TO DO SO. WHEN THE CIC IS INITIALLY INSTALLED, IT DOES NOT HAVE THE AUDIBLE ALARM TURNED ON. THE SECOND IS WHEN THE "NO COMM" MESSAGE IS DISPLAYED AND THE INFORMATION REGARDING THE EVENT IS ALSO DISPLAYED ACROSS THE TOP OF THE CIC IN A SCROLLING MESSAGE, NO AUDIBLE ALARM IS GENERATED, AND THE WAVEFORM WILL NOT BE DISPLAYED OR STORED. DEPENDING ON THE NUMBER OF ALARMS THAT ARE BEING PRODUCED ON THE CIC, THE SCROLLING MESSAGE MAY ONLY LAST A FEW SECONDS.GE HAS IDENTIFIED THAT THIS CONCERN IS DUE TO A FAILURE ON THE 1J AND 2J BOARD (CPU) IN UNITS WITH CERTAIN IDENTIFIED SERIAL NUMBERS.THIS ISSUE WAS IDENTIFIED IN JANUARY 2012 AND IS CONTINUING. AS OF THIS DATE THERE HAVE BEEN NO KNOWN ADVERSE OUTCOMES TO ANY PATIENTS. AS OF THE DATE OF THIS REPORT, 28 OCCURRENCES HAVE BEEN REPORTED. | 1
2607558 | THE CARESTREAM DRX-1 SYSTEM DETECTOR (FLAT PANEL DIGITAL IMAGER) USES A BATTERY. THE BATTERY WAS NOT INSTALLED IN THE DETECTOR WHEN THIS INCIDENT OCCURRED. THE BATTERY CHARGERS WERE MOUNTED ON THE WALL AND THE BATTERY WAS FOUND ON THE FLOOR, APPROXIMATELY 10 FEET FROM THE CHARGERS. THE BATTERY BECAME HOT ENOUGH TO MELT SOME OF THE PLASTIC CASING AND THE LABEL ON THE BATTERY CELL. THERE WERE NO KNOWN WITNESSES TO WHAT HAPPENED TO THE BATTERY. THE BATTERY WAS FOUND ON THE FLOOR, JUST BEFORE 5PM ON SATURDAY, AFTER A BURNING SMELL WAS NOTICED BY PERSONNEL IN THE VICINITY. IT IS NOT KNOWN WHEN OR IF THE BATTERY WAS PLACED INTO THE CHARGER. THE BATTERY CHARGER FROM THE SITE WAS ANALYZED AND NO RESIDUE WAS FOUND ON THE CHARGER. INVESTIGATION SHOWED THAT WHATEVER HAPPENED TO THE BATTERY OCCURRED OUTSIDE OF THE CHARGER. ONE CELL (#2 CELL) OF THIS BATTERY (THERE ARE 4 CELLS WITHIN EACH BATTERY) OVERHEATED AND PARTIALLY MELTED THE PLASTIC CASING OF THE BATTERY. A SMALL AREA ON THE FLOOR WAS BLACKENED - DUE TO INTERNAL THERMAL VENTING OF THE BATTERY PACK. THE BATTERY MANUFACTURE DATE IS 01/2012. THE BATTERY WAS NOT INSTALLED IN THE MEDICAL DEVICE (THE CARESTREAM DRX-1 SYSTEM) WHEN THE INCIDENT OCCURRED. THE SITE CONTACT DID NOT KNOW WHICH DRX-1 SYSTEM THE BATTERY HAD LAST BEEN USED INT. THE CARESTREAM DRX-1 SYSTEM WAS NOT RETURNED TO THE MANUFACTURER. THE BATTERY WAS RETURNED TO THE BATTERY MANUFACTURER FOR FAILURE ANALYSIS. A ROOT CAUSE HAS NOT YET BEEN IDENTIFIED. VARIOUS POSSIBLE CAUSES HAVE BEEN EVALUATED. | 1
2610436 | #2 PROBE FROM 6 PROBE SET WAS DEFECTIVE. PROBE PASSED TEST AT START OF CASE BUT WHEN USED DURING SURGERY THE WHOLE NEEDLE FROZE, CAUSING THE PERINEAL SKIN TO FREEZE. IT WAS IMMEDIATELY REMOVED AND REPLACED WITH A SINGLE WRAPPED PROBE. THE SKIN INITIALLY FROZE ON THE PATIENT BUT WE REMOVED THE PROBE QUICKLY ENOUGH THAT THE SKIN WAS OK. NO INJURY TO THE PATIENT. ALBERTA HEALTH SERVICES UNABLE TO RETURN THE PRODUCT FOR EVALUATION. DEVICE HISTORY RECORD REVIEW COMPLETED WITH NO RELATED ISSUES FOUND.REQUEST SUBMITTED FOR PATIENT INFORMATION AND CURRENT CONDITION. NOTIFIED IT IS AGAINST (B)(6) POLICY TO GIVE OUT PATIENT IDENTIFIERS AND PATIENTS CURRENT CONDITION AS THIS IS AGAINST PRIVACY AS WELL.PER ADDITIONAL INFORMATION RECEIVED FROM NURSE, "WE WERE ABLE TO CATCH BEFORE PERMANENT DAMAGE WAS DONE." | 0
2613808 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD DEVICE WAS NOT HOLDING A CHARGE. ATTEMPTS FOR PRODUCT RETURN AND ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2012 AND PRODUCT ANALYSIS HAS SINCE BEEN COMPLETED. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2614858 | ARTIC SUN WARMING DEVICE WAS PLACED ON PT WHOSE TEMPERATURE HAD FALLEN TO 33 DEGREES CELSIUS. THE PT EXPIRED DUE TO OTHER CO-MORBIDITIES. WHEN THE STAFF REMOVED THE ARCTIC SUN PADS FROM THE PT, THE SKIN WAS BLISTERED AND SCALDED. THE SKIN CAME OFF OF THE PT WHERE THE PADS HAD BEEN ATTACHED. | 0
2615684 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) COUNT IN THE PLATELET PRODUCT, ALONG WITH THE HIGHER THAN EXPECTED PLATELET PRODUCT VOLUME. THE REPORTED MEASURED PRODUCT VOLUME WAS 458 ML, WHEN THE TRIMA ACCEL SYSTEM REPORTED THE PRODUCT VOLUME TO BE 239 ML. THERE WAS NOT A TRANSFUSION RECIPIENT OR PT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PT INFO IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. POSSIBLE ROOT CAUSES FOR THE VOLUME DISCREPANCY WERE ALSO PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RDF DO NOT INDICATE A CLEAR CAUSE FOR THE HIGHER THAN EXPECTED PLATELET PRODUCT VOLUME AND THE ELEVATED WBC CONTENT IN THE PLATELET PRODUCT. BASED ON THE REPORTED PRODUCT VOLUME, THE DONOR WAS STILL WITHIN THE 15% TOTAL BLOOD VOLUME DONOR VOLUME LIMIT SPECIFIED IN THE MACHINE'S CONFIGURATIONS. FIFTEEN PERCENT OF THE TBV WOULD BE 569ML; 458ML WAS REPORTED AS COLLECTED, THEREFORE, MORE THAN THE 15% TBV VOLUME LIMIT WAS NOT COLLECTED. THE SIGNALS IN THE RDF DO INDICATE THE POSSIBILITY OF A SMALL DIVERSION OF FLUID FROM THE RETURN RESERVOIR TO THE PLATELET PRODUCT DURING EACH DRAW AND RETURN CYCLE, HOWEVER, THE AMOUNT OF VOLUME DIVERTED WAS NOT SIGNIFICANT ENOUGH TO TRIGGER A LEVEL SENSOR ERROR ALERT. IT CANNOT BE RULED OUT THAT A POSSIBLE MISLOADING OF THE CASSETTE, SPECIFICALLY THE LOADING OF THE PLATELET PUMP HEADER, COULD HAVE CONTRIBUTED TO A SMALL AMOUNT OF ADDITIONAL PRODUCT VOLUME BEING PUMPED TO THE PRODUCT BAGS EACH DRAW AND RETURN CYCLE. THIS VOLUME COULD HAVE ACCUMULATED OVER EACH DRAW AND RETURN CYCLE AND RESULTED IN THE HIGHER THAN EXPECTED PLATELET PRODUCT VOLUME. THERE IS NO CLEAR INDICATION IN THE RDF OF ANY POSSIBLE CAUSE FOR THE ELEVATED WBC CONTENT IN THE PLATELET PRODUCT. HOWEVER, IT CANNOT BE RULED OUT THAT THE ELEVATED WBC CONTENT COULD BE A POSSIBLE SIDE EFFECT OF A MISLOADED PLATELET PUMP HEADER. BASED ON THE AVAILABLE INFO, IT ALSO CANNOT BE RULED OUT THAT THE VOLUME AND WBC ISSUES COULD BE CAUSED BY WEIGHING, CALCULATION, OR OTHER PROCESS ERROR. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2615899 | THE CUSTOMER REPORTED THEY LOST MONITORING FOR FOUR PTS FOR 45 MINUTES WITH NO ALTERNATE PT MONITORING AVAILABLE. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2618686 | SPACELABS RECEIVED A REPORT THAT A TELEMETRY MONITOR FAILED TO ALARM DURING A PATIENT CODE AND THE PATIENT DIED. A SPACELABS FIELD SERVICE ENGINEER (FSE) TESTED THE SUBJECT DEVICES AT THE CUSTOMER'S SITE AND CONFIRMED THE DEVICES WORKED TO SPECIFICATIONS. THE PATIENT WAVEFORM SHOWS A SERIES OF RUNS AND LOW ECG VOLTAGE THAT WOULD HAVE CAUSED ALARMS IF ALARMS WERE TURNED ON IN THE MODULE LOCATED IN THE TELEMETRY RECEIVER HOUSING. THE ALARM REPORT FROM THE (B)(6) CLINICAL SUITE SYSTEM DOES NOT SHOW ANY ALARMS RECEIVED FROM THE MODULE. THE MODULE, WHICH STORES DATA, WAS REMOVED FROM THE RECEIVER HOUSING BY THE CUSTOMER AFTER THE EVENT. AS A RESULT OF BEING REMOVED AND THE TIME BETWEEN THE INCIDENT AND THE REPORT TO SPACELABS, ALL THE TRENDED DATA AND MODULE SETTINGS WERE LOST. AS A RESULT, WE ARE UNABLE TO DETERMINE IF THE ALARMS WERE TURNED OFF OR ON DURING THE EVENT. THE ROOT CAUSE OF THE ALLEGED FAILURE CANNOT BE IDENTIFIED. THE CUSTOMER HAS PLACED THE TELEMETRY MONITOR BACK INTO SERVICE AND CONTINUES TO USE THE DEVICE TO MONITOR PATIENTS. IT IS SPACELABS' POLICY TO SUBMIT A MEDICAL DEVICE REPORT WHENEVER WE BECOME AWARE OF THE DEATH OF A PATIENT CONNECTED TO ONE OF OUR DEVICES. WE HAVE CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AND CONSIDER THIS ISSUE CLOSED. | 1
2619037 | ON (B)(6) 2009, A (B)(6) MALE PT WAS SCHEDULED FOR SEED IMPLANTATION. THE REPORTER EXPERIENCED A HIGHER THAN BACKGROUND ACTIVITY IN A CARTRIDGE MANUFACTURED BY MICK RADIO-NUCLEAR INSTRUMENTS LOADED WITH IODINE I-125 SEEDS, MODEL 6711, 0.324 MCI CALIBRATED FOR 20(B)(6) 2009, MANUFACTURED BY GE HEALTHCARE. THE PRODUCT WAS SOLD AS A STERILE CONVENIENCE PACK (BATCH (B)(4)). THERE WAS NO CONTAMINATION FOUND PRIOR TO THE IMPLANT. THE UROLOGIST DID NOT EXPERIENCE ANY PROBLEMS (SUCH AS RESISTANCE) WITH THE APPLICATOR OR CARTRIDGE DURING THE IMPLANT PROCEDURE. AS THE SEEDS WERE BEING IMPLANTED INTO THE PT AND APPROXIMATELY 2/3 OF THE WAY THROUGH THE PROCEDURE, AN EMPTY CARTRIDGE WAS PASSED BACK AND SURVEYED. THE SURVEY INDICATED HIGHER (ACTIVITY NOT SPECIFIED) THAN BACKGROUND ACTIVITY. THE UROLOGIST WAS CONCERNED THAT ONE OF THE SEEDS LEAKED IN THE CARTRIDGE AND THE LEAKING SEED WAS NOW IMPLANTED INTO THE PT. THE PT WAS PLACED ON THYROID BLOCKING MEDICATION. BLOOD SAMPLES SHOWED A SLIGHTLY HIGHER THAN BACKGROUND ACTIVITY (ACTUAL ACTIVITY NOT SPECIFIED). | 0
2619925 | PATIENT WAS IN MRI MACHINE AND FELT BURNING SENSATION. PATIENT COULD NOT GET THE ATTENTION OF STAFF. 'PANIC BULB' WAS NOT GIVEN TO PATIENT AND SHE SUSTAINED A BURN, AT LEAST AT A STAGE 2 BURN OVER 3X3CM. THE MACHINE WAS TAKEN OUT OF SERVICE ONCE EVENT WAS DISCOVERED. GE WAS CONTACTED AND ARRIVED TO INSPECT THE MACHINE. MACHINE PASSED INSPECTION PER GE TECHNOLOGISTS. THEY PERFORMED AN ON-SITE INSPECTION & PASSED THE DEVICE FOR FURTHER USE. | 0
2620879 | THE CUSTOMER REPORTED THAT THE USER EXPERIENCED A SHOCK WHEN RE-SEATING THE POWER CORD ON THE BACK OF THE TOUCH DISPLAY. THERE WAS NO REPORTED INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. CUSTOMER UNCERTAIN OF WHICH SOLAR DEVICE THE DISPLAY WAS CONNECTED TO. THE VERSION OF SOLAR DEVICE/SOFTWARE IS UNKNOWN AT THIS TIME HENCE IDENTIFICATION OF THE 510(K) NUMBER WHICH DIRECTLY SUPPORTS THIS VERSION OF DEVICE IS UNCERTAIN. MANUFACTURER DATE UNKNOWN. | 1
2620902 | THE CUSTOMER REPORTED THAT DROP OUTS WERE OCCURRING INTERMITTENTLY, AFFECTING POTENTIALLY 96 PATIENTS WITH A DURATION LASTING A FEW HOURS. ALTERNATE PATIENT MONITORING WAS NOT OTHERWISE AVAILABLE. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2620917 | IT WAS REPORTED THAT A TECHNICIAN (ASSISTANT RADIOGRAPHIC PRACTITIONER) WAS INJURED AFTER HITTING A DISPLAY MONITOR WITH HER HEAD. THE CUT WAS ABOUT 2-3MM IN LENGTH AND REQUIRED MEDICAL INTERVENTION (SURGICAL GLUE). GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. PT'S WEIGHT WAS NOT PROVIDED. | 0
2621696 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE TRIPLE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RDF DO NOT INDICATE A CLEAR CAUSE FOR THE ELEVATED WBC CONTENT IN THE PLATELET PRODUCT. THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. THERE IS SOME INDICATION IN THE RDF OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER LATER IN THE PROCEDURE, HOWEVER THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATE WBC COUNTS. | 0
2621877 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONATION #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. (B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ANALYSIS OF THE RDF DID NOT SHOW SPECIFIC ROOT CAUSE OF THE ELEVATED WBC COUNTS MEASURED. IT IS POSSIBLE THAT THE FAILURE IS PART OF THE STATISTICAL DISTRIBUTION FOR LEUKOREDUCTION AT THIS CENTER. IT IS ALSO POSSIBLE THAT THIS FAILURE IS DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2621878 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONATION #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ANALYSIS OF THE RDF DID NOT SHOW SPECIFIC ROOT CAUSE OF THE ELEVATED WBC COUNTS MEASURED. IT IS POSSIBLE THAT THE FAILURE IS PART OF THE STATISTICAL DISTRIBUTION FOR LEUKOREDUCTION AT THIS CENTER. IT IS POSSIBLE THAT THE NUMEROUS ACCESS PRESSURE ALERTS COULD HAVE DISRUPTED THE STEADY STATE OF THE SYSTEM AND COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC COUNT. IT IS ALSO POSSIBLE THAT THIS FAILURE IS DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2622506 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONATION #:(B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2624531 | SAME SYSTEM AS REPORTED UNDER MFR #2184149-2012-00014. IT WAS REPORTED DURING AN ABLATION PROCEDURE, THE STIMULATOR WAS SENDING IMPROPER STIMULUS TO THE PT. THE PHYSICIAN WAS USING THE F3 PROTOCOL AND WANTED TO PACE AT 600 MS, BUT THE STIMULUS THAT WAS SENT WAS AT 465 MS. THE PHYSICIAN CHECKED THE SETTINGS AND STIMULATED AND THEN SAW THE VALUES WERE NOT WHAT HAD BEEN PROGRAMMED. THE PROCEDURE WAS STOPPED WITH NO CONSEQUENCES TO THE PT. THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, A SUPPLEMENTAL MEDWATCH WILL BE FILED. | 0
2625378 | THE CUSTOMER REPORTED FAINT CELL BUTTON IMAGE ON CELL 2 DURING A RUN ON TANGO OPTIMO. A FIELD SERVICE ENGINEER WAS DISPATCHED TO CHECK THE INSTRUMENT. AN IMAGE WAS THE REPORTED PROBLEM CLEARLY SHOWED AN INSUFFICIENT AMOUNT OF RED CELLS INSIDE THE AFFECTED CAVITY, THAT LED TO A FALSE POSITIVE IMAGE INTERPRETATION. OUR FIELD SERVICE IDENTIFIED AN OBSTRUCTED TUBING OF THE WASHER MODULE TO BE CAUSATIVE FOR THIS EVENT. THE AFFECTED WASHER MODULE WAS REPLACED. THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT. | 0
2625800 | THE PT WAS ADMITTED TO THE NEURO-INTENSIVE CARE UNIT FOR CEREBRAL BLEED. ON (B)(6) 2009, THE PT WAS PLACED ON THE ARCTIC SUN TEMPERATURE MGMT SYSTEM FOR NEUROGENIC FEVER MGMT. ON (B)(6) 2009 BLISTERS WERE NOTED ON THE PT'S BACK AND THIGHS, WHICH REQUIRED DEBRIDEMENT AND TREATMENT WITH SILVADENE. THE DEVICE WAS EVALUATED AT THE CUSTOMER SITE BY THE HOSPITAL BIOMEDICAL ENGINEERING STAFF IN CONSULTATION WITH MEDIVANCE TECHNICAL SUPPORT. THE HOSPITAL BIOMEDICAL STAFF REPORTED TO MEDIVANCE THAT THE DEVICE WAS FOUND TO BE PROPERLY CALIBRATED AND PERFORMING WITHIN SPECIFICATIONS. | 0
2625968 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2012, WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE HANDHELD AND FLASHCARD. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS; THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE LOCK BUTTON WAS IN THE LOCKED POSITION. THIS ANOMALY COULD NOT HAVE CONTRIBUTED TO THE EVENT THAT WAS REPORTED BY THE NURSE. ONCE THE LOCK BUTTON WAS MOVED TO THE UNLOCKED POSITION, THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED BY A COMPANY REPRESENTATIVE THAT A NURSE'S HANDHELD COMPUTER INTERMITTENTLY SWITCHES OFF AND THE BATTERY DRAINS FAIRLY RAPIDLY. A NEW HANDHELD WAS PROVIDED TO THE NURSE AND THE REPORTED HANDHELD WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS. . | 0
2626436 | THE PATIENT, WITH A HISTORY OF METASTATIC PHEOCHROMOCYTOMA, PRESENTED FOR A CT CRYOABLATION OF LUMBAR VERTEBRAL BODY AND SACRAL TUMORS. FOUR CRYOPROBES WERE INSERTED INTO THE LESIONS (S1, S2 AND L2). USING INTERMITTENT CT FLUOROSCOPIC GUIDANCE AN OSTEO-SITE BONE BIOPSY SET WAS PLACED THROUGH THE L2 VERTEBRAL BODY FROM THE LEFT. THE NEEDLE WAS REMOVED AND THE CRYOPROBE WAS PLACED. AN ADDITIONAL OSTEO-SITE NEEDLE WAS USED TO DEVELOP A TRACK TO THE LESION VIA THE RIGHT L2 PEDICLE. THE OSTEO-SITE WAS REMOVED AND A CRYOPROBE WAS INSERTED WHICH CROSSED THE PREVIOUSLY PLACED PROBE. THE CRYOPROBE THAT SUFFERED DAMAGE WAS THE PROBE PLACED FROM THE LEFT. OSTEO-SITE NEEDLES ARE USED TO CREATE THE TRACK FOR THE PLACEMENT OF THE PROBES. THE OSTEO-SITE NEEDLE IS THEN REMOVED AND THE PROBE IS INSERTED THROUGH THE TRACK. THE PLACEMENT OF THE PROBES WAS CONFIRMED BY CT IMAGING. THE ARGON GAS WAS TURNED ON AT 20% FOR THE CRYOPROBES PLACED IN THE L2 LESION. APPROXIMATELY 20-30 SECONDS LATER THERE WAS IMMEDIATE NEUROLOGIC AND VITAL SIGN DETERIORATION. THE PROCEDURE WAS STOPPED, THE PROBES WERE REMOVED, AND RESUSCITATION WAS ATTEMPTED UNSUCCESSFULLY. | 0
2626700 | DURING REVIEW OF THE VNS PATIENT'S PROGRAMMING HISTORY ON (B)(6), 2012 IT WAS DISCOVERED THAT A GENERATOR DIAGNOSTICS TEST WAS PERFORMED ON (B)(6), 2007 WHICH CHANGED THE PATIENT'S SETTINGS. NO FINAL INTERROGATION WAS PERFORMED AND IT WASN'T UNTIL THE PATIENT'S FOLLOWING VISIT ON (B)(6), 2007 THAT THE INCORRECT SETTINGS WERE NOTICED AND THE PATIENT WAS REPROGRAMMED TO THE CORRECT SETTINGS. ANALYSIS OF PROGRAMMING HISTORY. | 0
2626923 | THE CUSTOMER DECLINED TO PROVIDE THE PATIENT INFORMATION.><A CUSTOMER REPORTED THAT A DONOR WAS ASSIGNED TO A MACHINE, BUT THEN A DIFFERENT DONOR WAS RUN ON THAT MACHINE USING THE ORIGINAL DONOR'S INFO. THE DONOR WHO WAS ACTUALLY DRAWN ON THE MACHINE HAD A SMALLER TBV THAN THE ORIGINAL DONOR, BUT THE RUN WAS ABORTED IMMEDIATELY AFTER PRIME. THERE WAS NO DONOR REACTION AND NO BLOOD WAS DRAWN SUCCESSFULLY FROM THE DONOR. PATIENT INFORMATION IS UNAVAILABLE AT THIS TIME. THIS REPORT IS BEING FILED DUE TO AN OPERATOR ERROR THAT WOULD LIKELY CAUSE OR CONTRIBUTE TO A DEATH OR INJURY IF THIS SAME FAILURE WERE TO RECUR. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. INVESTIGATION: THE PROCEDURE WAS ABORTED RIGHT AFTER PRIME DUE TO AN INFILTRATION AND NO BLOOD WAS DRAWN FROM THE DONOR. THE DONOR DID NOT HAVE A REACTION DURING THE PROCEDURE, NOR DID THEY PRESENT ANY ISSUES AFTERWARD AND NO MEDICAL INTERVENTION WAS PERFORMED. THIS IS A SOFTWARE PRODUCT WHICH DOES NOT HAVE AN ASSOCIATED TECHNICAL FILE, SIMPLY A BATCH WORK ORDER WHICH WOULD DOCUMENT ANY PRODUCTION DEFECTS. A REVIEW OF THE PRODUCT DISPOSITION FILE SYSTEM FOR THIS CATALOG NUMBER WAS PERFORMED AND NO RECORDS WERE FOUND. ROOT CAUSE: OPERATOR ERROR. THE CUSTOMER'S SALES PERSON, WAS NOTIFIED SO THAT HE COULD FOLLOW UP WITH THE CUSTOMER AND POSSIBLY CONSIDER SOME RE-TRAINING FOR THE OPERATORS.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2627204 | AT OUR INSTITUTION, THE ELECTRONIC ORDERING AND EHR WAS DEPLOYED WITH FARCICAL TRAINING, AND NO WARNING TO AND PREPARATION FOR CLINICIANS WHO ATTEND THERE A FEW WEEKS PER MONTH. A PT WITH A MASSIVE GASTROINTESTINAL BLEED WAS ADMITTED TO ONE OF THESE DOCTORS. THE SIGN ONS CHANGED AND NO ONE AT THE HELP DESK (MANNED VIA (B)(6)) COULD OBTAIN THE CORRECT SIGN ON FOR THIS DOCTOR. THE CARE OF THE PT WAS LIMITED DUE TO LACK OF ACCESS TO THE DETAILS NEEDED IN THE MEDICAL RECORD TO ADEQUATELY MANAGE THIS CASE. THEY WERE NOT AVAILABLE ON ANY PAPER DOCUMENTS. THE ENTIRE HOSPITAL HAS BEEN COMPROMISED FOR DAYS, ACCORDING TO EYE WITNESS ACCOUNTS BY NURSES AND ANCILLARY STAFF. DEATHS WILL OCCUR FROM SUCH DELAYS AND CONFUSION IN ASSOCIATION WITH MEANINGFULLY USELESS HELP DESKS AND OTHER SUPPORT. | 1
2628477 | IT WAS REPORTED THAT DURING AN ISCHEMIC VT PROCEDURE, ERROR 260 WAS DISPLAYED ON THE CARTO 3 SYSTEM USED DURING THE PROCEDURE. THE ERROR MESSAGE OCCURRED JUST AFTER AN EXTERNAL EMERGENCY DEFIBRILLATION CAUSING PATIENT MOVEMENT ON THE TABLE. THE USER WAS ADVISED TO REBOOT THE SYSTEM WITHOUT CATHETERS CONNECTED TO PIU. WHEN THE STUDY WAS CONTINUED FROM REVIEW MODE. CASE WAS COMPLETED SUCCESSFULLY. (B)(4). IT WAS REPORTED THAT DURING AN ISCHEMIC VT PROCEDURE, ERROR 260 WAS DISPLAYED ON THE CARTO 3 SYSTEM USED DURING THE PROCEDURE. THE ERROR MESSAGE OCCURRED JUST AFTER AN EXTERNAL EMERGENCY DEFIBRILLATION CAUSING PATIENT MOVEMENT ON THE TABLE. THE USER WAS ADVISED TO REBOOT THE SYSTEM WITHOUT CATHETERS CONNECTED TO PIU. WHEN THE STUDY WAS CONTINUED FROM REVIEW MODE AND "LEARN NEW" FEATURE WAS SELECTED, NO ERRORS WERE PRESENT ANYMORE. THE ERROR 260 RELATES TO A DEFINED SOFTWARE BUG. THE TROUBLESHOOTING TO THIS SOFTWARE BUG IS EXPLAINED IN THE IFU FOR THE STATED ERROR MESSAGE. AS STATED, THE ERROR MESSAGE OCCURRED AFTER THE EXTERNAL EMERGENCY DEFIBRILLATION CAUSING PATIENT MOVEMENT ON THE TABLE. THE CUSTOMER COMPLETED THE CASE SUCCESSFULLY. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THIS DEVICE. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE PRIOR TO DISTRIBUTION.><INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. (B)(4). | 1
2629081 | IT WAS REPORTED THROUGH THE REVIEW OF PROGRAMMING HISTORY THAT A VNS PATIENT WAS FOUND TO BE AT UNINTENDED SETTINGS AT THE OFFICE VISIT OF (B)(6) 2007. THE PATIENT SETTINGS WERE FOUND TO BE 1/20/500/30/60/1.25/500. THE SETTINGS WERE CORRECTED AT THE OFFICE VISIT OF (B)(6) 2007 TO 1.25/20/500/30/5/1.25/500. NO SYSTEM OR NORMAL DIAGNOSTICS WERE RECORDED ON THE AVAILABLE PROGRAMMING HISTORY, ALTHOUGH A FAULTED DIAGNOSTIC IS SUSPECTED TO HAVE CONTRIBUTED TO THE RESET IN SETTINGS. AT THE MOMENT NO PATIENT ADVERSE EVENTS WERE REPORTED DUE TO THE UNINTENDED SETTINGS AND SETTINGS WERE CORRECTED AT THE FOLLOW-UP APPOINTMENT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2634131 | ATTEMPTING TO INTUBATE DURING A CODE BLUE EVENT - A #3 DISPOSABLE BLADE MADE AVAILABLE TO MD FOR INTUBATION BUT FAILED AFTER MULTIPLE ATTEMPTS, AIRWAY MANAGEMENT TAKEN OVER BY ANOTHER MD ATTEMPTS WITH THE DISPOSABLE #4 BUT COULD NOT GET A VIEW BECAUSE THE BLADE KEPT COLLAPSING AND THE VIEW WAS HORRIBLE KEPT THIS MD THEN WENT BACK TO THE #3 ORIGINALLY USED WHEN THE INITAL INTUBATION WAS ATTEMPTED. FINALLY, A NON-DISPOSABLE BLADE WAS USED WITH SUCCESSFUL INTUBATION. USE OF THIS DISPOSABLE PRODUCT ADDED TO DELAY WITH INTUBATION DUE TO THE SLENDER BLADE THE SUBGLOTTIS TISSUE FOLDED OVER THE BLADE OBSTRUCTING THE VIEW. | 0
2636092 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT FIND ANY CONCLUSIVE CAUSE OF THE ELEVATED RWBC CONTENT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND TRIMA ACCEL OPERATED AS INTENDED. WHILE THE TRIMA ACCEL SYSTEM IS TYPICALLY ROBUST AGAINST NUMEROUS DRAW PRESSURE ALERTS AND FLOW ADJUSTMENTS, IT IS POSSIBLE THAT THE HIGH NUMBER OF ALERTS AND ADJUSTMENTS THAT OCCURRED DISRUPTED THE STEADY STATE OF THE SYSTEM AND CONTRIBUTED TO THE HIGHER THAN EXPECTED RWBC CONTENT REPORTED FOR THIS PRODUCT. IT IS ALSO POSSIBLE THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2637720 | THE CUSTOMER IS UNABLE TO PROVIDE THE PATIENT'S INFORMATION. THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN AND INVESTIGATION.><THE CUSTOMER REPORTED THAT THEY RECEIVED THREE "HIGH PRESSURE IN CENTRIFUGE" ALARMS AND A ERYTHROCYTE SPILLOVER. NO VOLUME WAS COLLECTED. THIS HAPPENED AT THE BEGINNING OF THE COLLECTION. THE PATIENT'S INFORMATION IS NOT AVAILABLE. THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN AT THIS TIME. THIS REPORT IS BEING FILED DUE TO THE CUSTOMER'S FILING OF AN SAE REPORT.><THE CUSTOMER DECLINED TO PROVIDE THE PATIENT INFORMATION. PATIENT GENDER AND WEIGHT WERE OBTAINED FROM THE DATA LOGS FOR THIS PROCEDURE. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THR RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. AT 13 MINUTES INTO THE PROCEDURE, A NON RECOVERABLE `LEVEL SENSOR ERRO ' - TOO MANY CONSECUTIVE RESERVOIR DRAW ALARMS AND THE PROCEDURE ENDED. POSSIBLE CAUSES OF THE ALARM ARE A PARTIAL OBSTRUCTION TO THE ACCESS LINE, AN AIR BLOCK IN THE PLASMA LINE OR CHANNEL LINE PINCH CLAMPS THAT ARE OPEN. THE OPERATOR IS INSTRUCTED TO VERIFY THAT THE PLATELET AND PLASMA LINES ARE PRIMED WITH FLUID, THE TUBING IS CORRECTLY LOADED IN THE VALVES, THE CHANNEL LINE PINCH CLAMPS ARE OPEN AND THE PINCH CLAMP ON THE PLATELET ADDITIVE SOLUTION LINE IS CLOSED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. DEVICE WAS DISCARDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: BASED ON THE RDF ANALYSIS, POSSIBLE CAUSES OF THE LEVEL SENSOR ALARMS AND THE SPILLOVER INCLUDE, BUT ARE NOT LIMITED TO, THE TUBING NOT BEING FULLY OCCLUDED BY THE RBC VALVE DURING COLLECTION, A PLASMA LINE AIRBLOCK, AND/OR INCORRECT DONOR DATA ENTRY. CORRECTIVE/PREVENTIVE ACTION: DURING THE PROCEDURE, THE OPERATOR IS INSTRUCTED TO VERIFY THE PLATELET AND PLASMA LINES ARE PRIMED WITH FLUID, THE TUBING IS CORRECTLY LOADED IN THE VALVES, THE CHANNEL LINE PINCH CLAMPS ARE OPEN AND THE PINCH CLAMP ON THE PLATELET ADDITIVE SOLUTION LINE IS CLOSED. | 0
2639047 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2640192 | IT WAS REPORTED THAT DURING AN AFIB PROCEDURE, THE PATIENT HAD A SERIES OF COUGHING, WHICH RESULTED IN A MAP SHIFT/ ADVANCE CATHETER LOCALIZATION (ACL) RE-INITIALIZATION. AFTER THE RE-INITIALIZATION ERROR 8 APPEARED. THE FOLLOWING RF THERAPY RESULTED IN SATURATION OR DISAPPEARANCE OF BODY SURFACE ECG LEADS AND INTRACARDIAC SIGNALS ON THE CARTO 3 AND RECORDING SYSTEMS. AFTER THE ABLATION THE PATIENT INTERFACE UNIT WAS REBOOTED, THE ERROR WAS CLEARED AND THE ECG SIGNALS RETURNED. THEY ATTEMPTED A SECOND ABLATION WITH THE SAME RESULT. THE USER REPLACED THE LASSO CATHETER AND THE ABLATION CABLES AND ENSURED THE ECG LEADS WERE CONNECTED WELL TO THE PATIENT. HOWEVER, THE CASE ENDED UP BEING COMPLETED CONVENTIONALLY WITHOUT THE USE OF THE CARTO 3 SYSTEM. NO PATIENT INJURY WAS REPORTED. INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. CONCOMITANT BWI PRODUCTS USED DURING THE PROCEDURE: EZ STEER THERMOCOOL NAV 4MM CATHETER: MODEL #: D-1292-05-S, LOT #: UNKNOWN. LASSO W/ AUTO ID CATHETER: MODEL #: D-1220-81-S, LOT #.: UNKNOWN. C3 EXTERNAL REFERENCE PATCH: MODEL #: D-1283-02, LOT #.: UNKNOWN. (B)(4).><(B)(4). IT WAS REPORTED THAT DURING AN AFIB PROCEDURE, THE PATIENT HAD A SERIES OF COUGHING WHICH RESULTED IN A MAP SHIFT/ ADVANCE CATHETER LOCALIZATION (ACL) RE-INITIALIZATION. AFTER THE RE-INITIALIZATION ERROR 8 APPEARED. THE FOLLOWING RF THERAPY RESULTED IN SATURATION OR DISAPPEARANCE OF BODY SURFACE ECG LEADS AND INTRACARDIAC SIGNALS ON THE CARTO 3 AND RECORDING SYSTEMS. AFTER THE ABLATION THE PATIENT INTERFACE UNIT WAS REBOOTED, THE ERROR WAS CLEARED AND THE ECG SIGNALS RETURNED. THEY ATTEMPTED A SECOND ABLATION WITH THE SAME RESULT. THE USER REPLACED THE LASSO CATHETER AND THE ABLATION CABLES AND ENSURED THE ECG LEADS WERE CONNECTED WELL TO THE PATIENT. HOWEVER THE CASE ENDED UP BEING COMPLETED CONVENTIONALLY WITHOUT THE USE OF THE CARTO 3 SYSTEM. NO PATIENT INJURY WAS REPORTED. THE CARTO 3 SYSTEM ASSOCIATED WITH THE REPORTED COMPLAINT WAS BY BWI SERVICE ENGINEER. THE PATIENT INTERFACE UNIT (PIU) WAS REPLACED AND THE FAULTY PIU SENT TO THE MANUFACTURER. THE MANUFACTURER REPORTED THAT THE PIU WAS FOUND INOPERABLE DUE TO FAULTY BP CARD. THE DEFECTIVE CARD CAUSED THE ERROR 8. IT WAS FOUND THAT DIODE D1000 FAILED. IN ADDITION D11, D12, D53, D55 WERE DEFECTIVE AS WELL. DEFECTIVE DIODES OF BP CARD CAUSED THE ECG SIGNALS DISAPPEARANCE ON THE START OF ABLATION. ALL PARTS WERE SENT TO SUBCONTRACTOR FOR REPAIR. REGARDING THE ACL ISSUE, THE PATIENT COUGHS RESULTED IN THE ACL MAP SHIFT RE-INITIALIZATION. THIS IS A NORMAL SYSTEM BEHAVIOR DUE TO PATIENT MOVEMENT. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THIS DEVICE. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE PRIOR TO DISTRIBUTION. | 1
2645713 | THE ENTIRE EHR SYSTEM WAS UNAVAILABLE. DOCTORS COULD NOT FIND THEIR PTS, MEDICATIONS WERE DELAYED, AND IN AT LEAST TWO CASES, POTASSIUM LEVELS WERE AT HIGH RISK LEVELS, EITHER TOO HIGH OR TOO LOW. NO ONE KNEW, AND NO TREATMENTS WERE OFFERED BECAUSE NO ONE KNEW. THERE WAS NOT ANY WARNING THAT A PLANNED DOWNTIME WAS SCHEDULED. IT THEN EVOLVED INTO AN UNPLANNED DOWNTIME, CAUSING DISRUPTION AND UNSAFE CONDITIONS IN THE HOSPITAL AFFECTING SEVERAL HUNDRED PTS. | 1
2645718 | CRITICALLY ILL PT WITH MULTI-ORGAN DYSFUNCTION COMPLICATED BY URINARY TRACT INFECTION AND POSSIBLE OSTEOMYELITIS, WHO IS IN NEED OF MAJOR SURGERY OF THE HEART. THE EMR CONTAINS INNUMERABLE ORDERS, TREATMENTS, MEDICATIONS, AND MEDICATION ADMINISTRATION RECORDS THAT ARE OBFUSCATED WITH MEANINGFULLY IMMATERIAL VERBIAGE TO THE EXTENT THAT THE DEVICE IS POORLY USABLE AND THE COURSE OF THE PT IS IMPOSSIBLE TO TRACK AND COMPREHEND. WHAT HAPPENED IS THAT THE INFECTIOUS DISEASE CONSULTANT STOPPED ANTIBIOTICS, WHICH WAS REINFORCED WITH PROGRESS NOTES. ANOTHER CONSULTANT'S PARAPROFESSIONAL MISSED OR IGNORED THAT THE ANTIBIOTICS WERE DISCONTINUED AND ORDERED THEM, PUTTING THE PT AT RISK BY JEOPARDIZING SCHEDULED BONE BIOPSY CULTURES, WHICH, IN TURN, WOULD JEOPARDIZE THE DECISION TO OPERATE AND THE TIMING OF THE SURGERY. THE CASE IS AN EXAMPLE OF THE DYSFUNCTIONAL EVALUATION AND MANAGEMENT WORK FLOWS THAT ARE PART AND PARCEL OF CPOE AND EHR DEVICES WHEN USED FOR COMPLEX CASES. IT IS IMPOSSIBLE TO DETERMINE THE COURSE OF EVENTS AND RECONCILE ORDERS WITH CONSULTANT'S PROGRESS NOTES BECAUSE THERE IS NOT A USER FRIENDLY CONTEMPORANEOUS LISTING OF THE ORDERS AS THEY WERE ENTERED AND BY WHOM; AND THE DISCONTINUATION, OF ORDERS, DETAILS NOT LIST THE DOCTOR WHO ENTERED THAT; AND THE MEDICATION ADMINISTRATION RECORDS ARE PARTICULARLY UNUSABLE DUE TO THE MEANINGFULLY USELESS EXPANSIVE GRID LAYOUT AND ABSENCE OF A SEARCH FUNCTION. AS SUCH, THIS DEVICE IS DANGEROUS, SERVING AS AN IMPEDIMENT TO SAFE AND EFFECTIVE MEDICAL CARE. | 1
2646274 | PER BEDSIDE RN: "I HAD AN ORDER FOR A 12 LEAD EKG. I USED THE CABLES IN THE ROOM. AND WHEN I PLUGGED IN THE 12 LEAD EKG, THE SCREEN WENT BLANK. I SWITCHED OUT THE 12 LEAD CABLES, THE REGULAR EKG CABLES AND THE CORD FROM THE MONITOR TO THE CABLES AND THE SCREEN STILL WENT BLANK. I FINALLY HAD TO USE THE PORTABLE 12 LEAD EKG MACHINE AND I WAS ABLE TO OBTAIN A 12 LEAD." CLINICAL ENGINEERING ANALYSIS: THIS HAS BEEN A RECURRING PROBLEM. WHEN USING THE SOLAR 8000 IN THE 12 LEAD EKG MODE, IF THE SIGNAL IS POOR DUE TO ELECTRODE IMPEDANCE OR MOVEMENT ARTIFACT, THE EKG SECTOR ON THE DISPLAY GOES BLANK AND THE ERROR MESSAGES (IF THERE ARE ANY) DO NOT GUIDE THE USER TO THE PROBLEM. THIS RESULTS IN A CONSIDERABLE AMOUNT OF TIME WASTED TRYING TO TROUBLESHOOT THE SYSTEM DURING WHICH TIME THE PATIENT GOES ESSENTIALLY UNMONITORED FOR EKG. FREQUENTLY THE REASON FOR DOING THE 12-LEAD EKG IN THE FIRST PLACE IS DUE TO SOME DETERIORATION IN THE PATIENT'S CONDITION, SO THIS IS A CRITICAL TIME TO MAINTAIN MONITORING. THIS PROBLEM HAS BEEN ONGOING SPORADICALLY SINCE WE STARTED USING THE SOLAR 8000I MONITORS 4 YEARS. THE DEVICE SERIAL NUMBERS ARE FROM THE SAME BATCH. THE PROBLEM IS NOT ISOLATED TO A PARTICULAR MONITOR AND HAS BEEN OBSERVED ON SEVERAL DIFFERENT MONITORS.======================MANUFACTURER RESPONSE FOR BEDSIDE PATIENT MONITOR, GE (PER SITE REPORTER).======================INVESTIGATION IS ONGOING. GE SUGGESTS THE ISSUE IS CAUSED BY POOR SKIN PREP AND/OR PATIENT MOVEMENT. | 1
2646275 | STAFF DESCRIPTION: PATIENT DEVELOPED ST DEPRESSIONS ON THE MONITOR LEADS II & V1. CALLED RESIDENT AND WHILE OBTAINING A 12-LEAD EKG, WHEN V2-V6 LEADS CONNECTED THE EKG TRACING DISAPPEARED OFF THE SCREEN AND UNABLE TO OBTAIN A 12-LEAD STAT. CHANGED EKG TRUNK CABLE, PRIMARY 5 LEAD CONNECTION, V2-V6 CONNECTION CABLE CHANGED, AND PATIENT EKG PATCHES (ELECTRODES) CHANGED X2 AND RE-PREPPED TWICE BEFORE 12-LEAD OBTAINED. IT TOOK 10 MINUTESBETWEEN THE START OF EKG CHANGES AND WHEN 12-LEAD EKG OBTAINED.CARDIOLOGY CONTACTED FOR CONSULT BY PRIMARY TEAM AND CARDIAC ECHO DONE TO R/O MI/ ISCHEMIA.CLINICAL ENGINEERING ANALYSIS: THERE SEEMS TO BE SEVERAL FACTORS INVOLVED WITH THIS PROBLEM, INCLUDING ELECTRODE IMPEDANCE AND THE MONITOR'S RESPONSE TO HIGH IMPEDANCE / NOISY SIGNALS. STAFF REPORT INDICATES THEY PREPPED THE SKIN PRIOR TO PLACING ELECTRODES AND THE 5 LEAD MONITORING WAS WORKING OK BEFORE ATTEMPTING THE 12-LEAD. THIS PROBLEM HAS BEEN ONGOING SPORADICALLY SINCE WE STARTED USING THE SOLAR 8000I MONITORS 4 YEARS. THE DEVICE SERIAL NUMBERS ARE FROM THE SAME BATCH. THE PROBLEM IS NOT ISOLATED TO A PARTICULAR MONITOR AND HAS BEEN OBSERVED ON SEVERAL DIFFERENT MONITORS.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, SOLAR (PER SITE REPORTER).======================INVESTIGATION ONGOING. GE SUGGESTS PROBLEM IS DUE TO EXCESSIVE ARTIFACT/NOISE IN EKG SIGNAL. | 1
2651553 | A CUSTOMER REPORTED THEIR ADC BLOOD GLUCOSE METER WAS NOT RETAINING THE DATE AND TIME. THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. (B)(4). | 1
2654586 | DURING CRYOABALTION PROCEDURE OF A PARA SPINAL TUMOR, SYSTEM FAILED. SCREEN FLASHED AND ERROR MESSAGE DISPLAYED. PER OPERATORS MANUAL, MACHINE SHUT DOWN AND RESTARTED SEVERAL TIMES TO CLEAR ERROR WITH NO SUCCESS. CONTACTED TECH SUPPORT, WHO GAVE SUGGESTIONS ON CLEARING THE ERROR. THIS ALSO DID NOT WORK. AFTER APPROXIMATELY 30 MINUTES THE CASE WAS ABORTED WITHOUT BEING COMPLETED. NO PATIENT INJURY REPORTED, CASE RESCHEDULED. SERVICE TECHNICIAN SENT OUT TO HOSPITAL ON (B)(4) 2012 TO EVALUATE CRYO-SL. SERVICE TECH TESTED UNIT FOR 45 MINUTES AND COULD NOT DUPLICATE PROBLEM OR ERROR. DECISION MADE TO REPLACE TEMPERATURE AND I/O BOARD TO TRY AND ALLEVIATE INTERMITTENT FAULT. | 0
2655420 | ANALYSIS OF THE RETURNED HANDHELD COMPUTER AND THE ASSOCIATED SOFTWARE HAS BEEN COMPLETED. THE CRACKED SCREEN WAS CONFIRMED THAT RESULTED IN THE HANDHELD BEING UNRESPONSIVE. ANALYSIS DETERMINED THE EVENT WAS LIKELY RELATED TO USER MISHANDLING OR IMPROPER STORAGE. NO ISSUES WERE OBSERVED WITH THE SOFTWARE.><IT WAS INITIALLY REPORTED THAT THE COMPANY REPRESENTATIVE'S HANDHELD HAD A CRACKED AND THE SCREEN IS UNABLE TO BE SEEN. IT IS UNKNOWN HOW THE SCREEN GOT CRACKED. HE PULLED THE HANDHELD OUT OF HIS BAG AND FOUND IT CRACKED. THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. PRODUCT ANALYSIS IS PLANNED BUT HAS NOT BEEN COMPLETED. . | 0
2655804 | DISPLAY BOARD FAILED TO POWER ON. OUR EVAL DETERMINED THAT THE INVERTER BOARD SHORTED OUT CAUSING THE DISPLAY BOARD TO NOT POWER ON. THE EXACT ROOT CAUSE CANNOT BE DETERMINED. THERE WERE NO OBVIOUS MATERIAL DEFECTS OR MFG FLAWS THAT WOULD HAVE CONTRIBUTED TO THIS TYPE OF FAILURE. | 1
2656395 | A HEALTH PROFESSIONAL REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. (B)(4).><(B)(4). | 1
2659986 | ASCOM BECAME AWARE ON (B)(4) 2012 OF THE FOLLOWING EVENT: DURING INTERNAL SOFTWARE VERIFICATION FOR A NEW BUT NOT YET RELEASED VERSION OF SOFTWARE, ASCOM BECAME AWARE OF A MEMORY LEAK SOFTWARE ANOMALY. THE SOFTWARE MODULE IN QUESTION IS CURRENTLY USED IN BOTH THE ASCOM CARDIOMAX AND MMG SOFTWARE. WHEN ASSIGNMENTS ARE CHANGED FOR A LOCATION, THE LOCATION CONDITIONS ARE STORED IN A MEMORY CACHE; HOWEVER, THESE ENTRIES ARE NOT DELETED. EVENTUALLY THE MEMORY CACHE BECOMES FULL WHICH RESULTS IN A MODULE REBOOT. DURING THIS TIME, NO INFORMATION IS COMMUNICATED TO ANY DISPLAY DEVICES. TO ASCOM'S KNOWLEDGE, NO ADVERSE EVENT HAS BEEN ASSOCIATED WITH THIS SOFTWARE ANOMALY. BECAUSE THESE DEVICES ARE SECONDARY ALARM NOTIFICATION SYSTEMS ASCOM IS NOT CERTAIN THIS IS A REPORTABLE EVENT, BUT HAS CHOSEN TO FILE THIS MDR ANYWAY. ASSUMING NORMAL OPERATION, I.E. IF 1000 LOCATIONS ARE MODIFIED THREE (3) TIMES PER DAY AND EACH LOCATION HAS A SIMPLE CONDITION, THE MODULE IS EXPECTED TO OPERATE FOR SIX (6) TO 12 MONTHS BEFORE REBOOTING. THE FOLLOWING SOFTWARE IS AFFECTED: ASCOM CARDIOMAX: CATALOG #: FE3-G1ABAB, (B)(4) AND EARLIER. ASCOM MOBILE MONITORING GATEWAY (MMG): CATALOG#: FE3-D1ABAA AND FE3-D1ABAB, (B)(4). A SOFTWARE FIX HAS BEEN GENERATED, AND VERIFIED AND VALIDATED. NEW SOFTWARE VERSIONS WILL BE IMPLEMENTED AT THE CUSTOMER'S SITES, AND ARE PLANNED TO BE: ASCOM CARDIOMAX: (B)(4), ASCOM MMG: (B)(4). | 1
2661468 | THE DEVICE FAILED TO HOLD PRESSURE. A LEAK WAS FOUND WHERE THE LINE MEETS THE INFLATION SEAL. EIGHT DEVICES WERE PULLED AND CHECKED. SIX DEVICES OF THE SAME LOT LEAKED AT THE SAME SPOT.======================MANUFACTURER RESPONSE FOR INFUSABLE PRESSURE INFUSURE, INFUSABLE PRESSURE INFUSURE (PER SITE REPORTER).======================NONE. MANUFACTURER REPRESENTATIVE COMMENTED THAT "THEY MOVED THEIR MANUFACTURING PLANT TO A HIGHER ELEVATION WHICH CAN CAUSE THE GLUE TO COME APART". HE ALSO ADDED THAT "AS A RESULT OF THAT, THE GLUE WAS CHANGED TO A DIFFERENT BRAND." | 0
2661524 | AGFA IS SUBMITTING THIS SUPPLEMENTAL REPORT TO PROVIDE THE FOLLOWING: A SUMMARY OF THE ROOT CAUSE INVESTIGATION RESULTS. THE RADIOLOGIST'S DICTATION TEXT IS IN AGFA IMPAX, BUT SHOWS THE TEXT TRANSFER AS INCOMPLETE. AGFA HAS DETERMINED IT IS POSSIBLE WHEN BOOKMARKS ARE DELETED OR MOVED WITHIN THE STUDY REPORT, THE REPORT BODY CONTENT, WHEN NEEDED TO BE SENT TO THE 3RD PARTY (B)(4), IS NOT BETWEEN THE BEGINNING BOOKMARKS AND THE END BOOKMARKS AS DEFINED IN (B)(4). SINCE BOOKMARKS ARE NOT VISIBLE TO THE END USER, A BOOKMARK MAY BE COPIED/CUT/PASTED TO ANOTHER LOCATION WITHOUT THE END USER'S KNOWLEDGE OF THE OCCURRENCE. IT IS ALSO POSSIBLE THAT THE REPORT BODY CONTENT CAN BE MISSING BECAUSE THE END USER UNKNOWINGLY ENTERED THE TEXT OUTSIDE THE DEFINED BOOKMARKS. AGFA HAS NOT BEEN ABLE TO REPRODUCE THIS PROBLEM DURING ITS INVESTIGATION AND AS A RESULT, HAS NOT CONFIRMED A SOFTWARE QUALITY ISSUE WITHIN THE IMPAX CLIENT THAT COULD MOVE OR DELETE BOOKMARKS. AGFA IS PERFORMING FURTHER INVESTIGATION OF THE POTENTIAL QUALITY ISSUE. WHAT IS BEING DONE TO ADDRESS THE REPORTED ISSUE? CORRECTIVE ACTION IS UNDERWAY TO DEFINE THE BOOKMARKS IN THE TEMPLATE SO THEY CANNOT BE DELETED OR MOVED. THE END USER WILL BE UNABLE TO TYPE TEXT OUTSIDE THE BOUNDARIES DEFINED BY THE BOOKMARKS. BY PLACING THE BOOKMARKS IN PROTECTED SECTIONS, ONLY CERTAIN AREAS OF THE REPORT ARE EDITABLE AND THE PROTECTED SECTIONS CANNOT BE ALTERED BY THE END USER. BOOKMARKS CAN BE PROTECTED BY USING THE (B)(4) FUNCTIONALITY OF PROTECTED SECTIONS, BUT THE CURRENT ALGORITHM FOR PROTECTED SECTIONS IS NOT COVERING ALL SCENARIOS AND PROTECTING THE TEXT SECTIONS WITHIN THE REPORT. AGFA IS WORKING TO IMPLEMENT A NEW ALGORITHM WHICH NO LONGER USES THE BEGINNING AND ENDING BOOKMARKERS. THE NEW ALGORITHM WILL DEFINE THE PROTECTED SECTIONS IN THE TEMPLATE BY USING THE (B)(4) FUNCTIONALITY TO DEFINE THEM. THE NEW TEMPLATES WILL NO LONGER CONTAIN HIDDEN BOOK MARKERS, THUS DYNAMICALLY PROTECTING APPLICABLE SECTIONS. THE INITIAL REPORTER IN THIS EVENT IS (B)(6). THE AFFECTED SITE'S REPORTING GATEWAY HAS BEEN CORRECTED AND THE REPORT TEMPLATE MODIFIED SO THE DELETION OF BEGINNING AND ENDING BOOKMARKS IS NOT AS ACCESSIBLE. THE EXISTING REPORT TEMPLATE PROTECTION MECHANISM HAS BEEN USED FOR THIS MODIFICATION AND A MODIFIED TEMPLATE HAS BEEN ROLLED OUT AT THE SITE. AS AN ADDITIONAL CORRECTION, THE OUTBOUND INTERFACE AT THE SITE HAS BEEN UPDATED TO CHECK FOR SPECIFIC START TEXT IN DICTATION. (B)(4).><ON (B)(6) 2012, THE CUSTOMER CONTACTED AGFA TO REPORT, A RADIOLOGIST HAD DICTATED A STUDY AND ALL OF THE DICTATION TEXT WAS IN IMPAX, HOWEVER, ONLY THE FIRST TWO SENTENCES WERE TRANSFERRED TO (B)(4)). THE INCOMPLETE TRANSFER OF THE PATIENT REPORT WAS NOT REALIZED UNTIL TWO DAYS AFTER THE EVENT OCCURRED. THE PATIENT DID NOT SUFFER HARM DUE TO THE MISSING REPORT INFORMATION, BECAUSE THE TREATING PHYSICIANS WERE CONSULTED AT THE ONSET OF THE INCIDENT AND RECEIVED ALL OF THE RELEVANT INFORMATION ABOUT THE INCIDENT AND THE PATIENT RECEIVED ADEQUATE TREATMENT AT THIS TIME. AFTER THE EVENT OCCURRED AND THE PATIENT TREATED, THE RADIOLOGIST DICTATED THE STUDY. IT WAS NOT UNTIL TWO DAYS LATER, THE RADIOLOGIST NOTICED, THE PATIENT REPORT WAS MISSING INFORMATION AND THE DICTATION TEXT TRANSFER WAS INCOMPLETE. THE RADIOLOGIST KNEW THIS BECAUSE HE/SHE WAS THE ONE WHO DICTATED THE STUDY AND KNEW WHAT HAD OCCURRED WITH THE PATIENT. THE PATIENT'S REPORT WAS COMPLETE AND ACCURATE IN AGFA IMPAX, HOWEVER, THIS SITE HAS ANOTHER (B)(4) AND THE PATIENT INFORMATION IN THAT RIS WAS INCOMPLETE. THE AGFA HEALTHCARE RIS IS ONLY USED AS A REPORTING GATEWAY (AGFA'S RIS IS EMBEDDED IN (B)(4) TO ENABLE SPEECH RECOGNITION), THEN THE REPORT (BODY) IS SENT TO THE CUSTOMER'S MAIN RIS IN THIS INCIDENT (B)(4). EVEN THOUGH THERE WAS NO REPORTED HARM TO THE PATIENT DUE TO THE MISSING REPORT, THERE IS A PROBABILITY THAT SERIOUS INJURY COULD OCCUR TO OTHER PATIENTS IF THE EVENT WERE TO RECUR. AN INCOMPLETE RIS REPORT MAY BE OBVIOUS TO A READER; HOWEVER, THE MISSING INFORMATION MAY NOT BE OBVIOUS AND MAY CONTAIN CERTAIN INFORMATION THAT IS CRITICAL TO THE SUCCESSFUL OUTCOME TO A PATIENT, THEREFORE, THIS EVENT IS UNDER INVESTIGATION TO DETERMINE ROOT CAUSE AND CORRECTIVE ACTION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED AT A LATER DATE. .><(B)(4). | 1
2662540 | DR (B)(6) IMPLANTED THE ARROW-LOCK FIXATION DEVICE FOR A NON-ELECTIVE PROXIMAL PHALANX FRACTURE TO STABILIZE THE BONE. THE DATE OF THE PROCEDURE WAS (B)(6) 2012. DR (B)(6) PERFORMED AN EXPLANT ON (B)(6) 2012, AFTER THE SURGEON REVIEWED POST OPERATIVE X-RAYS AND DETERMINED THE PROCEDURE SHOWED INSTABILITY "FLOPPY TOE". UPON REMOVAL OF THE ARROWHEAD DEVICE THE SURGEON USED A K-WIRE, WHICH IS HIS NORMAL PROCEDURE. THE SURGEON IMPLANTED THE ARROW-LOK DEVICE OPPOSITE OF THE PRESCRIBED IMPLANTATION. THE ARROW-LOK FIXATION DEVICE WAS IMPLANTED AS FOLLOWS INTO THE PT: THE 2.5 MM DIA ARROWHEAD WAS IMPLANTED IN THE MIDDLE PHALANX AND THE 3.5 MM DIA ARROWHEAD TIP WAS IMPLANTED IN THE PROXIMAL PHALANX. DR (B)(6) COMMUNICATED IN HIS INTERVIEW THAT HE THOUGHT THE LARGER ARROW WENT PROXIMALLY BECAUSE THE PROXIMAL PHALANX IS LARGER. PRIOR TO THE SURGERY ON (B)(6) 2012, DR (B)(6) CONSULTED THE SURGEON DESIGNER, DR (B)(6) FOR INFO ON THE SURGICAL TECHNIQUE. DR (B)(6) WAS PROVIDED RADIOGRAPHS, SALES BROCHURE WITH RECOMMENDED PROCEDURE, AND INSTRUCTIONS FOR USE. PRESENT DURING THE IMPLANT WAS (B)(6), REP FOR ADD'L PROCEDURAL SUPPORT. MR HEALY RECOMMENDED THE USE OF A "C" ARM - THE SURGEON DECLINED THIS RECOMMENDATION. MR HEALY STATED IT WAS DIFFICULT TO OBSERVE THE IMPLANT PROCEDURE AND WAS UNABLE TO VALIDATE THE PROPER INSTALLATION OF THE ARROW-LOK FIXATION DEVICE. ARROWHEAD REVIEWED BOTH THE SALES BROCHURE AND THE INSTRUCTIONS FOR USE, TO VERIFY INSTRUCTIONAL ACCURACY. BOTH PIECES OF DOCUMENTATION CALL TO ATTENTION PROPER IMPLANTATION OF THE ANGLED ARROW-LOK FIXATION DEVICE. THE SALES BROCHURE INDICATES DIRECTIONAL ORIENTATION BOTH IN WRITTEN FORM AND AS A VISUAL ILLUSTRATION. DURING THE CONFERENCE CALL DR (B)(6) STATED THE ARROW-LOK PRODUCT WAS GOOD AND HE LIKED THE FIXATION ASPECT OF THE DEVICE. HE ALSO WOULD CONSIDER USING THE DEVICE IN THE FUTURE WITH ADD'L TRAINING. | 0
2662704 | THE HOSP REPORTED A PT EXPERIENCED DIFFICULTY BREATHING ONE HOUR AFTER A HEAT AND MOISTURE EXCHANGE FILTER (HMEF) WAS REPLACED. THE PT WAS UNDER MECHANICAL VENTILATION IN BI-LEVEL POSITIVE AIRWAY PRESSURE (BIPAP) MODE AND BEGAN TO HAVE DIFFICULTIES BREATHING. AUSCULTATION SHOWED NO LUNG PATHOLOGY TO CAUSE THE DIFFICULTY BREATHING SO THE HEALTHCARE PROFESSIONAL THOUGHT THE DIFFICULTY WAS DUE TO ANXIETY AND SUBSEQUENTLY RETURNED THE PT TO BED AND ADMINISTERED 25 MG OF TRANXILLIUM. FIFTEEN MINUTES LATER, THE PT BECAME SWEATY AND CYANOTIC. THE RESPIRATOR COULD NOT DELIVER BIPAP VENTILATION VOLUME AND SWITCHED TO INTERMITTENT POSITIVE PRESSURE VENTILATION (IPPV) MODE. THE PT WAS INDUCED TO SLEEP. THE HEALTHCARE PROFESSIONAL THOUGHT, THE DIFFICULTY BREATHING WAS DUE TO AN OBSTRUCTION OF THE TRACHEOSTOMY TUBE DUE TO SECRETIONS BECAUSE THE PT COULD NOT BE VENTILATED USING A MANUAL RESUSCITATOR. TREATMENT FOR BRONCHOSPASM WAS BEGUN BECAUSE THE PT'S CHEST DID NOT LIFT AND NO BREATH SOUNDS WERE HEARD. .. | 0
2663250 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2663518 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR (B)(6) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. (B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2663838 | DURING A PROGRAMMING HISTORY REVIEW ON (B)(4) 2012, IT WAS NOTED THAT A PARTIAL PROGRAMMING EVENT OCCURRED ON (B)(6) 2012. A FAULTED SYSTEMS DIAGNOSTIC OCCURRED EARLIER IN THE SAME APPOINTMENT; HOWEVER, IT IS UNCLEAR IF THE PATIENT'S SETTINGS WERE CHANGED AS A RESULT OF THE PARTIAL PROGRAM OR THE FAULTED DIAGNOSTIC AS THERE WAS NO INTERROGATION PERFORMED BETWEEN THE FAULTED DIAGNOSTIC AND THE PARTIAL PROGRAMMING EVENT. THIS EVENT RESULTED IN A CHANGE IN SETTINGS THAT WAS NOT CORRECTED AT THIS VISIT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS ISSUE. ANALYSIS OF PROGRAMMING HISTORY. | 0
2664355 | GE HEALTHCARE, AKA GE MEDICAL, OPTIMA XR220AMX PORTABLE X-RAY MACHINES. MULTIPLE PROBLEMS, MULTIPLE MACHINES FOUR AND MULTIPLE DATES. ON (B)(6) 2012, SHORT CIRCUIT IN DEVICE CAUSING USER INTERFACE SCREEN TO BUBBLE AND REQUIRING GE TO REPLACE SCREEN AND POWER SUPPLY. MULTIPLE USER INTERFACE LOCK UPS SINCE PRODUCT WAS PUT INTO USE ON (B)(6) 2012. THESE LOCK UPS REQUIRE USER TO REBOOT SYSTEM BEFORE PT IMAGES CAN BE ACQUIRED. ON (B)(6) 2012, DRIVE WHEEL FELL OFF OF AXLE. THE SIDE COVER WAS THE ONLY THING THAT KEPT WHEEL FROM FULLY FALLING OFF. GE REPLACED WHEEL AND COVER AND FOUND HARDWARE TO NOT BE TIGHTENED. ON (B)(6) 2012, POWER PLUG SPARKING INTERNALLY. GE FOUND WIRES TO BE LOOSE INSIDE OF POWER PLUG. ON (B)(6) 2012, ANOTHER MACHINE POWER PLUG SPARKING INTERNALLY. GE FOUND WIRES TO BE LOOSE INSIDE OF POWER PLUG. ON (B)(6) 2012, SCREEN TURNING PINK, UNABLE TO USE. GE FOUND CHARGER PLUG APPEARED TO HAVE ELECTRICALLY VAPORIZED INTERNALLY FROM POOR WIRING CONNECTION AND REPLACED POWER CORD PLUG, TOP LEVEL ASSEMBLY, AND GRID ASSEMBLY. | 1
2664496 | PT ON TELEMETRY COMPLAINED OF NOT FEELING WELL. REQUIRED RAPID RESPONSE TEAM EVALUATION. PHILIPS ENGINEER WAS PRESENT WHILE TELEMETRY WAS NOT TRANSMITTING TO CENTRAL STATION RESULTING IN INABILITY TO DOCUMENT RHYTHM STRIPS DURING THE EVENT. | 1
2664983 | INTERNAL PROGRAMMING HISTORY WAS REVIEWED AND IT WAS NOTED THAT DIAGNOSTIC DATA CONFIRMED THAT A LEAD TEST WAS INTERRUPTED ON (B)(6) 2009, WHICH INADVERTENTLY CHANGED THE PATIENT'S SETTINGS TO LEAD TEST PARAMETERS AND NO INTERROGATION FOLLOWED IN THIS VISIT. IT WASN'T UNTIL (B)(6) 2009, THAT THE OFF TIME SETTING IN QUESTION WAS FOUND AND CORRECTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2665694 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING THAT THEIR COULTER ACT DIFF ANALYZER GAVE ERRATIC CBC RESULTS FOR A (B)(6) PATIENT DRAWN VIA HEEL STICK. THE PATIENT WAS REDRAWN VIA HEEL STICK THE SAME DAY AND THE FOLLOWING DAY, ANOTHER HEEL STICK WAS PERFORMED WITH RESULTS THAT WERE RELEASED AS CORRECT. THE RESULTS PROVIDED BY THE CUSTOMER ARE SHOWN IN THE ATTACHMENT. THERE WAS NO AFFECT TO PATIENT TREATMENT WITH REGARD TO THIS EVENT. QC RAN WITHIN RANGE BEFORE THE INCIDENT AND THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATION. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE THE DAY OF THE EVENT AND INSPECTED THE INSTRUMENT. FSE DID NOT FIND ANY ISSUES WITH THE INSTRUMENT AND PERFORMED A REPRODUCIBILITY TEST, WHICH PASSED WITHIN SPECIFICATIONS. PER CUSTOMER, ALL OTHER SAMPLES RAN WITHOUT ANY PROBLEMS. THE CAUSE OF THIS EVENT IS LIKELY SAMPLE RELATED. | 0
2665998 | IT WAS REPORTED THAT, DURING AN EPISODE OF ASYSTOLY, THERE WAS NO "SEVERE" ALARM TRIGGERED (AUDIO NOR VISUAL). THE EKG WAS CONFIRMED TO THE ARTERIAL PRESSURE. IT WAS ALSO REPORTED THAT IT SEEMS THAT ON THIS PARAMETER THER'S NO "SEVERE" ALARM. THE INVESTIGATION HAS BEEN STARTED BUT IS NOT FINISHED. RESULTS WILL BE REPORTED IN THE FOLLOW-UP REPORT. | 0
2667005 | BREAKAGE OF THE CORD DURING IMPLANTATION. | 0
2670336 | THE CTS WENT DOWN WHICH CAUSED LOSS OF MONITORING FOR 4 PTS FOR 15 MINUTES. ALTERNATE PT MONITORING WAS AVAILABLE BUT IT TOOK THE USER 15 MINUTES TO START MONITORING. THERE WAS NO PT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2670835 | OUR AFFILIATE HAS REPORTED TO US THAT DURING AN ARTHROSCOPIC ELBOW PROCEDURE WITH THE USE OF AN FMS DUO PUMP FOR FLUID MANAGEMENT THERE WERE SOME ISSUES WHICH LEAD TO THE PATIENT'S ELBOW BECOMING INFLATED TO THE POINT THAT THE SURGEON WENT TO AN OPEN PROCEDURE TO COMPLETE THE REPAIR. IT APPEARS THAT THE "DAY SET TUBING CHAMBER" BECAME OVER FILLED; THIS WAS IMMEDIATELY REPLACED, AND AFTER APPROXIMATELY 8 MINUTES INTO USE THE "INFLOW" SUDDENLY INCREASED APPROXIMATELY TO A READING OF 200 ON THE UNIT'S CONSOLE, AS REPORTED BY THE SURGEON, AT WHICH POINT THE PATIENT'S ELBOW SWELLED CONSIDERABLY. THE PUMP WAS POWERED OFF AND THE SURGEON WENT OPEN FOR THE CONCLUSION. POST-OPERATIVELY, THE SURGEON STATES THAT THERE WERE NO PATIENT COMPLICATIONS; NO COMPARTMENT SYNDROME, NO FURTHER ISSUES. THE PUMP WAS SENT TO A SERVICE CENTER FOR EVALUATION; THE SERVICE CENTER COULD NOT DUPLICATE THE REPORTED DEVICE BEHAVIOR, NOR COULD THEY FIND ANYTHING WRONG WITH THE DEVICE IN GENERAL. THE PUMP HAS BEEN RETURNED TO THE USER FACILITY. ALSO SEE ASSOCIATED MDRS 1221934-2012-00190, 1221934-2012-00191, 1221934-2012-00192 AND 1221934-2012-00194 MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT.><THE COMPLAINT DEVICE IS NOT BEING RETURNED, WHICH PRECLUDES CONDUCTING AN EVALUATION; HOWEVER, A BATCH RECORD REVIEW HAS BEEN CONDUCTED TO DETERMINE IF THERE WERE ANY INTERNAL PROCESSING ISSUES WHICH WOULD HAVE CONTRIBUTED TO THE NATURE OF THE PRODUCT COMPLAINT. OUR RESULTS INDICATE THAT THIS BATCH OF PRODUCT WAS PROCESSED WITHOUT INCIDENT AND THEREFORE THERE IS NO INTERNALLY ASSIGNABLE CAUSE FOR THE REPORTED PROBLEM. (B)(4). WE CANNOT TELL ANYTHING FROM THIS; WE CANNOT DISCERN A ROOT OR UNDERLYING CAUSE FOR THE REPORTED FAILURE MODE. AT THIS POINT IN TIME, NO FURTHER ACTION IS WARRANTED. HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FORTHCOMING INFORMATION RECEIVED THAT IS PERTINENT AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
2671019 | THE CUSTOMER REPORTED A SURGERY WAS DELAYED WHEN IT WAS DISCOVERED THAT A PACS USER HAD FLIPPED ON IMAGE IN A CT HEAD SERIES, WHICH RESULTED IN SURGICAL PREPARATIONS BEING DONE FOR THE WRONG SIDE OF THE HEAD. THE ERROR WAS DISCOVERED PRIOR TO SURGERY. GE WAS CONTACTED TO IDENTIFY THE INDIVIDUAL WHO HAD PERFORMED THE IMAGE FLIP FROM INFO IN THE AUDIT LOG. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2672286 | THE CUSTOMER ALLEGED THAT THE COBAS IT 1000 WAS SENDING PATIENT RESULTS TO THE CUSTOMER'S LABORATORY INFORMATION SYSTEM (LIS) WITH THE WRONG PATIENT IDENTIFICATION. THE CUSTOMER STATED A COUPLE RESULTS WERE TRANSFERRED TO THE WRONG PATIENT IN THE LIS, AND THE RESULTS FROM THE LIS GO INTO THE HIS. THE CUSTOMER SEARCHED FOR THAT PATIENT AND DELETED THE RESULTS. THE CUSTOMER STATED THEY COULD NOT BE SURE THAT RESULTS WERE TRANSMITTED TO THE WRONG PATIENT, AND IT WAS DIFFICULT TO SAY IF THERE WERE RESULTS THAT WERE TRANSMITTED INCORRECTLY. CLARIFICATION HAS BEEN REQUESTED VERIFYING WHETHER ANY RESULTS WERE REPORTED OUTSIDE THE LABORATORY WITH THE WRONG PATIENT IDENTIFICATION NUMBER. THE CUSTOMER DID NOT REPORT ANY ADVERSE EVENTS. FURTHER INVESTIGATION DETERMINED THE COBAS IT 1000 WORKED AS INTENDED UP TO THE CORRUPTION OF THE EXTERNAL CUSTOMER'S DATABASE. NO EVIDENCE WAS FOUND TO PROVE THE COBAS IT 1000 APPLICATION PLAYED A ROLE IN CAUSING THE PATIENT MISMATCH REPORTED AT THE SITE. THE ROOT CAUSE OF THE REPORTED COMPLAINT WAS FOUND TO BE THE CORRUPTION OF THE CUSTOMER'S DATABASE.><IT WAS UNKNOWN IF THE INITIAL REPORTER SENT REPORT TO THE FDA. THIS EVENT OCCURRED IN (B)(6).><A ROOT CAUSE COULD NOT BE DETERMINED AS NECESSARY DATA IS NOT AVAILABLE FOR FURTHER INVESTIGATION. | 1
2672438 | THE CUSTOMER IS REPORTING THAT THE ORTHO PROVUE PROBE IS NOT DISPENSING ENOUGH RED CELLS IN THE DAT TEST. NO ALARM CONDITION OR ERROR MESSAGE WAS POSTED BY THE ORTHO PROVUE. NO ERRONEOUS RESULTS WERE REPORTED. ALL SAMPLES AFFECTED WERE REPEATED MANUALLY. THE SAMPLES REPORTED AS NEGATIVE BY THE PROVUE WERE NEGATIVE USING MANUAL GEL. AS PER THE OCD FIELD ENGINEER (FE) - THE FE CALLED THE CUSTOMER IN REGARDS TO THE PROVUE NOT DISPENSING ENOUGH RED CELLS. THE CUSTOMER STATED THAT THEY HAVE RESOLVED THE PROBLEM BY PRIMING AND RINSING THE PROBE AND RED CELLS ARE PIPETTING NORMAL. THE CUSTOMER IS NOT REQUESTING SERVICE. THE FE ASSISTED THE CUSTOMER WITH RESOLUTION VIA THE TELEPHONE. NO ONSITE SERVICING REQUIRED. INSTRUMENT IS OPERATING AS EXPECTED. (B)(4). | 0
2672449 | THE CUSTOMER IS REPORTING THAT THE ORTHO PROVUE PROBE IS NOT DISPENSING ENOUGH RED CELLS IN THE CROSSMATCH TEST. NO ALARM CONDITION OR ERROR MESSAGE WAS POSTED BY THE ORTHO PROVUE. NO ERRONEOUS RESULTS WERE REPORTED. ALL SAMPLES AFFECTED WERE REPEATED MANUALLY. THE SAMPLES REPORTED AS NEGATIVE BY THE PROVUE WERE NEGATIVE USING MANUAL GEL. AS PER THE OCD FIELD ENGINEER (FE) - THE FE CALLED THE CUSTOMER IN REGARDS TO THE PROVUE NOT DISPENSING ENOUGH RED CELLS. THE CUSTOMER STATED THAT THEY HAVE RESOLVED THE PROBLEM BY PRIMING AND RINSING THE PROBE AND RED CELLS ARE PIPETTING NORMAL. THE CUSTOMER IS NOT REQUESTING SERVICE. THE FE ASSISTED THE CUSTOMER WITH RESOLUTION VIA THE TELEPHONE. NO ONSITE SERVICING REQUIRED. INSTRUMENT IS OPERATING AS EXPECTED. (B)(4). | 0
2672985 | THE PATIENT WAS UNDERGOING AN ARTERIOGRAM WITH CONTRAST AND ON THE TABLE WITH A C-ARM IN SPECIAL PROCEDURES, RADIOLOGY. A CLEAR PLASTIC TABLE EXTENSION BROKE UNDER THE PATIENT'S WEIGHT AND THE PATIENT WAS EASED TO THE FLOOR BY ATTENDING DOCTORS AND TECHNICIANS. THE PATIENT WAS ON TABLE MATTRESS AND WITH PILLOWS, NO INJURY CONFIRMED BY CT OF NECK AND HEAD.======================MANUFACTURER RESPONSE FOR PLEXIGLAS TABLE EXTENSION TO FIT TABLE, IMAGING, INNOVA (PER SITE REPORTER).======================GE FIELD ENGINEER SUPPLIED CUSTOM MEASUREMENTS AND PLEXIGLASS THICKNESS REQUIREMENTS BASED UPON ANOTHER TABLE EXTENSION IN USE AT THE HOSPITAL AT THE TIME. | 0
2673032 | BIOMED TECHNICIAN WAS CALLED TO CICU TO TROUBLESHOOT THE MONITOR FAILURE OF NOT ABLE TO DETECT THE PRESENCE OF CARESCAPE, PATIENT DATA MODULE (PDM) PRIOR TO THE ADMITTING THE PATIENT. THE PDM CONNECTOR WAS DEFECTIVE (PLASTIC SURROUNDING PINS WAS DAMAGED AND MULTIPLE PINS WERE PUSHED IN NOT ALLOWING CONTACT TO BASE FEMALE CONNECTOR). AT THIS TIME ALL OTHER MONITORS (QTY 20) WERE INSPECTED AND FOUND 4 ADDITIONAL ROOMS DOWN (NOT IS USE) DUE TO THE SAME PROBLEM WHICH WERE NOT REPORTED TO BIOMED. OLD TRAM MODULES WERE USED TO KEEP ALL FIVE (5) CICU ROOMS OPERATIONAL.MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, CARESCAPE PDM (PER SITE REPORTER).======================NO RESPONSE AT THIS TIME YET. | 1
2674881 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD WAS REVIEWED. NOTHING WAS FOUND THAT WAS RELATED TO THIS EVENT. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. THE SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
2675755 | CARDIAC MONITOR DELETES PATIENT INFORMATION BY ITSELF AND ALSO THE PARAMETERS. | 1
2675947 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THERE WERE NO EVENTS NOTED IN THE DEVICE HISTORY RECORD (DHR) THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2675948 | THE CUSTOMER REPORTED THAT A DONOR DONATED ON (B)(6) 2012 AND DID NOT HAVE ANY SYMPTOMS OR PROBLEMS DURING THE DONATION PROCESS. AFTER THE DONOR HAD GONE HOME SHE DEVELOPED ITCHING ON THE LEFT SIDE OF HER HEAD AND FACE AND HAD REDNESS AND SWELLING OF THE LEFT EYELID, EYE BROW, CHEEK AND EAR. THE CUSTOMER STATED THAT THE DONOR CALLED THEM ON (B)(6) 2012 AND REPORTED THIS AND THAT HER LEFT EYE WAS SWOLLEN SHUT. THE DONOR ALSO REPORTED GOING TO HER DOCTOR ON (B)(6) 2012 AND SHE WAS PUT ON PREDNISONE TO TREAT THE ALLERGIC REACTION. THE CUSTOMER STATED THAT THE DONOR HAS NOT EXPERIENCED ANY OF THESE SYMPTOMS RELATED TO AN APHERESIS PROCEDURE BEFORE AND IS A LONG TIME DONOR. SHE ALSO REPORTED THAT THE DONOR DENIES ANY HISTORY OF ANY OTHER ALLERGIES AND HAD NOT HAD ANY RECENT BUG BITES. THE CUSTOMER STATED THAT THE DONOR DID NOT EXPERIENCE ANY RESPIRATORY DIFFICULTIES. THE DISPOSABLE SET AND THE ACD-A ARE NOT AVAILABLE FOR RETURN AS THEY WERE DISCARDED ON THE DAY OF DONATION. THIS REPORT IS BEING FILED DUE TO MEDICAL INTERVENTION VIA DOCTOR VISIT AND PRESCRIPTION. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. A REVIEW OF THE DEVICE HISTORY RECORD FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. BASED ON INFORMATION PROVIDED BY THE DONOR, IT IS DIFFICULT TO DETERMINE THE ACTUAL CAUSE OF THIS INCIDENT. IT IS POSSIBLE, BUT NOT CONCLUSIVE, THAT THE DONOR MAY HAVE EXPERIENCED A REACTION TO THE ACD-A OR ETHYLENE OXIDE.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE RDF SHOWED THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2676197 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RDF INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER TOWARDS THE END OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPOND WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2677753 | THERE WAS AN INSTRUMENTATION ERROR ON THE MICROSCOPIC URINALYSIS. THE TEST WAS REPEATED AND CORRECTIONS MADE TO THE RESULTS. | 0
2678312 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2678331 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR (B)(6) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. (B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2679856 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE ANALYSIS WAS REVIEWED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER TOWARD THE END OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2680478 | REPORTER INDICATED A PATIENT'S VNS GENERATOR WAS NOTED TO BE SET TO 0MA UPON INTERROGATION ON (B)(6) 2012. THE SETTINGS WERE CORRECTED THIS DAY. AT THE PREVIOUS APPOINTMENT ON (B)(6) 2012, VNS SETTINGS WERE AS INTENDED UPON INITIAL INTERROGATION. A SYSTEMS DIAGNOSTICS TEST WAS PERFORMED, BUT NO FINAL INTERROGATION WAS DONE. IT IS MOST LIKELY THAT THE SYSTEMS TEST FAULTED AND CHANGED THE OUTPUT CURRENT TO 0MA. IN ADDITION, THE PATIENT NO LONGER FELT STIMULATION AND HER DEPRESSION INCREASED SLIGHTLY TO BELOW PRE-VNS BASELINE LEVEL AS A RESULT OF THE LOSS IN THERAPY. AFTER THE VNS SETTINGS WERE READJUSTED, THE PATIENT EXPERIENCED VOICE ALTERATION AND FELT THE VNS STIMULATION. ATTEMPTS FOR PROGRAMMING HISTORY ARE IN PROGRESS.><THE PATIENT'S VNS PROGRAMMING HISTORY WAS REVIEWED. FAULTED SYSTEMS DIAGNOSTICS TESTS OCCURRED ON (B)(6) 2009, AND (B)(6) 2012. ALL SETTINGS WERE CORRECTED EXCEPT FOR ON (B)(6) 2012, WHEN A FINAL INTERROGATION WAS NOT PERFORMED. AT THE NEXT VISIT ON (B)(6) 2012, SETTINGS WERE 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF, WHICH IS INDICATIVE OF A PREVIOUS FAULTED SYSTEMS DIAGNOSTICS TEST. THE SETTINGS WERE NOT CORRECTED AT THIS VISIT, WHICH WAS PREVIOUSLY REPORTED BY THE PHYSICIAN. THE VNS SETTINGS WERE NOT CORRECTED UNTIL THE NEXT OFFICE VISIT ON (B)(6) 2012. .><DATE OF EVENT, CORRECTED DATA: THE CORRECT EVENT DATE WAS CONFIRMED FROM THE VNS PROGRAMMING HISTORY. ANALYSIS OF PROGRAMMING HISTORY. | 0
2680718 | ON JULY 10, 2012, DURING REVIEW OF THE PATIENT'S PROGRAMMING HISTORY, IT WAS DISCOVERED THAT A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED ON (B)(6) 2009 THAT CHANGED THE PATIENT'S SETTINGS. ALL THE PARAMETERS WERE CORRECTED EXCEPT FOR THE FREQUENCY PRIOR TO THE PATIENT LEAVING THE CLINICAL VISIT. NO ADVERSE EVENTS HAVE BEEN REPORTED DUE TO THIS PROGRAMMING ANOMALY. ANALYSIS OF PROGRAMMING HISTORY. | 0
2681519 | PATIENT ARRIVED FROM THE OR, WHEN HIS MONITOR WAS ADMITTED HIS VITAL SIGN (VS) PARAMETERS WERE NOT OUR NORMAL SETTINGS. IF WE HAD NOT BEEN PAYING ATTENTION, THE MONITOR WOULD NOT HAVE ALARMED FOR ANY UNSAFE VS. THE PATIENT PARAMETERS WERE AS FOLLOWS: HEART RATE: 1-300, SYSTOLIC BLOOD PRESSURE: 1-350, DIASTOLIC BLOOD PRESSURE: 1-350, MEAN BLOOD PRESSURE: 1-350. IN SPITE OF THE STRANGE PARAMETERS, THE MONITOR DISPLAYED THAT IT WAS IN THE 0-3 MONTH DEFAULT. | 1
2683402 | FAILURE OF ECG CONNECTOR PORT CAUSING NOISE IN THE ECG SIGNAL. MULTIPLE REPORTS WITH RISK/IMPACT: INABILITY TO READ/INTERPRET ECG, MONITOR MAY NOT BE ABLE TO PROCESS SIGNAL AND REGISTER AS ARTIFACT. WHEN INITIALLY REPORTED TO GE IN MAY, IT APPEARS THAT GE WAS AWARE OF THIS PROBLEM AND HAD NOT COMMUNICATED IT TO OUR INSTITUTION. PDM CONVERSION IS ON HOLD AT OUR HOSPITAL, PENDING RESOLUTION FROM GE. GE-UNRESOLVED, GE WORKING TO UNDERSTAND CAUSE AND SCOPE OF THE PROBLEM. OUR INSTITUTION IS IN COMMUNICATION WITH GE.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, MODULE, PATIENT DATA MODULE (PDM), (PER SITE REPORTER).======================UNRESOLVED. CAUSE AND SCOPE OF PROBLEM STILL UNDER INVESTIGATION BY GE. | 1
2684078 | PROCEDURE: REVISION DECOMPRESSION RIGHT L4-5, FUSION L4-5 WITH SEGMENTAL PEDICLE SCREW INSTRUMENTATION. DURING THE PROCEDURE THERE WAS AN ATTEMPT AT OPTICAL TRACKING. THE OPTICAL TRACKING PIN WAS PLACED INTO THE ILIAC CREST. O-ARM SPINS WERE THEN MADE AND THEN AN ATTEMPT TO TRANSFER THE IMAGES TO THE STEALTH SYSTEM THROUGH THE OPTICAL TRACKING FAILED. SEVERAL ATTEMPTS OF THIS WERE MADE AND THEY ALL FAILED, SO THE OPTICAL TRACKING WAS TERMINATED AND THE CASE AND PLACING OF PEDICLE SCREWS WAS PERFORMED BY C-ARM GUIDANCE. THERE WERE NO COMPLICATIONS OTHER THAN THE FAILURE OF THE O-ARM TO SEND SIGNALS TO THE STEALTH SYSTEM.FOLLOW UP ACTION: PER HOSPITAL'S BIOMEDICAL ENGINEER: THERE ARE USER CALIBRATIONS THAT NEED TO BE DONE INCLUDING A MOTION TEST. THE USERS WERE NOT AWARE OF THE NEED TO PERFORM THE MOTION TEST. IT WAS DECIDED THAT THE BIOMED DEPT WILL BE RESPONSIBLE FOR CONDUCTING THE CALIBRATIONS EVERY 30 DAYS AND THIS TESTING WILL INCLUDE PERFORMING THE MOTION TEST. | 1
2684079 | WHEN ATTEMPTING TO SET UP THE O-ARM FOR A SURGICAL MULTILEVEL POSTERIOR LUMBAR DECOMPRESSION AND INSTRUMENTED FUSION, AN ERROR MESSAGE DISPLAYED. "ERROR INITIALIZING COLLIMATORS" APPEARED ON THE FIRST LINE OF THE O-ARM. THE TECHNICIAN SHUT DOWN THE O-ARM, UNPLUGGED IT, DISCONNECTED THE MACHINE BEFORE REVERSING THAT, AND TURNED IT BACK ON. THE SAME MESSAGE WAS DISPLAYED. TECH SUPPORT FROM THE COMPANY WAS CONTACTED AND THE MONITOR WAS BOOTED UP AND DIAGNOSTIC TESTS WERE RAN AFTER REBOOTING. THE O-ARM STILL HAD THE SAME MESSAGE DISPLAYED. AN ADDITIONAL PHONE CALL TO A MEDTRONIC FIELD ENGINEER WAS MADE AND THE ERROR WAS RECOGNIZED AS BEING A HARDWARE ISSUE. THE O-ARM WAS REMOVED AND THE PROCEDURE WAS PERFORMED USING THE C-ARM. THE CASE WAS DELAYED BY ONE HOUR.THIS FACILITY HAS MET WITH REPRESENTATIVES FROM MEDTRONIC TO DISCUSS O-ARM PROBLEMS. MEDTRONIC IS WORKING WITH THIS FACILITY TO RESOLVE ANY ISSUES. | 0
2685477 | HARDWARE PROBLEM RELATED TO USE OF ECG TELEMETRY WITH DASH MONITOR IN COMBO MODE. OVERHEATING DUE TO HIGH NETWORK TRAFFIC CAN CAUSE ONE OF SEVERAL SYMPTOMS, MOST NOTABLY AND CLINICALLY SIGNIFICANT IS A DISTORTION OF THE ECG WAVEFORM (VERTICAL OR HORIZONTAL COMPRESSION). 4-5 RECENT REPORTS OF SEEING THIS PROBLEM IN PATIENT CARE AREAS. GE ACTION-PENDING. REPLACEMENT BOARD IDENTIFIED AND IN PRODUCTION BUT WITH LONG LEAD TIMES OF 4 MONTHS FOR ETHERNET CONTROLLER BOARD.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, DASH (PER SITE REPORTER). | 0
2685478 | FAILURE OF INVASIVE BLOOD PRESSURE (IBP) SPLITTER CABLE, SOMETIMES INTERMITTENT, RESULTING IN FLUCTUATING BLOOD PRESSURE READING OR INABILITY TO ZERO TRANSDUCER. RISK/IMPACT: POTENTIAL FOR MISDIAGNOSIS/MISTREAT BASED ON ERRONEOUS OR UNRELIABLE INFORMATION.RESOLVED AT HOSPITAL: HOSPITAL ENGINEER IDENTIFIED ANOTHER CABLE THAT COULD BE USED INSTEAD. IBP SPLITTER CABLES WERE REPLACED WITH BIFURCATED BP CABLES. DIFFICULTY GETTING GE REPLACEMENT CABLES INITIALLY BUT HAS BEEN RESOLVED AT HOSPITAL. | 1
2685553 | DURING REVIEW OF PROGRAMMING HISTORY OF A BATTERY LIFE CALCULATION IT WAS OBSERVED ON (B)(6) 2006 THE DEVICE WAS PROGRAMMED TO 0.5/30/500/30/3, THEN A SYSTEM DIAGNOSTIC WAS PERFORMED AND RESULTED IN "FAULT" COMMUNICATION. THE DEVICE WAS INTERROGATED AND THE SETTINGS HAD CHANGED TO 1/20/500/30/60 WHICH ARE INDICATIVE OF "FAULTED" DIAGNOSTICS. PRIOR TO THE PATIENT LEAVING THE VISIT THE DEVICE WAS NOT PROGRAMMED TO INTENDED SETTINGS. THE PATIENT SETTINGS WERE CORRECTED ON (B)(6) 2006. ANALYSIS OF PROGRAMMING HISTORY. | 0
2687485 | DURING THE REVIEW OF THE GENERATOR SOURCE CODE, IT WAS REVEALED THAT THE MAGNET SWIPE COUNTER (DIAGMAGNETCOUNTS) HAD BECOME MAXED OUT WITH 65,535 MAGNET SWIPES. THE DIAGMAGNETCOUNTS IS EXPECTED TO INCREASE BY 1 WITH EVERY MAGNET SWIPE PERFORMED UNTIL IT REACHES THE MAXIMUM OF 65,535 MAGNET SWIPES. WHEN PERFORMING A MAGNET MODE DIAGNOSTIC TEST, THE PROGRAMMING SOFTWARE READS THE DIAGMAGNETCOUNTS AND LOOKS FOR IT TO INCREMENT BY 1, INDICATING THAT A MAGNET SWIPE WAS PERFORMED. AS THE COUNTER HAD MAXED OUT, THE NUMBER WAS NO LONGER INCREMENTING. THEREFORE THE SOFTWARE ASSUMED THAT A MAGNET SWIPE HAD NOT OCCURRED AND PROVIDED THE "MAGNET SWIPE NOT DETECTED" MESSAGE TO THE USER. REVIEW OF THE SOURCE CODE DOES INDICATE THAT THE MAGNET SWIPES ARE REGISTERING AND THAT THE PROGRAMMED THERAPY IS BEING DELIVERED. THE SOFTWARE AND GENERATOR APPEAR TO BE FUNCTIONING AS INTENDED; HOWEVER DUE TO A DESIGN LIMITATION, MAGNET MODE DIAGNOSTICS CANNOT BE PERFORMED.><IT WAS REPORTED BY THE PHYSICIAN THAT HE WAS UNABLE TO PERFORM MAGNET MODE DIAGNOSTICS. HE INDICATED THAT HE WOULD SWIPE THE MAGNET AND THEN ATTEMPT TO PERFORM THE TEST, BUT THE TEST WOULD COME BACK WITH AN ERROR MESSAGE INDICATING THAT THE GENERATOR WAS NOT SENSING A MAGNET SWIPE. HE THEN WENT ON TO SAY THAT, HE KNOWS THAT A MAGNET SWIPE OCCURRED BECAUSE WHEN HE WENT BACK TO CHECK THE MAGNET ACTIVATIONS HISTORY, THE MAGNET SWIPE SHOWED UP. HE PROVIDED THE PATIENT'S MAGNET SETTINGS AND SWIPING TECHNIQUE WAS VERIFIED. HE INDICATED THAT HE WOULD SWIPE THE MAGNET AND THEN IMMEDIATELY RUN THE DIAGNOSTICS TEST. THE PHYSICIAN ATTEMPTED MAGNET MODE DIAGNOSTICS MULTIPLE TIMES AND WAS STILL RECEIVING THE SAME ERROR MESSAGE. THE MAGNET HISTORY WAS CHECKED AND, PER THE PHYSICIAN, ALL OF THE ACTIVATIONS SHOWED UP. THE PHYSICIAN REPORTED THAT HE WAS ABLE TO PERFORM SYSTEM MODE DIAGNOSTICS WITHOUT ISSUE, AND THAT EVERYTHING WAS OK, NO SPECIFICS PROVIDED. THE PHYSICIAN WENT ON TO INDICATE, THAT HE DIDN'T THINK ANYTHING WAS WRONG WITH THE MAGNET, AND BELIEVED THAT MAGNET STIMULATION WAS OCCURRING, HOWEVER STATED THAT THE PATIENT DOES NOT RESPOND ( PHYSICALLY/ VERBALLY) TO STIMULATION, AND NEVER HAS. HE ALSO INDICATED THAT PER THE PATIENT'S CAREGIVERS, THE MAGNET SWIPES HAVE NOT BEEN AS EFFECTIVE FOR THE PATIENT RECENTLY. HOWEVER THE PHYSICIAN FELT THAT THIS MAY HAVE BEEN RELATED TO A LOWER ON TIME AND DIFFERENT SETTINGS CHOICES AS THE PATIENT HAD BEEN SEEING A DIFFERENT VNS TREATING PHYSICIAN WHO HAD THE PATIENT PROGRAMMED TO 30SEC ON TIME FOR THE MAGNET AND NOT 60SEC. THE PROGRAMMING HISTORY FROM THE DATE OF THE EVENT HAS BEEN REQUESTED BUT HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE. NO ADDITIONAL INFORMATION IS KNOWN AT THIS TIME.><PRODUCT INFORMATION WAS RECEIVED ON (B)(6) 2012 AND ADDITIONAL PROGRAMMING HISTORY WAS RECEIVED AND REVIEWED ON (B)(6) 2012. ANALYSIS OF THE PROGRAMMING HISTORY REVEALED THAT THE MAGNET SWIPES WERE REGISTERING AND WERE PROVIDING THE PROGRAMMED AMOUNT OF OUTPUT CURRENT. IT APPEARS THAT THE MAGNET SWIPE ISSUE MAY BE THE RESULT OF A COMMUNICATION ERROR. IT IS CURRENTLY UNKNOWN IF A MAGNET MODE DIAGNOSTIC HAS BEEN PERFORMED SUCCESSFULLY SINCE THE INITIAL REPORT OF THIS EVENT. EVENT APPEARS TO BE RELATED TO A COMMUNICATION ISSUE BETWEEN THE PROGRAMMING SYSTEM AND THE GENERATOR. THE EXACT CAUSE OF THE COMMUNICATION ISSUE IS UNKNOWN.><. | 0
2688165 | THE HOOK BROKE INTO TWO PIECES DURING SURGERY. THE PIECES WERE REMOVED FROM THE AREA AND SAVED. NOTHING WAS RETAINED IN THE PATIENT. THERE WAS NO HARM TO THE PATIENT. | 0
2688507 | FALSE NEGATIVE RESULTS ON THE TRICYCLIC DRUG SCREEN. ALERE TRIAGE TOX DS LOT NUMBER 50995. THC - GAVE FALSE POSITIVE RESULTS ON (B)(6) 2012, WITH ALERE TRIAGE TOX DRUG SCREEN. | 0
2688532 | THE ANESTHESIA CIRCUIT HAD AN EDITH 1000 HEAT AND MOISTURE EXCHANGE FILTER WHICH WAS REPORTED AS AFFECTING THE ABILITY TO VENTILATE ON SEVERAL OCCASIONS. THE ANESTHESIOLOGIST TEAM REPORTED MEETING RESISTANCE WHEN ATTEMPTING TO VENTILATE. THE CIRCUIT FUNCTIONED AS EXPECTED WHEN THE FILTER WAS REMOVED. DATES OF USE: APPROX (B)(6) 2012. | 0
2688568 | AS THE SURGEON WAS JUST COMPLETING A DAVINCI TOTAL LAPAROSCOPIC HYSTERECTOMY, THE MEGA NEEDLE DRIVER JAWS FAILED WHILE IN USE. | 0
2688589 | DAVINCI MONOPOLAR SCISSOR BROKE OFF INTRA-ABDOMINALLY IN PATIENT. INSTRUMENT WAS UNABLE TO BE DETACHED FROM TROCAR. TROCAR HAD TO BE REMOVED FROM THE PATIENT TO REMOVE THE INSTRUMENT. NEW INSTRUMENT APPLIED AND THE PROCEDURE CONTINUED. THERE WAS NO HARM TO THE PATIENT. | 0
2689899 | WATER IS BEING FOUND INSIDE THE BLOOD PRESSURE CUFFS AND WATER IS BEING ASPIRATED INTO THE AUTOMATIC BLOOD PRESSURE MACHINES CAUSING DAMAGE. THESE ARE BEING USED IN HOSPITALS. STERILMED HAS BEEN NOTIFIED AND SUGGEST FACILITY PURCHASE NEW ONES. | 0
2689956 | ON (B)(6) 2012, THE PHARMACY DEPARTMENT BEGAN UTILIZING THE MEDCAROUSEL BY MCKESSON TO DISPENSE MEDICATIONS TO PATIENTS. THE DISPENSED CAROUSEL INVENTORY IS DRIVEN BY THE PRODUCT BARCODE, WHICH ENHANCES SAFETY BY ENSURING THAT THE MEDICATION IS CORRECTLY DISPENSED TO THE CORRECT PATIENT. THE PROBLEM IS THAT THE LABEL DOES NOT REFLECT THE DOSE OR THE FREQUENCY. EVEN THOUGH THE CORRECT PRODUCT IS DISPENSED, THE INCORRECT DOSE MAY INADVERTENTLY BE PROVIDED (1/2 TAB OF A MEDICATION VS. WHOLE TAB) BECAUSE THE PATIENT SPECIFIC DOSE IS NOT CLEAR ON THE LABEL. IN ADDITION, THE FREQUENCY IS NOT STATED ON THE LABEL, WHICH CAUSES CONFUSION WITH THE QUANTITY TO DISPENSE IN A 24 HOUR PERIOD. | 1
2693257 | IT WAS REPORTED THAT THE PT ECG DEVELOPED BRADYCARDIA WITH A SIGNIFICANT PAUSE AND THE MONITORING KAPPA XLT REPORTEDLY FAILED TO ALARM. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2693393 | ONE 5MM ROBOTIC CANNULA WAS DISCOVERED BENT AND REMOVED FROM CASE. ONE ROBOTIC MARYLAND WAS DISCOVERED HAVING SHORTED OUT AND WAS REMOVED FROM SERVICE. | 0
2696085 | A DOCTOR REPORTED THAT THE UNIT WAS NOT MEASURING CORRECTLY. THE MEASUREMENTS WERE COMPARED TO THOSE OF ANOTHER SYSTEM AND A DEVIATION OF 2 DIOPTERS WAS NOTED. THERE WAS NO HARM TO THE PATIENT INVOLVED AS THE MEASUREMENTS WERE NOT USED. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. (B)(4). | 1
2696703 | IT WAS REPORTED TO OUR COUNTRY MANAGER IN FRANCE THAT A VNS PROGRAMMING PHYSICIANS; (B)(4) HANDHELD COMPUTER BATTERY WAS FAULTY. IT WAS REPORTED THAT THE BACK COVER WAS BROKEN, THEREFORE THE BATTERY FALLS OUT. GOOD FAITH ATTEMPTS ARE UNDERWAY TO HAVE THE HANDHELD RETURNED FOR ANALYSIS.><THE HANDHELD AND SOFTWARE WERE RETURNED FOR ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD WHICH WOULD NOT AFFECT A BATTERY CHARGING. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE HANDHELD WAS RECEIVED WITHOUT AN AC ADAPTER AND BATTERY COVER. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY AND BATTERY LATCH WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE BATTERY COVER WAS MISSING. AS A RESULT OF THE MISSING BATTERY COVER, THE HANDHELD WOULD NOT POWER ON. ONCE A KNOWN GOOD COVER WAS INSTALLED, NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE MALFUNCTION OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
2697661 | PATIENT WAS CONNECTED TO AVALON FETAL MONITOR WITH US TRANSDUCER. AFTER APPROXIMATELY 30 MINUTES. THE PATIENT FELT THE TRANSDUCER WAS "HOT". NURSE CHECKED FOUND NOTHING WRONG WITH TRANSDUCER BUT SWAPPED OUT TRANSDUCER TO BE ON THE SAFE SIDE. TRANSDUCER WAS SENT TO BIOMED FOR EVALUATION.======================MANUFACTURER RESPONSE FOR FETAL MONITOR, AVALON US TRANSDUCER (PER SITE REPORTER).======================THE TRANSDUCER WAS CLEANED AND DISINFECTED EXTERNALLY AND CONNECTED TO FM 30 MONITOR. THE TESTING WAS PERFORMED AS PER THE PROCEDURE GIVEN IN THE USER MANUAL. THIS WAS CARRIED OUT BY 2 DIFFERENT TECHNICIANS SEVERAL TIMES INDEPENDENTLY AND FOUND NO PROBLEM. THEN WE CONNECTED A NEW TRANSDUCER FROM OUR STOCK AND COMPARED THE PERFORMANCE. THERE WAS NO DIFFERENCE. ALL SEVEN CRYSTALS SHOWED SENSITIVITY IN THE SAME RANGE USING A 20 CM WATER PHANTOM.WE LEFT THE TRANSDUCER CONNECTED TO MAACHINE AND TOOK SURFACE TEMPERATURE READING AT AVERAGE ROOM TEMPERATURE OF 72 DEGREES FAHRENHEIT AT AN INTERVAL OF AN HOUR DURING WORKING HOURS (7 AM TO 7 PM) AND IT WAS FOUND TO BE BETWEEN 71 TO 73 FAHRENHEIT. THERE WAS NO HEATING OF THE TRANSDUCER.THE TRANSDUCER WAS OPENDED AND INSPECTED UNDER MICROPSCOPE. ALL 7 CRYSTALS APPERARED PERFECTLY FINE WITHOUT ANY PHYSICAL DAMAGE. THERE WAS SLIGHT DUST ON BOTH THE PCBS AND WE DID ULTRASONIC CLEANING. DURING THIS PERIOD, THE TOP CASE WAS OPEN AND WE DID NOT FIND THE TRANSDCUER HEATING. IT WORKED EXACTLY THE SAME WAY THE NEW TRANSDUCER FROM OUR STOCK. ALL THE SMD COMPONENT'S SURFACE TEMPERATURE REMAINED WITHIN 2 DEGREES FAHRENHEIT COMPARED TO ROOM TEMPERATURE.THE TRANSDCUER TOP CASE WAS PUT BACK AND ONCE AGAIN IT WAS TESTED WITH 20 CM WATER PHANTOM. THERE WAS NO HEATING OF THE TRANSDUCER AND IT FUNCTIONED ABSOLUTELY NORMAL.WE FEEL THAT THE ROOM TEMPERATURE IN THE PATIENT ROOM MIGHT HAVE BEEN QUITE LOW AND HENCE THE PATIENT MIGHT HAVE FELT SOME DIFFERENCE.PLEASE CONFIRM THAT YOU HAVE RECEIVED THIS REPORT AND IF YOU WOULD LIKE US TO PERFORM ANY OTHER TEST. IT IS STILL CONNECTED TO THE MONITOR AND WE HAVE BEEN MONITORING THE PERFORMANCE, ESPECIALLY THE TEMPERATURE. IT CAN BE SHIPPED TOMORROW. | 0
2700433 | THIS X-RAY SYSTEM WILL INTERMITTENTLY EXPORT FLIPPED IMAGES TO THE PICTURE ARCHIVING COMPUTER SYSTEM (PACS). THESE FLIPPED IMAGES REFERS TO IMAGE INVERSION AND CAN BE DESCRIBED AS "MIRRORED" WHERE THE LEFT AND RIGHT ARE REVERSED. INVERTED WOULD APPLY TO THE TOP AND BOTTOM BEING REVERSED. THE RANDOM CHANGES IN THE IMAGE ORIENTATION COULD HAVE BEEN ANY OF THE FOUR POSSIBLE ORIENTATION ERRORS. A REBOOT OF THE SYSTEM WILL TEMPORALLY RESTORE NORMAL OPERATION. CONCLUSION: THE ROOT CAUSE OF THIS PROBLEM IS A SOFTWARE RELATED ISSUE. A FIELD CHANGE ORDER WILL BE RELEASED TO CORRECT THIS SOFTWARE ISSUE. | 1
2702257 | A VNS PROGRAMMING PHYSICIAN REPORTED THAT HE PERFORMED A SYSTEM DIAGNOSTIC TEST AND WHEN HE DID THE PATIENT COMPLAINED THAT SHE FELT A HORRIBLE PAIN UP HER NECK AT THE LEAD SITE AND HAS CONTINUED TO FEEL THE SAME PAIN EVERY HOUR SINCE THAT TIME. THE PATIENT'S DEVICE HAD BEEN PROGRAMMED OFF PRIOR TO THE TESTING. THE TEST IS PERFORMED AT 1MA. THE PHYSICIAN DID NOT PERFORM A FINAL INTERROGATION TO MAKE SURE THAT PATIENT WAS STILL TURNED OFF BEFORE SHE LEFT THE OFFICE. THE PATIENT WAS SO UPSET THAT SHE WOULD NOT ALLOW HIM TO GET THE PROGRAMMING WAND NEAR HER. THE PATIENT RETURNED TO CLINIC AND IT WAS DETERMINED THAT THEIR DEVICE HAD BEEN LEFT PROGRAMMED ON AND THEY WERE RECEIVING THERAPY AT 1MA ONCE AN HOUR. THE PATIENT WAS PROGRAMMED TO 1/20/500/30/60. THIS IS FROM A FAULTED SYSTEM DIAGNOSTIC TEST. THE PATIENT'S DEVICE WAS SET TO 0 MA. TEACHING WAS PROVIDED AS TO THE CAUSE OF THE EVENT AND HOW TO PREVENT IN THE FUTURE. ANALYSIS OF SETTING PROVIDED BY SITE. | 0
2703983 | THE CUSTOMER'S MOTHER WROTE AN EMAIL TO INQUIRE IF THERE WERE ANY PROMOTIONS FOR THE CARELINK USB. IN THE EMAIL, THE MOTHER REPORTED THAT THE CUSTOMER'S USB CABLE STOPPED WORKING, AND SHE CANNOT AFFORD ONE AT THIS TIME DUE TO ALL OF HER MEDICAL BILLS FROM INSULIN, TEST STRIPS AND HOSPITAL BILLS FROM HOSPITALIZATIONS THAT THE CUSTOMER HAS HAD FROM SEIZURES DUE TO LOW BLOOD GLUCOSE LEVELS. NOTHING FURTHER WAS REPORTED IN THE EMAIL. CURRENTLY, IT IS UNKNOWN WHETHER OR NOT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT AS NO PRODUCT HAS BEEN RETURNED. NO CONCLUSION CAN BE DRAWN AT THIS TIME. WE THEREFORE CONSIDER THIS REPORT COMPLETE TO THE BEST OF OUR KNOWLEDGE. | 0
2704655 | OVER THE PAST THREE YEARS, THERE HAVE BEEN MULTIPLE COMPLAINTS REGARDING THE LOSS OF COMMUNICATION BETWEEN THE GE TRANSPORT PRO AND THE ATTACHED TRAM MODULE. THE COMPLAINTS HAVE STATED THAT THE PATIENT DATA AND WAVEFORMS WOULD DROP OUT AND NO LONGER BE RECORDED WHILE THE PATIENT WAS IN TRANSPORT AROUND THE HOSPITAL. THIS CAN BE A PARTICULARLY HAZARDOUS EVENT, ESPECIALLY WHEN THE PATIENTS ARE UNDER INTENSIVE CARE AND REQUIRE A HIGH LEVEL OF MONITORING TO ENSURE THEIR HEALTH AND SAFETY.INITIALLY, THE UNITS WERE TESTED IN AN ATTEMPT TO REPRODUCE THE ISSUE, BUT NO APPARENT PROBLEM WAS IDENTIFIED, AND THE DEVICES WERE RETURNED TO SERVICE. AFTER ANOTHER COMPLAINT, THE UNIT WAS REMOVED FROM SERVICE AND REPLACED WITH A SPARE, AND NO FURTHER ISSUES WERE REPORTED. THE ONLY NOTICEABLE DIFFERENCE BETWEEN THE ORIGINAL UNIT AND THE SPARE WAS THE SOFTWARE VERSIONS. THE ORIGINAL DEVICE HAD A SOFTWARE VERSION 2.0, WHILE THE SPARE HAD VERSION 2.1.======================MANUFACTURER RESPONSE FOR MONITOR, PHYSIOLOGICAL, TRANSPORT PRO (PER SITE REPORTER).======================UNIT WAS RETURNED TO THE MANUFACTURE FOR EVALUATION. THEY COULD NOT REPRODUCE THE REPORTED PROBLEM BUT THEY RELOADED THE SYSTEM SOFTWARE ON THE UNIT. | 1
2704658 | TELEMETRY MONITOR BEGAN TO PICK UP WHAT LOOKED LIKE PACER SPIKES. A STRIP WAS PRINTED OUT AND PLACED IN THE BATTERY COMPARTMENT. TELEMETRY BOX WAS THEN SWITCHED AND PROBLEM WAS RESOLVED. | 1
2704659 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING ERRONEOUS LOW NRBC RESULTS WITH INSTRUMENT GENERATED FLAGS GENERATED BY THEIR UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM ON ONE SPECIFIC PATIENT ANALYZED TWICE ON THE SAME DAY ON THE SAME INSTRUMENT AS COMPARED TO RESULTS FROM A DIFFERENT INSTRUMENT (LH 750) AND THE MANUAL DIFFERENTIAL, WHICH WERE CONSIDERED CORRECT. THE RESULTS PROVIDED BY THE CUSTOMER ARE SHOWN IN THE ATTACHED DOCUMENT. SAMPLES ARE COLLECTED IN 3 ML GREINER BIOONE VACUETTE DIPOTASSIUM ETHYLENEDIAMINE TETRA-ACETIC ACID (K2EDTA) TUBES. THE SAMPLE TESTED WAS STORED AT ROOM TEMPERATURE AND WAS 1 TO 2 HOURS OLD. CONTROLS ANALYZED IN THE MORNING BEFORE THE INCIDENT RECOVERED WITHIN EXPECTED RANGES. THE DXH 800 INSTRUMENT IS CURRENTLY PERFORMING WITHIN QUALITY CONTROL (QC) SPECIFICATIONS WITH RESPECT TO CONTROLS (ACCURACY) AND REPRODUCIBILITY (PRECISION). IT WAS NOTED THAT THE NRBC FOR THE LH750 WAS REPORTED AS 600+ BECAUSE 600 IS THE UPPER END OF LINEARITY. THE MANUAL DIFFERENTIAL NRBC COUNT WAS REPORTED AS 350/ 15 WHITE BLOOD CELLS (WBC) BECAUSE OF THE LOW WBC COUNT. THIS WOULD NORMALLY BE REPORTED AS NRBC/100 WBCS AS THE INSTRUMENTS REPORT. ALL OTHER DXH 800 CBC AND DIFFERENTIAL RESULTS WERE COMPARED TO THE LH750 SINCE THESE RESULTS WERE STATED TO BE CORRECT AND A COMPLETE MANUAL DIFFERENTIAL WAS NOT PROVIDED. THE CUSTOMER STATED THAT IT WAS NOT POSSIBLE TO DO A MANUAL DIFFERENTIAL ON ALL SLIDES BECAUSE OF THE LOW WBC COUNT. FURTHER REVIEW OF THE PRINTOUT PROVIDED FOR THE INITIAL DXH 800 RUN SHOWED THAT NE% WAS HIGH, AND LY% AND MO% WERE LOW COMPARED TO THE LH750 BUT WERE ACCOMPANIED BY INSTRUMENT GENERATED FLAGS. IN ADDITION, RBC, HGB, HCT, AND MPV WERE LOW, AND PLATELET (PLT) WAS HIGH AS COMPARED TO THE LH750 WITH NO INSTRUMENT FLAGS. FURTHER REVIEW OF THE PRINTOUT PROVIDED FOR THE DXH 800 RERUN SHOWED THAT NE% WAS HIGH, AND LY% AND MO% WERE LOW COMPARED TO THE LH750 BUT WERE ACCOMPANIED BY INSTRUMENT GENERATED FLAGS. IN ADDITION, RBC AND HCT WERE HIGH, AND PLT AND MPV WERE LOW AS COMPARED TO THE LH750 WITH NO INSTRUMENT FLAGS. THE CAUSE OF THE ERRONEOUS NRBC RESULTS IS UNKNOWN. THE INSTRUMENT GENERATED SYSTEM MESSAGES WITH R FLAGS TO ALERT THE OPERATOR THAT OPERATION OF THE SYSTEM MAY BE AFFECTED AND TO REVIEW THE RESULTS (ERRONEOUS DIFFERENTIAL RESULTS WERE ALSO ACCOMPANIED BY INSTRUMENT GENERATED FLAGS; HOWEVER, ERRONEOUS CBC RESULTS WERE UNFLAGGED). PER LABELING FROM THE DXH 800 INSTRUCTIONS FOR USE: - THE DXH 800 SYSTEM MANAGER INCLUDES FLAGS, CODES, AND MESSAGES TO ALERT YOU TO ISSUES WITH PATIENT OR CONTROL RESULTS. YOU CAN ALSO CUSTOMIZE THE FLAGGING OF RESULTS AND DEFINE RULES FOR FLAGGING SAMPLE RESULTS. - ALL SYSTEM MESSAGES ARE ACCOMPANIED BY R (REVIEW) FLAGS. A SYSTEM MESSAGE INDICATES AN EVENT OCCURRENCE THAT MAY AFFECT THE OPERATION OF THE SYSTEM, REQUIRES OPERATOR NOTIFICATION, OR ENTRY INTO A HISTORY LOG. | 0
2704876 | THE FAULTY HANDHELD COMPUTER AND ASSOCIATED SOFTWARE HAVE BEEN RECEIVED AND UNDERGONE ANALYSIS. THE PROGRAMMING SYSTEM PERFORMED TO SPECIFICATION AND NO ANOMALIES WERE FOUND. NO ANOMALIES WERE OBSERVED WITH THE HANDHELD COMPUTER OR ANY OF THE CABLES AS DESCRIBED BY THE SITE.><IT WAS REPORTED THAT THE PHYSICIAN'S VNS HANDHELD PROGRAMMING COMPUTER WAS EXPERIENCING INTERMITTENT COMMUNICATION ISSUES AND WOULD ONLY WORK "WHEN YOU HELD THE CORD UP." A COMPANY REPRESENTATIVE INDICATED THE ISSUE APPEARED TO BE RELATED TO THE CONNECTION BETWEEN THE HANDHELD COMPUTER AND THE SERIAL CABLE ADAPTER. THE FAULTY HANDHELD WILL REPORTEDLY BE RETURNED HOWEVER HAS NOT BEEN RECEIVED TO DATE. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE ISSUE. DEVICE FAILURE IS SUSPECTED BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2707967 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR (B)(4) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. (B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2709778 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT T THEIR GLOUCESTER CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. (B)(4). THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS (35%) -LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS. POST-PURGE SATURATIONS (25%) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED. PLASMA LINE OCCLUSIONS (10%) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS. EVENT BASED, OTHER - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2711095 | PT IN HOSPITAL AT (B)(6). HE HAS DIABETES. AN ORDER FOR HIGH DOSE INSULIN COVERAGE WAS CORRECTLY ENTERED INTO THE SYSTEM ON OR BEFORE (B)(6) 2012. THE ORDER CLEARLY STATED IT WAS THE HIGH DOSE COVERAGE SCALE. HOWEVER, WHEN THIS ORDER WAS USED THE INSTRUCTIONS DISPLAYED THE LOW DOSE COVERAGE INSTRUCTIONS. THE PT GOT THE WRONG INSULIN. NO SIGNIFICANT HARM, BUT DID RUN HIGHER GLUCOSES. I CANCELED THE HIGH DOSE COVERAGE ORDER AND THEN RE-ENTERED THE IDENTICAL ORDER. IT NOW DISPLAYED THE PROPER INSTRUCTIONS. THIS WAS ENTERED IN THE HOSPITAL EHR SYSTEM CERNER POWER CHART. DATES OF USE: (B)(6) 2012. | 1
2711961 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT (B)(6). THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2711970 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT (B)(6). THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2713672 | COMPUTER DATABASE ISSUE: LAB RESULT (KETONE) POSTED IN APPROPRIATE CHART (PATIENT A) IN "SOARIAN" DATA BASE. DURING AUTOMATIC DATA TRANSFER INTO "IBEX" DATA BASE, THE KETONE RESULTS FROM PATIENT A POSTED INTO INCORRECT PATIENT RECORD (PATIENT B) IN THE "IBEX". | 1
2714382 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR LEEDS CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2714870 | AS A PT WAS BEING SENT INTO THE GANTRY, THE PT'S ARM WAS PINCHED BY THE MOVING COUCH TOP BETWEEN THE COUCH TOP AND THE COUCH. THE PT WAS ON THE COUCH HEAD FIRST AND SUPINE WITH ARMS DOWN BY THE SIDE AND HANDS ON CHEST. AS THE PT WAS BEING SENT INTO THE GANTRY, THE ARM WAS PINCHED BY THE MOVING COUCH TOP BETWEEN THE COUCH TOP AND COUCH. THE VINYL COVERING FOR THE COUCH TOP THAT SLIPS OVER THE COUCH TOP PAD DID NOT COVER THE GAP BETWEEN THE COUCH TOP AND THE COUCH. THERE WAS A SHEET UNDER THE PT, BUT IT WAS NOT COVERING THE ENTIRE COUCH TOP. THE PT WAS ONLY PINCHED. THERE WAS NO BREAKING OF THE SKIN OR NEED FOR MEDICAL TREATMENT. | 0
2717198 | SOFTWARE ISSUE IN THE MEDITECH ELECTRONIC MEDICAL RECORD SYSTEM BLOOD BANK MODULE. PREVIOUSLY IN BLOOD BANK BBK RESULT ENTRY, IT WAS POSSIBLE FOR EDITS TO ANTIGEN OR ANTIBODY IDENTIFICATION TEST RESULTS NOT TO FILE TO THE SPECIMEN RESULTS OR BBK HISTORY. THIS OCCURRED IF THE TOTAL LENGTH OF THE RESULT STRING FOR THE TEST EXCEEDED 20 CHARACTERS IN THE FORMAT "A=P, B=N, C=P..." WHERE A, B, AND C WERE ANTIGEN MNEMONICS, AND P AND N REPRESENTED POSITIVE AND NEGATIVE, RESPECTIVELY, OR THE FORMAT "X, Y, Z..." WHERE X, Y AND Z WERE ANTIBODY MNEMONICS AND IF THE EDIT BEING MADE DID NOT CAUSE THE CONTENTS OF THE RESULT FIELD ON THE SCREEN TO CHANGE. WITHIN BBK RESULT ENTRY, ANTIGEN AND ANTIBODY RESULTS WILL BE DISPLAYED ALPHABETICALLY IN A STRING FORMAT WITH IN THE AGID OR ABID RESULT FIELD. BECAUSE OF THE WAY ANTIGEN AND ANTIBODY RESULTS ARE HANDLED, AT A CERTAIN RESULT LENGTH A "..." IS APPENDED TO THE MAIN AGID OR ABID RESULT FIELD TO INDICATE MORE RESULTS ARE PRESENT BUT WILL FIT IN THE FIELD. FOR EXAMPLE, A STRING OF ANTIGEN WOULD DISPLAY AS "A=P, B=P, BG=P..." WHILE A STRING OF ANTIBODIES WOULD DISPLAY AS 'A, A1, AB, ACA, ADD, A..." WHEN NEW ANTIGENS OR ANTIBODIES ARE ADDED TO THESE RESULT STRINGS AFTER INITIAL VERIFICATION, THE CODE WILL CHECK TO SEE IF THIS RESULT STRING HAS CHANGED. IF THE NEW ANTIGENS OR ANTIBODIES BEING ENTERED DO NOT CAUSE THE PREVIOUSLY VERIFIED RESULT DISPLAY STRING TO BE ALTERED -A=P, B=P, BG=P, C=P... - THEN THOSE NEW ANTIGENS/ANTIBODIES WILL NOT BE FILED. IF THE NEW ANTIGENS OR ANTIBODIES DO CAUSE THE PREVIOUSLY VERIFIED RESULT DISPLAY STRING TO BE ALTERED THEN THOSE NEW ANTIGENS/ANTIBODIES WILL BE SAVED. THE MANUFACTURER DID NOTIFY OUR FACILITY THAT THE ISSUE HAS BEEN CORRECTED BY (B)(6) AND A REQUEST HAS BEEN MADE TO CHANGE THE CONTROL ONCE CODING HAS BEEN COMPLETED. OUR FACILITY REVIEWED ALL OF THE PT ANTIBODY FILES FROM (B)(6) 2012, WHEN THE MANUFACTURER STATED THE ISSUE APPEARED WHEN AN UPGRADE TO OUR MEDITECH SYSTEM OCCURRED. WE FOUND NO ERRORS AND ALL ANTIBODY/ANTIGEN IDENTIFICATIONS ARE SHOWING IN THE PT'S HISTORY REPORT. MOVING FORWARD UNTIL THE MEDITECH CORRECTION IS IN PLACE, OUR FACILITY WILL REVIEW/PRINT PT HISTORY WITH ALL ANTIBODY AND ANTIGEN IDENTIFICATIONS AND ATTACH THE PRINTOUT TO THE PAPER PANEL AND FILE IN THE ANTIBODY PANEL BINDER IN BLOOD BANK. | 1
2717876 | DURING ROBOTIC CASE A SMALL PIECE OF FOREIGN BODY WAS FOUND LOOSE IN CAVITY. IT WAS DEEMED TO BE A PIECE OF THE 'SPATULA' THAT WAS BEING USED AT THE TIME. | 0
2718045 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE COLLECTED PLATELET PRODUCT AT THEIR (B)(6) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2719328 | VENTANA MEDICAL SYSTEMS, INC. RECEIVED A CUSTOMER COMPLAINT REGARDING FLUID FROM A SOURCE INTERNAL TO A BENCHMARK ULTRA AUTOMATED SLIDE STAINER CONTACTING THE ELECTRICAL CIRCUITRY OF THE STAINER RESULTING IN A SHORT CIRCUIT AND ENOUGH HEAT TO CAUSE SOME MELTING OF WIRING INSULATION RESULTING IN SOME SMOKE BEING EMITTED. THERE WERE NO SERIOUS INJURIES OR DEATH ASSOCIATED WITH THE INCIDENT REPORTED. THE ROOT CAUSE OF THE INCIDENT ORIGINALLY REPORTED UNDER 2028492-2012-0001 WAS DETERMINED TO BE A BUILD-UP OF DEBRIS IN THE WASTE FLUID DISCHARGE SYSTEM. A NOTIFICATION HAS BEEN SENT TO CUSTOMERS TO INFORM THEM OF THE POTENTIAL FOR THIS TYPE OF EVENT TO OCCUR, HOW TO RECOGNIZE IF THE WASTE DISCHARGE SYSTEM IS BECOMING RESTRICTED AND TO REMIND THEM TO PERFORM PERIODIC CLEANING OF THE SYSTEM. A CAPA HAS BEEN OPENED TO PURSUE ADDITIONAL CORRECTIVE ACTION(S).><THE INSTRUMENT HAS BEEN RETURNED FROM THE CUSTOMER SITE IS BEING EXAMINED IN DETAIL. AS OF THE DATE OF THIS REPORT A DEFINITIVE CAUSE OF THE FLUID CONTACT WITH THE CIRCUIT BOARD THAT IS VISIBLY DAMAGED HAS NOT BEEN ESTABLISHED. A FOLLOW-UP REPORT WILL BE ISSUED PENDING THE OUTCOME OF THE INVESTIGATION INTO ROOT CAUSE OF THE LEAK. | 0
2721024 | THE USER FACILITY WAS UNABLE TO IDENTIFY OR DETERMINE WHAT EPS CYCLE THE SCOPE WAS PROCESSED WITH OR WHAT DAY THE SCOPE WAS PROCESSED ON. THE FACILITY STATED A PROCESS INDICATOR WAS PLACED AND IT PASSED. THE USER FACILITY HAS REPORTED THAT THE PATIENT IS FINE WITH NO SUSTAINING INJURIES.><THE CUSTOMER REPORTED THAT A PATIENT DEVELOPED COLITIS AFTER A PROCEDURE USING A COLONOSCOPE PROCESSED IN THE RELIANCE EPS. (B)(4) REVIEWED THE RELIANCE DG LOT AND CONFIRMED IT WAS MANUFACTURED TO SPECIFICATION WITH NO NON-CONFORMANCES. THE CUP AND DRY CHEMISTRY CONTENTS WERE FOUND TO BE IN GOOD CONDITION. THE CUP WAS PLACED INTO HOPKIN'S LABORATORY RELIANCE EPS WITH A CHEMICAL INDICATOR AND COMPLETED A CYCLE WITHOUT ISSUE; NO RESIDUAL OR CDS WAS OBSERVED IN THE CUP AND THE INDICATOR PASSED. A STERIS CLINICIAN COMPLETED AN IN-SERVICE ON (B)(4) 2012 REGARDING THE FACILITY'S REPROCESSING PRACTICES, INCLUDING PRE-CLEANING, PLACEMENT AND REMOVAL OF SCOPES INTO EPS, THE OPERATION OF EPS AND SCOPE STORAGE PRACTICE.><THE CUSTOMER REQUESTED STERIS TO INSPECT THE RELIANCE EPS AS PART OF THEIR ROOT CAUSE INVESTIGATION INTO THE REPORTED EVENT. WE WILL PROVIDE A FOLLOW-UP REPORT WHEN ADDITIONAL INFORMATION IS AVAILABLE. | 1
2722011 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE COLLECTED PLATELET PRODUCT AT (B)(6). THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2722012 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE COLLECTED PLATELET PRODUCT AT (B)(6). THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) -LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED.><(B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2722755 | THE 27.29 RESULTS ON A TOSOH AIA-900 WERE BEING RUN AND REPORTED VIA AN ORCHARD SOFTWARE LIS. THE LIS SOFTWARE WAS IMPROPERLY CONFIGURED TO REQUEST THE SPECIMEN AT A 1:1 DILUTION INSTEAD OF THE PROPER DILUTION OF 1:21 AS REQUIRED BY THE TEST PROTOCOL. ONE ERRONEOUS RESULT WAS REPORTED. SEVEN PTS RESULTS WERE TRANSMITTED AS >H (HIGH). THE INSTRUMENT TEST FILE WAS CONFIRMED AS CORRECT, AND UNTIL THE SOFTWARE FROM THE LIS IS CORRECTED TO REQUEST THE PROPER DILUTION, THE SPECIMENS ARE RUN BY-PASSING THE HOST QUERY FUNCTION. THE CORRECT RESULTS ARE OBTAINED AND REPORTED. | 1
2722895 | NOTE: THIS REPORT IS BEING ISSUED AS RESULT OF CHANGES TO COMPANY SOP'S AND DEFINITIONS OF MDR. THIS REPORTABLE EVENT WAS FOUND DURING FDA INSPECTION AND SUBSEQUENT REVIEW OF CRITICAL COMPLAINTS. IT WAS DETERMINED THAT THIS EVENT OCCURRED AS A RESULT OF A MALFUNCTION TO THE DEVICE POSSIBLY CAUSED BY USER ERROR (DROPPING THE DEVICE AND CRACKING THE CASE. THIS MDR IS BEING PROVIDED AS REQUESTED FROM THE WARNING LETTER ISSUED BY FDA AND AS INDICATED WILL DO IN COMPANY RESPONSE TO FDA RE: CMS CASE # (B)(4). FUKUDA DENSHI COMPLAINT NO. (B)(4) WAS AS THE RESULT OF AN END-USER AT (B)(6) HOSPITAL ON (B)(6) 2010 COMPLAINING OF A BURNING SMELL COMING OUT OF A CARDIAC OUT-PUT MODULE (HF-500) WHICH WAS ATTACHED TO A PT MONITOR FUKUDA DENSHI MODEL DS-5300. AS A RESULT OF THIS COMPLAINT, FUKUDA DENSHI PERFORMED A FULL INVESTIGATION AND ISSUED "HF-500 DETAILED REPORT OF FAILURE INVESTIGATION" DATED (B)(6) 2010. A MEDICAL DEVICES SAFETY INFORMATION REPORT WAS FILED BY THE USER WITH (B)(6) ON (B)(6) 2010 USING DOCUMENT NO. (B)(4), SENT TO FUKUDA DENSHI ON (B)(6) 2010. FUKUDA DENSHI THEN RESPONDED TO THE FILED REPORT WITH THE RESPONSE DATED (B)(6) 2010 BACK TO THE (B)(6). IT WAS DETERMINED THAT THIS INCIDENT OCCURRED AS A RESULT OF A BURNED VARISTOR. IT WAS ALSO OBSERVED THAT THE DEVICE CASE HAD CRACKED PERHAPS DUE TO AN EARLIER DROP AND THAT THIS DEVICE HAD BEEN IN USE FOR 12 YEARS AND 3 MONTHS. IT WAS DETERMINED THAT THE VARISTOR COULD HAVE BEEN DAMAGED DUE TO THE IMPACT THAT RESULTED IN THE CRACKED CASE. IN FURTHER INVESTIGATION, IT WAS DETERMINED THAT AT THE TIME FUKUDA DENSHI HAD AN ALREADY INSTALLED BASED OF 1226 UNITS AND THIS WAS THE FIRST CASE SEEN FOR THIS TYPE OF FAILURE OF THE VARISTOR. FUKUDA DENSHI VISITED THE CUSTOMER AND PROVIDED A THOROUGH EXPLANATION OF THE DEVICE MALFUNCTION, EXPLAINED THE CONSEQUENCES OF OPERATING A DEVICE AFTER AN APPARENT FALL AND SUBSEQUENT CRACKED CASE AND MADE THE DECISION TO REPLACED THE MONITORING SYSTEM. ADDITIONALLY, COMPLETE DOCUMENTATION AND CORRESPONDENCE WAS FORWARDED TO THE FDA IN THE FUKUDA DENSHI FDA 2ND RESPONSE LETTER DATED 06/14/2012. | 1
2724880 | THE USER FACILITY OBSERVED A FALSE HIGH PLATELET COUNT (PLT) ON A BURN PATIENT IN THE EMERGENCY ROOM. A COMPLETE BLOOD COUNT (CBC) AND NUCLEATED RED BLOOD CELLS (NRBC) WERE ANALYZED ON AN XE-5000 AUTOMATED HEMATOLOGY ANALYZER AT 13:11 IN THE AUTOMATED MODE. THE SAMPLE WAS JUDGED "NEGATIVE" WITH NO FLAGS. WORK AREA MANAGER (WAM) MIDDLEWARE HELD THE REPORT AS THE PLTS RESULTS WERE A CRITICAL HIGH VALUE. THE RESULTS WERE RELEASED TO THE EMERGENCY ROOM WITH A VALUE OF PLT=770. THE OPERATOR OF THE DEVICE WAS UNAWARE AT THE TIME THAT THE PT WAS BEING TREATED AS A BURN PT. THE USER INDICATED THAT AT AN UNSPECIFIED TIME, A PHASE MICROSCOPY PERFORMED WITH PLT=105. THE ORIGINAL SAMPLE WAS REPEATED, USING TWO ADDITIONAL CHANNELS FOR A CBC+ DIFFERENTIAL (DIFF) +NRBC+RETICULOCYTE AT 15:24 WITH A REPORTED PLT=751. THE SAMPLE RESULTS WERE IDENTIFIED AS "POSITIVE" WITH MORPHOLOGY AND COUNT ERRORS. A SECOND BLOOD SAMPLE WAS DRAWN IN THE INTENSIVE CARE BURN UNIT. THE SAMPLE WAS ANALYZED AT 17:10 FOR CBC AND DIFF +NRBC IN THE AUTOMATED MODE-PLT=611, "POSITIVE" WITH A COUNT ERROR. THE PT WAS REMOVED FROM LIFE SUPPORT AT AN UNSPECIFIED TIME AND EXPIRED. A REVIEW OF THE USER FACILITY XE-5000 CONTROLLER SETTING REVEALED THAT THE USER FACILITY CHOSE NOT TO SET FLAG FORMULAS ON THE ANALYZER. FLAG FORMULAS ARE DESIGNED FOR THE USER TO ESTABLISH JUDGMENT LIMITS FOR ABNORMAL BLOOD CONDITIONS. THE ACTUAL XE-5000 DEVICE WAS EVALUATED AT THE USER FACILITY. THE EVALUATION DID NOT REVEAL A ROOT CAUSE FOR THE REPORTED EVENT; HOWEVER, THE USER FACILITY, IN CHOOSING TO NOT SET FLAG FORMULAS FOR THE ANALYZER, THE RESULTS COULD NOT BE FLAGGED IN THE FINAL PRINTED RESULT. THE WORK AREA MANAGEMENT (WAM) SOFTWARE RULE THRESHOLD FOR THE USER FACILITY WAS SET FOR A VALUE LESS THAN 600 X 10^9/L; THEREFORE, THE SYSTEM WOULD NOT HAVE THE CAPABILITY TO CAPTURE A VALUE OVER 600. THERE ARE LIMITATIONS OF IMPEDANCE COUNTING AS NOTED IN THE USER MANUAL. GOOD LABORATORY PRACTICE WOULD INCLUDE UTILIZING FLAG FORMULAS IN CONJUNCTION WITH WAM RULES. THE ANALYZER AND WAM FUNCTIONED AS DESIGNED. BURN PTS WITH GREATER THAN 20% OF THE BODY AFFECTED MAY EXPERIENCE MORPHOLOGIC RBC CHANGES INCLUDING GLOBULAR FRAGMENTATION, BUDDING AND MICROSPHERES. OSMOTIC FRAGILITY IS INCREASED. | 0
2725157 | WE HAD REPLACED THIS ANESTHESIA DELIVERY UNIT (ADU) TWO DAYS BEFORE THIS EVENT. WE DID OUR STANDARD MACHINE CHECKOUT IN THE MORNING AND EVERYTHING CHECKED OUT SATISFACTORY. WE BROUGHT OUT OPEN HEART PATIENT TO OR, ATTACHED OUR MONITORS, AND DID OUR INDUCTION OF ANESTHESIA. AFTER THE PATIENT WAS INTUBATED WE ATTACHED HIM TO OUR VENTILATOR AND NOTICED THAT WE DID NOT HAVE ANY END TIDAL CO2, TIDAL VOLUMES, OR ANESTHETIC DELIVERY AGENT SHOWING UP ON OUR MONITOR. THIS OF COURSE CAN HAPPEN FOR OTHER REASONS. WE AUSCULTATED THE PATIENTS LUNGS AND KNEW THAT THE PATIENT HAD BEEN SUCCESSFULLY INTUBATED AS WE HAD USED THE GLIDESCOPE (SO WE HAD VISUALIZED THE ENDOTRACHEAL TUBE PASSING THRU THE VOCAL CORDS ON THE SCREEN ATTACHED TO THE GLIDESCOPE). PLEASE NOTE THAT THIS PATIENT HAD THE POTENTIAL TO BE A VERY DIFFICULT INTUBATION. WE DID QUICKLY GET OUR PORTABLE ETCO2 MONITOR TO MONITOR ETCO2. AFTER WE DETERMINED THAT THE PATIENT WAS SAFE, WE TRIED TO REPLACE THE GAS MODULE (THAT GIVES US ALL OF THIS INFORMATION) AND WE IMMEDIATELY CALLED BIOMED TO TAKE A LOOK AT THE MACHINE. IN THE MEANTIME WE HAD TO TELL OUR OPEN HEART SURGEON THAT WE WOULD HAVE TO DELAY THE START OF THE CASE UNTIL WE EITHER GOT THE CURRENT MACHINE OPERATING CORRECTLY OR WE BROUGHT IN A NEW MACHINE. THANKFULLY I HAD ANOTHER CRNA, MDA, AND ANESTHESIA TECHNICIAN TO HELP ME GET A NEW ANESTHESIA MACHINE INTO THE ROOM. HAD THIS BEEN ON A WEEKEND OR OFF SHIFT WHEN WE HAVE LIMITED STAFFING THE POTENTIAL FOR HARM TO THE PATIENT WOULD HAVE BEEN MUCH HIGHER. THIS WAS NOT EASY TO TRANSFER MACHINES WHILE THE PATIENT WAS ASLEEP AND REQUIRED EXTENSIVE MONITORING. WE GOT A NEW ADU INTO THE OPERATING ROOM AFTER THE PATIENT WAS HOOKED UP TO AN AUXILIARY MONITOR AND WAS ATTACHED TO AN AMBU BAG AND 100% OXYGEN AND WAS HAND VENTILATED UNTIL WE GOT THE NEW MACHINE INTO THE OR. THE MALFUNCTIONINGMACHINE WAS IMMEDIATELY TAKEN TO BE CHECKED OUT BY BIOMED. SUBSEQUENTLY THE NEW MACHINE THAT THEY BROUGHT IN TO REPLACE THE MALFUNCTIONING MACHINE ALSO HAD A CIRCUIT BOARD THAT ENDED UP NEEDING TO BE REPLACED AFTER OUR CASE WAS DONE. IN THE OPEN HEART ROOM WE HAVE MULTIPLE TV MONITORS THAT SHOW OUR VITAL SIGNS SO THAT THE PERFUSION TEAM, PHYSICIAN ASSISTANTS, AND SURGEONS CAN SEE EXACTLY WHAT IS HAPPENING WITH THE PATIENTS VITAL SIGNS AT ANY SECOND BASED ON HOW THEY ARE MANIPULATING THE HEART. THE 2ND MACHINE WOULD NOT DISPLAY THE VITAL SIGNS ONTO THE MONITORS. BIOMED WAS ABLE TO CONFIGURE SOMETHING TEMPORARILY TO GET US THRU THE CASE AND THEN REPLACED THE CIRCUIT BOARD AFTER THE CASE WAS OVER. BIOMED ASSESSMENT: THE MONITOR (DATEX OHMEDA ADU) WAS INTERMITTENTLY DISPLAYING" GAS MEASUREMENT REMOVED" WHICH PREVENTED ANY ETCO2 MEASUREMENTS FROM BEING DISPLAYED ON THE ADU. BIOMED REPLACED THE GAS MODULE DURING THE PROCEDURE BUT IT DID NOT REMEDIATE THE PROBLEM. THE ADU WAS REMOVED FOR SERVICE AND THE ERROR LOGS WERE CHECKED WHICH DID NOT INDICATE ANY SPECIFIC ERROR CODES THAT WOULD IDENTIFY THE CAUSE OF PROBLEM. AFTER THE MACHINE WAS POWERED OFF AND BROUGHT IN TO BIOMED THE PROBLEM COULD NOT BE DUPLICATED. BIOMED CONTACTED GE TECH SUPPORT AND FOLLOWED AND PERFORMED ALL THE RECOMMENDED TESTING PROCEDURES AND SELF DIAGNOSTICS. THE CAUSE OF THE PROBLEM COULD NOT BE IDENTIFIED. GE TECH SUPPORT THINKS IT'S POSSIBLE THAT RESETTING THE POWER TO THE MACHINE BY TURNING THE MONITOR OFF AND THEN ON AGAIN CLEARED THE PROBLEM AND WAS A TEMPORARY SOFTWARE GLITCH. THE 2ND ADU THAT WAS BROUGHT IN AS A REPLACEMENT FOR THE MACHINE NOTED ABOVE WAS MONITORING THE PATIENT AND DELIVERING ANESTHESIA PROPERLY ITSELF BUT IT WAS NOT SENDING DATA AND WAVEFORMS TO THE SLAVE DISPLAYS IN THE ROOM WHICH ARE ALSO VIEWED BY THE SURGEON AND PERFUSIONIST. THE PROBLEM WAS CAUSED BY A FAULTY OUTPUT DISPLAY BOARD (BDISP) WHICH WAS REPLACED AFTER THE PROCEDURE AND RESOLVED THE PROBLEM WITH THE MACHINE. NO HARM TO PATIENT. | 1
2727594 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE COLLECTED PLATELET PRODUCT AT THEIR (B)(6) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2727619 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE COLLECTED PLATELET PRODUCT AT THEIR (B)(6) CENTRE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS FROM (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
2730826 | THE CUSTOMER CALLED TO INQUIRE ABOUT THE CONSEQUENCES OF USING EXPIRED MATERIALS FOR MEASURING HEMAGLOBIN (HGB) FOR THE QUALIFICATION OF DONORS DONATING RED BLOOD CELLS (RBCS) ON THE TRIMA. THE CUSTOMRE DECLINED TO PROVIDE PATIENT INFORMATION FOR THIS EVENT. THIS REPORT IS BEING FILED DUE TO DEVICE MALFUNCTION IN THE FORM OF OPERATOR ERROR THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE TERUMO BCT SUPPORT SPECIALIST INFORMED THE CUSTOMER THAT POTENTIALLY THE DONOR WAS NOT QUALIFIED TO DONATE AN RBC PRODUCT AND THE RBC POTENCY MIGHT BE IN QUESTION. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT, THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. INVESTIGATION: INFORMATION WAS PROVIDED TO THE CUSTOMER INFORMING THEM THAT TERUMO BCT CONSIDERS THIS A HIGH PRIORITY EVENT AND ATTEMPTS WERE MADE TO OBTAIN FURTHER INFORMATION. THE CUSTOMER STATED THEY UNDERSTOOD AND CONSIDERED THIS ISSUE IS RELATED TO INTERNAL PROCESSING AND DID NOT CONSIDER THIS TO BE RELATED TO TERUMO (B)(4) PRODUCTS. BASED ON THIS VIEW THEY DECLINED TO PROVIDE ADDITIONAL INFORMATION REGARDING THE REPORTED CONDITION. THEREFORE FURTHER INVESTIGATION CANNOT BE COMPLETED. A TREND CODE REPORT DID NOT IDENTIFY ANY OTHER REPORTS OF POTENTIAL DATA ENTRY ERRORS DUE TO THE CUSTOMER USING EXPIRED PRODUCT TO CALCULATE DONOR HEMATOCRIT. ROOT CAUSE: THIS ISSUE WAS CLEARLY IDENTIFIED AS OPERATOR ERROR. THE EQUIPMENT PERFORMED AS SPECIFIED. | 0
2732484 | DURING AN ARTHROSCOPIC PROCEDURE USING A FLUID PUMP/MANAGEMENT SYSTEM, RATHER THAN GRAVITY, THERE WAS AN UNUSUAL AMOUNT OF FLUID EXTRAVASATION INTO THE TISSUE ABOVE THE OPERATIVE SITE. IN HOUSE TESTING SUGGESTS AN ISSUE WITH THE DISPOSABLE INFLOW TUBE SET RATHER THAN THE FLUID PUMP. | 0
2736028 | ATTEMPTS FOR THE RETURN OF THE FAULTY PROGRAMMING SYSTEM HAVE BEEN UNSUCCESSFUL TO DATE. THE SITE INDICATED THAT THEY WOULD NOT BE RETURNING ANY PROGRAMMING SYSTEMS AS THEIRS ARE FUNCTIONING NORMALLY.><INFORMATION WAS RECEIVED INDICATING THAT THE HANDHELD IN QUESTION WAS LIKELY AT A DIFFERENT FACILITY FOR THE SAME PHYSICIAN. ADDITIONAL ATTEMPTS FOR PRODUCT RETURN ARE IN PROGRESS.><IT WAS REPORTED THAT THE PHYSICIAN'S VNS HANDHELD PROGRAMMING COMPUTER WAS NOT ABLE TO HOLD A CHARGE FOR LONGER THAN 10-15 MINUTES. THE SITE INDICATED THAT THEY WERE ABLE TO USE THE PROGRAMMING SYSTEM WHICH WAS ABLE TO COMMUNICATE WITH PATIENT'S VNS DEVICES WITHOUT ISSUE BUT WERE INCONVENIENCED BECAUSE THEY COULD NOT REMOVE THE HANDHELD COMPUTER FROM THE POWER ADAPTER FOR LONG. THEY HAD BEEN EXPERIENCING THESE ISSUES SINCE ABOUT TWO WEEKS PRIOR TO THE REPORT. THE PHYSICIAN WAS SENT A REPLACEMENT HANDHELD COMPUTER. ATTEMPTS FOR THE RETURN OF THE FAULTY VNS HANDHELD PROGRAMMING COMPUTER ARE IN PROGRESS. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE ISSUE.><THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2012 AND PRODUCT ANALYSIS HAS SINCE BEEN COMPLETED. AN ANALYSIS WAS PERFORMED ON THE HANDHELD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE RETURNED AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS ALSO PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE FAILURE IS SUSPECTED BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2736247 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD PROGRAMMING COMPUTER WAS NOT CONSISTENTLY RESPONDING TO TAPS ON THE TOUCH SCREEN. THE ISSUE WAS ABLE TO BE RESOLVED WITH A HARD RESET AT FIRST, HOWEVER, THE ISSUES LATER PERSISTED AND WERE REPORTEDLY NOT ALWAYS ABLE TO BE RESOLVED WITH TROUBLESHOOTING. A COMPANY REPRESENTATIVE INDICATED THAT THE SCREEN RESPONSIVENESS ISSUE MAY HAVE BEEN THE HANDHELD COMPUTER FREEZING. THE SITE WAS SENT A NEW REPLACEMENT HANDHELD COMPUTER. ATTEMPTS FOR THE RETURN OF THE FAULTY HANDHELD COMPUTER ARE IN PROGRESS. NO PATIENT ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE ISSUE.><THE FAULTY HANDHELD COMPUTER AND ASSOCIATED SOFTWARE WERE RETURNED AND UNDERWENT ANALYSIS. ANALYSIS CONFIRMED THE TOUCH SCREEN WAS NOT FUNCTIONING CORRECTLY. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. NO ISSUES WERE OBSERVED WITH THE SOFTWARE. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE FAILURE IS SUSPECTED BUT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2736572 | VOLUNTARY REPORT WAS PREVIOUSLY SUBMITTED. AT APPROXIMATELY 1:00 AM ON (B)(4) 2012, THE GE CENTRAL STATION ON A MEDICAL-SURGICAL UNIT (5 WEST) INDICATED A LEADS FAILURE ALARM. THIS TELEMETRY ALARM WAS TO HAVE ACTIVATED THE ASSIGNED NURSES' ASCOM PHONE. THIS APPARENTLY DID NOT HAPPEN. WHEN THE NURSE RETURNED TO CHECK ON THE PT, SHE DISCOVERED THE PT UNRESPONSIVE AND CALLED A CODE BLUE. SUBSEQUENT INVESTIGATION OF THE EVENT INDICATED THAT THE ALARMS ORIGINATING FROM THE GE CENTRAL STATION CATEGORIZED AS "CRISIS", "WARNING", "ADVISORY", OR "MESSAGE". IN ORDER FOR AN ALARM TO ACTIVATE THE DESIGNATED PHONE CARRIED BY THE NURSE ASSIGNED TO THAT PT, THE ALARM LEVEL HAS TO BE SET AT EITHER "CRISIS" OR "WARNING". IN THIS CASE, THE LEADS FAILURE ALARM WAS DESIGNATED AS "ADVISORY" AND THEREFORE COULD NOT ACTIVATE THE PHONE. | 1
2738670 | (B)(4) SOFTWARE IS PLACING IMAGING IN THE WRONG PT CHART. THE CUSTOMER IS USING (B)(4). (B)(4) IS NOT COMPATIBLE WITH SOFTDENT V14.1 OR THE GENDEX GXS 700 SENSORS. ON (B)(4) 2009, GENDEX DENTAL SYSTEMS CONDUCTED A VOLUNTARY RECALL OF THE (B)(4), STATING ISSUES WITH THE PRACTICE MANAGEMENT SOFTWARE AS IT INTERFACES WITH SOFTDENT MANAGEMENT SYSTEM RELATING TO PT DATA MIX UPS. THIS RECALL WAS FDA TERMINATED ON (B)(4) 2011. THE CORRECTION TO THIS RECALL WAS TO UPDATE THE (B)(4). OUR DISTRIBUTOR (B)(4) WAS NOTIFIED ON TWO OCCASIONS AND ACKNOWLEDGED BOTH TIMES OF THIS RECALL. (B)(4) SENT A RECALL NOTIFICATION TO (B)(6) ON (B)(6) 2010 WITH INSTRUCTIONS ON HOW TO CORRECT THE ISSUE WITH AN UPGRADE TO (B)(4). IF ASSISTANCE WAS NEEDED, THE CUSTOMER WAS DIRECTED TO CALL GENDEX TECHNICAL SUPPORT AT (B)(4). THE CUSTOMER'S SOFTWARE WAS UPGRADED TO VERSION 2.0 ON (B)(4) 2012. TECHNICAL SUPPORT IS CURRENTLY WORKING WITH THIS CUSTOMER TO CORRECT THE ISSUE. THERE WAS NO PT OR USER INJURY ASSOCIATED WITH THIS EVENT. | 1
2739081 | AT THIS TIME THE HANDHELD HAS NOT BEEN RETURNED FOR ANALYSIS.><OUR CONSULTANT WENT TO THE SITE AND IT WAS REPORTED THAT NO ISSUES WERE SEEN WITH THE SITE'S HANDHELD COMPUTER. IT WAS REPORTED TO BE WORKING FINE.><IT WAS REPORTED TO OUR CONSULTANT BY A VNS TREATING PHYSICIAN THAT THEIR HANDHELD WAS NOT HOLDING A CHARGE AND IT HAS 2 LINES INTERFERING ON THE SCREEN. THE PRODUCT IS BEING RETURNED FOR ANALYSIS. AT THIS TIME THE HANDHELD HAS NOT BEEN RECEIVED AT THE MANUFACTURER. .><DATE RECEIVED BY MANUFACTURER (MO/DAY/YR) DATE SENT ON SUPPLEMENTAL 01 EMDR SHOULD HAVE BEEN (B)(6) 2012, NOT (B)(6) 2012. | 0
2739571 | IT WAS REPORTED THAT DURING AN A-FIB PROCEDURE, WHEN A CATHETER WAS PLUGGED INTO THE REF DECA PORT THE BODY SURFACE ECGS SIGNALS AND ALL THE IC (INTRACARDIAL) RECORDINGS HAD NOISE ON IT. WHEN THE SAME CATHETER AND CABLES ARE PLACED INTO POLE 20 B THE NOISE CEASED. UPON FURTHER FOLLOW-UP, IT WAS CONFIRMED THAT THE NOISE WAS ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BODY SURFACE ECGS AND ALL IC (INTRACARDIAL) RECORDINGS ON BOTH CARTO 3 AND EP RECORDING SYSTEMS SIMULTANEOUSLY. THE NOISE WASHED OUT ENTIRE SCREEN. THE PROCEDURE WAS COMPLETED BY MOVING THE CATHETER TO THE 20 POLE B PORT TO COMPLETE PROCEDURE. NO PATIENT INJURY WAS REPORTED. INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. (B)(4).><EVALUATION SUMMARY. (B)(4): IT WAS REPORTED THAT DURING AN A-FIB PROCEDURE, WHEN A CATHETER WAS PLUGGED INTO THE REF DECA PORT THE BODY SURFACE ECGS SIGNALS AND ALL THE IC (INTRACARDIAL) RECORDINGS HAD NOISE ON IT. WHEN THE SAME CATHETER AND CABLES ARE PLACED INTO POLE 20 B THE NOISE CEASED. UPON FURTHER FOLLOW-UP IT WAS CONFIRMED THAT THE NOISE WAS ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BODY SURFACE ECGS AND ALL IC (INTRACARDIAL) RECORDINGS ON BOTH CARTO 3 AND EP RECORDING SYSTEMS SIMULTANEOUSLY. THE NOISE WASHED OUT ENTIRE SCREEN. THE PROCEDURE WAS COMPLETED BY MOVING THE CATHETER TO THE 20 POLE B PORT TO COMPLETE PROCEDURE. NO PATIENT INJURY WAS REPORTED. THE CARTO 3 SYSTEM ASSOCIATED WITH THE REPORTED INCIDENT WAS INSPECTED BY BWI SERVICE ENGINEER. THE ECG 1 WAS REPLACED, THEN BWI SERVICE ENGINEER PERFORMED THE ATP TEST AND THE SYSTEM PASSED. SYSTEM WAS READY FOR USE. THE SUSPECTED CARD WAS TESTED BY THE CARTO MANUFACTURER. THE REPORTED PROBLEM WAS NOT REPRODUCED. THE NOISE LEVEL WAS WITHIN THE ALLOWED TOLERANCES. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. THE COMPLAINT COULD NOT BE CONFIRMED. | 1
2745150 | THE CUSTOMER ALLEGED THEY RECEIVED A QUESTIONABLE URINE PROTEIN RESULT ON THEIR URISYS 2400 ANALYZER. THE PATIENT'S INITIAL PROTEIN RESULT WAS NEGATIVE AND IT WAS REPORTED OUTSIDE THE LABORATORY. THE CUSTOMER NOTICED A PROBLEM WITH THE SAMPLE WHILE RUNNING A MICROSCOPIC TEST ON IT. BEFORE REPEAT TESTING, THE SAMPLE WAS REMIXED BASED ON THE LENGTH OF TIME IT HAD BEEN SITTING SINCE THE INITIAL RUN. THE SAMPLE WAS REPEATED ON THE URISYS 2400 AND THE RESULT WAS 70 MG/DL. THE CUSTOMER THEN TESTED THE SAMPLE ON A BAYER CLINITEK ANALYZER AND THE RESULT WAS 100 MG/DL. 100 MG/DL WAS MORE "CLINICALLY INDICATED" TO THE PHYSICIAN AND IT WAS ISSUED AS A CORRECTED REPORT. THE PATIENT WAS NOT TREATED BASED ON THE DISCREPANT RESULT AND THERE WERE NO ADVERSE EVENTS. THE URISYS 2400 CASSETTE REAGENT LOT NUMBER AND EXPIRATION DATE WERE NOT PROVIDED. THE CUSTOMER STATED THEY THOUGHT THE SAMPLE POSSIBLY HAD SOME TYPE OF CLOT IN IT BASED ON THE RESULT. THE DISCREPANT RESULT WAS THOUGHT TO BE SAMPLE SPECIFIC AND NO SERVICE VISIT WAS DISPATCHED. .><AN INVESTIGATION WAS NOT POSSIBLE AS THE CUSTOMER COULD NOT PROVIDE THE LOT NUMBER OF THE TEST STRIPS USED NOR RETURN ANY MATERIAL. A ROOT CAUSE COULD NOT BE DETERMINED. | 0
2748774 | CONSULTANT'S BRAND NEW COMPRESSION PROBE IS NOT WORKING PROPERLY AND IS READING MALFUNCTION. CONSULTANT FIRST TRIED USING THE PROBE ON (B)(6) 2012 AND IT READ MALFUNCTION. HE THEN TRIED USING IT AGAIN (B)(6) 2012 AND STILL READ MALFUNCTION. IT IS NOT GIVING ANY MEASUREMENTS AT ALL. THERE WAS NO ADVERSE EFFECT TO THE PT. THE INVESTIGATION COULD NOT BE COMPLETED; NO CONCLUSION COULD BE DRAWN, AS NO PRODUCT WAS RECEIVED. A REVIEW OF THE DEVICE HISTORY RECORDS HAVE BEEN REQUESTED. | 0
2749417 | HANDSET AT NURSE'S STATION WOULD NOT STOP ALARMING. FURTHER INVESTIGATION REVEALED 2 FLOORS NURSE CALL SYSTEM WERE NOT WORKING. RESETTING THE SYSTEM DID NOT FIX THE PROBLEM. A MANUFACTURER TECHNICIAN CAME ON SITE AND REPAIRED A CORRUPT SOFTWARE FILE ON THE MASTER COMPUTER (THE OTHER FAILED NURSE CALL IS A SLAVE SYSTEM).======================MANUFACTURER RESPONSE FOR NURSE CALL, NURSE COMMUNICATIONS MODULE (PER SITE REPORTER).======================MANUFACTURER TECHNICIAN REPAIRED A CORRUPT FILE IN THE SOFTWARE TO CORRECT THE FAILURE. | 1
2749420 | PUMP WOULD NOT TAKE THE JULY LIBRARY DOWNLOAD. THE SOFTWARE THAT MANAGES THE PUMP INDICATES THAT IT STILL HAS A FEBRUARY LIBRARY AND THE PUMP INDICATES THAT IT HAS AN APRIL LIBRARY. THE PUMP HAS AN ICON THAT IT IS COMMUNICATING WITH THE SOFTWARE BUT THE SOFTWARE HAS NOT RECOGNIZED THE PUMP SINCE APRIL. THIS IS THE FIRST REPORTED ISSUE OF THIS KIND OF EVENT AT THIS FACILITY.======================MANUFACTURER RESPONSE FOR INFUSION PUMP - GENERAL PURPOSE, PLUM A+ (PER SITE REPORTER).======================SERVICE REP IS NOT AVAILABLE FOR 5 WEEKS. I KNOW YOU WILL NEED FOLLOW UP ON THIS PUMP, SO I WANTED TO LET YOU KNOW THERE WILL BE A DELAY IN GETTING THIS INFORMATION TO YOU. THANKS FOR YOUR PATIENCE. | 1
2749653 | WHILE USING TENACULUM FORCEPS DURING A ROBOTIC PROCEDURE,ONE PIECE OF METAL CAME APART INTO THE PATIENT'S ABDOMEN. SURGEON WAS ABLE TO RETREIVE THE PIECE FROM THE PATIENT'S ABDOMEN WITHOUT ANY TRAUMA TO THE PATIENT. | 0
2753159 | THREE PATIENTS HAD DISCREPANCY OF DATE OF EXAM. CORRECT WORKLIST, BUT INCORRECT ON RECORDED EXAMS. DETERMINED THAT THE OPERATING SYSTEM HAD CRASHED RENDERING THE SYSTEM UNUSABLE. MANUFACTURER'S TECHNICIAN COMPLETED A "LOAD FROM COLD" PROCEDURE AND RELOADED THE OPERATING SYSTEM AND APPLICATIONS INCLUDING PATCH SOFTWARE. RESTORED CUSTOMER SETTINGS AND PROTOCOLS FROM SAVED DISKS AND TESTED SYSTEMS FUNCTIONALITY. ALSO CONFIRMED THAT TIME/DATE WAS CORRECT FOR RECORDED TEST EXAMS. SYSTEM PASSED ALL FUNCTION TESTS AND IS READY FOR USE. WE HAVE NOT HAD THIS KIND OF PROBLEM/EVENT IN THE PAST. | 1
2759719 | PATIENT WAS IN RECOVERY ROOM AWAKENING FROM GENERAL ANESTHESIA WHEN STATED SHE "COULD FEEL EVERYTHING" DURING THE PROCEDURE. SHE STATED SHE FELT AND COULD REMEMBER "EVERYTHING." THE PATIENT WAS SUCCESSFULLY INTUBATED ON THE FIRST ATTEMPT. THE PATIENT RECEIVED VERSED, FENTYL, LIDOCAINE WITH PROPOFOL, AND ANECTINE JUST PRIOR TO GOING INTO SURGERY AND DURING ANESTHESIA INDUCTION. ANESTHESIA WAS INDUCED WITH VERSED, FENTANYL, LIDOCAINE WITH PROPOFOL, AND ANECTINE BEING GIVEN AT THAT TIME. SURGERY STARTED APPROXIMATELY 7 MINUTES LATER AT WHICH TIME TRACRIUM WAS GIVEN. THE PATIENT RECEIVED NEOSTIGMINE WITH ROBINOL APPROXIMATELY 30 MINUTES AFTER SURGERY STARTED, AND ANOTHER DOSE OF PROPOFOL AND FENTANYL 3 MINUTES LATER. WHEN ANESTHESIA ENDED, THE PATIENT RECEIVED ANOTHER DOSE OF FENTANYL. NO DEVICES WERE CONNECTED TO THE ANESTHESIA MACHINE. MEDICATIONS GIVEN DURING THE PROCEDURE WERE PROVIDED VIA IV AND INHALATION (GAS). ALL OF THE IV PUMPS USED DURING THE CASE WERE FULLY FUNCTIONAL AND HAD NO PROBLEMS. NO HOSES/TUBING CONNECTED TO THE ANESTHESIA MACHINE OR THE IV PUMPS WERE BENT, KINKED, OR DISCONNECTED UNEXPECTEDLY DURING THE CASE. THE PATIENT'S VITAL SIGNS WERE STABLE THROUGHOUT THE CASE AND DID NOT HAVE SIGNIFICANT CHANGE. THE ANESTHESIOLOGIST IS EXPERIENCED. LIKEWISE, THE ANESTHESIOLOGIST IS FAMILIAR WITH THIS EQUIPMENT. THE STAFF, PHYSICIANS, AND ANESTHESIOLOGIST IN THE CASE DO NOT KNOW WHAT MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. THE PATIENT HAS REFUSED FURTHER TREATMENT/CARE RELATED TO THE SURGICAL AWARENESS. PRE-OP DOCUMENTATION INDICATED THAT THE PATIENT WAS APPREHENSIVE ABOUT THE SURGERY. NO OTHER PATIENT CHARACTERISTICS WERE PRESENT THAT MAY HAVE CONTRIBUTED TO THE EVENT. THIS IS THE FIRST TIME WE HAVE HAD THIS PROBLEM WITH THIS DEVICE.======================MANUFACTURER RESPONSE FOR ANESTHESIA MACHINE, AESTIVA-5 (PER SITE REPORTER).======================GE INSPECTED MACHINE AND RELEASED IT BACK INTO SERVICE. | 1
2759807 | SPECIFIC USER CIRCUMSTANCES ALLOWED A QUARANTINE STATUS APPLIED BY BLOOD CENTER STAFF TO BE INADVERTENTLY REMOVED. THE QUARANTINE STATUS WAS OVERWRITTEN WHEN MOBILE DONOR COLLECTION DATA WAS ATTEMPTED TO BE SYNCHRONIZED FROM A JUMP DRIVE A SECOND TIME. BLOOD CENTER STAFF APPLY QUARANTINE STATUS FOR DONATED PRODUCTS WHEN REVIEW OF DOCUMENTS OR RELATED INFO REQUIRES FURTHER INVESTIGATION (SUCH AS POST DONATION INFO). INADVERTENT RELEASE OF THESE UNITS COULD POTENTIALLY RESULT IN ADVERSE EVENTS. CODE REGRESSION ANALYSIS PERFORMED, IN-VITRO TESTING PERFORMED. SOFTWARE FUNCTIONED AS DESIGNED. ROOT CAUSE DETERMINED TO BE RELATED TO SELECTION OF INCORRECT DESIGN INPUT. ALL UNITS ASSOCIATED WITH ISSUE WERE RECONCILED, NO DEATH OR SERIOUS INJURY OCCURRED AS A RESULT OF SOFTWARE DESIGN ISSUE. A COMMUNICATION WAS PROVIDED TO ALL USERS TO CEASE RE-SYNCHRONIZATION OF DONOR INFO. ANALYSIS OF THE ISSUE WAS CONDUCTED AND IT WAS DETERMINED THAT NO OTHER FUNCTIONS WERE IMPACTED. A SOFTWARE MODIFICATION WAS PERFORMED, VERIFIED AND VALIDATED. THE SCOPE OF THE TESTING INCLUDED THE SOFTWARE MODIFICATION ITSELF, ALL POTENTIALLY IMPACTED AREAS OF THE SOFTWARE, AND OTHER REGRESSION TESTS TO ENSURE THAT THE FIX WAS EFFECTIVE AND DID NOT ADVERSELY AFFECT THE SOFTWARE. THE SOFTWARE MODIFICATION WAS IMPLEMENTED ON (B)(4) 2012. NOTIFICATION TO RESUME RE-SYNCHRONIZATION WAS PROVIDED SUBSEQUENT TO IMPLEMENTATION AND ACKNOWLEDGEMENT BY MEDICAL LEADERSHIP. CURRENTLY, THE BECS SYSTEM IS USED ONLY BY A SINGLE CORPORATE ENTITY. AS SUCH, ALL USERS AFFECTED BY ISSUE ARE CURRENTLY WORKING WITH THE UPDATED SOFTWARE APPLICATION. ALL BLOOD PRODUCTS IMPACTED BY THIS ISSUE HAVE EITHER BEEN WITHDRAWN AND/OR APPROPRIATE REGULATORY NOTIFICATION PROVIDED. | 1
2759860 | PT WAS PLACED ON FETAL MONITOR WHEN STRIP STARTED SHOWING THAT FETUS WAS IN DISTRESS. PT WAS A PLANNED C/S BUT DUE TO THE DISTRESS PT HAD STAT C/S DONE. WHEN MONITOR RECHECKED WITHOUT BEING CONNECTED TO ANYTHING, STRIP RAN AS IF A FETUS WAS IN DISTRESS. (B)(4). | 1
2765090 | THE CUSTOMER REPORTED A PT CAME TO THE EMERGENCY ROOM WITH A BROKEN LEG. THE PT WAS CONNECTED TO A DASH MONITOR. THE PT EXPERIENCED A REACTION TO A THERAPY AND A DASH DID NOT ALARM FOR A LOW SPO2 VALUE. THE PT CODED AS PEA (PULSES ELECTRIC ACTIVITY). CPR WAS PERFORMED AND THE PT WAS RESUSCITATED. THE MEDICAL DEVICES IN THE UNIT WERE FOUND TO BE CONFIGURED BY THE USER TO ALARM AT A 20% VOLUME AND AT AN ADVISORY LEVEL FOR VIOLATION OF SPO2. THE PT HAS SINCE BEEN DISCHARGED WITH NO SEQUELA. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2765465 | DURING THE SAMPLE PROCESSING, THE AUTO-VALIDATION OF A PATIENT SAMPLE WAS BLOCKED BY THE ABX PENTRA ML, ASKING FOR A SLIDE REVIEW. ONCE THE MANUAL RUN FOR THE SLIDE ENTRY WAS ENTERED BY THE USER, THE FIRST DISCIPLINE (DIFF REPORT) WAS CORRECTLY UPLOADED ON THE LIS (LABORATORY INFORMATION SYSTEM). AT THAT STEP, THE USER CREATED A SECOND MANUAL RUN TO ENTER A NEW WBC VALUE. AFTER VALIDATING THIS MANUAL RUN, THE FIRST DISCIPLINE (DIFF REPORT) WAS AGAIN CORRECTLY UPLOADED TO THE LIS. THEN THE USER MANUALLY VALIDATED THE SECOND DISCIPLINE (CBC) COMING FROM THE INSTRUMENT RUN. THE ABX PENTRA ML UPLOADED THE FULL REPORT (CBC + DIFF) TO THE LIS. THE DATA RECEIVED AT THE FINAL STAGE BY THE LIS WERE NOT THOSE THAT WERE JUST VALIDATED FOR THAT PATIENT AND DISPLAYED CORRECTLY ON THE SCREEN AS NOTED ABOVE. INSTEAD, THE DATA WAS THAT OF A PREVIOUS RUN FROM THE SAME PATIENT HISTORY. THERE WAS NO IMPACT TO THE PATIENT AS THE TECHNICIAN NOTICED THE INCORRECT DATA BEFORE THE RESULTS WERE PROVIDED TO THE DOCTOR. ADDITIONALLY, THE TECHNICIAN CHECKED THE RESULTS ON TH ABX PENTRA ML SCREEN AND ON THE INSTRUMENT SCREEN. UPON VERIFICATION OF THE CORRECT RESULTS, THE TECHNICIAN PROVIDED THE CORRECT INFORMATION TO THE DOCTOR. THE ROOT CAUSE OF THE REPORTED ISSUE HAS NOT BEEN DETERMINED. HOWEVER, AFTER CHECKING THE DATABASE AND THE FILES EXCHANGED BY THE ABX PENTRA ML AND THE LIS, IT WAS IDENTIFIED THAT A "GHOST RUN" WAS CREATED BY THE ABX PENTRA ML BETWEEN THE LAST MANUAL RUN AND THE REPORT OF THE CONCERNED PATIENT SAMPLE RUN. WHILE A ROOT CAUSE FOR THE "GHOST RUN" CONTINUES TO BE INVESTIGATED. THE PRELIMINARY INVESTIGATION RESULTS HAS CONFIRMED THAT TWO SCENARIOS MAY OCCUR DEPENDING ON THE ABX PENTRA ML SOFTWARE VERSION AND WHETHER THE CONTROL (DISCIPLINE) FEATURE IS AVAILABLE OR NOT. SOFTWARE VERSION V8.0.X AND BELOW: THE COHERENCE OF ALL RESULTS BEING MANUALLY RE-UPLOADED TO THE LIS MUST BE SYSTEMATICALLY CHECKED. SOFTWARE VERSIONS V9.0.1 AND V9.0.2: IN THE CASE WHERE THE "DISCIPLINE" FEATURE IS USED, AFTER AN ENTRY OF RESULT IN A MANUAL RUN, THE COHERENCE OF ALL DATA UPLOADED TO THE LIS MUST BE CHECKED FOR THE PATIENT FILE. | 1
2769190 | THE PATIENT WAS UNDERGOING A ROBOTIC-ASSISTED TOTAL LAPAROSCOPIC HYSTERECTOMY WITH BILATERAL SALPINGECTOMY. DURING THE PROCEDURE, ONE OF THE WIRES TO THE FENESTRATED BIPOLAR FORCEPS BROKE. THIS WAS THE FOURTH USE OF THIS INSTRUMENT; THEY ARE DESIGNED TO BE USED TEN TIMES BEFORE BEING REPLACED. | 0
2773130 | DURING DAVINCI HYSTERECTOMY, MONOPOLAR CURVED SCISSOR WAS BEING USED. END OF PLASTIC TIP COVER ACCESSORY BURNED/MELTED AND ARCING OF CAUTERY OCCURRED - SPARK HIT UTERUS. INSTRUMENT REMOVED, NEW TIP COVER ACCESSORY PUT ON BUT PHYSICIAN STILL SAW ARCING OCCUR UP INSTRUMENT. NEW MONOPOLAR INSTRUMENT OPENED AND USED. ONE MONOPOLAR HAD 1 LIFE LEFT AND ONE HAD 2 LIVES LEFT. NO HARM TO PATIENT. | 0
2777288 | IT WAS REPORTED THAT A "COMPONENT BURNT INSIDE. BATTERY IS UNPLUGGED FROM BOARD" AND THE USER SMELLED SMOKE. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY. (B)(4). SIGMA EVALUATED THE DEVICE IN QUESTION. FLUID RESIDUE WAS OBSERVED INSIDE THE WIRELESS BATTERY MODULE AT THE BOTTOM OF THE CASE. SIGMA'S ENGINEERING INVESTIGATION IS IN PROGRESS. WHEN COMPLETE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2781373 | OUR AFFILIATE IS REPORTING TO US THAT DURING AN ARTHROSCOPIC SHOULDER REPAIR WITH THE USE OF A MITEK FMS DUO PUMP FOR FLUID MANAGEMENT, THE PUMP, UNCONTROLLABLY AND WITHOUT WARNING, FILLED THE JOINT SPACE WITH FLUID. THE PATIENT'S SHOULDER WAS DRAMATICALLY INCREASED IN SIZE. THE PROCEDURE WAS ABANDONED BECAUSE OF THIS. THIS IS ALL OF THE INFORMATION THAT HAS BEEN MADE AVAILABLE TO MITEK AT THIS TIME. MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT.><THE COMPLAINT DEVICE WAS RECEIVED AND EVALUATED. UPON RECEIPT, THE DEVICE WAS GIVEN A THOROUGH VISUAL AND FUNCTIONAL EXAMINATION. THE 1ST PORTION OF THE EXAMINATION WAS VISUAL: THIS VISUAL IS PERFORMED AS A MATTER OF COURSE TO DISCERN IF THE DEVICE HAS ANY OBVIOUS PHYSICAL DAMAGE, SOMETHING THAT MAY OR COULD HAVE BEEN A FACTOR IN CONTRIBUTING TO THE REPORTED EVENT; IT IS NOTED THAT THE DEVICE WAS RECEIVED IN GOOD PHYSICAL CONDITION. THE 2ND PORTION OF THE EXAMINATION WAS THE FUNCTIONAL AND ELECTRICAL TESTING: A BATTERY OF TESTS WAS PERFORMED TO ASSESS THE DEVICE'S OPERATIONAL STATUS. FUNCTIONALLY AND ELECTRONICALLY, THE DEVICE PERFORMED WELL WITHIN ITS DESIGN AND MANUFACTURED PARAMETERS; THE UNIT PASSED ALL DIAGNOSTIC TESTS, FUNCTIONAL TESTS, AND IS FULLY OPERATIONAL, NO FAULT OR PERFORMANCE ANOMALY COULD BE FOUND WITH THE FMS PUMP. THE ROOT OR UNDERLYING CAUSE FOR THE REPORTED EVENT IS NOT ATTRIBUTABLE TO THE FMS PUMP, BUT LIES ELSEWHERE. AT THIS POINT IN TIME, NO FURTHER ACTION IS WARRANTED; HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FORTHCOMING INFORMATION RECEIVED THAT IS PERTINENT AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
2781924 | CLINICAL CHEMISTRY AUTOMATED LINE CONTINUES TO HAVE MULTIPLE FAILURES RESULTING IN SIGNIFICANTLY SLOWED TURNAROUND TIME FOR CHEMISTRY TESTS. SIEMENS SERVICE CONTINUES TO PROMISE REPAIRS BUT DOES NOT DELIVER. THE SIGNIFICANT DOWNTIME OF THE STREAMLAB SYSTEM SLOWS OVERALL TURNAROUND TIME FOR CHEMISTRY TESTS AND RESULTS IN SLOWER DIAGNOSES FOR PATIENTS, INCLUDING EMERGENT DIAGNOSES FOR TRAUMA PATIENTS. DATES OF USE: (B)(6) 2012. | 0
2789578 | IT WAS REPORTED THAT THE APPLICATION DISPLAYS THE "WRONG" VALUE FROM A COMPETITOR SAVED SESSION IMPORTED WITH THE RETRIEVE FROM PROGRAMMER (RFP) FEATURE IN THE APPLICATION. THE APPLICATION DISPLAYS "A/V PACE REFRACTORY" INSTEAD OF "PVARP" IT DISPLAYS THE ATRIAL REFRACTORY PERIOD UNDER THE PROGRAMMING TAB AS "190" BUT THE CLINICIAN ASSERTS THAT IT SHOULD BE "275." THE CLINICIAN STATED THIS HAPPENS FOR ALL NEW COMPETITOR DEVICES. THE CLINICIAN FAXED THE PROGRAMMER REPORT, A COPY OF THE SAVED SESSION FILE AND SCREEN SHOTS FROM THE APPLICATION. AN EMAIL WAS SENT TO THE CLIENT STATING THE APPLICATION DOES NOT CURRENTLY SUPPORT THE COMPETITOR DEVICE. FURTHER INVESTIGATION WILL BE CONDUCTED TO RESOLVE THE ISSUE. NO PATIENT COMPLICATIONS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2794050 | THE (B)(6) BECAME AWARE OF SEVERAL CASES OF SERIOUS INVASIVE INFECTIONS (FUNGEMIA) DUE TO 'GEOTRICHUM CLAVATUM' BY HEMATOLOGY PATIENTS. SINCE (B)(6) 2012, 27 CASES OF INFECTION WERE REPORTED, INCLUDING 17 DURING AN EPIDEMIC PEAK OCCURRED IN (B)(6) 2012. INVESTIGATIONS IN SEARCH OF A COMMON SOURCE WERE CONDUCTED BY (B)(6) ON THE FIRST 22 CASES. THEIR INVESTIGATION REVEALED THAT ALL CASES RECEIVED LABILE BLOOD PRODUCTS AND IN PARTICULAR BLOOD PRODUCTS PREPARED FROM SETS WITH REFERENCE CATALOG # 80484. THERE WERE 11 DIFFERENT LOTS INVOLVED IN THE REPORTED CASES. THIS REPORT COVERS THE INCIDENT THAT OCCURRED ON LOT # 01U3218, USED AT THE CUSTOMER SITE (B)(6). THE OTHER INCIDENTS HAVE BEEN REPORTED IN MEDWATCH REPORT #S 1722028-2012-00802 - 1722028-2012-00805, 1722028-2012-00807 - 1722028-2012-00812. PATIENT INFORMATION IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR DEATH OR INJURY OCCURRED. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. ROOT CAUSE: IT CAN BE CONCLUDED, AND SUPPORTED BY SEVERAL PAPERS REVIEWED, THAT THE INCIDENCE OF INVASIVE FUNGAL INFECTIONS IN PATIENTS WITH HEMATOLOGICAL MALIGNANCIES HAS RISEN OVER THE LAST TWO DECADES, MOST LIKELY AS A RESULT OF THE INCREASED USE OF INTENSIVE CYTOTOXIC THERAPY, ALLOGENEIC BLOOD STEM CELL TRANSPLANTATION, AND IMMUNOSUPPRESSIVE THERAPY. IT MUST BE CONCLUDED THAT THE TRIMA ACCEL SETS, WHICH ARE STERILIZED BY ETHYLENE OXIDE WITH A 12 LOG REDUCTION OF RESISTANT BACTERIA ENDOSPORES, ARE NOT THE SOURCE OF FUNGAL CONTAMINATION. CORRECTIVE/PREVENTIVE ACTION: BASED ON THE INVESTIGATION RESULTS, THE TRIMA ACCEL SET IS NOT THE SOURCE OF FUNGAL CONTAMINATION. AS A SECONDARY ACTIVITY, TERUMO BCT INITIATED ADDITIONAL MONITORING TO TEST FOR THE PRESENCE OF GEOTRICHUM CLAVATUM, WHICH NO PRESENCE HAS BEEN FOUND.><(B)(4). INVESTIGATION: A LITERATURE SEARCH REGARDING INVASIVE "GEOTRICHUM CLAVATUM" FOUND IN HEMATOLOGY WAS PERFORMED IN RELATION TO THIS INCIDENT. THE LITERATURE REVIEW REVEALED THAT "GEOTRICHUM CLAVATUM" INFECTIONS ARE VERY RARE. THE MAJORITY OF CASES ARE IN NEUTROPENIC PATIENTS HOSPITALIZED FOR TREATMENT OF ACUTE LEUKEMIA MYELOBLASTOMA. THEY WERE SUBJECTED TO PROPHYLACTIC OR EMPIRICAL ANTIFUNGAL TREATMENTS INCLUDING AZOLE AND CASPOFUNGIN, TO WHICH "G. CLAVATUM" IS NATURALLY RESISTANT. "GEOTRICHUM CLAVATUM" HAS ONLY BEEN REPORTED IN (B)(6). TAKING INTO CONSIDERATION THAT THE PARTS USED FOR CONSTRUCTION OF THE TRIMA DISPOSABLE SETS (INCLUDING THE (B)(6) SET) ARE COMMON TO ALL SETS MANUFACTURED BY TERUMO BCT AND ARE DISTRIBUTED WORLD-WIDE, IF THE MANUFACTURING PROCESS CAUSED FUNGAL CONTAMINATION, WE WOULD HAVE SEEN THIS REPORTED FROM MULTIPLE SOURCES. THE TERUMO BCT TRIMA PRODUCT CATALOG 80484 IS TERMINALLY STERILIZED USING AN ETHYLENE OXIDE STERILIZATION PROCESS. THE ONLY KNOWN FUNGUS (MOLD) RESISTANT TO ETHYLENE OXIDE IS "PYRONEMA DOMESTICUM". HOWEVER, AT THE LEVELS OF ETO USED DURING OUR STERILIZATION CYCLE, EVEN THIS MICROORGANISM IS DESTROYED. DEVICE HISTORY RECORDS WERE REVIEWED. NO ABNORMALITIES WERE NOTED IN THE MANUFACTURE OR STERILIZATION PROCESS OF THE DISPOSABLE SET. STORAGE AND SHIPPING METHODS WERE ALSO REVIEWED, WITH NO ABNORMALITIES FOUND. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2795605 | PATIENT WAS SUPPOSED TO HAVE WATER COOLED RFL (RADIOFREQUENCY LESION ABLATION)OF SI(SACRO-ILIAC)JOINT UNDER (MONITORED ANESTHESIA)MAC. PATIENT WAS POSITIONED, NEEDLES WERE PLACED FOR THE PROCEDURE AND RFL MACHINE ALTHOUGH TESTED OKAY, DID NOT FUNCTION PROPERLY AS THERE WAS A CABLE THAT WAS DEFECTIVE AND WE DID NOT HAVE A REPLACEMENT FOR THE CABLE. ALL OTHER COMPONENTS FOR THE SYSTEM WERE CHECKED AND CHANGED. | 0
2795758 | HARMONIC WORKING, THEN ERROR MESSAGE AND WOULD NOT WORK. | 0
2798596 | CLINICAL CHEMISTRY AUTOMATED LINE CONTINUES TO HAVE MULTIPLE FAILURES RESULTING IN SIGNIFICANTLY SLOWED TURNAROUND TIME FOR CHEMISTRY TESTS. SIEMENS SERVICE CONTINUES TO PROMISE REPAIRS BUT DOES NOT DELIVER. THE SIGNIFICANT DOWNTIME OF THE STREAMLAB SYSTEM SLOWS OVERALL TURNAROUND TIME FOR CHEMISTRY TESTS AND RESULTS IN SLOWER DIAGNOSES FOR PTS, INCLUDING EMERGENT DIAGNOSES FOR TRAUMA PTS. DATES OF USE: (B)(6) 2012. | 0
2803857 | THIS IS REGARDING SOFTWARE USED IN AN EMERGENCY COMMUNICATION CENTER BY DISPATCHERS TO PINPOINT CORRECT LOCATIONS FOR DISPATCH OF EMERGENCY VEHICLES. A 911 CALL FROM AN ACCIDENT SCENE RESULTED IN DISPATCH OF AN AMBULANCE AND HELICOPTER TO AN INCORRECT LOCATION. THE STREET ADDRESS AND CITY NAME HAD BEEN ENTERED CORRECTLY INTO THE SYSTEM WHICH RETURNED TWO IDENTICAL ADDRESSES THAT RANKED EQUALLY. THE FIRST ADDRESS OF THE TWO LISTED WAS SELECTED BECAUSE BOTH RANKED 99% AND BOTH ADDRESSES IN THE GEOCODER WERE IDENTICAL. UPON ARRIVAL TO THAT ADDRESS, AN ACCIDENT SCENE WAS NOT ABLE TO BE FOUND. IT WAS THEN THAT THE DISPATCHER REALIZED THE TWO EXACT ADDRESSES WITH THE SAME MATCH RANKING WERE PHYSICALLY DIFFERENT LOCATIONS. DUE TO ERRONEOUS MAP DATA THAT HAD BEEN SUPPLIED FROM THE COUNTY AND THE DEFAULT SORTING CONFIGURATION OF THE COLUMN, THE ADDRESS THAT CAME UP FIRST AND CHOSEN BY THE DISPATCHER, WAS INCORRECT. AS A RESULT, THE RESPONSE TO THE SCENE WAS DELAYED. THE PATIENT EXPIRED AT THE SCENE, HOWEVER, DUE TO THE PATIENT'S EXTENSIVE INJURIES, IT WAS DETERMINED THE DELAY DID NOT AFFECT THE OUTCOME. A REVIEW OF THIS IDENTIFIED THERE WAS A DIFFERENCE IN INTERSECTION/RANGE LIMITS BETWEEN THE TWO ADDRESSES BUT THE ACTUAL GEOGRAPHICAL LOCATION OF THE ADDRESS WAS WITHIN BOTH THE INTERSECTION/RANGE LIMITS. THE MANUFACTURER OF THE SOFTWARE HAS BEEN CONTACTED AND EXTENSIVE DISCUSSIONS OCCURRED REGARDING THE FUNCTIONALITY OF THE SOFTWARE AND THE INTEGRATION OF MAP OVERLAYS WITH THEIR SOFTWARE. SINCE THE EVENT DESCRIBED ABOVE THE ECC HAS ENCOUNTERED THE GEOCODER PROVIDING AN ADDITIONAL ADDRESS WITH IDENTICAL STREET INFORMATION, IDENTICAL 99% RANKING AND IDENTICAL INTERSECTION/RANGE LIMITS BUT WITH TWO PHYSICALLY DIFFERENT LOCATIONS. THIS SITUATION DID NOT RESULT IN ANY DELAYS IN PATIENT CARE. ERRONEOUS MAP DATA SUPPLIED BY THE COUNTY CONTRIBUTED TO THE GEOCODER MATCHES.OUR CONCERNS ARE:1. TWO ADDRESSES WITH DIFFERENT INTERSECTION/RANGE LIMITS WERE RANKED EQUALLY. 2. THE ADDRESS THAT WAS THE "BEST" MATCH, WAS LISTED SECOND IN BOTH SITUATIONS DESCRIBED HERE. 3. TWO DISTINCT PHYSICAL ADDRESSES CAN APPEAR IN THE GEOCODER WITH EXACT ADDRESS INFORMATION, EXACT INTERSECTION/RANGE LIMITS AND EQUAL RANKING. 4. THE "PREVIEW" FUNCTION WILL SHOW DIFFERENT GEOGRAPHIC LOCATIONS IN THESE SITUATIONS BUT IS AN OPTIONAL FIELD. 5. DEFAULT GEOCODER CONFIGURATION OF DATA WAS NOT PROVIDED TO ECC STAFF. 6. A MANUFACTURER PUBLISHED USERS GUIDE FOR THE GEOCODING TOOL HAS NOT BEEN PROVIDED TO THE ECC.SINCE THIS EVENT, ZOLL HAS PROVIDED THE RECOMMENDED GEOCODER CONFIGURATION. THE GEOCODER HAS BEEN CHANGED BASED ON THE RECOMMENDATION. A SUPER USER OF THE DISPATCH APPLICATION HAS BEEN USING THIS CONFIGURATION FOR OVER THREE WEEKS. WITH THE RECOMMENDED CONFIGURATION THE BEST ADDRESS FOR BOTH EVENTS IS RANKED AS THE FIRST MATCH. | 1
2803906 | AT THE BEGINNING OF THE SURGERY, WHILE EXPLORING THE PATIENT'S ABDOMEN, THE SOUND OF THE MONOPOLAR ESU WAS HEARD BY THE SURGEON, STAFF RN, AND DAVINCI REPRESENTATIVE. THE SURGEON STATED THAT SHE WASN'T DEPRESSING THE BIPOLAR FOOT PEDAL. PHYSICIAN PHONE STATEMENT NOTED THAT THIS INCIDENT OCCURRED AT THE VERY START OF THE CASE, AND THE FOOT PEDAL WAS NOT BEING ACTIVATED WHEN THE ESU WAS HEARD GOING OFF BY ITSELF; OBSERVE THE DAVINCI REP UNPLUGGING THE FOOT PEDALS CONNECTIONS FROM THE ROBOT AT THE REAR OF THE ESU. HEARD IT ACTIVATE AGAIN AFTER THIS WAS DONE. NURSE WHO WROTE UP INCIDENT STATEMENT SAYS SHE WAS ON THE OTHER SIDE OF THE ROOM, BUT DID HEAR THE ESU GOING OFF. SAYS SHE WAS THE PERSON WHO UNPLUGGED THE CABLES FROM THE FRONT OF THE DEVICE. SHE WAS NOT QUITE SURE IF THIS HAPPENED AT THE BEGINNING OR THE MIDDLE OF THE CASE. NURSE WITNESS STATEMENT: TWO DOCTORS WERE USING THE DAVINCI ROBOT CONSOLE; THEY WERE USING MONO-POLAR SETTING AT THAT MOMENT, HEARD WHEN PHYSICIAN SAY "WHY IS THE MONO-POLAR STILL GOING," SAID THAT SURGEON SAW SMOKE IN THE FIELD.SAW DAVINCI REP UN-PLUGGED THE CABLES FROM THE BACK OF THE FORCE FX ESU, SWITCHED OUT FORCE FX TO THE NEWER FORCE TRIAD ESU, CONTINUE WITH CASE.AS A RESULT OF WHAT APPEARS TO BE AN UNINTENDED DELIVERY OF MONOPOLAR ENERGY THROUGH A BIPOLAR MARYLAND CLAMP, THE PATIENT SUSTAINED A PERFORATED BOWEL, AND DESPITE CAREFUL EXAMINATION OF THE SURGICAL AREA, IT WAS NOT DISCOVERED UNTIL THE PATIENT CAME BACK TO THE HOSPITAL APPROXIMATELY A WEEK LATER WITH SIGNS OF INFECTION. THIS REQUIRED AN EXPLORATORY LAPAROTOMY, WHICH REVEALED A PERFORATED BOWEL. A PORTION OF COLON WAS REMOVED AND SENT TO SURGICAL PATHOLOGY. PATIENT SUBSEQUENTLY DEVELOPED BILATERAL PULMONARY EMBOLI AND REMAINS HOSPITALIZED TO THIS DAY. | 0
2804961 | OUR AFFILIATE IS REPORTING TO US THAT DURING AN ARTHROSCOPIC SHOULDER REPAIR WITH THE USE OF A MITEK FMS DUO PUMP FOR FLUID MANAGEMENT, THERE WERE SOME SWELLING ISSUES FOR THE PATIENT. IT APPEARS THAT THE PRESSURE ON THE PUMP FLUCTUATED; WHEN THE "LAVAGE" BUTTON WAS ACTUATED THE PRESSURE INCREASED 50% BUT DID NOT RETURN TO THE DEFAULT SETTING. THIS HAPPENED THROUGHOUT THE PROCEDURE RESULTING IN TOO MUCH FLUID BEING PUMPED INTO THE SHOULDER. THERE WAS EXTRAVASATION AND POSSIBLY SOME INTERVENTION NEEDED (MASSAGE OR AN EXTRA PORTAL, IT IS NOT YET DEFINED). THIS IS ALL OF THE INFORMATION MADE AVAILABLE TO MITEK AT THIS TIME AS THE SURGEON IS ON HOLIDAY, REACH OUTS FOR FURTHER INFORMATION AND CLARITY WILL GO OUT THE WEEK OF (B)(6) 2012 WHEN THE SURGEON RETURNS TO ROUTINE. MITEK IS AT THIS POINT IN TIME IN THE INFORMATION GATHERING MODE. WHEN ALL THAT CAN BE HAD, IS HAD AND THOROUGHLY INVESTIGATED AND EVALUATED, THOSE RESULTS WILL BE THE SUBJECT MATTER IN A FOLLOW-UP REPORT. AWAITING RETURN OF COMPLAINT DEVICE. | 0
2806103 | I HAD AN ANGIOGRAPHY BY ROUTE OF THE FEMORAL ARTERY IN (B)(6) 2008. AFTER THE PROCEDURE, THE NURSE HANDED ME A BROCHURE AND SAID THAT I RECEIVED THE STARCLOSE DEVICE. NO CONSENT WAS GIVEN FOR THIS DEVICE PRIOR TO THE PROCEDURE. AFTER ABOUT 30 MINUTES WHEN I GOT UP, MY RIGHT LEG WAS MOSTLY NUMB. I ASKED IF HIS WAS NORMAL AND WAS TOLD BY THE NURSE: IT CAN HAPPEN. I HAD DIFFICULTY WALKING TO THE CAR AFTER BEING RELEASED. SEVERAL HOURS LATER, I STARTED HAVING SEVERE PAIN DOWN THE INSIDE OF MY RIGHT LEG. SINCE THE PAIN PERSISTED FOR SEVERAL DAYS, A SONOGRAM WAS ORDERED TO SEE IF I HAD ANEURYSM AT THE SITE OF INSERTION. NONE WERE FOUND. AS A PHYSICIAN, I KNEW THAT MY FEMORAL NERVE WAS BEING IMPINGED UPON, SO I HAD A MRI DONE. IT SHOWED THAT THE STARCLOSE DEVICE WAS IN CLOSE PROXIMITY TO THE FEMORAL NERVE. IT BECAME OBVIOUS THAT THE DEVICE WAS COMPRESSING UPON THE FEMORAL NERVE WITH CERTAIN MOVEMENTS OF MY LEG. SPECIFICALLY WITH ABDUCTION, EXTERNAL ROTATION AND HIP FLEXION. I SAW THE CARDIOLOGIST WHO PERFORMED THE PROCEDURE, DR (B)(6), AND HE SAID THAT THIS PROBLEM WAS NOT POSSIBLE AND DISMISSED THE COMPLICATION. I ENDED UP SEEING A VASCULAR SURGEON IN 2009 TO DISCUSS RESECTING THE FEMORAL ARTERY AND REMOVING THE STARCLOSE DEVICE, BUT HE SAID THAT IT MIGHT BE LIKE JUMPING FROM THE FRYING PAN INTO THE FIRE. IN OTHER WORDS REMOVING IT WAS NOT POSSIBLE. AS A RESULT, I APPARENTLY WILL HAVE TO LIVE WITH LIMITED ACTIVITY FOR THE REST OF MY LIFE, ENDURE A LOT OF PAIN BECAUSE DOCTORS LIKE DR (B)(6) FROM (B)(6) DON'T WANT TO TARNISH THEIR REPUTATION BY REPORTING COMPLICATIONS. REPORTING THESE PROBLEMS IS HOW DANGEROUS PRODUCTS GET REMOVED FROM THE MARKET. UNFORTUNATELY, FEAR OF LAWSUITS AND PERHAPS EGOIC BEHAVIOR, ALONG WITH FORMS OF KICKBACKS MAY PREVENT PROPER REPORTING. HOPEFULLY, MY TIME SPENT ON THIS TOPIC WILL HELP TO PREVENT OTHERS FROM SUFFERING LIKE I HAVE AND WILL. NO OTHER COMPLICATING CONDITIONS EXIST HAT MAY HAVE INFLUENCED THE OUTCOME OF THIS PROCEDURE. THE STARCLOSE DEVICE WAS SIMPLY PLACED AGAINST THE FEMORAL NERVE AND IS NOW PERIODICALLY COMPRESSING UPON IT CAUSING A PERMANENT IATROGENIC FEMORAL NEUROPATHY. DIAGNOSIS OR REASON FOR USE: CHEST PAIN LEAD TO ANGIOGRAM. | 0
2811200 | IT WAS REPORTED THAT THERE WAS A DISCREPANCY IN THE PATIENT'S BATTERY VOLTAGE MEASUREMENT BETWEEN THE SYSTEM AND THE APPLICATION FOR THE SAME TRANSMISSION. THE ISSUE WILL BE ESCALATED FOR FURTHER INVESTIGATION. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
2814143 | SINCE EARLY SUMMER 2012, THE PICIS ANESTHESIA MANAGER APPLICATION WOULD PERIODICALLY FREEZE ON VARIOUS PC'S (PERSONAL COMPUTERS) AT DIFFERENT TIMES OF THE DAY AND NOT EVERYDAY, NOT EVEN EVERY WEEK. ON OCCASION COMPUTERS WOULD NOT DISPLAY A PERIODIC BLOOD PRESSURE VALUES FROM THE DATABASE SERVER, BUT ALL OTHER PARAMETERS WOULD COME ACROSS THE INTERFACE FINE. IN LATE FALL OF THIS YEAR, WE EXPERIENCED A PROBLEM WITH PICIS POPULATING DATA FIELDS WITH INCORRECT PATIENT DATA. THE IT DEPT WAS INVOLVED AND FOUND A PC IN OR7 HAD ITS IP (INTERNET PROVIDER) ADDRESS CHANGED AND THE DNS (DOMAIN NAME SYSTEM) FILE WAS NOT UPDATED ON SERVER. IT SHOULD BE NOTED THE HOSPITAL IS JOINING A NEW DOMAIN AND ALL IP ADDRESSES WILL EVENTUALLY CHANGED. THIS PARTICULAR WORKSTATION WAS CHANGED OUT OF SCHEDULE. OR7 COMPUTER WAS TAKEN OFF-LINE AND THE CORRECT DATA FOR EACH PATIENT RESUMED IN PICIS DATA FIELDS. THE DNS FILE WAS UPDATED APPROPRIATELY AND OR7 COMPUTER WAS PUT BACK ONLINE WITHOUT ISSUE. AS FOR OCCASIONAL WORK STATION FREEZING ISSUE, THE OPERATING SYSTEM WAS INSTALLED ON ALL HOSPITAL COMPUTERS IN EARLY JUNE. IT IS SUSPECTED THAT THE OPERATING SYSTEM IS NOT COMPATIBLE WITH PICIS VERSION 8.1 AND CAUSING OCCASIONAL PC FREEZES. NO PATIENTS WERE HARMED. | 1
2817259 | DURING THE RELEASE OF THE CAPSULE THE CAPSULE HOLDER BROKE AND THE CAPSULE CONTINUED DOWN THE SMALL BOWEL INSIDE THE CAPSULE HOLDER. AN ATTEMPT WAS MADE TO RETRIEVE THE SHEATH BUT WAS UNSUCCESSFUL DUE TO BEING IN THE SECOND PORTION OF DUODENUM. LIMITED STUDY DUE TO MALFUNCTION OF CAPSULE RELEASER. | 0
2817567 | IT WAS REPORTED THAT EVLW (EXTRA VASCULAR LUNG WATER) VALUES WERE NOT REFLECTIVE OF THE PATIENT'S CLINICAL STATUS. WHEN THE FIRST SERIES OF BOLUSES WERE PERFORMED AT THE BEGINNING OF CORONARY BYPASS SURGERY, THE CARDIAC OUTPUT VALUES CORRELATED WITH THE PATIENT'S CONDITION; HOWEVER, SUSPECT HIGH VALUES FOR EVLW (22 ML/KG) WERE ALSO OBSERVED. NO ISSUES WERE NOTED REGARDING BOLUS TEMPERATURES OR DELTA OF TEMPERATURES, AND NO ALARMS WERE REPORTED. THE DEVICE WAS REMOVED FROM SERVICE AND RETURNED TO EDWARDS FOR EVALUATION. THE REFERENCED DEVICE HAS NOT BEEN RETURNED.THE REFERENCED DEVICE WAS NOT RETURNED FOR EVALUATION, NOR WAS A DOWNLOAD OF THE CLINICAL DATA PROVIDED. WITH SUCH LIMITED INFORMATION, IT IS NOT POSSIBLE TO DETERMINE IF A PRODUCT FAILURE OCCURRED, OR IDENTIFY POTENTIAL CONTRIBUTING CLINICAL FACTORS. THE FILE WILL BE RE-OPENED AND A FOLLOW UP COMMUNICATION SUBMITTED SHOULD THE INFORMATION BECOME AVAILABLE.><THE CUSTOMER'S COMPLAINT OF EVLW (EXTRA VASCULAR LUNG WATER) VALUES FAILING TO CORRELATE TO THE PATIENT'S STATUS WAS UNABLE TO BE CONFIRMED THROUGH EVALUATION. EXAMINATION OF THE RETURNED EV1000 WAS UNABLE TO DETECT ANY FAILURE OF THE DEVICE TO FUNCTION AS INTENDED. THE EV1000 WAS TESTED PER THE ESTABLISHED SOP REQUIREMENTS USING A PLUM SIMULATOR AND THE RETURNED SYSTEM AC ADAPTOR. THE DEVICE APPROPRIATELY ZEROED CVP/AP AND MONITORED CO, SVI, AND SCVO2 FOR OVER 5 HOURS. DURING THE 5-HOUR BURN-IN TESTING, RESULTS WERE COLLECTED AND CONFIRMED TO BE WITHIN EXPECTED RANGE. BT/IT RANGE WAS ALSO ANALYZED AND FOUND TO BE WITHIN PRODUCT SPECIFICATIONS. THE SYSTEM WAS TURNED ON AND OFF OVER 5 TIMES AND POWERED UP NORMALLY. NO OTHER DEFECTS OR DAMAGE WAS OBSERVED OR NOTED ON THE DATABOX OR THE PC PANEL DURING THE VISUAL INSPECTION. THE MONITOR WAS MANUFACTURED 28-APRIL-2011 AND REVIEW OF THE DEVICE HISTORY RECORD SUPPORTS THAT THERE WERE NO NON-CONFORMANCES NOTED DURING THE MANUFACTURING OF THE DEVICE FOR ANY REASON. IT IS UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE CUSTOMER'S EXPERIENCE, AS NO FAULT COULD BE DETERMINED AND THERE WAS NO EVIDENCE OF ANY FAILURE OF THE REFERENCED DEVICE TO FUNCTION AS INTENDED.><THE DEVICE EVALUATION WAS NOT YET COMPLETED AT THE TIME OF THIS SUBMISSION. EVALUATION, DEVICE HISTORY RECORD REVIEW, INVESTIGATION AND CONCLUSION WILL BE COMMUNICATED IN A FOLLOW-UP SUBMISSION. | 1
2820202 | WHILE USING THE NEEDLE DRIVER, IT STOPPED WORKING. IT APPEARED THE CABLE WAS LOOSE AT THE TIP OF THE INSTRUMENT. | 0
2820731 | PT WAS ATTACHED TO TELEMETRY BOX; MONITOR DISPLAYED BATTERY NEEDED REPLACEMENT. WHEN NURSE WENT TO CHANGE BATTERY IT WAS HOT. THE HEAT MADE ONE OF THE TWO BATTERIES COVER LOOSEN. PT STATES THAT THE BOX DID FEEL WARM, BUT HE DID NOT HAVE ANY SENSATION OF A BURN. PT'S CHEST AREA EXAMINED AND NO INDICATION OF BURN. NURSE STATES THAT BATTERY DID FEEL HOT WHILE SHE WAS ATTEMPTING TO REPLACE IT, BUT THAT SHE DID NOT EXPERIENCE BURN. NURSE'S NAME (B)(6) WORK ORDER PLACED ON THIS TELEMETRY BOX. | 0
2820889 | AFTER POSITIONING PT LATERAL NON-INVASIVE BLOOD PRESSURE-NIBP-STOPPED CYCLING. UNABLE TO USE MANUAL OR AUTO ICONS ON MONITOR; GRAYISH IN COLOR/FROZEN. ABLE TO ACCESS NIBP MODULE SCREEN BUT ALL ICONS WERE FROZEN. AFTER REBOOTING ALL ICONS FOR NIBP IT FROZE AGAIN. AFTER APPROXIMATELY 10 MINUTES, THE STAFF WAS ABLE TO GENERATE BLOOD PRESSURE AFTER MANUAL ICON BECAME UNFROZEN IN NIBP MODULE SCREEN AND SELF CONVERTED TO AUTO MODE. NO ADVERSE OUTCOMES TO PATIENT. NUMBER: (B)(4) -GE HEALTHCARE-. | 1
2822291 | A TECH AT THE LAB INDICATED THAT TISSUES IN 5 CASSETTES FROM 2 CASES PROCESSED ON THE SUKURA VIP 5 INSTRUMENT WERE NOT PROCESSED ADEQUATELY, HOWEVER, THE REMAINDER OF TISSUES IN THIRTY-FOUR (34) CASSETTES WERE PROCESSED CORRECTLY. BASED ON CUSTOMER SURVEY, IT WAS DETERMINED THAT THE PUMP HAD A CRACKED DIAPHRAGM. THE PUMP DIAPHRAGM IS INTENDED TO BE REPLACED EVERY 12 MONTHS AND ITS REPLACEMENT IS INCLUDED AS PART OF THE RECOMMENDED ANNUAL PREVENTIVE MAINTENANCE CHECKLIST. THE LAB IN QUESTION DOES NOT HAVE A SVC ARRANGEMENT WITH SAKURA FINETEK, USA, NOR DO OUR RECORDS INDICATE THAT SAKURA FINETEK WAS CONTACTED BY THE SITE TO PERFORM SUCH SERVICES. THEREFORE, SAKURA FINETEK, USA, CANNOT DOCUMENT THAT THE RECOMMENDED PREVENTIVE MAINTENANCE(S) HAD BEEN PERFORMED AND IF SO, WHETHER THE INDIVIDUAL PERFORMING THE SERVICE WAS PROPERLY TRAINED. THE FACT THAT ONLY 5 CASSETTES FROM 2 CASES WERE NOT PROCESSED CORRECTLY AND MIXTURE OF SMALL AND LARGE SAMPLES WERE PROCESS AT THE SAME TIME RAISES CONCERN OF NON-INSTRUMENT FACTORS. IT IS BEST CLINICAL PRACTICE TO NOT MIX LARGE AND SMALL TISSUE SAMPLES DURING PROCESSING. | 0
2822305 | ON (B)(6), THE CUSTOMER CALLED INTO BAXA TECHNICAL SUPPORT TO STATE A PT HAD RECEIVED AN UNDER DOSE OF CALCIUM GLUCONATE. THE ERROR IN THE PT¿S ORDER WAS A RESULT OF THE PHARMACIST ENTERING THE INCORRECT CALCIUM CONCENTRATION IN THE FORMULARY; 100 MG/ML INSTEAD OF 9.3 MG/ML. THEREFORE, WHEN THE PHARMACIST REQUESTED 35 MG OF CALCIUM, ONLY 3.5 MLS WAS DELIVERED TO THE PT. UPON DISCOVERY OF THE USER ENTRY ERROR, THE PT¿S TOTAL PARENTERAL NUTRITION (TPN) THERAPY WAS HALTED. A REVIEW OF THE CUSTOMER¿S FORMULARY REVEALED THAT THE ABACUS SOFTWARE OPERATED AS INTENDED. NO ADVERSE EVENTS OR PATIENT INJURY OCCURRED AS A RESULT OF THE PARTIAL INFUSION. RESULTS - BAXA TECHNICAL SUPPORT REVIEW OF THE FACILITY¿S DATABASE SHOWS THAT UPON THE INSTALLATION OF THE ABACUS SOFTWARE, THE CALCIUM GLUCONATE 100 MG/ML INGREDIENT WAS NOT IN THE FORMULA. DATA SHOWS THAT WHEN THE CALCIUM GLUCONATE 100 MG/ML INGREDIENT WAS ADDED TO THE CUSTOMER¿S FORMULA BY THE CUSTOMER AT SOME POINT FOLLOWING THE BAXA SUPPLIED INSTALLATION AND CONFIGURATION; IT WAS CREATED AND SET TO ¿UNAVAILABLE¿, WHICH PREVENTED USERS FROM ACCESSING THE INGREDIENT. ONLY IN (B)(6) 2011, WAS THE ERRANT ENTRY MADE ¿AVAILABLE TO ROUTINE USERS AND WAS NOT REPORTED TO BAXA UNTIL THIS (B)(6) INCIDENT. THIS SUBMISSION IS BEING RETROACTIVELY SUBMITTED SINCE THE INITIAL SUBMISSION WAS MISTAKENLY SUBMITTED TO THE FDA ELECTRONIC SUBMISSIONS GATEWAY TEST ENVIRONMENT RATHER THAN THE PRODUCTION ENVIRONMENT. DEFECT AWARENESS HAS BEEN CONDUCTED FOR ALL EMPLOYEES SUBMITTING ELECTRONIC SUBMISSIONS TO PREVENT FUTURE RECURRENCE OF THIS ISSUE. IN ADDITION, WE WILL ALSO SEND PAPER COPIES TO ENSURE ALL SUBMISSIONS ARE SUCCESSFULLY RECEIVED PER REGULATORY REQUIREMENTS. DURING A REVIEW OF THE FACILITY¿S DATABASE THE FOLLOWING WAS DISCOVERED. WHEN THE USER FACILITY INITIALLY ADDED THE CALCIUM GLUCONATE 100 MG/ML INGREDIENT TO THEIR FORMULA, IT WAS ENTERED WITH THE INCORRECT CONCENTRATION OF CALCIUM (100 MG/ML INSTEAD OF 9.3 MG/ML) BUT WAS MADE UNAVAILABLE FOR COMPOUNDING. THE INCORRECT INGREDIENT APPEARS IN TWO OF THE FACILITY¿S TEMPLATES, ¿ION-BALANCING¿ AND ¿NICU FLUIDS (TEST)¿. EVERY TPN THAT WAS CREATED USING EITHER TEMPLATE WAS COMPOUNDED WITH A LOWER THAN ORDERED AMOUNT OF CALCIUM GLUCONATE SINCE(B)(6) 2011. THIS EVENT ON (B)(6) IS THE ONLY INSTANCE OF MEDICAL INTERVENTION, TO DATE, NO FURTHER PATIENT INVOLVED CASES HAVE BEEN REPORTED FROM THIS FACILITY RELATED TO THIS ERROR. BAXA TECHNICAL SUPPORT RESOLVED THE ISSUE WITH THE CUSTOMER DURING THE INITIAL CALL, BY CORRECTING THE CONCENTRATION OF THE CALCIUM GLUCONATE INGREDIENT. A REVIEW OF THE CUSTOMER¿S FORMULA REVEALED THAT THE ABACUS SOFTWARE OPERATED AS INTENDED AND THE CAUSE OF THE ISSUE WAS RELATED TO THE INCORRECT INGREDIENT CONCENTRATION ENTERED BY THE CUSTOMER IN THEIR FORMULA. A REVIEW OF THE RISK DOCUMENTATION WAS PERFORMED, AND IT HAS BEEN CONCLUDED THAT THIS ISSUE IS NOT OCCURRING WITH GREATER FREQUENCY OR SEVERITY THAN ANTICIPATED. ION BALANCING: ION BALANCING IS A PROCESS THAT ACCEPTS DOSE SPECIFICATIONS FOR ELECTROLYTES IN TERMS OF INDIVIDUAL IONS. THE IONS CONSIDERED ARE: SODIUM (NA+), POTASSIUM (K+), CALCIUM (CA2+), MAGNESIUM (MG2+), PHOSPHATE (PO43-), ACETATE (AC-), AND CHLORIDE (CI-). ABACUS DEVELOPS THE APPROPRIATE COMBINATION OF SALT SOLUTIONS TO ACHIEVE THE NET CHARGE BALANCE USING THE CHOSEN ANION BALANCE METHOD. THIS ¿FLOAT¿ ION SERVES TO ACHIEVE THE FINAL BALANCE, BASED ON THE DEMAND FOR OTHER IONS IN THE SOLUTION. THE FLOAT ION WILL BE EITHER CHLORIDE, ACETATE OR A RATIO OF THE TWO. SALT BASED ORDERING: METHOD THAT WILL BE USED TO ORDER SALTS. IF USERS ELECT THE SALT-BASED ORDERING METHOD, THEY WILL ORDER THE AMOUNTS OF THE ELECTROLYTES BASED ON THE SALT SOLUTIONS DESCRIBED IN THE FORMULA. FORMULA: CONTAINS ALL INGREDIENTS THAT ARE AVAILABLE FOR COMPOUNDING TPN SOLUTIONS. THERE ARE MANY INGREDIENTS IN THE FORMULA BUT ONLY A SUBSET ARE NEEDED BY MOST ACCOUNTS. TO LIMIT ACCESS TO INGREDIENTS THEY CAN BE MARKED ¿UNAVAILABLE¿ WHICH LEAVES THE INGREDIENT HIDDEN FROM ORDER ENTRY USERS. ¿UNAVAILABLE¿ INGREDIENTS CAN BE EDITED, ADDED OR REMOVED WITHOUT IMPACT TO THE ROUTINE USERS. IF AN INGREDIENT IS MADE ¿AVAILABLE¿ THEN IT IS INCUMBENT UPON THE SYSTEM ADMINISTRATOR TO EXAMINE THE DATA ENTERED IN THE FORMULA INGREDIENT LIST TO ENSURE THAT IT IS CONSISTENT WITH THE FACILITY¿S PRACTICES. EACH INGREDIENT HAS A NUMBER OF ESSENTIAL COMPONENTS THAT ARE ADDRESSED IN DETAIL IN THE FORMULA EDITOR. USERS MUST CREATE AN INGREDIENT WITH A CONCENTRATION THAT ADDRESSES EACH ORDER METHOD THEY WILL USE. IF THE USER WISHES TO ORDER AN INGREDIENT BY MASS (GRAMS) AND BY MEQS, THEY MUST CREATE TWO INGREDIENTS; ONE DESCRIBED AS MEQ/ML AND ONE DESCRIBED AS MG/ML. WHEN ORDERED AS A SALT, CALCIUM GLUCONATE IS TYPICALLY FOUND IN FORMULARIES AS MILLIGRAMS OF THE COMBINED CALCIUM AND GLUCONATE SALT BUT WHEN USED AS THE SOURCE FOR CALCIUM IN ION BASED ORDERS, ABACUS REQUIRES THAT THE CALCIUM GLUCONATE INGREDIENT CONTAIN AN ACCURATE ELEMENTAL VALUE OF THE CALCIUM CONTAINED IN THE COMBINED SALT. THIS ENABLES AN ION BASED ORDER TO BE PLACED FOR THE CALCIUM ION RATHER THAN THE SALT CALCIUM GLUCONATE. ONE HUNDRED MG OF CALCIUM GLUCONATE CONTAINS 9.3 MG OF ELEMENTAL CALCIUM. THE ABACUS SOFTWARE PERFORMS CALCULATIONS BASED ON ION PROPERTIES, AND CALCULATES THE MOLAR CONCENTRATION TO THE EQUIVALENT GRAM WEIGHT OF THE ION THAT IS DELIVERED BY A SALT. USING CALCIUM GLUCONATE 10% AS AN EXAMPLE, IF THE USER REQUESTS THAT 100 MG OF THE PRODUCT BE DELIVERED, ABACUS WILL MULTIPLY THE 0.465 MEQ/ML TIMES THE MOLECULAR WEIGHT OF CALCIUM DIVIDED BY ITS VALENCE TO ARRIVE AT THE MILLIGRAMS OF ELEMENTAL CALCIUM THAT WILL BE DELIVERED PER ML; 0.465 MEQ/ML X 40.08 MG/MMOL/ (VALENCE) = 9.3 MG OF CA2+/ML. USERS CANNOT ORDER THE AMOUNT OF A SALT IN AN ION-BASED FORMULA, BECAUSE THE SOFTWARE WILL CONVERT THE MILLIEQUIVALENT CONCENTRATION TO THE AMOUNT OF THE ELEMENTAL ION. THEREFORE, IF 100 MG/KG OF CALCIUM IS ORDERED IN THE ABOVE EXAMPLE, THE SOFTWARE WILL CALCULATE THE DOSE TO BE 10.75 ML OF CALCIUM GLUCONATE SOLUTION PER KILOGRAM OF BODY WEIGHT. | 1
2823484 | THE CUSTOMER REPORTED THAT THE TELEMON WAS NOT PROVIDING AUDIBLE ALARMS. A PHILIPS FIELD SUPPORT ENGINEER WENT ON SITE AND REPLACED THE FAILED SPEAKER AND TESTED THE DEVICE POST REPAIR. NORMALLY, BOTH VISUAL AND AUDIBLE ALARMS ARE GENERATED AT BOTH THE CENTRAL STATION AND THE TELEMON BEDSIDE AND A SITUATION WHERE THERE WAS NO AUDIBLE ALARMS AT THE TELEMON WOULD NOT JEOPARDIZE PT SAFETY AS THE CENTRAL STATION ALARM WOULD SOUND. HOWEVER, BECAUSE TELEMON CAN BE USED AS AN ISOLATED MONITORING DEVICE, AND BECAUSE LABELING DOES NOT SPECIFY CLOSE PERSONAL OBSERVATION WHEN IN USE, RECURRENCE OF THIS FAILURE REMAINS IN USE AT THE CUSTOMER SITE. | 1
2823737 | PT UNDERGOING THERAPEUTIC PLASMA EXCHANGE (TPE) TREATMENT. WITHIN 40 MINS. OF THE PROCEDURE, PT COMPLAINED OF "FEELING FUNNY" WITH SOME TINGLING AROUND THE MOUTH. THE PROCEDURE WAS STOPPED; 1500 ML OF ALBUMIN AT A FLOW RATE OF 68ML/MIN WAS INFUSED DURING THE 40 MIN BEFORE THE PROCEDURE WAS STOPPED. THE PT WAS MONITORED POST PROCEDURE BY STAFF AND EVALUATED BY A PHYSICIAN. AFTER EVALUATION BY A PHYSICIAN, PT WAS GIVEN LASIX AND DISCHARGED HOME THAT SAME DAY. THE UNITED STATES AGENT FOR THE PRODUCT LISTED IS FRESENIUS KABI, USA LLC. THE TPE DISPOSABLE SETS (PART #9400451, LOT #BAT051) USED WITH THE NAME DEVICE HAVE BEEN RETURNED TO THE FRESENIUS MFG SITE FOR INVESTIGATION. ALL RESULTS FROM THE INVESTIGATION WILL BE SENT AS A FOLLOW-UP TO THIS MDR REPORT. | 0
2834118 | IT WAS REPORTED THAT THE GAMMA XXL MONITOR ON BED 1 IN THE MEDIUM CARE AREA DID NOT GENERATE A VF ALARM. THE PT HAD A PACEMAKER, BUT THE USERS DID NOT SET THE PACEMAKER DETECTION TO ON. SPIKES WERE COUNTED AS QRS COMPLEXES. THE PT STARTED TO GO INTO FIBRILLATION, BUT THE MONITOR REPORTEDLY DID NOT ALARM. PT REMAINED IN FIBRILLATION AND EVERY ATTEMPT TO RESUSCITATE THE PT FAILED. IT WAS REPORTED THAT THE PT EVENTUALLY WENT INTO ASYSTOLY AND DIED. DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
2835140 | AS MD WAS USING THE NEEDLE DRIVER, SHE NOTICED THE JAWS WOULD NOT OPEN CORRECTLY. IT WAS NOTED THAT THE WIRES AT THE TIP OF THE INSTRUMENT WERE BROKEN. | 0
2839734 | PER RADIOLOGY MANAGER, UNIT WAS TAKEN TO THE OR, WOULD NOT POWER UP. AFTER BIO MED STAFF INVESTIGATED, FOUND HARD DRIVE HAD FAILED. DISCOVERED CPU WAS AT FAULT. GE CAME TO HOSPITAL TO EVALUATE, REPLACED HARD DRIVE AND CPU. | 1
2840096 | DURING INTERNAL DEVELOPMENT AND TESTING, IT WAS FOUND THAT THE VALUE FOR THE SUVBSA MEASUREMENTS IN RADSUITE IS TOO HIGH BY A FACTOR OF 10. (B)(4). THE DEVICE WAS CLEARED FOR USE WITH SUV CALCULATION BY FDA IN 12/2005. THIS ISSUE HAS EXISTED SINCE THE INTRODUCTION OF THE CALCULATION. THIS HAS NEVER BEEN REPORTED FROM THE FIELD (NO RELATED COMPLAINTS). THIS ISSUE WAS DISCOVERED DURING INTERNAL TESTING IN A DEVELOPMENT CYCLE. THERE ARE FOUR PT SIZE FACTORS THAT MAY BE USED TO CALCULATE SUV AND IN EACH CASE, EXCEPT FOR BSA, THE CALCULATION IS CORRECT. THIS ISSUE IS PLANNED TO BE CORRECTED IN TWO SEPARATE RELEASES CURRENTLY IN DEVELOPMENT (VERSIONS 5.36 AND 8.30.6). A PRODUCT BULLETIN WILL ALSO BE DISTRIBUTED TO ALL CURRENT CUSTOMERS TO MAKE THEM AWARE OF THE ISSUE. | 1
2840606 | THE CUSTOMER REPORTED THAT THE DONOR HAD A MILD VASOVAGAL REACTION DURING A COLLECTION PROCEDURE. MEDICATION (ETILEFRIN) WAS GIVEN TO THE DONOR. PATIENT AGE IS NOT AVAILABLE AT THIS TIME. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN FOR EVALUATION. THIS REPORT IS BEING FILED DUE TO MEDICAL INTERVENTION VIA ETILEFRIN. (B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE REPORTED DONOR REACTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. THE TOTAL VOLUMES COLLECTED WERE WITHIN THE SAFETY LIMIT OF 15% OF THE DONOR'S TOTAL BLOOD VOLUME. BASED ON THE RDF ANALYSIS, THE TRIMA ACCEL SYSTEM FUNCTIONED AS INTENDED. VASOVAGAL REACTION IS A KNOWN POSSIBLE SIDE EFFECT FOR APHERESIS PROCEDURES. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4).><(B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE DONOR REACTION REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
2840803 | AS REPORTED IN THE LITERATURE, THE FLOTRAC/VIGILEO FAILED TO TRACK RAPID HEMODYNAMIC CHANGES IN AN UNSTABLE CARDIAC SURGICAL PATIENT; "THE PATIENT WAS SCHEDULED TO UNDERGO MECHANICAL PROSTHETIC AORTIC VALVE REPLACEMENT FOR SEVERE AORTIC STENOSIS. HIS PREOPERATIVE TRANSTHORACIC ECHOCARDIOGRAPHY (TTE) WAS PERTINENT FOR AN AORTIC VALVE AREA=0.8 CM2, MEAN AND PEAK PRESSURE GRADIENTS ACROSS THE AORTIC VALVE OF 60 MM HG AND 106 MM HG, RESPECTIVELY, AND PRESERVED LEFT VENTRICULAR EJECTION FRACTION (>60%) " INDUCTION OF GENERAL ANESTHESIA AND INTUBATION WERE UNEVENTFUL; THE PATIENT WAS MONITORED USING STANDARD MONITORS IN COMBINATION WITH A LEFT RADIAL ARTERIAL LINE CATHETER, A CENTRAL VENOUS CATHETER, AND TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). TEE CONFIRMED A RELATIVELY SMALL AORTIC ANNULUS (1.9 CM) TO BODY SURFACE AREA (BSA) (2.35 M2), NORMAL RIGHT VENTRICULAR FUNCTION AND SIZE, MILD TRICUSPID REGURGITATION, AND MILD DIASTOLIC DYSFUNCTION. WHILE TECHNICALLY DIFFICULT, A MECHANICAL AORTIC PROSTHESIS (ST. JUDE SIZE #21) WAS SUCCESSFULLY IMPLANTED IN THE AORTIC POSITION. INTRAOPERATIVE TEE DEMONSTRATED A WELL-SEATED VALVE AND NORMAL LEFT VENTRICULAR FUNCTION AFTER UNEVENTFUL WEANING FROM CARDIOPULMONARY BYPASS. THE MEAN PRESSURE GRADIENT ACROSS THE AORTIC PROSTHESIS WAS 20 MMHG; FOLLOWING INTRAOPERATIVE MANAGEMENT, THE PATIENT WAS TRANSFERRED TO THE CARDIAC SURGICAL ICU ON A DEXMEDETOMIDINE INFUSION IN PREPARATION FOR FAST-TRACK EXTUBATION. THE PATIENT WAS SUPPORTED ON A LOW DOSE NOREPINEPHRINE INFUSION (0.03 MCG/KG/MIN). UPON ARRIVAL TO THE ICU, THE FLOTRAC / VIGILEO (3.01) MONITOR WAS APPLIED (NOTE: THE DEVICE WAS NOT USED IN THE OPERATING ROOM). NINETY MINUTES AFTER ARRIVAL TO THE ICU, A SPONTANEOUS BREATHING TRIAL (CONTINUOUS POSITIVE AIRWAY PRESSURE CPAP OF 5 CM H2O) WAS INITIATED IN PREPARATION FOR EXTUBATION. READINGS ON THE FLOTRAC/VIGILEO DENOTED NORMAL CARDIAC FUNCTION THROUGHOUT THE SPONTANEOUS BREATHING TRIAL. DESPITE THE NORMAL READINGS ON THE FLOTRAC/VIGILEO, ARTERIAL BLOOD PRESSURE RECORDINGS SHOWED A GRADUAL DECLINE IN THE SYSTOLIC, DIASTOLIC AND MEAN ARTERIAL PRESSURES. APPROXIMATELY ONE HOUR LATER, AND WHILE STILL INTUBATED AND BREATHING SPONTANEOUSLY, THE PATIENT WAS NOTED TO BE UNRESPONSIVE AND DEVELOPED PULSELESS VENTRICULAR TACHYCARDIA. CARDIOPULMONARY RESUSCITATION WAS EMPLOYED INTERMITTENTLY FOR APPROXIMATELY 1 HOUR (A TOTAL OF INTRAVENOUS BOLUSES OF EPINEPHRINE 1 MG, AMIODARONE 300 MG WERE ADMINISTERED COMBINED WITH CHEST COMPRESSIONS). FOLLOWING THIS PERIOD, AN EKG REVEALED ACUTE RIGHT VENTRICULAR INFARCTION. CONCOMITANT TEE REVEALED RIGHT VENTRICULAR DILATATION, GLOBAL RIGHT VENTRICULAR WALL AKINESIS, MILD TRICUSPID REGURGITATION AND LEFTWARD BOWING OF THE INTERVENTRICULAR SEPTUM. PHYSIOLOGIC DATA POINTS FOR CARDIAC OUTPUT PER THE FLOTRAC/VIGILEO WERE REPORTED AS NORMAL LIMITS THROUGHOUT THE EPISODE OF CARDIAC ARREST. THE PATIENT WAS RUSHED INTO THE CARDIAC ANGIOGRAPHY SUITE AND FOUND TO HAVE TOTAL RIGHT CORONARY ARTERY OCCLUSION NOT AMENABLE TO STENTING. THE PATIENT WAS SUBSEQUENTLY TAKEN BACK TO THE OPERATING ROOM FOR EMERGENT RIGHT CORONARY BYPASS GRAFTING. ENSUING SEVERE GLOBAL RIGHT VENTRICULAR AKINESIS AND RECALCITRANT HEMODYNAMIC INSTABILITY REQUIRED USE OF THE TANDEMHEART RIGHT VENTRICULAR ASSIST DEVICE (RVAD). THE PATIENT WAS TRANSFERRED TO THE CARDIAC ICU IN CRITICAL-SEVERE CONDITION. HIS POSTOPERATIVE COURSE WAS COMPLICATED BY PERSISTENT CARDIOGENIC SHOCK UNRESPONSIVE TO HIGH DOSES OF IONOTROPIC AGENTS. HE WAS NOT AMENABLE TO WEANING FROM RVAD, AND RAPIDLY DEVELOPED MULTI-ORGAN FAILURE REQUIRING THE INSTITUTION OF CONTINUOUS VENOVENOUS HEMODIAFILTRATION (CVVHDF). FOUR DAYS LATER HIS CARE WAS DEEMED FUTILE AND THE FAMILY ELECTED TO WITHDRAW LIFE SUPPORT. NO PRODUCTS WERE RETURNED FOR EVALUATION AFTER REVIEW OF THE INFORMATION, THE CAUSE OF THE COMPLAINT CANNOT BE DETERMINE WITH ANY CERTAINTY WITHOUT THE OPPORTUNITY TO EXAMINE THE DEVICES, IT IS NOT POSSIBLE TO DETERMINE IF ANY DAMAGES OR DEFECTS WERE PRESENT. NO ACTIONS WILL BE TAKEN AT THIS TIME. | 1
2841249 | THE DOCTOR WOULD ONLY GET A PARTIAL CUT, THE MANUFACTURER REP SUGGESTED IT WAS THE DEVICE AND A NEW DEVICE WAS GIVEN AND IT PERFORMED ACCURATELY. THE NEW DEVICE WAS FROM THE SAME LOT NUMBER.UNFORTUNATELY THE DOCTOR HAD LIMITED RECOLLECTION OF THE EVENT WITH THE TIME THAT HAS PASSED SINCE THE INCIDENT. THE DAVINCI REPRESENTATIVE WAS PRESENT AND REPLACED THE DEVICE WITH THE NEW VESSEL SEALER WHICH PERFORMED WITHOUT DIFFICULTY. IT IS DIFFICULT TO DETERMINE FROM THE MD NOTES EXACTLY WHEN THE DEVICE FAILED DURING THE PROCEDURE. THERE WAS NO PATIENT HARM.THERE WERE 3 VESSEL SEALER RELATED ISSUES AT OUR FACILITY. THE DAVINCI DEVICE IS NEW TO OUR FACILITY AND PHYSICIAN USE WOULD HAVE BEEN LOW AT THE TIME OF THE EVENT. | 0
2842748 | THE ROBOTIC FORCEP DEVICE REFUSED TO OPEN. THE FAILURE OCCURRED AFTER THE LAST USE OF THE DEVICE AND NO OTHER DEVICE WAS NEEDED FOR THIS CASE. THE FORCEP WAS REPLACED. THERE WAS NO PATIENT HARM. | 0
2842805 | CUTTING BLADE WOULD NOT RETRACT. | 0
2842811 | NOTED A HOLE IN THE INTERNAL PACKAGE. | 0
2842816 | THE MEDICAL DEVICE WOULD NOT OPEN. THE DEVICE WAS CHECKED AFTER THE CASE, THERE WERE 2 USES STILL LEFT ON DEVICE. | 0
2843522 | THE PATIENT WAS UNDERGOING A ROBOTIC TOTAL LAPAROSCOPIC HYSTERECTOMY WITH BILATERAL SALPINGECTOMY. THE WIRES ON THE INSTRUMENT BROKE. THE FRAGMENTS WERE RETRIEVED FROM THE PATIENT'S ABDOMEN. | 0
2847065 | ON (B)(6) 2012 A PT WAS DISCOVERED BY HIS FAMILY. HE HAD APPARENTLY BEEN DECEASED FOR OVER 1 HOUR. AT THE TIME OF THIS EVENT, THE PT WAS BEING MONITORED BY TELEMETRY (GE CIC PRO CENTRAL STATION, SW VERSION 5.1) WHICH HAD GONE INTO "ARRHYTHMIA SUSPEND" MODE. WHILE IT WAS STAFF RESPONSIBILITY TO CONTINUE TO MONITOR THIS PT, IT WAS NOTED THERE IS NO ESCALATION OF THE ALARM MECHANISM, EVEN IF THIS MODE PERSISTS FOR AN EXTENDED PERIOD OF TIME. THE CONCERN IDENTIFIED BY THE TEAM INVESTIGATING THIS EVENT, IS THAT ALTHOUGH THE EQUIPMENT WAS FUNCTIONING AS DESIGNED, ENCHAINED DESIGN RESULTING IN A RE-ALERT MECHANISM MAY BE A DESIGN ENHANCEMENT WORTH PROMOTING. THIS IS BEING SENT FOR CONSIDERATION OF OTHERS WHO MAY HAVE ENCOUNTERED SIMILAR EXPERIENCES. A LETTER REQUESTING SOFTWARE ENHANCEMENT WAS ALSO BEEN SENT TO GE BY THE CLINICAL ENGINEERING TECHNICAL SITE LEADER. | 1
2849232 | CUSTOMER REPORT OF SMOKE FROM THE UPS POWER SUPPLY ON THE PAS DEVICE. CUSTOMER DENIES OBSERVING ANY SPARK OR FLAME. PATIENT WAS PRESENT. THERE WAS NO HARM TO THE PATIENT OR ANY USER. (B)(4). ADDITIONAL DATA / FAILURE INVESTIGATION. CAUSE OF FAILURE IS DETERMINED TO BE CUSTOMER CAUSED. CUSTOMER SPILLED IV FLUIDS ONTO THE UPS POWER SUPPLY, WHICH INGRESSED INTO THE ELECTRICAL COMPONENTS. | 1
2852318 | DA VINCI THORACIC GRASPER BLACK PROTECTIVE COATING RIPPED BY METAL DA VINCI TROCAR. INSTRUMENT INSPECTED BY TWO DOCTORS AND STATED INSTRUMENT COATING APPEARED TO BE COMPLETE; WOUND INSPECTED NO COATING PIECES NOTED. | 0
2853194 | WE RECEIVED NOTIFICATION FROM OUR CUSTOMER THAT OUR SECONDLOOK, COMPUTER-AIDED DETECTION (CAD) SYSTEM FOR MAMMOGRAPHY HAD ALLEGEDLY MALFUNCTIONED. THE CUSTOMER REPORTED THAT THE ICAD SYSTEM IMAGES WERE BEING SENT TO A HOLOGIC WORKSTATION THAT APPEARED SLIGHTLY SKEWED IN THE STRUCTURED WORKSTATION REPORT. THE CUSTOMER INDICATED THAT THIS EVENT OCCURRED AFTER HOLOGIC UPGRADED THEIR WORKSTATION SOFTWARE. (B)(4). ICAD RECEIVED INFORMATION FROM THE CUSTOMER INDICATING THAT HOLOGIC HAD IDENTIFIED THAT THE PROBLEM WAS BEING CAUSED BY THEIR DEVICE. ICAD CONTACTED HOLOGIC FOR MORE INFORMATION ON THIS ISSUE. OUR INVESTIGATION HAS IDENTIFIED THE ROOT CAUSE OF THE REPORTED CONDITION AS A SOFTWARE ERROR USED IN THE HOLOGIC WORKSTATION THAT RESULTED IN THE SKEWED CAD MARKS THAT OUR CUSTOMER REPORTED TO ICAD. ON (B)(4), 2008, ICAD BECAME AWARE OF A CLASS 3 RECALL ON THE HOLOGIC FULL FIELD DIGITAL MAMMOGRAPHY SYSTEM. THIS SYSTEM WAS IDENTIFIED BY OUR CUSTOMER AS BEING USED WITH THE ICAD SECONDLOOK DIGITAL SERVER IN THE (B)(4) 2008 REPORTED EVENT. ICAD CONTACTED HOLOGIC TO DETERMINE IF THE CORRECTIVE ACTIONS IDENTIFIED IN THE HOLOGIC RECALL NOTIFICATION HAD TAKEN PLACE. THE HOLOGIC REPRESENTATIVE INDICATED THAT THE WORKSTATION SERIAL NUMBERS LISTED ON THE RECALL WERE BEING CORRECTED. ICAD MADE A FOLLOW-UP PHONE CALL TO OUR CUSTOMER TO CHECK THE STATUS OF THE HOLOGIC RECALL ACTIONS. OUR CUSTOMER REPORTED THAT AFTER HOLOGIC CORRECTED THE REPORTED SOFTWARE ISSUE, THE PROBLEM HAD NOT REOCCURRED. NO ADDITIONAL ACTIONS ARE PLANNED BY ICAD HOWEVER, WE WILL CONTINUE TO MONITOR FOR THE REPORTED CONDITION. | 1
2854238 | SINGLE SITE MEDIUM-LARGE CLIP APPLIER-BOTH IN SET, HELD CLIPS TOO LOOSELY, CLIPS WERE FALLING OFF, EVEN THOUGH IT HASN'T EXPIRED YET. | 0
2854922 | (B)(6). DATE RECEIVED: (B)(4) 2012. ACTUAL ERROR: YES. DESCRIPTION: SITUATION: A FENTANYL INFUSION WAS ORDERED USING EPIC SOFTWARE FOR A (B)(6) PT AT A RATE OF 2 MCG/KG/HR, FOR A TOTAL OF 30 MCG/HR. THE ALARIS PUMP USED TO ADMINISTER THE INFUSION DID NOT HAVE AN OPTION TO SELECT 30 MCG/HR. THE PUMP WAS THEN INADVERTENTLY PROGRAMMED TO ADMINISTER 30 MCG/KG/HR. BACKGROUND: FENTANYL IS A STRONG AGONIST AT MU-OPIOID RECEPTORS, WITH A RELATIVELY RAPID ONSET AND SHORT DURATION OF ACTION. INTRAVENOUS FENTANYL IS COMMONLY USED FOR ANALGESIA AND FOR ANESTHESIA, OFTEN IN COMBINATION WITH A BENZODIAZEPINE. AS A POTENT OPIOID AGONIST, ITS SIDE EFFECTS CAN BE POTENTIALLY FATAL, ESPECIALLY IF ADMINISTERED AT HIGH, INAPPROPRIATE DOSES. IN PARTICULAR, THESE SIDE EFFECTS INCLUDE SOMNOLENCE, RESPIRATORY DEPRESSION, SEIZURES, AND COMA. CONTINUOUS SEDATION OR ANALGESIA WITH FENTANYL FOR PEDIATRIC PATIENT IS RECOMMENDED WITH AN INITIAL IV BOLUS OF 1-2 MCG/KG AND THEN 1 MCG/KG/HR. TITRATION TO DESIRED EFFECT IS APPROPRIATE WITH USUAL DOSING AT 1-3 MCG/KG/HR, THOUGH SOME PTS REQUIRE 5 MCG/KG/HR. ASSESSMENT: A MEDICATION ERROR OCCURRED IN OUR HOSPITAL AND WAS RECORDED AS A NEAR MISS MEDICATION EVENT IN OUR VOLUNTARY ONLINE REPORTING SYSTEM, QUANTROS. THIS ERROR COULD HAVE RESULTED IN A PEDIATRIC PT RECEIVING AN OVERDOSE OF MEDICATION (FENTANYL INFUSION), POTENTIALLY RESULTING IN SEVERE HARM HAD THE ERROR NOT BEEN DISCOVERED PRIOR TO ADMINISTRATION. THE MEDICATION SAFETY OFFICER BECAME AWARE OF THE SITUATION AFTER REVIEWING QUANTROS REPORTS AND FURTHER EVALUATED THE OPTIONS FOR ORDERING PEDIATRIC INFUSION IN THE EPIC SYSTEM. IT WAS FOUND THAT THE UNITS AVAILABLE TO SELECT WHEN ORDERING A MEDICATION DID NOT MATCH THE UNITS AVAILABLE TO SELECT FOR ADMINISTRATION OF THE MEDICATION VIA ALARIS PUMPS FOR ALL PEDIATRIC INFUSIONS. AT THE TIME OF THE ERROR, PROVIDERS COULD NOT APPROPRIATELY ORDER FENTANYL INFUSIONS (AND CERTAIN OTHER MEDICATIONS) FOR PEDIATRICS. THESE INFUSIONS COULD ONLY BE ORDERED IN MCG/HR OR ML/HR, HOWEVER, THEY SHOULD BE ORDERED IN MCG/KG/HR (PER THE JOINT COMMISSION AND TO COMPLY WITH GUARDRAILS ON ALARIS PUMPS). ALARIS PUMPS USED TO ADMINISTER THESE INFUSION TO PTS ARE PROGRAMMED WITH THE UNITS, MCG/KG/HR. IN THIS SITUATION, THE UNITS FOR FENTANYL DID NOT CORRESPOND BETWEEN THE ALARIS PUMP AND THE EPIC SYSTEM, SO THE PUMP WOULD NOT ALLOW THE DOSE TO BE SELECTED AS ORDERED IN EPIC. SINCE THE DOSE OF 30 MCG/HR WAS NOT A PROGRAMMABLE OPTION, THE PUMP WAS INADVERTENTLY PROGRAMMED FOR 30 MCG/KG/HR. THE PT RECEIVED A BOLUS THAT WAS WITHIN NORMAL RANGE AND THE ERROR WAS DISCOVERED PRIOR TO ADMINISTERING THE INCORRECT INFUSION, THEREFORE, THE PT WAS NOT HARMED. RECOMMENDATION: UNITS SHOULD CORRESPOND BETWEEN THE EPIC SYSTEM AND ALARIS PUMPS TO PREVENT POTENTIAL MEDICATION ERRORS FROM OCCURRING. WITHOUT MATCHING UNITS, NURSES HAVE THE RESPONSIBILITY OF CONVERTING THE DOSES TO APPROPRIATE UNITS, WHICH CAN RESULT IN A DELAY OF MEDICATION ADMINISTRATION AND A MEDICATION ERROR DUE TO MISCALCULATION, OR INAPPROPRIATE DOSING UNITS MAY BE SELECTED, WHICH COULD CAUSE SERIOUS HARM TO PTS. THE UNITS DISPLAYED ON ALARIS PUMPS CANNOT BE ALTERED, SO THE EPIC SYSTEM REQUIRES MODIFICATION. THE OPTIONS AVAILABLE TO SELECT UNITS FOR MEDICATIONS SHOULD BE MODIFIED IN THE EPIC SYSTEM SO THAT THEY MATCH THE ALARIS PUMPS. THESE UNITS SHOULD BE ADDED AS AN OPTION AS A DROP-DOWN BOX FROM THE MAGNIFYING GLASS ICON. HIGHLIGHTING THE UNITS AVAILABLE ON THE ALARIS PUMPS IN EPIC MAY ALSO HELP MAKE PRESCRIBERS AWARE THAT THOSE ARE THE ONES THAT CORRESPOND TO THE PUMPS. NURSES AND PROVIDERS SHOULD BE EDUCATED ABOUT THIS NEAR MISS MEDICATION ERROR, AS WELL OF THE NEW FEATURES THAT WILL BE IMPLEMENTED INTO THE SYSTEM. IF ANY MEDICATION IS MISSED AND THIS SITUATION SHOULD BE IMPLEMENTED. THE PHARMACY SHOULD BE EQUIPPED WITH TOOLS AND CONVERSION CHARTS TO ASSIST WITH UNIT AND DOSE CONVERSIONS THAT MAY BE COMMONLY USED. WHENEVER THIS PROCEDURE OCCURS, A BRIEF EXPLANATION OF THE STEPS TAKEN TO ENSURE CORRECT DOSAGE ADMINISTRATION SHOULD BE DOCUMENTED. THESE INDIVIDUAL CASES SHOULD BE REPORTED TO THE MEDICATION SAFETY OFFICER WHO SHOULD F/U WITH THE MAKING ADD'L CHANGES IN THE EPIC SYSTEM AS NECESSARY. MEDICATION ADMINISTERED TO OR USED BY THE PT: NO. OUTCOME: THE PT WAS NOT HARMED. WHEN AND HOW WAS ERROR DISCOVERED: THE PT RECEIVED A BOLUS THAT WAS WITHIN RANGE AND THE ERROR WAS DISCOVERED PRIOR TO ADMINISTERING THE INCORRECT INFUSION, THEREFORE, THE PT WAS NOT HARMED. PT COUNSELING PROVIDED: UNK. REPORTER'S RECOMMENDATION: UNITS SHOULD CORRESPOND BETWEEN THE EPIC SYSTEM AND ALARIS PUMPS TO PREVENT POTENTIAL MEDICATION ERRORS FROM OCCURRING. WITHOUT MATCHING UNITS, NURSES HAVE THE RESPONSIBILITY OF CONVERTING THE DOSES TO APPROPRIATE UNITS, WHICH CAN RESULT IN A DELAY OF MEDICATION ADMINISTRATION AND A MEDICATION ERROR DUE TO MISCALCULATION, OR INAPPROPRIATE DOSING UNITS MAY BE SELECTED, WHICH COULD CAUSE SERIOUS HARM TO PTS. THE UNITS DISPLAYED ON ALARIS PUMPS CANNOT BE ALTERED, SO THE EPIC SYSTEM REQUIRES MODIFICATION. THE OPTIONS AVAILABLE TO SELECT UNITS FOR MEDICATIONS SHOULD BE MODIFIED IN THE EPIC SYSTEM SO THAT THEY MATCH THE ALARIS PUMPS. THESE UNITS SHOULD BE ADDED AS AN OPTION AS A DROP-DOWN BOX FROM THE MAGNIFYING GLASS ICON. HIGHLIGHTING THE UNITS AVAILABLE ON THE ALARIS PUMPS IN EPIC MAY ALSO HELP MAKE PRESCRIBES AWARE THAT THOSE ARE THE ONES THAT CORRESPOND TO THE PUMP. NURSES AND PROVIDERS SHOULD BE EDUCATED ABOUT THIS NEAR MISS MEDICATION ERROR, AS WELL OF THE NEW FEATURES THAT WILL BE IMPLEMENTED INTO THE SYSTEM. IF ANY MEDICATION IS MISSED AND THIS SITUATION COMES UP AGAIN, A PROCEDURE OF HAVING THE NURSE CHECK WITH A PHARMACIST TO VERIFY THE UNIT CONVERSION SHOULD BE IMPLEMENTED. THE PHARMACY SHOULD BE EQUIPPED WITH TOOLS AND CONVERSION CHARTS TO ASSIST WITH UNIT AND DOSE CONVERSIONS THAT MAY BE COMMONLY USED. WHENEVER THIS PROCEDURE OCCURS, A BRIEF EXPLANATION OF THE STEPS TAKEN TO ENSURE CORRECT DOSAGE ADMINISTRATION SHOULD BE DOCUMENTED. THESE INDIVIDUAL CASES SHOULD BE REPORTED TO THE MEDICATION SAFETY OFFICER WHO SHOULD F/U WITH MAKING ADD'L CHANGES IN THE EPIC SYSTEM AS NECESSARY. RELEVANT MATERIALS PROVIDED: NONE. | 1
2869565 | CASE WAS PERFORMED ON (B)(6) 2012. IT WAS NOTICED THAT DURING THE SECOND THAW THERE SEEMED TO BE A LOT OF ICE AROUND THE URETHRA. THE URETHRAL WARMER WAS RUNNING AT 42 DEGREES, AND MOBILE. THE DOCTOR DECIDED TO LEAVE A CATHETER IN LONGER AND FOLLOW UP AT A LATER DATE POST-OP. WHILE WORKING WITH THE PHYSICIAN TODAY (B)(6) 2012, THE DOCTOR INFORMED ME THAT THE PATIENT HAS SUFFERED SIGNIFICANT URETHRAL DAMAGE, IS SUFFERING FROM HEMATURIA AND SLOUGHING, AND IS STILL CATHETERIZED. I PROVIDED THE NUMBER OF SOME OF ENDOCARE'S PROCTORS TO CONSULT WITH AND THE DOCTOR IS CURRENTLY CONTENT TO SEEK OUTSIDE ADVISE AND PROCEED ACCORDINGLY. UPON INFORMING ME OF THE SITUATION, HE AND I AGREED THAT IT SHOULD BE DOCUMENTED AND REPORTED ACCORDINGLY. NEITHER OF US CAN RECALL ANY SPECIFIC REASON FOR INJURY TO HAVE OCCURRED AND FREEZE TIMES SEEM WELL WITHIN NORMAL SPECIFICATIONS. THE TEMP PROBES APPEAR TO HAVE BEEN WITHIN THE GLAND, BUT THAT WOULD NOT EXPLAIN THE URETHRAL DAMAGE. AS OF THIS MOMENT, THE PHYSICIANS AND I HAVE NO EXPLANATION FOR THE EVENT. SYSTEM NOT RETURNED. HOWEVER, A SERVICE TECHNICIAN WAS SENT OUT TO INSPECT THE UNIT FOLLOWING THE REPORT OF THE INJURY. UNIT PERFORMING AS INTENDED. CONCOMITANT PROCEDURE KIT WAS DISCARDED BY THE HOSPITAL. A DEVICE HISTORY RECORD REVIEW FOR THIS LOT OF PRODUCT WAS CONDUCTED WITH NO ISSUES FOUND. HEALTHTRONICS CLINICAL EDUCATION SPECIALIST SPOKE WITH THE DOCTOR, WHO STATED HE HAS CONSULTED WITH ANOTHER DOCTOR ON THE ISSUE. THE CLINICAL EDUCATION SPECIALIST PROVIDED THE DOCTOR WITH THE CONTACT INFORMATION FOR ANOTHER DOCTOR SHOULD HE WANT A SECOND OPINION. | 0
2876346 | ON (B)(6) 2012 DURING REVIEW OF THE VNS PATIENT'S PROGRAMMING HISTORY, IT WAS DISCOVERED THAT ON (B)(6) 2003 A GENERATOR DIAGNOSTIC TEST WAS PERFORMED, WHICH RESULTED IN A FAULT READING AND CHANGED THE PATIENT'S SETTINGS. IT WAS NOT UNTIL (B)(6) 2004, THAT THESE INCORRECT SETTINGS WERE OBSERVED AND THE PATIENT WAS PROGRAMMED OFF. ANALYSIS OF PROGRAMMING HISTORY. | 0
2876966 | WHILE SURGEON WAS SUTURING MESH WITH THE DA VINCI INSTRUMENT, THE WIRE WAS NOTED TO BE FRAYED. | 0
2877986 | A CUSTOMER REPORTED EXPERIENCING THE LOST BOTTLE ISSUE RELATED TO FSCA 1494. WHILE REVIEWING DATA IN OBSERVA, THE CUSTOMER NOTICED A LAB NUMBER MISSING IN A SEQUENCE OF NEGATIVE BOTTLES. THE CUSTOMER BELIEVES THE BOTTLE WAS IMPROPERLY LOADED WITHOUT TWO BEEPS, WHICH MAY RESULT IN ONE BOTTLE BECOMING LOST AND ONE BOTTLE BECOMING MISIDENTIFIED IN THE SYSTEM. THE MISIDENTIFIED BOTTLE CAUSED THE PEDIATRIC PATIENT'S PRELIMINARY NEGATIVE RESULT TO NOT BE SENT OUT AS A 36 HOUR PRELIMINARY REPORT. PER (B)(6) GUIDELINE "CG149 - ANTIBIOTICS FOR EARLY ONSET NEONATAL INFECTION" PEDIATRIC PATIENTS WHOSE BLOOD CULTURE IS NEGATIVE AND NO LAB EVIDENCE OF INFECTION AT 36 HOURS CAN HAVE ANTIBIOTIC THERAPY DISCONTINUED. THE CUSTOMER HAD TO WAIT FOR THE CSF RESULT WHICH TOOK AN ADDITIONAL 48 OURS DURING WHICH THE PATIENT WAS LEFT ON ANTIBIOTIC THERAPY. BASED ON THE INFORMATION PROVIDED BY THE CUSTOMER, THERE IS NO KNOWN CONSEQUENCE OF THE LOST BOTTLE. THE CUSTOMER REPORTED THIS ISSUE TO BIOMERIEUX ON (B)(6) 2012. BASED ON THE BACKUP PROVIDED BY THE CUSTOMER, THE BOTTLE HISTORY INDICATES THE MISIDENTIFIED BOTTLE WAS LOADED ON (B)(6) 2012 AND FLAGGED A FINAL NEGATIVE STATUS ON (B)(6) 2012. BASED ON THE 36 HOUR GUIDELINE, THE BOTTLE WOULD HAVE INDICATED THE PATIENT ANTIBIOTIC THERAPY COULD HAVE BEEN STOPPED ON (B)(6) 2012, THUS THIS IS CONSIDERED THE DATE OF THE EVENT. THE CUSTOMER'S SYSTEM WAS MANUFACTURED (B)(6) 2000, BASED ON THE SERIAL NUMBER AND MANUFACTURING RECORD. THE SYSTEM WAS UPGRADED TO B.40 FIRMWARE ON (B)(6) 2012, WHICH WAS USING A LBC-586 CPU. THE CPU WAS UPGRADED TO A LX800 CPU ON (B)(6) 2012. FSCA 1494 ADDRESSED AN IMPROPER LOADING ISSUE WHERE A USER PERFORMS ANOTHER ACTION (SUCH AS LOADING A BOTTLE OR SCANNING ANOTHER BARCODE) BEFORE THE BARCODE SCANNER RESPONSE BEEPS HAVE BEEN ISSUED TO INDICATE COMPLETION OF THE PREVIOUS BOTTLE LOADING PROCESS (SCAN AND LOAD). THIS IS A KNOWN SOFTWARE MALFUNCTION AND THE FSCA WAS SENT TO THE FIELD ON (B)(6) 2012, WITH INSTRUCTIONS FOR ALL SUBSIDIARIES AND DISTRIBUTORS TO NOTIFY AFFECTED CUSTOMERS OF THE ISSUE BY SENDING THEM THE CUSTOMER LETTER THAT CONTAINED UPDATED INSTRUCTIONS TO PREVENT THE ISSUE FROM OCCURRING. AN INVESTIGATION HAS BEEN OPENED FOR THIS INCIDENT. THE ORIGINAL BOTTLE HAS BEEN DISPOSED AND IS NOT AVAILABLE FOR INVESTIGATION. THE INSTRUMENT AND SOFTWARE CONFIGURATION IS KNOWN AND THE INVESTIGATION BASED ON THE NATURE OF THE COMPLAINT THE INVESTIGATION WILL FOCUS ON THE HARDWARE AND SOFTWARE COMPONENTS OF THE BACT/ALERT SYSTEM. THE EVENT IS REPORTABLE TO THE FDA AS AN ADVERSE EVENT BECAUSE THE PATIENT WAS KEPT ON ANTIBIOTIC THERAPY THAT RESULTED IN PROLONGED HOSPITALIZATION DUE TO THE LOST BOTTLE MALFUNCTION. A VOLUNTARY RECALL ON BACT/ALERT 3D B.40 FIRMWARE WAS IMPLEMENTED BY BIOMERIEUX ON (B)(4) 2012, THAT INCLUDED A CUSTOMER LETTER (URGENT PRODUCT CORRECTION NOTICE) AND FAX BACK ACKNOWLEDGEMENT FORM. THE URGENT PRODUCT CORRECTION NOTICE CONTAINED UPDATED INSTRUCTIONS FOR USE THAT EXPLAINED HOW TO FOLLOW THE SYSTEM RESPONSE BEEPS AND INCLUDED A CAUTION STATEMENT WARNING CUSTOMERS OF THE IMPACT OF IMPROPER USE. THE RECALL IS ONGOING AT THIS TIME. INVESTIGATION IS ONGOING.><A VOLUNTARY RECALL ON BACT/ALERT 3D B.40 FIRMWARE WAS IMPLEMENTED BY BIOMERIEUX ON OCTOBER 16, 2012 THAT INCLUDED A CUSTOMER LETTER (URGENT PRODUCT CORRECTION NOTICE) AND FAX BACK ACKNOWLEDGEMENT FORM. THE URGENT PRODUCT CORRECTION NOTICE CONTAINED UPDATED INSTRUCTIONS FOR USE THAT EXPLAINED HOW TO FOLLOW THE SYSTEM RESPONSE BEEPS AND INCLUDED A CAUTION STATEMENT WARNING CUSTOMERS OF THE IMPACT OF IMPROPER USE. THE RECALL IS ONGOING AT THIS TIME. COMPLAINT INVESTIGATION: THE CUSTOMER'S BACKUP WAS REVIEWED AND CONFIRMED THE PATTERN OF LOADED BOTTLES DOES INDICATE THE LOST BOTTLE ISSUE OCCURED WITH THEIR INSTRUMENT. THIS IS A KNOWN ISSUE AND IS ASSOCIATED WITH FIELD SAFETY CORRECTIVE ACTION Z-0299/0329-2013. THE LOG FILE INDICATES ON (B)(6) 2012 A BOTTLE WAS MOVED OUT OF CELL 4C09 TO 4C10 DURING A STRING OF 10 BOTTLES BEING LOADED INTO THE INSTRUMENT. THIS BOTTLE WAS THE LAST BOTTLE LOADED IN THE 10 BOTTLE SEQUENCE AND THE BOTTLE WAS "FOUND" ON (B)(6) 2012 BY THE CUSTOMER WHEN THE SYSTEM WAS POWER CYCLED. A REVIEW OF THE CUSTOMER LOG INDICATES THE 10 BOTTLES WERE LOADED IN 35 SECONDS. AS PART OF THE CAPA INITIATED TO FURTHER INVESTIGATE THIS LOST BOTTLE ISSUE FOR FSCA Z-0299/0329-2013, A CUSTOMER STUDY WAS CONDUCTED THAT DETERMINED THE AVERAGE CUSTOMER BOTTLE LOAD TIME WAS BETWEEN 3.58 AND 5.40 SECONDS. THE COMPLAINT INVESTIGATION CONCLUDED THE SYSTEM WAS NOT ABLE TO KEEP UP WITH THIS CUSTOMER'S LOAD SPEED AND RECOMMENDED THE CUSTOMER FOLLOW THE UPDATED INSTRUCTIONS FOR USE PROVIDED IN THE URGENT PRODUCT CORRECTION NOTICE TO WAIT FOR THE RESPONSE BEEPS ISSUED BY THE INSTRUMENT DURING BOTTLE LOADING PROCESS TO ENSURE THE BOTTLES ARE PROPERLY LOADED. THESE INSTRUCTIONS WERE PROVIDED TO THE KC DISTRICT RECALL COORDINATOR AS THE SHORT-TERM CORRECTION FOR THIS ISSUE. THE LONG-TERM SOLUTION IS BEING ADDRESSED WITH A SOFTWARE UPDATE AND THE TIMELINE ASSOCIATED HAS ALSO BEEN PROVIDED TO THE KC DISTRICT RECALL COORDINATOR. | 0
2887003 | THE HOSPITAL REPORTED THAT, AT THE START OF A CASE, PRIOR TO DELIVERING ANESTHESIA, THE UNIT WAS ALARMING HIGH POSITIVE END-EXPIRATORY PRESSURE (PEEP). THE CUSTOMER REPORTED THEY WERE GETTING TIDAL VOLUME READINGS AND THE PATIENT'S CHEST WAS NOTED TO BE RISING. THE CLINICIAN HAD DIFFICULTY VENTILATING IN MECHANICAL AND MANUAL MODES. THE PATIENT DESATURED (40% RANGE), WAS SWITCHED TO AN AMBU BAG, AND WAS RESUSCITATED. A SECOND CRNA CAME IN THE OPERATING ROOM. THE PATIENT WAS REPORTEDLY PLACED BACK ON THE MACHINE, AND DIFFICULTY IN VENTILATION WAS NOTED. THE PATIENT WAS SWITCHED TO AN AMBU BAG AGAIN. AN ANESTHESIOLOGIST CAME IN THE ROOM, THE PATIENT WAS PLACED BACK ON THE MACHINE, AND THE SAME REPORTED DIFFICULTY IN VENTILATION WAS NOTED. THE MACHINE WAS REPLACED AND THE CASE WAS FINISHED WITH NO REPORTED PATIENT INJURY. GE HEALTHCARE'S INVESTIGATION IN ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0

## 2013
3165734 | IMAGES THAT ARE FLIPPED WHILE STREAMING HAVE THE POTENTIAL TO INCORRECTLY DISPLAY IMAGE ORIENTATION MARKERS. WE HAVE ALREADY CREATED A CORRECTION TO PREVENT THE MALFUNCTION FROM OCCURRING AND WILL INCLUDE THIS CORRECTION IN ALL FUTURE UPGRADES. WE HAVE ALSO NOTIFIED OUR CUSTOMERS OF THE MALFUNCTION AND PROVIDED THEM WITH STEPS THEY CAN TAKE TO PREVENT THE MALFUNCTION FROM OCCURRING. SEE SCANNED PAGES. | 1
3422154 | GE ALERT INCIDICATING THE INFANT WARMERS WITH BLENDER KITS MAY HAVE THE AIR AND O2 WALL FITTINGS INVERTED. WE HAVE 1 OF THESE INFANT WARMERS IN OUR NICU, BUT DOES NOT HAVE THE BLENDER KIT. | 0
3499032 | THERE HAVE BEEN MULTIPLE DEFECTIVE LEAKING UNITS WITH THE MULTI ABSORBER MEDISORB. THESE MULTIPLE DEFECTIVE LEAKING UNITS HAVE CONTRIBUTED TO ANESTHESIA LEAKAGE. THE MANUFACTURER HAS PULLED THE ASSOCIATED LOT FROM THE HOSPITAL. | 0
3147051 | DURING TESTING, IT WAS FOUND THAT IF THE TOUCHGUARD OVERRIDE BUTTON ON THE HAND HELD CONTROLLER IS STUCK, THEN MOVEMENTS AND TREATMENTS ARE STILL POSSIBLE. THE INVESTIGATION INTO THIS INCIDENT IS ON-GOING AT THIS TIME. ONCE THE INVESTIGATION IS COMPLETE, ELEKTA WILL FORWARD ADDITIONAL INFORMATION TO THE FDA. | 0
3300008 | HOSPITAL NETWORK FAILED. MCKESSON ACUDOSE MEDICATION CABINET WAS IN USE IN CARDIAC SURGICAL SUITE. USER REBOOTED THE MACHINE BY MISTAKE. ACCESS TO FURTHER MEDICATIONS IN THE CABINET WAS NOT POSSIBLE AFTER THAT BECAUSE THE CABINET PERSISTED IN TRYING TO ACCESS THE NETWORK DOMAIN. THE POTENTIAL WAS THERE FOR A CLINICAL DISASTER. LUCKILY THE NETWORK RESUMED OPERATION BEFORE ANY ADDITIONAL MEDICATIONS WERE NEEDED. | 1
3434604 | I WRITE TO YOU TO IDENTIFY A DESIGN FLAW WITH RESPECT TO A RETROFIT TO THE GE ADU ANESTHESIA WORKSTATION THAT I FEEL CAN COMPROMISE PT SAFETY IN THE OPERATING ROOM. OUR INSTITUTION PURCHASED SEVERAL EARLY EDITIONS OF THE GE ADU IN THE LATE 1990'S. THEY WERE FACTORY FITTED WITH 2 CRT MONITORS, THE LEFT ONE FOR VENTILATORS PARAMETERS, THE RIGHT FOR ECG, NIBP, SPO2, ETCO2, ET GAS AND OTHER PT PHYSIOLOGIC PARAMETERS. OVER A PERIOD OF 13-14 YRS THE BRIGHTNESS OF THE CRT MONITORS BEGAN TO FADE. WITHIN THE LAST 12 MONTHS, GE HAS RETROFITTED 2 OF OUR ADU MACHINES WITH NEW CRT MONITORS. THESE MONITORS HAVE A NEW ON/OFF SWITCH ON THE UPPER LEFT BORDER OF THE MONITOR THAT WAS NOT PREVIOUSLY PRESENT. THIS NEW ON/OFF SWITCH HAS A VERY LOW PRESSURE REQUIRED FOR ACTIVATION/DEACTIVATION AND IS THE ONLY POWER SWITCH ON THE ADU THAT IS NOT RECESSED. AS A RESULT, I HAVE HAD AT LEAST 6 SEPARATE EPISODES WHERE STAFF HAS BRUSHED AGAINST THE NON-RECESSED SWITCH INTRAOPERATIVELY, RESULTING IN THE SHUT DOWN OF THE MONITOR DISPLAYING PT PHYSIOLOGIC PARAMETERS DURING SURGERY. THE MONITOR TAKES AT LEAST 1-2 MINS TO REBOOT WITH SOME PARAMETERS TAKING SEVERAL MINS BEFORE THEY ARE AGAIN DISPLAYED. I PERSONALLY HAD A DIFFICULT INTUBATION IN AN EMERGENCY SITUATION, WHERE THIS SWITCH WAS INADVERTENTLY BRUSHED TO THE "OFF" POSITION, WITH LOSS OF SPO2/ECG/ETCO2 WHILE I WAS INTUBATING THE PT. PARALYSIS HAD ALREADY BEEN INDUCED. I HAD NO ETCO2 TO CONFIRM SUCCESS OR FAILURE OF THE EMERGENCY TUBE PLACEMENT AND NO SPO2 MONITOR TO CONFIRM SUCCESS/FAILURE OF OTHER FORMS OF VENTILATION. STAFF LITERALLY HAD TO RUN TO ICU FOR A COLORIMETRIC CO2 DETECTOR AND TO PACU FOR A PORTABLE SPO2 MONITOR WHILE THE ADU MONITOR POWERED BACK UP AND WENT THROUGH SELF TESTING. THE SIMPLE FIX TO THIS SAFETY PROBLEM WOULD BE TO RECESS THIS SWITCH SO IT COULD NOT BE INADVERTENTLY PRESSED. A THICK, HOLLOW PLASTIC WASHER GLUED OVER THE SWITCH COULD WORK, EFFECTIVELY RECESSING THE SWITCH. AT THE PRESENT TIME I AM TAPING A ROLL OF TAPE OVER THE SWITCH (ACTING LIKE A LARGE WASHER) TO PREVENT ACCIDENTAL DEACTIVATION OF THE SWITCH. THE MAIN POWER SWITCH AND O2 FLUSH VALVE ON THE ADU ARE CURRENTLY RECESSED, AND THE ADU IS LIKELY ONLY GOING TO BE SUPPORTED FOR ANOTHER FEW YRS BY THE MFR. | 1
3135322 | DURING THE LAST 1-2 MINS OF PROCEDURE, THE ANESTHESIOLOGIST FELT THE VENTILATOR PORTION OF THE ANESTHESIA MACHINE WAS NOT AS SMOOTH AS BEFORE AND "WAS ACTING DIFFERENT." PT WAS THEN BAGGED BY THE ANESTHESIOLOGIST FOR THE LAST 1-2 MINS AND THEN TAKEN OFF THE MACHINE AND MOVED TO RECOVERY. NO ADVERSE REACTION TO PT. THE ANESTHESIA MACHINE WAS TAKEN OUT OF USE UNTIL IT WAS EVALUATED. THE FOLLOWING DAY DOCTORS OXYGEN SERVICE INC. EVALUATED THE MACHINE AND REPLACED A STICKY DIAPHRAGM ASSEMBLY. TESTING AND OK FOR USE. | 0
3354671 | CALLER REPORTS BOTH THE SPRINT FIDELIS LEAD AND INTEGRITY ALERT SYSTEM HAVE A LOT OF ISSUES AND HAD BEEN RECALLED BY THE FDA. IN (B)(6), HER LEAD BROKE. THE DOCTOR HAD ALSO PROGRAMMED HER ALERT SYSTEM WRONG. BECAUSE OF THIS, SHE RECEIVED MORE THAN 13 INAPPROPRIATE SHOCKS. SHE CALLED MEDTRONIC TO COMPLAIN, BUT RECEIVED NO ASSISTANCE TO PREVENT THE RECURRENCE OF THESE ADVERSE EVENTS.><ADDITIONAL INFORMATION RECEIVED FROM REPORTER ON (B)(6) 2013 FOR REPORT # (B)(4). CALLER IS REQUESTING ADVERSE EVENT IS CHECKED "YES" AND DEVICE IS AN IMPLANT CHECKED "YES". SHE ALSO STATED SHE RECEIVED 12 INAPPROPRIATE SHOCKS, NOT MORE THAN 13 AS STATED IN INITIAL REPORT. SHE ALSO EXPLAINED THAT MEDTRONIC DID GIVE HER SOME ASSISTANCE. SHE SAID THE SOFTWARE FOR HER DEVICE WAS UPGRADED IN 2008, BUT WAS NOT PROGRAMMED CORRECTLY OR CLOSELY MONITORED, WHICH LED TO ALARM ISSUES AND ALARM FAILURES. SHE STATED THAT UNFORTUNATELY, COMMUNICATION BETWEEN DOCTORS WAS LACKING WHICH ALSO CAUSED ISSUES RELATING TO HER ADVERSE EVENTS. | 0
3004502 | THE CUSTOMER REPORTED THAT DPM CENTRAL STATION'S TRANSMITTER FAILED, WHICH MAY HAVE RESULTED IN A LOSS OF AMBULATORY TELEMETRY MONITORING. NO PT INJURY WAS REPORTED. | 1
3231801 | IT WAS REPORTED THE USER FORGOT TO CONNECT THE FRESH GAS HOSE TO THE VENTILATION UNIT BECAUSE THE PREVIOUS CASE IT WAS OPEN SYSTEM CASE, THEN USER START NEXT OPERATION WITHOUT CONNECTION OF THE FRESH GAS HOSE TO THE VENTILATION UNIT COSY, WHICH LED TO A BRAIN DEATH OF THE PATIENT. IF THE FRESH GAS HOSE IS NOT CONNECTED THE PRE USE CHECK WILL FAIL. THIS TEST IS REQUIRED TO BE PROCEEDED ACCORDING TO THE INSTRUCTIONS FOR USE ON A DAILY BASIS. IF THE FRESH GAS HOSE IS DISCONNECTED DURING A CASE THE VENTILATION WILL BE PERFORMED WITH ROOM AIR. DEPENDING ON THE OXYGEN UPTAKE FROM THE PATIENT AN INSPIRATORY OXYGEN LOW ALARM WILL BE POSTED IF THE SET ALARM LIMITS ARE EXCEEDED. THE LOWEST LIMIT THAT CAN BE SET IS 18%. THE LOGFILE OF THE FABIUS WAS SENT TO THE MFR FOR EVAL. THE LOGFILE INDICATED THAT AT 1:07 IN THE NIGHT A PRE USE CHECK WAS PERFORMED AND FAILED. LATER THAT DAY AT THE TIME OF THE REPORTED OCCURRENCE- THE FABIUS LOG FILE INDICATED A LOW OXYGEN CONCENTRATION (BELOW 10%). THIS ITEM IS RECORDED TWICE IN THE LOG WITHIN 2 MINUTES INTERVALS AND WILL LEAD TO A HIGH PRIORITY ALARM. AFTER FURTHER 8 MINUTES THE MESSAGE O2 LOW INOP IS ENTERED IN THE LOG. THERE EXISTS NO HINT TO A TECHNICAL DEVICE ERROR THAT ASSOCIATES TO THE REPORTED EVENT. PICTURES FROM THE ALARM LOGS WERE TAKEN AND ALSO SENT TO THE MFR. THE PICTURES WITH THE ENTRIES OF THE ALARM LOG WERE ANALYZED. THE ALARM LOG BOOK WAS CLEARED BEFORE THE INVESTIGATION COULD HAVE BEEN STARTED BY DRAGGER. | 0
3039983 | DURING OPERATING ROOM CASE, ANESTHESIA MACHINE FAILED. THE PT HAD TO BE MANUALLY BAGGED WHILE THE ANESTHESIA MACHINES WERE TRADED OUT. | 0
3471959 | ANESTHESIA MACHINE MADE A 'FAN FAILED' ERROR DURING A SURGICAL PROCEDURE, MACHINE HAD TO BE TAKEN OUT OF SERVICE DURING PROCEDURE. | 0
2964142 | HIGH VOLUME FAILURE OF ANESTHETIC VAPORIZERS WITHIN A SHORT PERIOD OF TIME - 17 DEVICES BETWEEN (B)(6) 2013. ALL RESULTED IN A FAILED LOW PRESSURE LEAK TEST ON THE DEVICE. TEC 6 PLUS - SERIAL #'S: (B)(4). | 0
3358285 | CHILDREN'S MEDICAL VENTURES (CHMV) RECEIVED A REPORT FROM A DURABLE MEDICAL EQUIPMENT (DME) ASSOCIATED WITH THE SMARTMONITOR 2 PS (PROFESSIONAL SERIES) DEVICE. THE REPORT STATED THAT THERE WAS A LONG DELAY BETWEEN PATIENT EVENT AND THE DEVICE ALARMING FOR THE PATIENT EVENT. THE SMARTMONITOR DEVICE WAS REPORTEDLY IN USE AT THE TIME OF THE REPORTED ISSUE. NO PATIENT HARM HAS BEEN REPORTED. THE DME WAS CONTACTED FOR A FOLLOW-UP CONVERSATION. HOWEVER, THE DME WAS UNABLE TO PROVIDE THE DATA DOWNLOAD OR THE PRESCRIBED SETTINGS OF THE DEVICE IN QUESTION BUT STATED THAT THE DELAY WAS LONGER THAN 5 SECONDS. (B)(4). THE COMPLAINT ISSUE ALLEGED BY THE CUSTOMER WAS NOT ABLE TO BE CONFIRMED BECAUSE THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DME STATED THAT IT COULD NOT CONFIRM WHETHER OR NOT THE DEVICE HAD BEEN SENT IN FOR EVALUATION. THE SMARTMONITOR2 PS AND SMARTMONITOR2 PSL ARE INTENDED FOR USE IN THE CONTINUOUS MONITORING OF RESPIRATION, HEART RATE, AND SPO2 LEVELS (SMARTMONITOR2 PS ONLY) OF INFANT, PEDIATRIC, AND ADULT PATIENTS. THE DEVICE DETECTS AND ALARMS FOR PERIODS OF HIGH OR LOW HEART RATE, HIGH OR LOW BREATH RATE, AND HIGH OR LOW SATURATION (SMARTMONITOR2 PS ONLY). WHEN USED AS AN INFANT MONITOR, IT IS INTENDED FOR USE IN A HOME OR HOSPITAL ENVIRONMENT. FOR INFANTS ONLY, IT MONITORS AND ALARMS FOR CENTRAL APNEAS. WHEN USED AS A PEDIATRIC OR ADULT MONITOR, IT IS INTENDED FOR USE IN A HOSPITAL ENVIRONMENT. PATIENT ALARM LIMITS ARE SET BY THE HEALTH CARE PROFESSIONAL BEFORE THE SMARTMONITOR 2 PS IS DELIVERED TO THE PATIENT AND ARE TYPICALLY SET UP WITH PRESCRIBED SETTINGS WHICH INCLUDE A DELAY BEFORE RECORDING APNEAS AND A DELAY BEFORE ANNUNCIATING AN ALARM FOR AN APNEA CONDITION. THE USER OF THIS EQUIPMENT IS RESPONSIBLE FOR READING, UNDERSTANDING, AND FOLLOWING THE WARNING AND CAUTION STATEMENTS THROUGHOUT THIS MANUAL. BASED ON A COMPLETE REVIEW OF THE COMPLAINT ALLEGATION, CHMV HAS NO REASON TO BELIEVE THE DEVICE IN QUESTION DID NOT PERFORM TO SPECIFICATION; EFFECTIVELY NOTIFYING THE CAREGIVER OF A POTENTIAL EVENT. IT IS DETERMINED THAT NO FURTHER INVESTIGATION INTO THE ALLEGED COMPLAINT ISSUE IS APPROPRIATE AT THIS TIME. HOWEVER, IF THE DEVICE IS RETURNED FOR INVESTIGATION AND IT IS CONCLUDED THAT A MALFUNCTION TOOK PLACE, A FOLLOW-UP, ADDITIONAL INFORMATION REPORT WILL BE FILED. SEE SCANNED PAGE. | 1
3358266 | CHILDREN'S MEDICAL VENTURES (CHMV) RECEIVED A REPORT FROM A DURABLE MEDICAL EQUIPMENT (DME) ASSOCIATED WITH THE SMARTMONITOR 2 PS (PROFESSIONAL SERIES) DEVICE. THE REPORT STATED THAT THERE WAS A LONG DELAY BETWEEN PATIENT EVENT AND THE DEVICE ALARMING FOR THE PATIENT EVENT. THE SMARTMONITOR DEVICE WAS REPORTEDLY IN USE AT THE TIME OF THE REPORTED ISSUE. NO PATIENT HARM HAS BEEN REPORTED. THE DME WAS CONTACTED FOR A FOLLOW-UP CONVERSATION. HOWEVER, THE DME WAS UNABLE TO PROVIDE THE DATA DOWNLOAD OR THE PRESCRIBED SETTINGS OF THE DEVICE IN QUESTION BUT STATED THAT THE DELAY WAS LONGER THAN 5 SECONDS. (B)(4). THE COMPLAINT ISSUE ALLEGED BY THE CUSTOMER WAS NOT ABLE TO BE CONFIRMED BECAUSE THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DME STATED THAT IT COULD NOT CONFIRM WHETHER OR NOT THE DEVICE HAD BEEN SENT IN FOR EVALUATION. THE SMARTMONITOR2 PS AND SMARTMONITOR2 PSL ARE INTENDED FOR USE IN THE CONTINUOUS MONITORING OF RESPIRATION, HEART RATE, AND SPO2 LEVELS (SMARTMONITOR2 PS ONLY) OF INFANT, PEDIATRIC, AND ADULT PATIENTS. THE DEVICE DETECTS AND ALARMS FOR PERIODS OF HIGH OR LOW HEART RATE, HIGH OR LOW BREATH RATE, AND HIGH OR LOW SATURATION (SMARTMONITOR2 PS ONLY). WHEN USED AS AN INFANT MONITOR, IT IS INTENDED FOR USE IN A HOME OR HOSPITAL ENVIRONMENT. FOR INFANTS ONLY, IT MONITORS AND ALARMS FOR CENTRAL APNEAS. WHEN USED AS A PEDIATRIC OR ADULT MONITOR, IT IS INTENDED FOR USE IN A HOSPITAL ENVIRONMENT. PATIENT ALARM LIMITS ARE SET BY THE HEALTH CARE PROFESSIONAL BEFORE THE SMARTMONITOR 2 PS IS DELIVERED TO THE PATIENT AND ARE TYPICALLY SET UP WITH PRESCRIBED SETTINGS WHICH INCLUDE A DELAY BEFORE RECORDING APNEAS AND A DELAY BEFORE ANNUNCIATING AN ALARM FOR AN APNEA CONDITION. THE USER OF THIS EQUIPMENT IS RESPONSIBLE FOR READING, UNDERSTANDING, AND FOLLOWING THE WARNING AND CAUTION STATEMENTS THROUGHOUT THIS MANUAL. BASED ON A COMPLETE REVIEW OF THE COMPLAINT ALLEGATION, CHMV HAS NO REASON TO BELIEVE THE DEVICE IN QUESTION DID NOT PERFORM TO SPECIFICATION; EFFECTIVELY NOTIFYING THE CAREGIVER OF A POTENTIAL EVENT. IT IS DETERMINED THAT NO FURTHER INVESTIGATION INTO THE ALLEGED COMPLAINT ISSUE IS APPROPRIATE AT THIS TIME. HOWEVER, IF THE DEVICE IS RETURNED FOR INVESTIGATION AND IT IS CONCLUDED THAT A MALFUNCTION TOOK PLACE, A FOLLOW-UP, ADDITIONAL INFORMATION REPORT WILL BE FILED. | 1
2968928 | THE CUSTOMER REPORTED AN INCORRECT INTERPRETATION OF A CONTROL ON THE TANGO OPTIMO ANALYZER. THE PICTURES OF THE WELL REACTIONS WERE CLEARLY 4+/4+/-. HOWEVER, THE TANGO INTERPRETED THE RESULT AS 4+/-/-. UPON RE-RUN THE RESULT WAS INTERPRETED AS EXPECTED. AN INVESTIGATION AND AN INSPECTION OF THE AFFECTED TANGO OPTIMO IS STILL ONGOING.><THE CUSTOMER REPORTED AN INCORRECT INTERPRETATION OF A CONTROL ON THE TANGO OPTIMO ANALYZER. THE PICTURES OF THE WELL REACTIONS WERE CLEARLY 4+/4+/-. HOWEVER, THE TANGO INTERPRETED THE RESULT AS 4+/-/-. UPON RE-RUN THE RESULT WAS INTERPRETED AS EXPECTED. A FIELD SERVICE ENGINEER CHECKED THE AFFECTED TANGO OPTIMO THOROUGHLY BUT NOTHING INDICATES INTO THE DIRECTION OF PERFORMANCE ISSUES RELATED TO THE AUTOMATED IMAGE QUALIFICATION. DESPITE SEVERAL INQUIRIES, COMPLAINT RESULT IMAGES HAVE NOT BEEN PROVIDED. BASED ON THE INFORMATION PROVIDED WITH THE FIELD SERVICE REPORT WE HAVE NO INSTIGATION TO SUSPECT INSTRUMENT PERFORMANCE ISSUES TO BE CAUSATIVE FOR THE REPORTED PROBLEM OF VISUALLY 4 + REACTIONS BEING INTERPRETED AS NEGATIVE BY THE AFFECTED TANGO OPTIMO. THE INSTRUMENT IS WORKING ACCORDING TO ITS SPECIFICATIONS. AS THE COMPLAINT RESULT IMAGES HAVE NEITHER BEEN PROVIDED AS SCREENSHOT NOR AS PART OF AN INSTRUMENT DATABASE, WE ARE NOT IN THE POSITION TO PROVIDE A POTENTIAL ROOT CAUSE FOR THIS INCIDENT. THIS IS OUR FINAL REPORT ON THIS INCIDENT.><THIS IS OUR INITIAL REPORT ON THIS INCIDENT. | 0
3266895 | THE CUSTOMER REPORTED FALSE POSITIVE REACTIONS PRESUMEDLY CAUSED BY CARRY OVER OF A SAMPLE WITH SEVEN ANTIBODIES. THE CORRECT FUNCTION OF THE ALLEGEDLY DEFECTIVE REAGENTS WERE PROVED BY TESTING THE RETAINED SAMPLES OF OUR QUALITY CONTROL LABORATORY ON TANGO OPTIMO. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DID NOT OBSERVE ANY FALSE POSITIVE REACTIONS. A REVIEW OF THE BATCH RECORD DOCUMENTATION SHOWED NO IRREGULARITIES WHICH MIGHT HAVE NEGATIVELY AFFECTED THE QUALITY OF THE COMPLAINED LOTS. THE CARRYOVER EVENT COULD BE REPRODUCED AT THE MANUFACTURING UNIT AND AFFECTED THE RESULTS OF TWO SUBSEQUENTIALLY PIPETTED SAMPLES TO REACT POSITIVELY INSTEAD OF NEGATIVELY. FURTHERMORE IT WAS SHOWN THAT A KNOWN NEGATIVE DONOR SAMPLE YIELDED A POSITIVE RESULT IN AN INDEPENDENT RUN AFTER BEING TESTED IMMEDIATELY DOWNSTREAM THE ANTIBODY CONTAINING SAMPLE, INDICATING THE PRESENCE OF DETECTABLE MOLECULES TRANSFERRED FROM THERE. A FINAL TITRE WAS NOT DETERMINED. WE STRONGLY BELIEVE THIS CARRY-OVER ISSUE TO BE ASSOCIATED TO SAMPLE-SPECIFICITIES. THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT. | 0
2757033 | THE CUSTOMER REPORTED HAVING PROBLEMS WITH BACK GROUND PICTURE IMAGES ON THE TANGO OPTIMO. THE CUSTOMER REPORTED THAT SHE WAS RUNNING MULTIPLE TYPE AND SCREEN SAMPLES AND ALL OF THE RESULTS WERE FINE EXCEPT THE LAST THREE PATIENTS. THE CUSTOMER STATES THE BACKGROUND HAS TURNED WHITE AND IS NOT A BLUE GREEN COLOR AS NORMAL. THESE THREE PATIENT SAMPLES ALSO HAD BUBBLES IN THE IMAGES. THE LAST THREE PATIENTS OF AN UNKNOWN NUMBER OF REQUESTED PATIENTS (ASSAY ABSD3) HAVE BEEN PROCESSED WITHOUT COOMBS REAGENT BEING ADDED. THIS NUMBER EQUALS NINE CAVITIES, HENCE INVOLVEMENT OF TWO INDIVIDUAL STRIPS. INVESTIGATION BY THE CUSTOMER REVEALED THAT THERE WAS AN ENTRY IN THE EVENT LOG LIST SAYING THAT AHG WAS EMPTY, BUT THE RESULTS WERE NOT FLAGGED. OUR FIELD SERVICE ENGINEER FOUND THE INSTRUMENT WORKING PROPERLY. THE LEVEL DETECTION BOARD WAS REPLACED PREVENTIVELY. THERE IS NO FURTHER INDICATION FOR ANY INSTRUMENT MALFUNCTIONS. SINCE THE AMOUNT OF COOMBS REAGENT THAT IS ASPIRATED USUALLY SUFFICES FOR EIGHT CAVITIES (= ONE STRIP), WE CAN CONCLUDE THAT THE ASPIRATION COMPLETELY FAILED TWICE IN A ROW. TOGETHER WITH THE INFORMATION THAT THE REAGENT VIAL WAS QUALIFIED TO BE EMPTY BY THE SYSTEM AND SUPPORTED BY THE FACT THAT THE ISSUE IS NOT EXISTENT WHEN APPLYING VIALS WITH HIGHER VOLUME LEVEL, THIS STRONGLY SUGGESTS A SUBOPTIMAL TEACHING OF THE Z-MAX POSITION OF THE RIGHT PIPETTOR OF THE INSTRUMENT IN QUESTION. THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT. | 0
2761455 | THE CUSTOMER REPORTED ONE FALSE NEGATIVE ANTIBODY SCREENING TEST WITH BIOTESTCELL 3 ON TANGO OPTIMO. THE CUSTOMER REPORTED THAT THE REACTIONS LOOKED GRAINY BUT HAVE BEEN EVALUATED NEGATIVELY BY TANGO OPTIMO. FURTHERMORE THE CUSTOMER REPORTED THAT THE ANTIBODY COULD BE IDENTIFIED IN GEL METHOD AS AN ANTI-D. THE CUSTOMER HAS NEITHER RETURNED THE SUPPOSEDLY DEFECTIVE PRODUCT NOR THE PATIENT SAMPLE THAT HAD CAUSED THE FALSE NEGATIVE TEST RESULT. THEREFOR,E OUR QUALITY CONTROL LABORATORY TESTED THE RETAINED SAMPLE WITH DIFFERENT POSITIVE AND NEGATIVE CONTROLS, AMONGST OTHERS AN ANTI-D REFERENCE REAGENT WITH 0.05 IU / ML. THIS REFERENCE REAGENT CORRECTLY REACTED POSITIVELY WITH CELL #1 AND #2 OF BIOTESTCELL 3. ALL POSITIVE AND NEGATIVE REACTIONS WERE CORRECT. WE DID NOT OBSERVE ANY FALSE NEGATIVE REACTIONS. THE AFFECTED TANGO OPTIMO WAS INSPECTED AND THE RESULT CAMERA RGB VALUES HAVE BEEN FOUND TO BE OUT OF RANGE. THEY HAVE BEEN CORRECTED BY THE FIELD SERVICE ENGINEER IN CHARGE. THIS IS OUR COMBINED INITIAL AND FINAL REPORT ON THIS INCIDENT. | 0
3113388 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIVITY ON THE ANTIBODY SCREENING ASSAY ON GALILEO ECHO M01258. A REVIEW OF THE ECHO RESULT FILES FOR THE NEGATIVE ANTIBODY SCREEN SHOWED THAT CELL II APPEARED TO HAVE A FUZZY CELL BUTTON. THE CUSTOMER WAS ASKED TO CLEAN THE WASH MANIFOLD AND REPEAT TESTING. POSITIVE RESULTS WERE OBTAINED IN ALL THREE TEST WELLS. CONTROLS PERFORMED AS EXPECTED. THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. | 0
2960531 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIVITY ON THE CROSSMATCH ASSAY ON GALILEO NEO (B)(4). UPON REVIEW OF THE IMAGE RESULT FILES, REACTIONS APPEARED AS REPORTED BY THE INSTRUMENT. NO ADDITIONAL SAMPLE OR DONOR SEGMENTS WERE AVAILABLE FOR INVESTIGATION. AN IMMUCOR FIELD SERVICE ENGINEER PERFORMED THE UNEXPECTED REACTION CHECKLIST. THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. | 0
2881572 | A CUSTOMER REPORTED THAT AN UNEXPECTED NEGATIVE RESULT WAS OBTAINED WITH THE 3-CELL SCREENING ASSAY ON THE GALILEO ECHO (B)(4)) FOR A PATIENT WITH A HISTORY OF ANTI-E AND ANTI-C. A REVIEW OF THE IMAGE FILES FOR INITIAL TESTING SHOWED THAT CELLS 1 AND 3 APPEARED NEGATIVE AS EXPECTED. CELL 2 RESULTED AS NEGATIVE BUT VISUALLY APPEARED POSITIVE (1+). RESULTS FOR REPEAT TESTING VISUALLY APPEARED AS REPORTED BY THE INSTRUMENT. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS. | 0
3113394 | A CUSTOMER OBTAINED UNEXPECTED NEGATIVE REACTIVITY DURING VALIDATION TESTING OF THE MODIFIED ANTIBODY IDENTIFICATION ASSAY ON THE GALILEO NEO (90168). THE ASSAY WAS MODIFIED TO DECREASE UNEXPECTED POSITIVE REACTIVITY AND EQUIVOCAL REACTIONS AND TO PROVIDE BETTER SPECIFICITY. PROVIDED THE ANTIBODY SCREEN WAS POSITIVE, ADDITIONAL TESTING WOULD NEED TO BE PERFORMED PRIOR TO TRANSFUSION. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
3358134 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIVITY ON GALILEO ECHO (B)(4). A REVIEW OF THE IMAGE RESULT FILES FOR THE UNEXPECTED NEGATIVE REACTIVITY SHOWED THAT SOME TEST WELLS THAT WERE REPORTED AS NEGATIVE BY THE ECHO VISUALLY APPEARED POSITIVE. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS. | 0
2916693 | A CUSTOMER REPORTED THAT UNEXPECTED NEGATIVE REACTIVITY WAS OBTAINED WITH THE ANTIBODY SCREENING ASSAY ON GALILEO ECHO (B)(4). NO ADDITIONAL INFORMATION HAS BEEN RECEIVED. NO SAMPLE RETURNED FOR INVESTIGATION. | 0
2844356 | A CUSTOMER STATED THAT AN ABO DISCREPANCY OCCURRED BETWEEN MANUAL TUBE TESTING AND TESTING PERFORMED ON GALILEO ECHO INSTRUMENT (B)(4). THE CUSTOMER WAS REFERRED TO THE LIMITATION SECTION OF THE OPERATOR MANUAL WHICH STATES THE GALILEO CANNOT RELIABLY DETECT HEMAGGLUTINATION REACTIONS THAT ARE GRADED AS 1+ OR LESS IN TEST TUBE METHODOLOGY. THE GALILEO DOES NOT GENERATE AN INTERPRETATION OF MIXED-FIELD. SUCH A MIXED-FIELD REACTION WILL BE INTERPRETED AS POSITIVE, NEGATIVE, OR EQUIVOCAL. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
3123867 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS ON THE ANTIBODY SCREENING ASSAY ON GALILEO ECHO (B)(4). A REVIEW OF THE COLOR CHECK IMAGES SHOWED THAT PLASMA WAS NOT PIPETTED INTO THE TEST WELLS WHICH CAN LEAD TO UNEXPECTED NEGATIVE INTERPRETATIONS. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) REGARDING LIQUID LEVEL DETECTION AND SAMPLE DISPENSING ON THE ECHO. THE INSTRUMENT IS OPERATING AS EXPECTED. | 0
2992654 | A CUSTOMER REPORTED THAT GALILEO ECHO (B)(4) ASSIGNED THE INCORRECT SAMPLE IDENTIFIER TO A PATIENT SAMPLE. A REVIEW OF THE RESULT FILES SHOWED THAT REACTIONS APPEARED AS REPORTED BY THE ECHO FOR EACH SAMPLE. ON (B)(4) 2013, SAMPLES WERE RECEIVED FROM THE CUSTOMER. A REVIEW OF SAMPLE BARCODE LABELS DID NOT SHOW ANY ARTIFACTS OR PRINT QUALITY ISSUES. THE TUBES WERE PLACED IN A SAMPLE RACK AND LOADED ONTO THE ECHO. THE BARCODE FOR SAMPLE TUBE 069526 WAS SEATED LOW IN THE SAMPLE RACK. THERE WAS MINIMAL WHITE SPACE BETWEEN THE BARCODE AND THE SAMPLE RACK. AS PART OF THE INSTALLATION PROCESS, CUSTOMERS ARE ASKED TO SEND IN EXAMPLES OF THEIR BARCODE LABELS TO DETERMINE THE BARCODE SYMBOLOGY AND IF ANY CONFIGURATIONS ARE NEEDED FOR THE ECHO TO READ THE BARCODES. THE CUSTOMER PROVIDED EXAMPLES OF THEIR BARCODES AND IT WAS DETERMINED THAT THE BARCODE SYMBOLOGY WAS CODE 39 WITHOUT CHECK DIGITS. CUSTOMERS ARE ADVISED THAT THE RECOMMENDED CORRECTION IS IMPLEMENTATION OF CHECK DIGITS IN THE CODE 39 BARCODES. A CHECK DIGIT IS A NUMERIC DIGIT USED TO ENSURE THAT A BARCODE NUMBER IS ENTERED ACCURATELY INTO THE COMPUTER. THE LABEL POSITION IN CONJUNCTION WITH HAVING CODE 39 BARCODES WITHOUT CHECK DIGITS ARE POSSIBLE CAUSES OF THE MISREAD SAMPLE ID. | 0
3262497 | A CUSTOMER REPORTED OBTAINING UNEXPECTED POSITIVE REACTIVITY AND FAILED QC DURING ANTIGEN TYPING ON GALILEO ECHO M01342. A REVIEW OF THE IMAGE RESULT FILES SHOWED THAT RED CELL IMAGES OF THE POSITIVE REACTIONS APPEARED TO HAVE UNSUSPENDED RED CELLS. AN IMMUCOR FIELD SERVICE ENGINEER MADE ADJUSTMENTS TO THE SHAKE GAP. THE PHENOTYPE QC WAS PERFORMED WITH ACCEPTABLE RESULTS. THE INSTRUMENT IS PERFORMING WITHIN SPECIFICATIONS. | 0
3062562 | A CUSTOMER REPORTED THAT UNEXPECTED NEGATIVE REACTIVITY WAS OBTAINED ON THE ANTIBODY SCREENING ASSAY ON GALILEO ECHO (B)(4) WITH A KNOWN ANTI-E POSITIVE SAMPLE. A REVIEW OF THE RESULT FILES FOR THE UNEXPECTED NEGATIVE REACTIVITY SHOWED THAT CELL 2, WHICH IS A HOMOZYGOUS E POSITIVE CELL, VISUALLY APPEARED POSITIVE. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS. | 0
3062654 | A CUSTOMER REPORTED THAT AN UNEXPECTED NEGATIVE RESULT WAS OBTAINED ON GALILEO ECHO (B)(4) WITH A SAMPLE KNOWN TO CONTAIN ANTI-K. A REVIEW OF THE COLOR CHECK IMAGES SHOWED THAT SEVERAL TEST WELLS CONTAINED BUBBLES. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) REGARDING LIQUID LEVEL DETECTION AND THAT THE PRESENCE OF BUBBLES MAY CAUSE THE SAMPLE TO BE PIPETTED INCORRECTLY LEADING TO UNEXPECTED NEGATIVE INTERPRETATIONS. | 0
2754416 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE RESULTS ON GALILEO ECHO M01232. AN IMMUCOR TECHNICAL SUPPORT SPECIALIST REVIEWED THE IMAGE RESULT FILES FOR THE ECHO INSTRUMENTS. RESULTS OF THE REPEAT ANTIBODY SCREEN WERE REPORTED AS NEGATIVE. VISUALLY, CELL 2 APPEARED WEAKLY POSITIVE. AN IMMUCOR FIELD SERVICE ENGINEER FOUND THAT THE RESIDUAL VOLUME WAS LOW. THE RESIDUAL VOLUME WAS RESET. QUALITY CONTROL WAS PERFORMED AND SAMPLES WERE TESTED. THE EXPECTED RESULTS WERE OBTAINED. THE SYSTEM IS OPERATING ACCORDING TO SPECIFICATIONS. THE CUSTOMER IS AWARE OF TECHNICAL COMMUNICATION (B)(4) WHICH STATES TO VERIFY ALL NEGATIVE SCREENING AND IDENTIFICATION ASSAY RESULTS ON THE ECHO PRIOR TO RELEASING RESULTS. NO ADDITIONAL SAMPLE WAS AVAILABLE FOR INVESTIGATION. | 0
3036155 | A CUSTOMER OBTAINED UNEXPECTED NEGATIVE REACTIVITY ON GALILEO NEO (B)(4). THE IMAGE RESULT FILES WERE REVIEWED AND INITIAL TESTING APPEARED NEGATIVE AS REPORTED BY THE INSTRUMENT. THE CUSTOMER DID NOT PERFORM ADDITIONAL TESTING FOR FURTHER INVESTIGATION. NO ADDITIONAL INFORMATION WAS PROVIDED. THE INSTRUMENT IS OPERATING WITHIN SPECIFICATIONS. | 0
2881573 | A CUSTOMER REPORTED THAT AN UNEXPECTED NEGATIVE RESULT WAS OBTAINED WITH THE ANTIBODY IDENTIFICATION PANEL ON THE NEO. A REVIEW OF THE IMAGE FILES AT THE READER CHECK STAGE SHOWED EMPTY WELLS WITH A FEW BUBBLES VISIBLE. THE CUSTOMER WAS MADE AWARE OF TECHNICAL COMMUNICATION (B)(4) REGARDING LIQUID LEVEL DETECTION AND THAT THE PRESENCE OF BUBBLES MAY CAUSE THE SAMPLE TO BE PIPETTED INCORRECTLY LEADING TO UNEXPECTED NEGATIVE INTERPRETATIONS. | 0
2992502 | A SIEMENS HEALTHCARE DIAGNOSTICS FIELD SERVICE ENGINEER, WHO WAS REPLACING THE SAMPLE PROBE OF A SYSMEX CA-1500 ANALYZER, SERIAL NUMBER (B)(4), WAS STUCK BY THE SAMPLE NEEDLE WHILE DISPOSING OF THE PART. HIS SKIN WAS PUNCTURED AND HE SUFFERED POTENTIAL EXPOSURE TO BLOOD-BORNE PATHOGENS. IT IS NOT KNOWN IF HE WAS WEARING APPROPRIATE PERSONAL PROTECTIVE EQUIPMENT (PPE), SUCH AS GLOVES, AT THE TIME. THE FSE WAS TREATED AT THE ER AND GIVEN ANTI-VIRAL PROPHYLAXIS. IT IS NOT KNOWN IF THE FSE OR ANY POTENTIAL SOURCE PTS WERE TESTED FOR EXPOSURE TO BLOOD-BORNE PATHOGENS. THE EVENT OCCURRED IN (B)(6) AND WAS REPORTED TO SYSMEX CORP (B)(4) 2013. NO MALFUNCTION OF THE SYSMEX CA-1500 ANALYZER WAS REPORTED. THE FSE WAS SERVICING THE ANALYZER AND IT WAS NOT IN OPERATION AT THE TIME OF THE EVENT. A REVIEW OF THE SYSMEX CA-1500 SERIES INSTRUCTIONS FOR USE INCLUDES NUMEROUS WARNINGS REGARDING SAFE HANDLING OF THE SAMPLE PROBE. A USER IS ADVISED TO HANDLE ALL INSTRUMENTS PARTS AS BIOLOGICALLY HAZARDOUS, TO WEAR GLOVES WHILE HANDLING THE PROBES AND TO WASH HANDS WITH DISINFECTANT AFTER CLEANING. WHEN CLEANING, THE PROBE SHOULD BE HANDLED FROM TOP TO BOTTOM TO PREVENT PIERCING A FINGER. A USER ALSO IS WARNED THAT THE PIERCER IS EXTREMELY SHARP AND DANGEROUS AND THAT THE USED PIERCER BE DISCARDED IN A SAFE PLACE TO PREVENT STICK INJURY. CURRENT STATUS OF THE FSE IS NOT KNOWN. THIS OCCURRENCE IS REPORTED IN REGARD TO THE (B)(4) NEEDLE-STICK MEDICAL DEVICE REPORTING GUIDELINE, AS ANTI-VIRAL POST-EXPOSURE PROPHYLAXIS WAS ADMINISTERED IN ADDITION TO FIRST AID. NONE OF THE COMPLAINT NARRATIVE STATES OR IMPLIES THAT THE DEVICE FAILED IN ANY WAY. SPECIFIC CORRECTIVE ACTION ON THE INSTRUMENT, OTHER THAN REPLACEMENT OF THE SAMPLE PROBE, IS NOT KNOWN. | 0
3091032 | THE PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL. INCORRECT RESULTS WERE REPORTED. THERE WAS NO HARM TO ANY PATIENT. ON (B)(4) 2013 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND PERFORMED A PM AND A CHECK OF FLUIDICS AND THE READER CAMERA. POST SERVICE THE CUSTOMER RAN THE SAMPLE THAT WAS IN QUESTION WHICH FIRST WAS REPORTED AS A NEGATIVE IN THE SCREEN CELLS, AND WAS 1+ ON ANOTHER PROVUE ANALYZER. THIS TIME THIS ANALYZER REPORTED THE SAMPLE AS 1+. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
2900086 | THE CUSTOMER REPORTS 2 SEPARATE EVENTS OF FALSE POSITIVE REACTIONS TO THE ANTI-A MICRO WELL IN THE MTS ABD/ABD GEL CARD ISSUED BY THE PROVUE (1 EVENT WITH A DONOR SAMPLE AND 1 EVENT WITH A QC SAMPLE). THE EVENT INVOLVING THE DONOR SAMPLE IS REPORTABLE. NO INCORRECT RESULTS WERE REPORTED. CTS CONFIRMED THAT THE MTS DILUENTS AND THE LISTED GEL CARDS ALL HAVE NORMAL APPEARANCE AND HAVE BEEN STORED ACCORDINGLY. ALL FOILS TO THE GEL CARDS CONFIRMED TO BE SEALED. CTS GUIDED THE CUSTOMER INTO THE LOG FILES AND CONFIRMED THE GRIPPER TO BE HOLDING THE GEL CARDS INCORRECTLY BY REVIEWING RANDOM .TIF IMAGES. ON (B)(4) 2013 AN OCD FIELD SERVICE ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND FOUND THE CARDS TO BE SLIGHTLY CROOKED AND THE RIGHT REAR GRIPPER PIN TO BE SLIGHTLY BINDING. THE FE REPLACED ALL FOUR GRIPPER PINS AND THE GRIPPER PIN SPRINGS. THE FE PERFORMED ALL GRIPPER ADJUSTMENTS. THE FE VERIFIED PROPER OPERATION BY PERFORMING THE DIAGNOSTICS CARD READ PROCEDURE 3 TIMES WITHOUT ERROR. ALL IMAGES WERE AS EXPECTED. QC PERFORMED AND ACCEPTED BY THE CUSTOMER. REPAIRS HAVE RETURNED THE INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
3247569 | THE ORTHO PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. ON (B)(6), 2013 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND FOUND THE CARDS WERE NOT LEVEL DURING READING. THE FE LEVELED THE GRIPPER AND PERFORMED THE ADJUSTMENTS. THE FE PERFORMED THE READER CAMERA ADJUSTMENTS. THE FE REPLACED THE SYRINGE AFTER FINDING WHITE CRYSTALS AT THE BASE OF THE GLASS EXTERIOR OF THE SYRINGE. THE CUSTOMER THEN RAN CONTROLS AND ALL CONTROLS PASSED. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4). | 0
2868263 | THE CUSTOMER STATES THAT 2 SAMPLES TESTED (SAMPLES FROM SAME BATCH RUN) ON THE ORTHO PROVUE IN THE FORWARD GROUPING PORTION OF THE ABO BLOOD GROUPING TESTS RESULTED IN UNEXPLAINED HEMOLYSIS WHILE THE SAME 2 SAMPLES WERE REPEATED BY THE MANUAL GEL METHOD AND RESULTS WERE SATISFACTORY WITH NO HEMOLYSIS IN ANY OF THE FORWARD TYPING. NO ERRONEOUS RESULTS WERE REPORTED. 1 OF 2 EVENTS. ON (B)(4) 2012 AN OCD FIELD ENGINEER (FE) WENT TO THE CUSTOMER SITE AND VERIFIED THE ERRORS IN THE LOG FILES. THE FE RAN A SAMPLE AND WAS ABLE TO SEE THE SAME RESULTS AS STATED FOR SERVICE (UNEXPLAINED HEMOLYSIS). THE FE THEN REPLACED THE CAVRO DILUTOR AND PROBE. THE FE PRIMED THE SYSTEM SEVERAL TIMES. THE FE RAN THE SAME SAMPLE AGAIN WITH GOOD RESULTS. THE CUSTOMER RAN AND ACCEPTED QC. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. ALL REQUIRED DOCUMENTATION WILL BE GIVEN TO THE CUSTOMER. (B)(4). | 0
3021003 | THE CUSTOMER STATES THAT THEY RECEIVED A FALSE NEGATIVE RESULT USING THE ORTHO PROVUE ON A SAMPLE THAT WAS POSITIVE IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. THE CUSTOMER CONFIRMED WITH CTS THAT THE DAILY QC IS ACCEPTABLE. THE CUSTOMER CONFIRMED WITH CTS THAT VISUAL INSPECTION OF THE GEL CARDS SHOWED THAT THE REACTIONS WERE CLEARLY NEGATIVE. WHEN THE CUSTOMER TESTED THE SAME SAMPLE IN THE MANUAL GEL METHOD USING THE SAME LOT OF REAGENTS A 2+, REACTIVITY WAS NOTED WITH CELL #1 AND CELL #2. THE CUSTOMER PROCEED TO WORK UP THE ANTIBODY USING THE TUBE METHOD AND DETERMINED THAT THE PATIENT HAS A WARM AUTO WITH NO SPECIFICITY. (B)(4). SERVICE NOT REQUESTED. | 0
3236712 | THE ORTHO PROVUE GEL CAMERA MISREAD HAZE AND RESULTED IT AS NEGATIVE. NO ERRONEOUS RESULTS WERE REPORTED. ON (B)(4) 2013 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND REVIEWED BOTH OF THE IMAGES WITH THE CUSTOMER. THE FE PERFORMED THE READER CAMERA ADJUSTMENTS AND CREATED A NEW REFERENCE IMAGE. THE CUSTOMER RAN QC AND ACCEPTED THE RESULTS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. COMPLAINT : (B)(4). | 0
3346765 | THE CUSTOMER IS REPORTING A FALSE NEGATIVE REACTION WITH ONE SAMPLE IN THE ANTI-D MICROTUBE OF THE ABD/REV GEL CARD LOT 042313037-30 EXP 11-APR-2014. THE SAME PATIENT SAMPLE HAS TESTED IN THE MANUAL GEL METHOD AND FOUND TO BE RH POSITIVE (2+). THIS ISSUE IS OCCURRING ON BOTH PROVUES AT THIS FACILITY. BOTH PROVUES HAVE BEEN TAKEN OUT OF SERVICE. CUSTOMER REFUSES TO TEST PATIENT SAMPLES ON THE INSTRUMENTS UNTIL A FIELD ENGINEER PERFORMS DIAGNOSTIC CHECKS. ISSUE STARTED ON: (B)(6) 2013. FREQUENCY: 1 SAMPLE; SAMPLE ID: NOT PROVIDED; MICROTUBES/WELLS OR CELL (DONOR #) AFFECTED: ANTI-D; METHODOLOGY USED: PROVUE; INCUBATION TIME (FOR MANUAL TEST ONLY): N/A; ERROR CODE: NO ERRORS DURING ANY TESTING; REACTION GRADE OBTAINED: NEGATIVE; CUSTOMER WAS EXPECTING: POSITIVE; TEST REPEATED: YES. BOTH PROVUES INITIALLY GAVE NEGATIVE RESULTS IN THE D MICROWELL. PROVUE# (B)(4) GAVE A 1+ POSITIVE UPON RETEST. THE SAMPLE WAS NOT RETESTED ON PROVUE (B)(4). MANUAL GEL TESTING USING THE SAME LOT OF GEL CARDS WAS DONE TWICE AND PRODUCED A 2+ AND 3+. TUBE METHOD AT IMMEDIATE SPIN GAVE A 3+ REACTION USING IMMUCOR ANTI-D LOT 506132 EXP 9-25-14. METHOD USED TO REPEAT: PROVUE AND TUBE. OTHER RELEVANT DETAILS: PATIENT IS A PREGNANT FEMALE, TYPED AS GROUP B AND DETERMINED TO BE RH POSITIVE WITH D.O.B. (B)(6). PATIENT WAS NOT TRANSFUSED WITH IN A YEAR. NO FURTHER INFORMATION IS KNOWN. NO ERRONEOUS RESULTS WERE REPORTED. NO TRANSFUSIONS WERE GIVEN. CUSTOMER IS NOT CONFIDENT THE PROVUE IS OPERATING AS EXPECTED, ALTHOUGH THE QC TESTING (DONE WITH PATIENT SAMPLES) PASSED. CUSTOMER IS CONCERNED THE PROVUE MISSED VERY WEAKLY POSITIVE REACTIONS WITH THE ABD/REVERSE GEL CARD AND FEARS THE SAME MAY HAPPEN WITH ANTIBODY SCREENS. CUSTOMER FEELS THAT THE RESULT THAT WAS REPORTED AS NEGATIVE BY THE PROVUE SHOULD HAVE BEEN POSITIVE, OR AT LEAST FLAGGED BY THE PROVUE FOR REVIEW. THE TIF IMAGE PRODUCED BY THE PROVUE SHOWING SMALL AGGLUTINATION WAS SCANNED AND EMAILED TO CTS BY THE CUSTOMER. THIS IMAGE WAS CALLED NEGATIVE BY THE PROVUE. THIS CUSTOMER DEMANDED SERVICE WITHOUT PRIOR TROUBLESHOOTING. A SERVICE ORDER WAS CREATED AND AN FE WAS DISPATCHED TO THE CUSTOMER SITE. CUSTOMER STATES BOTH PROVUES AT THIS FACILITY MADE INCORRECT INTERPRETATIONS IN REPORTING THE REACTIONS IN THE ANTI-D MICROTUBES AS NEGATIVE SINCE SMALL AGGLUTINATION IS VISIBLE. CUSTOMER EMAILED THE FILE TO CTS. CUSTOMER STATES ONE IMAGE WAS OBTAINED FOR EACH PROVUE. INITIAL REPORT SENT WITH MFG # 2250051 INSTEAD OF MFG #1056600. THIS REPORT SHOULD NEVER HAVE BEEN SENT USING MFG # 2250051. FILE WILL BE RESUBMITTED UNDER THE CORRECT MFG NUMBER.><THE ROOT CAUSE THAT LED TO THIS EVENT CANNOT BE DETERMINED BASED ON THE INFORMATION REPORTED BY THE CUSTOMER. THE FE RESPONDED TO THE SITE AND VERIFIED THAT THE INSTRUMENT IS PERFORMING TO SPECIFICATIONS. QUALITY CONTROL TESTING BEFORE AND AFTER THE SERVICE CALL PASSED. THERE WAS NO REPORTED HARM TO ANY PATIENTS. (B)(4). | 0
2834931 | CUSTOMER RUN LACTATE SAMPLES ABOVE 10 MMOL/L ON TWO DIFFERENT ABL825'S AS STANDARD PROCEDURE. ON ONE PATIENT SAMPLE THE RESULT WAS 18 MMOL/L ON ONE ABL825 AND 14 MMOL/L ON THE OTHER. THE PATIENT WAS BEING MONITORED FOR LACTATE IN THE ICU. THE CUSTOMER STATED THAT THIS DIFFERENCE COULD HAVE EFFECTED TREATMENT IF THEY HAD NOT RUN THE DOUBLE CHECK, AND THAT THE CLINICIANS COULD HAVE FALSELY THOUGHT THAT THE PATIENTS' TREATMENT WERE WORKING AND LOWERING THE LACTATE VALUE. NO PATIENT MEDICAL INTERVENTION WAS NEEDED. THE FIELD SERVICE ENGINEER INSTALLED THE CUSTOMERS ELECTRODE WITH A NEW MEMBRANE UNIT AND RAN WHOLE BLOOD COMPARISONS. THE INSTRUMENT WAS NOW MEASURING WITHIN THE SPECIFIED PERFORMANCE. RESULTS WERE WITHIN 1 MMOL/L ON ALL VALUES ABOVE 12 MMOL/L, AND QC'S WERE RUNNING CLOSER TO THE TARGET VALUE. RADIOMETER BECAME AWARE OF THE ISSUE THROUGH INTERNAL TESTING IN CONNECTION WITH A CAPA. ON (B)(4) 2012 A FIELD SAFETY NOTICE WAS DISTRIBUTED TO WARN THE USERS OF A POTENTIAL NEGATIVE BIAS THAT MAY OCCUR AFTER EXCHANGE OF THE LACTATE MEMBRANE UNIT. THE CORRECTION WAS REPORTED IN ACCORDANCE WITH 21 CFR 806.10. (B)(4). | 0
3015389 | THE CUSTOMER STATED AN ARCHITECT C16000 ANALYZER GENERATED A FALSELY ELEVATED SODIUM RESULT FOR ONE PATIENT SAMPLE. THE ANALYZER GENERATED AN INITIAL SODIUM RESULT OF 198 AND A REPEAT RESULT OF 140. THE FALSELY ELEVATED SODIUM RESULT WAS NOT REPORTED OUTSIDE THE LABORATORY AND THERE WAS NO ADVERSE IMPACT TO PATIENT MANAGEMENT REPORTED. .><(B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. THE CUSTOMER HAD REPORTED CONTROLS OUT OF RANGE FOR MULTIPLE ASSAYS INCLUDING SODIUM TWO DAYS PRIOR TO THIS COMPLAINT. IN RESPONSE TO THAT ISSUE, THE POPPETT VALVES WERE REPLACED. IT CANNOT BE DETERMINED BASED ON THE AVAILABLE INFORMATION IF THIS ISSUE WAS A FACTOR IN THE CURRENT COMPLAINT. NO ADVERSE TRENDS WERE IDENTIFIED FOR THE C16000. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
3112060 | IT WAS REPORTED THAT THE ANALYZER IS UNRELIABLE IN ITS ABILITY TO DELIVER AND DETECT THE REAGENT. IT WAS ALSO REPORTED THAT THE ANALYZER CONTINUES TESTING W/O REAGENT AND ERROR FLAGS ARE NOT CONSISTENTLY GENERATED. THERE WERE NO PT INJURIES ASSOCIATED WITH THIS REPORT. AN ON-SITE INVESTIGATION WAS CONDUCTED AND THE PROBLEM WAS REPRODUCED. AMONG OTHER THINGS, THE ORIGINAL LEVEL DETECTION STRIP WAS REMOVED AND VISUALLY INSPECTED, WITH NO DEFECTS FOUND. THE ORIGINAL LEVEL DETECTION STRIP WAS REPLACED AND THE SCREW THAT GROUNDS THE STRIP TO THE REAGENT COMPARTMENT WAS TIGHTENED. THE PROBLEM WAS RESOLVED. NO PARTS WERE REPLACED. THE MFR IS ALSO EVALUATING THE PROBLEM BUT HAS BEEN UNABLE TO REPRODUCE IT. THE MFR DETERMINED THAT IT IS POSSIBLE THE PROBLEM IS ASSOCIATED WITH THE TIGHTNESS OF THE SCREW. AS PREVENTIVE ACTION, THE ASSEMBLY MANUAL IS BEING REVISED TO ADD THE FOLLOWING ADD'L CHECKS. THE SCREW THAT SECURES THE LEVEL DETECTION STRIP TO THE R1 REAGENT COMPARTMENT MUST BE TIGHT. MEASURE THE RESISTANCE OF FG TO ENSURE IT IS < 1 OHM. IN ADDITION, NOTIFICATION WILL BE SENT TO THE FIELD SERVICE ENGINEERS. THE NOTIFICATION WILL INSTRUCT THEM TO CHECK THE RESISTANCE OF THE FG TO ENSURE IT IS < 1 OHM AND TIGHTEN THE SCREW IF THIS TYPE OF EVENT IS OBSERVED ON INSTALLED ANALYZERS. | 0
2944136 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) AND REPORTED THAT APPROXIMATELY 5 ML OF UNKNOWN LIQUID HAD LEAKED FROM THEIR COULTER LH 500 HEMATOLOGY ANALYZER. THE OPERATOR WAS WEARING A LAB COAT AND GLOVES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND CONFIRMED A LEAK AT THE CONNECTOR FOR TUBING FROM FEED THROUGH FITTING (FF29) TO T-FITTING FT31 IN THE SAMPLE PATH. FSE FOUND THE CONNECTOR TO BE MOLDED AND REPLACED IT WHICH RESOLVED THE LEAK. THE CAUSE OF THE LEAK IS RELATED TO WORN TUBING FROM FF29 TO FT31. (B)(4). | 0
2981689 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT THAT A COULTER LH 500 HEMATOLOGY ANALYZER WAS LEAKING FROM PINCH VALVE (PV) 3. PV3 IS THE PATHWAY FROM THE WBC APERTURE TO THE WASTE CHAMBER VC1 IN THE CBC MODULE. THE VOLUME OF THE LEAK IS UNKNOWN. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) INCLUDING A GOWN, GLOVES AND GOGGLES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND OBSERVED THAT THE SOURCE OF THE LEAK WAS THE I- BEAM TUBING AT PV3. FSE REPLACED THE TUBING AND PERFORMED A MINI PREVENTIVE MAINTENANCE (PM) ON THE FRONT PART OF THE CBC MODULE. NO FURTHER EVIDENCE OF LEAKING WAS OBSERVED. THE CAUSE OF THE LEAK IS ATTRIBUTED TO THE TUBING AT PV3. (B)(4). | 0
3044677 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT A LEAK OF FAINT GREENISH BLUE FLUID ON THE FRONT CENTER AROUND THE SAMPLE PROBE OF A COULTER LH 780 HEMATOLOGY ANALYZER. THE VOLUME OF THE LEAK WAS DESCRIBED BY THE CUSTOMER AS 1ML AND THE SOURCE OF THE LEAK WAS UNKNOWN. PER CUSTOMER, THE INSTRUMENT GENERATED SWEEPFLOW TANK NOT FULL ERRORS AND THE LEAK HAPPENED AT SHUTDOWN. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED AND THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON-SITE AND CONFIRMED THE LEAK WAS CLENZ (CLEANER) DUE TO A CUT PIECE OF TUBING ON THE SWEEP FLOW MODULE. THE FSE REPLACED THE ENTIRE SWEEP FLOW MODULE WHICH RESOLVED THE LEAK. THE FSE THEN PERFORMED TWO SHUTDOWNS AND STARTUPS WITH NO FURTHER LEAKS OBSERVED. THE CAUSE OF THE LEAK IS ATTRIBUTED TO CUT TUBING IN THE SWEEP FLOW MODULE. THE INSTRUMENT GENERATED SWEEP FLOW TANK NOT FULL ERRORS THEREBY ALERTING THE OPERATOR TO AN INSTRUMENT PROBLEM. (B)(4). | 0
3058858 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT THAT THE BLOOD SAMPLING VALVE (BSV) ON A COULTER LH 500 HEMATOLOGY ANALYZER WAS LEAKING CLEANER. THE VOLUME OF THE LEAK WAS REPORTED AS APPROXIMATELY 0.5 ML. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) INCLUDING A LAB COAT AND GLOVES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON-SITE AND FOUND THAT THE TUBING THROUGH PINCH VALVE PV11 WAS LEAKING CLENZ (CLEANER) AROUND THE FITTING. THE FSE REPLACED THE TUBING WHICH RESOLVED THE LEAK. THE CAUSE OF THE LEAK IS ATTRIBUTED TO THE TUBING THROUGH PV11. (B)(4). | 0
3007834 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT THAT A UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM WAS LEAKING CLEANER. THE VOLUME OF THE LEAK WAS UNKNOWN. THE LEAK WAS CONTAINED WITHIN THE INSTRUMENT. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A GLOVES, EYE PROTECTION AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND OBSERVED THAT CLEANER WAS LEAKING OUT OF THE WASH COLLAR AND NOTED THAT THE WASH COLLAR WAS OUT OF ALIGNMENT. FSE ADJUSTED THE WASH COLLAR AND PERFORMED THE ALIGNMENT. FSE THEN PERFORMED TWO SHUTDOWNS AND CONFIRMED THAT THERE WAS NO MORE LEAKING. THE CAUSE OF THE LEAK IS ATTRIBUTED TO THE WASH COLLAR BEING OUT OF ALIGNMENT. (B)(4). | 0
2981777 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT A LEAK OF CLEAR FLUID AT THE RINSE BLOCK OF A COULTER ACT DIFF 2 ANALYZER. THE VOLUME OF THE LEAK WAS ABOUT 5 ML. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES, GOGGLES AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. NO PATIENT SAMPLES OR RESULTS WERE AFFECTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THAT THE PINCH VALVE LV10 WAS NOT CLOSING PROPERLY. PINCH VALVE LV10 CONTROLS THE DILUENT PUMP TO THE PROBE WASH AND THE DILUENT PUMP TO THE DILUENT RESERVOIR. FSE REPLACED PINCH VALVE LV10 AND THE LEAK WAS RESOLVED. THE REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES. THE CAUSE OF THE LEAK IS ATTRIBUTED TO THE PINCH VALVE LV10 NOT CLOSING PROPERLY. (B)(4). | 0
3028526 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT A LEAK INSIDE A COULTER HMX HEMATOLOGY ANALYZER. THE INSTRUMENT WAS GIVING 'FLOW CELL CLOG' ERRORS WHEN THE LEAK WAS NOTICED. THE VOLUME OF THE LEAK WAS ABOUT 50 ML. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES. NO ERRONEOUS RESULTS WERE GENERATED OR REPORTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND DID NOT FIND ANY LEAKS INSIDE THE INSTRUMENT. THE CUSTOMER TOLD FSE THAT THE TUBING HAD POPPED OFF PORT 11 OF THE BLOOD SAMPLING VALVE (BSV). THE CUSTOMER HAD BEEN ABLE TO REPLACE THE TUBING WITH HELP FROM BEC CUSTOMER TECHNICAL SUPPORT (CTS) AND THIS HAD RESOLVED THE LEAK. THE INSTRUMENT WAS STILL GIVING FLOW CELL CLOG ERRORS. FSE OBSERVED THAT THE BSV WAS NOT MOVING PROPERLY. FSE REPLACED THE BSV ACTUATOR AND THE INSTRUMENT THEN RAN WITHOUT ANY LEAKS OR ERRORS. FSE ALSO REPLACED THE DIFFERENTIAL MIXING CHAMBER AND FLUSHED THE VACUUM LINES AS A PREVENTIVE MEASURE. THE REPAIRS WERE VERIFIED PER ESTABLISHED PROCEDURES. THE CAUSE OF THE LEAK WAS THAT TUBING POPPED OFF OF PORT 11 OF THE BSV. THE INSTRUMENT OPERATED AS INTENDED BY GENERATING FLOW CELL CLOG ERRORS, ALERTING THE OPERATOR TO AN INSTRUMENT PROBLEM. (B)(4). | 0
3036146 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT 2ML OF WHAT APPEARED TO BE SPLATTERED DILUTED BLOOD BEHIND THE TRANSPORT SHIELD IN A UNICEL DXH 800 COULTER CELLULAR ANALYSIS SYSTEM. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT AND GLOVES AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. MEDICAL ATTENTION WAS NOT SOUGHT. NO ERRONEOUS PATIENT RESULTS WERE GENERATED. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON-SITE AND CONFIRMED THE LEAK FROM A BROKEN SPRING INSIDE THE WASH COLLAR ASSEMBLY WHICH WAS NOT ALLOWING THE WASH COLLAR TO EXTEND DOWNWARD COVERING THE END OF THE PROBE. THE FSE REPLACED THE SPRING WHICH RESOLVED THE LEAK. THE CAUSE OF THE LEAK IS ATTRIBUTED TO A BROKEN SPRING INSIDE THE WASH COLLAR. (B)(4). | 0
3092604 | A CUSTOMER CONTACTED BECKMAN COULTER (BEC) TO REPORT A LEAK IN A COULTER HMX HEMATOLOGY ANALYZER. THE VOLUME OF THE LEAK WAS UNKNOWN. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) INCLUDING A LABCOAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THAT THE LEAK WAS COMING FROM THE BSV (BLOOD SAMPLING VALVE) ASSEMBLY. FSE TOOK APART THE BSV ASSEMBLY, CLEANED IT AND REATTACHED BUT IT CONTINUED TO LEAK. THE FSE THEN REPLACED THE BSV ASSEMBLY WHICH RESOLVED THE LEAK AND THE UNIT WAS CONFIRMED AS OPERATIONAL. THE CAUSE OF THE LEAK IS ATTRIBUTED TO THE BSV ASSEMBLY. (B)(4). | 0
2745783 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT WHEN THEY TAKE SAMPLES OUT OF THE CLOSED VIAL STATION, THERE ARE A FEW DROPS OF DILUENT ON TOP OF THE CAP DUE TO A LEAK AT THE PROBE OF THEIR COULTER ACT DIFF 2 ANALYZER. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE INCIDENT AND NO INJURY OR EXPOSURE WAS REPORTED. NO PATIENT SAMPLES WERE AFFECTED BY THE LEAK. A SERVICE REQUEST WAS SET UP; HOWEVER, THE CUSTOMER CALLED BACK ON (B)(6) 2012 AND CANCELLED THE SERVICE REQUEST BECAUSE THE INSTRUMENT WAS NO LONGER LEAKING. THE CUSTOMER DID NOT KNOW HOW THE LEAK WAS REPAIRED. THE CAUSE OF THE LEAK IS UNKNOWN. (B)(4). | 0
2927559 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) AND REPORTED THAT ONE PATIENT SAMPLE RUN ON THEIR COULTER ACT 8/10 ANALYZER GAVE A HGB RESULT OF 7.3G/DL WITHOUT AN INSTRUMENT FLAG; THE PATIENT WAS SENT TO THE HOSPITAL AND REDRAWN WHERE THE HGB WAS 8.0G/DL AND WAS CONSIDERED CORRECT. BEC REVIEW OF THE CUSTOMER SUPPLIED DATA SHOWED THAT THE INSTRUMENT GAVE HIGHER PLATELET (PLT) RESULTS COMPARED TO THE REFERENCE INSTRUMENT. HOWEVER, THE INSTRUMENT GAVE AN INSTRUMENT GENERATED FLAG ON PLT TO ALERT THE OPERATOR TO FURTHER REVIEW THE SAMPLE RESULTS. THE PATIENT WAS NOT ADMITTED TO THE HOSPITAL, JUST REDRAWN AND RUN ON A REFERENCE INSTRUMENT. THERE WAS NO IMPACT OR AFFECT TO PATIENT TREATMENT WITH REGARD TO THIS EVENT. CONTROLS WERE WITHIN LIMITS AT THE TIME OF THE EVENT AND THE INSTRUMENT IS CURRENTLY PERFORMING WITHIN QC SPECIFICATIONS. SAMPLE COLLECTION DETAILS WERE NOT PROVIDED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND NO INSTRUMENT PROBLEMS. CONTROLS RECOVERED WITHIN THE MEAN AND REPRODUCIBILITY WAS WITHIN SPECIFICATIONS. FSE NOTED THE TECHNOLOGY DIFFERENCES BETWEEN INSTRUMENTS AND STATED THAT THE RESULT OF 7.3G/DL ON THE ACT8/10 INSTRUMENT AND THE 8.0G/DL ON THE REFERENCE INSTRUMENT WERE BOTH BELOW THE PATIENT LOW LIMIT. THE CAUSE OF THE EVENT IS UNKNOWN. | 0
2857199 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A LEAK AT THE PROBE OF THEIR COULTER ACT DIFF ANALYZER. THE VOLUME OF THE LEAK WAS A FEW DROPS AND WAS NOT CONTAINED WITHIN THE INSTRUMENT. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, GOGGLES AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED AND NO PATIENT SAMPLES OR RESULTS WERE AFFECTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THAT THE RINSE BLOCK WASTE LINE WAS CLOGGED. FSE REPLACED THE RINSE BLOCK ASSEMBLY AND THE LEAK WAS RESOLVED. THE REPAIR WAS VERIFIED PER ESTABLISHED PROCEDURES. THE CAUSE OF THE LEAK WAS THAT THE RINSE BLOCK WASTE LINE WAS CLOGGED. (B)(4). | 0
2798079 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A 5 ML CLEAR FLUID LEAK IN THEIR COULTER LH 500 HEMATOLOGY ANALYZER AFTER THE SHUTDOWN/STARTUP WAS COMPLETED. NO PATIENT SAMPLES OR RESULTS WERE AFFECTED. THE LEAK WAS NOT CONTAINED WITHIN THE INSTRUMENT. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. THE CUSTOMER WAS UNABLE TO LOCATE THE SOURCE OF THE LEAK AND A SERVICE VISIT WAS SET UP. A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED ON-SITE THE DAY OF THE EVENT AND REATTACHED TUBING THAT HAD POPPED OFF FITTING AT FEED THROUGH (FF42) ON THE BLOOD SAMPLE VALVE (BSV) MODULE WHICH RESOLVED THE ISSUE. THE CAUSE OF THE LEAK IS THE TUBING AT FF42 - THE ASPIRATION PATH FROM THE BSV TO PRIMARY-MODE ASPIRATION PUMP (PM4). (B)(4). | 0
2908957 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING THAT THEY HAD COMPLETED A SHUTDOWN/STARTUP AND WERE PREPARING TO RUN QC WHEN A HALF CUP FLUID LEAKED FROM THEIR COULTER LH 500 HEMATOLOGY ANALYZER AND FLOWED ONTO THE COUNTER. THE INSTRUMENT DID NOT GIVE ANY ERRORS AND THEY WERE UNABLE TO LOCATE THE SOURCE OF THE LEAK. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, FACE SHIELD AND GLOVES AT THE TIME OF THE INCIDENT. NO INJURY OR EXPOSURE WAS REPORTED AND THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND FOUND THAT BLUE TUBING FROM FF (FEED THROUGH FITTING) 174 TO FF 180 ON VALVE PV49 HAD SPLIT. FSE REPLACED ALL TUBING THROUGH PV49 WHICH REPAIRED THE LEAK. FSE THEN COMPLETED SHUT DOWN AND VERIFICATION WITH NO LEAKS. THE CAUSE OF THE LEAK WAS A SPLIT IN THE TUBING FROM FF 174 TO FF 180 ON PV49. (B)(4). | 0
2927560 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THERE WAS RESIDUE THAT RESEMBLED DRY DILUENT UNDER THEIR COULTER ACT 5 DIFF CAP PIERCE (CP) INSTRUMENT. THE INSTRUMENT WAS NOT IN USE. THE SOURCE OF THE LEAK COULD NOT BE DETERMINED. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED AND THERE WAS NO AFFECT TO PATIENT SAMPLES OR RESULTS. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE AND OBSERVED THAT THE REAGENT BLOCK WAS CRACKED ON THE HEMOGLOBIN (HGB) LYSE END. FSE REPLACED THE CRACKED REAGENT BLOCK AND PRIMED REAGENTS. FSE THEN RAN THE INSTRUMENT AND NO FURTHER EVIDENCE OF LEAKING WAS FOUND. THE CAUSE OF THE LEAK WAS A CRACKED REAGENT BLOCK. (B)(4). | 0
2823382 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) STATING THAT THEIR COULTER ACT DIFF ANALYZER LEAKED 10 DROPS OF CLEAR FLUID DURING STARTUP. THE CUSTOMER DID NOT KNOW THE SOURCE OF THE LEAK. THE LEAK WAS NOT CONTAINED AND THE INSTRUMENT RECOVERED A FATAL ERROR 7 (WHICH ALERTED THE OPERATOR TO AN INSTRUMENT PROBLEM). NO ERRONEOUS RESULTS WERE GENERATED. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF GLOVES AND A LAB COAT AT THE TIME OF THE EVENT. NO INJURY OR EXPOSURE WAS REPORTED. A FIELD SERVICE ENGINEER WENT ON-SITE A DAY AFTER THE EVENT AND CONFIRMED THE LEAK. FSE REPLACED THE ASPIRATION PROBE AND THE PROBE WIPE WHICH RESOLVED THE LEAK. THE FATAL ERROR 7 WAS CAUSED BY THE FACT THAT THE PROBE SENSOR GOT CAUGHT ON THE SAMPLE TUBING AND POPPED OFF. FSE REPLACED THE PROBE SENSOR AND ADJUSTED THE SAMPLE LINE WHICH RESOLVED THE ERROR MESSAGE. PREVENTATIVE MAINTENANCE WAS PERFORMED TO OPTIMIZE THE ANALYZER. THE CAUSE OF THE LEAK IS RELATED TO THE ASPIRATION PROBE AND THE PROBE WIPE WHICH NEEDED TO BE REPLACED. (B)(4). | 0
2745782 | A CUSTOMER CONTACTED BECKMAN COULTER INC. (BEC) REPORTING A LEAK OF DILUENT AND BLOOD FROM THE MANUAL ASPIRATION PROBE ON THEIR COULTER HMX INSTRUMENT. THE VOLUME OF THE LEAK WAS LESS THAN 3CC. ALL OPERATORS WERE WEARING PERSONAL PROTECTIVE EQUIPMENT (PPE) CONSISTING OF A LAB COAT, GLOVES AND FACE PROTECTION WHEN THE LEAK WAS DISCOVERED. NO INJURY OR EXPOSURE WAS REPORTED AND NO PATIENT SAMPLES OR RESULTS WERE AFFECTED. A FIELD SERVICE ENGINEER (FSE) WENT ON-SITE A DAY AFTER THE EVENT AND DISCOVERED LOSS OF RETRACT ABILITY IN THE ASPIRATION PADDLE SPRING WHICH RESULTED IN SLOW RESPONSE TO RETURN TO HOME POSITION; DURING THIS TIME THE PROBE WIPE WOULD COME DOWN AND STRIKE THE PADDLE RESULTING IN FLUID LEAK FROM THE IMPACT. THE FSE REPLACED ASPIRATION PADDLE SPRING WHICH RESOLVED THE ISSUE. THE FSE THEN VERIFIED REPAIR PER ESTABLISHED PROCEDURES AND RESULTS MET ESTABLISHED SPECIFICATIONS. THE CAUSE OF THE LEAK WAS THE ASPIRATION PADDLE SPRING. (B)(4). | 0
3031248 | CUSTOMER REPORT THAT THE DEVICE WAS WITNESSED SMOKING FROM BACK. NO PT PRESENT. NO USER HARM. (B)(4). ADDITIONAL DATE / FAILURE INVESTIGATION. FIELD SVC TECH DISCOVERED DAMAGED CABLE FROM UNK CAUSE IN THE BACK OF THE DEVICE. NO CHARRING WAS FOUND. NO EVIDENCE OF A THERMAL EVENT. PART REPLACED AND TESTING OF THE DEVICE FOUND NO FURTHER ISSUES, FUNCTIONING NORMALLY. | 0
2815628 | THE USER FACILITY REPORTED THAT A HOSE CONNECTION LOOSENED AND WATER LEAKED THROUGHOUT THE DEPARTMENT AND INTO THE HALLWAY. NO INJURIES, PROCEDURAL DELAYS OR CANCELLATIONS OR PROPERTY DAMAGE WAS REPORTED. A STERIS SERVICE TECHNICIAN INSPECTED THE UNIT AND FOUND A LOOSE CLAMP AT THE CONNECTION. HE REPAIRED THE UNIT, RAN TEST CYCLES AND RETURNED THE UNIT TO SERVICE. NO FURTHER ISSUES HAVE BEEN REPORTED. | 0
2907590 | COMPLAINANT ALLEGED THAT DURING FUNCTIONAL TESTING, THE DEVICE PROMPTED A "DO NOT USE" MESSAGE. COMPLAINANT INDICATED THAT THERE WAS NO PATIENT INVOLVEMENT IN THE REPORTED MALFUNCTION. THE DEVICE HAS NOT BEEN RECEIVED FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION. | 0
3105855 | THERE WAS NO PATIENT INVOLVED IN THIS EVENT. THIS DEVICE MALFUNCTIONED BECAUSE IT WAS SWITCHING ITSELF ON AUTOMATICALLY. A DEVICE SWITCHING ITSELF ON AUTOMATICALLY, IF LEFT UNDETECTED COULD RESULT IN THE BATTERY BECOMING DEPLETED. ROUTINE TESTING CONFIRMED THAT THIS DEVICE WAS INSTALLED BY THE CUSTOMER IN (B)(6) 2008 AND PERFORMED TO SPECIFICATION UP TO (B)(6) 2010. THE DEVICE FAILED THE WEEKLY AUTO SELF-TEST DUE TO A LOW BATTERY ON (B)(6) 2010. A FRESH PAD-PAK WAS INSTALLED ON (B)(6) 2010 AND THE DEVICE PERFORMED TO SPECIFICATION UNTIL (B)(6) 2013. THERE IS NO EVIDENCE WITHIN THE HISTORY OF THE DEVICE OR DURING TESTING TO INDICATE THE DEVICE WAS SWITCHING ITSELF ON. INVESTIGATION DID NOT CONFIRM A FAULT WITH THE DEVICE OR ANY COMPONENT IN THE DEVICE. THE RETURNED PAD-PAK FROM LOT A867 WAS FOUND TO HAVE PERFORMED AS EXPECTED FOR 30 MONTHS BEFORE GIVING A LOW BATTERY WARNING. THE PAD-PAK, WHICH CONTAINS THE ELECTRODES AND BATTERIES, IS LABELED FOR SINGLE USE BUY THE SAMARITAN PAD 300 AND 300P DEVICES ARE FOR MULTI-USE. | 0
2911645 | THE ACCOUNT STATED THAT ONE PT SPECIMEN TESTED CELL-DYN 1800 HEMOGLOBIN = 8.3 G/DL BUT TESTED HEMOGLOBIN = 10.8 G/DL AT A REFERENCE LABORATORY. THE ACCOUNT SUSPECTED AN ASPIRATION ISSUE ON THE CELL-DYN 1800, AS THE SAMPLE HAD TO BE HELD UNDER THE PROBE FOR AN EXTENDED PERIOD BEFORE SAMPLE ASPIRATION BEGAN. ADDITIONALLY, THE ACCOUNT NOTICED SALT BUILD UP ON THE DILUENT SYRINGE. THE CONTROLS WERE RUNNING WITHIN RANGE AND THE MAINTENANCE WAS UP TO DATE ON THE CELL-DYN 1800. SERVICE WAS REQUESTED FOR THE CELL-DYN 1800. THE ACCOUNT DID NOT REPORT THE CELL-DYN 1800 HEMOGLOBIN RESULT OUTSIDE OF THE LABORATORY. NO IMPACT TO PT MANAGEMENT WAS REPORTED. (B)(4). RESULT, OTHER: SILICON TUBING (S1 - BUBBLE MIX, S2 - PINCH TUBING, S4 TUBING), RETAINER FOR PROBE, WASTE OUTLET TUBE ASSEMBLY (TUBING AND SENSOR WIRE). THE FIELD SERVICE REP (FSR) VISITED ON (B)(4) 2008 AND OBSERVED ISSUES WITH FLOW/CLOGS. SEVEN PARTS WERE REPLACED. THE S1 (SILICON) TUBING FOR THE RBC/PLT BUBBLE MIX WAS REPLACED AS CLOGGED/OBSTRUCTED (HARD FAILURE). S2 SILICON TUBING FOR THE FLOW PANEL PINCH TUBING WAS REPLACED AS WORN OUT (PREVENTIVE MAINTENANCE). S4 SILICON TUBING WAS CHANGED AS A PROACTIVE REPLACEMENT (PREVENTIVE MAINTENANCE). WORN OUT SYRINGES WERE REPLACED (100 UL SAMPLE SYRINGE-PREVENTIVE MAINTENANCE, LEAKING 10 ML DILUENT SYRINGE-HARD FAILURE) THE WASTE CABLE/TUBING WAS REPLACED AS ONE WIRE WAS BROKEN (HARD FAILURE). THE RETAINER FOR THE ASPIRATION PROBE WAS MISSING AND ALSO REPLACED (HARD FAILURE). PRECISION AND CONTROLS WERE RUN AND THE INSTRUMENT WAS PERFORMING WITHIN LABELING CLAIMS. THE CELL-DYN SYSTEM 1800 OPERATOR'S MANUAL, 07H80-01, REV D, PROVIDES A TROUBLESHOOTING GUIDE IN THE INTRODUCTION. A REVIEW OF THE COMPLAINT ANALYSIS TRENDING SYSTEM (CATS) AND TRENDING ANALYSIS SUPPORT SYSTEM (TASS) REPORTS, FOR THE PERIOD JUNE 2007 THROUGH JANUARY 2008, DID NOT INDICATE ANY ADVERSE TREND WITH THE CELL-DYN 1800, LIST BASE 07H77-01, FOR ISSUES WITH DISCREPANT RESULTS ON MULTIPLE PARAMETERS. THERE WAS NO SYSTEMIC ISSUE. THS IS A FINAL REPORT. END OF REPORT. | 0
3368450 | THE USER OF A SYSMEX XS-1000I ANALYZER, SERIAL NUMBER (B)(4), REPORTED TO (B)(6), THAT THE OPERATOR REPORTED AN INCORRECT, LOW HEMOGLOBIN RESULT FOR (B)(6) MALE WHO WAS ADMITTED TO THE HOSPITAL FOR A SURGERY OF "COLICISTITE B VIDEO" ON (B)(6) 2013. THE SURGERY WAS PERFORMED SUCCESSFULLY, AND THE PT WAS DISCHARGED. ON THIS DAY, HE FELT DIZZY AND FELL. AT THAT MOMENT, THE DOCTOR MADE THE REQUEST OF A COMPLETE BLOOD COUNT (CBC) THE RED BLOOD CELL (RBC) DATA WAS VERY LOW, INCLUDING A HEMOGLOBIN (HGB)=4 MG/DL. THESE LOW RESULTS WERE RELAYED BY PHONE TO THE SAMPLE BEING REPEATED. DUE TO THE CLINICAL PICTURE OF THE PT, THE DOCTOR DECIDED TO PERFORM A MICROSURGERY TO VERIFY IF THERE WAS SOME INTERNAL BLEEDING, WHICH WAS NOT APPARENT. AFTER THIS PROCEDURE, THE PT DIED AND AN AUTOPSY SHOWED DEATH DUE TO PULMONARY THROMBOEMBOLISM. THE SAMPLE WAS REPEATED AND THE RBC RESULTS WERE "NORMAL." NO DATA WAS SUPPLIED WITH THIS REPORT TO DETERMINE WHETHER FLAGS OR ERRORS WERE IDENTIFIED ON THE RESULTS. THE USER FACILITY IS (B)(6) AND IS SERVICED BY (B)(6). THE CLIENT DID NOT HAVE A E-CHECK COMMERCIAL CONTROLS TO ANALYZE ON THE DAY OF THE EVENT, BUT RAN A PT SAMPLE FROM THE PREVIOUS DAY, AND ALL WERE WITHIN THE ACCEPTABLE RANGE. THE DAY AFTER THE EVENT, E-CHECK CONTROLS WERE USED RESULTS WERE WITHIN RANGE. A (B)(6) FIELD SERVICE REPRESENTATIVE (FSR) VISITED THE SITE AND PERFORMED A PRECISION CHECK AND THE PRECISION WAS WITHIN ACCEPTABLE RANGE FOR ALL PARAMETERS. THE SAMPLE PATH FOR CONTROLS AND SAMPLES IS THE SAME. THE FSR SUSPECTED A POSSIBLE CLOG DURING ASPIRATION OF THIS SPECIFIC SAMPLE, AND DISCOVERED THE SAMPLE DETECTION SENSOR WAS TURNED OFF. THE MODE OF ANALYSIS (AUTOMATED OR MANUAL) WAS NOT PROVIDED. MANUAL ANALYSIS REQUIRES THE OPERATOR TO MIX AND ASSURE COMPLETE ASPIRATION OF THE SAMPLE. THE FSR CLEANED THE SAMPLE PATH FOR THE HGB CHANNEL, INCLUDING CLEANING/VERIFICATION OF TUBING, CHAMBERS AND VALVES, AND VERIFIED DILUTION. A FAILURE WAS NOT FOUND, AND THE ANALYZER WILL BE MONITORED BY THE USER. | 0
3289971 | POSSIBLE FALSE NEGATIVE. NO TRIGGERS PRESENT. ONE OBJECT OF INTEREST OUTSIDE CELL SPOT. INSTRUMENT OPERATIONAL. PERFORMED ALL REQUIRED SETUPS PER TECHNICAL DOCUMENTATION. CLEANED OPTIC PATH. RAN VERIFICATION SLIDE AND SHORT BATCH. | 0
3346786 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING FALSE POSITIVE STAINING OF TISSUE WITH TTR ANTIBODY (DAKO A002). THE RESULT WAS NEGATIVE WHEN STAINING OF THE SAME TISSUE SAMPLE WAS SUBSEQUENTLY PERFORMED MANUALLY. THE LAB ADVISED OF MISDIAGNOSIS FOR THE PATIENT FROM WHOM THE TISSUE SAMPLE WAS DERIVED; AND THAT FURTHER SPECIFIC INFO REQUIRED BY LEICA BIOSYSTEMS TO CONDUCT AN INVESTIGATION INTO THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT MAY NOT BE ABLE TO BE PROVIDED FOR SEVERAL WEEKS AS THE RESPONSIBLE STAFF MEMBER IS UNAVAILABLE. INVESTIGATION OF THIS COMPLAINT BY LEICA BIOSYSTEMS IS IN PROGRESS. | 0
2756907 | VENTANA MEDICAL SYSTEMS, INC. RECEIVED A CUSTOMER COMPLAINT REGARDING LABORATORY TECHNICIANS SMELLING BURNING WIRES AND SEEING SMOKE BEING EMITTED FROM A BENCHMARK ULTRA STAINER MODULE. FLUID FROM A SOURCE INTERNAL TO THE BENCHMARK ULTRA INSTRUMENT CONTACTED ELECTRICAL COMPONENTS RESULTING IN A SHORT CIRCUIT EVENT AND ENOUGH HEAT TO CAUSE SOME MELTING OF WIRING INSULATION CAUSING SMOKE TO BE EMITTED. THERE WERE NO SERIOUS INJURIES OR DEATH ASSOCIATED WITH THE INCIDENT REPORTED. THE INSTRUMENT (SN (B)(4)) HAS BEEN RETURNED FROM THE CUSTOMER SITE AND IS BEING EXAMINED IN DETAIL. THERE IS EVIDENCE OF A SHORT CIRCUIT OF MULTIPLE ELECTRICAL COMPONENTS DUE TO CONTACT WITH FLUIDS. THE ROOT CAUSE OF THE FLUID COMING IN CONTACT WITH THE ELECTRICAL COMPONENTS IS DUE TO A RESTRICTION IN THE WASTE DRAINAGE SYSTEM FROM A BUILD-UP OF DEBRIS.><THE ROOT CAUSE OF THE INCIDENT ORIGINALLY REPORTED UNDER 2028492-2012-00003 WAS DETERMINED TO BE A BUILD-UP OF DEBRIS IN THE WASTE FLUID DISCHARGE SYSTEM. A NOTIFICATION HAS BEEN SENT TO CUSTOMERS TO INFORM THEM OF THE POTENTIAL FOR THIS TYPE OF EVENT TO OCCUR, HOW TO RECOGNIZE IF THE WASTE DISCHARGE SYSTEM IS BECOMING RESTRICTED AND TO REMIND THEM TO PERFORM PERIODIC CLEANING OF THE SYSTEM. A CAPA HAS BEEN OPENED TO PURSUE ADDITIONAL CORRECTIVE ACTION(S). | 0
2969882 | REPORTER STATED THAT THE URISYS 1100 DEVICE HAS PRODUCED MULTIPLE NEGATIVE RESULTS FOR LEUKOCYTES, WHILE A VISUAL READING OF THE ASSOCIATED TEST STRIP SHOWS LEUKOCYTE RESULTS OF +1 AND +2. NO ACTIONS BASED ON THE DEVICE RESULTS WERE REPORTED AND NO ADVERSE EVENT OCCURRED. THE MANUFACTURER REQUESTED THE RETURN OF THE SUSPECT PRODUCT FOR EVALUATION. IT IS NOT KNOWN IF THE INITIAL REPORTER HAS OR INTENDS TO REPORT THE EVENT TO FDA. | 0
2883756 | REPORTER STATED THAT A PATIENT SAMPLE WAS TESTED ON THE SUSPECT URISYS 1100 WHICH PRODUCED A FALSE NEGATIVE RESULT FOR PROTEIN. AN ADDITIONAL TEST WAS REPORTEDLY PERFORMED, USING THE SAME STRIP LOT, ON A DIFFERENT URISYS INSTRUMENT WHICH REPORTED A PROTEIN RESULT OF 500 MG/DL. NO ADVERSE EVENT WAS REPORTED. THE MANUFACTURER REQUESTED THE RETURN OF THE SUSPECT PRODUCT FOR EVALUATION. IT IS NOT KNOWN IF THE INITIAL REPORTER HAS REPORTED OR INTENDS TO REPORT THE EVENT TO F.D.A. | 0
3251970 | ON (B)(6) 2013, BIOMERIEUX BECAME AWARE OF A PATIENT DEATH ASSOCIATED WITH A FALSE NEGATIVE BACT/ALERT BPA RESULT FOR A POOLED PLATELET SAMPLE. THE PATIENT HAD BEEN DIAGNOSED WITH PANCYTOPENIA, NEUTROPENIA AND LEUKEMIA, AND WAS UNDERGOING CHEMOTHERAPY FOR >1 MONTH. THE PATIENT RECEIVED A TRANSFUSION OF FOUR-DAY-OLD POOLED PLATELETS (B)(6) 2013 AT 22:00, AND WENT INTO SEPTIC SHOCK. TESTING OF THE PATIENT'S BLOOD DETECTED STREPTOCOCCI GROUP A. THE PATIENT DIED (B)(6) 2013 AT 12:50. INTERNAL INVESTIGATION AT (B)(6) ISOLATED THE CAUSE TO AN ASYMPTOMATIC BLOOD DONOR WHO HAD A RESPIRATORY INFECTION 1-2 WEEKS PRIOR TO BLOOD DONATION. THE DONOR HAD STREPTOCOCCI GROUP A. FURTHER TESTING OF ALL PRODUCTS PRODUCED WHICH CONTAIN THIS DONORS BLOOD WERE POSITIVE FOR STREPTOCOCCI GROUP A. THE CUSTOMER'S SAMPLE PROCESSING PROCEDURE INDICATES TO WAIT 24 HOURS BEFORE REMOVING THE SAMPLE FROM THE POOLED PLATELETS; HOWEVER, IN THIS CASE, THE SAMPLE OF 8-10ML FROM A FIVE-DONOR PLATELET POOL WAS REMOVED AT 18 HOURS. THE ASSOCIATED BACT/ALERT BPA BOTTLE CONTAINING THE POOLED PLATELETS REMAINED "NEGATIVE" FOR THE MAXIMUM MONITORING PERIOD OF 7 DAYS. AS INDICATED IN THE PACKAGE INSERT, A REPORT OF "NEGATIVE" SHOULD NOT BE INTERPRETED AS MEANING THAT THE ORIGINAL PRODUCT IS STERILE. THE NEGATIVE STATUS COULD BE DUE TO UNDER-INOCULATION OF THE BOTTLE, NO ORGANISMS PRESENT IN THE INOCULUM, THE NUMBER OF ORGANISMS WERE TOO SMALL FOR DETECTION, OR A CULTURE BOTTLE/MEDIUM THAT DOES NOT SUPPORT THE GROWTH OF THE ORGANISM. FOR BEST OVERALL RECOVERY WHEN CULTURING PLATELET SPECIMENS IT IS STRONGLY RECOMMENDED THAT MORE THAN ONE TYPE OF CULTURE BOTTLE BE UTILIZED (E.G., ONE AEROBIC AND ONE ANAEROBIC). THE PACKAGE INSERT ALSO STATES THAT MORE THAN ONE TYPE OF CULTURE BOTTLE SHOULD BE USED WHEN CULTURING PLATELET SPECIMENS (E.G. ONE AEROBIC, ONE ANAEROBIC CULTURE BOTTLE) FOR BEST OVERALL RECOVERY. IN THIS CASE, THE CUSTOMER ONLY USED AN AEROBIC (BPA) BOTTLE WHEN CULTURING THE PLATELET SPECIMENS. AS PART OF THE BIOMERIEUX INTERNAL INVESTIGATION, THE SUBMITTED BACT/ALERT INSTRUMENT DATABASE WAS REVIEWED AND THERE WERE NO INDICATIONS OF INSTRUMENT ERRORS OR OPERATOR ERRORS THAT COULD HAVE CONTRIBUTED TO THE EVENT. (B)(4). A REVIEW OF THE BATCH RECORD AND QC TESTING RESULTS FOR BACT/ALERT BPA CULTURE BOTTLE, LOT NUMBER 1032201, BY BIOMERIEUX, DID NOT REVEAL ANY ABNORMALITIES OR INDICATORS OF PRODUCT NON-CONFORMITY. THERE WERE NO DISCREPANCIES IN THE QC RELEASE DATA, AND THE RESULTS WERE WITHIN THE IDENTIFIED SPECIFICATIONS. THE RESULTS OF THE INVESTIGATION DID NOT DETECT A PRODUCT MALFUNCTION OR PRODUCT ISSUE ASSOCIATED WITH THE BACT/ALERT INSTRUMENT OR BACT/ALERT BPA CULTURE BOTTLE, LOT NUMBER 1032201 ASSOCIATED WITH THIS COMPLAINT. THE INSTRUCTIONS IN THE PACKAGE INSERT ARE ADEQUATE AND CORRECTLY INFORM THE CUSTOMER OF PRECAUTIONS TO REDUCE THE POTENTIAL OF AN ADVERSE EVENT. THIS EVENT IS BEING REPORTED AS AN ADVERSE EVENT BECAUSE THERE IS NO INFORMATION AVAILABLE TO CLEARLY SEPARATE THE BPA RESULT AND THE PATIENT DEATH. | 0
3403264 | PT EXPERIENCED AN UNWITNESSED FALL FROM BED RESULTING IN SUBDURAL HEMATOMA WITH DECREASED LEVEL OF CONSCIOUSNESS. PRODUCT REQUIRES FACILITY TESTING PRIOR TO USE; PRODUCT NOT DESIGNED TO BE FOLDED. PAD IN QUESTION TESTED AT HOSP BY ACCOUNT REP IN FRONT OF NURSING STAFF AND FOUND TO BE OPERATIONAL. PAD TEST DATE (DATE IN USE) WAS NOT FILLED OUT ON PAD. ACCOUNT REP: (B)(4). PADS SENT REP (B)(4) INTO THE FACILITY TO VERIFY PERFORMANCE ISSUES OF PAD. THE PAD IN QUESTION WAS OBSERVED AND TESTED WITH MONITOR AND FOUND TO WORK. REPORTER, (B)(6) (NURSES), ALL WITNESSED THE PADS CORRECT OPERATION WITH THE MONITOR. ADDITIONALLY, A PAD FROM THE SAME SHIPMENT TO (B)(6) OF THE SAME LOT AND SHIP DATE WAS TESTED AND FOUND TO BE OPERATIONAL WITH MONITOR, WITH NO COMPLAINT. ONE WOULD HAVE TO CONCLUDE THAT THE PAD WAS NOT CORRECTLY CONNECTED TO THE MONITOR. PADS REP WILL INITIATE THE REPLACEMENT OF ALL STANLEY TABS MONITOR WITH NEW PADS MONITORS, PROVIDE ADDITIONAL IN-SERVICE TRAINING ON OPERATION OF MONITOR/PAD CONNECTION, TESTING OF MONITOR BEFORE USE, AND OTHER INSTALLATION REQUIREMENTS AS LAID OUT BY PADS MONITOR AND MANUAL. | 1
3174303 | FACILITY REPORTED A PRODUCT PROBLEM WITH NO INJURIES RESULTING. REPORT STATED THAT THE TABS ALARM WAS FOUND TO HAVE A "FAULTY" CONNECTION FROM THE PAD TO THE MONITOR. RELATED TO MANUFACTURER REPORT NUMBER 1929691-2013-0003. THE FACILITY WAS FOUND TO HAVE A NUMBER OF MONITORS REQUIRING REPAIR OF MONITOR CONNECTORS, INCLUDING THIS REPORTED MONITOR. BECAUSE OF THE LARGE NUMBER OR MONITORS IN USE BY THE FACILITY, AND LACK OF RETURNS FOR MAINTENANCE, PRODUCT DISTRIBUTOR AND MANUFACTURER UNDERTOOK AN INSPECTION OF ALL TABS MONITORS AT THE FACILITY TO LOCATE ADDITIONAL MONITORS REQUIRING MAINTENANCE OR REPLACEMENT AS APPROPRIATE. AFTER REPEATED INSERTIONS AND REMOVALS OF A PAD AND/OR NURSE CALL CONNECTOR, THE PINS OF THE RJ-TYPE CONNECTORS OFTEN NEED MAINTENANCE AND/OR REPLACEMENT. FACILITY WAS INSTRUCTED AS TO PROPER TESTING BEFORE EACH USE, AS PRESCRIBED IN PRODUCT USER MANUALS. USER MANUALS ALSO INSTRUCT THAT ANY PROBLEMS NOTED DURING TESTING SHOULD TRIGGER A RETURN FOR INSPECTION AND MAINTENANCE. SIGNS OF PRODUCT REQUIRING MAINTENANCE ARE A MONITOR THAT WILL NOT "ARM" FOR USAGE AS IN THIS CASE, OR, A MONITOR THAT ALARMS WHEN THE CONNECTOR IS PULLED AFTER PROPER ARMING. | 1
3370545 | COULD NOT DEFLATE BALLOON. HAD TO CUT VALVE AND ALLOW WATER TO DRAIN OUT BEFORE REMOVAL. PT WAS IN PAIN. | 0
2892127 | GE REFERENCE NUMBER M1148290, LOT NUMBER 123004 HAD AN OBSTRUCTION DUE TO MANUFACTURING QUALITY CONTROL ISSUE. THE ANESTHESIOLOGIST'S VIGILANCE AND TROUBLESHOOTING AVERTED A NEGATIVE OUTCOME TO THE PATIENT. | 0
3226098 | ANESTHESIA ACUDOSE DEVICE SHOWS "FATAL ERROR" AND A RE-BOOT IS REQUIRED DURING A PROCEDURE. NO PATIENT HARM. THIS TYPE OF EVENT CREATES A DELAY OBTAINING MEDICATIONS FROM THE MEDICATION CABINET. AFTER RE-BOOTING, MEDICATION CAN BE OBTAINED FROM THE DEVICE. THERE ARE 12 DEVICES THAT HAVE BEEN AFFECTED. THE DEVICES WERE ORIGINALLY ON SOFTWARE VERSION 8.0.2. THIS FACILITY EXPERIENCED ISSUES WITH THE DEVICES ALMOST DAILY FOR APPROXIMATELY 2 WEEKS. THE MANUFACTURER THEN RELEASED A PATCH THAT SEEMED TO CORRECT THE ISSUE. HOWEVER, RECENTLY, THIS FACILITY HAS HAD SIMILAR ISSUES WITH THE PATCHED DEVICES FOR TWO CONSECUTIVE DAYS.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?NA.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #2IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #3IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #4IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #5IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #6IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #7IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #8IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #9IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #10IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #11IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #12IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3107964 | THIS IS A FOLLOW UP REPORT.><A CUSTOMER CONTACTED BECKMAN COULTER (BEC) STATING THAT WHILE A MEDICAL LABORATORY ASSISTANT WAS CLEANING A SERUM SPILLAGE IN THE AREA OF THE DRAWER OF THE AUTOMATE 2550 SAMPLE PROCESSING SYSTEM, SHE INJURED HER HAND. THE OPERATOR LIFTED THE BASE PLATE WITH ONE HAND WHILE CLEANING THE AREA WITH THE OTHER. UPON LOWERING THE PLATE BACK DOWN, IT FELL ONTO HER HAND, INJURING HER RIGHT INDEX FINGER. THE SPILLAGE SPREAD AS FAR AS UNDERNEATH THE BASE PLATE OF THE INPUT DRAWER WHICH HOLDS THE SAMPLE RACKS. THE LABORATORY ASSISTANT WENT TO THE ACCIDENT AND EMERGENCY (A&E) DEPARTMENT WHERE HER HAND WAS X-RAYED AND HER FINGER WAS STRAPPED, BEING DIAGNOSED AS BEING SOFT TISSUE DAMAGE. THE OPERATOR SAW HER GENERAL PRACTITIONER (GP) WHO SIGNED HER OFF FROM WORK FOR TWO WEEKS AS HER INJURIES WOULD PREVENT HER FROM PERFORMING THE PREDOMINANTLY MANUAL TASKS WITHIN THE LABORATORY ENVIRONMENT. ADDITIONAL INFORMATION: AS OF (B)(4) 2013, IT WAS CONFIRMED THAT THE SPLASH GUARD WAS REPLACED.><THE INJURY WAS CAUSED BY THE ABSENT SPLASH GUARD THAT WOULD HAVE PROTECTED HER AND IN TURN GAVE HER ACCESS TO THE BASE FRAME WITHOUT PULLING THE DRAWER OUT. IF THE SPLASH GUARD HAD BEEN IN PLACE, THE CUSTOMER WOULD HAVE NEEDED TO PULL THE DRAWER OUT TO ACCESS THE BASE FRAME AND THEREFORE, IT WOULD NOT HAVE BEEN ABLE TO FALL ON HER HAND. IT WAS DETERMINED THAT THE SPLASH GUARD WAS BROKEN BY THE CUSTOMER DURING NORMAL USE PRIOR TO THIS EVENT. THE CUSTOMER REMOVED THE REMAINING PART AND CONTINUED TO USE THE INSTRUMENT. THE CUSTOMER DID NOT REPORT THE DAMAGED SPLASH GUARD TO BEC. THE OPERATOR WAS NOT FOLLOWING CLEANING PROCESS OF THE USER GUIDE AND INJURED HERSELF BECAUSE SHE TRIED TO BY PASS THE CLEANING OPERATION. INSTEAD OF REMOVING THE INPUT BASEPLATE TO FACILITATE CLEANING, SHE SUPPORTED IT WITH ONE HAND WHILST CLEANING UNDERNEATH WITH THE OTHER. | 0
2756699 | CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT THEY RECEIVED A PRINTOUT FROM THE DXC INSTRUMENT FOR A SAMPLE ID THAT CONTAINED INCORRECT PATIENT DEMOGRAPHICS COMPARED TO THE TUBE BARCODE. CUSTOMER REPORTED THAT NO INCORRECT RESULTS WERE REPORTED OUT OF THE LABORATORY. CUSTOMER REPORTED THEY MANUALLY PROGRAMMED AND RERAN THE SAMPLE. CUSTOMER REPORTED THE REPEATED RESULTS WERE REPORTED OUT OF THE LABORATORY. BEC INVESTIGATION INDICATED THE DATALINK/DL2000 DATA MANAGER COMMAND CENTRAL (DL 2000) DOWNLOADED THE OLD PROGRAM TO THE DXC. WHEN THE LABORATORY INFORMATION SYSTEM DOWNLOADED THE NEW PROGRAM FOR THE SAMPLE ID TO THE DL2000, THE DL2000 ALREADY HAD PROGRAMMING FOR THAT SAMPLE ID FROM 2010 AND DOWNLOADED THE EXISTING 2010 PROGRAMMING TO THE DXC. BEC INVESTIGATION INDICATED THE DL2000 WAS FAILING TO DELETE OLD PATIENT DEMOGRAPHICS. BEC REPLACED THE DATABASE. BEC TESTED THE NEW DATABASE AND CONFIRMED THAT PATIENT DEMOGRAPHICS ARE NOW DELETING AS REQUESTED. CUSTOMER REPORTED THEY REUSE SAMPLE IDS. (B)(4). | 1
3221232 | WHILE PERFORMING ORAL CARE, A 2 TO 2.5 INCH PIECE OF THE SUBGLOTTAL SUCTION TUBING BROKE OFF AND WAS LODGED IN THE PATIENT'S THROAT. RN WAS UNABLE TO VISUALIZE THE BROKEN PIECE TO REMOVE THE OBJECT. UPON INSPECTION, THE SUCTION CATHETER APPEARED TO BE PERFORATED WHERE IT BROKE OFF. PT WAS INTUBATED AT THE TIME. THE SURGICAL ICU RESIDENT WAS CALLED TO THE BEDSIDE AND THEY WERE UNSUCCESSFUL AT FINDING THE OBJECT. CHEST XRAY AND KUB WERE ORDERED AND NO FOREIGN OBJECTS WERE SEEN ON THESE FILMS. ALL SUBGLOTTAL SUCTION TUBINGS FROM KIT WERE REMOVED AND DISCARDED. NO APPARENT PATIENT HARM. THE PIECE MISSING FROM THE END OF THE TUBING WAS UNABLE TO BE LOCATED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ORAL CARE FOR AN INTUBATED PATIENTDEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO | 0
3354293 | CUSTOMER REPORTED WHILE A PATIENT WAS BEING MONITORED BY THE HYPERVISOR VI CENTRAL MONITORING SYSTEM WITH AN AMBULATORY TRANSMITTER, THE CENTRAL STATION DETECTED A VENTRICULAR FIBRILLATION EVENT, BUT NO ECG WAVEFORM WAS DISPLAYED. THE PATIENT EXPIRED A DAY AFTER THE REPORTED EVENT. THE EVENT IS CURRENTLY UNDER INVESTIGATION. | 1
2887194 | THE CUSTOMER REPORTED THE CLINICAL INFORMATION CENTER (CIC) PRO MP100 WAS NOT AUDIBLY ALARMING, THOUGH VISUAL PARAMETERS AND MESSAGES WERE AVAILABLE ON THE DISPLAY. THE CUSTOMER STATED THE ISSUE WAS NOT IMMEDIATELY NOTICED BY THE MEDICAL STAFF. FOLLOWING A REBOOT OF THE CIC, THE AUDIBLE ALARMS BEGAN WORKING. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2954073 | THE PT HAD AN INTRA CRANIAL PRESSURE (ICP) DISCONNECT ALARM. THE NURSE PAUSED THE ALARM WITHOUT CHECKING ON THE PT FOR 3 MINS. THE CUSTOMER STATED AFTER THE 3 MIN ALARM PAUSE, THE ALARM DID NOT SOUND AND THE PT BLED AND REQUIRED A TRANSFUSION OF 2 UNITS OF BLOOD. THE HOSPITAL BIOMED WAS CONTACTED POST-INCIDENT AND HE STATED THAT THE ICE DISCONNECT INOP WAS GENERATED BY THE INVOLVED MONITOR. THE NURSE HAD PAUSED THE ALARM AND LEFT THE ROOM. A VISITOR OF THE PT NOTIFIED STAFF THAT THERE WAS BLOOD ON THE PT. THE PT WAS ATTENDED TO AND RECOVERED FROM THE LOSS OF BLOOD. THE HOSPITAL BIOMED TESTED THE INVOLVED MONITOR POST-INCIDENT AND COULD NOT DUPLICATE THE PROBLEM WITH THE ALARMS. THE MONITOR LABELING STATES THAT WHEN A DISCONNECT ALARM IS PAUSED, THE MEASUREMENT IN QUESTION IS SWITCHED OFF SINCE THE COMMON EXPECTATION IS FOR THE PROVIDER TO CORRECT THE SITUATION AND RESTART THE PAUSED ALARM. THIS INFO WILL BE COMMUNICATED TO THE CUSTOMER VIA A WRITTEN LETTER. THE INVOLVED DEVICE REMAINS IN USE AT THE CUSTOMER SITE AND IS CONSIDERED FULLY FUNCTIONAL. | 1
3023612 | IT WAS REPORTED THAT THE PHYSICIAN TRIED TO USE THE VNS PROGRAMMING COMPUTER BUT WAS UNABLE TO DUE TO THE COMPUTER GIVING HER A WARNING MESSAGE THAT THE BATTERY LATCH IS OPEN. TROUBLE SHOOTING WAS PERFORMED TO MAKE SURE THAT THE BATTERY LATCH IS SECURED BUT THE SAME WARNING MESSAGE APPEARS. THE PRODUCT WAS SENT BACK TO THE MANUFACTURER AND WAS RECEIVED ON (B)(4) 2013 AND NOW PRODUCT ANALYSIS IS BEING PERFORMED.><PRODUCT ANALYSIS WAS PERFORMED ON THE RETURNED PROGRAMMING COMPUTER. DURING ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WOULD POWER OFF UNEXPECTEDLY. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A LOOSE MOUNTING SCREW ON THE BATTERY LATCH ASSEMBLY. ONCE THE SCREW WAS TIGHTENED, NO FURTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3026448 | ON (B)(4) 2013 THE PROGRAMMING COMPUTER WAS RECEIVED BY THE MANUFACTURER FOR PRODUCT ANALYSIS. AFTER PRODUCT ANALYSIS WAS PERFORMED ON THE VNS PROGRAMMING COMPUTER IT WAS FOUND THAT THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH DEBRIS. ONCE THE DISPLAY WAS CLEANED, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE VNS PROGRAMMING COMPUTER STOPPED WORKING AND THE TOUCH SCREEN IS NOT WORKING. TROUBLE SHOOTING HAVE BEEN PERFORMED BUT WERE UNSUCCESSFUL. PRODUCT ANALYSIS ON THE VNS COMPUTER WILL BE PERFORMED AFTER THE PRODUCT ARRIVES TO THE MANUFACTURER. DEVICE MALFUNCTION IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2897169 | WHEN WE IMPLEMENTED MAG VIEW AS OUR NEW MAMMOGRAPHY REPORTING SYSTEM BACK ON FEBRUARY 06, 2012; THE SYSTEM HAD A GAIL RISK ASSESSMENT TOOL THAT WAS ALREADY A PART OF THE SOFTWARE. THIS RISK ASSESSMENT TOOL ALLOWS THE RADIOLOGISTS TO DETERMINE WHAT PERCENTAGE OF RISK BASED ON FAMILY, AND PERSONAL HISTORY OF A PATIENT. THIS RISK PERCENTAGE IS USED TO DETERMINE IF A PATIENT IS ELIGIBLE FOR MRI AS A HIGH RISK SCREENING TOOL.WE ADDED A SOFTWARE CALLED TYRER CUSZIK TO MAG VIEW THAT WOULD ALLOW US TO CAPTURE 100 PERCENT OF OUR SCREENING PATIENT POPULATION.ONCE WE IMPLEMENTED THIS MODEL IN ADDITION TO THE GAIL MODEL IN MAG VIEW, WE STARTED USING THIS DATA TO IDENTIFY ANY PATIENTS THAT NEED FOLLOW UP CALLS FROM HER FOR PREVENTATIVE TREATMENT.ONCE WE STARTED RUNNING REPORTS WE NOTICED THAT SOME OF THE DATA WAS NOT MAKING SENSE WHEN COMPARING PATIENTS' HISTORY WITH RISK FACTOR PERCENTAGES GENERATED BY THE GAIL SOFTWARE, SO WE SPOKE TO THE RADIOLOGIST WHO AGREED AND MAG VIEW WHO THEN STATED THAT THE ISSUE WAS A USER ISSUE NOT A SOFTWARE ISSUE. WE THEN PULLED 20 RANDOM CASES AND COMPARED THEM TO NCI GAIL VS. MAG VIEW GAIL VS. BRACA GAIL ONLY TO FIND THAT MAG VIEW DID NOT MATCH EITHER NCI OR BRACA.WHEN WE POINTED THIS OUT TO MAG VIEW WE WERE TOLD THAT WE ARE NOT READING THE NUMBERS CORRECTLY, OR WE DID NOT ENTER THE DATA CORRECTLY. WE THEN ASKED MAG VIEW WHAT VERSION OF GAIL ARE THEY USING NCI OR BRACA, AND THEY HAVE YET TO ANSWER THAT QUESTION. | 1
3267872 | WHEN THE ROSA DEVICE IS USED IN STEREOTACTIC SURGERY, THE "DISTANCE TO TARGET" FUNCTION THAT IS PROVIDED IN "COOPERATIVE MODE" DURING THE GUIDANCE PHASE CAN GIVE THE USER AN ERRONEOUS VALUE UNDER CERTAIN CONDITIONS. WHEN THE USER REQUESTS THIS DISTANCE TO TARGET INFORMATION VIA THE SOFTWARE INTERFACE, THE VALUE IS DISPLAYED CORRECTLY. HOWEVER, IF THE DISTANCE TO TARGET WINDOW REMAINS OPEN AND THE ROBOTIC ARM IS MOVED, THE VALUE IS NOT AUTOMATICALLY REFRESHED IN THE CORRESPONDING WINDOW THUS THE CORRECT DISTANCE TO TARGET IS NOT PROVIDED. THIS PROBLEM IS PRESENT WITH THE VERSIONS 2.5.0, 2.5.3 AND 2.5.4 OF THE ROSA SOFTWARE. THIS POTENTIAL ISSUE WAS IDENTIFIED BY OUR ENGINEERING TEAM WHILE PERFORMING INTERNAL TESTING ON THE DEVICE AND UNDER VERY SPECIFIC CONDITIONS. A FIELD SAFETY NOTICE WAS ISSUED AND TRANSMITTED TO ALL OF OUR CUSTOMERS. THE FIELD SAFETY NOTICE CONTAINS SPECIFIC RECOMMENDATIONS FOR OUR USERS TO PREVENT ANY POTENTIAL INJURIES TO A PATIENT. | 0
3129032 | IN THE MIDDLE OF THE CARDIAC CATHETERIZATION THE MAC-LAB COMPUTER SYSTEM SHUT DOWN.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?CARDIAC CATHETERIZATION. | 1
3428969 | PATIENT UNDERGOING DIAGNOSTIC EP STUDY WITH ST JUDE EQUIPMENT WHERE WE ARE THE 1ST HOSPITAL TO USE THIS EQUIPMENT IN US, RECENTLY FDA APPROVED. AFTER 75% OF THE CASE HAD BEEN COMPLETED, THE INTRACARDIAC IMAGES WENT DOWN. THE ST JUDE REP DID NOT KNOW CAUSE AND RECOMMENDED THAT THE PROCEDURE BE ABORTED. THE PATIENT WAS NOT HARMED. WE HAVE HAD FOLLOW UP MEETINGS WITH VENDOR AND THEY ARE WORKING ON IDENTIFYING ROOT CAUSE: IT MAY BE RELATED TO DISLODGING OF RIGHT LEG LEAD BUT WHY WOULD IT TAKE OUT ALL THE INTRACARDIAC IMAGES. NO WARNING OR EXPLANATION-IT APPEARED AMPLIFIER WAS DEAD. WE DO NOT PLAN TO CONTINUE TO USE OF THIS 1ST IN USE SYSTEM UNTIL WE HAVE A BETTER UNDERSTANDING OF ROOT CAUSE. | 1
2946773 | IT WAS REPORTED THAT THE SVO2 JUMPS AROUND WITHOUT A STABLE VALUE. THE CABLE AND PHILLIPS MODULE WERE CHANGED OUT WITHOUT SUCCESS. ONCE THE VIGILANCE II WAS CHANGED OUT THE PROBLEM WAS RESOLVED. NO PATIENT INJURY WAS REPORTED. | 1
3002862 | IT WAS REPORTED THAT THE CVP ON THE EV1000 DID NOT MATCH THE BEDSIDE CVP BY 8-10 POINTS. THE EV1000 UNIT WAS CHANGED OUT AND THE CVP ON EV1000 MATCHED THE BEDSIDE. ALL CABLES USED WERE KNOWN WORKING CABLES. THE CVP WAS ZEROED THE EXACTLY THE SAME ON BOTH EV1000 UNITS. NO OTHER FAULT MESSAGES WERE SEEN. NO OTHER PROBLEMS EXPERIENCED WITH DEFECTIVE EV1000 ASIDE FROM THE CVP NOT COINCIDING WITH BEDSIDE. NO PATIENT INJURY WAS REPORTED. EXAMINATION OF THE RETURNED EV1000 WAS UNABLE TO DETECT ANY FAILURE OF THE DEVICE TO FUNCTION AS INTENDED. THE EV1000 WAS TESTED USING A PLUM SIMULATOR. THE DEVICE APPROPRIATELY ZEROED CVP/AP AND MONITORED CO, SVI, AND SCVO2 FOR 24 HOURS. DURING THE 24 HOURS BURN-IN TESTING, RESULTS WERE COLLECTED AND CONFIRMED TO BE WITHIN EXPECTED RANGE. "BT/IT RANGE WAS ALSO TESTED AND FOUND TO BE WITHIN PRODUCT SPECIFICATIONS AND THE DEVICE PASSED THE EXTERNAL CVP TEST." NO OTHER DEFECTS OR DAMAGE WAS OBSERVED OR NOTED ON THE DATABOX OR PC PANEL DURING THE VISUAL INSPECTION. THE COMPLAINT COULD NOT BE CONFIRMED DURING THE EVALUATION; THE FAILURE COULD NOT BE REPLICATED AND THE NO DEFICIENCIES WERE NOTED ON THE UNIT." IT COULD NOT BE DETERMINED IF SOME CLINICAL FACTOR MAY HAVE CONTRIBUTED TO THE EVENT REPORTED." NO ACTIONS WILL BE TAKEN AT THIS TIME. | 1
2914671 | IT WAS REPORTED THAT DURING VT ABLATION PROCEDURE, THE PHYSICIAN PACED THE PATIENT THRU THE CARTO SYSTEM AND SUDDENLY THE BODY SURFACE ECG SIGNALS AND IC EXCEPT FOR THE RVA SIGNAL WENT FLAT LINE. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT'S CONSEQUENCES. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION OF THE EVENT, ON (B)(6) 2012 IT WAS CONFIRMED THAT THE SIGNALS FROM THE CARTO MONITOR AND EP SYSTEM DISAPPEARED AS SOON AS THE PHYSICIAN STIMULATED THE PATIENT VIA CARTO SYSTEM. WHEN THE PHYSICIAN STIMULATED THE PATIENT DIRECTLY VIA BIOTRONIK, THE BS ECG'S AND IC SIGNALS WERE VISIBLE ON THE MONITOR. DESPITE THE LOSS OF SIGNALS WERE FIXED BY USING A NEW CABLE, THE LOST OF SIGNALS OCCURRED ON BOTH SYSTEMS (CARTO AND EP RECORDING) SIMULTANEOUSLY MAKING THIS EVENT REPORTABLE DATING ALERT DATE AS (B)(6) 2012. INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE PRODUCT ANALYSIS INVESTIGATION IS COMPLETED. (B)(4).><(B)(4). IT WAS REPORTED THAT DURING VT ABLATION PROCEDURE, THE PHYSICIAN PACED THE PATIENT THRU THE CARTO SYSTEM AND SUDDENLY THE BODY SURFACE ECG SIGNALS AND IC EXCEPT FOR THE RVA SIGNAL WENT FLAT LINE. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT'S CONSEQUENCES. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION OF THE EVENT, ON (B)(4) 2012 IT WAS CONFIRMED THAT THE SIGNALS FROM THE CARTO MONITOR AND EP SYSTEM DISAPPEARED AS SOON AS THE PHYSICIAN STIMULATED THE PATIENT VIA CARTO SYSTEM. WHEN THE PHYSICIAN STIMULATED THE PATIENT DIRECTLY VIA BIOTRONIK, THE BS ECG'S AND IC SIGNALS WERE VISIBLE ON THE MONITOR. DESPITE THE LOSS OF SIGNALS WERE FIXED BY USING A NEW CABLE, THE LOST OF SIGNALS OCCURRED ON BOTH SYSTEMS (CARTO AND EP RECORDING) SIMULTANEOUSLY MAKING THIS EVENT REPORTABLE DATING ALERT DATE AS (B)(6) 2012. UPON RECEIPT, THE DEVICE WAS VISUALLY INSPECTED AND IT WAS OBSERVED THERE WAS DAMAGE ON P2 END CENTER PIN. THE CABLE FAILED ELECTRICAL TEST DUE TO LEAKAGE WAS IDENTIFIED. FURTHER INVESTIGATION REVEALED THAT THE PHYSICAL DAMAGE TO THE CONNECTOR CONTRIBUTED TO THE ELECTRICAL FAILURE DUE TO POTENTIAL MOISTURE BUILT UP. THE DEVICE HISTORY RECORD WAS REVIEWED AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE DHR REVIEW VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. THE CUSTOMER COMPLAINT HAS BEEN VERIFIED. | 1
3060135 | IT WAS REPORTED THAT DURING AN EP STUDY AND C3 MAPPING PROCEDURE THE PATIENT SUFFERED A LOCALIZED ALLERGIC CONTACT REACTION TO THE C3 EXTERNAL REFERENCE PATCHES. IT IS UNKNOWN IF AN ABLATION WAS PERFORMED. THE CUSTOMER REQUESTS A FULL LIST OF INGREDIENTS FOR THE ADHESIVE C3 PATCHES. THE IFU DID NOT PROVIDE THE INFORMATION THAT THE CALLER IS REQUESTING. ADDITIONAL INFORMATION WAS REQUESTED ON THE EVENT, BUT NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. PRODUCT INVESTIGATION WILL NOT BE PERFORMED SINCE THE CATHETER WAS NOT RETURNED FOR ANALYSIS. SINCE NO LOT NUMBER WAS PROVIDED, NO DEVICE HISTORY RECORD REVIEW COULD BE PERFORMED. CONCOMITANT PRODUCT: CARTO 3 SYSTEM, MODEL #: M-4800-0,1 SERIAL #: (B)(4). | 0
2898210 | IT WAS REPORTED DURING AN IDIOPATHIC VENTRICULAR TACHYCARDIA PROCEDURE THERE WAS A MAP SHIFT OF 38MM WITHOUT ANY ERROR OR ALERT. THE CASE WAS COMPLETED AND THERE WAS NO PATIENT INJURY REPORTED. INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4).><ADDITIONAL INFORMATION WAS RECEIVED FROM THE BWI FIELD REPRESENTATIVE ON THE EVENT. AFTER A PLANNED CARDIOVERSION, THE MAP AND CATHETER POSITION WERE NOT ALIGNED. THE REASON FOR THE CARDIOVERSION WAS A VT WHICH COULD NOT BE OVER STIMULATED. THE ACL FUNCTION WAS IN USE DURING THE CASE. THE NAVISTAR CATHETER AND THE IBI DECAPOLAR ACL CATHETER MOVED RELATIVE TO THE MAP. BOTH THE ACL CATHETER AND THE MAGNETIC CATHETER SHIFTED IN THE SAME DIRECTION AND IT WAS THE SAME AMOUNT OF SHIFT. THE IBI DECAPOLAR WAS CONNECTED TO THE REF/DECA IN THE PATIENT INTERFACE UNIT. THE REFERENCE PATCH DID NOT MOVE OR GET LOOSE BEFORE THE MAP SHIFT. THE CASE WAS COMPLETED CONVENTIONALLY. EVALUATION SUMMARY. (B)(4). IT WAS REPORTED DURING AN IDIOPATHIC VENTRICULAR TACHYCARDIA PROCEDURE THERE WAS A MAP SHIFT OF 38MM WITHOUT ANY ERROR OR ALERT. THE CASE WAS COMPLETED AND THERE WAS NO PATIENT INJURY REPORTED. THE ISSUE WAS NOT DUPLICATED; HOWEVER, ACCORDING TO ADDITIONAL INFORMATION RECEIVED FROM THE BWI FIELD REPRESENTATIVE, THE MAP SHIFT OCCURRED AFTER CARDIOVERSION. MAP SHIFTS AFTER CARDIOVERSION OCCURS BECAUSE THE PATIENT'S HEART CAN PHYSICALLY MOVE WITH RESPECT TO THE INITIAL PATCH REFERENCE POSITIONS. CARTO IS NOT INTENDED TO FIX SUCH RELATIVE MOTION. IFU STATES: "IN RARE CASES, THE RELATIVE POSITION BETWEEN THE PATCHES REMAINS THE SAME BUT THE POSITION OF THE HEART RELATIVE TO THE BACK PATCHES HAS CHANGED (FOR EXAMPLE, AFTER REPOSITIONING THE PATIENT'S HEAD ON A PILLOW OR MOVING THE PATIENT'S ARM WITHOUT CHANGING THE PATIENT'S POSITION ON THE MATTRESS, OR AFTER THE PATIENT IS CARDIOVERTED). DUE TO SUCH INTERNAL ANATOMIC DISPLACEMENTS, THERE MIGHT BE MISALIGNMENT OF THE MAP TO THE HEART LOCATION, "THE SYSTEM WILL NOT GIVE A WARNING AND AN INCORRECT MAP (MAP SHIFT) MIGHT BE GENERATED. "WHEN IN DOUBT, CLOSE THE CURRENT MAP AND BEGIN A NEW ONE. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. | 1
2808689 | IT WAS REPORTED THAT IN THE MIDDLE OF LEFT VENTRICULAR TACHYCARDIA (WITHOUT ARTERY ACCESS) CASE, THE PHYSICIAN REALIZED A BIG DISTORTION WITH THE EP RECORDER SYSTEM. AFTER FOLLOW UP WITH THE CUSTOMER TO KNOW DETAILED INFORMATION OF THE EVENT, ON (B)(6, 2012, ADDITIONAL INFORMATION WAS RECEIVED. THE REPORTER STATED THAT THE CARTO3 REQUESTED THE ECG REPOSITION BECAUSE IT WAS IN THE WRONG PLACE. THE ECG AND THE PATCH WERE OK BUT IT WAS NOT POSSIBLE TO SEE ANY CATHETER IN THE SYSTEM. THE PIU WAS REBOOTED AND AFTER RESTART IT, THE PROBLEM CONTINUED AND A RED LIGHT IN THE BACK OF THE PIU (MAGNETIC LOCATION CARD) AND A MESSAGE IN THE WORKSTATION WITH: PUI TASTED FAILED WAS SHOWED. THE SYSTEM WAS REBOOTED 3 TIMES WITH THE SAME RESULT. THE CASE WAS CANCELLED DUE TO THIS ISSUE WITHOUT ANY PROBLEM TO THE PATIENT. TRANSSEPTAL PUNCTURE WAS PERFORMED AT THE TIME OF THE CASE CANCELLATION AND EXTENDED HOSPITALIZATION WAS NEEDED MAKING THIS EVENT REPORTABLE. (B)(4). IT WAS REPORTED THAT IN THE MIDDLE OF LEFT VENTRICULAR TACHYCARDIA (WITHOUT ARTERY ACCESS) CASE, THE PHYSICIAN REALIZED A BIG DISTORTION WITH THE EP RECORDER SYSTEM. THE SYSTEM WAS REBOOTED 3 TIMES WITH NO SUCCESS. THE CASE WAS CANCELLED DUE TO THIS ISSUE WITHOUT ANY PROBLEM TO THE PATIENT. TRANSSEPTAL PUNCTURE WAS PERFORMED AT THE TIME OF THE CASE CANCELLATION AND EXTENDED HOSPITALIZATION WAS NEEDED MAKING THIS EVENT REPORTABLE. THE SYSTEM WAS EVALUATED AND THE REPORTED FAILURE COULD NOT BE REPRODUCED. PREVENTIVE MAINTENANCE AND FULL SYSTEM ACCEPTANCE TEST PROCEDURES (ATP) WERE PERFORMED AND PASSED SUCCESSFULLY. THE COMPLAINT WAS NOT CONFIRMED. THE DHR ASSOCIATED WITH CARTO 3 # 11633 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON IT'S RELEASE.><INVESTIGATION IS STILL IN PROGRESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR INVESTIGATION IS COMPLETED. (B)(4). | 1
2878724 | IT WAS REPORTED DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE THERE WAS A SHIFT WITHOUT PATIENT MOVEMENT. UPON REQUEST, THE BWI FIELD REPRESENTATIVE PROVIDED ADDITIONAL INFORMATION ON THE EVENT. THE PATIENT WAS SEDATED. THERE WAS NO INFORMATION WITH REGARD TO HOW MUCH THE MAP SHIFTED. THERE WAS NO ERROR MESSAGE. THE MAP SHIFT DID NOT OCCUR AFTER THE USER MOVED THE HEAD OF FLUOROSCOPY. THE ACL FUNCTION WAS ON DURING THE CASE. THE PHYSICIAN DID NOT PERFORM A CARDIOVERSION PRIOR TO THE MAP SHIFT. THE REFERENCE PATCH DID NOT MOVE BEFORE THE MAP SHIFT. NO INFORMATION IS AVAILABLE REGARDING WHETHER THE MAP SHIFT OCCURRED DURING RF ABLATION. IT IS ALSO UNKNOWN WHICH CATHETER SHIFTED NOR IN WHICH DIRECTION. THE PROBLEM WAS NOT SOLVED BUT NO CATHETER OR CABLE WAS CHANGED. (B)(4). IT WAS REPORTED DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE THERE WAS A SHIFT WITHOUT PATIENT MOVEMENT. UPON REQUEST, THE BWI FIELD REPRESENTATIVE PROVIDED ADDITIONAL INFORMATION ON THE EVENT. THE PATIENT WAS SEDATED. THERE WAS NO INFORMATION WITH REGARD TO HOW MUCH THE MAP SHIFTED. THERE WAS NO ERROR MESSAGE. THE MAP SHIFT DID NOT OCCUR AFTER THE USER MOVED THE HEAD OF FLUOROSCOPY. THE ACL FUNCTION WAS ON DURING THE CASE. THE PHYSICIAN DID NOT PERFORM A CARDIOVERSION PRIOR TO THE MAP SHIFT. THE REFERENCE PATCH DID NOT MOVE BEFORE THE MAP SHIFT. NO INFORMATION IS AVAILABLE REGARDING WHETHER THE MAP SHIFT OCCURRED DURING RF ABLATION. IT IS ALSO UNKNOWN WHICH CATHETER SHIFTED NOR IN WHICH DIRECTION. THIS PROBLEM WAS NOT RESOLVED. NO CATHETER OR CABLE WAS CHANGED. THE CUSTOMER CALLED AFTER THE CASE WHEN SHE WAS OUTSIDE OF THE HOSPITAL. THE CUSTOMER WAS ADVISED TO OBSERVE FURTHER CASES AND PAY ATTENTION TO PATIENT SETUP AND TO PROVIDE THE METAL VALUES AND CALI VALUES. THE ISSUE WAS NOT DUPLICATED DURING FURTHER PROCEDURES. THE COMPLAINT WAS NOT CONFIRMED. NO CONCLUSION CAN BE DRAWN BASED ON THE AVAILABLE INFORMATION. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE.><INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4). | 1
3317863 | IT WAS REPORTED THAT IN THE MIDDLE OF A PAROXYSMAL A-FIB PROCEDURE, THE CARTO 3 SYSTEM BEGAN RE-INITIALIZING ON ITS OWN AND THERE WERE BLUE LIGHTS NEXT TO ECG CARD 1 AND MAG LOC CARD. THIS ISSUE OCCURRED TWICE. THE FIRST TIME, IT WAS FULLY INITIALIZED ON ITS OWN AND CASE WAS CONTINUED. DURING THE SECOND OCCURRENCE THE CUSTOMER WAS ADVISED TO REBOOTING THE SYSTEM, DISCONNECTING THE CATHETERS AND BS ECG CABLE. BWI REPRESENTATIVE, REBOOTED AND RECONNECTED THE BS ECG CABLE ONLY TO CARTO 3 SYSTEM, ALLOWED TO INITIALIZE AND ERRORS 911 AND 1201 APPEARED. THE CATHETERS WERE RECONNECTED AND THE ERROR 1201 WAS RESOLVED UNTIL CS PACING WAS ATTEMPTED. THE ERROR RETURNED AND ALL BS AND IC SIGNALS ON BOTH RECORDING SYSTEM AND C3 BECAME EXTREMELY NOISY. THE MICROPACE STIMULATOR WAS REBOOTED. THE LIMB LEAD CONNECTIONS AT PATIENT AND ECG LEADS INTO ECG BLOCK WERE CHECKED AND ALL WERE SECURE. THE REPORTER WAS TRANSFERRED TO BWI FIELD ENGINEER WHO ADVISED TO DISCONNECTING MAP CATHETER WHILE PACING BUT ERROR 1201 AND NOISE CONTINUED. THE CUSTOMER DISCONNECTED THE LASSO CATHETER AND ERROR 1201 PERSISTED DURING CS PACING. BY DISCONNECTING SOUNDSTAR CATHETER THE ERROR 1201 WAS RESOLVED AND ALL SIGNALS IMPROVED. SHORTLY AFTER, THE ERROR 1201 REAPPEARED INTERMITTENTLY AND THE SIGNALS REMAINED CLEAR. ON-CALL BWI FIELD ENGINEER CONTINUED TO TROUBLESHOOT WITH BWI REPRESENTATIVE IN THE ACCOUNT TO DETERMINE THE SOURCE OF NOISE. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN CLARIFICATION OF THE EVENT, ON (B)(6) 2013, IT WAS CONFIRMED THAT IT WAS 60 CYCLE NOISE ON LASSO AND BS CHANNEL ONLY WHEN PACING WAS PERFORMED, IT OCCURRED WHEN LASSO WAS MOVED FROM LSPV (LEFT SUPERIOR PULMONARY VEIN) TO RSPV (RIGHT SUPERIOR PULMONARY VEIN). THE PHYSICIAN WAS ABLE TO INTERPRET THE BOTH BS AND ALL IC RECORDINGS ONLY WHEN NOT PACING WAS PERFORMED. (B)(4). IT WAS REPORTED THAT IN THE MIDDLE OF A PAROXYSMAL A-FIB PROCEDURE, THE CARTO 3 SYSTEM BEGAN RE-INITIALIZING ON ITS OWN AND THERE WERE BLUE LIGHTS NEXT TO ECG CARD 1 AND MAG LOC CARD. THIS ISSUE OCCURRED TWICE. THE FIRST TIME IT WAS FULLY INITIALIZED ON ITS OWN AND CASE WAS CONTINUED. DURING THE SECOND OCCURRENCE THE CUSTOMER WAS ADVISED TO REBOOTING THE SYSTEM, DISCONNECTING THE CATHETERS AND BS ECG CABLE. BWI REPRESENTATIVE, REBOOTED AND RECONNECTED THE BS ECG CABLE ONLY TO CARTO 3 SYSTEM, ALLOWED TO INITIALIZE AND ERRORS 911 AND 1201 APPEARED. THE CATHETERS WERE RECONNECTED AND THE ERROR 1201 WAS RESOLVED UNTIL CS PACING WAS ATTEMPTED. THE ERROR RETURNED AND ALL BS AND IC SIGNALS ON BOTH RECORDING SYSTEM AND C3 BECAME EXTREMELY NOISY. THE MICROPACE STIMULATOR WAS REBOOTED. THE LIMB LEAD CONNECTIONS AT PATIENT AND ECG LEADS INTO ECG BLOCK WERE CHECKED AND ALL WERE SECURE. THE REPORTER WAS TRANSFERRED TO BWI FIELD ENGINEER WHO ADVISED TO DISCONNECTING MAP CATHETER WHILE PACING BUT ERROR 1201 AND NOISE CONTINUED. THE CUSTOMER DISCONNECTED THE LASSO CATHETER AND ERROR 1201 PERSISTED DURING CS PACING. BY DISCONNECTING SOUNDSTAR CATHETER THE ERROR 1201 WAS RESOLVED AND ALL SIGNALS IMPROVED. SHORTLY AFTER, THE ERROR 1201 REAPPEARED INTERMITTENTLY AND THE SIGNALS REMAINED CLEAR. ON-CALL BWI FIELD ENGINEER CONTINUED TO TROUBLESHOOT WITH BWI REPRESENTATIVE IN THE ACCOUNT TO DETERMINE THE SOURCE OF NOISE. FIELD ENGINEER RECOMMENDED TO THE BWI REPRESENTATIVE TO PERFORM A TROUBLESHOOTING IN THE CARTO 3 SYSTEM. THE BWI REPRESENTATIVE CHECKED ALL THE LIMB LEADS AND REPLACED ELECTRODES ON THE PATIENT, DISCONNECTED ALL SENSOR BASED CATHETERS, DISCONNECTED ECG OUT CABLE BUT THE INTERMITTENT ERROR 1201 - ECG CONNECTION ERROR PERSISTED. THE FIELD ENGINEER SUGGESTED REPLACING THE BS - ECG CABLE. AFTER ECG CABLE REPLACEMENT, THE SYSTEM WAS FOUND FUNCTIONAL AND READY FOR USE. THE DHR ASSOCIATED WITH CARTO 3 RMT SYSTEM # (B)(4) WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE.><INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4) WHEN THE LASSO WAS UNPLUGGED AND A DUO DECA POLAR CATHETER WAS PLUGGED INTO THE LASSO PORT (20 POLE A), THE NOISE AND PACING ISSUES WERE RESOLVED. THE DUO DECA WAS A SJM LIVEWIRE (COMPETITORS CATHETER) THAT WAS CONNECTED THROUGH A CARTO 3 PIN BOX INTO THE 20 POLE A. DUE TO THE FACT THAT SEVERE NOISE OCCURRED DURING PACING, THE EVENT BECOMES REPORTABLE. AWARENESS DATE WAS CHANGED TO (B)(6) 2013. | 1
2914681 | IT WAS REPORTED THAT DURING A PAROXYSMAL A-FIB PROCEDURE, A MAP SHIFT OCCURRED. THE SHIFT WAS APPROXIMATELY 0.5 CM. CARTO 3 SYSTEM DIDN'T GIVE ANY WARNING MESSAGE. THE CASE COULD BE FINISHED WITHOUT ANY PROBLEMS. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION REGARDING THE EVENT, IT WAS CONFIRMED THAT NO WARNING WAS SHOWN IN THE SYSTEM BECAUSE OF THE MAP SHIFT. THE ACL FUNCTION WAS TURNED ON DURING THE CASE. THE ACL CATHETER AND MAGNETIC CATHETER SHIFT IN THE SAME DIRECTION AND SAME AMOUNT. THERE WERE NO CARDIOVERSION PERFORMED BY THE PHYSICIAN AND NO REFERENCE PATCH MOVEMENT WAS OBSERVED BEFORE THE MAP SHIFT. TO IDENTIFY THE CATHETER LOCATION SHIFTS, A SJM ESOPHAGUS CATHETER AND A CS CATHETER WERE USED. THE SJM CS CATHETER WAS CONNECTED TO 10P. INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4).><(B)(4). IT WAS REPORTED THAT DURING A PAROXYSMAL A-FIB PROCEDURE, A MAP SHIFT OCCURRED. THE SHIFT WAS APPROXIMATELY 0.5 CM. CARTO 3 SYSTEM DIDN'T GIVE ANY WARNING MESSAGE. THE CASE COULD BE FINISHED WITHOUT ANY PROBLEMS. THE ISSUE WAS IDENTIFIED AS A WRONG WORKFLOW. AFTER THE CUSTOMER WAS INSTRUCTED HOW TO PLACE PROPERLY THE PATCH AND USE THE SID VALUES RECOMMENDED, THERE WAS NO REOCCURRENCE OF THE MAP SHIFT SINCE THEN. THE DHR ASSOCIATED WITH CARTO 3 (B)(4) WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. | 1
3260173 | DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, IT WAS REPORTED THERE WAS AN UNIT ERROR DURING INITIALIZATION. ERROR NUMBERS 258,268, 278 WERE CONSISTENTLY SHOWN. THE FOLLOWING TROUBLE SHOOTING STEPS WERE TAKEN TO RESOLVE THIS ISSUE: PIU WAS REBOOTED SEVERAL TIMES; HOWEVER, ERRORS 258,268,278 WERE STILL SHOWN; WORK STATION WAS REBOOTED SEVERAL TIMES; HOWEVER, ERRORS 258,268,278 WERE STILL SHOWN; SWITCH THE POWER SUPPLY TO THE PIU WITH WORK STATION OFF AND ON AGAIN; SOFTWARE RE-INSTALLATION AFTER WORKSTATION DATA FULL-FORMAT ON (B)(6) 2013. ERRORS 258,268,278 WERE STILL SHOWN. ADDITIONAL INFORMATION RECEIVED STATED THERE WAS NO PATIENT CONSEQUENCE. THE PATIENT WAS UNDER GENERAL ANESTHESIA FOR 3 HOURS. THE CASE WAS CANCELED DUE TO THIS ISSUE. PHYSICIAN DID NOT CONSIDER DELAYING OF THE PROCEDURE CAUSED A POTENTIAL RISK TO THE PATIENT. PATIENT DID NOT REQUIRE EXTENDED HOSPITALIZATION. TRANSSEPTAL PUNCTURE WAS PERFORMED PRIOR TO CASE CANCELATION. DUE TO THIS, THIS IS NOW AN MDR REPORTABLE EVENT. (B)(4).><(B)(4) ADDITIONAL INFORMATION RECEIVED STATED THERE WAS NO PATIENT CONSEQUENCE. THE PATIENT WAS UNDER GENERAL ANESTHESIA FOR 3 HOURS. THE CASE WAS CANCELED DUE TO THIS ISSUE. PHYSICIAN DID NOT CONSIDER DELAYING OF THE PROCEDURE CAUSED A POTENTIAL RISK TO THE PATIENT. PATIENT DID NOT REQUIRE EXTENDED HOSPITALIZATION. TRANSSEPTAL PUNCTURE WAS PERFORMED PRIOR TO CASE CANCELATION. DUE TO THIS, THIS IS NOW A MDR REPORTABLE EVENT. AFTER UNIT EVALUATION IT WAS DETERMINED THE MAIN CARD OF THE PIU WAS DEFECTIVE AND IT WAS REPLACED BY THE FIELD SERVICE ENGINEER. AFTER REPAIR, UNIT WAS FUNCTIONING PER SPECIFICATION. THE DEFECTIVE MAIN MODULE WAS SENT TO CARTO3 MANUFACTURER FOR FURTHER EVALUATION. PER MANUFACTURER¿S EVALUATION, IT WAS REPORTED ERRORS 258, 268, 278 WERE VERIFIED. AFTER NEW ALTERA&DSP BURNING (MAIN MODULE IMBEDDED SOFTWARE BURNING), ISSUE WAS RESOLVED. THE CUSTOMER¿S COMPLAINT WAS CONFIRMED. DHR REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. | 1
3246205 | IN REQUESTING ADDITIONAL INFORMATION FROM THE BWI LEGAL DEPARTMENT FOR COMPLAINT (PI1-DOXOVP ¿ REPORT # 3008203003-2013-00069), ON (B)(4) 2013, THE COMPLAINTS MANAGEMENT DEPARTMENT RECEIVED INFORMATION FOR ANOTHER DIFFERENT EVENT. AFTER VERIFICATION, WE WERE UNABLE TO CONFIRM THAT THERE IS A PRE-EXISTING COMPLAINT FOR THIS EVENT. THEREFORE, PI1-FCI9DT FOR THIS EVENT WAS CREATED ON (B)(6) 2013. ON OR BEFORE (B)(6) 2010, THE PATIENT WAS EXPERIENCING ATRIAL FIBRILLATIONS. ON (B)(6) 2010, THE PATIENT UNDERWENT AN ABLATION PROCEDURE AT (B)(6). AT OR NEAR THE CONCLUSION OF THE SURGICAL ABLATION PROCEDURE PERFORMED, THE PATIENT¿S BLOOD PRESSURE DROPPED AND CREATED A CONDITION OF HYPOTENSION. AFTER OVER ONE HOUR, THE PHYSICIAN SOUGHT INTERVENTION OF OTHER PHYSICIANS WHO THEN PERFORMED A PERICARDIOCENTESIS AND DETERMINED THE PRESENCE OF BLOOD WITHIN THE PERICARDIUM, CREATING A TAMPONADE WHICH CAUSED THE PATIENT TO GO INTO CARDIAC ARREST. THEREAFTER, THE PATIENT UNDERWENT OPEN HEART SURGERY. UPON PERFORMING SURGERY, IT WAS DISCOVERED A PERFORATION OF THE RIGHT ATRIUM. THE PATIENT SUFFERED FROM A REDUCED FLOW OF BLOOD AND OXYGEN TO THE BRAIN, THEREBY PROXIMATELY CAUSING A CEREBRAL INJURY ALL OF WHICH HAS RENDERED HER UNABLE TO PHYSICALLY CARE FOR HERSELF AND RESULTED IN HER SUBSEQUENT ADMISSION TO A NURSING HOME FOR LONG-TERM AND PERMANENT CARE. ADDITIONAL CLARIFICATION WAS REQUESTED, BUT NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. MANUFACTURER'S REF. NO: (B)(4).><(B)(4). ON OR BEFORE (B)(6) 2010, THE PATIENT WAS EXPERIENCING ATRIAL FIBRILLATIONS. ON (B)(6) 2010 THE PATIENT UNDERWENT AN ABLATION PROCEDURE AT (B)(6). AT OR NEAR THE CONCLUSION OF THE SURGICAL ABLATION PROCEDURE PERFORMED, THE PATIENT¿S BLOOD PRESSURE DROPPED AND CREATED A CONDITION OF HYPOTENSION. AFTER OVER ONE HOUR, THE PHYSICIAN SOUGHT INTERVENTION OF OTHER PHYSICIANS WHO THEN PERFORMED A PERICARDIOCENTESIS AND DETERMINED THE PRESENCE OF BLOOD WITHIN THE PERICARDIUM, CREATING A TAMPONADE WHICH CAUSED THE PATIENT TO GO INTO CARDIAC ARREST. THEREAFTER, THE PATIENT UNDERWENT OPEN HEART SURGERY. UPON PERFORMING SURGERY, IT WAS DISCOVERED A PERFORATION OF THE RIGHT ATRIUM. THE PATIENT SUFFERED FROM A REDUCED FLOW OF BLOOD AND OXYGEN TO THE BRAIN, THEREBY PROXIMATELY CAUSING A CEREBRAL INJURY ALL OF WHICH HAS RENDERED HER UNABLE TO PHYSICALLY CARE FOR HERSELF AND RESULTED IN HER SUBSEQUENT ADMISSION TO A NURSING HOME FOR LONG-TERM AND PERMANENT CARE. THE EVENT OCCURRED IN THE PAST. AFTER REVIEWING THE HISTORY OF THE CUSTOMER COMPLAINTS ASSOCIATED WITH THIS SPECIFIC SYSTEM, IT WAS FOUND THAT ONLY TWO MONTHS AFTER THIS EVENT, A PREVENTATIVE MAINTENANCE WAS PERFORMED. THE BWI FIELD SERVICE ENGINEER HAD IMAGED THE DELL HARD DRIVE AND RELOADED THE MERGE MODULE. A FULL SYSTEM PREVENTATIVE MAINTENANCE WAS ALSO PERFORMED. THE SYSTEM WAS TESTED AND READY FOR USE. NO FAILURE WAS FOUND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. | 1
2875989 | IT WAS REPORTED DURING AN ATRIAL FLUTTER (AFL) PROCEDURE THE CARTO 3 SYSTEM WAS EXPERIENCING INTERMITTENT NOISE ACROSS ALL CHANNELS INCLUDING THE 12-LEAD. THE INDIFFERENT ELECTRODE WAS MOVED FROM THE BACK TO THE LEG WITH THE WIRE REMAINING ISOLATED FROM CARTO 3 CABLING. THE NAVISTAR CABLE WAS CHANGED. NEITHER OF THESE EFFORTS RESULTED IN NOISE REDUCTION. THE 12-LEAD PATCHES WERE CHANGED AND THE NOISE DID NOT REPEAT. THE RIGHT LEG LEAD PATCH MAY NOT HAVE BEEN FULLY ADHERED. THE CASE WAS COMPLETED WITH NO FURTHER NOISE ISSUES. UPON REQUEST, ADDITIONAL INFORMATION WAS PROVIDED ON THE EVENT BY THE BWI FIELD REPRESENTATIVE. THE NOISE STARTED OCCURRING HALF WAY THROUGH THE CASE. THERE WAS NOISE ON THE BODY SURFACE (BS) ECG'S AND INTRACARDIAC (IC) SIGNALS ON BOTH THE CARTO 3 SYSTEM AND THE EP RECORDING SYSTEM. THE PHYSICIAN WAS NOT ABLE TO INTERPRET THE BS ECGS AND ALL IC RECORDINGS BECAUSE OF THE NOISE. THERE WAS HIGH FREQUENCY NOISE WHERE ALL EGMS ON BOTH THE BODY SURFACE AND IC DISAPPEARED. IT WAS THOUGHT THAT THE NOISE WAS ONLY DURING ABLATION, BUT IT WAS OCCURRING AT ALL TIMES. WE WERE ABLE TO RESOLVE THE NOISE ISSUE SO THE CASE WAS COMPLETED SUCCESSFULLY. (B)(4). IT WAS REPORTED DURING AN ATRIAL FLUTTER (AFL) PROCEDURE THERE WAS NOISE ON THE BODY SURFACE (BS) ECG'S AND INTRACARDIAC (IC) SIGNALS ON BOTH THE CARTO 3 SYSTEM AND THE EP RECORDING SYSTEM. THE PHYSICIAN WAS NOT ABLE TO INTERPRET THE BS ECGS AND ALL IC RECORDINGS BECAUSE OF THE NOISE. THE 12-LEAD PATCHES WERE CHANGED AND THE NOISE DID NOT REPEAT. THE RIGHT LEG LEAD PATCH MAY NOT HAVE BEEN FULLY ADHERED. THE CASE WAS COMPLETED WITH NO FURTHER NOISE ISSUES. AN ADDITIONAL OEM MANUFACTURER DEVICE HISTORY RECORD (DHR) WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE.><INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4). | 0
3143204 | IT WAS REPORTED BY AN ATTORNEY THAT HE IS INVOLVED IN A LAWSUIT REGARDING AN ABLATION PROCEDURE. THE ALLEGATION IS THAT HIS CLIENTS ESOPHAGUS WAS PERFORATED WHICH SUBSEQUENTLY LED TO HIS CLIENT¿S DEATH. ADDITIONAL INFORMATION ON THE EVENT WAS REQUESTED, BUT NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. EVALUATION SUMMARY. (B)(4). IT WAS REPORTED BY AN ATTORNEY THAT HE IS INVOLVED IN A LAWSUIT REGARDING AN ABLATION PROCEDURE. THE ALLEGATION IS THAT HIS CLIENTS ESOPHAGUS WAS PERFORATED WHICH SUBSEQUENTLY LED TO HIS CLIENT¿S DEATH. HIS UNDERSTANDING IS THAT THERE WAS A 3D MAPPING OF THE PROCEDURE IN WHICH IT WOULD PROVIDE LOCATIONS AND AMOUNT OF ELECTRICITY APPLIED. THE HOSPITAL IS TELLING HIM THAT THEY CAN NOT RETRIEVE THE INFORMATION. THE ATTORNEY IS TRYING TO EDUCATE HIMSELF ON HOW THAT MIGHT BE RETRIEVED. THE USER DID NOT BACKUP THE PATIENT FILES AS IT IS RECOMMENDED IN THE USER MANUAL. IN ADDITION, THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH THIS SPECIFIC SYSTEM WAS REVIEWED AND THERE WAS NOT ANY COMPLAINT THAT MAY BE RELATED TO THE REPORTED ISSUE. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE.><INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4).><ADDITIONAL INFORMATION WAS PROVIDED ON THE EVENT ON OCTOBER 9, 2013. IT WAS THOUGHT THAT THERE WAS AN ESOPHAGUS PROBE/TEMPERATURE MONITOR USED WITH THE CARTO XP SYSTEM IN THIS PROCEDURE. INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4).><INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. THE BWI LEGAL DEPARTMENT PROVIDED ADDITIONAL INFORMATION ON THE EVENT ON (B)(4) 2013. THE PATIENT UNDERWENT AN ABLATION PROCEDURE FOR HIS ATRIAL FIBRILLATION ON (B)(6) 2010, AND DEVELOPED ESOPHAGEAL FISTULA AND INFECTION. THE PATIENT SOUGHT TREATMENT ON (B)(6) 2010. DESPITE TREATMENT, THE INFECTION SPREAD THROUGHOUT THE PATIENT¿S BODY AND HE DIED OF FAILURE OF HIS ORGANS ON (B)(6) 2010. THE PHYSICIAN WHO PERFORMED THE PROCEDURE ON THIS PATIENT WAS ALLEGED PERFORATING THE PATIENT¿S PERICARDIUM AND THE WALL OF ESOPHAGUS DURING THE ABLATION PROCEDURE. THE ALLEGATIONS TO THE PHYSICIAN WERE FAILING TO ADEQUATELY MONITOR THE PLACEMENT OF ELECTRODES DURING ABLATION, ADMINISTERING EXCESSIVE ABLATION AND USING TOO MUCH POWER TO PATIENT¿S HEART IN THE AREA NEAR THE ESOPHAGUS, AND FAILING TO PERFORM TESTING AFTER THE PROCEDURE TO DETECT THE PERFORATION OF PERICARDIUM AND ESOPHAGUS. THE BWI LEGAL DEPARTMENT ALSO PROVIDED INFORMATION STATING IT WAS NOT THOUGHT THAT THERE WAS AN ISSUE WITH THE CARTO XP SYSTEM. (B)(4). | 1
2975732 | IT WAS REPORTED THAT DURING A PAROXYSMAL A-FIB PROCEDURE, IT WAS NOT POSSIBLE TO PACE FROM CARTO 3 SYSTEM. THE CASE WAS FINISHED USING PACING FROM THE EP RECORDING SYSTEM. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION OF THE EVENT IT WAS STATED THAT IT WAS A PLANNED PACING FOR ASSESSING CONDUCTION BLOCK. THE PHYSICIAN TRIED TO PACE FROM ALL THE DIAGNOSTIC CATHETERS CONNECTED TO THE PIU (ALSO IN EMERGENCY MODE). BY CHANGING THE PACING CABLE AND REBOOTING THE WORKSTATION DID NOT SOLVE THE PROBLEM. BY DISCONNECTING ALL THE CATHETERS AND CABLES FROM THE PIU AND RECONNECTING THEM TO THE EP RECORDING SYSTEM IT WORKED FINE. THE MICROPACE STIMULATOR DIDN'T SHOW ANYMORE THE ERROR MASSAGE OF HIGH IMPEDANCE. THE PHYSICIAN DIDN'T CONSIDER THAT THERE WAS ANY POTENTIAL RISK TO THE PATIENT FROM THIS PACING ISSUE. DUE TO THE FACT THAT THE PHYSICIAN WAS NOT ABLE TO USE THE EMERGENCY PACING PORT FROM THE PIU, THIS EVENT BECOMES REPORTABLE. INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4).><(B)(4). IT WAS REPORTED THAT DURING A PAROXYSMAL A-FIB PROCEDURE, IT WAS NOT POSSIBLE TO PACE FROM CARTO 3 SYSTEM. THE CASE WAS FINISHED USING PACING FROM THE EP RECORDING SYSTEM. THE CARTO 3 SYSTEM AND THE EP RECORDING SYSTEM WERE CHECKED BY THE FSE (FIELD SERVICE ENGINEER) AND IT WAS NOTICED THAT THERE WAS A SETTING ISSUE ON THE BARD SYSTEM. THE PACING CIRCUIT WAS NOT ENABLED. BY PERFORMING THE CORRECT SETTINGS, THE PACING WORKS PROPERLY. THE ISSUE WAS NOT CARTO 3 SYSTEM RELATED. IT WAS FOUND OPERATIONAL. THE DHR ASSOCIATED WITH CARTO 3 (B)(4) WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. | 1
3273090 | IT WAS REPORTED THAT DURING AN ATRIAL FLUTTER RIGHT (R-AFL) PROCEDURE THE SIGNALS DISAPPEARED. TROUBLESHOOTING HAD BEEN DONE BY REPLACING VARIOUS PARTS WITH NO SUCCESS. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION IT WAS CONFIRMED THAT THE SIGNAL LOSS OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS ECGS AND ALL IC (INTRACARDIAL) RECORDINGS ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME. THE PHYSICIAN WAS NOT ABLE TO INTERPRET THE SIGNALS WITH THE SIGNAL LOSS ISSUE. THE CASE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4).><(B)(4) IT WAS REPORTED THAT DURING AN ATRIAL FLUTTER RIGHT (R-AFL) PROCEDURE THE SIGNALS DISAPPEARED. TROUBLESHOOTING HAD BEEN DONE BY REPLACING VARIOUS PARTS WITH NO SUCCESS. THE PHYSICIAN WAS NOT ABLE TO INTERPRET THE SIGNALS WITH THE SIGNAL LOSS ISSUE. THE CASE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. BWI REPRESENTATIVE STATED THAT WHEN THE PHYSICIAN STARTED TO ABLATE, THEY FIRST SAW A BIG SPIKE FOLLOWED FOR A LARGE NOISE BY FLAT LINES ON ALL IC SIGNALS AND THE BS ECG AND THE CARTO 3 SYSTEM WAS SHOWING ERROR 8. THE BWI ENGINEER ARRIVED ON SITE FOR TROUBLESHOOTING AND IT WAS POSSIBLE TO REPRODUCE THE NOISE AND ERROR 8. BY CHANGING THE BACKPLANE, IT WAS POSSIBLE TO SOLVE THE NOISE ISSUE BUT AFTER ABOUT 20 MIN ERROR 8 APPEARED AGAIN. THE ENGINEER REPLACED THE ECG BOARDS AND PERFORMED THE FULL SYSTEM TEST WITHOUT ANY FURTHER ISSUE. AFTER THIS, THE SYSTEM WAS FOUND OPERATIONAL. THE REPLACED PARTS WERE SENT TO THE MANUFACTURER FOR FURTHER INVESTIGATION AND THEY CONFIRMED THAT THE ISSUE WAS CAUSED BY THE FAULTY BACKPLANE CARD. DIODE TVS D1000 WAS FOUND FAILED AT ISOLATION TEST. THE ECG CARDS WERE FOUND OPERATIONAL. THE DHR ASSOCIATED WITH CARTO 3 # R10063 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. | 1
3297849 | IT WAS REPORTED THAT DURING AN ATRIOVENTRICULAR NODAL REENTRANT TACHYCARDIA (AVNRT) PROCEDURE, WHILE THE PHYSICIAN WAS APPLYING RF ENERGY, THE CATHETER TIP WENT FROM RED TIP TO GREEN WHILE ABLATING. THE REPORTER WAS ADVISED TO CHANGE OUT THE CATHETER CABLE AND TO RELOAD THE CARTO APPLICATION. THE BWI REPRESENTATIVE CALLED BACK AND STATED THAT THE PHYSICIAN DIDN¿T PERFORM ANOTHER RF APPLICATION SINCE THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. THE PHYSICIAN PERFORMED A SECOND PROCEDURE AFTER REBOOTING THE WORK STATION AND THE ISSUE DIDN¿T REOCCUR. (B)(4). IT WAS REPORTED THAT DURING AN ATRIOVENTRICULAR NODAL REENTRANT TACHYCARDIA (AVNRT) PROCEDURE, WHILE THE PHYSICIAN WAS APPLYING RF ENERGY, THE CATHETER TIP WENT FROM RED TIP TO GREEN WHILE ABLATING. THE REPORTER WAS ADVISED TO CHANGE OUT THE CATHETER CABLE AND TO RELOAD THE CARTO APPLICATION. THE BWI REPRESENTATIVE CALLED BACK AND STATED THAT THE PHYSICIAN DIDN¿T PERFORM ANOTHER RF APPLICATION SINCE THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. THE PHYSICIAN PERFORMED A SECOND PROCEDURE AFTER REBOOTING THE WORK STATION AND THE ISSUE DIDN¿T REOCCUR. THE SYSTEM WAS CHECKED BY THE MANUFACTURER (HTC) AND THEY TRIED TO REPRODUCE THE ISSUE WITH NO SUCCESS, IN ADDITION, SIMILAR ISSUE WAS NEVER FOUND IN CARTO 3 SYSTEM TESTING OR KNOWN FROM THE FIELD. THIS WAS A SINGLE EVENT THAT DID NOT REOCCUR AFTER REBOOTING THE WORKSTATION AS IT WAS PREVIOUSLY MENTIONED. COMPLAINT CONDITION WAS NOT CONFIRMED. THE DHR ASSOCIATED WITH CARTO 3 # 11677 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE.><INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4). | 1
3197169 | IT WAS REPORTED THAT DURING AN ATRIAL FLUTTER (AFL) PROCEDURE, ALL THE BODY SURFACE (BS) AND INTRACARDIAC (IC) ECG FLATLINED ON THE CARTO 3 SYSTEM AND THE BARD RECORDING SYSTEM. THE ECG SIGNALS RETURNED APPROXIMATELY 5 SECONDS LATER. THE ECG ON THE ANESTHESIA CART WAS NOT AFFECTED AND CONTINUED TO SHOW NORMAL RHYTHMS. PACING WAS ATTEMPTED DURING THE 5 SECONDS BUT THE USER WAS UNABLE TO TELL IF PACING ENERGY WAS DELIVERED. PACING WAS SHOWN TO BE DELIVERED NORMALLY WHEN THE ECG CHANNELS DISPLAYED AGAIN. NO PATIENT INJURY WAS REPORTED. NO FURTHER DETAILS WERE AVAILABLE AT THE TIME OF THE CALL. THE SIGNAL LOSS ON ALL THE BODY SURFACE (BS) AND INTRACARDIAC (IC) ECG ON THE CARTO 3 SYSTEM AND THE BARD RECORDING SYSTEM AT THE SAME TIME, IS INDICATIVE OF A REPORTABLE EVENT. UPON REQUEST, THE BWI FIELD REPRESENTATIVE PROVIDED ADDITIONAL INFORMATION ON THE EVENT. THE SIGNAL LOSS WAS ONLY ABOUT 5 SECONDS. THE SIGNALS CAME BACK AND THEN THE CASE PROCEEDED. PACING WAS FOR ASSESSING THE CONDUCTION BLOCK. THEY HAD TRIED A COUPLE OF DIFFERENT SITES FOR PACING AND THEN IT WAS PACING/CAPTURING. THIS OCCURRED DURING THAT 5 SECOND LOSS OF SIGNALS. THE PATIENT WAS IN A STABLE NORMAL SINUS RHYTHM DURING THIS TIME WHICH WAS DOCUMENTED BY THE MONITORING RN'S SIGNALS AND THE ANESTHESIOLOGIST MONITORING SCREEN. THE PHYSICIAN WAS NOT TRYING TO PACE AND ABLATE FORM THE SAME CATHETER SIMULTANEOUSLY. THE STIMULATOR BEING USED WAS THE BARD MICROPACE. THE PHYSICIAN DID NOT THINK THAT THERE WAS ANY POTENTIAL RISK TO THE PATIENT FROM THIS PACING ISSUE. THE RISK TO THE PATIENT IS LOW FROM THIS PACING ISSUE THEREFORE, THIS ISSUE IS NOT REPORTABLE. (B)(4). IT WAS REPORTED THAT DURING AN ATRIAL FLUTTER (AFL) PROCEDURE, ALL THE BODY SURFACE (BS) AND INTRACARDIAC (IC) ECG FLATLINED ON THE CARTO 3 SYSTEM AND THE BARD RECORDING SYSTEM. THE ECG SIGNALS RETURNED APPROXIMATELY 5 SECONDS LATER. THE ECG ON THE ANESTHESIA CART WAS NOT AFFECTED AND CONTINUED TO SHOW NORMAL RHYTHMS. PACING WAS ATTEMPTED DURING THE 5 SECONDS BUT THE USER WAS UNABLE TO TELL IF PACING ENERGY WAS DELIVERED. PACING WAS SHOWN TO BE DELIVERED NORMALLY WHEN THE ECG CHANNELS DISPLAYED AGAIN. NO PATIENT INJURY WAS REPORTED. NO FURTHER DETAILS WERE AVAILABLE AT THE TIME OF THE CALL. THE PROBLEM DID NOT OCCUR IN THE NEXT CASES AND WAS NOT DUPLICATED. A BAD RL CONNECTION COULD CAUSE THE REPORTED PROBLEM. THE SIGNAL LOSS WAS ONLY APPROXIMATELY 5 SECONDS. THE SIGNALS CAME BACK AND THE PHYSICIANS PROCEEDED WITH THE ABLATION. THEY HAD TRIED A COUPLE OF DIFFERENT SITES AND THEN IT WAS PACING/CAPTURING. THIS WAS DURING THAT 5 SECOND LOSS OF SIGNALS. THE PATIENT WAS IN A STABLE NSR DURING THIS TIME WHICH WAS DOCUMENTED BY THE MONITORING RN'S SIGNALS AND THE ANESTHESIOLOGIST MONITORING SCREEN. COMPLAINT CONDITION WAS NOT CONFIRMED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE.><INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4). | 1
3160589 | IT WAS REPORTED THAT DURING AN ISCHEMIC VENTRICULAR TACHYCARDIA (ISVT) PROCEDURE, THE CARTO 3 SYSTEM LOST 12-LEAD ECG¿S SIGNALS FOR APPROXIMATELY 3-4 SECONDS AFTER CARDIOVERSION WAS PERFORMED IN THE PATIENT. IT WAS ALSO REPORTED THAT THE 12-LEAD AND INTRACARDIAC (IC) SIGNALS WERE FLAT WHEN THE PHYSICIAN WAS PACING FROM THE RV CATHETER CONNECTED TO REF/DECA. A BLUE PIN BOX WAS IN USE WITH REF/DECA. MOVING THE RV CATHETER DIRECTLY TO THE RECORDING SYSTEM RESOLVED THE PACING ISSUE. IT WAS RECOMMENDED TO TRY A DIFFERENT PIN BOX OR TO MOVE THE RV CATHETER TO PORT 20B WHEN ROUTED THROUGH CARTO 3 FOR TROUBLESHOOTING. THE REPORTER WAS NOT ABLE TO PERFORM A TROUBLESHOOTING AT THE TIME OF THE CALL. ADDITIONAL INFORMATION WAS REQUESTED TO THE CUSTOMER TO OBTAIN DETAILED INFORMATION REGARDING THE EVENT AND IT WAS CONFIRMED THAT THE SIGNAL LOSS OCCURRED ON BOTH CARTO AND THE RECORDING SYSTEM AT THE SAME TIME. THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE WITH THE SIGNAL LOSS ISSUE BY BYPASSING THE CARTO 3 SYSTEM. THE PACING PERFORMED IN THE PATIENT WAS PLANNED AND ACCORDING TO CUSTOMER¿S OPINION, IT APPEARED TO BE A HARDWARE ISSUE. BLOOM STIMULATOR WAS USED BY THE CUSTOMER IN THIS PROCEDURE. THE CASE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. (B)(4). IT WAS REPORTED THAT DURING AN ISCHEMIC VENTRICULAR TACHYCARDIA (ISVT) PROCEDURE, THE CARTO 3 SYSTEM LOST 12-LEAD ECG¿S SIGNALS FOR APPROXIMATELY 3-4 SECONDS AFTER CARDIOVERSION WAS PERFORMED IN THE PATIENT. IT WAS ALSO REPORTED THAT THE 12-LEAD AND INTRACARDIAC (IC) SIGNALS WERE FLAT WHEN THE PHYSICIAN WAS PACING FROM THE RV CATHETER CONNECTED TO REF/DECA. A BLUE PIN BOX WAS IN USE WITH REF/DECA. THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE WITH THE SIGNAL LOSS ISSUE BY BYPASSING THE CARTO 3 SYSTEM. THE CASE WAS COMPLETED WITHOUT PATIENT CONSEQUENCES. THE BWI REPRESENTATIVE REPORTED LATER THAT THE ACCOUNT HAS DONE TWO CASES SINCE THE REPORTED INCIDENT AND THE PROBLEM DIDN'T DUPLICATE ITSELF DURING THESE CASES. IN ADDITION, THE COMPLAINT HISTORY OF THE SYSTEM WAS REVIEWED AND IT WAS FOUND THAT NO MORE SIMILAR PROBLEMS WERE REPORTED SO FAR. DHR REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE. THE DHR ASSOCIATED WITH CARTO 3 #13377 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE.><INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4). | 1
3023770 | IT WAS REPORTED THAT DURING AN AFIB PROCEDURE, THE CARTO 3 WAS NOT SHOWING A RED NAVISTAR TIP DURING RF. A THERMOCOOL SF NAVISTAR WAS BEING USED TO ABLATE AT 30 WATTS. THE STOCKERT APPEARED TO BE FUNCTIONING NORMALLY AND WAS SENDING ABLATION DATA TO THE RECORDING SYSTEM. THE NAVISTAR ICON TIP REMAINED GREEN ON CARTO 3. THERE WAS NO IMPEDANCE CHANGE OR ELECTROGRAM REDUCTION WAS SEEN. RF WAS BEING DELIVERED EVEN THOUGH THE GREEN TIP WAS NOT TURNING RED. PACING WAS SET TO NONE AND THE ABLATION CHANNEL ON THE RECORDING SYSTEM WAS DISABLED. NORMAL IMPEDANCES WERE DISPLAYED ON STOCKERT. CHANGING THE REDEL CABLE, REDEL ADAPTOR, NAVISTAR CABLE, NAVISTAR AND THE STOCKERT DID NOT RESOLVE THE ISSUE. THE CASE WAS COMPLETED BY CHANGING TO ANOTHER CARTO 3 AND THE NAVISTAR ICON TIP TURNED RED WITH ABLATION. NO PATIENT ADVERSE CONSEQUENCE IS REPORTED. (B)(4).><(B)(4). IT WAS REPORTED THAT DURING AN AFIB PROCEDURE, THE CARTO 3 WAS NOT SHOWING A RED NAVISTAR TIP DURING RF. A THERMOCOOL SF NAVISTAR WAS BEING USED TO ABLATE AT 30 WATTS. THE STOCKERT APPEARED TO BE FUNCTIONING NORMALLY AND WAS SENDING ABLATION DATA TO THE RECORDING SYSTEM. THE NAVISTAR ICON TIP REMAINED GREEN ON CARTO 3. THERE WAS NO IMPEDANCE CHANGE OR ELECTROGRAM REDUCTION WAS SEEN. RF WAS BEING DELIVERED EVEN THOUGH THE GREEN TIP WAS NOT TURNING RED. PACING WAS SET TO NONE AND THE ABLATION CHANNEL ON THE RECORDING SYSTEM WAS DISABLED. NORMAL IMPEDANCES WERE DISPLAYED ON STOCKERT. BIOSENSE WEBSTER FIELD SERVICE ENGINEER REPLACED THE BACKPLANE BOARD AND SET UP A WET BATH. FIELD SERVICE ENGINEER RAN FULL SYSTEM TEST AND THE SYSTEM PASSED IT SUCCESSFULLY AND THE SYSTEM WAS READY TO USE. THE DEFECTIVE BACKPLANE WAS REPLACED AND WAS SENT TO CARTO MANUFACTURER (HTC) FOR FURTHER INVESTIGATION. THE CARTO MANUFACTURER FOUND THAT A PIECE OF SOLDER SHORTED TWO PINS OF THE CONNECTOR P2 AND THIS CAUSED THE REPORTED ISSUE. THE BACKPLANE WAS SENT TO SUBCONTRACTOR FOR FURTHER REPAIR. THE DEVICE HISTORY RECORD REVIEW WAS PERFORMED. NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICE OF THIS DEVICE. | 1
3005190 | DURING THE PROCEDURE IT WAS REPORTED THERE WAS A SIGNAL FLAT LINE DURING ABLATION. ADDITIONAL INFORMATION PROVIDED STATED THE SIGNAL LOSS OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS ECGS AND ALL IC (INTRACARDIAL) RECORDINGS. THE SIGNAL LOSS OCCURRED ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME. THE PHYSICIAN WAS NOT ABLE TO COMPLETE THE PROCEDURE. THE PHYSICIAN STATED THE PATIENT WAS DOING FINE. NO PATIENT CONSEQUENCE DUE TO CANCELED PROCEDURE. (B)(4). | 1
2829433 | IT WAS REPORTED THAT DURING A PAROXYSMAL A-FIB PROCEDURE, A MAP SHIFT (4-5 MM) ON THE CARTO 3 SYSTEM WAS NOTICED. THE REPORTER STATED THAT THE ABLATION CATHETER AND LASSO WERE NOT ALIGNED. THE FLUORO CLEARLY SHOWED THE NAV THERMOCOOL TIP TOUCHING A SPECIFIC ELECTRODE ON THE NAV LASSO BUT THEY WERE A DISTANCE APART ON THE CARTO SYSTEM. BY RE-SETTING VISUALIZATION FUNCTION FIXED THE DISCREPANCY. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES. INVESTIGATION IS STILL IN PROCESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED TO THE FDA ONCE THAT THE REPAIR RECORD INVESTIGATION IS COMPLETED. (B)(4).><(B)(4). IT WAS REPORTED THAT THERE WAS A MAP SHIFT ON THIS CARTO 3. CALLER WAS TRANSFERRED TO CSS. THE CALLER STATED THAT THE ABLATION CATHETER AND LASSO ARE NOT ALIGNED UNTIL THE RESET VISUALIZATION FUNCTION WAS USED AND THE LASSO WAS NOT SHAPED PROPERLY. THE FIELD ENGINEER ARRIVED ON SITE AND MAGNETIC SYSTEM TESTS WERE PERFORMED ON THE SYSTEM AND IT PASSED. IN ADDITION, THE RECOMMENDED SID VALUES WERE RE-MEASURED. IT WAS INFORMED TO THE LAB STAFF THAT THE FLUOROSCOPE (GE INNOVA 2100 IQ) SHOULD ONLY BE UTILIZED IN AP OR RAO/LAO AT THE FOLLOWING TABLE HEIGHTS. MOVING BETWEEN AP AND RAO/LAO CAUSES MAGNETIC DISTORTION WITH THIS FLUOROSCOPE. NO PROBLEM WAS FOUND WITH THE SYSTEM. IN ADDITION, THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH THIS SPECIFIC SYSTEM WAS REVIEWED AS WELL AND THERE WAS NOT ANY ADDITIONAL COMPLAINT RELATED TO THE REPORTED ISSUE. THE SYSTEM WAS FOUND OPERATIONAL. THE DHR ASSOCIATED WITH CARTO 3 #13073 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. | 1
3304807 | DURING THE PREPARATION FOR AN ELECTIVE HEART PROCEDURE A RAID DRIVE ERROR MESSAGE CAME UP ON THE HORIZON CARDIOLOGY SYSTEM. I SPOKE TO OUR INFORMATION SYSTEMS (IS) DEPARTMENT ABOUT THE ERROR WHO INDICATED IT SHOULD BE A LOCAL DRIVE ISSUE. THE PATIENT WAS MOVED TO LAB 1 AND WE RECEIVED THE SAME ERROR. AT THE SAME TIME, THE TEAM IN LAB 3 RECEIVED THE ERROR. WHILE SPEAKING TO IS DEPARTMENT ABOUT THE PROBLEM A ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) CASE WAS CALLED FOR THE PATIENT. PREPARATIONS ARE BEING MADE CURRENTLY TO DO THE CASE UTILIZING AN HP PORTABLE MONITOR. ALL ELECTIVE CASES ARE CURRENTLY ON HOLD.UPDATE: (B)(6) WAS ABLE TO FORCE ALL LABS INTO LOCAL MODE. DUE TO A DELAY IN THE STEMI PATIENT COMING TO THE DEPARTMENT (PATIENT WAS BEING STABILIZED) WE WERE ABLE TO GET THE STEMI LAB INTO LOCAL MODE AND DID NOT NEED TO PERFORM THE CASE UTILIZING THE PORTABLE MONITOR. WE WERE LATER NOTIFIED THAT THE SEQUEL DATABASE WAS BACK UP. THE ISSUE AROSE BECAUSE WHEN THE SEQUEL DATABASE FAILED IT DID NOT AUTOMATICALLY ROLL OVER TO THE REDUNDANT SERVER AS IT IS SUPPOSED TO. SYSTEMS ADMINISTRATOR IS MEETING WITH (B)(6) AND OUR HOSPITAL TEAM TO IDENTIFY WHY THIS DID NOT WORK AS DESIGNED. WHILE THERE WERE DELAYS IN THE CASES.LABS HAVE ALL BEEN REBOOTED AND ARE CURRENTLY FUNCTIONING PROPERLY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
2945430 | AFTER FIRST EXPOSURE, TECHNICIAN GOT AN IMAGE WITH NO ANATOMY. PORTABLE MACHINE WAS RESTARTED, A TEST EXPOSURE WAS TAKEN, AND EXAM WAS COMPLETED. (B)(4). | 0
3365722 | ATRIAL FIBRILLATION WAS BEING ANTICOAGULATED WITH RIVEROXABAN, A DRUG THAT IS TO BE GIVEN IN THE EVENING ACCORDING TO THE LABEL. THE DRUG WAS HELD FOR A PROCEDURE AND THEN, THE DRUG CONTINUED BY ENTERING THE APPROPRIATE CLICKS ON THE CPOE SYSTEM IN THE A.M. THAT DAY, THE DRUG WAS NOT ADMINISTERED BECAUSE THE CPOE SYSTEM DID NOT TRANSFER THE "RESUME" ORDER TO THE RN WORK LIST. IT WAS NOT UNTIL THE NEXT DAY THAT THE PT RECEIVED THE MEDICATION. WHEN IT WAS GIVEN, IT WAS GIVEN IN THE MORNING, AT THE WRONG TIME. DANGEROUS DEFAULTS IN THE SYSTEM AND INTERFACE COMMUNICATION FAILURE PUT THIS PT AT RISK FOR STROKE AND OTHER EMBOLIC COMPLICATIONS. | 1
3289303 | PATIENT WITH PERMANENT ATRIAL FIBRILLATION AND LABILE RATE WAS ON A NEGATIVE CHRONOTROPIC MEDICATION CALLED DIGOXIN THAT WAS TITRATED AS AN OUT PATIENT IN THE CLINIC TO A DOSE OF .125 MG ON MONDAY, WEDNESDAY, AND FRIDAY TO KEEP THE HEART RATE BETWEEN 60 AND 85 BEATS PER MINUTE. THE PATIENT WAS ADMITTED TO THE HOSPITAL FOR A MEDICAL CONDITION. THE SUNRISE CPOE DEVICE HAS A DROP-DOWN MENU TO ORDER DOSING INTERVALS, WITHOUT A MANUAL OVER RIDE. THE DROP-DOWN MENU AND OVER 50 OPTIONS FOR DOSING INTERVALS BUT NONE WERE FOR MONDAY, WEDNESDAY, AND FRIDAY, OR THREE DAYS PER WEEK. AS A RESULT, THE PATIENT WAS EITHER OVERDOSED OR UNDERDOSED, WHICH IS SIGNIFICANT IN VIEW OF HIS AGE AND THE NARROW THERAPEUTIC RATIO OF DIGOXIN. THIS REPRESENTS A DEFECT IN THE DEVICE AND THE FDA SHOULD REQUIRE A MANUAL OVER RIDE OPTION IN ORDER TO INDIVIDUALIZE MEDICATION DOSES. | 1
3502101 | THE DEVICE CALLED CPOE INVOLVES ELECTRONIC ORDERING OF TESTS AND MEDICATIONS AND TREATMENTS. BLOOD TESTS ARE ROUTINELY AS STANDARD OF CARE ORDERED TODAY FOR THE NEXT MORNING, EG "PROTHROMBIN TIME IN THE AM." ALL STAFF INVOLVED WITH THE PATIENT CARE KNOW EXACTLY WHAT THIS MEANS. ON THIS DEVICE, THERE IS NOT AN OPTION IN THE MENUS FOR THAT SPECIFIC ORDER FOR ANY TEST OR TREATMENT, AND EVEN WHEN IT IS SO STATED IN THE COMMENTS, THE BUILT IN CALENDAR DIRECTS THE BLOOD TEST TO BE DONE "TODAY" IN THE AM RATHER THAN IN THE AM TOMORROW, EVEN THOUGH THE BLOOD TODAY HAS ALREADY BEEN TESTED. THIS CAUSES DOUBLE TESTING WHEN IT IS NOT NEEDED AND NO TEST WHEN IT IS NEEDED, RESULTING IN DELAYS AND CONFUSION AS TO THE MONITORING OF DISEASE CRITICAL PARAMETERS IN CRITICALLY ILL PATIENTS. THESE DELAYS IN CARE ARE LIFE THREATENING AND COSTLY, ESPECIALLY PERTAINING TO POTASSIUM LEVELS THAT ARE OUT OF RANGE AND HEMOGLOBIN LEVELS THAT ARE DANGEROUSLY LOW. | 1
3339289 | THE PT WAS DEFICIENT IN A VITAMIN. IT WAS ORDERED BY THE ATTENDING PHYSICIAN FOR IT TO BE REPLACED. A CONSULTANT CAME ALONG AND ORDERED A DIFFERENT FORM OF THE SAME VITAMIN WITHOUT DISCONTINUING THE OTHER. THIS HAPPENS OFTEN AND IS SOMETIMES MISSED BY THE PHARMACY THAT DELIVERS BOTH DRUGS, OFTEN ADMINISTERED AT DIFFERENT TIMES. THE NURSES DO NOT SEE THE OVERALL LIST. THEY ONLY SEE WHAT THEY ARE TASKED WITH. THE DOCTOR DOES NOT KNOW THAT A DUPLICATE WAS ORDERED BECAUSE THE DOCTOR IS NOT ABLE TO SEE AN ALPHABETIZED LIST OF MEDICATIONS, AND MISSES THE VITAMIN ORDERED PREVIOUSLY BECAUSE OF THE MONOTONY OF SCROLLING THOUGH 2-3 PAGES OF MEDICATIONS PRESENTED IN RANDOM ORDER. THE CPOE DEVICE IS NOT SMART ENOUGH TO KNOW THAT THE SEPARATELY NAMED VITAMINS ARE THE SAME. OVER TIME, THIS MAY RESULT IN A HYPERVITAMINOSIS SYNDROME. | 1
3339297 | THE PT WAS RECOVERING FROM SEPSIS, ACUTE KIDNEY, RESPIRATORY, AND HEART FAILURE. SUBTLE, BUT SUDDEN CLINICAL REGRESSION OCCURRED WITHOUT THE TRADITIONAL SIGNS OF SEPSIS OR GENITOURINARY INFECTION. URINE AND BLOOD CULTURES WERE ORDERED. THE CPOE OF THE EHR COMPUTER HAD CLINICAL DECISION SUPPORT (CDS). THE CDS DEVICE AUTOMATICALLY DISCONTINUED THE URINE CULTURE AND ORDERED A URINALYSIS ON THE DOCTORS' TWO ATTEMPTS TO ORDER THE DISEASE CRITICAL DIAGNOSTIC TESTS. APPROX 4 DAYS PASSED AND THERE STILL WAS NOT A URINE CULTURE, WHILE THE PT SLIPPED BACK IN TO KIDNEY FAILURE, RESPIRATORY FAILURE, AND PULMONARY EDEMA AND WAS PLACED BACK ON THE VENTILATOR. FINALLY, SEVERAL MORE DAYS LATER, A URINE CULTURE WAS DONE SHOWING >100,000 GRAM NEGATIVE RODS, THAT THEN GREW OUT OF THE BLOOD CULTURE. THE DELAY IN DIAGNOSIS WAS DIRECTLY A RESULT OF THE CDS USURPING THE ROLE OF DOCTOR, COUNTERMANDING TESTS. THIS CDS DEVICE DOES NOT HAVE ANY FUNCTIONAL OVER RIDE TO GET THE URINE CULTURE DONE. DEATH FROM THE DELAY IN DIAGNOSING AND TREATING SEPSIS WAS NARROWLY AVOIDED. THE CDS DEVICE IS DEFECTIVE IN SO FAR AS IT DOES NOT PERMIT AN OVER RIDE, EVEN IF IN THE BEST JUDGEMENT OF THE PHYSICIAN. | 1
3339308 | HEART FAILURE WAS TREATED WITH DIURETICS INCLUDING LOOP AND DISTAL TUBULE K+ SPARING DIURETIC. HOLD DIURETIC PARAMETER FOR A POTASSIUM OF GREATER THAN 4.8 MG/DL WAS PART OF THE DAILY ORDER ISSUED TO PHARMACY AND NURSING. THE K+ SPARING DIURETIC WAS ADMINISTERED TO THE PT MULTIPLE TIMES WHEN THE POTASSIUM EXCEEDED THAT PARAMETER, CAUSING THE SERUM POTASSIUM TO REACH A LIFE THREATENING LEVEL, REQUIRING EMERGENCY THERAPY FOR HYPERKALEMIA. THE DESCRIPTION OF THE ORDER INCLUDES NOTICES THAT THE DRUG CAUSES CANCER, THAT "UNNECESSARY USE OF THIS DRUG SHOULD BE AVOIDED", AND THE PILL SIZE, AND ADDITIONAL IRRELEVANT INFO. AMIDST THAT IS THE HOLD PARAMETER. THE EXCESSIVE VERBIAGE MUST HAVE CONFUSED NO FEWER THAN SEVEN DIFFERENT NURSES AND THE PHARMACIST(S). | 1
2948908 | AT OUR INSTITUTION, THERE ARE VEXING INTERFACE PROBLEMS BETWEEN ELECTRONIC ORDERING DEVICE CPOE AND THE LAB. THE HEALTH CARE PROFESSIONALS CAN NOT TRUST THE SYSTEM OF DEVICES AND HUMAN INTERACTIONS WITH THEM TO OBTAIN THE RIGHT TESTS AT THE RIGHT TIME. LIKEWISE, LIFE CRITICAL TEST COME BACK TO AN EMR SILO SILENTLY W/O NOTICE, WASTING PRECIOUS MINS AND DELAYING CARE TO THE MOST VULNERABLE. FOR INSTANCE, THE PT HAD CARDIO-RESPIRATORY ARREST. AN ARTERIAL BLOOD GASS WAS OBTAINED FROM THE PT, YET, DURING THE ONGOING CRITICAL CARE MANAGEMENT, NO ONE WAS PROVIDED THE RESULTS WHEN THEY WERE READY, BUT THEY WERE SILENTLY IN THE EMR W/O WARNING. THE PH WAS LESS THAN 7. CARE WAS DELAYED. | 1
3305468 | PT'S MEDICATION LIST AND OTHER ACTIVE ORDERS DID NOT APPEAR ON THE DOCTOR'S ORDER SECTION ON THE CPOE SYSTEM RENDERING IT IMPOSSIBLE FOR THE DOCTOR TO CONFIRM, ALTER, AND RECONCILE THE MEDICATION LIST. THE UP TO DATE MEDICATION LIST APPEARED ON THE NURSES' CPOE SYSTEM UNDER ORDERS. IT IS NOT CLEAR WHY THE MEDICATION LIST ON THIS PT WAS SEQUESTERED. FOR OBVIOUS REASONS, THIS DEFECT IN THE CPOE IS POTENTIALLY LIFE THREATENING WHEN THE DOCTOR(S) DO NOT HAVE ACCESS TO THE CURRENT MEDICATION LIST. | 1
3305444 | THE PT WAS SUFFERING FROM INTERNAL BLEEDING. THE CPOE ORDER WAS FOR HCT/HGB TO BE DONE EVERY 6 HOURS ON A 6-12 O'CLOCK SCHEDULE TO DETERMINE STABILITY OR NEED FOR MORE AGGRESSIVE INTERVENTION OR BLOOD TRANSFUSION. THE H AND H HAD BEEN 28% AND 9.1 GM/DL. AT THE TIME OF THE ANTICIPATED 12 O'CLOCK BLOOD DRAW, THE PATIENT WAS NOT IN THE ROOM BECAUSE TESTS WERE BEING DONE. SEVERAL HOURS LATER, THE PATIENT COLLAPSED IN SYNCOPE WITH LIFE THREATENING SLOW HEART RATE AND HYPOTENSION. IT WAS DISCOVERED THEN THAT THE 12 O'CLOCK BLOOD TEST RESULT WAS NOT AVAILABLE BECAUSE THE TEST WAS NEVER DRAWN. WITH ALL OF THE HYPE OVER HOW THESE CPOE AND EHR DEVICES WERE GOING TO MAKE CARE SAFER, THESE DEVICES FAILED TO WARN OR INDICATE TO THE NURSE OR THE PHLEBOTOMIST THAT THE DISEASE CRITICAL BLOOD TESTS WAS NOT YET DONE, WITH THE NURSES CHAINED TO THE USER UNFRIENDLY EHR, THEY TENDED TO NEGLECT THE PTS. THIS PATIENT SUFFERED FROM THAT NEGLECT FOR THE H AND H WAS 21% AND 6.9 GM/DL WHEN IT WAS DONE AFTER THE COLLAPSE, REQUIRING EMERGENCY TRANSFUSION AND ICU CARE. THE FAILURE OF THE EHR DEVICE TO PROVIDE A WARNING THAT TESTS WERE NOT DONE AS ORDERED REPRESENTS A POTENTIAL LIFE THREATENING DESIGN FLAW IN THE DEVICE. | 1
3354787 | SEVERE LABILE HYPERTENSION IN THE SETTING OF KIDNEY FAILURE WAS NOT TREATED DESPITE ORDERS THAT WERE ON THE CPOE DEVICE FOR PRN THERAPY. THE NURSES ARE OFTEN BLAMED FOR THE OMISSION, BUT THE PRN MEDICATION IS NOT CLEARLY DELINEATED ON THEIR TASK SCHEDULE GENERATED BY THE EHR CPOE DEVICE. THE FAILURE TO ADMINISTER MEDICATIONS THAT MEET THE PHYSIOLOGICAL PARAMETERS IN THE PRN ORDER IS EXTENSIVE. IN THIS CASE, THE PATIENT'S KIDNEYS WERE FURTHER DAMAGED FROM UNTREATED SEVERE HYPERTENSION REQUIRING HEMODIALYSIS. | 1
3342067 | POTASSIUM CHLORIDE WAS PRESCRIBED TWICE PER DAY AS TREATMENT FOR HYPOKALEMIA. THE LAB TESTING REVEALED A JUMP IN THE POTASSIUM LEVEL, BUT THE RESULT CAME TO THE EHR WITHOUT ALERT OR WARNING, AND THE NURSES CONTINUED TO GIVE THE PT POTASSIUM ANYWAY. THE NURSE DID NOT KNOW THAT THE POTASSIUM LEVEL WAS HIGH, A WIDESPREAD PROBLEM OF EHR ASSOCIATED AUTOMATION, WITH RESULTS BEING DEPOSITED ELECTRONICALLY AND SILENTLY IN THE EHR SILO. THOUGH THIS PT DID NOT DIE, OTHERS HAVE FROM THIS TYPE OF DEFECT. | 1
2825339 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. (B)(4). | 1
3290030 | THE CUSTOMER'S FATHER REPORTED INCORRECT TIME/DATE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 1
2810294 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR PRECISION XTRA BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE (B)(4) LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. (B)(4). | 1
3007396 | A DISCORDANT, FALSELY LOW BRAIN NATRIURETIC PEPTIDE (BNP) RESULT WAS RELEASED FOR A PATIENT SAMPLE AFTER ERROR MESSAGES FROM THE ADVIA CENTAUR CP INSTRUMENT WERE TRANSMITTED THROUGH THE (B)(4) LABORATORY INFORMATION SYSTEM (LIS), WHERE THEY WERE MISINTERPRETED. THE ADVIA CENTAUR CP INSTRUMENT PRODUCED ERROR MESSAGES INDICATING THAT THE THERE WAS INSUFFICIENT SAMPLE VOLUME AND THAT THE TASK WAS NOT COMPLETE. THE LIS MISINTERPRETED THESE ERROR MESSAGES, AND A RESULT WAS GENERATED AND RELEASED. IT IS UNKNOWN IF THE DISCORDANT RESULT WAS REPORTED TO THE PHYSICIAN(S). THE SAMPLE WAS RERUN ON AN ALTERNATE INSTRUMENT, AND A HIGHER RESULT WAS OBTAINED. THE LABORATORY THEN RAN TWO "TEST" SAMPLES FOR TROPONIN AND BNP TO SEE IF THE ISSUE WAS REPRODUCIBLE. FOR BOTH SAMPLES, RESULTS WERE GENERATED DESPITE THE INSTRUMENT TRANSMITTING ERROR MESSAGES TO THE LIS. THERE ARE NO KNOWN REPORTS OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCES DUE TO THE DISCORDANT BNP AND TROPONIN RESULTS. THE LIS VENDOR HAS CHANGED HOW THE LIS PROCESSES NON-NUMERIC RESULTS FROM THE ADVIA CENTAUR CP INSTRUMENT. THERE HAVE BEEN NO REOCCURRENCES OF THIS ISSUE SINCE THE LIS PROGRAMMING WAS CHANGED. THE SIEMENS INSTRUMENTS ARE PERFORMING ACCORDING TO SPECIFICATIONS. NO FURTHER EVALUATION OF THIS DEVICE IS REQUIRED. | 1
3393531 | THE DIGITAL FIXATION DEVICE WAS IMPLANTED ON (B)(6) 2013 BY DR (B)(6) DPM. THE SURGEON REPORTED THAT THE PT COMPLAINED OF PAIN A COUPLE OF WEEKS FOLLOWING THE PROCEDURE. A RADIOGRAPH REVEALED THAT THE IMPLANT DISLOCATED FROM THE MIDDLE PHALANX AND TWISTED. THE SURGEON SHARED IMMEDIATE POST-OP FLUORO AND THE FOLLOW-UP RADIOGRAPH. THE PT HAD VERY SMALL MIDDLE PHALANGES. THE MIDDLE PHALANX WAS SO SHALLOW THAT BARELY 2MM OF BONE WAS TRAILING THE DISTAL ARROWHEAD TIP. THE SHALLOW PLACEMENT RESULTED IN INADEQUATE ENGAGEMENT AND PULLOUT. THE PROXIMAL ARROW TWISTED AND DID NOT GO DEEP ENOUGH, ENGAGING THE DENSE SUBCHONDRAL BONE AT THE BASE OF THE PROXIMAL PHALANX. PLACING THE IMPLANT LONG AND DEEP IN THE PROXIMAL PHALANX AS STATED IN THE OPERATIVE TECHNIQUES SHOULD RESULT IN ENOUGH FIXATION TO PREVENT THE IMPLANT FROM ROTATING. AFTER REVIEW OF THE PT POST-OP FLUORO AND THE FOLLOW-UP RADIOGRAPH IT WAS DETERMINED THAT THE SURGEON DID NOT SELECT THE PROPER LENGTH IMPLANT TO ACHIEVE ADEQUATE FIXATION. FOLLOWING THE PUBLISHED OPERATIVE TECHNIQUE WOULD HAVE RESULTED IN THE SELECTION OF THE PROPER LENGTH OF IMPLANT AND IT'S POSITION IN THE BONE. | 0
3514066 | X-RAY TECHNICIAN (CRT) WAS POSITIONING PATIENT FOR AN X-RAY OF THE WRIST. THE TECHNICIAN TRIPPED OVER A CORD AND FELL TO THE FLOOR LANDING ON HER LEFT SIDE. THE TECHNICIAN FELT PAIN IN LEFT HIP, AND DISCOMFORT TO LEFT WRIST AND ELBOW. THE TECHNICIAN ELECTED TO NOT GO BY AMBULANCE TO THE EMERGENCY ROOM, BUT RATHER HAD X-RAY'S OF THE HIP TAKEN ON-SITE. THE X-RAY SHOWED A HIP FRACTURE. THE TECHNICIAN WAS THEN TAKEN TO (B)(6) HOSPITAL FOR ADMISSION AND SURGERY. (B)(4). THE UDI DIGITAL UPGRADE KIT INCLUDES A DIGITAL FLAT PANEL DETECTOR AND CABLE WHICH IS CONNECTED TO AN INTERFACE BOX DURING INSTALLATION AND USE. IT IS THIS CABLE THAT THE USER TRIPPED OVER WHILE PREPARING TO X-RAY A PATIENT. (B)(4) IS THE DISTRIBUTOR THAT SOLD AND INSTALLED THE UPGRADE KIT TO THE USER FACILITY: (B)(6). | 0
3513830 | PADDLE USED FOR MAMMOGRAPHY NOTED TO BE CRACKED PRIOR TO USE ON THE PATIENT. A NEW PADDLE WAS RECEIVED. | 0
3339418 | AN INCORRECT RESPONSE OCCURRED WHEN A BLOOD DONOR USED HAEMONETICS SOFTWARE SOLUTIONS DONOR DOC SOFTWARE TO COMPLETE A DONOR HEALTH QUALIFICATION FORM. THE DONOR DOC SOFTWARE DID NOT PERFORM AS INTENDED WHICH ALLOWED THE DONOR TO BE DEEMED ELIGIBLE TO DONATE. SPECIFICALLY, QUESTION #2 WAS LEFT UNANSWERED WHICH SHOULD HAVE PROHIBITED THE DONOR FROM DONATING UNTIL THAT QUESTION WAS ANSWERED CORRECTLY. BLOOD PRODUCTS ASSOCIATED WITH THE DONATION WERE DISTRIBUTED BASED ON INACCURATE DONOR HISTORY INFORMATION. | 1
3075014 | .><.><AGFA IS VOLUNTARILY REPORTING THE FOLLOWING EVENT: CUSTOMER, (B)(6), HAS FIVE (5) DX-D 100 WIRELESS UNITS USED WITHIN TWO (2) OF THEIR SITES. THE UNITS HAVE FUNCTIONED WELL WITH NO PROBLEMS. ONE OF THE UNITS INSTALLED IN (B)(4) 2012, HAS FUNCTIONED WELL WITHOUT PROBLEMS FOR 10 MONTHS, BUT HAS RECENTLY SHOWN UNEXPECTED BEHAVIOR. AGFA WAS MADE AWARE OF THIS BEHAVIOR ON (B)(6) 2013, WHEN THE CUSTOMER CONTACTED AGFA. THERE WERE THREE (3) EVENTS LEADING UP TO AGFA'S AWARENESS: EVENT ONE: ON (B)(6) 2013, TWO HOSPITAL TECHNICIANS WERE IN AN ELEVATOR AND EXITING WITH THE DX-D 100 UNIT THAT WAS INSTALLED IN (B)(4) 2012. AS THEY STARTED TO MOVE FORWARD OUT OF THE ELEVATOR, THE UNIT SUDDENLY TURNED A HARD LEFT ON ITS OWN, THEN STRAIGHTENED OUT AND MOVED BACKWARDS TOWARDS THE TECHNICIANS. THE TECHNICIAN DRIVING THE UNIT LET GO OF THE HANDLE AND THE UNIT CONTINUED TO MOVE BACKWARDS UNTIL THE UNIT "PARTIALLY PINNED" THE TECHNICIANS TO THE ELEVATOR WALL. AS SOON AS THE UNIT CAME IN CONTACT WITH THE TECHNICIANS IT CAME TO A COMPLETE STOP. THE TECHNICIAN WAS ABLE TO DRIVE THE UNIT FORWARD AND OUT OF THE ELEVATOR. NO ONE WAS INJURED BY THIS INCIDENT. THIS UNIT OPERATED WITH COMPLETE NORMAL MOVEMENT BEFORE AND AFTER THE INCIDENT. NO COMPLAINT WAS FILED BY THE CUSTOMER. EVENT TWO: ON (B)(6) 2013, TWO HOSPITAL TECHNICIANS WERE RETURNING THE SAME UNIT AS DESCRIBED IN EVENT ONE, TO ITS PARKING AREA FOR CHARGING. THE TECHNICIAN DRIVING THE UNIT, SUCCESSFULLY TURNED RIGHT AND THEN STRAIGHT FOR A FEW METERS, WHEN THE UNIT SUDDENLY ACCELERATED AND TURNED A HARD RIGHT. THE TECHNICIAN TRIED TO PULL THE UNIT BACK IN ORDER TO STRAIGHTEN THE UNIT'S DIRECTION, BUT THE UNIT GRAZED THE WALL. WHEN THIS OCCURRED, THE TECHNICIAN LET GO OF THE SWITCH AND THE UNIT STOPPED MOVING. NO ONE WAS INJURED BY THIS INCIDENT. THIS UNIT OPERATED WITH COMPLETE NORMAL MOVEMENT BEFORE AND AFTER THE INCIDENT. NO COMPLAINT WAS FILED BY THE CUSTOMER. EVENT THREE: ON (B)(6) 2013, A TECHNICIAN HAD BEEN WORKING ON THE EVENING SHIFT AND WAS USING THE SAME UNIT DESCRIBED IN EVENTS ONE AND TWO, FOR NEARLY 15 MINUTES. THE TECHNICIAN NEEDED TO ENTER THE ELEVATOR WITH THE UNIT AND PROCEEDED TO MAKE A RIGHT TURN INTO THE ELEVATOR. IMMEDIATELY AFTER THE REAR WHEELS OF THE UNIT CROSSED THE ELEVATOR THRESHOLD, THE UNIT SUDDENLY MOVED FORWARD ON ITS OWN AND THEN ACCELERATED UNEXPECTEDLY. THE TECHNICIAN ATTEMPTED TO STOP THE UNIT BY PRESSING THE HANDLE AND FORCEFULLY PULLING BACK, BUT THE UNIT PROCEEDED TO MOVE FORWARD AT AN ACCELERATED RATE. THE UNIT TURNED SLIGHTLY RIGHT AND HIT THE REAR WALL OF THE ELEVATOR WITH FORCE AND STOPPED. THE TECHNICIAN RESTARTED THE UNIT AND ATTEMPTED TO DRIVE IT BACKWARDS OUT OF THE ELEVATOR. THE UNIT HAD NO FORWARD MOTION CAPABILITY AND ONLY MOVED BACKWARDS INTERMITTENTLY, AND IN SHORT SPURTS OF MOVEMENT. THE TECHNICIAN SMELLED SMOKE SOMETIME AFTER THE UNIT HIT THE ELEVATOR WALL. NO ALARMS OR WARNINGS WERE VISIBLE ON THE CONSOLE SCREEN. NO ONE WAS INJURED BY THIS INCIDENT. THE UNIT WAS REMOVED BY THE CUSTOMER AND NOT USED ANYMORE AFTER THIS 3RD EVENT. THE CUSTOMER CONTACTED AGFA THE NEXT DAY ON (B)(4) 2013 AND REQUESTED A COMPLAINT BE OPENED FOR THE EVENTS INVOLVED WITH THE PROBLEM DX-D 100 UNIT. AGFA CUSTOMER SUPPORT SPECIALIST PERFORMED INITIAL ONSITE TESTING AT THE CUSTOMER SITE AND DISCOVERED A DEFECTIVE MOTION CONTROL BOARD AND A DEFECTIVE MOTOR BATTERY PACK ON THE DX-D100 UNIT. AGFA'S SUPPLIER OF THE DX-D100, (B)(4), ALONG WITH AGFA REPRESENTATIVES HAVE INVESTIGATED THE UNIT TOGETHER AT THE CUSTOMER SITE IN (B)(4). PARTS WERE REPLACED ON THE PROBLEM UNIT IN ORDER TO MAKE THE UNIT OPERATIONAL WITH THE ABILITY TO PERFORM ADDITIONAL TESTING FOR THE INVESTIGATION. THE COMPLETE DRIVE PART OF THE UNIT WAS CHECKED AND TESTED. TWO DEFECT BOARDS WERE REPLACED; THOSE FOR MOTOR BATTERY LOADING AND DIGITAL MOTION. THE VOLTAGE ADJUSTMENT OF THE LEFT GAUGE WAS INCORRECT AND THERE WAS NOT A GOOD CONNECTION TO THE ENCODER OF THE RIGHT MOTOR. THERE WAS ALSO A BAD CONNECTION, A BURNED TRANSISTOR OF THE LEFT MOTOR. IT HAS BEEN DETERMINED THAT BOTH THE RIGHT AND LEFT DRIVE SYSTEM SHOWED FAILURES THAT CAN INFLUENCE THE DRIVING OF THE UNIT AND ALL OF THE RELATED PARTS CAUSING THE PROBLEMS HAVE BEEN REPLACED. THESE PROBLEM PARTS WILL BE SHIPPED BACK TO (B)(4). THE CUSTOMER'S DX-D100 UNIT IS NOW WORKING PROPERLY AND THE CUSTOMER HAS ACCEPTED USING THE UNIT AGAIN. NO INJURIES TO ANYONE NOR HARM TO PATIENTS OCCURRED DURING THIS DESCRIBED EVENT. .><(B)(4). AS ROOT CAUSE HAS NOT YET BEEN DETERMINED, AGFA IS WORKING CLOSELY WITH , OUR OEM FOR THE DX-D100 UNITS. ONCE ROOT CAUSE IS DETERMINED, A PATH FORWARD FOR CORRECTIVE ACTION WILL BE IDENTIFIED AND PROPERLY COMMUNICATED TO DESCRIBE THE PLAN. A MANDATORY SERVICE BULLETIN WILL BE CREATED TO ADDRESS THE CORRECTIVE ACTION/S REQUIRED TO ELIMINATE THE UNINTENDED MOVEMENTS OF THE DX-D100 AND DX-D100 WIRELESS UNITS. ANOTHER SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE ROOT CAUSE AND CORRECTIVE ACTION IS IMPLEMENTED. THESE ACTIONS, ONCE KNOWN, WILL ALSO BE REPORTED WITHIN OUR REPORTABLE CORRECTION TO THE FDA.><THIS IS A 2ND FOLLOWUP SUPPLEMENT TO REPORT THAT AGFA HAS DETERMINED, ALONG WITH (B)(4), THE ROOT CAUSE OF THE PROBLEM OF SPORADIC MOVEMENT WITH THE DX-D100 PRODUCT WAS DEFECTIVE FIRMWARE ON THE DIGITAL MOTION CONTROL BOARD. A MANDATORY SERVICE BULLETIN WAS ISSUED VIA OUR REPORTABLE CORRECTION COMMUNICATED TO THE FDA (FDA REFERENCE # - Z-1487-13). FIRMWARE OF THE DIGITAL MOTION CONTROL BOARD WILL BE UPGRADED TO VERSION V5R6B0. AS A LONG TERM CORRECTIVE ACTION, THE DEFECTIVE MOTION CONTROL BOARD (DMCB) WILL BE UPDATED AND A 2ND MANDATORY SERVICE BULLETIN WILL BE ISSUED TO CHANGE OUT THE OLD DMCB. ALL FURTHER REPORTS WILL BE COMMUNICATED VIA FDA REFERENCE # - Z-1487-13. | 0
3289330 | POST OPERATIVE PATIENT WAS BEING INFUSED NORMAL SALINE AT A RATE OF 125 PER HOUR. SERUM SODIUM RESULT WAS SENT ELECTRONICALLY TO THE EHR BUT THERE WAS NO WARNING OR NOTICE THAT THE RESULT WAS THERE. A HEMOGLOBIN LEVEL WAS LESS THAN 7 WHICH ALSO WAS SENT TO THE EHR REPOSITORY, BUT NO ONE KNEW THAT THERE WAS A NEW RESULT THERE. THIS RESULTED IN A DELAY OF TREATMENT WITH LIFE THREATENING POTENTIAL CONSEQUENCES. | 1
3493682 | BLOOD PRESSURE WAS LOW HOURS BEFORE A SCHEDULED DIALYSIS IN AN ALERT FUNCTIONAL HOSPITALIZATION PATIENT WITH RECENT ONSET KIDNEY FAILURE. THE HYPOTENSION WAS TREATED. FOR 24 HOURS, THE VITAL SIGNS WERE LOST/UNAVAILABLE ON THE EHR, THAT PREVENTED THE DIALYSIS NURSES FROM BEING AWARE OF THE RECENT UNEXPLAINED HYPOTENSION. THE BLOOD PRESSURE BECAME LOW NORMAL BEFORE DIALYSIS WAS STARTED. THE DIALYSIS WAS STARTED WITH AN AGGRESSIVE FLUID REMOVAL GOAL, THAT PRESUMABLY WOULD HAVE BEEN ALTERED (REDUCED) HAD THE DIALYSIS NURSES BEEN AWARE AND INFORMED THE NEPHROLOGIST WHO HAD ORDERED THE AGGRESSIVE FLUID REMOVAL. THE PATIENT BECAME HYPOTENSIVE WITHOUT WARNING, LOST CONSCIOUSNESS AND DIED, DESPITE CPR. OF ADDITIONAL IMPORTANCE, THE EHR HAD LISTED THE PATIENT AS A "DO NOT RESUSCITATE", YET THE RESUSCITATION WENT ON DESPITE THAT. IT IS A COMMON PROBLEM THAT THE EHR IS NOT READ FOR INNUMERABLE REASONS, INCLUDING TOO MUCH INFO AND TOO SMALL A FONT, AND WHEN IT IS READ, DISEASE CRITICAL DATA IS MISSING, LOST, OR IN THE WRONG SILO. | 1
3289329 | THE PATIENT WAS SUFFERING FROM MULTI-ORGAN FAILURE. A CHEST RADIOGRAPH WAS ORDERED USING THE HOSPITAL'S CPOE DEVICE. THE CHEST RADIOGRAPH WAS NOT DONE BECAUSE THE ORDER WAS NEVER RECEIVED BY THE INTENDED RECIPIENT, THAT BEING THE RADIOLOGY DEPARTMENT. SUCH INTEROPERABILITY FAILURES ARE WIDESPREAD AND RESULT IN DELAYS OF TREATMENT OF THE CRITICALLY ILL PATIENTS. THE RADIOLOGY SYSTEM IS A GENERAL ELECTRIC MEDICAL SYSTEM DEVICE KNOWN AT THE CENTRICITY RA 600. THE CPOE DEVICE IS SPECIFIED BELOW. REASON FOR USE: RADIOLOGY SYSTEM. | 1
3387578 | THIS CUSTOMER REPORTED TO THE (B)(4) THAT A PT HAD AN UNNECESSARY EMERGENCY CARDIO ANGIOGRAM WITH POST-OPERATIVE COMPLICATIONS. THIS WAS A RESULT OF THE COPY BUTTON BEING ACCIDENTLY USED TO PRINT THE 12 LEAD ECG REPORT WHICH RESULTED IN A PREVIOUS PT'S REPORT BEING USED TO UNDERTAKE SURGICAL INTERVENTION. THE COMPLAINT REPORTS IT APPEARS NO SECONDARY / VERIFICATION OF THE DATA TOOK PLACE BEFORE COMMENCING SURGERY. THE PT WAS POSITIVELY IDENTIFIED FROM THEIR WRISTBAND. PT DATA NOT AVAILABLE. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3417304 | OVER THE PAST 60 DAYS, SEVERAL PATIENTS ON THIS PARTICULAR DEVICE WERE OVER-HEATING PERIODICALLY FOR REASONS NOT UNDERSTOOD. TESTS WERE DONE ON WARMERS BUT THE PROBLEM COULD NOT BE IDENTIFIED UNTIL IT WAS DISCOVERED AN ISSUE WITH THE ROUND PIECE OF TAPE ON THE WARMER SENSOR WHICH ATTACHES TO THE PATIENT. THE TAPE WHICH ACTS AS AN ADHESIVE TO THE CHILD AS WELL AS A HEAT BLOCKER FOR THE SENSOR WOULD INEXPLICABLY DETACH FROM CERTAIN PATIENTS. WE BELIEVE THE TAPE COMES LOOSE DUE TO A PATIENT'S SKIN HUMIDITY. WHEN THIS OCCURS THE SENSOR DOES NOT ADEQUATELY DETECT THE PATIENT'S TEMPERATURE (READS COOLER) THUS THE PATIENT BEGINS TO OVERHEAT. THESE EVENTS HAVE NOT LED TO PATIENT HARM. | 0
3300972 | DURING INSPECTION OF THE RAILS AND SCREWS, IT WAS FOUND THAT MORE THAN ONE LATERAL OUTER BLADE TO LATERAL CART SCREWS WERE LOOSE. NO DETECTOR FALL EVENT AND NO INJURY WAS REPORTED. THE LOOSE SCREWS IN QUESTION MAY IMPACT THE LATERAL DRIVE BY CREATING A GAP BETWEEN THE LINEAR BLADE AND THE LATERAL CART THAT MIGHT LEAD TO BALL SCREW STRESS WHICH IN TURN MAY RESULT IN A MECHANICAL FAILURE. THIS ISSUE WAS IDENTIFIED AS PRESENTING A DIFFERENT FAILURE MODE THAN THE ONE IDENTIFIED IN MDR # 9613299-2013-00001. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. | 0
3289331 | THE PATIENT HAD DIARRHEA. A STOOL ANALYSIS RESULT WAS RETURNED ELECTRONICALLY TO THE ELECTRONIC MEDICAL RECORD. TWENTY FOUR HOURS LATER, A HEALTHCARE PROFESSIONAL SAW THE RESULT NECESSITATING IMMEDIATE TREATMENT. BECAUSE THERE WAS NOT ANY NOTIFICATION THAT A NEW RESULT HAD ARRIVED AT THE ELECTRONIC MEDICAL RECORD, THE TREATMENT WAS DELAYED BY 24 HOURS CAUSING INCREASED MORBIDITY IN A PATIENT ALREADY DEBILITATED BY SEVERE ILLNESS. THE FAILURE OF THE DEVICE TO INDICATE THAT A NEW RESULT HAS ARRIVED IS A DEFECT IN THE DESIGN OF THE DEVICE. DELAYS IN TREATMENT HAVE LIFE THREATENING CONSEQUENCES. ALL ELECTRONIC MEDICAL RECORD DEVICES RECEIVING RESULTS WITHOUT SUCH WARNING SHOULD BE REDESIGNED TO PROVIDE SAID NOTIFICATION. AND, WARNING LETTER SHOULD BE SENT TO ALL USERS ALERTING THEM OF THIS DEFECT. | 1
3149555 | A MISIDENTIFICATION OCCURRED WHEN THE PT'S NAME WAS LOST FROM THE ATTENDING DOCTOR AND BECAME LISTED UNDER THE NAME OF ANOTHER PHYSICIAN WITHOUT ANYONE KNOWING ABOUT IT FOR MORE THAN 24 HOURS. THE WRING HEALTH CARE PROFESSIONALS WERE CALLED AND A TEST WAS PERFORMED THAT HAD LIFE THREATENING RISK, BUT WAS DONE BECAUSE THE ATTENDING PHYSICIAN WAS NO LONGER RUNNING THE CASE, ACCORDING TO THE (B)(4). THE PT DEVELOPED ACUTE RENAL FAILURE FROM THE CONTRAST AGENT USED DURING THE TEST. | 1
3487540 | OLYMPUS WAS INFORMED THAT A PEDIATRIC COLONOSCOPE CULTURED POSITIVE FOR AN UNK MICROORGANISM. OLYMPUS FOLLOWED UP WITH THE USER FACILITY TO OBTAIN ADDITIONAL INFO REGARDING THE EVENT WITH NO RESULTS. THE DEVICE WAS RETURNED TO OLYMPUS FOR EVALUATION. A PHYSICAL EVALUATION WAS PERFORMED ON THE DEVICE. THE DEVICE WAS INSPECTED USING A SMALL DIAMETER BAROSCOPE AND SLIGHT RESIDUE AND DEBRIS WERE FOUND IN THE BIOPSY CHANNEL, CHANNEL MOUNT, SUCTION CHANNEL, AND THE AIR/WATER CYLINDER UNIT. THE DEVICE HAD SCRAPES AND SCRATCHES INSIDE THE BIOPSY CHANNEL WALL, WHICH COULD HAPPEN WHEN AN OPEN OR BROKEN FORCEPS WAS INSERTED INSIDE THE CHANNEL. IN ADDITION, THE BENDING SECTION GLUE WAS CRACKED AND THE GLUE AROUND THE LIGHT GUIDE LENS WAS CHIPPED. THERE WERE NICKS AND DENTS ON THE DISTAL END COVER, AND A SMALL CHIP ON THE OBJECTIVE AND LIGHT GUIDE LENS. THE DEVICE WAS SERVICED AND RETURNED TO THE USER FACILITY. THE USER'S EXPERIENCE CANNOT BE CONCLUSIVELY DETERMINED; HOWEVER, THE REPROCESSING PRACTICES COULD NOT BE RULED OUT AS A CONTRIBUTORY FACTOR TO THE REPORTED EVENT. THE USER FACILITY HAD OLYMPUS REPROCESSING TRAINING IN OCTOBER AND IS IN THE PROCESS OF IMPLEMENTING CHANGES TO THEIR REPROCESSING PROCESS. | 0
3472050 | OLYMPUS WAS INFORMED THAT DURING A DIAGNOSTIC COLONOSCOPY, FLUID CAME OUT OF THE AUXILIARY CHANNEL OF THE DISTAL END OF THE ENDOSCOPE AND INTO THE PT'S COLON. THE FLUID WAS IDENTIFIED AS ENZYMATIC CLEANER. THERE WAS NO PT INJURY REPORTED. THE PT IS DOING FINE. OLYMPUS FOLLOWED UP WITH THE REPORTER TO OBTAIN ADD'L INFO AND WAS INFORMED THAT THE COLONOSCOPE WAS MANUALLY SOAKED WITH MCKESSON MULTI-ENZYMATIC CLEANSER, AND THEN REPROCESSED IN THE AUTOMATED ENDOSCOPE REPROCESSOR (AER). HOWEVER THE AUXILIARY WATER CHANNEL WAS NOT CONNECTED TO THE AER, HENCE IT WAS NOT FLUSHED. THE USER FACILITY HAS TAKEN CORRECTIVE ACTION AND NOW TESTS AND FLUSHES THE AUXILIARY WATER PORT PRIOR TO A PROCEDURE. THE DEVICE WAS NOT RETURNED TO OLYMPUS FOR EVAL. HOWEVER, THE USER FACILITY'S REPROCESSING PRACTICES COULD NOT BE RULED OUT AS A CONTRIBUTORY FACTOR TO THE REPORTED EVENT. | 0
3471593 | THE USER FACILITY REPORTED THAT DURING A COLONOSCOPY PROCEDURE A CLIP THAT WAS USED ON A PREVIOUS PT WAS EXPELLED THROUGH THE TIP OF THE SAME COLONOSCOPE AND FELL INTO THE PT. IT IS UNK IF THE CLIP WAS RETRIEVED FROM THE PT. THE PROCEDURE WAS COMPLETED WITH THE SAME DEICE. THERE WAS NO PT INJURY REPORTED. THE DEVICE REFERENCED IN THIS REPORT WAS RETURNED TO OLYMPUS FOR EVAL. THE INTERNAL CHANNELS OF THE DEVICE WAS EXAMINED WITH THE USE OF A BORESCOPE AND RIGID TELESCOPE AND FOUND MINOR SCRAPES AND SCRATCHES IN THE BIOPSY CHANNEL NEAR THE DISTAL END. THERE WAS YELLOWISH RESIDUE BUILD-UP IN THE SUCTION CYLINDER UNIT. THE DEVICE PASSED THE LEAK TEST. THERE WERE NICKS AND DENTS NOTED ON THE DISTAL END COVER. THERE WERE BROKEN FRAYED WIRES ON THE BENDING SECTION MESH. THE DEVICE WAS SERVICED AND RETURNED TO THE USER FACILITY. AN OLYMPUS ENDOSCOPY SUPPORT SPECIALIST (ESS) OFFERED A REPROCESSING IN-SERVICE TO THE USER FACILITY, BUT HAS NOT HEARD BACK FROM THE USER FACILITY. THE EXACT CAUSE OF THE USER'S EXPERIENCE COULD NOT BE CONCLUSIVELY DETERMINED, BUT INSUFFICIENT CLEANING OF THE DEVICE COULD NOT BE RULED OUT AS A CONTRIBUTORY FACTOR TO THE REPORTED EVENT. | 0
3181356 | SURGEON NEEDED TO USE CYSTOSCOPE. NO IMAGE WAS AVAILABLE ON THE VIDEO MONITOR. EQUIPMENT AND INSTRUMENT CONNECTIONS WERE VERIFIED. NOTED THAT TOWER IS DEFECTIVE AND NECESSITATES TURNING THE MONITOR OFF AND ON TO OBTAIN AN IMAGE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3144877 | THIS INSTRUMENT, A COMPONENT OF THE TRIGEN INTRAMEDULLARY NAILING SYSTEM, HAS BEEN USED DURING SURGICAL PROCEDURE ON 16 PTS SINCE 2011. IT WAS DISCOVERED THAT THE INSTRUMENT COULD BE DISMANTLED FURTHER THAN PREVIOUSLY KNOW FOR MORE EXTENSIVE CLEANING. | 0
2880254 | THE USER FACILITY REPORTED AN EMPLOYEE FELL WHILE OPERATING AN EQUIPMENT MANAGEMENT SYSTEM (EMS). THE CUSTOMER REPORTED THE EMS WAS DIFFICULT TO MANEUVER, SO THE EMPLOYEE PULLED THE BRAKE HANDLE LOCATED ON THE PANEL WITH FULL FORCE, CAUSING THE PANEL TO COME OFF AND THE EMPLOYEE AND PANEL TO FALL TO THE FLOOR. THE EMPLOYEE REPORTED PAIN IN THE RIGHT BUTTOCK. THE EMPLOYEE WAS GIVEN A PAIN KILLER AFTER THE EVENT AND WENT HOME. THE EMPLOYEE IS NOW BACK TO WORK AND AT PRESENT DOES NOT SHOW ANY INDICATION OF LASTING EFFECTS FROM THE EVENT. THE CUSTOMER DECLINED TO PROVIDE FURTHER DETAILS ON TREATMENT CITING HIPAA. A STERIS SERVICE TECHNICIAN INSPECTED THE UNIT AND FOUND THE PANEL TAPED TO THE UNIT. THE PANEL WAS TAPED TO THE UNIT AFTER THE REPORTED EVENT OCCURRED. THE TECHNICIAN PROPERLY REATTACHED THE PANEL. UPON INSPECTION, THE TECHNICIAN FOUND THAT THE LOWER BRAKE RELEASE SWITCH WAS WORKING INTERMITTENTLY. A NEW LOWER BRAKE RELEASE SWITCH WAS INSTALLED AND BOTH UPPER AND LOWER BRAKES WERE TESTED AND VERIFIED TO BE OPERATIONAL. THE CUSTOMER HAD PREVIOUS SERVICE CALLS RELATED TO THE BRAKES ON THIS UNIT WHERE THE BRAKE BLADDER WAS 'BLOWN'. UPON HIS INSPECTION OF THE UNIT, THE TECHNICIAN FOUND DAMAGE TO THE LOWER BRAKE RELEASE SWITCH. DAMAGE TO THE LOWER BRAKE RELEASE SWITCH CAN OCCUR IF THE FACILITY IS INCORRECTLY REPOSITIONING THE UNIT WITHOUT DISENGAGING THE PNEUMATIC BRAKE. THIS MAY HAVE RESULTED IN THE SWITCH PERFORMING INTERMITTENTLY. THE OPERATOR MANUAL STATES (PP 3-5): THE CS DF80 SYSTEM HAS A PNEUMATIC BRAKE. THIS BRAKE IS CONSTANTLY ENGAGED. NEVER REPOSITION THE UNIT WITHOUT DISENGAGING THE BRAKE FIRST. PRESS THE BRAKE BUTTON TO DISENGAGE THE BRAKE. CAUTION: AVOID DAMAGE TO PNEUMATIC BRAKES. RELEASE BRAKE BUTTONS ONLY AFTER SYSTEM HAS COME TO A COMPLETE STANDSTILL". A STERIS ACCOUNT MANAGER WILL CONTACT THE FACILITY TO OFFER AN IN-SERVICE ON HOW TO CORRECTLY MANEUVER THE UNIT AS WELL AS THE IMPORTANCE OF ENGAGING AND PROPERLY DISENGAGING THE PNEUMATIC BRAKES. | 0
3010738 | UPON ROUNDING AND ASSESSMENT OF PT, RN NOTED CLEAR DRAINAGE IN EVD CONTAINER VENTRICULAR DRAIN AND THEN NOTED FLUIDS IV DILANTIN WERE INFUSING THROUGH THE EVD TO THE PT. FLUIDS WERE STOPPED AND NEUROSURGEON IMMEDIATELY NOTIFIED. DRAIN WAS FLUSHED PER PHYSICIAN; NO CHANGE IN NEURO STATUS NOTED. RN STATED HE DID NOT TRACE LINE AND INADVERTENTLY CONNECTED IV DILANTIN INFUSION TO THE STOPCOCK ON THE EVD TUBING. THE EVD TUBING DID HAVE THE BLUE LINE AND WAS LABELED APPROPRIATELY. INTERIM MEASURE INITIATED: PORTS ON EVD TUBING WERE TAPED OFF; LOOKING INTO PURCHASING MEDTRONIC DELTA VALVES TO PLACE ON PORTS THAT WOULD ALLOW ASPIRATION BUT NOT INJECTION. DATES OF USE: (B)(6) 2013 (10 DAYS). | 0
3311022 | THE ISSUES INVOLVES DISPLAY FUNCTIONALITY WITHIN THE CERNER FETALINK SOFTWARE SOLUTION AND OCCURS SPORADICALLY WHEN THE APPLICATION IS IN LIVE VIEW. THE FETAL HEART RATE WAVEFORM DISPLAYED WITHIN THE APPLICATION'S LIVE VIEW SPORADICALLY STOPS DISPLAYING DATA TO THE USER. DATA IS NOT ACTUALLY LOST AND ALARMS FOR HIGH AND LOW OUT OF RANGE VALUES CONTINUE TO SOUND. THE PATIENT'S VITAL SIGNS, SUCH AS HEART RATE, CONTINUE TO BE DISPLAYED. ADDITIONALLY, THE ACTUAL PHYSICAL MONITORS THAT CERNER FETALIK SUPPLEMENTS ARE NOT AFFECTED BY THIS ISSUE AND WILL CONTINUE TO PROVIDE ONGOING PATIENT DATA, INCLUDING THE FETAL HEART RATE AND CORRESPONDING PAPER STRIP, IF UTILIZED. THE USER CAN RETRIEVE THE DISPLAY OF THE FETAL HEART RATE WAVEFORM WITHIN CERNER FETALINK BY LOGGING OUT OF THE APPLICATION AND LOGGING BACK IN. NOT ALL CERNER FETALINK CLIENT SITES ARE AFFECTED BY THIS ISSUE AND NOT ALL USERS OR PATIENTS AT AFFECTED CLIENT SITES EXPERIENCE THIS ISSUE; IT HAPPENS SPORADICALLY. CERNER HAS NOT RECEIVED ANY REPORTS TO DATE OF PATIENT INJURY OR DEATH IN RELATION TO THIS ISSUE. ALTHOUGH CLINICIANS WILL BE EVALUATING MONITORED PATIENTS' DATA AT REGULAR INTERVALS AND WOULD BE AWARE THAT THE FETAL HEART RATE WAVEFORM IS MISSING, PERIODIC CHANGES TO THE FETAL HEART RATE TRACING THAT SIGNAL A CHANGE IN PATIENT CONDITION, BUT DO NOT TRIGGER ALERTING, AS WELL AS TRENDS OVER TIME, COULD POTENTIALLY BE MISSED BY USERS AS A RESULT OF THIS ISSUE. CARE COULD POTENTIALLY BE DELAYED WHILE THE CLINICIAN RE-LAUNCHES THE APPLICATION IN ORDER TO RETRIEVE THE WAVEFORM DISPLAY. CERNER IS WORKING CLOSELY WITH CLIENT SITES AFFECTED BY THIS ISSUE AND HAS PROVIDED A DIAGNOSTIC PACKAGE TO AID IN FURTHER TROUBLESHOOTING. UNTIL A SOFTWARE MODIFICATION IS AVAILABLE TO ADDRESS THIS ISSUE, CERNER WILL PUBLISH A PRIORITY REVIEW FLASH TO ALERT OTHER CLIENTS TO THE ISSUE AND PROVIDE THE CURRENTLY AVAILABLE WORKAROUNDS THAT CAN BE USED TO RETRIEVE THE FETAL HEART RATE WAVEFORM DISPLAY. CERNER CORPORATION WILL PROVIDE A FOLLOW-UP REPORT WHEN THE SOFTWARE MODIFICATION IS AVAILABLE. | 1
3378722 | OUR AFFILIATE IS REPORTING TO US THAT THE SURGEON NOTICED AT THE END OF THE PROCEDURE, EXCESSIVE FLUID BUILDUP WITHIN THE BREAST AND LOWER NECK AREA ON TWO PATIENTS ON THE SAME DAY WITH THE SAME FMS PUMP. THESE PROCEDURES WERE ARTHROSCOPIC SAD¿S. THE DEVICE HAS BEEN SENT TO A SERVICE CENTER. THERE IS NO FURTHER INFORMATION MADE AVAILABLE TO MITEK. ALSO SEE ASSOCIATED MDR 1221934-2013-00272. THE DEVICE WAS RECEIVED AND WAS THOROUGHLY EVALUATED VISUALLY, ELECTRICALLY AND FUNCTIONALLY. VISUALLY, IT IS NOTED THAT THE COMPLAINT DEVICE ARRIVED IN GOOD CONDITION, NO ANOMALIES NOTED. THE DEVICE WAS THEN TESTED FOR ELECTRICAL AND FUNCTIONAL CONFORMANCE; THE PUMP PASSED ALL DIAGNOSTIC AND FUNCTIONAL TESTS WITHIN ITS DESIGNED AND MANUFACTURED PARAMETERS; NO FAULTS NOTED. WE CANNOT TELL ANYTHING FROM THIS; CANNOT DISCERN THE UNDERLYING OR ROOT CAUSE FOR THE REPORTED ISSUE. AT THIS POINT IN TIME, NO CORRECTIVE OR FURTHER ACTION IS WARRANTED, HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FUTURE INFORMATION RECEIVED THAT IS RELATIVE AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 0
3285525 | OUR REP IS REPORTING TO US THAT DURING AN ARTHROSCOPIC KNEE PROCEDURE (SCOPE, PARTIAL MEDIAL MENISECTOMY) WITH THE USE OF AN FMS SOLO PUMP FOR FLUID MANAGEMENT, THERE WAS SOME EXTRAVASATION THAT WAS NOTICED AT THE END OF THE SURGERY. ALTHOUGH THE PROCEDURE WAS CONCLUDED SUCCESSFULLY, WHEN THEY WERE APPLYING DRESSING, THEY NOTED THAT THERE WAS SWELLING (EXTRAVASATION) DOWN TO THE BOTTOM OF THE CALF. MEASUREMENTS OF THE AFFECTED AREA WERE TAKEN, LEG WAS ELEVATED AND ICED WHICH RENDERED NOTICEABLE DECREASE IN THE EDEMA, CLOSE OBSERVATION DURING POST ¿OP RECOVERY AND FOLLOW-UP PHONE CALLS WITH PATIENT WERE CONDUCTED WITH THE PATIENT REPORTING CONTINUED DECREASE IN CALF EDEMA. FOR PROGNOSIS, THEY ANTICIPATE NO ADVERSE EFFECTS. THE COMPLAINT DEVICE WAS RECEIVED, AND WAS EVALUATED AND TESTED EXTENSIVELY. VISUALLY IT IS NOTED THAT THE DEVICE WAS RECEIVED IN GOOD COSMETIC CONDITION, OTHER THAN SOME MINOR CHASSIS SCRATCHES, THERE WAS NO APPARENT PHYSICAL DAMAGE OR ANOMALIES. FUNCTIONALLY AND ELECTRONICALLY, THE DEVICE FUNCTIONED WELL WITHIN ITS DESIGN AND MANUFACTURED PARAMETERS; COULD FIND NO FAULT WITH THE FMS PUMP. WE CANNOT DISCERN THE UNDERLYING CAUSE FOR THE REPORTED ISSUE; THE ROOT OR UNDERLYING CAUSE FOR THE REPORTED EVENT IS NOT ATTRIBUTABLE TO THE FMS PUMP, BUT LIES ELSEWHERE. AT THIS POINT IN TIME, NO FURTHER ACTION IS WARRANTED; HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FORTHCOMING INFORMATION RECEIVED THAT IS PERTINENT AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD.><ATTEMPTS HAVE BEEN MADE TO RETRIEVE ADDITIONAL INFORMATION ABOUT THE EVENT AND DEVICE. THE ADDITIONAL INFORMATION WILL REPORTEDLY BE FORWARDED TO DEPUY MITEK HOWEVER IS NOT KNOWN IF IT WILL BE RECEIVED WITHIN THE 30 DAY REPORTING REQUIREMENT, THEREFORE, DEPUY MITEK WOULD LIKE TO FILE THIS INITIAL MEDWATCH REPORT AT THIS TIME. WHEN AND IF ADDITIONAL INFORMATION IS RECEIVED, IT WILL BE REFLECTED IN A FOLLOW-UP MEDWATCH REPORT. DEVICE EVAL: AWAITING RETURN. | 0
2958737 | DURING A LEFT HEART CATH THE SALINE TUBING FROM THE MEDRAD AVANTA FLUID MANAGEMENT SYSTEM WHEN FLUSHED WITH SALINE WOULD NOT WORK/WOULD NOT FLUSH. THE TUBING HAD TO BE CHANGED OUT DURING THE CASE. THE NEW TUBING WAS FROM THE SAME LOT NUMBER, AND IT WORKED FINE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?LEFT HEART CATH. | 0
2954014 | DURING SET UP, IT WAS NOTICED THAT THE TUBING WAS KINKED IN TWO PLACES WHICH WOULD NOT ALLOW THE CONTRAST TO TRAVEL THROUGH. ONE PLACE WAS KINKED RIGHT AT THE RESERVOIR AND THE OTHER KINK WAS WHERE THE CONTRAST EXCHANGES THROUGH THE TUBING. TRIED TO STRAIGHTEN TUBING OUT, BUT WAS UNABLE. | 0
3347799 | IT WAS REPORTED THAT WHILE A TECHNOLOGIST WAS INSTALLING A MAGNIFICATION STAND ON THE SYSTEM, SHE HEARD THE GRID ENGAGE MOTOR NOISE, SO SHE DID A LITTLE TUG ON THE MAGNIFICATION STAND TO TEST IF IT WAS ON. IT WASN'T AND IT STARTED TO FALL. SHE DIDN'T WANT IT TO HIT THE FLOOR, SO SHE TRIED STOPPING IT. WHEN IT FELL IT HIT HER THIGH AND LOWER LEG. SHE HAS A HUGE HEMATOMA ON HER THIGH AND MANY BRUISES. SHE APPLIED ICE TO HER THIGH AND THEN PROCEEDED TO CONTINUE WORKING. NO PATIENT WAS INVOLVED IN THIS EVENT. THE MAGNIFICATION STAND WAS EVALUATED BY AN HOLOGIC FIELD ENGINEER. THE MAGNIFICATION ASSEMBLY WAS INSPECTED FOR MECHANICAL AND LEVER/SWITCH FAILURES. THE MAGNIFICATION STAND WAS FOUND TO BE WORKING PROPERLY. THE MAGNIFICATION STAND WAS MOUNTED AND REMOVED SEVERAL TIMES LOCKING IN PLACE EACH TIME. WE CAN ONLY CONCLUDE THAT THE TECHNOLOGIST DID NOT PROPERLY LOCK THE MAGNIFICATION STAND TO THE SYSTEM BEFORE TUGGING ON IT, CAUSING IT TO FALL OFF. | 0
3052985 | IT WAS REPORTED THAT DURING A BREAST EXAMINATION ON THE MAMMOMAT INSPIRATION SYSTEM THE UNIT FAILED DUE TO A DEFECTIVE FOOTSWITCH. THE TECH PRESSED THE EMERGENCY RELEASE BUTTON BUT DID NOT HAVE ENOUGH STRENGTH TO ACTIVATE THE BUTTON AND WAS UNABLE TO RELEASE COMPRESSION. THE TECH HAD TO LEAVE THE COMPRESSED PATIENT IN THE EXAM ROOM AND GET HELP FROM HER COLLEAGUES TO RELEASE THE PATIENT. AS A RESULT THE PATIENT COMPLAINED OF A SORE BREAST BUT NO MEDICAL INTERVENTION WAS REQUIRED. THE REPORTED FOOTSWITCH WAS EXCHANGED AT THIS FACILITY AND THE REPLACED PART WAS SENT TO THE FACTORY FOR FURTHER INVESTIGATION. IT IS YET UNCLEAR WHY THE USER WAS UNABLE TO PRESS THE EMERGENCY RELEASE BUTTON. THE INVESTIGATION IS ON-GOING. MAXIMUM COMPRESSION FORCE APPLIED BY THE MOTOR IS 200N. PRESSING EMERGENCY BUTTON WILL IMMEDIATELY STOP ALL DEVICE MOVEMENT. FORCE TO PRESS THE EMERGENCY MECHANICAL RELEASE BUTTON ON TOP OF THE COMPRESSION UNIT IS SPECIFIED WITH A MAXIMUM VALUE OF 100N, DEPENDING ON THE APPLIED COMPRESSION FORCE. CHAPTER "SYSTEM OVERVIEW", PAGE 8, IN THE OPERATOR MANUAL (SPB7-330.620.01.02.02) DESCRIBES MANUAL RELEASE OF COMPRESSION. AN ADDENDUM TO THE OPERATOR MANUAL (XPW7-330.621.02.01.02) DESCRIBING IN DETAIL USE OF THE EMERGENCY MECHANICAL RELEASE WAS DISTRIBUTED WITH UPDATE INSTRUCTION XP043/12/P. (B)(6). | 0
3054476 | IT WAS REPORTED TO (B)(4) THAT DURING A PROCEDURE, TWO GENERATORS FAILED. THE SURGEON HAD NO TONE DOWN AND NO COAGULATION ON BOTH DEVICES. THE SURGEON TRIED A THUNDERBEAT DEVICE WHICH FUNCTIONED FINE, BUT THEN DECIDED BECAUSE OF THE PATIENT'S ANATOMY TO SWITCH TO AN OPEN PROCEDURE. THERE WAS NO PATIENT HARM. (SEE MFR REPORT #9617070-2013-00008, FOR SECOND GENERATOR). THE UNIT WAS SERVICED AND REPAIRED. THE UNIT WAS INSPECTED AND TESTED, DURING TESTING 200 REF 53 AND 200 REF 36 POPPED-UP DUE TO FAULTY PLASMABLEND BOARD. ALSO THE UNIT IS RUNNING ON OLD VERSION SOFTWARE AND THE LEFT HAND HANDLE WAS BROKEN ON THE INSIDE. THE UNIT NEEDS TO REPLACE THE PLASMABLEND, FRONT PANEL, FULL CALIBRATION, UPGRADE TO NEW VERSION SOFTWARE, AND TESTED. | 0
3242318 | DURING TRANSHIATAL ESOPHAGECTOMY UNDER GENERAL ANESTHESIA, ANESTHESIA MACHINE MONITOR SUDDENLY WENT BLACK GIVING NO INFORMATION ON VENTILATOR AND ANESTHETIC GASES NOR ABILITY TO CHANGE SETTINGS. VENTILATOR BELLOWS CONTINUED TO RISE AND FALL. WE SWITCHED FROM ANESTHESIA VENTILATOR TO AMBU BAG W/ PEEP VALVE, AND CHANGED PRIMARY ANESTHETIC FROM INHALED ISOFLURANE TO IV PROPOFOL. WE WERE ABLE TO TURN OFF VENTILATOR AND REMOVE ISOFLURANE CASSETTE. BIOMED PAGED STAT WHO CAME AND QUICKLY DETERMINED THAT A NEW ANESTHESIA MACHINE WOULD NEED TO REPLACE MALFUNCTIONING MACHINE. BIOMED REMOVED EXISTING MACHINE AND BROUGHT IN NEW MACHINE. PATIENT REMAINED STABLE AND CASE WAS COMPLETED WITH REPLACEMENT MACHINE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TRANSHIATAL ESOPHAGECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2900842 | ANESTHESIA TECH NOTICED MOLD IN THE EXPIRATORY SIDE OF THE UNIT 12 HOURS AFTER IT WAS LAST USED ON A PATIENT. IT WAS DETERMINED THAT THE GROWTH IS A RESULT OF LONG CASES WITH LOW FLOW. ============MANUFACTURER RESPONSE.============MANUFACTURER RECOMMENDS THAT INSPIRATORY-EXPIRATORY VALVE MANIFOLD BE DISASSEMBLED AND AUTOCLAVED QUARTERLY. THE MANUFACTURER ALSO RECOMMENDS VISUALLY INSPECTING THE VALVES MONTHLY AND POSSIBLY FLUSHING THE SYSTEM WITH O2 OR AIR AT THE END OF EACH DAY TO MINIMIZE THE POSSIBILITY OF MOLD GROWTH. | 0
3329473 | THE PATIENT WAS BEING INDUCED UNDER ANESTHESIA WHEN THE CRNA AND MDA WENT TO VENTILATE THE PATIENT THEY WERE UNABLE TO SO BECAUSE THERE WAS AN AIR LEAK. THE VALVE ON THE CONDENSATION TRAP APPEARS TO HAVE STUCK OPEN THUS ALLOWING AIR TO ESCAPE; PREVENTING PROPER VENTILATION. STAFF REVIEWED THE NORMAL PROCESS AND ALGORITHM AND FOUND NO PRACTICE DEVIATIONS FROM THE STANDARD WORK. AS A RESULT OF THIS INCIDENT THEY HAD TO WAKE THE PATIENT, CHANGED ROOMS AND EQUIPMENT AND RE-INDUCED THE PATIENT. | 0
3111698 | CERVICAL SPECIMEN WAS ANALYZED FOR (B)(6) ON THE (B)(4) GENEXPERT SYSTEM. (B)(6) RESULT ANALYZED AS (B)(6). RESULT CALLED TO PROVIDER. PROVIDER QUESTIONED RESULT AND REQUEST TESTING. UPON REPEAT TESTING USING THE SAME SAMPLE, RESULT WAS (B)(6). MICROBIOLOGY SUPERVISOR CONTACTED VENDOR, (B)(4), AND ASKED THEM TO INVESTIGATE THE REASON FOR THIS DISCREPANCY. (B)(4) EXPLAINED THAT THIS WAS DUE TO A SOFTWARE GLITCH. THE INITIAL RESULT SHOULD HAVE BEEN CALLED INVALID BY THE SOFTWARE. HOWEVER, IT WAS CALLED (B)(6) BY THE SOFTWARE. VENDOR STATES THAT THEY SAW A FEW EXAMPLES OF THIS DURING THEIR VALIDATION AND THAT THEY ARE WORKING ON UPDATING THE SOFTWARE MOVING FORWARD. VENDOR RECOMMENDED ADDITIONALLY CHECKING THE RESULT CURVE TO COMPARE TO THE (B)(6) RESULT AND THE ENDPOINT RESULT TO VERIFY RESULTS. DATES OF USE: (B)(6) 2013. | 0
2982677 | DATE OF REPLY: SOURCE (B)(6), GIRAFFE AND PANDA T-PIECE RESUSCITATION SYSTEMS AND MASK RESUSCITATION SYSTEMS: CLASS 1 RECALL - POTENTIAL REVERSAL OF OXYGEN/AIR: MEDICAL LOGISTICS CONDUCTED A SEARCH ON THE EQUIPMENT DATABASE. THIS PARTICULAR EQUIPMENT IS NOT IN THE (B)(6) INVENTORY. | 0
3058180 | HEMOCUE (B)(4) HAS REC'D TWO COMPLAINTS CONCERNING THAT HEMOCUE HB 201 WAS MEASURING HIGHER THAN EXPECTED, COMPLAINTS WERE REC'D (B)(6) 2013. NO PT IMPACT WAS REPORTED AN NO MEASUREMENT RESULTS FROM THE HB 201 SYSTEM OR COMPARATIVE METHOD WERE PROVIDED BY THE CUSTOMERS. AT THIS POINT, NO MDR WAS NEEDED BASED ON INITIAL RISK ASSESSMENT PERFORMED ON INFO FROM CUSTOMERS. DURING INTERNAL INVESTIGATION OF THE COMPLAINTS IT WAS FOUND THAT THE HIGHER MEASUREMENTS WAS DUE TO DEFORMATION OF THE LID OF THE VIAL WHERE THE MICROCUVETTES ARE STORED. THIS COULD POTENTIALLY LEAD TO INCORRECTLY HIGH MEASUREMENT RESULTS WHEN ANALYZING HIGHLY REDUCED BLOOD. THE DEVIATION IN MEASUREMENT COULD BE UP TO +1.1 G/DL AT A HB LEVEL OF 8 G/DL, ACCORDING TO THE INTERNAL INVESTIGATION. THE INVESTIGATION WAS FINALIZED (B)(6) 2013 AND IT WAS BASED ON THE RESULTS FROM INVESTIGATION THE MDR REPORTING WAS DECIDED. ACCORDING TO THE RISK ASSESSMENT AND CLINICAL IMPACT ANALYSIS, AN INCORRECT MEASUREMENT AS DESCRIBED ABOVE COULD POTENTIALLY LEAD TO A MISSED MEDICAL INTERVENTION, I.E. A BLOOD TRANSFUSION WHEN NECESSARY. HOWEVER, IT IS HIGHLY UNLIKELY TO OCCUR SINCE IT ONLY AFFECTS VENOUS BLOOD SAMPLES WITH EXTREMELY LOW OXYGENATED HEMOGLOBIN. IT IS RARE THAT A BLOOD TRANSFUSION IS INITIATED BASED ON A SINGLE POINT-OF-CARE HEMOGLOBIN MEASUREMENT RESULT, BUT IT COULD OCCUR IN SURGERY ROOM ON PTS WITH MASSIVE BLEEDING. A ROOT CAUSE INVESTIGATION OF THE DEFORMATION OF THE LID ON THE IDENTIFIED UNITS IS ONGOING. EVAL SUMMARY: PROBLEM DEFINITION FOR INVESTIGATION (CUSTOMER COMPLAINT): THE HEMOCUE HB (B)(4) MEASUREMENT RESULTS WERE HIGHER THAN EXPECTED WHEN TESTED WITH A MICROCUVETTES FROM THE VIAL. THE CUSTOMER FOUND THAT THE LID WAS DEFORMED. INVESTIGATION: REVIEW OF BATCH DOCUMENTATION (DHR): NO DEVIATIONS FOUND. RESULTS FROM INVESTIGATION (PRODUCTION): THE DAMAGE PROBABLY OCCURRED WHEN THE LID WAS MOUNTED ON THE VIAL. RESULTS FROM THE INVESTIGATION (PRODUCT SUPPORT SYSTEM): THE INNER PART OF THE LID IS PARTLY DEFORMED. CONCLUSION: THE HIGHER MEASUREMENT RESULTS ARE PROBABLY DUE TO THE DEFORMATION OF THE LID OF THE VIAL. ROOT CAUSE INVESTIGATION IS HANDLED WITHIN A CAPA. | 0
2822633 | SYNTHES CONSULTANT REPORTED: MATRIX NEURO SET LID HAS PEELING GRAPHICS. THE PEELED GRAPHICS ARE REPORTEDLY GETTING INTO THE PLATES AND SCREWS WHEN WASHED. THE GRAPHIC CASE WAS ADDED TO THIS COMPLAINT AS A REPORTABLE EVENT DUE TO; GRAPHICS REPRESENT THE POTENTIAL FOR UNRETRIEVED DEVICE FRAGMENTS OF NON-IMPLANT GRADE MATERIAL. THIS IS 3 OF 3 REPORTS. WITHOUT A LOT NUMBER THE DEVICE HISTORY RECORDS REVIEW COULD NOT BE COMPLETED. THE INVESTIGATION COULD NOT BE COMPLETED, NO CONCLUSION COULD BE DRAWN, AS NO PRODUCT WAS RECEIVED. | 0
2911631 | DURING PREVENTATIVE MAINTENANCE, THE IMAGE ON THE CENTRAL CONTROL MONITOR OF THE HEART LUNG CONSOLE WOULD MOVE UP AND DOWN INTERMITTENTLY. THERE WAS NO ADVERSE CONSEQUENCE TO THE PT AS A RESULT OF THIS EVENT. EVAL IN PROGRESS BUT NOT CONCLUDED. | 1
2941461 | I HAD MY MCKESSON HEATING PAD ON THE SMALL OF MY BACK. I STARTED SMELLING SOMETHING BURNING. I NOTICED THAT A WHOLE BURNT THROUGH THE BACK OF MY HEATING PAD THROUGH THE COVER AND INTO MY COUCH. RETAILER: (B)(6) PURCHASE DATE: (B)(6) 2012, THIS DATE IS AN ESTIMATE. THE PRODUCT WAS NOT DAMAGED BEFORE THE INCIDENT. THE PRODUCT WAS NOT MODIFIED BEFORE THE INCIDENT. DOCUMENT NUMBER: (B)(6), REPORT NUMBER: 20121124-EF7B7-2147461630. | 0
3464456 | NO PT INVOLVEMENT. PHILIPS HEMODYNAMIC PT MONITOR CAUGHT FIRE IN VACANT PT ROOM WITHIN THE HOSPITAL WHILE NOT IN USE. IT APPEARS TO BE THE BATTERIES THAT CAUSE THE FIRE. THE FIRE DEPT WAS CALLED AS THERE WAS DIFFICULTY EXTINGUISHING THE FIRE. HALF OF THE PTS ON FLOOR HAD TO BE EVACUATED DUE TO SMOKE AND ODOR ISSUES. THIS WAS ALSO REPORTED TO PHILIPS MEDICAL AND ASSIGNED CASE # (B)(4). | 0
2917301 | BABY MONITORED WITH EFM AND ST ANALYSIS FROM 08.00 AM. OXYTOCIN INFUSION ONGOING, MATERNAL TEMP 38.2C. FETAL BLOOD SAMPLING PH (B)(6) AT 08.45 AM. DELIVERY TERMINATED AT 11.06 AM WITH MIDCAVITY VACUUM EXTRACTION, (B)(6). BABY RESUSCITATED FOR 30 MINUTES, NO CORD BLOOD GASES AVAILABLE. FIRST PH AT ARRIVAL AT NEONATAL UNIT: 6.30. ANTIBIOTICS GIVEN TO BABY PRIOR TO BLOOD CULTURE: NEGATIVE FOR MOTHER AND BABY. BABY DIED AT (B)(6). THE INCIDENT WAS REPORTED TO NEOVENTA OFFICE IN (B)(6). THE DEVICE REMAINS IN THE HOSPITAL AND NO FAILURE TO THE DEVICE HAS BEEN REPORTED. THERE IS NO OBVIOUS CAUSE OF THE OUTCOME IN THE EFM RECORDING. THERE ARE NO CORD GASES AVAILABLE UNTIL THE ARRIVAL TO THE NEONATAL UNIT AFTER MORE THAN 30 MIN OF RESUSCITATION: PH (B)(6) THE OBSTETRIC TEAM PERFORMED CORRECTLY ACCORDING TO GUIDELINES. SUSPICION OF INFECTION DUE TO HIGH CRP VALUES, AS ANTIBIOTICS WAS ADMINISTERED TO BABY PRIOR TO BLOOD CULTURE. THE PARENTS HAVE DENIED POST-MORTEM EXAMINATION. | 0
3168503 | IT WAS REPORTED THAT A HOMECHOICE DEVICE EXPERIENCED A SYSTEM ERROR 2240 / 2367. THIS OCCURRED WHILE THE PATIENT WAS PERFORMING PERITONEAL DIALYSIS THERAPY. THE PATIENT STATED THAT THEY DISCONNECTED PRIOR TO THE ALARM AND DID NOT USE PROPER DISCONNECT PROCEDURE. AFTER THEY RECONNECTED USING ASEPTIC TECHNIQUE, THE DEVICE ALARMED. NO PATIENT INJURY OR MEDICAL INTERVENTION WAS REPORTED. ADDITIONAL INFORMATION WAS REQUESTED AND IS NOT AVAILABLE. (B)(4). USE ERRORS AND PROPER USER INSTRUCTIONS ARE ADDRESSED IN ?THE HOMECHOICE AND HOMECHOICE PRO APD SYSTEMS PATIENT AT-HOME GUIDE? (07-19-63-293, JULY 28, 2010), WHICH IS SHIPPED WITH EVERY HOMECHOICE DEVICE. PAGE 10-25 PROVIDES INSTRUCTIONS FOR PROPERLY OPENING THE TRANSFER SET. USERS ARE NOT INSTRUCTED TO DISCONNECT DURING THERAPY UNLESS FOR AN EMERGENCY DISCONNECT PROCEDURE (SECTION 15.10). A FORMAL REVIEW OF THE LABEL FOR THE PRODUCT FAMILY WILL BE CONDUCTED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, THEN A FOLLOW UP REPORT WILL BE SUBMITTED. | 0
3101338 | VITAL SIGNS - GE- HYPERINFLATION BAG SYSTEM- 1663X3LP, LOT 6199E. UNABLE TO PRESSURIZE BAG ABOVE 5CM H2O AND THUS UNABLE TO VENTILATE PATIENTS WITH POSITIVE PRESSURE. SINCE THIS IS THE INTENDED USE OF THE DEVICE, ITS FAILURE REPRESENTS A LIFE-SAFETY HAZARD. THIS HAS BEEN FOUND TO BE A PROBLEM ON ALL INSTANCES OF THIS PRODUCT. DATE OF USE: (B)(6) 2013. REASON FOR USE: ELECTROCONVULSIVE THERAPY. | 0
3068310 | THE PT WAS IN A LUNG TREATMENT AND WAS TREATED WITH A VARIAN 2100CD, 80 MLC DIRECTLY WITH SKIN ON THE COUCH TOP. THE KIND OF TREATMENT WAS STATIC (AP/PA). DURING THE TREATMENT, THE COUCH TOP CAUSED THE PT TO GOT A INCORRECT DOSE TREATMENT AND HAS MOIST DESQUAMATION. TREATMENT REGION WAS LUNG. RADIATION ENERGY AND MODALITY OF 18 MEV X-RAY. FIELD SIZE WAS 15 X 15 CM2. PT RECEIVED A CUMULATIVE DOSE FROM 7020 CGY, THIS IS A DOSE PER FRACTION OF 180 CGY. DOES PER FIELD WAS 90 CGY. THE MEDICAL INTELLIGENCE COUCH TOP IS, LIKE E.G. SIEMENS, SINMED AND Q-FIX, A SOLID CARBON FIBRE COUCH TOP WITHOUT A TENNIS RACKET AREA. IT FULFILLS THE REQUIREMENTS OF THE ALUMINUM EQUIVALENT ACCORDING TO 21 CFR SEC 1020.30. ALL CARBON FIBER TABLETOPS RESULT IN AN INCREASE IN PT SKIN DOSE AND ALL ARE AWARE OF THE INCREASE IN DOSE. THIS DESIGN ENABLE S A HIGHER STABILITY FOR THESE COUCH TOPS AND MINIMIZES ANY SAGGING OF THE PT, WHICH CAN OCCUR WITH TRADITIONAL TENNIS RACKET INSERTS OR FOIL AND A NOTABLY IMPROVEMENT IMAGING QUALITY, COMPARED TO TENNIS RACKET TYPES. OVER THE COMPLETE TREATMENT AREA, SOLID CARBON FIBER COUCH TOPS DO NOT HAVE A HIGHER DOSE ATTENUATION COMPARED TO TENNIS RACKET AREAS. AS A CONSEQUENCE OF HIGHER ATTENUATION CAUSED BY THE SOLID CARBON FIBRE COUCH TOPS THERE IS A, IN COMPARISON, NOTABLY HIGHER SKIN DOSE. EVEN VERY THIN CARBON FIBER SOLID INSERTS, TO BE PLACED IN THE TENNIS RACKET AREA, HAVE AN NOTABLE EFFECT ON SKIN DOSE UP TO 70% RELATIVE SURFACE DOSE. SAME CAN BE NOTED WITH SEVERAL STATE OF THE ART IMMOBILIZATION DEVICES. CORRESPONDING MONTE CARLO CALCULATIONS DEMONSTRATE THAT THIS EFFECT TO SKIN DOSE CAN BE DECISIVELY COMPENSATED ONLY BY TENNIS RACKETS. THEREFORE, ALL SOLID CARBON FIBRE COUCH TOPS CAN BE STATED TO SHOW SIMILAR EFFECTS. THE EFFECT HAS TO BE CONSIDERED FOR ADAPTION OF TREATMENT TECHNIQUES. SHOULD THE USER EMPLOY THE SAME TECHNIQUES AS FOR A COUCH TOP WITH TENNIS RACKET, SKIN REACTIONS ARE MORE LIKELY TO OCCUR. OPPOSITE, COUCH TOPS WITH A TENNIS RACKET AREA HAVE MASSIVE FRAME STRUCTURES, WHICH HAVE TO BE CONSIDERED IN TREATMENT PLANNING AS WELL. THIS IMPOSES THAT COUCHTOPS MAY NOT BE TREATED "AS AIR" . | 0
3252425 | IMAGING EQUIPMENT IN THE CATH LAB FAILED DURING A PROCEDURE. PT ON THE TABLE AND PROCEDURE IN PROGRESS. JUST BEFORE FINAL IMAGING AND DEPLOYMENT OF CAROTID STENT IMAGING FAILED AND COULD NOT/WOULD NOT RE-BOOT. CASE CANCELED/RESCHEDULED. NO ADVERSE PT EVENT. | 0
3120086 | THE EXAM LIGHT ON THE GE HEALTHCARE, GIRAFFE OMNIBED INCUBATOR BROKE ON THE MOUNTING ARM NEAR THE PIVOT POINT WHEN THE LIGHT WAS ADJUSTED. THIS COULD HAVE RESULTED IN THE LIGHT FALLING ON THE BABY. OTHER LIGHTS ON THE OMNIBED INCUBATORS WERE INSPECTED AND 8 OTHER LIGHTS WERE FOUND TO BE CRACKED. THE DEFECTIVE LIGHTS WERE REMOVED FROM SERVICE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?VIEWING THE BABY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3297768 | DURING SURGERY, THE PHYSICIAN WAS TRYING TO COAGULATE A TISSUE AND NOTICED THAT THE INTUITIVE SURGICAL PK DISSECTING FORCEPS(ELECTROSURGICAL INSTRUMENT) WEREN'T WORKING. HE NOTICED THE PK WIRE FROM THE DEVICE WAS LOOSE/BROKEN. BROKEN DEVICE WAS REMOVED AND REPLACED WITH NEW ONE. NO HARM WAS NOTED TO PATIENT AND NO PARTS OF INSTRUMENT WAS LEFT INSIDE.NO INJURY TO PATIENT OR STAFF.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?GYN SURGERY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3167136 | DURING SURGERY, THE VESSEL SEALER STOPPED WORKING. THE ALARM ON THE MONITOR SAID, "BLADE MAY BE EXPOSED".NO HARM TO PATIENT.======================MANUFACTURER RESPONSE FOR VESSEL SEALER, ENDOWRIST ONE (PER SITE REPORTER).======================THE LOCAL REP AT THIS FACILITY PICKED UP PREVIOUS INSTRUMENTS REPORTED (4). HE REPORTS " USER ERROR" DESPITE DIFFERENT USERS. I HAVE YET TO RECEIVE ANY WRITTEN EVALUATION FROM THE MANUFACTURER.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?GYN PROCEDURE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3113354 | DAVINCI VESSEL SEALER MALFUNCTIONED DURING SURGERY. BLADE WOULD NOT RETRACT BACK. NO HARM WAS DONE TO PATIENT. DEVICE WAS INTACT UPON INSPECTION.MANUFACTURER REQUESTS DEVICE(S)FOR EVALUATION. THIS IS THE 4TH DEFECTIVE DEVICE. LOCAL REP IS SUPPOSED TO PICK UP PRODUCTS FOR EVALUATION.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?GYN PROCEDURE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3049343 | DA VINCI VESSEL SEALER WAS BEING USED DURING SURGERY WHEN AN ERROR CODE APPEARED ON VISION CART MONITOR: "BLADE EXPOSED". STAFF WAS UNABLE TO REUSE THE INSTRUMENT SO IT WAS TAKEN OUT AND SAVED TO BE SUBMITTED TO THE MANAGER.THIS IS THE THIRD OF THESE DEVICES WE HAVE REPORTED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC GYN SURGERY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3203851 | WHILE IN THE MIDDLE OF ROBOTIC CASE, VESSEL SEALER ALARMED AND MONITOR DISPLAYED, "VESSEL SEALER MULFUNCTION". NEW VESSEL SEALER OPENED TO CONTINUE CASE. USED ONE SEQUESTERED.======================MANUFACTURER RESPONSE FOR VESSEL SEALER FOR DA VINCI S 1, ENDOWRIST ONE (PER SITE REPORTER).======================THIS IS THE 6TH REPORTED EVENT AT OUR FACILITY WITH THIS DEVICE. THERE HAVE BEEN OTHERS NOT REPORTED. PREVIOUS DEVICES HAVE BEEN RETURNED TO MFG WITH NO RESPONSE AS YET.OUR FACILITY HAS DECIDED THAT UNTIL FURTHER NOTICE AND TO SEE IF THIS IS USER OR DEVICE ERROR, THE DEVICE REP WILL BE IN ALL PROCEDURES USING THE DEVICE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?GYN PROCEDURE. | 0
3034411 | VESSEL SEALER HANDPIECE MALFUNCTIONED WITH ERROR CODE. BLADE EXPOSED ON VISION CART DURING PROCEDURE. TWO VESSEL SEALERS WERE OPENED WITH THE SAME ERROR CODES AND WERE UNABLE TO BE USED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED LAPAROSCOPIC SUPRACERVICAL HYSTERECTOMY ANDBILATERAL SALPINGO-OOPHORECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3349349 | WHEN SURGEON WAS ABOUT TO USE THE DEVICE, MACHINE ALARMED AND THE MONITOR SHOWED: "BLADE EXPOSE". DEVICE REPLACED WITH NEW ONE. | 0
3284649 | PATIENT DEVELOPED CANALICULITIS AFTER INSERTION OF AN OASIS MEDICAL FORM FIT HYDROGEL CANALICULAR PLUG. REPORTED EVENT: (B)(6) FEMALE HAD FORM FIT PLUGS INSERTED IN THE UPPER AND LOWER PUNCTUM IN (B)(6) 2013 WHO PRESENTED TO THE OFFICE WITH MUCUS DISCHARGE AND MILD PAIN IN (B)(6) 2013. PRODUCT DESCRIPTION: OASIS MEDICAL FORM FIT HYDROGEL PLUG. PRODUCT REFERENCE: 6303. PRODUCT LOT: NOT REPORTED. DATE OF INITIAL PROCEDURE: (B)(6) 2013. DATE OF COMPLICATION: (B)(6) 2013. DATE REMOVED: (B)(6) 2013. DATE REPORTED TO OASIS: (B)(6) 2013. DISCUSSION: THE PUNCTUM WAS IRRIGATED IN THE OFFICE WITH FULL RESOLUTION IN (B)(6) 2013 AFTER ANTIBIOTIC AND STEROID TREATMENT WAS UNSUCCESSFUL. THE DEVICE WAS NOT RETURNED FOR EVALUATION. NO CONCLUSION CAN BE DRAWN AS TO THE DIRECT CAUSE OF THE REPORTED ADVERSE EVENT. | 0
3129821 | IT WAS REPORTED THAT THE DEVICE FAILED IN THE MIDDLE OF THE PROCEDURE. IT DID NOT HAVE ANY IMAGE DISPLAY ON THE MONITOR. THE USER PERFORMED ALL THE RECOMMENDED CHECKS, AND WAS UNABLE TO IMPROVE THE CONDITIONS. THE PT WAS SEDATED THROUGHOUT THE TROUBLESHOOTING PERIOD. THE PROCEDURE WAS EVENTUALLY ABORTED DUE TO THE CAMERA UNIT THAT FAILED WITHOUT ANY PT INJURY. AT THE TIME OF THIS REPORT, MULTIPLE ATTEMPTS TO OBTAIN FURTHER INFO FROM THE HOSPITAL HAVE BEEN UNSUCCESSFUL, AND THE DEVICE HAS NOT YET BEEN RETURNED FOR EVALUATION. AS A RESULT, A DETERMINATION CANNOT BE MADE AT THIS TIME. IF FURTHER INFO BECOMES AVAILABLE, GYRUS SCMI WILL CONTINUE THE INVESTIGATION AND UPDATE THE AGENCY ACCORDINGLY. | 1
2770178 | IT WAS REPORTED THAT DURING A URETEROSCOPY PROCEDURE USING THE DIGITAL FIBEROPTIC SCOPE. THE CCU STOPPED MIDWAY THROUGH THE PROCEDURE AND THE HOSP HAD TO PROCEDURE ANOTHER CONTROLLER FROM NEW HAMPSHIRE TO COMPLETE THE PROCEDURE. THE PT WAS AWAKENED FROM ANESTHESIA AS IT TOOK 1 1/2 HOURS TO PROCURE THE NEW CONTROLLER. THE PT WAS PLACED BACK UNDER ANESTHESIA AND THE PROCEDURE WAS COMPLETED WITH NO FURTHER PROBLEMS OR HARM TO THE PT. THE CUSTOMER'S COMPLAINT OF THE INSTRUMENT NOT WORKING WAS UNCONFIRMED. THE EVAL OF THE RETURNED IDC-1500 FOUND THE SCOPE HAS AN OLDER SOFTWARE VERSION OF 4.20.7 INSTALLED. THE SCOPE RAN CONNECTED WITH THE TEST DUR-D FOR ONE HOUR WITH NO INDICATION OF LOSS OF IMAGE. | 1
3427183 | LOW OXYGEN LEVEL ALARM, UNIT WAS REPAIRED BY GE HEALTHCARE IN (B)(6) 2013, OXYGEN CELL WAS REPLACED BY THE REPAIR DEPOT. THERE ARE CONSISTENT PROBLEMS WITH THESE VENTILATORS. THE BATTERIES DO NOT HOLD UP UNDER NORMAL USE. THEY REQUIRE REPEATED RECALIBRATION TO KEEP THE O2 CELL FROM DRIFTING. THE BATTERY CELL AND BATTERY GAUGE THAT MONITORS THE CHARGE REQUIRES DISCHARGE AND RECALL EVERY 3 MONTHS. FOR A LIFE SUPPORT DEVICE IT IS NOT DEPENDABLE. [MANUFACTURER: VERSAMED/GE]. | 0
2913061 | ISSUE: WHEN USING THE ROCHE COBAS E411 (RCHCE4LI) DRIVER VERSION 8.00.0001 OR 8.00.0002, WITH INSTRUMENT MANAGER, IF THERE IS A BARCODE READ ERROR, THE CURRENT SAMPLE MAY BE IDENTIFIED AS THE LAST SAMPLE ID THAT WAS IN THAT RACK/DISK POSITION, POTENTIALLY CAUSING SAMPLE MISMATCH. IF THE LAST SAMPLE HAD ANY PENDING TESTS IN THE ORDERS DATABASE, TEST SELECTION WOULD BE DOWNLOADED TO THE INSTRUMENT WITH THE INCORRECT SAMPLE ID, AND RESULTS WOULD REPORT ON THE WRONG SAMPLE. ISSUE DETAILS: LIS DOWNLOADED A TEST ORDER (TEST ID (B)(6)) IN DUPLICATE FOR SAMPLE 1. SAMPLE 1 WAS LOADED TO POSITION 3 OF A RACK/DISK 123, THE INSTRUMENT QUERIED IM AND RECEIVED AN ORDER FOR TEST ID (B)(6). THE INSTRUMENT RAN THE TEST AND CORRECTLY GENERATED AND REPORTED A RESULT. LATER, SAMPLE 2 IS PLACED IN POSITION 3 OF THE RACK/DISK 123, BUT WHEN THE INSTRUMENT TRIED TO READ IT, THERE WAS A BARCODE READ ERROR AND IM DOWNLOADED THE PENDING TEST FOR SAMPLE 1. THIS BARCODE ERROR CAUSED IM TO INCORRECTLY IDENTIFY THE SAMPLE BASED ON THE RACK/DISK AND POSITION WHICH SUBSEQUENTLY CAUSED THE INSTRUMENT TO RUN TESTS ON THE SAMPLE INCORRECTLY AND RESULTS REPORTED WITH THE WRONG SPECIMEN. THERE WAS A BUG IN V8.00.0001 AND V8.00.0002 OF THE DRIVER. THIS BUG WAS IDENTIFIED, CORRECTED AND RELEASED BEFORE THIS CUSTOMER COMPLAINT OCCURRED. WHEN THE ORIGINAL BUG WAS REPORTED, NO PATIENT SAFETY CONCERNS WERE RAISED AT THAT TIME. THE PATIENT SAFETY ISSUE WAS NOT BROUGHT TO DATA INNOVATION¿S ATTENTION UNTIL AFTER THE PATCH WAS RELEASED. THIS RESULTED IN A DELAY IN REPORTING THE ISSUE TO THE FDA IN A TIMELY MANNER. A CAPA INCIDENT IS BEING OPENED TO DETERMINE WHY THE PATIENT SAFETY ISSUE WAS NOT IDENTIFIED SOONER. ACTIONS TAKEN: SITE REPORTING THE PROBLEM HAS BEEN CONTACTED WITH INFORMATION ABOUT THE ISSUE, AND INFORMED OF THE SOLUTION TO USE TO PREVENT THE PROBLEM. THE DRIVER HAS BEEN CORRECTED AND WAS RELEASED SEPTEMBER 14, 2012. THE SOFTWARE PATCH IS AVAILABLE FOR DOWNLOAD FROM OUR WEBSITE, AS PER OUR ESTABLISHED PROCEDURES. AN EMAIL DESCRIBING THIS ISSUE AND RESOLUTION WAS SENT TO THE CUSTOMER BASE AFFECTED BY THIS ISSUE ON (B)(4) 2012. DATA INNOVATIONS WILL ASSIST ANY CUSTOMER WITH QUESTIONS OR NEEDING ASSISTANCE WITH APPLYING THE PATCH. CAPA INCIDENT OPENED TO RESEARCH WHY PATIENT SAFETY WAS NOT IDENTIFIED EARLIER IN THE COMPLAINT HANDLING PROCESS. | 1
3056604 | PER REPORT FROM LITHOTRIPSY TECHNICIAN - DURING ROUTINE LITHOTRIPSY CASE, WATER OVER TEMPERATURE SAFETY SWITCH FAILED, POSSIBLY CAUSING PATIENT SCALDING. TECH NOTICED BLISTERING ON PATIENT SKIN AFTER LITHOTRIPSY CASE. EQUIPMENT ISSUE RESOLVED ON THE SAME DAY AND DOCUMENTED ON SERVICE REPORT. PER INFORMATION RECEIVED FROM HOSPITAL DIRECTOR OF RISK MANAGEMENT ON (B)(4) 2013 - BURN TO FLANK AREA NOTICED POST-PROCEDURE. PATIENT ADMITTED TO HOSPITAL. PER CONVERSATION WITH HOSPITAL RISK MANAGER ON (B)(4) 2013 - PATIENT RECEIVED A 3RD DEGREE BURN TO THE FLANK AREA AND WAS ADMITTED TO THE HOSPITAL. BURN WAS NOTICED POST-PROCEDURE. PER EMAIL FROM HOSPITAL RISK MANAGER ON (B)(4) 2013 - THE BURN SIZE IS APPROXIMATELY 3% TOTAL BODY SURFACE AREA AND MEASURES 22.5CM X 17.3CM. (B)(6), A THIRD PARTY MOBILE SERVICE PROVIDER WHO OWNS THIS LITHOTRIPTOR, REQUESTED SERVICE. SERVICE REQUEST (B)(4) OPENED AND HEALTHTRONICS SERVICE TECHNICIAN DISPATCHED (B)(4) 2013. PER SERVICE REPORT: FOUND SHORT IN OVERTEMP SWITCH. REPAIRED SHORT AND CHECKED HEATER CONTROLLER FOR PROPER OPERATION. IN ADDITION, SERVICE REQUEST (B)(4) REPLACED HEATER CONTROLLER AND CHECKED FOR PROPER PROGRAMMING. TESTED OVERTEMP SWITCH BY INCREASING TEMPERATURE. SAFETY SWITCH ENGAGED AND UNIT GENERATED AN ERROR MESSAGE. VERIFIED WATER TEMPERATURE DECREASED. ALSO REPLACED RB1 MODULE TO ENSURE HEATER WOULD NOT STAY ON DUE TO STICKING RELAY. RELEASED UNIT FOR USE. | 0
3535321 | VENDOR LOWERED EQUIPMENT ONTO POWER CORD INADVERTENTLY. CORD WAS DAMAGED. EQUIPMENT PLUGGED INTO POWER SOURCE, RESULTING IN BURNING OF ELECTRICAL INSULATION ON CORD AND SPARKING. EQUIPMENT UNPLUGGED. UNIT IS OWNED AND OPERATED BY VENDOR. | 0
3239402 | OUR HOSP RECENTLY PURCHASED SEVERAL GE MACLAB DEVICES TO MONITOR VITAL SIGNS DURING CATH LAB PROCEDURES. AS PART OF THE INSTALL, IT IS THE FACILITIES RESPONSIBILITY TO INSTALL ANTIVIRUS SOFTWARE ON THESE DEVICES TO PROTECT THEM FROM MALICIOUS SOFTWARE. GE PROVIDES INSTALL INSTRUCTIONS AND VALIDATED VERSIONS OF SOFTWARE TO USE. UNFORTUNATELY THE VERSIONS THEY VALIDATE ARE NO LONGER AVAILABLE FOR PURCHASE, AND THEY DO NOT VALIDATE CURRENT VERSIONS OF ANTIVIRUS SOFTWARE FOR THIS SOFTWARE VERSION, EVEN THOUGH THEY SUPPORT IT. IF WE LOAD THE CURRENT VERSION OF ANTIVIRUS SOFTWARE AVAILABLE ON THE DEVICE, IT CAN INVALIDATE OUR WARRANTY, IF WE DON'T AND THE DEVICE GETS INFECTED AND IMPACTS PT SAFETY, IT WILL BE OUR RESPONSIBILITY SINCE WE WERE SUPPOSED TO LOAD ANTIVIRUS SOFTWARE ON. NOT BEING ABLE TO LOAD ANTIVIRUS SOFTWARE CAN IMPACT PT SAFETY IF THE DEVICE GETS INFECTED AND IS NOT AVAILABLE FOR PT USE. IT WOULD HELP IF GE WOULD VALIDATE VERSIONS OF ANTIVIRUS THAT CAN BE PURCHASED AND NOT VERSIONS THAT DON'T EXIST ANYMORE, FOR EQUIPMENT THAT THEY SUPPORT. WOULD LIKE TO COMPLY WITH THE FDA'S JUN 13 SAFETY COMMUNICATION ON CYBER SECURITY FOR MEDICAL DEVICES. | 1
2932116 | LEFT EAR SUDDEN PARTIAL HEARING LOSS AND TINNITUS SYMPTOM OCCURRING AFTER CERVICAL SPINE MRI. | 0
3181660 | THE MAGNET MOVER MOVED BACKWARD AT THE PT SAFETY CHECK POINT (PSCP) WHEN PRESSING THE FORWARD BUTTON. THEN IT STOPPED MOVING. OPERATING ROOM PERSONNEL CRANKED IT HOME AND RESET IT. NO PT WAS INVOLVED WITH THE VENT. IMRIS CUSTOMER SUPPORT SPECIALIST TESTED THE MAGNET MOVER BACK AND FORTH INTO OPERATING ROOM AND COULD NOT DUPLICATE THE FAILURE. IMRIS CUSTOMER SUPPORT AND DEVELOPMENT SPECIALIST REMOTELY COLLECTED THE LOGS FROM THE MATRIX COMPUTER FOR FURTHER ROOT CAUSE ANALYSIS. IMRIS DETERMINED THAT THE ERRONEOUS REVERSE MOVEMENT OCCURRED ON (B)(6) 2013, AT THE TIME 12:00:23.938 IN OPERATING ROOM 1. THE FWD BUTTON WAS PUSHED, BUT SERVO COUNTS WERE SHOWN TO BE DECREASING, WHICH DEMONSTRATES THAT THE SERVO WAS MOVING THE MAGNET IN THE REVERSE DIRECTION. THIS MATCHES THE RACE CONDITION SIGNATURE THAT WAS OBSERVED IN COMPLAINT CASE #(B)(4) IN WHICH A MEDICAL DEVICE REPORT WAS FILED AND APPEARS TO BE THE SAME ROOT CAUSE MANIFESTING AS THE INCORRECT SERVO INDEX. ALL LOG ENTRIES WERE CHECKED FOR ERRONEOUS FORWARD MOVEMENT AND ERRONEOUS REVERSE MOVEMENTS (29,269 LOG ENTRIES). NO OTHER ERRONEOUS MOVEMENTS WERE OBSERVED. IMRIS DETERMINED THAT THE ROOT CAUSE OF THE INCIDENT WAS THAT THE MOTOR SERVO USED TO CONTROL THE MAGNET WAS IN AN INCORRECT STATE, CAUSING REVERSE MOTION WHEN THE FORWARD BUTTON WAS PRESSED BY THE OPERATOR . ON (B)(4) 2013, THE CUSTOMER SITE WAS UPGRADED WITH THE MMC 2.1 SOFTWARE RELEASE TO RESOLVE ALL THE IDENTIFIED CAUSES OF ERRONEOUS MOVEMENT. (REF MDR 3003807210-2012-00002). ALL OTHER SIMILAR SITES WILL ALSO BE UPGRADED TO MMC 2.1 THAT HAVE NOT BEEN UPGRADED ALREADY. | 0
2861828 | THE HOSPITALS STERILIZING FACILITY HAS REPORTED THAT THEY HAVE CONSISTENTLY IDENTIFIED STAINING ON THE INSIDE OF THE KIMGUARD WRAPPING AROUND THE MATRIX DEGENERATIVE TRAYS SUBSEQUENT TO STERILIZING. THE DARK STAINS ON THE WRAPPING APPEAR MOST PROMINENT ON KIMGUARD USED TO WRAP TRAYS THAT HOUSE THE SCREWS OR OTHER IMPLANTS. THE STAINS ARE SUSPECTED TO COME FROM THE BLACK ALUMINUM MATERIAL THAT MAKE UP THE VARIOUS IMPLANT CADDIES. THESE OBSERVATIONS WERE MADE DURING THE VALIDATION PROCESS THAT THE STERILIZING FACILITY CONDUCTS TO ENSURE THAT SETS USED MEET THEIR STERILITY GUIDELINES. THIS IS THE 1ST OF 2 REPORTS SUBMITTED ON THIS EVENT. THE STERILIZING FACILITY HAS BEEN PROVIDED WITH SYNTHES GUIDELINES FOR THE CARE AND MAINTENANCE OF IMPLANTS, INSTRUMENTS & CASES. NOTABLY, THE PH LEVEL USED IN THEIR STERILIZING FACILITY IS 10. ACCORDING TO THE GUIDELINES THIS FALLS OUTSIDE OF THE RECOMMENDED 6-9.5 PH RANGE FOR DEALING WITH ALUMINUM. INVESTIGATION IS ON GOING; NO CONCLUSION COULD BE DRAWN AS NO DEVICE WAS RETURNED OR LOT NUMBER PROVIDED. | 0
3421089 | SUBSCRIBER FELL IN HER HOME AND HER HEALTH ALERT PERSONAL EMERGENCY RESPONSE BUTTON DID NOT WORK WHEN SHE PRESSED IT. | 0
2982889 | ON (B)(6) 2013, A (B)(6) FEMALE PATIENT (IDENTIFICATION WITHHELD) WAS TRANSFERRED FROM (B)(6) HOSPITAL TO (B)(6) HOSPITAL (BOTH IN (B)(6)) FOR FURTHER TREATMENT OF AN UNSPECIFIED CARDIAC/CHEST CONDITION. AT THE TIME OF TRANSFER, THE PATIENT'S COMPUTED TOMOGRAPHY (CT) STUDY OF THE CHEST ACQUIRED AT (B)(6) WAS ELECTRONICALLY TRANSMITTED TO (B)(6). AT THE TIME OF THE STUDY TRANSMISSION, (B)(6) WAS EXPERIENCING AN EPISODE OF DOWNTIME WHICH CAUSED THE STUDY TRANSMISSION TO BE DELAYED. WHEN THE (B)(6) STAFF DISCOVERED THAT THE STUDY IN QUESTION COULD NOT BE ACCESSED IN IMPAX, THEY DECIDED TO PERFORM ANOTHER EXAM ON THE PATIENT, BUT UNFORTUNATELY THE PATIENT EXPIRED DURING THIS TIME. THE AFFECTED STUDY BECAME AVAILABLE IN IMPAX APPROXIMATELY TWO HOURS FOLLOWING THE PATIENT'S DEMISE AFTER THE EPISODE OF DOWNTIME OCCURRED. PRIOR TO THIS EVENT, (B)(6) HAD CONTACTED AGFA EARLIER ON (B)(6) 2013 TO CHECK THE DOWNTIME ISSUE. AGFA ADDED NEW DATAFILES, RESTARTED (B)(4) AND CONFIRMED IMPAX WAS OPERATIONAL ON (B)(6) 2012, BUT AGAIN THIS WAS BEFORE THE EVENT OCCURRED WITH THE PATIENT. THE DATE AGFA BECAME AWARE OF THE EVENT WAS NOT UNTIL (B)(4) 2013, WHEN THE CUSTOMER, (B)(6), CONTACTED AGFA AND TOLD AGFA THEY WERE CONDUCTING AN INTERNAL INVESTIGATION AND REQUESTED ASSISTANCE FROM AGFA TO LOCATE THE SPECIFIC MISSING AUDIT LOGS RELATED TO THE REMOTELY ACQUIRED STUDY THAT WAS DELAYED FROM (B)(6) ON (B)(6) 2013. AN AGFA SERVICE ENGINEER IMMEDIATELY OPENED A PRODUCT COMPLAINT ON (B)(6) 2013 AND ALSO SUCCESSFULLY OBTAINED AND DELIVERED THE REQUESTED LOGS TO THE CUSTOMER. AGFA HAS SINCE MADE MULTIPLE ATTEMPTS TO THE CUSTOMER TO OBTAIN SPECIFIC CLINICAL DETAILS SURROUNDING THE PATIENT DEATH, BUT AT THIS TIME, THE CUSTOMER HAS NOT MADE THESE DETAILS AVAILABLE. THE PRODUCT IMPAX 5.2 WAS END OF LIFE (B)(6) 2008. THE END OF SERVICE LIFE WAS (B)(6) 2010. AGFA AND THE SITE ARE CURRENTLY IN PREPARATION TO UPGRADE THEIR SYSTEM FROM IMPAX 5.2 TO IMPAX 6.5. AT THIS TIME AGFA CONSIDERS ROOT CAUSE TO BE RELATED TO SITE NETWORK ISSUES, 3RD PARTY ISSUES AND THAT THE PRODUCT HAS BEEN END OF LIFE FOR OVER 4 YEARS AND END OF SERVICE LIFE FOR OVER 2 YEARS, REQUIRING SIGNIFICANT MAINTENANCE. IF WE OBTAIN ADDITIONAL INFORMATION FROM THE CUSTOMER, WE WILL SUBMIT A SUPPLEMENTAL REPORT.><(B)(4). THE SLOW PERFORMANCE / SERVER DOWNTIME EXPERIENCED AT THE SITE HAVE BEEN DETERMINED TO BE RELATED TO MULTIPLE ROOT CAUSES, LISTED BELOW. THE SITE IS RUNNING A TIVOLI BACKUP DURING WORKING HOURS THAT IS HEAVILY IMPACTING CPU LOADS. IT WAS RECOMMENDED THAT THIS BACKUP BE RESCHEDULED FOR OFF PEAK HOURS. THE ORACLE VERSION RUNNING AT THE SITE (10.1.0.4.0) HAS A KNOWN DEFECT. THE ISSUE WAS RESOLVED IN ORACLE 10.2.0.2 AND THE SITE HAS BEEN ADVISED TO UPGRADE TO THIS VERSION OR HIGHER. AS RECOMMENDED BY ORACLE THE SITE IS TO INSTALL ORACLE PATCH 3650599. CHANGE THE FILE SYSTEM LAYOUT OF ¿/DBASE¿ AND ¿/DBASE1¿ SO THAT IT WILL MATCH AGFA¿S REQUIREMENTS FOR IMPAX-DATABASES AND BEST PRACTICE GUIDELINES. MIGRATE TO LMT (LOCALLY-MANAGED-TABLE SPACES). THE SITE HAS A TEMP-DATA FILE THAT IS 30GB BIG. THIS IS NOT RECOMMENDED AND CANNOT BE CREATED WITH THE TOOLS THAT ARE SUGGESTED TO BE USED BY AGFA. THIS ISSUE WILL BE REVIEWED AS IT VIOLATES AGFA'S BEST PRACTICE. UPGRADE TO IMPAX 6.5, SCHEDULED IN SEPT 2013. ON SEPTEMBER 19, 2013, AGFA HEALTHCARE VOLUNTARY REPORTED THIS ISSUE TO THE FDA & HEALTH (B)(4). CORRECTIONS RECOMMENDED TO SITE: RESCHEDULE TIVOLI BACKUP, UPGRADE ORACLE, CHANGE FILE SYSTEM LAYOUT TO MATCH AGFA'S REQUIREMENTS, MIGRATE TO LMT, & REVIEW SITES' TEMP-DATA FILE. AGFA AND (B)(6) HOSPITAL ARE IN PREPARATION FOR THE UPGRADE TO IMPAX 6.5 IN OCTOBER 2013. ALL OTHER DOCUMENTATION WILL BE REPORTED TO THE FDA VIA (B)(4). DEVICE MANUFACTURER DATE HAS BEEN CORRECTED FROM 06/01/2011 TO 12/17/2004.><. | 1
3250736 | HOSPITAL IDENTIFIED MEDICATIONS WERE INADVERTENTLY APPEARING ON PATIENTS' DISCHARGE MEDICATION RECONCILIATION PAPERWORK. THESE MEDICATIONS WERE NOT PRESCRIBED NOR INTENDED FOR THESE PATIENTS. REVIEW IDENTIFIED THIS WAS RELATED TO THE FACT THAT THE GENERIC ID WAS CHANGED ON A MEDICATION TEMPLATE IN HORIZON MEDS MANAGER (HMM). THE GENERIC ID ON THIS TEMPLATE IS NORMALLY -1. THIS CHANGE IN HMM ALSO CHANGED MED TABLES IN HORIZON CLINICALS. THIS ATTACHED A DIFFERENT MEDICATION TO THIS TEMPLATE. THIS TEMPLATE WAS CHANGED BACK TO THE CORRECT GENERIC ID IN HMM BUT THAT DID NOT UPDATE THE MEDICATION TABLE IN HORIZON CLINICALS. AS A RESULT, THE TEMPLATE REMAINED ATTACHED TO THE DRUG NORGESTREL-ETHINYL ESTRADIOL IN HORIZON CLINICALS TABLES. THE GENERIC ID WAS LATER CHANGED IN HMM TO MATCH THE MEDICATION LATUDA. THIS UPDATED THE HORIZON CLINICALS TABLES. AGAIN THE GENERIC ID WAS CHANGED IN HMM BACK TO -1 BUT DID NOT UPDATE THE HORIZON CLINICAL TABLES. AS A RESULT OF THE ABOVE MENTIONED CASE, 19 PATIENTS WERE AFFECTED. SIX WERE EITHER IDENTIFIED PRIOR TO DISCHARGE OR AT RECEIVING FACILITY UPON PATIENT TRANSFER. TEN WERE DISCHARGED HOME AND NEVER RECEIVED THE MEDICATION AT HOME. TWO WERE DISCHARGED HOME, NEVER RECEIVED THE MEDICATION AT HOME, BUT WERE SUBSEQUENTLY READMITTED AND ON READMISSION RECEIVED THE MEDICATION CAUSING NO HARM. ONE PATIENT WAS DISCHARGED TO A SENIOR NURSING FACILITY, RECEIVED THE MEDICATION AT THE SENIOR NURSING FACILITY, AND WAS SUBSEQUENTLY READMITTED WITH CHANGE IN MENTAL STATUS.======================MANUFACTURER RESPONSE FOR MEDICATION RECONCILIATION SOFTWARE (HMR), HMR (PER SITE REPORTER).======================HOSPITAL INFORMATION TECHNOLOGY LEADERSHIP CONTINUE TO WORK WITH VENDOR TO IDENTIFY ROOT CAUSE AND DATE OF SOFTWARE INTERFACE OCCURRENCE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ELECTRONIC MEDICATION RECONCILIATION SOFTWARE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3446866 | AN ISSUE WAS DISCOVERED AT A CUSTOMER SITE (HIGHLANDS ORTHO). THE OFFICEPACS MIGRATION TOOL INCORRECTLY IDENTIFIED PT FOR PRIOR STUDIES DURING THE RECONCILIATION MIGRATION PROCESS. IMAGES WERE IDENTIFIED FOR A SPECIFIC PT IN ORTHOPACS POST-MIGRATION BUT THE IMAGES WERE DIFFERENT THAN THE ORIGINAL VISIT IN OFFICEPACS. THE OFFICEPACS MIGRATION TOOL THAT CAUSED THE MALFUNCTION IN THIS CASE, IS NOT A DEPLOYED FEATURE OF THE OFFICEPACS PRODUCT. THE TOOL IS USED BY THE INTERNAL PROFESSIONAL SERVICES TEAM AT MERGE TO PERFORM PACS MIGRATIONS. ONCE THE ISSUE WAS REPORTED, THE PROFESSIONAL SERVICES TEAM WAS ABLE TO EXAMINE LOG FILES ON THE DESTINATION DEVICE TO FIND THE MISMATCHED VALUES THAT DID NOT RECONCILE. THOSE STUDIES WERE MANUALLY DELETED THEN FROM THE SYSTEM AND RE-MIGRATED. (B)(4). | 1
3237050 | WHEN THE HOSPITAL STARTED CONVERTING TO ELECTRONIC CHARTING THEY PURCHASED (B)(4) MOBILE COMPUTER CARTS TO PUT THRU OUT THE FACILITY INCLUDING PT ROOMS, RECOVERY AND SURGICAL AREAS. UPON STARTUP, WE WERE HAVING POWER ISSUES WITH SAID CARTS. FURTHER INVESTIGATION SHOWED THAT THE 20 AMP FUSE HOLDER HAD A DESIGN FLAW THAT CAUSED THE FUSE HOLDER ITSELF TO OVERHEAT AND MELT. BECAUSE OF THE POTENTIAL FIRE HAZARD, THE HOSPITAL SHUT DOWN USE OF ALL CARTS UNTIL THIS PROBLEM WAS RESOLVED. THE MFR UNDERSTOOD THE GRAVITY OF THIS ISSUE AND CAME OUT AND REPLACED ALL WIRING HARNESSES, SAYING THAT THE PROBLEM WAS A PRODUCTION LOT ISSUE AND THE NEW HARNESSES SHOULD TAKE CARE OF ANY FURTHER OVERHEATING ISSUES. IT DID NOT. TWO WEEKS LATER, WE WERE SEEING THE SAME PROBLEMS WHICH IN TURN RESULTING AGAIN OF THE COMPLETE SHUT DOWN OF ALL CARTS AND RAISING HUGE ALARMS WITHIN OUR ADMINISTRATION. RUBBERMAID WAS NOTIFIED OF THE SITUATION AND RESPONDED WITH A REFINED DESIGN OF THE WIRING HARNESS. THEY ALSO STARTED TO MAKE THE HARNESSES IN THEIR FACILITY INSTEAD OF SUBCONTRACTING THEM OUT FOR BETTER QUALITY CONTROL. WE HAVE SINCE HAD TWO AUDITS OF THE WIRING HARNESSES ON ALL CARTS AND HAVE HAD NOT FURTHER ISSUES. | 0
3115259 | AT APPROXIMATELY 0036 PT TELEMETRY MONITOR FAILED TO ALARM FOR FAST HEART RATE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TELEMETRY MONITORING.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3115267 | PATIENT WAS BEING EVALUATED WHEN THIS RN NOTICED A DISCREPANCY IN THE HR AS HE VISUALIZED BOTH THE TELEMETRY MONITOR AND THE PATIENT AT REST WHILE BEING QUESTIONED. THIS RN REVIEWED TELEMETRY AND REVEALED THE RATE THAT HAD BEEN RECORDED PER TELEMETRY WAS ACTUALLY BETWEEN 250-300 BEATS/MINUTE, ALSO NOTING PRIOR TELEMETRY HAD ALSO BEEN READING THE HR INADEQUATELY. MD NOTIFIED OF THIS DISCREPANCY AND STAT EKG DONE. BIOMED PAGED TO TELEMETRY, UNKNOWN CAUSE AT THIS TIME, REQUIRES FOLLOW-UP. PATIENT WAS ASYMPTOMATIC THROUGHOUT MORNING, HOWEVER THIS PATIENT HAS COME IN THIS ADMISSION FOR SYNCOPE AND COLLAPSE. AT LEAST EIGHT SIMILAR INCIDENTS INVOLVING INCORRECT INFORMATION FROM TELEMETRY WERE RECORDED ON THE UNIT WITHIN A SHORT TIME, AND CLINICAL USERS EXPRESSED THAT SUCH EVENTS HAVE HAPPENED BEFORE; IT IS A RECURRING PROBLEM. FURTHER REPORTS INCLUDE ISSUES: TELEMETRY MONITOR PICKED UP THE INCORRECT HEART RATE - IT SAID THE HEART RATE WAS 101 BUT IT WAS REALLY 150; TELEMETRY MONITOR SHOWED VTACH AND VFIB WHILE PATIENT HAD PULSE AND WAS NOT HAVING ANY PROBLEMS. THE LEADS WERE TAKEN OFF AND MONITOR STILL SHOWED VFIB. TELEMETRY MONITOR ALARMED FOR HR 39 BUT PT HEART RATE WAS 114. TELEMETRY MONITOR SHOWING SP02 OF 29%- NOT ACCURATE, RN IN ROOM TALKING WITH PT AT TIME OF EVENT;INTERFERENCES WITH TELEMETRY MONITORING, HEART RYTHM FADED; OTHERS.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TELEMETRY MONITORING.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3120102 | RECURRING TELEMETRY PROBLEM PRODUCING FOURTEEN PATIENT SAFETY REPORTS AND AFFECTING A NUMBER OF PATIENTS. THE MAIN ISSUE IS DESCRIBED AS COMPRESSED WAVEFORMS, WITH SPECIFIC COMMENTS INCLUDING: AT APPROXIMATELY 1307 PT TELEMETRY RHYTHM BECAME COMPRESSED; AT APPROXIMATELY 0610 PATIENT TELEMETRY RHYTHM WAS COMPRESSED DUE TO COMPUTER GLITCH; TELEMETRY MONITORS EXPERIENCING "COMPRESSION," MAKING THE PATIENT'S SIX SECOND RHYTHM STRIP INDETERMINATE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TELEMETRY MONITORING.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3115219 | INTERRUPTION IN TELEMETRY MONITORING, AFFECTING NUMEROUS PATIENTS AND PRODUCING 35 PATIENT SAFETY REPORTS. THESE REPORTS ARE RELATED TO DROP OUTS IN THE SYSTEM. SPECIFIC REPORT TEXTS INCLUDE: AT APPROXIMATELY 0146 PT TELEMETRY MONITORING WAS INTERRUPTED; AT APPROXIMATELY 2005, 2017, AND 2219 PT HAD AN INTERRUPTION IN TELEMETRY MONITORING; AT APPROXIMATELY 1724, 1757, 1758, 2043, 2100, AND 2132 PT TELEMETRY MONITORING WAS INTERRUPTED; AT APPROXIMATELY 1547, 1930, AND 2354 PT TELEMETRY MONITORING WAS INTERRUPTED DUE TO COMPUTER GLITCH; AT APPROXIMATELY 1547, 1930, AND 2354 PT TELEMETRY MONITORING WAS INTERRUPTED DUE TO COMPUTER GLITCH; TELEMETRY DROPOUT OF INFORMATION (THE TRACING ITSELF) INTERMITTENTLY FROM 15:52:37 THROUGH 15:52:56, ALSO FROM 15:20:10 THROUGH 15:20:44. E-MAIL SENT TO LEADERSHIP OF THE TWO UNITS AFFECTED. MULTIPLE "DROPOUT" OCCURENCES ON TELEMETRY, NO ADVERSE EVENTS, PATIENT DISCHARGED THIS DAY. OF NOTE, THE SAME PATIENT TELEMETRY WAS INVOLVED IN YESTERDAY'S DROPOUTS; CHANGED BATTERIES, CONTINUED TO OCCUR ON SAME PATIENT -TELEMETRY MONITOR HAD A "DROP OUT" OCCURRENCE. PT'S 6 SECOND RHYTHM STRIP WAS INTERRUPTED BY THE SYSTEM MAKING THE RHYTHM UNDETERMINED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TELEMETRY MONITORING.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 1
3504581 | PATIENT HERE FOR A VERTIBULOPLASTY AND WAS LAID PRONE FOR THE PROCEDURE. ANESTHESIA PUT THE BP CUFF ON THE LEG. WHEN IT WAS REMOVED THE PATIENT HAD A LACERATION WITH ADIPOSE TISSUE SHOWING. PLASTIC SURGERY HAS BEEN CONTACTED FOR EVALUATION. PATIENT'S MEDICAL HISTORY IS THAT SHE HAS ARTHRITIS, MYASTHENIA GRAVIS, AND IS ON STEROIDS. THE PROCEDURE WAS ABORTED. | 0
3308371 | THE EVENT WAS AN ARRHYTHMIA THAT REQUIRED MEDICATION INTERVENTION. WHEN THE NURSE TRIED TO PRINT THE ECG STRIP, A MESSAGE POPPED UP THAT SAID "SAVING" AND THE UNIT DID NOT PRINT A STRIP. STAFF REPORTED THAT THIS HAPPENS FREQUENTLY WITH THE OTHER CARDIAC MONITORS AS WELL. | 1
3390296 | THE CENTRAL TELEMETRY MONITORING SYSTEM HAD BEEN SHUTTING DOWN INTERMITTENTLY OVER THE COURSE OF THREE DAYS. THE COMPANY WAS NOTIFIED AND HAD BEEN ON SITE MAKING REPAIRS. ON THE THIRD DAY, THE CENTRAL MONITORING SYSTEM SHUT DOWN. IMMEDIATELY PRIOR TO SHUTTING DOWN, THE PT WAS NOTED TO HAVE A NORMAL SINUS RHYTHM. WHEN THE MONITORS WENT BACK UP, THE PT WAS NOTED TO BE IN V-FIB. THE PT LATER EXPIRED. | 1
3305690 | SPACELABS RECEIVED A REPORT OF A TELEMETRY PATIENT WITH A 12 BEAT RUN OF V-TACH THAT DID NOT ALARM OR PRINT OUT ON THE CENTRAL MONITOR. THE PATIENT DIED. IT IS SPACELABS POLICY TO REPORT EACH EVENT ASSOCIATED WITH A PATIENT DEATH. SPACELABS SENT A FIELD SERVICE ENGINEER (FSE) TO THE CUSTOMER SITE TO EVALUATE THIS REPORTED EVENT. THE FSE TESTED THE CENTRAL MONITOR AND DETERMINED THE DEVICE WAS OPERATING AS DESIGNED. THE HOSPITAL REFUSED TO ALLOW THE FSE TO COPY OR EVALUATE THE PATIENT DATA IN THEIR INTESYS CLINICAL SUITE (ICS IS A RETROSPECTIVE REVIEW DATABASE) THAT STORES PATIENT DATA AND WOULD PROVIDE DETAILS SURROUNDING THE INCIDENT THAT OCCURRED. ONLY ONE ECG STRIP OF THE 12 BEAT VRUN EVENT WAS AVAILABLE FOR ANALYSIS. THE HOSPITAL PROVIDED THE DEFAULT USER SETTINGS AND SUPPLEMENTAL PATIENT WAVEFORM DATA FOR THE TIME FRAME BEFORE THE REPORTED EVENT. HOWEVER, WITH NO DATA PROVIDED FOR THE TIME OF THE EVENT OTHER THAN THE ECG STRIP, SPACELABS WAS UNABLE TO REACH A CONCLUSION REGARDING THE EVENT THAT OCCURRED. THERE HAVE BEEN NO FURTHER REPORTS OF FAILURE TO ALARM OR PRINT OUT FOR V-TACH FROM THIS CUSTOMER. SPACELABS WILL CONTINUE TO MONITOR THIS ISSUE AS PART OF OUR TRENDING PROCESS. SPACELABS CONSIDERS THIS ISSUE CLOSED. | 1
2852756 | THE CUSTOMER REPORTED THE CARESCAPE B650 MONITOR WAS REBOOTED FOR THREE CYCLES RESULTING IN AN ERROR MESSAGE. THE MONITOR WAS NOT ABLE TO REGAIN FUNCTIONALITY DURING THE EVENT. THE CRITICALLY ILL PATIENT WAS TRANSFERRED TO ANOTHER MONITOR. THE PATIENT WAS A DO NOT RESUSCITATE STATUS AND EXPIRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2947521 | THE PT UNDERWENT A LUMBAR INTERBODY FUSION PROCEDURE AT THE L3 - L5 SPINE LEVELS IN WHICH THE NEUROVISION DEVICE WAS UTILIZED. PT REPORTED SIGNIFICANT LEFT HIP/LEG PAIN IMMEDIATELY FOLLOWING SURGERY WHICH DID NOT ABATE IN THE SUBSEQUENT DAYS. LUMBAR PLEXUS INJURY AND FEMORAL NERVE PALSY WERE REPORTED. NO ADDITIONAL PT HISTORY IS KNOWN AT THIS TIME. (B)(4). THE DEVICE WAS EVALUATED BY NUVASIVE ON (B)(4) 2010 AND WAS FOUND TO OPERATE WITHIN ESTABLISHED SPECIFICATIONS. A MALFUNCTION IS NOT KNOWN TO HAVE OCCURRED AND NO ROOT CAUSE OF THE REPORTED EVENT HAS BEEN IDENTIFIED. REVIEW OF SERVICE AND COMPLAINT HISTORY INDICATES NO FAULTS WERE NOTED FOLLOWING THE SURGICAL PROCEDURE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE A SUBSEQUENT REPORT WILL BE FILED. | 0
3200929 | PATIENT WAS ANESTHETIZED FOR CRANIOTOMY PROCEDURE FOR A BRAIN MASS. AFTER HER SCAN WAS LOADED WITHOUT DIFFICULTY, AN ERROR MESSAGE WAS NOTED INDICATING THAT THE NAVIGATION SOFTWARE WAS CORRUPTED. NUMEROUS REBOOTS WERE ATTEMPTED WITHOUT SUCCESS. THE LOCAL SALES REPRESENTATIVE WAS NOT IMMEDIATELY AVAILABLE AND TECHNICAL SUPPORT FROM THE COMPANY ADVISED THAT THE SOFTWARE REP WOULD HAVE TO COME IN TO FIX THE PROBLEM. ANOTHER PATIENT DISK WAS ATTEMPTED TO DETERMINE IF THE FAILURE WAS SPECIFIC TO THE PATIENT'S DISK. THE FAILURE WAS DETERMINED TO NOT BE RELATED TO THE DISK, BUT SOFTWARE RELATED. PROCEDURE WAS CONTINUED WITHOUT THE USE OF THE BRAINLAB NEURONAVIGATION. SALES REP CAME TO TROUBLESHOOT THE ISSUE LATER AND HAD TO RELOAD SOFTWARE. SHE FELT THAT THE MALFUNCTION MAY HAVE OCCURRED DUE TO MULTIPLE INAPPROPRIATE SHUTDOWN PROCEDURES.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?CRANIOTOMY USING NEURONAVIGATIONDEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO | 1
3239943 | A (B)(6) YEAR OLD MAN, PRESENTED TO HOSPITAL WITH CHEST PAIN AND UNDERWENT CARDIAC CATHETERIZATION FOR A LESION LOCATED IN THE PROXIMAL RAMUS INTERMEDIUS. DURING THE PROCEDURE, INTRAVASCULAR ULTRASOUND WAS PERFORMED USING THE TVC IMAGING SYSTEM. POST REMOVAL OF TVC SYSTEM, AN INFLATABLE DEVICE WAS USED ON THE LESION. UPON REMOVAL OF THE INFLATABLE DEVICE, IT WAS DETERMINED THAT THE PT HAD A DISSECTION IN THE DISTAL RAMUS INTERMEDIUS. RESULTS FROM THE INFRAREDX INVESTIGATION OF THE ANGIOGRAPHIC FILMS AS WELL AS DISCUSSIONS WITH THE ATTENDING PHYSICIAN SUGGEST THAT THE ROOT CAUSE OF THIS EVENT MOST MORE LIKELY RELATED TO THE FIRST INFLATABLE DEVICE. DURING THIS FIRST INFLATION, THE WIRE WAS RETRACTED SLIGHTLY CAUSING THE TIP OF THE WIRE TO RESIDE IN THE TIGHT BEND OF AN ARTERY THAT EXPERIENCES A LARGER DEGREE OF CONTRACTION DURING EACH CARDIAC CYCLE. | 0
3122631 | INOVENT CIRCUIT WRONGLY CONNECTED [DEVICE MISUSE]. OXYGEN LEVEL DECREASED [OXYGEN SATURATION DECREASED]. PATIENT CONDITION DETERIORATED [GENERAL PHYSICAL HEALTH DETERIORATION]. CASE DESCRIPTION: THE INITIAL REPORT SPONTANEOUS REPORT WAS RECEIVED ON (B)(6) 2013 BY IKARIA (B)(4) FROM A MEDICAL CENTER IN (B)(6). IT WAS REPORTED THAT WHEN THE INFANT WAS TEMPORARILY SWITCHED FROM INOVENT (B)(4) TO MANUAL INOFLO (NITRIC OXIDE) DELIVERY NO WAS NOT DELIVERED BECAUSE A "CIRCUIT WAS WRONGLY CONNECTED". FOLLOW-UP INFORMATION RECEIVED ON (B)(6) 2013 IS ALSO INCLUDED IN THIS INITIAL REPORT. MEDICAL HISTORY: FEMALE, (B)(6), POTTER'S SYNDROME, HYPOXIC RESPIRATORY FAILURE (HRF) WITH PULMONARY HYPERTENSION. CONCOMITANT MEDICATIONS: MIDAZOLAM, NOBELBAR (PHENOBARBITAL SODIUM), HUMULIN R (INSULIN HUMAN), MYOCOR (GLYCERYL TRINITRATE), CRITPAN (DOPAMINE HYDROCHLORIDE), AND BICARBONATE. A FEMALE NEONATE WITH POTTER'S SYNDROME WAS ON A BABY LOG 8000 VENTILATOR WITH THE FOLLOWING SETTINGS: PEAK INSPIRATORY PRESSURE (PIP) 28 CM H2O, POSITIVE END-EXPIRATORY PRESSURE (PEEP) 8 CM H2O AND FRACTIONAL INSPIRED OXYGEN CONCENTRATION (FIO2) 1.0. ON (B)(6) (HER DATE OF BIRTH), SHE WAS STARTED ON "INTUBATED NITRIC OXIDE (NO) DELIVERY" AT 20 PPM VIA INOVENT (B)(4) FOR HYPOXIC RESPIRATORY FAILURE WITH PULMONARY HYPERTENSION; SHE WAS ALSO STARTED ON MYOCOR (GLYCERYL TRINITRATE) FOR PULMONARY HYPERTENSION. AT 16:30 ON (B)(6) 2013 THE INFANT WAS TEMPORARILY SWITCHED TO MANUAL NO DELIVERY (NOS) AND "A BAGGING WAS DONE WITH ONLY OXYGEN DELIVERY WITHOUT NO DELIVERY" BECAUSE "A CIRCUIT WAS WRONGLY CONNECTED". FOLLOWING THE INTERRUPTION OF NO DELIVERY, THE NEONATE'S SPO2 DECREASED TO 50% (BASELINE 72%). IT WAS ALSO REPORTED THAT HER CONDITION DETERIORATED (NOS). SHE WAS "IMMEDIATELY INTUBATED AND NO DELIVERED", BUT DUE TO THE "TOUCHING OF THE PATIENT FOR SUCTIONING" IT TOOK 12 HOURS FOR HER OXYGEN SATURATION LEVELS TO REMAIN STABLE AT 70%. ON (B)(6) 2013, THE NEONATE WAS ALSO STARTED ON MIDAZOLAM FOR SEDATION. ON (B)(6) 2013 THE USE OF VASOPRESSOR CRITPAN (DOPAMINE HYDROCHLORIDE) WAS INITIATED AND THE NEONATE RECEIVED BICARBONATE FOR ACIDOSIS. ON (B)(6) 2013, SHE WAS STARTED ON HUMULIN R (INSULIN HUMAN), AND ON (B)(6) 2013 SHE AGAIN RECEIVED BICARBONATE FOR ACIDOSIS. TREATMENT WITH INOFLO (NITRIC OXIDE) WAS DISCONTINUED ON (B)(6)2013 AT 14:00 (NOS) FOR UNSPECIFIED REASONS. SHE REMAINED ON MIDAZOLAM, HUMULIN R, MYOCOR (GLYCERYL TRINITRATE), AND CRITPAN (DOPAMINE HYDROCHLORIDE). OUTCOME: THE PATIENT "RECOVERED SAFELY" AND "THERE WERE NO PARTICULAR HEALTH HAZARDS". IT WAS DETERMINED THAT "AN OXYGEN DISS HOSE WAS ATTACHED TO THE OXYGEN FLOWMETER AT THE END OF PRE-USE INSPECTION, BUT THE OXYGEN DISS HOSE WAS REMOVED FOR SOME REASON DURING THE MANUAL NO ADMINISTRATION, AND A TUBE OF JACKSON LEASE WAS CONNECTED TO THE OXYGEN FLOWMETER BY USER'S MISTAKE". IN RESPONSE TO RECEIVING THE REPORT, THE PERSON IN CHARGE FROM IKARIA (B)(4) CHECKED THE OXYGEN FLOWMETER OF THE INOVENT AND DISS HOSE AND FOUND NO PARTICULAR PROBLEM. IT WAS POSTULATED THAT "IT DID NOT COME OFF BY ITS OWN" AND "PROBABLY, IT WAS DROPPED BY NURSES". THE PRE-USE INSPECTION WAS COMPLETED BY NURSES AND THE DOCTOR. THE PREVIOUS MONTHLY INSPECTION WAS COMPLETED BY REPORTER. THIS INCIDENT WAS CONSIDERED TO BE A MEDICAL ACCIDENT. THE REPORTER SUGGESTS PREVENTIVE MEASURES BE TAKEN SUCH AS A CHANGE IN THE METHOD OF PRE-USE INSPECTION AND A CHANGE OF OXYGEN TUBE SHAPE. IN THE REPORTER'S OPINION, "PROCEDURE FAILURE" WHICH RESULTED IN A "BAGGING DONE WITH OXYGEN ONLY AND NO NO" WAS THE CAUSE OF THE INFANT'S CONDITION DETERIORATION AND OXYGENATION LEVEL DECREASE. THE ROOT CAUSE WAS DETERMINED BY USERS AT THE USER FACILITY TO BE USE ERROR, THEREFORE TESTING OF THE ACTUAL DEVICE WAS NOT REQUIRED. THE MANUAL DELIVERY CIRCUIT WAS MIS-CONFIGURED DURING SETUP BY THE USER AND, AS A RESULT, NITRIC OXIDE DRUG WAS NOT DELIVERED DURING ATTEMPTED MANUAL DELIVERY. RATHER, ONLY OXYGEN WAS DELIVERED AND THE PATIENT'S OXYGEN SATURATION LEVEL FELL. FOLLOW-UP INVESTIGATION INTO THE INCIDENT WITH THE USER FACILITY LED TO THE CONCLUSION THAT IT WAS ATTRIBUTABLE TO USE ERROR. THE INITIAL SPONTANEOUS REPORT WAS RECEIVED ON (B)(6) 2013 BY IKARIA (B)(6) FROM A MEDICAL CENTER IN (B)(6). IT WAS REPORTED THAT WHEN THE INFANT WAS TEMPORARILY SWITCHED FROM INOVENT (B)(4) TO MANUAL NO DELIVERY NO WAS NOT DELIVERED BECAUSE A "CIRCUIT WAS WRONGLY CONNECTED". FOLLOW-UP INFORMATION RECEIVED ON (B)(4) 2013, REVEALED THAT THE INFANT EXPERIENCED AN OXYGEN SATURATION DECREASE WHICH WAS CONSIDERED SERIOUS. | 1
3143192 | DASH 5000 MONITORS UNABLE TO: PROVIDE CONTINUOUS CARDIAC TELEMETRY MONITORING, UNABLE TO PERFORM A 12 EKG. THIS HAS BEEN AN INTERMITTENT ISSUE WITH THE MONITORS SINCE PURCHASE AND INSTALLMENT IN OUR FACILITY IN (B)(6) 2012. OUR FACILITY DID NOT HAVE ANY DASH 5000 MONITORS WITH PRODUCT CODE SHQ. IN (B)(6) 2012 THE COMPANY REPLACED SOME MONITOR MODULES AND OUR FACILITY CONTINUED TO HAVE INTERMITTENT ISSUES WITH EKG FUNCTION OF THE MONITORS. PER OUR FACILITY CLINICAL ENGINEERING DEPARTMENT, THEY CONTACTED THE COMPANY DURING THE WEEK OF (B)(6) 2013 HAD REQUESTED IF THE COMPANY WOULD BE ABLE TO TELL US WHETHER THE COMPANY HAD INADVERTENTLY PLACED SHQ CODE PRODUCTS IN OUR MONITORS LAST (B)(6) (2012) BECAUSE WE HAVE CONTINUED TO EXPERIENCE ISSUES WITH THE MONITORS. NO RESPONSE HAS BEEN REC'D AS OF THIS DATE. | 1
3313868 | NURSE CALL SYSTEM ON POST-PARTUM NOT WORKING. POWER DOWN, UNINTERRUPTABLE POWER SUPPLY (UPS) IS NOT FUNCTIONING PROPERLY. | 0
3229546 | IT WAS REPORTED THAT WHEN A TREATMENT PLAN WITH MULTIPLE ISOCENTERS IS CREATED AND TRANSFERRED TO MOSAIQ, ALL FIELDS ARE MERGED INTO ONE ISOCENTER. THERE WAS NO WARNING OF THIS ACTION. THERE WAS NO PT INVOLVED AT THE TIME OF THE EVENT. NO FURTHER INFO WAS REPORTED. | 1
3519419 | I WAS GIVEN THE ONETOUCH DIABETES GLUCOSE MONITOR BY MY DOCTOR AND I DOWNLOADED THEIR ONETOUCH DIABETES MANAGEMENT SOFTWARE V2.3 TO MY PC TO USE TO DOWNLOAD AND REPORT ON READINGS COLLECTED FROM MY METER. I NOTICED A COUPLE OF MATERIAL DEFECTS WITH THE SOFTWARE THAT WOULD ALLOW FOR INCORRECT RESULTS TO BE SHARED WITH MY DOCTOR. I SHARED THIS WITH THEIR CUSTOMER SERVICE STARTING ON (B)(6) 2013. THE SOFTWARE REQUIRES THAT YOU TELL IT TIME SLOTS BEFORE AND AFTER MEAL TIMES. THIS THEN IS USED TO SLOT ANY READINGS THAT ARE DOWNLOADED FROM THE ONETOUCH DEVICE INTO THOSE FIXED SLOTS. IF YOU TRAVEL (MULTIPLE TIME ZONES) OR EAT MEALS OUTSIDE OF THOSE SLOTS, YOUR READINGS ARE INCORRECTLY SLOTTED AND YOU CANNOT EDIT/FIX THE SLOTS FOR ANY OF THE READINGS. THIS MEANS YOU COULD BE REPORTING READINGS TO YOUR DOCTOR AS BEING PRE-MEAL WHEN THEY ARE REALLY POST-MEAL AND REPORTING OTHER READINGS AS POST-MEAL WHEN THEY ARE REALLY PRE-MEAL. THE SECOND DEFECT THAT IMPACTS REPORTING ACCURATE DATA TO YOUR PHYSICIAN IS THAT THE SOFTWARE, ALTHOUGH IT CLAIMS TO SUPPORT (B)(4) FILE FORMAT DATA EXPORT AND EXPORT, DOESN'T WORK . AS A SUGGESTED "WORK AROUND" THEY SUGGESTED EXPORTING THE DATA TO A COMMA SEPARATED FILE, EDITING THE DATA, AND RE-IMPORTING IT. THERE ARE SEVERAL ISSUES WITH THIS METHOD INCLUDING A SECOND SOFTWARE DEFECT. FIRST OF ALL, IF YOU SIMPLY EXPORT TO (B)(6), OPEN IN (B)(6), SAVE BACK AS (B)(6) WITH NO CHANGES TO THE DATA, AND TRY TO RE-IMPORT THE SOFTWARE CLAIMS THAT THE FILE IS NOT SUPPORTED (B)(6) FORMAT. THE SECOND ISSUE IS THAT EVEN IF YOU COULD EDIT AND RE-IMPORT THE FILE, YOU HAVE TO KNOW WHAT THE NUMBERS 1 THROUGH 7 MEAN IN ONE OF THE COLUMNS TO MANUALLY UPDATE THE MEAL SLOTS. THE THIRD ISSUE IS THAT I'VE BEEN IN THE COMPUTER INDUSTRY FOR OVER 25 YEARS AND IF I CANNOT WORK AROUND THIS SOFTWARE ISSUE, HOW WOULD ANYBODY EXPECT SOMEBODY LIKE MY MOM BEING ABLE TO? BOTH ISSUES WERE ACKNOWLEDGED BY THEM AND PASSED ON THEIR DATA MANAGEMENT TEAM ON (B)(6) 2013. THAT WAS MY LAST CONTACT FROM ONETOUCH. I SWITCHED TO THE BAYER CONTOUR NEXT USB DEVICE (A COMPETITOR'S PRODUCT) AND IT WORKS PERFECTLY. WITH EVERY READING I CAN SIMPLY TELL IT IF IT IS A FASTING, BEFORE MEAL OR AFTER MEAL READING AS I TAKE MY READINGS. THEIR PC SOFTWARE IS ALSO OF MUCH HIGHER QUALITY AS WELL. SINCE REPORTING THESE DEFECTS IN THE ONETOUCH SOFTWARE THAT WOULD RESULT IN INCORRECT DATA BEING SHARED WITH MY DOCTOR, I HAVE BEEN NOTIFIED BY (B)(6) WILL NO LONGER COVER BAYER GLUCOSE TEST STRIPS AND WILL ONLY COVER ONETOUCH STRIPS. I REFUSE TO GO BACK TO A PRODUCT WHOSE ACCOMPANYING SOFTWARE HAS MATERIAL DEFECTS THAT LEADS TO INACCURATE DATA BEING SHARED AND WILL THUS PAY FOR MY BAYER TEST STRIPS OUT OF POCKET IN 2014. | 1
2940958 | (B)(6) 2006, 10:00 AM: ARTHROSCOPY PUMP MALFUNCTIONED DURING PROCEDURE CAUSING INCREASED PRESSURE IN RIGHT KNEE JOINT. (B)(6) 2006, 10:00 AM: WHILE PATIENT WAS UNDERGOING A KNEE DA PROCEDURE THERE WAS A MALFUNCTION OF THE FMS/MITEK IRRITATION PUMP KNOWN AS A FMS DUO+. THE PRESSURE INCREASED FROM 50 TO 150 CAUSING INCREASE PRESSURE AND SWELLING IN THE PATIENT'S RIGHT KNEE JOINT AND SURROUNDING TISSUE. THE ORTHOPEDIC SURGEON FOLLOWED UP WITH THE PATIENT POSTOPERATIVE. | 0
3313757 | ON (B)(4) 2013 CARESTREAM HEALTH RECEIVED A REPORT OF LOST IMAGES FROM FOUR (4) STUDIES. THE CUSTOMER IDENTIFIED FIFTY FIVE (55) STUDIES THAT HAD NOT AUTOMATICALLY ARCHIVED TO THE GE PACS. FIFTY ONE (51) STUDIES WERE FOUND STILL ONLINE AND WERE SAVED ON THE GE PACKS. ONE (1) STUDY WAS FOUND TO BE DELETED BY THE USER BECAUSE THE USER BELIEVED IT WAS ON THE CARESTREAM BACKUP SERVER. THREE (3) STUDIES WERE FOUND TO BE NO LONGER AVAILABLE ON THE CARESTREAM BACKUP SERVER. CARESTREAM HEALTH, INC. BEGAN AN INVESTIGATION OF THE PROBLEM. THE INVESTIGATION INCLUDED REVIEW OF LOGS AND A VISIT TO THE SITE TO EVALUATE THE SYSTEM. THE SYSTEM AT THE SITE WAS CONFIGURED TO COPY INCOMING STUDIES TO THE GE PACS. THE THREE (3) LOST STUDIES, NOT DELETED BY THE USER, WERE PREVIOUSLY AVAILABLE ON THE CARESTREAM BACKUP SERVER AS OF (B)(6) 2012. CARESTREAM'S INVESTIGATION OF THE LOGS REVEALED THE STUDIES WERE SUCCESSFULLY STORED ON THE CARESTREAM BACKUP SERVER, BUT THERE WAS NO RECORD OF STUDIES BEING SENT TO THE GE PACS. CARESTREAM BELIEVES THAT THE INFO ROUTER JOB WAS NOT SUCCESSFUL, AND BECAUSE IT WAS NOT NOTICED BEFORE THE INFO ROUTER PURGES ITS JOBS AFTER 14 DAYS, THE CUSTOMER WAS UNAWARE THAT THE STUDIES WERE NOT ARCHIVED. THE SYSTEM WAS ALSO CONFIGURED TO ALLOW THE DELETION OF UN-ARCHIVED STUDIES, THEREFORE, WHEN THE SYSTEM REACHED THE HIGH WATER MARK, THE SYSTEM ALLOWED THE STUDIES TO BE DELETED. THE AFFECTED STUDIES WERE DATED (B)(6) 2012 (PRIOR YEAR MAMMOGRAPHY IMAGES). WHILE THE STUDY IMAGES WERE NO LONGER AVAILABLE, THE REPORTS AND DIAGNOSES FOR THESE STUDIES WERE STILL AVAILABLE ON THE (B)(4). | 1
3151900 | ON (B)(6) 2013, CARESTREAM HEALTH RECEIVED A REPORT OF A 5 YEAR OLD PASSING AWAY IN THE OPERATING ROOM. CARESTREAM WAS INFORMED THAT CT IMAGES WERE UNAVAILABLE FOR THE CHILD WHILE THE CHILD WAS IN THE OPERATING ROOM. CARESTREAM HEALTH IS CURRENTLY INVESTIGATING THIS ISSUE TO DETERMINE IF THE VUE PACS CONTRIBUTED TO THE ADVERSE EVENT. IT APPEARS THAT THE PT CAME IN AND WAS ASSIGNED A NEW MEDICAL RECORD NUMBER AS IF HE WAS A NEW PT. CD IMPORT CT IMAGES WERE ASSIGNED TO THE NEW MEDICAL RECORD NUMBER. UNK TO THE TECHNICIAN AT THE TIME, THERE HAS BEEN AN EXISTING MEDICAL RECORD NUMBER WITH PRIOR STUDIES FROM 2008 ALREADY IN THE SYSTEM. A MERGE PROCESS WAS CARRIED OUT TO MERGE THE TWO (2) MEDICAL RECORDS FOR THE PT SO THAT THE IMAGES WERE EVENTUALLY AVAILABLE. THE RADIOLOGIST WAS ABLE TO READ THE IMAGES FOR THE SURGEON. | 1
2956377 | A BUG WAS DISCOVERED DURING TESTING INTERNALLY IN THE MANUFACTURER COMPANY. THE BUG WAS EVALUATED TO HAVE A PATIENT RISK. THE BUG IS LOCATED IN THE PART OF THE SOFTWARE THAT HANDLES VISUALIZATION OF NERVE FIBER TRACTS THAT ARE EXTRACTED FROM A DIFFUSION TENSOR IMAGING MR SERIES. THE DETECTED ISSUE CAN CAUSE EXPORT AND VISUALIZATION OF FIBERS WITH WRONG POSITIONING IN RELATION THE UNDERLYING STRUCTURAL DATASET. THIS CAN OCCUR WHEN ALL THE FOLLOWING CRITERIA ARE MET: USING TWO DIFFERENT STRUCTURAL DATASETS THAT HAVE IDENTICAL RESOLUTION AND NUMBER OF SLICES; THE STRUCTURAL DATASETS ARE NOT ORIGINALLY ALIGNED BUT ARE THEN ALIGNED THROUGH CO-REGISTRATION IN NORDICBRAINEX. IT OCCURS WHEN CHANGING UNDERLAY IN THE 3D VIEWER FROM ONE TO THE OTHER OF THESE STRUCTURAL DATASETS AFTER PERFORMING THE FIBER TRACKING, AND THEN CREATING AND EXPORTING/VISUALIZING FIBER GROUPS. THE ISSUE IS CAUSED BY THE FIBERS NOT BEING REALIGNED TO THE CO-REGISTERED POSITION OF THE SECOND DATASET WHEN CHANGING THE UNDERLAY, BUT WHEN FIBERS ARE EXPORTED OR VISUALIZED IN THE MPR THE CO-REGISTRATION INFORMATION IS TAKEN INTO ACCOUNT, CAUSING WRONG POSITIONING OF THE FIBERS WHEN LATER VISUALIZED ON A STRUCTURAL DATASET. THE MISALIGNMENT OF THE FIBERS WILL THEN BE IN DIRECT PROPORTION TO THE ALIGNMENT PERFORMED BY THE NORDICBRAINEX COREGISTRATION BETWEEN THE TWO STRUCTURAL DATASETS. TO AVOID THIS ISSUE ONE SHOULD NOT CHANGE UNDERLAY IN THE 3D VIEWER WHEN WORKING WITH DTI AND/OR FIBER GROUPS, UNLESS ONE IS ABSOLUTELY SURE THE DATASETS DIFFER IN RESOLUTION AND/OR NUMBER OF SLICES. THIS BUG IS POTENTIALLY DANGEROUS IF THE MISALIGNMENT IS NOT SEEN BY THE USER. IN MOST CASES IT WILL PROBABLY BE DISCOVERED BY THE USER BECAUSE THE FIBERS DON'T MATCH THE ANATOMICAL STRUCTURES, BUT IF THE MISALIGNMENT IS NOT SO LARGE AND/OR THE FIBER GROUP IS LOCATED IN A PART OF THE BRAIN WHERE IT IS NOT SO EASY TO COMPARE WITH ANATOMICAL STRUCTURES, THE MISALIGNMENT IS MAYBE NOT DISCOVERED. THIS COULD THEREFORE, BE HARMFUL TO THE PATIENT IF THE INFORMATION IS USED IN FOR EXAMPLE NEUROSURGERY PLANNING. | 1
3318840 | INTERFACE USING SCIMAGE/PACS SYSTEM CREATES PROBLEM IN VIEWING ALL LEADS OF EKG WHEN THE EKG REACHES THE POWER CONSOLE. LEAD III, AVR, AVL, AVF AND SPORADIC V LEADS DROP FROM VIEW WHEN INTERFACE OCCURS. EKG MACHINE PRINTS/STORES COMPLETE DATA, SO IT IS AN INTERFACE ISSUE. THIS INTERFACE PROBLEM HAS BEEN OCCURRING AT LEAST ONCE/DAY FOR THE PAST MONTH OR SO. APPROX 50-100 PTS' EKGS HAVE BEEN AFFECTED. IT HAS BEEN REPORTED TO SCIMAGE ON AT LEAST 5 OCCASIONS. THE MOST RECENT BEING FRIDAY (B)(6) 2013. POTENTIAL FOR DELAY IN DIAGNOSTIC INTERVENTION WHEN EKG'S NEED TO BE ACCESSED REMOTELY. TWO OTHER RECENT PTS INVOLVED: SJ (B)(6) 2013. THESE WERE ALSO REPORTED TO SCIMAGE. ALSO SEE (B)(4). | 1
3322188 | INTERFACE USING SCIMAGE/PACS SYSTEM CREATES PROBLEM IN VIEWING ALL LEADS OF EKG WHEN THE EKG REACHES THE POWER CONSOLE. LEAD III, AVR, AVL, AVF AND SPORADIC V LEADS DROP FROM VIEW WHEN INTERFACE OCCURS. EKG MACHINE PRINTS/STORES COMPLETE DATA, SO IT IS AN INTERFACE ISSUE. THIS INTERFACE PROBLEM HAS BEEN OCCURRING AT LEAST ONCE/DAY FOR THE PAST MONTH OR SO. APPROX 50-100 PTS' EKGS HAVE BEEN AFFECTED. IT HAS BEEN REPORTED TO SCIMAGE ON AT LEAST 5 OCCASIONS. THE MOST RECENT BEING FRIDAY (B)(6) 2013. POTENTIAL FOR DELAY IN DIAGNOSTIC INTERVENTION WHEN EKG'S NEED TO BE ACCESSED REMOTELY. TWO OTHER RECENT PTS INVOLVED: (B)(6) 2013. THESE WERE ALSO REPORTED TO SCIMAGE. ALSO SEE (B)(4). | 1
3322189 | INTERFACE USING SCIMAGE/PACS SYSTEM CREATES PROBLEM IN VIEWING ALL LEADS OF EKG WHEN THE EKG REACHES THE POWER CONSOLE. LEAD III, AVR, AVL, AVF AND SPORADIC V LEADS DROP FROM VIEW WHEN INTERFACE OCCURS. EKG MACHINE PRINTS/STORES COMPLETE DATA, SO IT IS AN INTERFACE ISSUE. THIS INTERFACE PROBLEM HAS BEEN OCCURRING AT LEAST ONCE/DAY FOR THE PAST MONTH OR SO. APPROX 50-100 PTS' EKGS HAVE BEEN AFFECTED. IT HAS BEEN REPORTED TO SCIMAGE ON AT LEAST 5 OCCASIONS. THE MOST RECENT BEING FRIDAY, (B)(6)2013. POTENTIAL FOR DELAY IN DIAGNOSTIC INTERVENTION WHEN EKG'S NEED TO BE ACCESSED REMOTELY. TWO OTHER RECENT PTS INVOLVED: (B)(6) 2013 AND (B)(6) 2013. THESE WERE ALSO REPORTED TO SCIMAGE. (B)(4). | 1
2957289 | IT WAS REPORTED THAT WHEN USING ADVANCED VESSEL ANALYSIS, THE AUTOMATIC TRACKING AND MEASUREMENT FOR ABDOMINAL AORTIC ANEURYSM WERE INACCURATE. THE ANEURYSM LENGTH INDICATED BY THE MEASUREMENT OF AUTOMATIC TRACKING WAS LONGER THAN THE REAL ANEURYSM LENGTH. THE AUTOMATIC TRACKING CAN SOMETIMES BE A ZIGZAG LINE INSTEAD OF A SMOOTH LINE, AND THUS INCREASE THE LENGTH OF THE CENTERLINE. THE USER CAN MANUALLY EDIT THE CENTERLINE. THE SURGEON AT THIS SITE DETECTED THIS WRONG MEASUREMENT BEFORE THE SURGERY. NO INJURY WAS REPORTED. | 1
2975132 | IT WAS REPORTED THAT BIOMED WAS CHARGING THE M300 IN THE BIOMED SHOP AND PLUGGED THE UNIT INTO A POWER BAR. IT LATER ALARMED LOW BATTERY SO THE BIOMED MOVED THE PLUG TO ANOTHER OUTLET AND SHORTLY AFTER THEY HEARD A "POP" SOUND. WHEN THE BIOMED LOOKED AT THE CHARGER, THERE WAS SMOKE REPORTED COMING OUT OF THE CASING. THE BIOMED OPENED THE UNIT AND DISCOVERED THAT THE POWER SUPPLY BOARD WAS DAMAGED. THERE WERE NO PATIENT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
3330617 | USER HAVE REPORTED THAT THE MONITOR HAS DISPLAYED SPO2 BETWEEN 89 AND 99 FOR TWO HOURS. THERE WAS A PT INJURY REPORTED. (B)(4). THE DELTA SOFTWARE LOGS DID NOT INDICATE ANY ABNORMAL FUNCTION OF THE DELTA MONITOR. PHOTOS INDICATE THAT THE POD WAS AN OLDER MASIMO SPO2 POD HARDWARE REVISION, MANUFACTURED IN DECEMBER 2009. THE SPO2 POD WAS REQUESTED, BUT NOT PROVIDED TO DRAEGER FOR ANALYSIS. THE ROOT CAUSE COULD NOT BE CONCLUSIVELY DETERMINED BASED ON THE INFORMATION PROVIDED. DRAEGER ASSESSED THE RISK FOR THE REPORTED ISSUE AND DETERMINED THAT THERE ARE NO NEW OR REVISED RISKS. IF THERE IS A COMMUNICATION FAILURE WITH A POD, AN ERROR MESSAGE IS DISPLAYED TO ALERT THE USER. NO ACTIONS ARE PLANNED BY DRAEGER AT THIS TIME. DRAEGER WILL CONTINUE TO MONITOR FOR SIMILAR REPORTS OF THIS ISSUE. | 1
3066576 | IT WAS REPORTED THAT THE M300 INTERMITTENTLY SHUT DOWN WHILE IN USE. THE PT WOULD REPORTEDLY BE DISCHARGED WHEN THE M300 SHUT DOWN. IT WAS FURTHER REPORTED THAT THIS SYMPTOM WAS USUALLY OBSERVED IF THE PT WAS AMBULATORY. THE M300 WAS EXAMINED BY THE HOSPITAL BIOMED WHO REPORTED THE BATTERY APPEARED TO BE FULLY CHARGED AND THAT THE M300 WOULD SOMETIMES DISPLAY A 'CHECK INTERNAL BATTERY' MESSAGE WHEN POWERED BACK UP AFTER A SHUT DOWN. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
3060579 | IT WAS REPORTED THAT THE INTERNAL BATTERY DOES NOT HOLD THE SPECIFIED 3 HOURS WHEN DISCONNECTED FROM MAIN POWER. THERE WAS NO PT INJURY REPORTED. (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
3415944 | IT WAS REPORTED THAT DURING A CERVICAL ARTERIOVENOUS FISTULA SURGERY, CABLES WERE LOCATED IN THE THORACIC REGION OF THE CHILD AND A BURN REPORTEDLY HAPPENED IN A WHITE CABLE LOCATED IN THE STERNUM. DURING SURGERY, NOTHING WAS NOTICED. AT THE END OF SURGERY, AT THE TIME THE CABLES WERE REMOVED FROM THE PT, A BURN WAS NOTED BELOW THE ECG CABLE. THERE WAS NO BURNING DETECTED AT THE ELECTRODE, JUST BELOW THE CABLE THAT WAS ON TOP OF THE PT. THE SURGERY WAS CLOSE TO THE NECK AND THE PLATE WAS PLACED IN THE RIGHT THIGH. THE CABLE WAS REMOVED AND TAKEN TO ENGINEERING. THE ENGINEER MADE SOME CONDUCTIVITY TESTS TO SEE IF IT DETECTED SOME VANISHING POINT ON THE CABLE. NO ANOMALIES WERE OBSERVED. DRAEGER REFERENCE NUMBER: (B)(4). DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 0
3099068 | THE CUSTOMER REPORTED THAT THE S/5 ECG ANESTHESIA MONITOR WAVEFORM WOULD NOT SYNC TO THE INTRA AORTIC BALLOON PUMP (IABP). THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
3147356 | CUSTOMER REPORTED TH PASSPORT 2 MONITOR DISPLAYED INTERMITTENT SPO2 READINGS, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND REPLACED THE UNIT'S SPO2 BOARD AND CONNECTOR. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECS. | 0
3147354 | CUSTOMER REPORTED AN ISSUE WITH THE SPECTRUM MONITOR, WHICH MAY HAVE AFFECTED THE ECG MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE SYSTEM, CORRECTED, INCLUDED REPLACEMENT OF THE UNIT'S CPU. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS. | 0
3147357 | CUSTOMER REPORTED NO DISPLAY ON THE DOM 4 MONITOR. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE DISPLAY AND REPLACED THE LCD ASSEMBLY. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS. | 1
3091678 | CUSTOMER REPORTED THE DOCKING STATION ON A V SERIES MONITOR FAILED WHICH MAY HAVE IMPACTED MONITORING. NO PT INJURY WAS REPORTED. | 1
3263886 | CUSTOMER REPORTED THE DPM 5 MONITOR SUT DOWN, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S POWER SUPPLY. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091448 | THE CUSTOMER REPORTED THERE WERE NO GAS READINGS ON THE DPM 6/7 MONITOR, WHICH MAY HAVE RESULTED IN LOSS OF GAS MONITORING. NO PT WAS REPORTED. MINDRAY SERVICE REPS CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECS. | 1
3172851 | THE CUSTOMER REPORTED A FAILED DPM 6 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3264072 | CUSTOMER REPORTS DPM 6 HAS INTERMITTENT ECG MONITORING WHICH MAY HAVE AFFECTED ECG MONITORING. NO PT INJURY REPORTED. COMPANY REP REPAIRED THE SOLDER ON THE FRONT PANEL CONNECTORS. MONITOR WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3247259 | CUSTOMER REPORTED NO SPO2 READINGS FROM THE DPM 7 MONITOR. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND RESOLDERED THE UNIT'S FRONT PANEL CONNECTORS. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3172845 | THE CUSTOMER REPORTED A FAILED DPM 6 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3091653 | CUSTOMER REPORTED A FAILED DISPLAY ON THE SPECTRUM MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPS REPLACED THE MONITORS MAIN BOARD. PERFORMED ALL FUNCTIONAL AND SAFETY CHECKS. | 1
3091698 | CUSTOMER REPORTED A SPECTRUM MONITOR'S DISPLAY FAILED, RESULTING IN A POSSIBLE LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPS REPLACED THE DISPLAY'S BACKLIGHT. | 1
3091068 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6/7 MONITOR'S MPM MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF ECG MONITORING. NO PT INJURY WAS REPORTED. SERVICE REP REPLACED FRONT PANEL. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3004460 | THE CUSTOMER REPORTED A FAILED PASSPORT V MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3091156 | CUSTOMER REPORTED THE ACCUTORR V MONITOR LOCK UP WHICH MAY HAVE RESULTED IN A LOSS OF MONITORING. NO PT INJURY WAS REPORTED. SERVICE REPS REPLACED THE UNITS CPU BOARD, CALIBRATED AND TESTED THE UNIT TO FACTORY SPECIFICATIONS. | 1
3263906 | CUSTOMER REPORTED THE DPM 6/7 MONITOR DISPLAYED A LOW GAS READINGS, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND SUGGESTED TO THE CUSTOMER TO EXCHANGE THE UNIT'S MODULE. THE CUSTOMER DECLINED TO EXCHANGE. | 1
3172854 | THE CUSTOMER REPORTED THE DPM 7 MONITOR'S MPM MODULE FAILED, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3192537 | THE CUSTOMER REPORTED THAT THE PASSPORT V MONITOR FAILED, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3091140 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR SHUTS DOWN INTERMITTENTLY WHICH MAY RESULT IN A LOSS OF MONITORING. NO PATIENT INJURY WAS REPORTED. SERVICE REPRESENTATIVES REPLACED THE SYSTEM PUMP, BOARD CHIP AND PERFORMED THE REQUIRED PREVENTATIVE MAINTENANCE . CALIBRATED AND SAFETY TESTED THE MONITOR TO FACTORY SPECIFICATIONS. | 1
3091300 | CUSTOMER REPORTED INTERMITTENT LOSS OF THE ECG SIGNAL ON THE PASSPORT V AND TRIO MONITOR WHEN THESE MONITORS ARE USED IN ONE SPECIFIC ROOM IN THEIR FACILITY, WHICH MAY HAVE RESULTED IN A LOSS OF ECG MONITORING. NO PATIENT INJURY AS REPORTED. MINDRAY REPRESENTATIVES WERE UNABLE TO CONFIRM THE REPORTED ISSUE. PROPER OPERATION WAS VERIFIED. (WO - (B)(4)). | 1
3091275 | CUSTOMER REPORTED INACCURATE GAS READINGS FROM THE DPM 6/7 MONITOR. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPS CALIBRATED GAS MODULE. VERIFIED PROPER OPERATION, AND MADE ALL CHECKS WITHIN FACTORY SPECIFICATIONS. | 1
3359985 | CUSTOMER REPORTED THE GAS MODULE III FAILED, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND REPLACED GAS BENCH. PERFORMED GAS CALIBRATION. THE UNIT WAS TESTED, CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091091 | CUSTOMER REPORTED THE PASSPORT V MONITOR HAD NO CO2 WAVEFORMS. NO PT INJURY WAS REPORTED. CUSTOMER REPRESENTATIVE REPLACED THE GAS MODULE. | 1
3360000 | CUSTOMER REPORTED THAT THE ACCUTORR PLUS MONITOR SHUTS DOWN WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. THE CORRECTIVE ACTION WAS REPLACING THE PCB CPU BOARD. | 1
3091771 | CUSTOMER REPORTED THE DPM 6/7 MONITOR'S "LA LEAD" WAS NOT WORKING WHICH MAY HAVE RESULTED IN A LOSS OF ECG MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPAIRED THE SOLDER ON ALL CONNECTORS ON THE FRONT PANEL. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3263801 | CUSTOMER REPORTED INACCURATE READINGS FROM THE PASSPORT V MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP COULD NOT DUPLICATE THE PROBLEM. CORRECTIONS INCLUDED RESEATING ALL CABLES AS A PRECAUTION. UPGRADED SOFTWARE. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY'S SPECIFICATIONS. | 1
3091308 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPLACED THE MAIN POWER SWITCH, CABLE CPU TO SPO2 BOARD AND SPO2 BOARD. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3091708 | CUSTOMER REPORTED THE SPECTRUM MONITOR'S VIEW 12 CARD FAILED, RESULTING IN A POSSIBLE LOSS OF 12 LEAD MONITORING. NO PT INJURY WAS REPORTED. | 1
3091235 | CUSTOMER REPORTED THE V SERIES MONITOR KEPT REBOOTING, WHICH MAY HAVE RESULTED IN LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. THE UNIT IS BEING RETURNED TO MINDRAY'S REPAIR CENTER FOR REPAIR. | 1
3247706 | CUSTOMER REPORTED THE ACCUTORR V MONITOR FROZE, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE UNIT'S CPU BOARD. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3247319 | CUSTOMER REPORTED PASSPORT 2 MONITOR SHUT DOWN, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENTS ON THE CPU BOARD. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3263893 | CUSTOMER REPORTED THE VSERIES MONITOR'S DISPLAY FROZE, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CARRIER BOARD AND CPU SERVICE KIT. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3359758 | CUSTOMER REPORTED THE GAS MODULE III DISPLAY AN ERROR MESSAGE, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S GAS MODULE BENCH. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091093 | CUSTOMER REPORTED THE DPM 6/7 MONITOR'S GAS MODULE WOULD NOT ZERO, WHICH MAY HAVE IMPACTED GAS MONITORING. NO PT INJURY WAS REPORTED. SERVICE REPRESENTATIVE REPLACED THE UNITS GAS MODULE. CALIBRATED UNIT TO FACTORY SPECIFICATIONS. | 1
3091213 | CUSTOMER REPORTED THE ACCUTORR V MONITOR WAS INOPERABLE, WHICH MAY HAVE RESULTED IN LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP REPLACED CPU BOARD AND CALIBRATED THE UNIT TO FACTORY SPECIFICATIONS. | 1
3264016 | CUSTOMER REPORTS THAT ACCUTORR V MONITOR DOES NOT READ PROPERLY, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED UNIT'S SYSTEM CONFIGURATION RESET AND FACTORY DEFAULTS RESTORED. UNIT CALIBRATED, SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091445 | CUSTOMER REPORTED THE DPM 6/7 MONITOR DISPLAYED AND O2 ACCURACY ERROR RESULTING IN A POSSIBLE LOSS OF O2 MONITORING. NO PT INJURY WAS REPORTED. (B)(4) REPS CALIBRATED MODULE TO FACTORY SPECIFICATIONS AND RESOLVED THE ISSUE. | 1
3091566 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPLACED THE CPU BOARD AND RESEATED THE SPO2 CHIPS. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3247533 | CUSTOMER REPORTED LOSS OF COMMUNICATION BETWEEN V SERIES MONITOR AND VPS MODULE, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE MONITOR AND REPLACED CARRIER BOARD. CALIBRATED AND TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3054826 | THE CUSTOMER REPORTED THE DPM 6/7 MONITOR'S GAS MODULE FAILED, WHICH MAY HAVE RESULTED IN A LOSS OF GAS MONITORING. NO PT INJURY WAS REPORTED. | 1
3247322 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE SPO2 CABLE. | 1
3147320 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE SPO2 BOARD. | 1
3091511 | CUSTOMER REPORTED THE DPM 6/7 MPM MODULE DID NOT DISPLAY WAVEFORMS WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPAIRED SOLDER ON THE FRONT PANEL CONNECTORS OF THE DPM 6/7 MONITOR. TESTED AND CALIBRATED TO SPECIFICATION. | 1
3174746 | CUSTOMER REPORTED THE DPM 6/7 MONITOR'S MPM MODULE FAILED WHICH MAY HAVE RESULTED IN A LOSS OF MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3360089 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED ECG MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND REPLACED THE ECG CABLE. | 1
3360018 | CUSTOMER REPORTED THE DPM 6 MONITOR IS PROVIDING FALSE ASYSTOLE ALARMS. NO PT INJURY REPORTED. MINDRAY SERVICE REP EVALUATED THE UNIT AND REPLACED THE MPM MODULE. | 1
3091838 | CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR V MONITOR, WHICH MAY HAVE RESULTED IN LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE CPU BOARD. UNIT CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091266 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR INTERMITTENTLY SHUTS DOWN WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPS REPLACED THE POWER SUPPLY ASSEMBLY. PERFORMED TESTS AND CALIBRATIONS PER SERVICE MANUAL. | 1
3091562 | CUSTOMER REPORTED THE DPM5 MONITOR DISPLAYED A CO2 COMMUNICATION ERROR WHICH MAY HAVE RESULTED IN A POSSIBLE LOSS OF CO2 MONITORING. NO REPORTS OF PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES RESET THE MONITOR TO THE CORRECT SETTINGS AND RESOLVED THE ISSUE. PERFORMED ALL FUNCTIONAL AND SAFETY TESTS. | 1
3247294 | CUSTOMER REPORTED AN ISSUE WITH THE SPECTRUM MONITOR, WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PT INJURY WAS REPORTED COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTION INCLUDED REPLACEMENT OF THE CO2 MODULE. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091543 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT V MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE SPO2 MODULE. THE UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS. IT FUNCTIONED NORMALLY AND WAS RETURNED TO USE. | 1
3091173 | CUSTOMER REPORTED THE DISPLAY ON A SPECTRUM MONITOR FAILED WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. SERVICE REPRESENTATIVES REPAIRED THE UNIT AND RESEATED THE CABLES. CALIBRATED SAFETY TESTED THE MONITOR TO FACTORY SPECIFICATIONS. | 1
3091210 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR HAD A BLANK SCREEN, WHICH MAY HAVE RESULTED IN LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED AND REPROGRAMMED THE CP CHIP AND CPU BOARD. UNIT WAS SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091319 | CUSTOMER REPORTED AN ISSUE WITH THE PANORAMA CENTRAL STATION DISPLAY WHICH MAY HAVE RESULTED IN A LOSS OF AMBULATORY TELEMETRY MONITORING. NO PATIENT INJURY WAS REPORTED. (B)(4) SERVICE REP REPLACED THE HARD DRIVE AND UPDATED SYSTEM SOFTWARE. PASSES ALL FUNCTIONAL CHECKS AND IS READY FOR PATIENT USE. | 1
3091184 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. N O PT INJURY REPORTED. MINDRAY SERVICE REPS REPLACED THE SPO2 BOARD. TESTED, CALIBRATED AND SAFETY CHECKED UNIT TO FACTORY SPECIFICATIONS. | 1
3091904 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR HAD A BLANK SCREEN, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENTS OF THE CPU BOARD AND POWER SUPPLY. UNIT WAS CALIBRATED AND TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091565 | CUSTOMER REPORTED THE SPECTRUM MONITOR'S CO2 FUNCTION WAS NOT WORKING WHICH MAY HAVE RESULTED IN A LOSS OF CO2 MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVE REPLACED THE UNITS CO2 MODULE, PERFORMED FUNCTIONAL TESTS AND RETURNED THE UNIT TO SERVICE. | 1
3264026 | CUSTOMER REPORTS THAT POSITION TWO OF VSERIES MONITOR WILL NOT COMMUNICATE WITH ITS ATTACHED MODULE, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CUSTOMER WAS PROVIDED WITH A LOANER UNIT, WHILE IT IS BEING RETURNED TO THE COMPANY'S REPAIR CENTER. | 1
3131195 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6/7 MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND COULD NOT DUPLICATE THE PROBLEM. UNIT WAS CALIBRATED AND TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091895 | CUSTOMER REPORTED A PASSPORT 2 MONITOR WOULD NOT READ SPO2 DATA, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REP REPLACED THE UNITS SPO2 ASSEMBLY. | 1
3091317 | CUSTOMER REPORTED A FAILED DISPLAY ON THE SPECTRUM MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REP REPLACE THE UNITS CPU. | 1
3091306 | CUSTOMER REPORTED THE SPECTRUM MONITOR SHUT DOWN WHICH MAY HAVE RESULTED IN A LOSS OF MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE CPU AND MOTOR PUMP. CALIBRATED AND SAFETY TESTED THE MONITOR TO FACTORY SPECIFICATIONS. | 1
3263944 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR DISPLAYED A BLANK SCREEN, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND CORRECTED THE PROBLEM BY ADJUSTING A LOOSE CONNECTION ON THE UNIT. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3147440 | CUSTOMER REPORTED A LOSS OF POWER FROM THE V SERIES DOCKING STATION WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND CONFIRMED THE REPORTED PROBLEM. DOCK WAS REPLACED. | 1
3091563 | CUSTOMER REPORTED THE SPECTRUM MONITOR WOULD NOT DISPLAY AN ECG OR HEART RATE WHICH MAY HAVE RESULTED IN A POSSIBLE LOSS OF MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES RECONNECTED THE MAIN BOARD DATA CABLE. PERFORMED ALL FUNCTIONAL AND SAFETY TESTS. | 1
3165760 | THE CUSTOMER REPORTED THAT THE DPM 7'S MPM MODULE FAILED, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3165755 | THE CUSTOMER REPORTED A FAILED ACCUTORR V MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3054886 | THE CUSTOMER REPORTED THAT THE V-SERIES MONITOR FAILED, WHICH MAY HAVE RESULTED IN LOSS OF PRIMARY MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3263800 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR'S DISPLAY, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE FRONT PANEL ASSEMBLY. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091717 | CUSTOMER REPORTED A FAILED DISPLAY ON THE PASSPORT 2 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPLACED THE CPU AND PERFORMED FUNCTIONAL TESTS. | 1
3091181 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6/7 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF GAS MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPS REPLACED THE WATER TRAP RECEPTACLE. CALIBRATED AND UNIT TO FACTORY SPECIFICATIONS. | 1
3091056 | CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR V MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES RESET THE SPO2 CABLE, PERFORMED CALIBRATION, SAFETY AND DIAGNOSTICS TO FACTORY SPECIFICATIONS. | 1
3054788 | THE CUSTOMER REPORTED A FAILED MPM MODULE ON THE DPM 6/7 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3247464 | CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR VPS MODULE THAT MAY HAVE IMPACTED PATIENT MONITORING. THIS INCLUDED THE TEMPORARILY LOSS OF PATIENT DATA. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND THE ISSUE RESULTED FROM A LOOSE CABLE CONNECTION BETWEEN THE V SERIES EXTERNAL MODULE, RACK AND THE UNITS. CHECKED ALL OTHER CABLE CONNECTIONS AND TESTED UNIT PER FACTORY SPECIFICATIONS. UNIT WAS RETURNED TO SERVICE. | 1
3091779 | CUSTOMER REPORTED A BURNING SMELL FROM THE ACCUTORR PLUS MONITOR. NO SPECIFIC SOURCE OF THE BURNING SMELL WAS REPORTED AND NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND REPLACED THE UNIT'S POWER SUPPLY. | 1
3091318 | CUSTOMER REPORTED AN ISSUE WITH THE PANORAMA TELEMETRY SYSTEM, WHICH MAY HAVE RESULTED IN A LOSS OF AMBULATORY TELEMETRY MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REP EVALUATED THE TELEMETRY SYSTEM. CORRECTIONS INCLUDED SETTING UP THE CORRECT TELEMETRY DEVICE IDENTIFICATION AT THE CENTRAL STATION AND THE PROBLEM RESOLVED. | 1
3091493 | CUSTOMER REPORTED A FAILED DISPLAY ON THE PASSPORT 2 MONITOR, WHICH MAY HAVE RESULTED IN A LOSS OF MONITORING. NO PATIENT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPLACED THE HIGH VOLTAGE BOARD ON THE DISPLAY OF THE PASSPORT 2 MONITOR. TESTED UNIT ACCORDING TO FACTORY SPECIFICATIONS. | 1
3147434 | CUSTOMER REPORTED LOSS OF ECG AND SPO2 MONITORING ON THE DPM 7 MONITOR. NO PT INJURY WAS REPORTED. CUSTOMER HAD AN EXPIRED SENSOR. MASIMO CORP, THE OEM FOR THE SENSOR, HAS BEEN INFORMED OF THE REPORTED ISSUE AND IS WORKING DIRECTLY WITH THE CUSTOMER TO REPLACE THE PART. THE LOSS OF ECG MONITORING IS UNDER INVESTIGATION. | 1
3247330 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT V MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE SPO2 MODULE ASSEMBLY. MONITOR WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3091315 | CUSTOMER REPORTED THE SPECTRUM MONITOR DISPLAY WAS NOT WORKING WHICH MAY HAVE RESULTED IN A LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. MINDRAY SERVICE REPRESENTATIVES REPLACED THE DISPLAY CHECKED AND TESTED THE SYSTEM. | 1
3264045 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR WHICH MAY HAVE AFFECTED ECG MONITORING. NO PT INJURY REPORTED. COMPANY REP COULD NOT REPEAT THE REPORTED PROBLEM BUT REWORKED THE CPU. CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS. | 1
3247510 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT XG MONITOR WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT. CUSTOMER WAS PROVIDED WITH A LOANER, WHILE THE MONITOR IS BEING RETURNED TO THE COMPANY'S REPAIR CENTER. | 1
3263980 | CUSTOMER STATES "O2 WILL NOT REGISTER," WHICH MAY HAVE AFFECTED O2 MONITORING. NO PT INJURY REPORTED. COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACING O2 BENCH. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS. | 1
3359961 | CUSTOMER REPORTED AN ISSUE WITH THE BENEVIEW MONITOR, WHICH MAY HAVE AFFECTED O2 MONITORING. NO PT INJURY WAS REPORTED. COMPANY REPRESENTATIVE EVALUATED THE UNIT AND WAS UNABLE TO CORRECT THE ISSUE. CUSTOMER WAS PROVIDED WITH A REPLACEMENT UNIT. | 1
3247267 | CUSTOMER REPORTED AN ISSUE WITH THE BENEVIEW MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. COMPANY REP EVALUATED THE UNIT AND REPLACED THE GAS MODULE. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3091727 | CUSTOMER REPORTED A BENEVIEW MONITOR'S MPM MODULE FAILED, RESULTING IN A POSSIBLE LOSS OF PRIMARY MONITORING. NO PT INJURY WAS REPORTED. | 1
3453240 | WE HAVE HAD A MONITORING DEVICE MANUFACTURED BY LIFE SENSING INSTRUMENT LOCATED IN (B)(4), IN OUR CARDIAC REHAB AND WE BELIEVE THEY ARE NOT A SOUND COMPANY. THE SOFTWARE FAILS/LOCKS UP QUITE OFTEN, THEY HAVE ALMOST NO SUPPORT FOR THE EQUIPMENT, THEY WERE SENDING US SOFTWARE UPGRADES ALMOST EVERY WEEK FOR A WHILE, AND NEVER DID GET OUR LAPTOP WORKING RIGHT WHICH WE PURCHASED FROM THEM. IT HAS IMPAIRED OUR ABILITY TO SAFELY MONITOR PTS. THEY ALSO TOLD US OUR TRANSMITTERS WERE DIGITAL AND IT TURNS OUT THEY ARE ANALOG. I CAN'T BELIEVE THEY ARE ALLOWED TO MAKE MEDICAL EQUIPMENT AND I CAN'T BELIEVE THEY ARE FOLLOWING ALL THE FDA MFG REGULATIONS AND QUALITY CONTROL. THEY ALSO SAY YOU GET FREE LOANERS AND FREE SOFTWARE UPGRADES FOR LIFE ON ALL THEIR QUOTES, BUT THEN WE GOT A BILL FOR THE LOANER, AND THEY SAY WE CAN'T HAVE THE NEW SOFTWARE UPGRADES ON OUR SYSTEM. | 1
3369763 | GE HEALTHCARE HAS RECEIVED NOTIFICATION OF A DEATH OF A FETUS. THE FOLLOWING INFORMATION WAS OBTAINED BY GE HEALTHCARE THROUGH THE LEGAL PROCEEDINGS RELATED TO THIS CASE. MULTIPLE CLINICAL SIGNS PRESENTED THAT INDICATED FETAL DISTRESS DURING THE MONITORING SESSION TO WHICH THE OB TEAM DID NOT TAKE APPROPRIATE ACTIONS: LOW FETAL HEART RATE VARIABILITY- A HEALTHY FETUS HAS HIGH HEART RATE VARIABILITY. MATERNAL HEART RATE WAS VERY CLOSE TO FETAL HEART RATE; THE STRIP CHART INDICATED THE MATERNAL HEART RATE AND FETAL HEART RATE ACTUALLY OVERLAPPED 5 TIMES DURING THE MONITORING SESSION, AS SHOWN BY THE HBC INDICATION- A HEALTHY FETAL HEART RATE IS TYPICALLY HIGHER THAN THE MATERNAL HEART RATE. FETAL HEART RATE ACCELERATED DURING CONTRACTIONS- A HEALTHY FETUS WOULD HAVE HEART RATE DECELERATIONS DURING MATERNAL CONTRACTIONS. MULTIPLE MATERNAL PARAMETERS TO INDICATE THE MOTHER WAS AT HIGH RISK FOR A COMPLICATED BIRTHING PROCESS, I.E. SHE WAS OVERWEIGHT, TACHYCARDIC, AND FEVERISH. THE OB TEAM ATTEMPTED TO USE A FETAL SCALP ELECTRODE (FSE) TO OBTAIN A FETAL HEART RATE DIRECTLY ON FOUR OCCASIONS WITH TWO DIFFERENT ELECTRODES. THEY OBSERVED NO GOOD SIGNAL FROM THE FSE ATTEMPTS AND CONCLUDED THE ELECTRODES WERE DEFECTIVE, INSTEAD OF CONCLUDING THERE WAS A PROBLEM WITH THE FETUS. ONE OF THE OBSTETRICIANS COMMUNICATED THEY DO NOT LOOK AT THE MONITOR STRIP CHART AT ALL. ANOTHER MEMBER OF THE OB TEAM COMMUNICATED THEY DIDN'T KNOW ABOUT OVERLAPPING HEARTRATES AND, THEREFORE, DID NOT UNDERSTAND THE INDICATIONS. THE OBSTETRICIANS COMMUNICATED THEY DID NOT READ THE MONITOR USER'S MANUAL AND DID NOT UNDERSTAND HEARTBEAT COINCIDENCE. THE OBSTETRICIANS APPLIED, REMOVED, AND THEN APPLIED SPO2 AGAIN. WHEN SPO2 WAS APPLIED AND INDICATED A MATERNAL HEART RATE THAT WAS OVERLAPPING WITH THE FETAL HEART RATE, APPROPRIATE ACTIONS WERE NOT TAKEN. THE OBSTETRICIANS DID NOT CONSIDER THE TWO HEART BEATS WERE BOTH FROM THE MOTHER AND THAT THE FETAL HEART BEAT WAS NOT BEING DETECTED. GE HEALTHCARE PROVIDED TRAINING ON THE DEVICE WHICH WAS ATTENDED BY HOSPITAL STAFF SEVERAL MONTHS PRIOR TO THE INCIDENT. THE HOSPITAL IS RESPONSIBLE FOR ENSURING THEIR CLINICIANS ARE PROPERLY TRAINED ON THE DEVICE THROUGH PERSONNEL TURNOVER CYCLES, ETC. THE GE TRAINING WAS PROVIDED ONE TIME ONLY AND WAS NOT PURCHASED AGAIN FOR NEW CLINICIANS THAT DID NOT RECEIVE THE ORIGINAL TRAINING. IT WAS CONCLUDED THAT THE ROOT CAUSE WAS USER ERROR INTERPRETING THE CLINICAL SITUATION OF THE MOTHER AND THE FETUS.><THE LEGAL COMPLAINT THAT GENERAL ELECTRIC COMPANY RECEIVED ALLEGES THE FOLLOWING: "THE COROMETRICS MONITOR MADE IT APPEAR TO THE HEALTHCARE PROVIDERS THAT THE FETAL HEART RATE WAS BEING MONITORED THROUGHOUT LABOR AND WAS NORMAL." "HOWEVER, AT SOME POINT HOURS BEFORE DELIVERY, HIS FETAL HEART RATE BECAME DISTRESSED. INSTEAD OF PICKING UP THIS DISTRESS, THE COROMETRICS FETAL HEART MONITOR MADE A SMOOTH TRANSITION TO THE MATERNAL HEART RATE, CONFUSING THE HEALTHCARE PROVIDERS INTO BELIEVING THE BABY'S HEART WAS FINE." GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED.><NO INFORMATION HAS BEEN PROVIDED TO GE HEALTHCARE BY THE HOSPITAL ON THE STATUS OF THE UNIT. AT THIS TIME, IT IS NOT KNOWN IF THE UNIT WAS TAKEN OUT OF SERVICE OR IF IT CONTINUES TO BE USED WITH PATIENTS. NO FURTHER DETAILS ABOUT THE ALLEGED DEVICE OR THE EVENT HAVE BEEN PROVIDED TO GE HEALTHCARE. THERE ARE NO SERVICE RECORDS FOR THE DEVICE IN THE GE HEALTHCARE DATABASE. THEREFORE, WITHOUT SUFFICIENT INFORMATION ABOUT THE EVENT, OR THE EVALUATION OF THE ALLEGED DEVICE, IT IS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THE ALLEGED ISSUE. | 1
850560 | A PD PT CONTACTED BAXTER TECHNICAL SVC CTR REGARDING A LOW DRAIN VOLUME ALARM DURING INITIAL DRAIN OF THERAPY USING THE HOME CHOICE SYS. THE DRAIN VOLUME WAS 116ML. THE SVC REP ASKED THE HOME PT TO CHECK PT LINE, BUT THE HOME PT WAS NOT ABLE TO CHECK. THE HOME PT WAS IN THE HOSP AND TRIED TO GET THE NURSE. HE COULD NOT GET HELP. THE HOME PT HUNG UP THE PHONE. ADD'L INFO OBTAINED FROM A PD NURSE. THE NURSE WAS CONTACTED IN ORDER TO OBTAIN ADD'L INFO REGARDING THE LOW DRAIN ALARM AND THE STATUS OF THE HOME PT. THE NURSE STATED THAT THE PT HAD EXPIRED AT MOREHEAD HOSP DUE TO A COMPLICATION OF SEPSIS. THE NURSE, ALSO, INDICATED THAT THE PT WAS ADMITTED TO THE HOSP IN 2007 BECAUSE PERITONITIS WAS SUSPECTED. REGARDING THE PT MEDICAL HISTORY THE NURSE INDICATED THAT THE PT SUFFERED FROM DIABETES AND HAD A LEFT UNILATERAL AMPUTATION LOCATED BELOW THE KNEE. NO OTHER INFO AVAILABLE AT THIS TIME. BAXTER INVESTIGATION REVEALED THAT THE PT EXPIRED DUE TO SEPSIS. A COPY OF THE FACILITY'S ESRD DEATH NOTIFICATION WAS PROVIDED. FURTHER INFO REGARDING MEDICAL HISTORY, PRIMARY AND SECONDARY CAUSE OF DEATH HAS BEEN REQUESTED FROM MOREHEAD HOSP, BUT NOT REC'D YET. BAXTER REQUESTED THE INSTRUMENT FOR EVAL. THE RESULTS WILL BE PROVIDED UPON COMPLETION. | 0
2755717 | THE CUSTOMER REPORTED THE GE MEDICAL DEVICE DID NOT HAVE ANY AUDIO. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2798567 | THE CUSTOMER REPORTED THERE WAS NO AUDIO ON THE DINAMAP PROCARE 100 PT MONITOR DUE TO SPEAKER HARDWARE FAILURE. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. INVESTIGATION REVEALED THE ROOT CAUSE OF THE FAILURE WAS DUE TO THE SPEAKER BEING 6 YEARS OLD AND BEYOND ITS INTENDED LIFE. ACCORDING TO THE SPEAKER MANUFACTURER, JEMA, THE INTENDED LIFE OF THE SPEAKER IS 1-2 YEARS. SPEAKER ASSEMBLY WAS REPLACED WITH SPEAKER FIELD REPLACEABLE UNIT (B)(4) AND THIS CORRECTED THE PROBLEM OF NO SOUND. | 1
2948447 | THE CUSTOMER REPORTED THE PT DATA MODULE (PDM) DISCONNECTED FROM THE MODULE RACK AND WHEN THE HOSPITAL BIOMED REINSERTED THE PDM, IT WAS NOTED THAT THE LATCH WAS NOT WORKING. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. OCCUPATION OF REPORTER IS UNK. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2779527 | A POSSIBLE MALFUNCTION OF THE DEFIBRILLATION SYNC FUNCTION OF THE DEVICE WAS REPORTED. INITIAL VISUAL INSPECTION OF THE DEVICE SHOWED DAMAGE TO THE UNIT AND NEED FOR REPAIR. NO DEATH OR SERIOUS INJURY IS ASSOCIATED WITH THIS RECORD. GE HEALTHCARE¿S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
3308633 | THE CUSTOMER REPORTED AN SPO2 ESCALATION ALARM ISSUE ON THE GE COMPACT MONITOR F-CM1. THE COMPLAINT REFERS TO THE USE OF A (B)(4) (NON-GEHC) CABLE AND A PROBE WITH THE GEHC DEVICE. IT IS REPORTED A CHILD DID NOT GET ENOUGH OXYGEN AND RECEIVED A BRAIN DAMAGE. UNDER (B)(4) LAW, PT INFORMATION IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSPITAL. MANUFACTURING DATE IS UNKNOWN. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3525439 | THE CUSTOMER ALLEGES THE PROCARE MONITOR B20 DID NOT VISUALLY OR AUDIBLY ALARM AND THE PT EXPIRED. DEVICE LOGS AND STRIPS WERE REQUESTED FROM THE CUSTOMER. AT THIS TIME, LOGS FOR THE TIME IN QUESTION HAVE NOT BEEN RECEIVED NOR THE RESULTS FROM THE DEVICE BEING TESTED. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2978463 | RADIOLOGIST ATTEMPTED TO OPEN CASE (B)(6) BUT RECEIVED A MESSAGE THAT SOMEONE ELSE WAS ALREADY READING IT. SHE MOVED ON AND OPENED ANOTHER CASE (B)(6). THE CASES AND REPORT PANE LOADED AS NORMAL WITH BOTH PACS DISPLAYS LOADED WITH IMAGES FROM THE CURRENT CASE. IN THE TOP LEFT OF PACS VIEWER, THE CURRENT CASE WAS LISTED AS WELL AS WHAT LOOKED LIKE A PRIOR CT FOR THE SAME PT. HOWEVER ON CLOSER EXAMINATION, WHEN THE RADIOLOGIST DRAGGED THE SERIES DOWN FORM THE TOP LEFT OF THE PACS VIEWER WHERE PRIORS ARE LISTED, THE PRIOR ON THE LIST WAS ACTUALLY THE CASE FROM THE PREVIOUS PT, (B)(6). PICOM POWERCONSOLE 5.24.0.12 AND XYZ VIEWER 5.5.4.8 - ERROR REPORTED (B)(4) 2013. | 1
3403666 | AGFA SUBMITTED MDR REPORT # 1225058-2013-00001 TO THE FDA ON (B)(4) 2010 FOR A CUSTOMER IN THE U.S.A AND OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT), BUT WITH A CUSTOMER IN (B)(6). IN (B)(6) 2012, THE (B)(6) CUSTOMER REPORTED TO AGFA THAT SOME FIELDS FOR CUSTOM FINDING "MORPHOLOGY" UNDER "AORTIC VALVE" WERE NOT BEING SAVED. SPECIFICALLY, THE FOLLOWING REPORT OPTIONS: "TRANSESOPHOGEAL ECHO (TEE) REPORT --> AORTIC VALVE --> MORPHOLOGY: WHEN CHOOSING "YES" FOR RHEUMATIC VALUE, AND "PROBABLE RHEUMATIC AV'" FOR PROBABILITY RHEUMATIC, AND "COMMISSURAL FUSION, IN ROLLING.. (INCOMPLETE)... FOR RHEUMATIC MORPHOLOGY, THEN WHEN CLICKING ON SHOW REPORT, NONE OF THESE GET UPDATED ON THE TEE REPORT. WHEN CLICKING ON THE AORTIC VALVE AGAIN, ALL THE CHOSEN VALUES DISAPPEAR. IN (B)(6) 2012, AN AGFA CARDIO CLINICAL REPORTING SPECIALIST DETERMINED THE SITE USERS WERE SELECTING SPECIFIC PORTIONS OF COMPOUND SENTENCES WHEN CREATING A CARDIOVASCULAR REPORT, BUT WHEN REVIEWED, THE USER REALIZED THOSE SPECIFIC COMPOUND SENTENCES WERE NOT POPULATING IN THE CARDIOVASCULAR REPORT. THE CAUSE OF THE ISSUE WAS DUE TO A COMPOUND SENTENCE THAT WASN'T CONFIGURED CORRECTLY. AGFA DOES NOT SUPPORT COMPOUND STATEMENTS IN TEE AND COULD NOT MAKE ANY CHANGES TO ANY OF THE COMPOUND STATEMENTS. AGFA RECOMMENDED THE SITE TURN OFF THE COMPOUND STATEMENT FOR RHEUMATIC AORTIC VALVE BY CHANGING THE GUI TYPES FROM THE COMPOUND FEATURE TO THE RADIO FEATURE. THE COMPOUND SENTENCE WAS DISABLED. NO REPORTS WERE AFFECTED DUE TO THIS EVENT AND NO HARM TO PTS HAS BEEN REPORTED. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. (B)(4).><. CORRECTED 510K # FROM K050858 TO K050228 ADDED PRODUCT CODE : LLZ - | 1
2996784 | A RADIOLOGIST LAUNCHED AN ABDOMEN STUDY OF A (B)(6) AND WAS INTERRUPTED WHILE DICTATING. HE WAS REQUESTED TO REVIEW A SECOND ABDOMEN STUDY, THIS ONE OF A (B)(6) PATIENT ((B)(6)). WHEN THE RADIOLOGIST RETURNED TO COMPLETE HIS DICTATION OF PATIENT (B)(6)'S CASE, PATIENT (B)(6)'S IMAGES WERE STILL BEING DISPLAYED AND THE RADIOLOGIST COMPLETED HIS DICTATION OF PATIENT AGAINST PATIENT (B)(6)'S IMAGES. AS A CONSEQUENCE, PATIENT (B)(6) UNDERWENT AN UNNECESSARY SURGICAL PROCEDURE, I.E., APPENDECTOMY. | 1
3207503 | IT WAS REPORTED THAT A PHYSICIAN IN ESTONIA ENCOUNTERED WHAT HE SUSPECTED TO BE AN INCORRECT TRISOMY RISK CALCULATION ON VIEWPOINT 6.3.1 WHICH CORRELATED TO A (B)(6) WOMAN WHO IN REALITY WAS (B)(6). THE ISSUE WAS IMMEDIATELY OBVIOUS TO THE CAREGIVER, AND THERE WAS NO REPORT OF THE PT OR FETUS INJURY. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE INVESTIGATION RESULTS ARE AVAILABLE. THE DEVICE MANUFACTURE IS UNK AT THIS TIME. | 1
3040415 | THE MULTI-CHAMBER CFA ISSUE WAS REPORTED INTERNALLY. IT RELATES TO OUT OF SYNCHRONIZATION SEGMENTATION MODIFICATIONS THAT DO NOT ALIGN TO OUTPUT NUMERICAL VALUES. IT WAS DISCOVERED AND IN CERTAIN WORKFLOW EDITING THE CHAMBER SEGMENTATION DID NOT RESULT IN CORRESPONDING CHANGES TO DATA SHOWN WITHIN TIME VOLUME GRAPHS, INDEXES, AND RESULTS. THE RISK ASSOCIATED WITH THIS ISSUE IS THAT PT MONITORING, THERAPIES OR TREATMENT MAY BE PROVIDED ERRONEOUSLY BASED ON INCORRECT DATA REPORTED IN THE VITREA MULTI-CHAMBER CFA APPLICATION. A RARE BUT IRREVERSIBLE SERIOUS INJURY MAY OCCUR UNDER NORMAL USE CONDITIONS. THE PRODUCT HAS BEEN IN COMMERCIAL USE FOR APPROXIMATELY FOUR YEARS, WITH 102 SYSTEMS LICENSED WITHIN THE VITREA USER BASE. NO COMPLAINT OR FIELD REPORT HAS BEEN REPORTED. | 1
3020754 | THE CERNER MILLENNIUM POWER CHART FAILS TO TRACK LIFE CRITICAL IMPLANTED DEVICES SUCH AS PACEMAKERS, WHOSE BATTERIES HAVE FINITE LIFE. THIS PT WAS HOSPITALIZED FOR TWO WEEKS WITH CERNER EHR AND CPOE SYSTEMS DEPLOYED FOR THE CARE. AT NO POINT WAS THERE A WARNING THAT THE BATTERY WAS AT END OF LIFE. AFTER THE PT WAS DISCHARGED FROM THAT HOSPITAL AND THE CARE MANAGED WITH PAPER RECORDS WAS IT NOTED BY AN ASTUTE PHYSICIAN THAT THE PULSE GENERATOR MUST BE URGENTLY REPLACED, THE CERNER SYSTEM FAILS THE PT'S WITH PACEMAKERS BECAUSE IT HAS NO FUNCTIONALITY FOR TRACKING THE LIFE OF THE BATTERY OR ITS IMPLANT DATE. IT IS URGED THAT THERE BE A WARNING ON THE SYSTEM THAT THE EHR DOES NOT RECORD PACEMAKER OR AICD DEFIBRILLATOR SPECIFIC TRACKING INFO. | 1
2874021 | THE PROBE FOR COMPARTMENTAL PRESSURE MONITORING SYSTEM STOPPED FUNCTIONING WHILE IN USE. THE PHYSICIAN MOVED THE TIP AROUND AND WAS ABLE TO GET A READING WHEN THE DEVICE STOPPED FUNCTIONING AGAIN. THE TEST WAS CANCELLED WHEN THE PHYSICIAN COULD NOT ACHIEVE ACCURATE READINGS FROM THE DEVICE. THE PATIENT WAS TRANSFERRED TO ORTHOPEDICS. THE PHYSICIAN WAS AWARE THAT THE PATIENT'S PRESSURE WAS ELEVATED PRIOR TO TESTING. MANUFACTURING ENGINEER EVALUATED THE DEVICE AND INDICATED THAT ((B)(4)) MANUFACTURED THE PROBE FOR COMPARTMENTAL PRESSURE MONITORING SYSTEM. THE LOT WAS PROCESSED PER SPECIFICATION REQUIREMENTS, AND PREVIOUSLY INSPECTED / CONFORMED TO THE SYNTHES INSPECTION SHEET. THE COMPLAINT ISSUE WAS DUE TO AN UNKNOWN CAUSE. A RESPONSE FROM (B)(4) DATED (B)(4) 2013 AND FURTHER COMMUNICATION, REVEALED NO FAILURE WAS FOUND DURING THE EVALUATION, AND (B)(4) WAS UNABLE TO FIND ANY FAULT WITH THE PROBE.><THIS DEVICE IS USED FOR DIAGNOSIS NOT TREATMENT. SUBJECT DEVICE HAS BEEN RECEIVED AND IS CURRENTLY IN THE EVALUATION PROCESS. INVESTIGATION IS ON GOING; NO CONCLUSION COULD BE DRAWN. A REVIEW OF THE DEVICE HISTORY RECORDS SHOWED THAT THE PROBE FOR COMPARTMENTAL PRESSURE MONITORING SYSTEM WAS PROCESSED PER SPECIFICATION REQUIREMENTS. THE LOT WAS INSPECTED AND ACCEPTED PER SPECIFICATION. THEY WERE NO COMPLAINT RELATED ISSUES WITH THIS LOT. | 0
3185619 | A MEDICAL DEVICE, MODEL "KOS" (PROPERTY OF THE DISTRIBUTOR), WAS PRESENT IN THE FACILITY TO TEMPORARILY SUBSTITUTE THE FINAL USER'S KOS UNIT, WHICH WAS AT DISTRIBUTOR FACILITY FOR SVC REPAIR. THE SUBSTITUTING KOS WAS NOT PROPERLY CALIBRATED CAUSING THE TEMP TO EXCEED THE DESIGNED LEVEL. THIS RESULTED IN THE PARAFFIN HEATING EXCESSIVELY AND EVENTUALLY BOILING OFF. ONCE PROPERLY CALIBRATED THE KOS UNIT WORKED PROPERLY, AS DESIGNED. TISSUES (45 SPECIMENS) WERE SUB-OPTIMIZED DUE TO THE ADDITIONAL HEATING DURING PARAFFIN INFILTRATION, HOWEVER USER INDICATED THAT DIAGNOSIS WAS POSSIBLE. CUSTOMER CONFIRMED ON (B)(6) 2013 THAT SOME/ALL OF AFFECTED SPECIMENS WERE USED FOR PT DIAGNOSIS. THEREFORE WE HAVE DESIGNATED THIS EVENT AS A "PRODUCT PROBLEM" AND NOT AN "ADVERSE EVENT". | 0
2915696 | REPORTER INDICATED THAT A VNS DELL X50 COMPUTER WAS NOT HOLDING A CHARGE. USING A DIFFERENT KNOWN GOOD CHARGING CORD DID NOT RESOLVE THE ISSUE. THE COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE PENDING ANALYSIS.><PRODUCT ANALYSIS WAS COMPLETED ON THE RETURNED VNS COMPUTER AND FLASHCARD. AN ANALYSIS OF THE RETURNED HANDHELD COMPUTER WAS ABLE TO VERIFY THAT THE MAIN BATTERY WAS ABLE TO HOLD A CHARGE. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE MAIN BATTERY WAS INSERTED BACKWARDS. AS A RESULT, THE HANDHELD WAS UNABLE TO BE POWERED USING THE RETURNED BATTERY. NO FURTHER ANOMALIES WERE IDENTIFIED.NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3021285 | ON (B)(6) 2013 PRODUCT ANALYSIS WAS COMPLETED ON THE FLASHCARD AND HANDHELD. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND IT WAS IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A SWOLLEN MAIN BATTERY THAT PREVENTED MAIN BATTERY COVER FROM SEATING PROPERLY AND REGISTERING A FALSE OPEN BATTERY LATCH CONDITION (REPORTED ON MFR. REPORT # 1644487-2013-00521). ALTHOUGH THE MAIN BATTERY WAS SWOLLEN, IT WAS ABLE TO HOLD A CHARGE AND POWER THE HANDHELD FOR OVER AN HOUR. ALSO DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD AN OPEN ELECTRICAL CONNECTION IN THE POWER CABLE. AS A RESULT OF THE OPEN WIRE CONNECTION, THE HANDHELD WOULD RECEIVE POWER FROM THE AC ADAPTER INTERMITTENTLY. PRODUCT ANALYSIS DETERMINED THAT THE CAUSE FOR THE BREAK IN THE SERIAL DATA CABLE IS UNKNOWN. IT WAS INITIALLY REPORTED THAT THE HANDHELD DOES NOT HOLD A CHARGE ANYMORE AND MUST BE PLUGGED INTO THE OUTLET IN ORDER TO USE IT. THE MANUFACTURER'S CONSULTANT ATTEMPTED TO CHARGE THE HANDHELD USING HIS POWER ADAPTER FOR QUITE SOME TIME BUT THE DEVICE STILL WOULD NOT HOLD A CHARGE. IT WAS NOTED THAT NO USER MISHANDLING OCCURRED; THE HANDHELD IS ALWAYS KEPT SECURELY IN THE OFFICE WHEN NOT IN USE. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2966510 | IT WAS REPORTED ON (B)(6) 2013 THAT A VNS TREATING PHYSICIAN'S HHD WAS OVERHEATING WHILE CHARGING. THE LEVEL OF HEATING OBSERVED WAS HIGH WHILE THE HANDHELD WAS CHARGING AND IT COULD BE FELT WHEN THE HANDHELD WAS TOUCHED. THERE WAS NO INDICATION OR REPORT THAT ANYONE WAS BURNED OR INJURED WHEN TOUCHING THE HANDHELD. IT WAS ALSO INDICATED THAT THERE WAS NO USER MISHANDLING OR TRAUMA THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE HEATING AND THE HANDHELD WAS STORED NORMALLY INSIDE A SHELF. THE HANDHELD WAS RETURNED BACK TO THE MANUFACTURER ON (B)(6) 2013 AND IS CURRENTLY UNDERGOING PRODUCT ANALYSIS.><PRODUCT ANALYSIS OF THE RETURNED DELL X5 HANDHELD DEVICE WAS COMPLETED AND APPROVED ON (B)(6) 2013 AND THE OVERHEATING WHILE CHARGING WAS NOT CONFIRMED. THE HANDHELD'S MAIN BATTERY WAS FULLY RECHARGED SUCCESSFULLY USING A POWER SUPPLY ADAPTER. WHILE CHARGING, THE HANDHELD DEVICE TEMPERATURE WAS MONITORED. THE HANDHELD DEVICE TEMPERATURE MEASURED BETWEEN 74.3F AND 77.4F, WHICH DEMONSTRATES THERE IS NO OVERHEATING CONDITION PRESENT. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2950376 | ON (B)(6) 2013, IT WAS REPORTED THAT THE PHYSICIAN'S PROGRAMMING SYSTEM "NEEDS REPAIR". HE PHYSICIAN'S NURSE STATED THAT THE PROGRAMMING SYSTEM IS "NOT PICKING UP A SIGNAL AT ALL". THE PHYSICIAN HAS OTHER PROGRAMMING SYSTEMS THAT THEY HAVE BEEN ABLE TO SUCCESSFULLY USE TO INTERROGATE PATIENTS. ATTEMPTS FOR FURTHER INFORMATION ARE UNDERWAY BUT NO ADDITIONAL INFORMATION HAS BEEN RECEIVED TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2013 WHEN IT WAS REPORTED THAT THE PROGRAMMING SYSTEM IS NOW WORKING PROPERLY. ONE OF THE PHYSICIANS AT THE OFFICE FOUND A COUPLE OF THINGS WRONG WITH THE HANDHELD; THE BACK LATCH WAS OPEN AND THE FLASHCARD WASN'T INSERTED ALL THE WAY. HE FIXED THESE COUPLE OF THINGS AND NOW THE PROGRAMMING SYSTEM WORKS FINE AND THE OFFICE WANTS TO KEEP IT. . | 0
3260419 | AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. DURING THE ANALYSIS IT WAS VERIFIED THAT THE TOUCHSCREEN DISPLAY WAS UNRESPONSIVE. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED.><ON (B)(6) 2013 INFORMATION WAS RECEIVED FROM THE REPORTER THAT THE PHYSICIAN¿S HANDHELD COMPUTER WAS NOT RESPONDING TO ANY TAPS OR TOUCHES ON THE SCREEN. A HARD RESET WAS PERFORMED, BUT WHEN THE HANDHELD COMPUTER REACHED THE SCREEN ALIGNMENT SCREEN, THE PHYSICIAN ATTEMPTED TO ALIGN THE SCREEN BUT IT WOULD NOT RESPOND. THE SCREEN WAS WIPED DOWN WITH A DRY CLOTH IN AN ATTEMPT TO CLEAN THE SCREEN AND THE CREVICE AROUND THE SCREEN. AFTER BEING CLEANED ANOTHER HARD RESET WAS PERFORMED ON THE HANDHELD COMPUTER AND THE SCREEN WOULD STILL NOT RESPOND. A REVIEW OF THE MANUFACTURER¿S DEVICE RECORDS WAS PERFORMED AND CONFIRMED THAT THE HANDHELD COMPUTER PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. THE PHYSICIAN¿S PROGRAMMING WAND, HANDHELD COMPUTER, AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER AND ARE PENDING ANALYSIS. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS CONFIRMED ALL FUNCTIONAL TEST WERE PASSED BY THE HANDHELD COMPUTER PRIOR TO DISTRIBUTION. DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2966067 | ON (B)(6) 2013, IT WAS REPORTED THAT A HANDHELD DEVICE WAS NOT WORKING ANYMORE. DESPITE BEING PLUGGED IN AND CHARGED, THE DEVICE DID NOT WORK. THE DEVICE WOULD NOT TURN ON. THE HANDHELD DEVICE, POWER SUPPLY, SERIAL CABLE, AND SOFTWARE FLASHCARD WERE RETURNED ON (B)(6) 2013 AND ARE CURRENTLY UNDERGOING PRODUCT ANALYSIS.><AN ANALYSIS WAS PERFORMED ON THE HANDHELD, AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD, AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
3012768 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(4) 2013 WHEN THE PHYSICIAN REPORTED THAT THE DATE IT WAS FIRST NOTICED THAT THE PROGRAMMING HANDHELDS WERE NOT HOLDING A CHARGE WAS (B)(4) 2013 (OTHER TWO PROGRAMMING SYSTEMS NOT HOLDING A CHARGE ARE CAPTURED ON MFR. REPORT # 1644487-2013-00743 AND 1644487-2013-00745). THE SAME WALL OUTLET WAS USED TO CHARGE ALL THREE HANDHELDS. THE SAME POWER CHARGER WAS NOT USED TO CHARGE ALL THREE HANDHELDS AS THE PHYSICIAN¿S OFFICE HAD TWO CHARGING UNITS. THE HANDHELDS WERE REPORTED TO NOW BE ABLE TO HOLD A CHARGE SINCE A NEW POWER CHARGER WAS SENT TO THEM. PRODUCT ANALYSIS ON THE SERIAL CABLE WAS COMPLETED ON (B)(4) 2013. AN ANALYSIS WAS PERFORMED ON THE RETURNED AC ADAPTER AND THE NO ANOMALIES ASSOCIATED WITH THE AC ADAPTER WERE NOTED DURING TESTING USING A KNOWN GOOD HANDHELD AND SERIAL CABLE. THE AC ADAPTER PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(4) 2013 IT WAS REPORTED THAT THE PHYSICIAN'S THREE PROGRAMMING SYSTEMS ARE NOT HOLDING A CHARGE (OTHER TWO PROGRAMMING SYSTEMS NOT HOLDING A CHARGE ARE CAPTURED ON MFR. REPORT # 1644487-2013-00743 AND 1644487-2013-00745). IT WAS UNKNOWN IF THE SAME POWER CHARGER WAS BEING USED TO CHARGE ALL THREE HANDHELDS BUT IT WAS BELIEVED THAT THEY USED THE SAME OUTLET TO CHARGE THEM. THE PHYSICIAN LATER REPORTED THAT ALL THREE HANDHELDS ARE WORKING NOW. IT WAS STATED THAT THEY HAVE ONE POWER CORD THAT WAS GOING TO BE RETURNED TO THE MANUFACTURER FOR PRODUCT ANALYSIS. THE POWER CORD WAS RECEIVED FOR PRODUCT ANALYSIS ON (B)(4) 2013. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT YET BEEN COMPLETED. GOOD FAITH ATTEMPTS FOR FURTHER INFORMATION FROM THE PHYSICIAN'S OFFICE HAVE BEEN UNSUCCESSFUL. . | 0
2880184 | THE PROGRAMMING SYSTEM WAS RETURNED ON (B)(6) 2013 AND PRODUCT ANALYSIS HAS SINCE BEEN COMPLETED. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS ALSO PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. PRODUCT ANALYSIS ON THE RETURNED PROGRAMMING WAND ALSO REVEALED NO ANOMALIES. THE WAND PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED BY THE PHYSICIAN THAT HE WAS HAVING ISSUES WITH HIS PROGRAMMING SYSTEM. HE DID INDICATE THAT HE HAD A SECOND SYSTEM, SO NO PATIENTS WERE AFFECTED BY THIS ISSUE, BUT HE WAS HAVING TROUBLE COMMUNICATING WITH THE PATIENT'S GENERATORS. THE HANDHELD WAS CONFIRMED TO BY UNPLUGGED AT THE TIME OF USE AND THE WAND BATTERY LIGHT STAYED ON FOR >25 SECONDS INDICATING THAT THE WAND BATTERY HAD SUBSTANTIAL BATTERY LIFE. THE ISSUE WAS SAID TO HAVE STARTED ON (B)(6) 2012 THERE WERE NO OTHER DETAILS AVAILABLE ON THE MATTER. TROUBLESHOOTING WAS PERFORMED BY A MANUFACTURER'S REPRESENTATIVE. IT WAS INDICATED THAT WHEN THE WAND WAS SUBSTITUTED A DEMO GENERATOR COULD BE COMMUNICATED WITH; HOWEVER IT WAS ALSO NOTED THAT THE PHYSICIAN'S WAND WAS ABLE TO COMMUNICATE INITIALLY, BUT THEN STOPPED. IT WAS NOTED THAT THAT THE CONNECTION BETWEEN THE HANDHELD AND SERIAL CABLE WAS LOOSE, SO THE COMMUNICATION APPEARED TO BE SUCCESSFUL WHEN HELD IN PLACE WITH THE PHYSICIAN'S WAND. A REPLACEMENT WAS SENT TO THE PHYSICIAN, HOWEVER THE SUSPECT PROGRAMMING SYSTEM HAS NOT BEEN RETURNED TO DATE. . | 0
2785335 | IT WAS REPORTED BY A REPRESENTATIVE, THAT DURING THE UPGRADE OF THE PHYSICIAN'S HANDHELD HE WAS RECEIVING A MESSAGE STATING "UNRECOGNIZED CARD, ENTER DEVICE DRIVER FOR CARD." OR "ENTER NAME OF DEVICE." TWO DIFFERENT FLASHCARDS WERE TRIED WITH THE SAME RESULT. THE SOFTWARE COULD NOT BE INSTALLED, SO THE HANDHELD WAS RETURNED. ANALYSIS ON THE HANDHELD HAS SINCE BEEN COMPLETED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE VNS SOFTWARE WOULD NOT INSTALL INTO THE HANDHELD WHEN THE RETUNED FLASHCARD WAS INSERTED. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH 2 BENT PINS IN THE HANDHELD FLASHCARD SLOT. ONCE THE PINS WERE STRAIGHTENED, NO FURTHER ANOMALIES WERE IDENTIFIED. DURING ANALYSIS OF THE RETURNED FLASHCARDS, NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE RETURNED FLASHCARDS AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL.><FOLLOW UP WAS PERFORMED WITH THE MANUFACTURER'S REPRESENTATIVE; HOWEVER HE WAS UNABLE TO SAY IF THE FLASHCARD INADVERTENTLY INSERTED BACKWARDS BEFORE OR DURING THE UPGRADE ATTEMPT. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. ADDITIONALLY WHILE DEVICE FAILURE OCCURRED, THE LIKELY CAUSE IS USER ERROR/ MISHANDLING OF THE DEVICE. | 0
3318968 | IT WAS ALSO INITIALLY REPORTED ON THE RETURN PRODUCT FORM, PRIOR TO THE INITIAL MDR SUBMISSION, THAT THE SCREEN WAS FREEZING WHILE TRYING TO INTERROGATE AND CHANGE SETTINGS. HOWEVER, PRODUCT ANALYSIS WAS PERFORMED ON BOTH THE HANDHELD AND SOFTWARE AND FOUND NO ANOMALIES. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(6) 2013, IT WAS REPORTED THAT THE PROGRAMMING SYSTEM WAS WORKING PROPERLY IN THE MORNING WHEN A PATIENT WAS SEEN; HOWEVER, WHEN TRYING TO REVIEW THE PAST DATA ON THE HANDHELD, THE SCREEN FROZE. A HARD RESET WAS PERFORMED, BUT UPON REBOOTING THE DEVICE, THE SOFTWARE BECAME STUCK IN THE ALIGN SCREEN LOOP. ANOTHER HARD RESET DID NOT RESOLVE THE ISSUE. THE PHYSICIAN ATTEMPTED TO CLEAN THE SCREEN AND REMOVE ANY DEBRIS STUCK IN THE SIDES; HOWEVER, THE PROBLEM STILL PERSISTED. THE HANDHELD WAS RETURNED TO THE MANUFACTURER AND A NEW ONE WAS SHIPPED OUT FOR THE PHYSICIAN. REVIEW OF THE HHD DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. TYPE OF DEVICE, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED WRONG DEVICE TYPE. DEVICE MANUFACTURE DATE, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED THE MANUFACTURE DATE FOR THE INCORRECT SUSPECT MEDICAL DEVICE.><REPORT SOURCE, CORRECTED DATA: FOLLOW-UP REPORT #1 INADVERTENTLY DID NOT INDICATE WHEN THE NEW INFORMATION CAME FROM. COMPANY REPRESENTATIVE AND USER FACILITY SHOULD HAVE BEEN MARKED. DATE RECEIVED BY MANUFACTURER (MO/DAY/YR), CORRECTED DATA: FOLLOW-UP REPORT #1 INADVERTENTLY DID NOT INCLUDE THE DATE THE INFORMATION WAS RECEIVED BY THE MANUFACTURER. THE DATE LISTED SHOULD HAVE BEEN (B)(4) 2013.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3119360 | IT WAS INITIALLY REPORTED THAT THE PHYSICIAN HANDHELD WAS NOT WORKING PROPERLY. IT WILL FUNCTION WHEN IT IS PLUGGED IN BUT AS SOON AS IT IS UNPLUGGED THE SCREEN GOES BLACK AND IT IS UNABLE TO BE DETERMINED WHO MUCH BATTERY POWER THERE IS. THE HANDHELD IS STORED IN NORMAL OFFICE CONDITIONS. THE HANDHELD IS NOT LEFT PLUGGED IN. A HARD RESET WAS ATTEMPTED BUT AS SOON AS THE HANDHELD WAS UNPLUGGED IT DIES. IT HAD TO BE PLUGGED IN FOR A FEW MINUTES BEFORE IT WOULD TURN BACK ON. GOOD FAITH ATTEMPTS FOR PRODUCT RETURN HAVE BEEN UNSUCCESSFUL TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR REVIEW. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT HANDHELD WAS UNABLE TO CHARGE THE MAIN BATTERY. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH BROKEN SOLDER CONNECTIONS ON THE HANDHELD MAIN BOARD. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
3311465 | ANALYSIS OF THE HANDHELD AND SOFTWARE WERE COMPLETED ON (B)(4) 2013. NO SOFTWARE ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE SCREEN ALIGNMENT UTILITY. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WILL NOT GO PAST THE FIRST SCREEN WHERE IT SAYS TO TAP THE SCREEN. A HARD RESET WAS PERFORMED WHICH DID NOT RESOLVE THE ISSUE. THE LOCK BUTTON WAS NOT ENGAGED AND IT WAS REPORTED THAT THE HANDHELD IS CHARGED ALL OF THE TIME; HOWEVER, THE GREEN LIGHT IS NOT SHOWING AT THE TOP OF THE HANDHELD. IT WAS REPORTED THAT THE SERIAL CORD WAS NOT LOOSE. A NEW HANDHELD WAS PROVIDED TO THE PHYSICIAN'S OFFICE. THE HANDHELD AND FLASHCARD WERE RETURNED TO DEVICE MANUFACTURER FOR ANALYSIS ON (B)(4) 2013. THE RETURNED PRODUCT FORM INDICATED THAT THE HANDHELD WAS RETURNED DUE TO "NOT WORKING - FROZEN". ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3012769 | ON (B)(4) 2013 IT WAS REPORTED THAT THE PHYSICIAN'S THREE PROGRAMMING SYSTEMS ARE NOT HOLDING A CHARGE (OTHER TWO PROGRAMMING SYSTEMS NOT HOLDING A CHARGE ARE CAPTURED ON MFR. REPORT # 1644487-2013-00744 AND 1644487-2013-00745). IT WAS UNKNOWN IF THE SAME POWER CHARGER WAS BEING USED TO CHARGE ALL THREE HANDHELDS BUT IT WAS BELIEVED THAT THEY USED THE SAME OUTLET TO CHARGE THEM. THE PHYSICIAN LATER REPORTED THAT ALL THREE HANDHELDS ARE WORKING NOW. IT WAS STATED THAT THEY HAVE ONE POWER CORD THAT WAS GOING TO BE RETURNED TO THE MANUFACTURER FOR PRODUCT ANALYSIS. THE POWER CORD WAS RECEIVED FOR PRODUCT ANALYSIS ON (B)(4) 2013. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT YET BEEN COMPLETED. GOOD FAITH ATTEMPTS FOR FURTHER INFORMATION FROM THE PHYSICIAN'S OFFICE HAVE BEEN UNSUCCESSFUL.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2013 WHEN THE PHYSICIAN REPORTED THAT THE DATE IT WAS FIRST NOTICED THAT THE PROGRAMMING HANDHELDS WERE NOT HOLDING A CHARGE WAS (B)(6) 2013 (OTHER TWO PROGRAMMING SYSTEMS NOT HOLDING A CHARGE ARE CAPTURED ON MFR. REPORT # 1644487-2013-00744 AND 1644487-2013-00745). THE SAME WALL OUTLET WAS USED TO CHARGE ALL THREE HANDHELDS. THE SAME POWER CHARGER WAS NOT USED TO CHARGE ALL THREE HANDHELDS AS THE PHYSICIAN¿S OFFICE HAD TWO CHARGING UNITS. THE HANDHELDS WERE REPORTED TO NOW BE ABLE TO HOLD A CHARGE SINCE A NEW POWER CHARGER WAS SENT TO THEM. PRODUCT ANALYSIS ON THE SERIAL CABLE WAS COMPLETED ON (B)(6) 2013. AN ANALYSIS WAS PERFORMED ON THE RETURNED AC ADAPTER AND THE NO ANOMALIES ASSOCIATED WITH THE AC ADAPTER WERE NOTED DURING TESTING USING A KNOWN GOOD HANDHELD AND SERIAL CABLE. THE AC ADAPTER PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2925478 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2013 WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE RETURNED HANDHELD AND FLASHCARD. AN ANALYSIS WAS PERFORMED ON THE HANDHELD AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(6) 2013, IT WAS REPORTED THAT THE PHYSICIAN'S OFFICE WAS HAVING ISSUES WITH THEIR PROGRAMMING SYSTEM COMMUNICATING WITH VARIOUS PATIENTS, WHICH STARTED A FEW DAYS PRIOR. IT WAS NOTED THAT THE PROGRAMMING SYSTEM WAS NOT PLUGGED INTO THE WALL DURING INTERROGATIONS. THE MANUFACTURER'S CONSULTANT WAS USED HER GOOD WAND WITH THE PHYSICIAN'S HANDHELD AND WAS UNABLE TO COMMUNICATE WITH THE DEMO GENERATOR. THE CONSULTANT'S PROGRAMMING SYSTEM WAS SAID TO BE OPERATING PROPERLY. THE PHYSICIAN'S HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2013 FOR PRODUCT ANALYSIS. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT YET BEEN COMPLETED. . | 0
3006811 | IT WAS REPORTED THAT THE HANDHELD WILL NOT "USE" THE VERSION 8.1 FLASHCARD. IT WAS LATER REPORTED THAT A MANUFACTURER EMPLOYEE VISITED THE SITE TO TROUBLESHOOT THE EVENT. DURING TROUBLESHOOTING THE MANUFACTURER EMPLOYEE NOTED THAT WHEN TRYING TO UPGRADE THE HANDHELD FROM THE VERSION 7.1 FLASHCARD TO THE VERSION 8.1 FLASHCARD THE SOFTWARE DID NOT LOAD ON THE HANDHELD. THE EMPLOYEE REINSERTED THE VERSION 7.1 FLASHCARD INTO THE HANDHELD AND PERFORMED A HARD RESET ON THE HANDHELD SUCCESSFULLY. THE HANDHELD DISPLAYED THE VERSION 7.1 SOFTWARE. THE MANUFACTURER EMPLOYEE UPDATED THE DATE AND TIME ON THE HANDHELD AND THEN REPLACED THE VERSION 7.1 FLASHCARD WITH THE VERSION 8.1 FLASHCARD; HOWEVER, THE INSTALLATION PROCESS NEVER BEGAN. THE EMPLOYEE ONCE AGAIN INSERTED THE VERSION 7.1 FLASHCARD, WHICH AGAIN LOADED THE VERSION 7.1 SOFTWARE ON THE HANDHELD. ANOTHER VERSION 8.1 FLASHCARD WITH THE SAME LOT NUMBER WAS INSERTED INTO THE HANDHELD, WHICH RESULTED IN A MESSAGE "UNRECOGNIZED CARD IN SOCKET 1". A NEW HANDHELD WAS SHIPPED TO THE PHYSICIAN'S OFFICE. THE HANDHELD AND BOTH VERSION 8.1 FLASHCARDS ARE EXPECTED TO RETURN TO MANUFACTURER FOR PRODUCT ANALYSIS; HOWEVER, HAVE NOT BEEN RECEIVED TO DATE.><ATTEMPTS TO HAVE THE DEVICE RETURNED TO MANUFACTURER FOR ANALYSIS HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
2959623 | ADDITIONAL INFORMATION WAS RECEIVED THAT THE WAND, HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. NO VISUAL OR MECHANICAL ANOMALY WAS IDENTIFIED WITH THE WAND. CONTINUITY TESTING OF THE SERIAL DATA CABLE AND THE BATTERY CABLE PASSED. THE PROGRAMMING WAND PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS INITIALLY REPORTED THAT THE PHYSICIAN'S HANDHELD IS NOT HOLDING A CHARGE. THIS HAS BEEN OCCURRING FOR A NUMBER OF YEARS AND HE IS ONLY ABLE TO USE THE HANDHELD IF IS IT PLUGGING INTO AN OUTLET. THE PHYSICIAN HAS TRIED TO PLUG THE HANDHELD INTO DIFFERENT OUTLETS WHICH HAS NOT RESOLVED THE SITUATION. GOOD FAITH ATTEMPTS FOR PRODUCT RETURN HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
3262740 | AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE HANDHELD COULD NOT COMPLETE THE SCREEN ALIGNMENT UTILITY DURING THE INITIAL SETUP PROCESS. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH DEBRIS THAT WAS TRAPPED UNDER THE CASE. ONCE THE DISPLAY WAS CLEANED, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. REVIEW OF THE DEVICE MANUFACTURING RECORDS INDICATED THAT THE DEVICE PASSED ALL FINAL TESTS PRIOR TO DISTRIBUTION. NO OTHER INFORMATION HAS BEEN PROVIDED.><IT WAS REPORTED THAT THE VNS TREATING PHYSICIAN¿S VNS HANDHELD WAS NOT WORKING. THE COMPANY REPRESENTATIVE PERFORMED TROUBLESHOOTING, AND ONCE IT RESETS, IT SAYS TO TAP THE SCREEN BUT WILL NOT RESPOND. AFTER TRYING TO TAP THE SCREEN AS PROMPTED, THE SCREEN DID NOT RESPOND. THE REPRESENTATIVE TRIED PERFORMING A HARD RESET MULTIPLE TIMES BUT WAS UNABLE TO GET PAST THAT SCREEN. A REPLACEMENT HANDHELD DEVICE WAS PROVIDED TO THE PHYSICIAN AND FAULTY PROGRAMMING SYSTEM WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE. . | 0
2818492 | ON (B)(6) 2012 IT WAS REPORTED THAT THE PLASTIC AROUND THE PLUG OF THE CONSULTANT'S HANDHELD SERIAL CABLE BROKE A COUPLE OF WEEKS AGO, EXACT DATE UNKNOWN BUT LIKELY (B)(6) 2012. THE SERIAL CABLE WAS REPORTED TO BE WORKING FINE BUT IT HAD TO BE HELD IN A CERTAIN WAY FOR IT TO WORK. IT WAS REPORTED THAT THE SERIAL CABLE WOULD BE RETURNED FOR PRODUCT ANALYSIS HOWEVER IT HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2012 WHEN THE SERIAL CABLE WAS RETURNED FOR PRODUCT ANALYSIS. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT YET BEEN COMPLETED.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2012 WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE SERIAL CABLE. PRODUCT ANALYSIS CONFIRMED MECHANICAL DAMAGE TO CONNECTOR HOUSING OF DATA CABLE, RESULTING IN DIFFICULTY WITH INSERTION INTO A HANDHELD; CAUSE UNKNOWN. A VISUAL ANALYSIS OF THE CABLE WAS ABLE TO VERIFY THAT THE HOOD WAS DAMAGED. BECAUSE THE HOOD WAS DAMAGED, THE CONNECTOR IN THE HOOD ASSEMBLY WAS LOOSE AND DIFFICULT TO FULLY INSERT INTO THE HANDHELD. NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
2824869 | THE EVENT OF THE COMPUTER SCREEN BEING UNABLE TO CALIBRATE THE SCREEN BRIGHTNESS IS AN INTENDED OUTCOME OF THE UPGRADE TO 8.1 SOFTWARE. THE BUTTON ON THE SIDE OF THE COMPUTER WILL STILL CONTROL THE BRIGHTNESS LEVEL, WHICH IS ONLY ADJUSTABLE FROM VERY BRIGHT TO DIM ONLY. NO MALFUNCTION OCCURRED, AND THE EVENT WAS REPORTED AS A MALFUNCTION IN ERROR. THE COMPUTER WILL NOT BE RETURNED.><REPORTER INDICATED A VNS (B)(4) COMPUTER WITH NEWLY-INSTALLED (B)(4) WAS NOW UNABLE TO HAVE THE SCREEN CALIBRATED AND THE SCREEN BRIGHTNESS CAN ONLY BE ADJUSTED WITH THE POWER BUTTON, NOT THE SCREEN PROMPTS. PERFORMING A HARD RESET DID NOT RESOLVE THE ISSUES. ATTEMPTS FOR FURTHER INFORMATION AND POSSIBLE RETURN OF THE COMPUTER AND FLASHCARD ARE IN PROGRESS. THE EVENT OF THE COMPUTER SCREEN BEING UNABLE TO BE CALIBRATED TO ADJUST BRIGHTNESS IS NOT A MALFUNCTION, AND IS AN INTENDED OUTCOME OF THE MANUFACTURER 8.1 SOFTWARE UPGRADE. THE EVENT WAS REPORTED AS A MALFUNCTION IN ERROR.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
2833843 | THE COMPANY REPRESENTATIVE FOLLOWED UP WITH THE PHYSICIAN. THE BLANK SCREEN REPORTEDLY ONLY OCCURRED ON THE DIAGNOSTIC SCREEN. THE COMPANY REPRESENTATIVE EXPLAINED TO THE PHYSICIAN REGARDING WAND MOVEMENT AND THE RISK FOR INTERRUPTED COMMUNICATION, AS DISCUSSED IN MANUFACTURER LABELING. THE PHYSICIAN DID NOT ANSWER WHAT THE APPROXIMATE BATTERY POWER REMAINING ON THE HANDHELD COMPUTER SCREEN WAS WHEN THE SCREEN WENT BLANK, BUT THE PHYSICIAN REPORTED THAT HE KEPT THE HANDHELD COMPUTER PLUGGED INTO THE POWER OUTLET UNTIL USE. THE PHYSICIAN DID NOT INDICATE IF A MESSAGE WAS OBSERVED, BUT HE REPORTED THAT AFTER THE SCREEN WENT "BLANK," HE WAS ABLE TO "START OVER" AND PROCEED. THE PHYSICIAN REPORTED ON (B)(6) 2012 THAT THE HANDHELD COMPUTER WAS STILL NOT WORKING PROPERLY AND REQUESTED A REPLACEMENT PROGRAMMING SYSTEM WHICH WAS SUBSEQUENTLY PROVIDED.><A VNS TREATING PHYSICIAN REPORTED ON (B)(6) 2012 THAT THE PAST WEEK HE REPROGRAMMED FOUR OF HIS VNS PATIENT VNS SETTINGS, BUT WHEN HE TRIED TO PERFORM DIAGNOSTICS, IT SHOWED A "BLANK SCREEN." IT IS UNCLEAR WHAT WAS MEANT BY BLANK SCREEN WITH THE INFORMATION PROVIDED. ADDITIONALLY, THE PHYSICIAN REPORTED THAT WITH ONE OF THE PATIENTS IT RESULTED IN "RESETTING" THE GENERATOR AND IT "TURNED OFF." THE GENERATOR RESET AND BEING TURNED OFF IS REPORTED IN MANUFACTURER REPORT NUMBER: 1644487-2012-03021. A REPLACEMENT PROGRAMMING SYSTEM WAS PROVIDED TO THE PHYSICIAN. HOWEVER, THE PROGRAMMING SYSTEM USED THE WEEK OF (B)(6) 2012 HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE. ATTEMPTS TO TROUBLESHOOT THE SYSTEM BY THE COMPANY REPRESENTATIVE AND PRODUCT RETURN ARE IN PROGRESS BUT HAVE BEEN UNSUCCESSFUL TO DATE.><THE PROGRAMMING SYSTEM WAS RECEIVED FOR PRODUCT ANALYSIS BY THE MANUFACTURER ON (B)(4)2012. THE PROGRAMMING WAND PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THERE WERE NO VISUAL OR MECHANICAL ANOMALY WAS IDENTIFIED. CONTINUITY TESTING OF THE SERIAL DATA CABLE AND THE BATTERY CABLE PASSED. IN ADDITION, NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. .><. | 0
3354452 | DURING THE PERIODIC PROGRAMMING HISTORY REVIEW, IT WAS FOUND THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED WHICH CAUSED AN UNINTENDED CHANGE IN THE PATIENT¿S VNS SETTINGS. THE SETTINGS WERE NOT CORRECTED AT THE TIME OF THE EVENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS OCCURRENCE. ANALYSIS OF PROGRAMMING HISTORY. | 0
3354964 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT DURING INTERROGATION ON OFFICE VISIT ON (B)(6) 2011, THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE SAME OFFICE VISIT. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2011. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE DEVICE OFF TIME WAS NOT CORRECTED. THE DEVICE WAS INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2011 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; HOWEVER, THE PHYSICIAN DID NOT CORRECT THE OFF TIME BACK TO AN EFFICACIOUS SETTING. THE SETTING WAS CORRECTED ON (B)(6) 2011. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2957805 | IT WAS REPORTED THAT THE VNS PROGRAMMING SYSTEM COULD NOT INTERROGATE THE PATIENT'S VNS GENERATOR SUCCESSFULLY AFTER IT HAS INTERROGATED IT THE FIRST TIME DURING THE OFFICE VISIT ON (B)(6) 2013. AFTER THE VISIT TROUBLE SHOOTING WAS PERFORMED BY THE USER OF THE VNS PROGRAMMING SYSTEM, AND IT WAS FOUND THAT THE WAND BATTERY WAS MOST LIKELY DEAD. THE VNS PATIENT RETURNED TO THE OFFICE ON (B)(6) 2013 AND AFTER THE USER CHANGED THE BATTERY THE PROGRAMMING SYSTEM WORKED WITH NO ISSUES, BUT THE PATIENT WAS FOUND PROGRAMMED AT OUTPUT CURRENT= 0 MA/ FREQUENCY= 20 HZ/ PULSE WIDTH= 500 'SEC/ON TIME= 30 SEC/OFF TIME= 60 MIN / MAGNET OUTPUT CURRENT= 0 MA/ ON TIME= 60SEC/ PULSE WIDTH= 500'SEC, WHICH IS AN INDICATION OF FAULTED SYSTEM DIAGNOSTICS. THE PATIENT'S SETTINGS WERE THEN CORRECTED. THE FAULTED DIAGNOSTICS WAS MOST LIKELY DUE TO THE FAILED COMMUNICATION ON (B)(6) 2013. ANALYSIS OF PROGRAMMING HISTORY. | 0
2775201 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS NOTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED ON (B)(6) 2006 WHICH CHANGED THE SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE CORRECTED THEN NEXT DAY ON (B)(6) 2006. ANALYSIS OF PROGRAMMING HISTORY. | 0
2903912 | UPON REVIEW OF THE IN-HOUSE PROGRAMMING HISTORY DATABASE, IT WAS OBSERVED THAT A GENERATOR DIAGNOSTICS TEST WAS PERFORMED ON (B)(6) 2009 WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE NOT CORRECTED UNTIL THE PATIENT'S NEXT VISIT ON (B)(6) 2010. NO PATIENT ADVERSE EVENTS WERE REPORTED. MANUFACTURER LABELING INDICATES THAT GENERATOR DIAGNOSTIC TESTS SHOULD ONLY BE PERFORMED IN THE OPERATING ROOM TO TEST THE DEVICE AND SHOULD NOT BE PERFORMED AT FOLLOW-UP VISITS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3221740 | DURING A REVIEW OF PROGRAMMING HISTORY ON (B)(6) 2013 IT WAS NOTED THAT ON ONE OCCASION A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2009. THE CHANGE IN SETTINGS WERE FOUND AND CORRECTED ON (B)(6) 2010. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2950859 | WHILE REVIEWING PROGRAMMING HISTORY FOR THE PATIENT IT WAS DISCOVERED THAT THE PATIENT LEFT AN APPOINTMENT AT UNINTENTIONAL SETTING DUE TO AN INCOMPLETE DIAGNOSTICS. THE PATIENT CAME INTO THE APPOINTMENT AT CORRECT SETTINGS AND INCOMPLETE DIAGNOSTICS OCCURRED DURING THE APPOINTMENT. AS A RESULT OF THE INCOMPLETE DIAGNOSTICS THE PATIENT SETTINGS WERE CHANGED. THE SETTING CHANGED WAS NOTICED DURING INTERROGATION AND PARTIALLY CORRECT BUT THE OFF TIME WAS LEFT AT 60 MINUTES. THE SETTINGS WERE NOT CORRECT UNTIL A LATER APPOINTMENT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2743817 | DURING REVIEW OF THE IN-HOUSE DATABASE FOR THE PATIENT'S PROGRAMMING/DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON THE DATE OF IMPLANT WHICH CHANGED THE PATIENT'S SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING ALTHOUGH IT WAS ATTEMPTED TO REPROGRAM THE SETTINGS. HOWEVER, A FINAL INTERROGATION WAS NOT PERFORMED TO CONFIRM THE SETTINGS. THE SETTINGS WERE CORRECTED ON THE FOLLOWING VISIT ON (B)(6) 2006. ANALYSIS OF PROGRAMMING HISTORY. | 0
3355160 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2011 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2011. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2011. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2011 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2011. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
3071574 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2007 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2007. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2007. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2007 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2007. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
3069182 | DURING REVIEW OF A VNS PATIENT'S PROGRAMMING HISTORY, IT WAS NOTED ON (B)(6) 2009 THE DEVICE WAS INTERROGATED AND THE SETTINGS WERE 1MA/30HZ/500PULSEWIDTH/30SEC ON/5MIN OFF, SYSTEMS DIAGNOSTICS WERE PERFORMED WHICH RESULTED IN "FAULT". A FINAL INTERROGATION WAS NOT PERFORMED TO VERIFY SETTINGS. ON THE NEXT RECORDED ON (B)(6) 2009, THE FIRST INTERROGATION OF THE DEVICE REVEALED THE SETTINGS WERE 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF; THE SETTINGS WERE CORRECTED THIS DAY. ANALYSIS OF PROGRAMMING HISTORY. | 0
2881868 | DURING REVIEW OF INTERNAL PROGRAMMING HISTORY ON (B)(6) 2012, A COMPUTER SOFTWARE ERROR WAS OBSERVED ON ADJUSTED DATE OF (B)(6)2010). WHILE PERFORMING THE REVIEW IT WAS NOTED ON (B)(6) 2010 (THE ADJUSTED DATE - ORIGINAL DATE IS (B)(6) 2010) THE DEVICE WAS INTERROGATED AND THE SETTINGS WERE 2.25/30/250/30/3, A SYSTEM DIAGNOSTIC WAS PERFORMED AND RESULTED IN "FAULT" THE DEVICE WAS INTERROGATED AND THE SETTINGS WERE CHANGED TO 0/30/250/30/3 AND WERE NOT PROGRAMMED PROGRAMMED TO INTENDED SETTINGS PRIOR TO THE PATIENT LEAVING THE VISIT. THE PATIENT RETURNED TO CLINIC ON (B)(6) 2010 AND THEIR SETTINGS WERE CORRECTED AT THAT TIME. ANALYSIS OF PROGRAMMING HISTORY. | 0
3237800 | REVIEW OF THE MANUFACTURER'S PROGRAMMING HISTORY DATABASE REVEALED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2011. A FINAL INTERROGATION WAS PERFORMED WHICH SHOWED THE DEVICE SETTINGS WERE CHANGED TO UNINTENDED PARAMETERS, BUT THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE CLINIC. THE OFF-TIME AND MAGNET OUTPUT CURRENT WERE LATER CORRECTED ON (B)(6) 2011. HOWEVER, THE PULSE WIDTH WAS NOT CORRECTED UNTIL (B)(6) 2012. NO PATIENT ADVERSE EVENTS WERE REPORTED. MANUFACTURER LABELING INDICATES TO PERFORM A FINAL INTERROGATION PRIOR TO PATIENTS LEAVING THE CLINIC TO VERIFY THE DEVICES ARE AT INTENDED PARAMETERS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3353216 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS. THE SETTINGS WERE NOT CORRECTED AT THE TIME OF THE EVENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3023379 | A NEUROLOGIST REPORTED TO OUR CONSULTANT THAT THEIR HANDHELD COMPUTER IS NO LONGER ABLE TO HOLD A CHARGE AND THAT HE HAS ALREADY TRIED TO PLUG IT INTO SEVERAL DIFFERENT WALL OUTLETS WITH THE SAME RESULT. HE STATED THAT THE AC POWER CORD WAS CHECKED VISUALLY AND NOTHING APPEARED TO BE WRONG WITH IT. HE WAS NOT SURE WHEN THIS ISSUE FIRST BEGAN BUT EXPLAINED THAT IT HAS BEEN GOING ON FOR A WHILE. THE HANDHELD WAS RETURNED FOR ANALYSIS. ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE SYNC CABLE CONNECTOR ON THE BOTTOM OF THE HANDHELD WAS DAMAGED. BECAUSE OF THE DAMAGE, THE HANDHELD WAS UNABLE TO RECEIVE POWER FROM THE AC ADAPTER. THE CAUSE FOR THE DAMAGE TO THE CONNECTOR IS MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE AS IT APPEARS THE CONNECTOR DETENTS ARE NOT BEING COMPRESSED WHEN REMOVING OR MANIPULATING THE SERIAL DATA CABLE, THUS PLACING AN EXTENSIVE AMOUNT OF FORCE ON THE PCB AND ATTACHED CABLE RECEPTACLE. NO OTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS . | 0
2900369 | REPORTER INDICATED A VNS DELL X5 COMPUTER BECAME FROZEN ON THE "INTERROGATION SUCCESSFUL" SCREEN WHILE USING A VNS DEMONSTRATOR GENERATOR. A HARD RESET RESOLVED THE ISSUE, AND THE COMPUTER WAS ABLE TO BE USED SUCCESSFULLY WITHOUT FURTHER ISSUES. THE COMPUTER WILL NOT BE RETURNED AT THIS TIME. . | 0
3062684 | DURING REVIEW OF THE PATIENT'S PROGRAMMING HISTORY ON (B)(6) 2013 IT WAS DISCOVERED THAT ON (B)(6) 2011 A SYSTEM DIAGNOSTICS TEST WAS PERFORMED WHICH CHANGED THE PATIENT'S SETTINGS. SOME OF THE SETTINGS WERE CORRECTED AFTERWARDS BUT THEN A SECOND FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED WHICH CHANGED THE PATIENT'S SETTINGS AGAIN. THIS TIME THERE WAS NO FINAL INTERROGATION. IT WAS NOT UNTIL THE PATIENT'S NEXT VISIT ON (B)(6) 2011 THAT THE SETTINGS WERE PARTIALLY CORRECTED. ON (B)(6) 2012 THE PATIENT'S SETTINGS WERE COMPLETELY CORRECTED. NO PATIENT ADVERSE EVENTS WERE REPORTED TO HAVE OCCURRED DUE TO THESE SETTINGS CHANGES. ANALYSIS OF PROGRAMMING HISTORY. | 0
2866684 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2010 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2009. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2009 WAS WITHIN NORMAL LIMITS. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2009 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2010. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
3351164 | DURING REVIEW OF PROGRAMMING HISTORY ON (B)(6) 2013, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS. THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE APPOINTMENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3351124 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS. THE SETTINGS WERE NOT CORRECTED AT THE TIME OF THE EVENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3069159 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS NOTED THAT THE PATIENT'S VNS SETTINGS WERE CHANGED ON THE VISIT OF (B)(6) 2009 TO 0MA/30HZ/500PULSEWIDTH/30SEC ON/60MIN OFF. SETTINGS WERE CHANGED TO 1.5MA/30HZ/500PW/30SEC ON/60MIN OFF BEFORE PATIENT LEFT THE OFFICE. THE OFF TIME WAS NOT CORRECTED UNTIL (B)(6) 2010. ANALYSIS OF PROGRAMMING HISTORY. | 0
3306100 | A REVIEW OF THE PATIENT'S PROGRAMMING HISTORY INDICATES THAT UPON INITIAL INTERROGATION ON (B)(6) 2005, THE PATIENT'S SETTINGS WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST. ALTHOUGH THERE IS NO HISTORY OF A FAULTED TEST BEING PERFORMED PRIOR TO THIS DATE, IT IS LIKELY THE FAULT OCCURRED ON THE DAY OF IMPLANT ((B)(6) 2013). AS IT IS RECOMMENDED THAT PATIENTS BE PROGRAMMED OFF (TO 0MA) FOLLOWING SURGERY, AND THE PATIENT'S SETTINGS WERE DECREASED ON (B)(6) 2013, IT IS LIKELY THESE CHANGED SETTINGS FOUND ON (B)(6) 2013 WERE UNINTENDED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2787243 | IT WAS REPORTED THAT THE PHYSICIAN'S VNS PROGRAMMING HANDHELD COMPUTER WAS EXPERIENCING SCREEN FREEZES, TYPICALLY AT THE INTERROGATION SUCCESSFUL SCREEN. A HARD RESET USUALLY RESOLVES THE ISSUE. THE PHYSICIAN INDICATED THAT HE WAS FRUSTRATED HAVING TO PERFORM TROUBLESHOOTING DURING SURGERY AND REQUESTED A REPLACEMENT PROGRAMMING SYSTEM. THE FAULTY HANDHELD COMPUTER PROGRAMMING SYSTEM WAS RETURNED AND UNDERWENT ANALYSIS. THE SCREEN FREEZING ISSUE WAS CONFIRMED AS REPORTED. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE SCREEN FREEZING ISSUE. . | 0
2792826 | DURING REVIEW OF THE PATIENT'S PROGRAMMING HISTORY AVAILABLE IN THE IN-HOUSE DATABASE, IT WAS OBSERVED THAT THE SETTINGS ON THE DATE OF IMPLANT WERE THE SHIPPING SETTINGS AND THEN UPON INITIAL INTERROGATION ON (B)(6) 2005, THE SETTINGS WERE AT UNINTENDED PARAMETERS. THERE WERE SEVERAL SYSTEM DIAGNOSTIC TESTS PERFORMED AFTER THE FINAL INTERROGATION ON THE DATE OF IMPLANT. A SYSTEM DIAGNOSTIC TEST DID FAULT WHICH CHANGED THE SETTINGS TO UNINTENDED PARAMETERS. THE TWO SUBSEQUENT SYSTEM DIAGNOSTIC TESTS ON THIS DATE RESULTED IN OKAY RESULTS WITHIN NORMAL LIMITS. AS A RESULT OF THE FAULTED SYSTEM DIAGNOSTIC TEST ON THE DATE OF IMPLANT AND A FINAL INTERROGATION NOT BEING PERFORMED, THE PATIENT'S DEVICE WAS SET AT UNINTENDED SETTINGS. NO PATIENT ADVERSE EVENTS WERE REPORTED. THE PATIENT'S SETTINGS WERE REPROGRAMMED ON (B)(6) 2005. ANALYSIS OF PROGRAMMING HISTORY. | 0
3334721 | DURING REVIEW OF PROGRAMMING HISTORY IT WAS SEEN THAT THE PATIENT CAME INTO AN APPOINTMENT AND UNEXPECTED SETTINGS. REVIEW OF PROGRAMMING HISTORY SHOWS THAT THERE WAS A GENERATOR DIAGNOSTICS THAT WAS PERFORMED AT THE PRIOR APPOINTMENT. AT THE COMPLETION OF GENERATOR DIAGNOSTICS THE PATIENT IS PROGRAMMED TO SETTINGS OF OUTPUT CURRENT 0 MA, FREQUENCY 30 HZ, PULSE WIDTH 500 USEC, ON TIME 30 SEC AND OFF TIME 5 MINUTES. THERE WAS NOT FINAL INTERROGATION AND THE PATIENT SETTINGS WERE NOT CORRECT PRIOR TO THE PATIENT LEAVING THE APPOINTMENT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2800581 | DURING A REVIEW OF PROGRAMMING HISTORY FOR THIS VNS PATIENT ON (B)(6) 2012, IT WAS NOTED THAT AN INTERRUPTED DIAGNOSTIC TEST ON (B)(6) 2007 ALTERED THE PATIENT'S SETTINGS. THE PATIENT LEFT THE APPOINTMENT AT UNINTENDED SETTINGS. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS ISSUE. REVIEW OF PROGRAMMING HISTORY. | 0
2938068 | DURING A REVIEW OF PROGRAMMING HISTORY ON (B)(6) 2013 FOR FILE (B)(4), IT WAS NOTED THAT ON TWO OCCASIONS, THIS PATIENT'S DEVICE WAS INTERROGATED AT SETTINGS INDICATIVE OF A FAULTED SYSTEM DIAGNOSTIC TEST: (B)(6) 2010 (EVENT OCCURRED ON (B)(6) 2010) AND (B)(6) 2010 (EVENT OCCURRED ON (B)(6) 2010).THE CHANGE IN SETTINGS WAS NOT CORRECTED AT THE SAME TIME. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. REVIEW OF PROGRAMMING HISTORY. | 0
3135295 | A PROGRAMMING HISTORY REVIEW FOUND THAT A FAULTED SYSTEM DIAGNOSTIC TEST WAS PERFORMED ON (B)(6) 2010 WHICH CHANGED THE PATIENT'S SETTINGS FROM 2 MA OUTPUT CURRENT, 20 HZ FREQUENCY, 500 MICROSECONDS PULSE WIDTH, 30 SECONDS ON TIME, 3 MINUTES OFF TIME, 2.25 MA MAGNET OUTPUT CURRENT, 500 MICROSECONDS MAGNET PULSE WIDTH, 60 SECONDS MAGNET ON TIME TO 0 MA OUTPUT CURRENT, 20 HZ FREQUENCY, 500 MICROSECONDS PULSE WIDTH, 30 SECONDS ON TIME, 3 MINUTES OFF TIME, 0 MA MAGNET OUTPUT CURRENT, 500 MICROSECONDS MAGNET PULSE WIDTH, 60 SECONDS MAGNET ON TIME. THESE SETTINGS WERE NOT CHANGED UNTIL THE NEXT VISIT ON (B)(6) 2011. ANALYSIS OF PROGRAMMING HISTORY. | 0
2991507 | A SITE REPORTED THAT THEIR DELL X50 HANDHELD COMPUTER WAS FROZEN ON THE INTERROGATION SCREEN AND THEY COULD NOT INTERROGATE AND RUN DIAGNOSTICS. A HARD RESET WAS PERFORMED AND THE HANDHELD IS WORKING NOW. THE PATIENT WILL BE SCHEDULED TO COME BACK. . | 0
2818355 | DURING REVIEW OF THE VNS PATIENT'S PROGRAMMING HISTORY ON (B)(6) 2012, IT WAS DISCOVERED THAT A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED ON (B)(6) 2008 THAT CHANGED THE PATIENT'S SETTINGS. ALL THE PARAMETERS WERE CORRECTED PRIOR TO THE PATIENT LEAVING THE CLINIC VISIT EXCEPT FOR THE FREQUENCY, OFF TIME, MAGNET OUTPUT, AND MAGNET ON TIME. THE SETTINGS WERE NOT CORRECTED UNTIL THE NEXT VISIT ON (B)(6) 2008. NO ADVERSE EVENTS HAVE BEEN REPORTED DUE TO THIS SETTING CHANGE. ANALYSIS OF PROGRAMMING HISTORY. | 0
3056175 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT ALTERED GENERATOR SETTINGS. THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE APPOINTMENT. NO PATIENT ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. REVIEW OF PROGRAMMING HISTORY. | 0
2775164 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS NOTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED ON (B)(6) 2009, WHICH CHANGED THE VNS SETTING TO UNINTENDED PARAMETERS. THE CHANGE WAS NOT NOTED UNTIL (B)(6) 2009, WHEN THE SETTINGS WERE CORRECTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
3285743 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2006 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2005. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2005 WAS WITHIN NORMAL LIMITS. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2005 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2006; HOWEVER, THE DEVICE'S DUTY CYCLE WAS NOT CORRECTED UNTIL (B)(6) 2006. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2929194 | DURING REVIEW OF THE VNS PROGRAMMING HISTORY DATABASE, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2006 WHICH INADVERTENTLY CHANGED THE SETTINGS TO UNINTENDED PARAMETERS. THE NORMAL AND MAGNET OUTPUT CURRENT WAS PROGRAMMED BACK TO 1.0MA, BUT THE OTHER SETTINGS WERE NOT. THE DEVICE WAS TURNED OFF THE NEXT VISIT ON (B)(6) 2006 AND THEN REPROGRAMMED TO THE PROGRAMMING ANOMALY SETTINGS. THE OFF TIME WAS NOT CORRECTED FROM 60 MINUTES UNTIL (B)(6) 2008 BY A DIFFERENT PROGRAMMER. NO PATIENT ADVERSE EVENTS WERE REPORTED. MANUFACTURER LABELING LISTS THE IMPORTANCE OF PERFORMING FINAL INTERROGATIONS PRIOR TO PATIENT'S LEAVING THE CLINIC TO IDENTIFY AND CORRECT SUCH PROGRAMMING ANOMALIES. ANALYSIS OF PROGRAMMING HISTORY. | 0
3070801 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT A FAULTED GENERATOR DIAGNOSTIC TEST ALTERED DEVICE SETTINGS. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. ANALYSIS OF PROGRAMMING HISTORY. | 0
3070735 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS IDENTIFIED THAT A FAULTED SYSTEMS TEST OCCURRED ON (B)(6) 2010 AT 09:28:45, THE PATIENT WAS THEN PROGRAMMED TO 2MA/30HZ/500PULSEWIDTH/30SEC ON/5MIN OFF AT 09:27:43 AND THEN ANOTHER SYSTEMS TEST FOLLOWED AT 09:30:41 WHICH WAS SUCCESSFUL, BUT NO FINAL INTERROGATION WAS DONE TO VERIFY SETTINGS. AT THE NEXT OFFICE VISIT ON (B)(6) 2010, THE INITIAL INTERROGATION AT 11:33:41 HAS SETTINGS OF 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF/1MA MAG/500PULSEWIDTH/30SEC ON, WHICH IS INDICATIVE OF A FAULTED SYSTEMS DIAGNOSTICS TEST. THE SETTINGS WERE ALL CORRECTED ON (B)(6) 2010. ANALYSIS OF PROGRAMMING HISTORY. | 0
3180705 | ON (B)(6) 2013 DURING REVIEW OF THE PATIENT¿S PROGRAMMING HISTORY IT WAS OBSERVED THAT ON DATE OF IMPLANT, (B)(6) 2010, A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED THAT CHANGED THE PATIENT¿S SETTINGS. NO FINAL INTERROGATION WAS PERFORMED AND ALTHOUGH THE PATIENT WAS AT 0MA AND THEREFORE THERAPY AT THAT TIME WASN¿T AFFECTED, THE OFF TIME HAD BEEN CHANGED TO 60MIN. ON THE PATIENT¿S NEXT VISIT THIS WAS NOT NOTICED AND THEREFORE THE SETTINGS WERE ONLY PARTIALLY CORRECTED; THE OFF TIME WAS NOT CORRECTED UNTIL THE PATIENT¿S FOLLOWING VISIT ON (B)(6) 2011. ANALYSIS OF PROGRAMMING HISTORY. | 0
2910483 | REPORTER INDICATED THAT A PATIENT'S VNS SETTINGS WERE NOTED TO HAVE "RESET" ON (B)(6) 2012 FOLLOWING A VNS DIAGNOSTICS TEST. IT IS UNCLEAR IF THE "RESET" OCCURRED DURING THE OFFICE VISIT OR AT AN EARLIER DATE. ATTEMPTS FOR FURTHER INFORMATION AND VNS PROGRAMMING HISTORY ARE IN PROGRESS.><ALL ATTEMPTS TO THE REPORTER FOR ADDITIONAL INFORMATION AND VNS PROGRAMMING HISTORY HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
3090787 | WHILE REVIEWING THE PATIENT¿S VNS GENERATOR PROGRAMMING HISTORY A SYSTEM FAULTED DIAGNOSTICS WAS NOTED TO HAVE LEFT THE PATENT WITH UNINTENDED SETTINGS. ON (B)(6) 2006 THE PATIENT LAST SETTINGS WERE OUTPUT CURRENT= 1.75 MA/ FREQUENCY= 30 HZ/ PULSE WIDTH= 250 ¿SEC/ON TIME= 21 SEC/OFF TIME= 1.1 MIN / MAGNET OUTPUT CURRENT= 2.5 MA/ PULSE WIDTH= 500 ¿SEC / ON TIME= 60 SEC. A SYSTEM DIAGNOSTICS WAS THEN PERFORMED WHICH FAULTED. NO FINAL INTERROGATION WAS THEN PERFORMED TO CHECK THE PATIENT'S SETTINGS. THE NEXT RECORDED VISIT ON (B)(6) 2013 THE PATIENT SETTINGS WERE OUTPUT CURRENT= 1 MA/ FREQUENCY= 20 HZ/ PULSE WIDTH= 500 ¿SEC/ON TIME= 30 SEC/OFF TIME= 60 MIN / MAGNET OUTPUT CURRENT= 1 MA/ PULSE WIDTH= 500 ¿SEC / ON TIME= 30 SEC. THIS EVENT IS CAPTURED IN MDR REPORT #: 1644487-2013-01230. ON (B)(6) 2007 THE PATIENT'S SETTINGS WAS STILL AT OUTPUT CURRENT= 1 MA/ FREQUENCY= 20 HZ/ PULSE WIDTH= 500 ¿SEC/ON TIME= 30 SEC/OFF TIME= 60 MIN / MAGNET OUTPUT CURRENT= 1 MA/ PULSE WIDTH= 500 ¿SEC / ON TIME= 30 SEC. A GENERATOR DIAGNOSTICS WAS THEN PERFORMED WHICH RESULTED THE PATIENT¿S SETTINGS TO CHANGE TO OUTPUT CURRENT= 0 MA/ FREQUENCY= 30 HZ/ PULSE WIDTH= 500 ¿SEC/ON TIME= 30 SEC/OFF TIME= 5 MIN / MAGNET OUTPUT CURRENT= 0 MA/ PULSE WIDTH= 500 ¿SEC / ON TIME= 60 SEC. ON (B)(6) 2007 THE PATIENT'S SETTINGS WAS THEN CORRECTED TO OUTPUT CURRENT= 0.25 MA/ FREQUENCY= 30 HZ/ PULSE WIDTH= 250 ¿SEC/ON TIME= 30 SEC/OFF TIME= 5 MIN / MAGNET OUTPUT CURRENT= 0.5 MA/ PULSE WIDTH= 500 ¿SEC / ON TIME= 60 SEC. THIS EVENT IS CAPTURED IN THIS MDR REPORT. ANALYSIS OF PROGRAMMING HISTORY. | 0
3124489 | WHILE REVIEWING THE PATIENT'S PROGRAMMING HISTORY, IT WAS FOUND THAT FINAL INTERROGATION WAS NOT PERFORMED AFTER A SYSTEM DIAGNOSTIC TEST AND FAULTED GENERATOR DIAGNOSTIC TEST ON (B)(6) 2006. UPON INITIAL INTERROGATION AT THE NEXT OFFICE VISIT OF (B)(6) 2007, THE PATIENT WAS FOUND TO BE PROGRAMMED TO DIFFERENT SETTINGS THAN THOSE THE PATIENT WAS PROGRAMMED TO ON (B)(6) 2006. THE PATIENT'S SETTINGS WERE NOT ADJUSTED UNTIL (B)(6) 2007. ANALYSIS OF PROGRAMMING HISTORY. | 0
3351903 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS. THE SETTINGS WERE NOT CORRECTED AT THE TIME OF THE EVENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. ANALYSIS OF PROGRAMMING HISTORY. | 0
2980928 | ON (B)(6) 2010, AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT ALTERED THE NORMAL MODE OUTPUT CURRENT. THE SETTINGS WERE NOT RETURNED TO EFFICACIOUS VALUES PRIOR TO THE PATIENT LEAVING THE OFFICE. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. REVIEW OF PROGRAMMING HISTORY. | 0
2961157 | DURING REVIEW OF THE MANUFACTURER'S PROGRAMMING HISTORY DATABASE, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2009 WHICH CHANGED THE PATIENT'S SETTING TO UNINTENDED PARAMETERS. A FINAL INTERROGATION WAS NOT PERFORMED, SO THE SETTINGS WERE NOT IDENTIFIED OR CORRECTED PRIOR TO THE PATIENT LEAVING THE CLINIC. UPON INITIAL INTERROGATION ON (B)(6) 2010, THE SETTINGS WERE AT UNINTENDED PARAMETERS, BUT THEY WERE CORRECTED AT THIS VISIT. MANUFACTURER LABELING INDICATES TO PERFORM FINAL INTERROGATIONS PRIOR TO PATIENTS LEAVING THE CLINIC TO IDENTIFY SUCH PROGRAMMING ANOMALIES. NO PATIENT ADVERSE EVENTS WERE REPORTED. ANALYSIS OF PROGRAMMING HISTORY. | 0
2886388 | DURING A REVIEW OF PROGRAMMING HISTORY ON (B)(4) 2012, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2010 THAT RESULTED IN THE PATIENT'S DEVICE BEING PROGRAMMED TO UNINTENDED SETTINGS. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS EVENT. REVIEW OF PROGRAMMING HISTORY. | 0
3293559 | WHILE REVIEWING THE PATIENT'S PROGRAMMING HISTORY, IT WAS NOTED THAT A FAULTED SYSTEM DIAGNOSTIC TEST WAS PERFORMED ON THE IMPLANT DATE, (B)(6) 2003. ALTHOUGH SOME OF THE PATIENT'S SETTINGS WERE REPROGRAMMED AND A FINAL INTERROGATION WAS PERFORMED, THE PATIENT'S MAGNET OUTPUT CURRENT WAS LEFT ON AT 1MA. THE PATIENT'S SETTINGS WERE FOUND TO BE CORRECTED UPON INITIAL INTERROGATION AT THE NEXT OFFICE VISIT OF (B)(6) 2003. ANALYSIS OF PROGRAMMING HISTORY. | 0
2753558 | DURING REVIEW OF THE VNS PROGRAMMING HISTORY AVAILABLE TO THE MANUFACTURER, IT WAS NOTED THAT INTERROGATION OF THE PATIENT'S SETTINGS ON (B)(6) 2009 REVEALED THE PARAMETERS WERE INDICATIVE THAT A SYSTEM DIAGNOSTICS TEST HAD FAULTED RESULTING IN AN UNINTENDED CHANGE IN SETTINGS. THE PATIENT HAD PREVIOUSLY BEEN SEEN ON (B)(6) 2008 DURING WHICH TIME A SYSTEM DIAGNOSTIC TEST WAS PERFORMED. A FINAL INTERROGATION WAS NOT PERFORMED TO CONFIRM THE PATIENT WAS AT INTENDED SETTINGS. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THE CHANGE IN SETTINGS. THE PHYSICIAN BEGAN TITRATING THE PARAMETERS BACK TO THERAPEUTIC LEVELS UPON DISCOVERY. ANALYSIS OF PROGRAMMING HISTORY. | 0
3004989 | DURING REVIEW OF THE PATIENT'S PROGRAMMING HISTORY, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2010 WHICH CHANGED THE SETTINGS TO UNINTENDED PARAMETERS. THE SETTINGS WERE NOT CORRECTED UNTIL (B)(6) 2010. A FINAL INTERROGATION WAS PERFORMED ON (B)(6) 2010 BUT THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE CLINIC. NO ADVERSE EVENTS WERE REPORTED AS A RESULT. ANALYSIS OF PROGRAMMING HISTORY. | 0
2922225 | DURING MANUFACTURER REVIEW OF A PATIENT'S VNS PROGRAMMING HISTORY, IT WAS IDENTIFIED THAT THE PATIENT LEFT THE OPERATING ROOM AT UNTENDED VNS SETTINGS OF 1MA/20HZ/500PULSEWIDTH/30SEC ON/60MIN OFF AFTER INITIAL VNS IMPLANT SURGERY ON (B)(6) 2004. A FINAL INTERROGATION WAS NOT PERFORMED TO VERIFY SETTINGS FOLLOWING SYSTEMS DIAGNOSTICS TESTING. THE SYSTEMS DIAGNOSTICS ARE NOT REFLECTED IN THE HISTORY AS FAULTED, BUT THIS IS SUSPECTED TO HAVE OCCURRED. THE SETTINGS WERE LATER CORRECTED ON (B)(6) 2004. ANALYSIS OF PROGRAMMING HISTORY. | 0
2949019 | A PROGRAMMING ANOMALY WAS OBSERVED DURING PROGRAMMING HISTORY REVIEW. THE EVENT OCCURED AFTER THE PATIENT'S APPOINTMENT ON (B)(6) 2011 AND PRIOR TO THE PATIENT'S APPOINTMENT ON (B)(6) 2011 AS NOTED IN THE PROGRAMMING HISTORY. A FAULTED TEST LIKELY OCCURRED AND ALTERED THE PATIENT'S SETTINGS AND THEY WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. ANALYSIS OF PROGRAMMING HISTORY. | 0
3077637 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AT FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2010 THAT CHANGED DEVICE SETTINGS. UPON INTERROGATION AT THE NEXT APPOINTMENT ON (B)(6) 2011, THE DEVICE WAS FOUND TO BE OFF. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS. REVIEW OF PROGRAMMING HISTORY. | 0
2836259 | THE PATIENT HAD GENERATOR REPLACEMENT ON (B)(6) 2012. PRODUCT ANALYSIS FOR THE EXPLANTED GENERATOR WAS ON (B)(6) 2012 BY THE MANUFACTURER, AND THE AS-RECEIVED SETTINGS WERE INDICATIVE OF A FAULTED SYSTEM DIAGNOSTIC TEST OCCURING SOMETIME PRIOR TO THE DATE OF SURGERY. REVIEW OF THE IN HOUSE PROGRAMMING DATABASE REVEALS THAT THESE WERE LIKELY UNINTENDED SETTINGS. SINCE IT IS UNKNOWN WHAT DATE THE PROGRAMMING ANOMALY OCCURRED, THE EVENT DATE IS LISTED AS THE DATE OF SURGERY AT THIS TIME. ATTEMPTS FOR A COPY OF THE PHYSICIAN'S PROGRAMMING/DIAGNOSTIC HISTORY FOR THE PATIENT HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
2874363 | DURING A REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT A MAGNET MODE DIAGNOSTIC TEST ON (B)(6) 2010 RESULTED IN A LOW OUTPUT CURRENT. THE LOW OUTPUT CURRENT WARNING ON THE MAGNET MODE TEST IS A RESULT OF THE USER NOT SWIPING THE MAGNET BEFORE PERFORMING THE MAGNET MODE DIAGNOSTICS TEST. THE LOW OUTPUT CURRENT RESULT IS LIKELY DUE TO THE GENERATOR COMPARING THE LAST MAGNET SWIPE OUTPUT TO THE NEWLY PROGRAMMED MAGNET MODE OUTPUT CURRENT WHICH HAD NOT BEEN DELIVERED YET. THE "MAGNET COUNT" VARIABLE WAS NOT BEING UPDATED UPON INITIAL INTERROGATION OF THE GENERATOR. THE MAGNETCOUNT WAS OBTAINED FROM THE LAST PROGRAMMING SESSION ON THE HANDHELD FROM THE LAST GENERATOR. THE REASON THE WARNING MESSAGE WAS NOT DISPLAYED THAT A MAGNET SWIPE HAD NOT BEEN DETECTED IS DUE TO THE WAY THE HANDHELD SOFTWARE OBTAINS THE MAGNETCOUNT VARIABLE TO DETERMINE IF THE MESSAGE SHOULD BE DISPLAYED. THIS WAS IMPLEMENTED IN V8.0 SOFTWARE UPGRADE OF THE HANDHELD COMPUTERS. IT PROVIDED A SOFTWARE FIX FOR SCENARIOS IN WHICH PERFORMING A MAGNET MODE DIAGNOSTIC WITHOUT PERFORMING A MAGNET SWIPE WILL NOT RESULT IN THE WARNING MESSAGE THAT A MAGNET SWIPE WAS NOT DETECTED. THE HIGH IMPEDANCE EVENT IS CAPTURED IN MFR REPORT #1644487-2010-00815. ANALYSIS OF PROGRAMMING HISTORY. | 0
3244985 | PRODUCT ANALYSIS OF THE HANDHELD AND FLASHCARD WAS COMPLETED. NO SOFTWARE ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS OF THE HANDHELD. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE TOUCHSCREEN DISPLAY WAS UNRESPONSIVE. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><ON (B)(4) 2013, THE PHYSICIAN REPORTED THAT HE WAS HAVING PROBLEMS WITH HIS HANDHELD. TROUBLESHOOTING WAS PERFORMED; HOWEVER, THE ISSUE DID NOT RESOLVE. TROUBLESHOOTING WAS AGAIN PERFORMED ON (B)(6) 2013; HOWEVER, THE PHYSICIAN WAS UNABLE TO GET THE HANDHELD TO MOVE PAST THE ALIGN SCREEN. MULTIPLE HARD RESETS WERE PERFORMED, BUT THIS DID NOT RESOLVE THE ISSUE. IT WAS NOTED THAT THE OFFICE WAS HIT BY LIGHTENING ON (B)(6) 2013, AND SINCE THEN, MANY OF THE PHYSICIAN'S ELECTRONIC HARDWARE HAS NOT BEEN WORKING. THE PHYSICIAN LOST THREE COMPUTERS TO THE INCIDENT. THE HANDHELD WAS PLUGGED INTO THE WALL AT THE TIME OF THE LIGHTENING STRIKE. NO PATIENT'S WERE AFFECTED BY THIS ISSUE. THE PROGRAMMING SYSTEM HAS BEEN RETURNED AND IS PENDING PRODUCT ANALYSIS. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
3156995 | IT WAS INITIALLY REPORTED THAT THE PATIENT WAS REFERRED TO SURGERY DUE TO ¿LEAD IMPEDANCE¿. ADDITIONAL CLARIFICATION WAS RECEIVED THAT THE PATIENT HAD LOW IMPEDANCE. FOLLOW-UP INDICATED THAT THE PATIENT WAS HAVING AN INCREASE IN SEIZURES, BELOW PRE-VNS BASELINE WHICH WAS ATTRIBUTED TO THE LOW IMPEDANCE. FURTHER INFORMATION WAS RECEIVED THAT THE GENERATOR DIAGNOSTICS WAS THE SOURCE OF THE LOW IMPEDANCE. THE FUNCTION OF THE GENERATOR DIAGNOSTICS IS TO BE RUN IN THE OPERATING ROOM WITH THE TEST RESISTER TO EVALUATE THE FUNCTIONALITY OF THE GENERATOR. IF A GENERATOR DIAGNOSTICS IS RUN WITH AN IMPLANTED PATIENT IT WILL ALWAYS GIVE LOW IMPEDANCE IF RUN ON AN IMPLANTED PATIENT. PART OF THE PROGRAMMING FOR THE GENERATOR DIAGNOSTICS SET THE GENERATOR TO STANDARD SETTINGS WITH AN OUTPUT CURRENT OF 0MA. THE PATIENT BEING TURNED OFF BY THE GENERATOR DIAGNOSTICS IS BELIEVED TO BE THE CAUSE OF THE INCREASE IN SEIZURES. ADDITIONAL INFORMATION RECEIVED INDICATED THAT THE PATIENT¿S VNS GENERATOR WAS TURNED OFF AFTER IT WAS THOUGHT TO BE AN IMPEDANCE ISSUE. THE PATIENT¿S VNS GENERATOR AND LEAD WERE REPLACED ON (B)(6) 2013 AND WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2013. PRODUCT ANALYSIS WAS COMPLETED AND THERE WERE NO ANOMALIES FOUND WITH THE PULSE GENERATOR. ABRADED OPENINGS WERE IDENTIFIED IN THE OUTER AND THE INNER SILICONE TUBING OF BOTH LEAD COILS OF THE VNS LEAD. NO OTHER ANOMALIES WERE IDENTIFIED IN THE RETURNED LEAD PORTION OTHER THAN TYPICAL WEAR AND EXPLANT RELATED OBSERVATION. THE PRODUCT ANALYSIS OF THE VNS LEAD IS REPORTED IN MDR # 1644487-2013-01725. ANALYSIS OF PROGRAMMING/DIAGNOSTIC HISTORY PERFORMED.><INITIAL REPORT INADVERTENTLY MISSED PRODUCT INFORMATION. | 0
2881812 | OUR REGIONAL MANAGER IN (B)(6) REPORTED THAT A HANDHELD COMPUTER IS NOT WORKING. THE PROGRAMMING SOFTWARE DOES NOT START EVEN AFTER RESETTING. THEY PERFORMED SEVERAL RESETS AND ONLY THE WINDOWS SCREEN APPEARS. THE HANDHELD AND SOFTWARE WERE RETURNED FOR ANALYSIS.><AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD. A REVIEW OF THE SOFTWARE IN THE RETURNED HANDHELD GIVES THE INDICATION THAT AN UPGRADE WAS SUCCESSFULLY COMPLETED ON THE RETURNED HANDHELD. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A MISSING 2577 FOLDER THAT CONTAINS THE VNS INSTALLATION FILES. SINCE THE FILES WERE MISSING, THE SOFTWARE WAS UNABLE TO BE INSTALLED INTO THE HANDHELD. A ROOT CAUSE FOR THE MISSING FILES COULD NOT BE DETERMINED BASED ON THE AVAILABLE INFORMATION. NO OTHER ANOMALIES WERE IDENTIFIED. . | 0
3132015 | DURING REVIEW OF PROGRAMMING HISTORY IT WAS NOTED THAT THE PATIENT CAME INTO AN APPOINTMENT SET AT UNINTENTIONAL SETTINGS. IT IS SUSPECTED THAT THERE WAS AN INCOMPLETE DIAGNOSTICS AT THE PREVIOUS APPOINTMENT. THERE WAS NO FINAL INTERROGATION TO CONFIRM THE PATIENT LEFT AT THE INTENDED SETTINGS. THE PATIENT¿S SETTINGS WERE CORRECTED AT THE APPOINTMENT WHEN THE SETTING CHANGE WAS OBSERVED. THE BATTERY LIFE CALCULATION SHOWED 0.24 YEARS UNTIL NEOS. DATE OF EVENT, CORRECTED DATA: PREVIOUSLY SUBMITTED MDR INDICATED AN INCORRECT EVENT DATE. RELEVANT TESTS/LABORATORY DATA, INCLUDING DATES: PREVIOUSLY SUBMITTED MDR INDICATED AN INCORRECT DATA DATES.><ANALYSIS OF PROGRAMMING HISTORY. | 0
2894103 | ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2013 WHEN IT WAS REPORTED THAT THE HANDHELD AND FLASHCARD WERE RECEIVED BY THE MANUFACTURER FOR PRODUCT ANALYSIS ON (B)(6)2013. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT YET BEEN COMPLETED.><ON (B)(4) 2012 IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS FREEZING ON VARIOUS SCREENS THE PREVIOUS WEDNESDAY. MULTIPLE HARD RESETS WERE PERFORMED BUT THE SCREEN WAS STILL FREEZING. ATTEMPTS FOR THE RETURN OF THE HANDHELD AND FLASHCARD WERE MADE BUT THE PRODUCTS HAVE NOT BEEN RETURNED FOR PRODUCT ANALYSIS TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED ON (B)(4) 2013 WHEN PRODUCT ANALYSIS WAS COMPLETED ON THE HANDHELD AND FLASHCARD. AN ANALYSIS WAS PERFORMED ON THE HANDHELD AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO SPECIFICATIONS. DURING THE ANALYSIS, NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2799785 | ON (B)(6) 2012, IT WAS REPORTED THAT A PHYSICIAN'S HANDHELD DEVICE WAS CONTINUALLY FREEZING, TYPICALLY ON THE MAIN MENU BEFORE IT WAS USED TO INTERROGATE AND PROGRAM PATIENTS. A HARD RESET RESOLVED THE FROZEN SCREEN, BUT THE DEVICE WOULD CONTINUE TO FREEZE LATER. ATTEMPTS FOR PRODUCT RETURN HAVE BEEN UNSUCCESSFUL.><THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED ON (B)(6) 2012. PRODUCT ANALYSIS WAS COMPLETED. THE FLASHCARD WAS RETURNED FOR THE FOLLOWING ALLEGED REASON: "COMPUTER SOFTWARE ERROR, FROZEN SCREEN". AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE HANDHELD, AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
3131696 | REPORTER INDICATED A VNS DELL X5 COMPUTER WAS FREEZING AT THE INITIAL SCREEN EVERY TIME AN INTERROGATION WAS ATTEMPTED. PERFORMING A HARD RESET DOES NOT RESOLVE THE SCREEN FREEZING. THE COMPUTER AND FLASHCARD HAVE BEEN RETURNED AND ARE PENDING ANALYSIS.><ANALYSIS OF THE COMPUTER AND FLASHCARD WAS COMPLETED.NO ANOMALIES ASSOCIATED WITH THE HANDHELD COMPUTER WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. . | 0
2830901 | PRODUCT ANALYSIS WAS COMPLETED ON THE RETURNED COMPUTER, FLASHCARD, AND DEMONSTRATOR VNS GENERATOR. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS OF THE HANDHELD AND FLASHCARD SOFTWARE IDENTIFIED THAT THE (B)(6) DATABASE WAS CORRUPT. A ROOT CAUSE FOR FILE CORRUPTION COULD NOT BE IDENTIFIED BASED ON THE AVAILABLE INFORMATION. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE (B)(6) DATABASE WAS NOT CORRUPT, GIVING THE INDICATION THAT THE FILE CORRUPTION DID NOT OCCUR PRIOR TO THE 8.1 SOFTWARE UPGRADE. BECAUSE THE ARCHIVE DATABASES WERE ALSO CORRUPT, IT COULD NOT BE DETERMINED IF THE FILE CORRUPTION OCCURRED DURING OR AFTER THE V8.1 UPGRADE. NO FURTHER ANOMALIES WERE IDENTIFIED. RESULTS OF DIAGNOSTIC TESTING FOR THE GENERATOR INDICATED THE DEVICE WAS OPERATING PROPERLY AND COMMUNICATED PROPERLY. ELECTRICAL TEST RESULTS SHOWED THAT THE PULSE GENERATOR PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THERE WERE NO ADVERSE FUNCTIONAL, MECHANICAL, OR VISUAL ISSUES IDENTIFIED WITH THE RETURNED GENERATOR.><REPORTER INDICATED A VNS DELL X5 COMPUTER WAS DISPLAYING SQL MESSAGES FOLLOWING AS SOFTWARE UPGRADE TO VNS VERSION 8.1 SOFTWARE. THE SQL ERROR WAS ALSO DUPLICATED WHEN A DEMONSTRATOR VNS GENERATOR WAS TESTED. THE COMPUTER, POWER CORD, SERIAL CABLE, FLASHCARD, AND DEMONSTRATOR VNS GENERATOR HAVE BEEN RETURNED AND ARE PENDING ANALYSIS. . | 0
2985950 | THE HANDHELD COMPUTER AND FLASHCARD WERE RETURNED TO THE MANUFACTURER ON (B)(6) 2013 FOR ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE HANDHELD COMPUTER. NO ANOMALIES ASSOCIATED WITH THE HANDHELD COMPUTER WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS.><REPORTER INDICATED THAT A VNS DELL X5 COMPUTER WITH VERSION 8.1 SOFTWARE WAS CONTINUALLY FREEZING AT THE "INTERROGATION SUCCESSFUL" SCREEN FOR AN UNKNOWN AMOUNT OF TIME. TROUBLESHOOTING WITH THE COMPUTER NOTED THE "SCREEN FREEZING" WOULD OCCUR FOR 10-15 SECONDS AT A TIME ONLY. THE COMPUTER WILL NOT BE RETURNED AT THIS TIME. . | 0
3193066 | THE PATIENT'S VNS GENERATOR WAS RETURNED FOR PRODUCT ANALYSIS AFTER BEING EXPLANTED DUE TO THE DEVICE NEARING BATTERY DEPLETION. UPON RECEIPT OF THE DEVICE, IT WAS NOTED THAT THE SETTINGS WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST. FOLLOW UP WAS PERFORMED WITH THE SITE TO SEE IF THEY HAD THE PATIENT'S SETTINGS OR DIAGNOSTIC HISTORY PRIOR TO THE EXPLANT SURGERY; HOWEVER, THIS INFORMATION WAS UNAVAILABLE. NO INFORMATION WAS PROVIDED. A REVIEW OF PROGRAMMING HISTORY FOUND DATA UP TO THE YEAR 2010. IT IS UNKNOWN IF AND WHEN A FAULTED DIAGNOSTIC TEST MAY HAVE OCCURRED. NO OTHER INFORMATION HAS BEEN PROVIDED. THE INCORRECT DATE WAS INADVERTENTLY LISTED ON THE INITIAL REPORT. THE DATE OF THE REPORT WAS (B)(4) 2013 THE SAME DATE AS IN MANUFACTURER REPORT NUMBER 1644487-2013-01957. THE INFORMATION SUBMITTED ON THE INITIAL REPORT WAS INADVERTENTLY SENT, AS IT IS DUPLICATE INFORMATION TO THAT WHICH WAS REPORTED IN MANUFACTURER REPORT NUMBER 1644487-2013-01957. THE INCORRECT AWARE DATE WAS INADVERTENTLY LISTED ON THE INITIAL REPORT. THE AWARE DATE SHOULD HAVE BEEN (B)(4) 2013, THE SAME AWARE DATE AS IN MANUFACTURER REPORT NUMBER 1644487-2013-01957.><ANALYSIS OF PROGRAMMING HISTORY. | 0
2922427 | DURING REVIEW OF THE VNS PATIENT'S PROGRAMMING HISTORY ON (B)(6) 2012 IT WAS DISCOVERED THAT TWO FAULTED SYSTEM DIAGNOSTICS TESTS OCCURRED WHICH CHANGED THE PATIENT'S SETTINGS ON TWO DIFFERENT DATES. ON (B)(6) 2008 A FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED AND NO FINAL INTERROGATION WAS PERFORMED. THE PATIENT THEN PRESENTED AT THE NEXT VISIT ON (B)(6) 2008 WITH THE INCORRECT SETTINGS OF OUTPUT=0MA/FREQUENCY=20HZ/PULSE WIDTH=500USEC/ON TIME=30SEC/OFF TIME=60MIN/MAGNET OUTPUT=0MA/MAGNET PULSE WIDTH=500USEC/MAGNET ON TIME=30SEC. THE PATIENT WAS THEN REPROGRAMMED TO THE DESIRED SETTINGS. A FEW MONTHS LATER ON (B)(6) 2008 ANOTHER FAULTED SYSTEM DIAGNOSTICS TEST OCCURRED, AGAIN WITH NO FINAL INTERROGATION. ON THE PATIENT'S NEXT VISIT ON (B)(6) 2008 THE PATIENT WAS FOUND AT THE INCORRECT SETTINGS OF OUTPUT=1MA/FREQUENCY=20HZ/PULSE WIDTH=500USEC/ON TIME=30SEC/OFF TIME=60MIN/MAGNET OUTPUT=1MA/MAGNET PULSE WIDTH=500USEC/MAGNET ON TIME=30SEC. THE PATIENT WAS AGAIN REPROGRAMMED TO THE CORRECT SETTINGS. NO PATIENT HARM HAS BEEN REPORTED TO HAVE OCCURRED DUE TO THESE SETTINGS CHANGES. ANALYSIS OF PROGRAMMING HISTORY. | 0
2839283 | INTERNAL PROGRAMMING HISTORY WAS REVIEWED AND IT WAS NOTED THAT THERE WAS A FAULTED SYSTEM DIAGNOSTIC TEST. PROGRAMMING HISTORY WAS REVIEWED FOR THE AVAILABLE DATA FROM IMPLANT DATE, (B)(6) 2008, TO (B)(6) 2012. DURING THIS TIME, THREE FAULTED SYSTEM DIAGNOSTIC TESTS WERE PERFORMED. TWO OF THESE TESTS DID NOT AFFECT THE PATIENT'S PROGRAMMED SETTINGS; HOWEVER, THE TEST PERFORMED ON THE ADJUSTED DATE (B)(6) 2009 LEFT THE PATIENT WITH THE UNINTENDED SETTINGS 0/20/500/30/10. THESE SETTINGS WERE ADJUSTED AT THE NEXT VISIT ON (B)(6) 2009. NO OTHER ANOMALIES WERE OBSERVED. TRAINING WAS PROVIDED TO THE PHYSICIAN VIA TRAINING ON THE UPGRADED SOFTWARE ON (B)(4) 2011. PRODUCT ANALYSIS WAS PERFORMED ON THE HANDHELD AND SOFTWARE ON (B)(4) 2011 FOR A FROZEN SCREEN ERROR (REPORTED IN MFR #: 1644487-2011-00927). ANALYSIS OF PROGRAMMING HISTORY.><PATIENT IDENTIFIER, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED INCORRECT PATIENT IDENTIFIER. LOT #, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED INCORRECT LOT NUMBER. DEVICE AVAILABLE FOR EVALUATION, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED INCORRECT INFORMATION. NAME AND ADDRESS, CORRECTED DATA: INITIAL MDR INADVERTENTLY LISTED INCORRECT REPORTER. LABELED FOR SINGLE USE, CORRECTED DATA: INITIAL MDR INADVERTENTLY SELECTED INCORRECT CHOICE. USAGE OF DEVICE, CORRECTED DATA: INITIAL MDR INADVERTENTLY SELECTED INCORRECT CHOICE. | 0
3011413 | DURING MANUFACTURER REVIEW OF A VNS PATIENT'S PROGRAMMING HISTORY, IT WAS NOTED A FAULTED SYSTEMS DIAGNOSTICS TEST OCCURRED ON (B)(6) 2011 WHICH CHANGED THE SETTINGS. THE SETTINGS WERE CORRECTED IMMEDIATELY EXCEPT FOR THE MAGNET MODE ON TIME, WHICH REMAINED AT 30 SECONDS WHEN IT WAS FORMALLY 60 SECONDS. .><ANALYSIS OF PROGRAMMING HISTORY. | 0
2887469 | AN ANALYSIS WAS PERFORMED ON THE HANDHELD. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE ALLEGED SCREEN FREEZE COMPLAINT IS A KNOWN ISSUE THAT HAS BEEN EVALUATED AND IS READILY CONFIRMED. NO OTHER ISSUES WERE OBSERVED WITH THE RETUNED HANDHELD. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT A NEUROLOGIST'S (B)(4) HANDHELD COMPUTER WAS FROZEN. OUR CONSULTANT WAS GOING TO GO TO THE NEUROLOGY OFFICE AND INVESTIGATE. THE HANDHELD COMPUTER WAS FROZEN ON THE INTERROGATE SUCCESSFUL SCREEN WHEN SHE PICKED IT UP. WHEN TROUBLESHOOTING THE HANDHELD IT WAS ON THE CYBERONICS MAIN MENU. THE HANDHELD WAS NOT PLUGGED IN WHEN THE EVENT OCCURRED, NOR DID SHE PERFORM A HARD RESET. THE CONSULTANT WAS ABLE TO INTERROGATE AND PERFORM DIAGNOSTICS WITH HER DEMO GENERATOR WITHOUT A PROBLEM. THE HANDHELD IS NOT BEING RETURNED FOR ANALYSIS AS THE EVENT APPEARS TO HAVE RESOLVED. SOFTWARE FIRMWARE CAUSED EVENT.><ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD WAS CONTINUING TO HAVE THE SAME PROBLEMS AND WOULD BE RETURNED. THE EVENT RESOLVES WITH A HARD RESET BUT OCCURS AGAIN. THE HANDHELD COMPUTER IS PENDING COMPLETION OF ANALYSIS. . | 0
3181853 | ON (B)(4) 2013 PRODUCT ANALYSIS WAS COMPLETED ON THE HANDHELD AND FLASHCARD. AN ANALYSIS WAS PERFORMED ON THE HANDHELD NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><THE HANDHELD AND SOFTWARE WERE RETURNED ON JUNE 27, 2013 AND ARE PENDING PRODUCT ANALYSIS.><ON (B)(6) 2013, IT WAS REPORTED THAT THE PHYSICIAN WAS HAVING PROBLEMS WITH HIS PROGRAMMING SYSTEM. IT WAS CLARIFIED THAT THE SCREEN HAS BEEN CONTINUALLY LOCKING UP OVER THE PAST FEW MONTHS. TROUBLESHOOTING STEPS INCLUDED CONFIRMING THE LOCK BUTTON WAS NEVER ENGAGED AND PERFORMING HARD RESETS TO GET THE SCREEN TO UNFREEZE. THE PHYSICIAN STATED THAT HE HAS BEEN ABLE TO PROGRAM AND INTERROGATE PATIENTS, BUT IS TIRED OF THE TIME IT TAKES TO DO SO BECAUSE OF THE PROGRAMMING SYSTEM. THE FROZEN SCREEN ISSUE OCCURRED ON THE PARAMETERS SCREEN. THE PROGRAMMING SYSTEM IS PENDING RETURN FOR PRODUCT ANALYSIS. . | 0
2833678 | AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. AFTER 7 DAYS, CHECKED THE HANDHELD DATE AND TIME WERE CHECKED AND VERIFIED THAT THEY WERE CORRECT. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><A VNS PROGRAMMING PHYSICIAN REPORTED THAT HIS HANDHELD COMPUTER WOULD NOT KEEP TIME. IT WAS REPORTED THAT THE DEVICE CHANGES THE DATE AND TIME BACK CONSTANTLY AND WHEN THE PHYSICIAN CORRECTS THE DATE AND TIME, IT WOULD REMAIN CORRECTED FOR A FEW DAYS AND THEN RANDOMLY CHANGE AGAIN A FEW DAYS LATER. THE PHYSICIAN KEEPS THE HANDHELD COMPUTER PLUGGED IN NEAR HIS DESK AND USES IT ON A FREQUENT BASIS. THE HANDHELD HAS BEEN RETURNED FOR ANALYSIS AND COMPLETION IS PENDING. . | 0
3229629 | THE PHYSICIAN'S OFFICE REPORTED THAT THEY WERE EXPERIENCING PROBLEMS WITH THE DELL X5 HANDHELD COMPUTER GETTING STUCK WHEN THEY TRYING TO INTERROGATE. AFTER SELECTING INTERROGATION, IT LOOKED LIKE IT IS TRYING TO INTERROGATE BUT IT DID NOT SHOW THE PATIENT'S SETTINGS. SOMETIMES IT WOULD WORK IF THEY TRY IT AGAIN, BUT THEY REQUESTED A REPLACEMENT COMPUTER/SOFTWARE BECAUSE IT WAS TAKING TOO LONG TO RETRY. A REPLACEMENT WAS PROVIDED TO THE OFFICE. THE FAULTY COMPUTER WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS, BUT ANALYSIS HAS NOT BEEN COMPLETED TO DATE. THE SOFTWARE WAS NOT RETURNED WITH THE COMPUTER. ATTEMPTS FOR RETURN OF THE SOFTWARE HAVE BEEN UNSUCCESSFUL TO DATE.><THE PHYSICIAN'S OFFICE REPORTED THAT IT WAS BELIEVED THE SOFTWARE (FLASHCARD) WAS RETURNED WITH THE HANDHELD DEVICE. SINCE IT WAS NOT, THE OFFICE PLANNED TO LOOK FOR THE FLASHCARD. THE RETURN PRODUCT FORM INDICATED THE REASON FOR PRODUCT RETURN WAS "NOT FUNCTIONING.' AN ANALYSIS WAS PERFORMED ON THE HANDHELD, AND NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE REPORTED EVENT COULD NOT BE CONFIRMED SINCE THE SOFTWARE WAS NOT RETURNED WITH THE HANDHELD. . | 0
3272838 | THE HANDHELD AND FLASHCARD WERE RETURNED FOR ANALYSIS ON (B)(6) 2013. ANALYSIS OF THE SOFTWARE WAS COMPLETED ON (B)(6) 2013. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO SPECIFICATIONS. THE AVAILABLE INFORMATION IN THE ISSUE FILE GIVES NO INDICATION THAT THE ALLEGED SCREEN FREEZE COMPLAINT IS RELATED TO THE EVENT WHICH HAS BEEN PREVIOUSLY INVESTIGATED BY THE MANUFACTURER. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE HANDHELD WAS REPORTED IN MFR. REPORT # 1644487-2013-02354/02><ON (B)(6) 2013, A PHYSICIAN REPORTED ISSUES WITH HIS PROGRAMMING SYSTEM BEGINNING SIX WEEKS PRIOR. THE DEVICE WAS REPORTEDLY SLOWING DOWN AND HAVING SCREEN FREEZES. THE PHYSICIAN ALSO REPORTED THAT THE BACKLIGHT WAS VERY DIM AND THAT THE DEVICE WOULD NOT RESPOND TO SCREEN TOUCHES. CONNECTIONS WERE CHECKED, NO PINS WERE BENT, THE SCREEN WAS CLEANED, THE WAND BATTERY WAS REPLACED, AND HARD RESETS WERE PERFORMED. THE ISSUES WERE CONSISTENT AS THEY WERE HAPPENING WITH MULTIPLE PATIENTS. THE SOFTWARE EVENTS ARE CAPTURED IN MFR REPORT #1644487-2013-02355. THE PROGRAMMING COMPUTER EVENTS ARE CAPTURED IN MFR REPORT #1644487-2013-02354.><ON (B)(6) 2013, THE PHYSICIAN CALLED THE MANUFACTURER TO FOLLOW UP ON THE ISSUES WITH HIS PROGRAMMING SYSTEM. HE STATED THAT THE HANDHELD WAS GOING OUT (MFR #: 1644487-2013-02354) AND THE WAND WOULD NO LONGER HOLD A CHARGE. THE WAND ISSUE IS BEING INVESTIGATED IN ANOTHER COMPLAINT FILE. NO ADDITIONAL INFORMATION WAS PROVIDED ON THE PREVIOUS SCREEN FREEZE AND SLOWNESS ISSUES. ATTEMPTS FOR PRODUCT RETURN ARE BEING MADE; HOWEVER, THE PRODUCT HAS NOT YET BEEN RETURNED. . | 0
3224352 | FOLLOW UP CONFIRMED THAT THE NURSE PROVIDED ALL OF THE INFORMATION SHE COULD. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED.><ON (B)(6) 2013, IT WAS REPORTED THAT THE PATIENT'S GENERATOR SETTINGS WERE INTERROGATED AND FOUND TO BE STIMULATING FOR 30 SECONDS EVERY 60 MINUTES, WITH A MAGNET OUTPUT CURRENT OF 1MA. HOWEVER, THE PATIENT SHOULD HAVE BEEN RECEIVING 30 SECONDS OF STIMULATION EVERY 5 MINUTES, WITH A 0MA OUTPUT CURRENT. IT WAS STATED THAT IT WAS KNOWN A FAULTED SYSTEM DIAGNOSTIC TEST COULD LEAD TO THESE UNINTENDED SETTINGS; HOWEVER, THE NURSE STATED THAT SHE INTERROGATED THE PATIENT'S DEVICE AT THE END OF THE LAST VISIT TO MAKE SURE THE SETTINGS WERE OK. NO OTHER INFORMATION WAS PROVIDED AS THE NURSE DID NOT GIVE ANY ADDITIONAL DETAILS, EXCEPT THAT THE PATIENT IS IMPLANTED WITH A MODEL 102 GENERATOR. . | 0
2957665 | ADDITIONAL INFORMATION WAS RECEIVED THAT AN INCOMPLETE DIAGNOSTIC DID OCCUR AND RESULTED IN THE PATIENT BEING PROGRAM TO 1.0 MA FOR A SHORT TIME. THE PHYSICIAN DID NOT NOTICE THIS CHANGE PRIOR TO THE PATIENT LEAVING THE OFFICE AND THE PATIENT LEFT WITH HIS GENERATOR TURNED ON.><IT WAS INITIALLY REPORTED THAT THE PATIENT HAD PAINFUL STIMULATION IN HIS JAW DURING A SYSTEMS DIAGNOSTICS. THE PATIENT HAD BEEN PROGRAMMING OFF PREVIOUSLY AND THE PAINFUL STIMULATION WAS BELIEVED TO BE RELATED TO THE STIMULATION FROM DIAGNOSTICS. THE PATIENT LATER CONTACT THE MANUFACTURER AND REPORTED THE INTENSE PAIN IN THE JAW CONTINUED AFTER THE APPOINTMENT AND SEEMED TO BE OCCURRING EVERY HOUR. THE PATIENT FOLLOWED UP WITH THE PHYSICIAN AND IT WAS DETERMINED THAT THE PATIENT WAS ACCIDENTALLY LEFT ON. THE PATIENT WAS THEN PROGRAMMING OFF AGAIN AND IS NOT HAVING ANY ADDITIONAL INTENSE PAIN BUT STILL HAS SOME RESIDUAL PAIN IN THE JAW. SINCE THE PATIENT WAS REPORTING THAT THE INTENSE PAIN WAS OCCURRING EVERY HOUR THERE IS A SUSPECTED INCOMPLETE DIAGNOSTICS THAT MAY HAVE OCCURRED THAT LEFT THE PATIENT PROGRAMMED ON AND TO AN OFF TIME OF 60 MINUTES. THE PHYSICIAN STATES THAT THE ISSUE OF PAIN HAS RESOLVED, AND WAS A RESULT OF THE VNS BEING PROGRAMMED ON FOR A SHORT TIME BUT IT IS TURNED OFF NOW AND THE PATIENT HAS COMPLETED RECOVERED AFTER ALLOWING TIME FOR HEALING. IT IS STILL UNCLEAR IF AN INCOMPLETE DIAGNOSTICS OCCURRED. GOOD FAITH ATTEMPTS FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. . | 0
3388264 | ON (B)(6) 2013, GE HEALTHCARE RECEIVED INFORMATION CONCERNING A DEVICE MALFUNCTION WITH NO ASSOCIATED ADVERSE EVENT. AN ORDER OF RAPIDSTRAND RX WAS SENT TO A USER FACILITY WITH INCORRECTLY LOADED BRACHYTHERAPY IMPLANT NEEDLES (ERROR IN THE CONFIGURATION OF LOADED SEEDS). BEFORE IMPLANTATION THE FACILITY WAS ADVISED BY GE HEALTHCARE OF THE ERROR AND WAS ABLE TO REVIEW THE CONFIGURATION OF THE IMPLANT. SINCE ALL NEEDLES WERE SYMMETRICALLY LOADED, THIS IMPLANT WAS NOT AFFECTED AND IT WAS SUCCESSFULLY CARRIED OUT ACCORDING TO THE TREATMENT PLAN. ALTHOUGH IN THIS INSTANCE THE IMPLANT WAS NOT AFFECTED, DUE TO THE NATURE OF THE ERROR WHICH RESULTED IN TWO OTHER INCORRECTLY LABELED IMPLANTS FOR DIFFERENT PATIENTS, THIS MALFUNCTION WAS ASSESSED AS HAVING THE POTENTIAL TO CAUSE DEATH OR SERIOUS INJURY SHOULD IT RECUR. QUALITY ASSURANCE INVESTIGATION (B)(4) IS ONGOING. CONCLUSIONS: ERROR BY CONTRACT MANUFACTURER COMPUTER OPERATOR WHEN ENTERING THE TREATMENT PLAN DETAILS INTO THE ORDER FULFILLMENT SOFTWARE. (E.G. ADDITIONAL LOT# 130247A). EVALUATION SUMMARY: PRELIMINARY EVALUATION RESULTS: THIS FAILURE OCCURRED BECAUSE THE PATIENT DOES PLAN PROVIDED WITH EACH ORDER WAS INCORRECTLY ENTERED BY GE HEALTHCARE'S CONTRACT MANUFACTURER, THERAGENICS, INTO THEIR OR (B)(4) FULFILLMENT SOFTWARE SYSTEM (SYSTEM OPERATOR ERROR) PRIOR TO ASSEMBLY. THE AFFECTED USER FACILITY PROVIDES PATIENT-SPECIFIC NEEDLE LOADING PLANS GENERATED FROM A SIMUPLAN PLATFORM. THIS PLATFORM GENERATES AND DISPLAYS THE LOADING PLAN AS INDIVIDUAL SEALED SOURCE IMPLANT NEEDLES PLAN IN THE OPPOSITE ORDER (FROM RIGHT TO LEFT AND NOT LEFT TO RIGHT) TO MOST COMMERCIALLY AVAILABLE SYSTEMS. THIS WAS NOT TAKEN INTO ACCOUNT BY THE NEEDLES BEING INCORRECTLY LOADED AS DETAILED BELOW. PATIENT DOES PLANT CONSISTED OF 66 IODINE-125 BRACHYTHERAPY SOURCES TO BE LOADED IN TO 23 IMPLANT NEEDLES. ALL NEEDLES HAD ASYMMETRICAL LOADING PLANS AND WERE CORRECTLY LOADED. THE CUSTOMER (BOTH DOCTORS AND PHYSICS DEPARTMENT) WAS ADVISED IN WRITING AND BY TELEPHONE BY GE HEALTHCARE CUSTOMER SERVICE, LOCAL SALES REPRESENTATIVE, AND THE GE HEALTHCARE MEDICAL DIRECTOR ON THE DAY BEFORE THE IMPLANT DATE. NO ADJUSTMENTS WERE NECESSARY FOR THE ORDER WHICH WAS SUBSEQUENTLY IMPLANTED SUCCESSFULLY ACCORDING TO THE TREATMENT PLAN. | 0
3050264 | REMOTE TELEMETRY SYSTEM ISSUES WITH LACK OF TRANSMISSION TO CENTRAL MONITORING STATION. PHILIPS PRODUCT PICIX AND MX40 TELEMETRY BOX. WHOLE SECTIONS OF THE STATION HAVE GONE DOWN FOR AS LONG AS TWO HOURS AS WELL AS VARIOUS TELEMETRY BOXES FOR SHORTER TIME PERIODS. WE HAVE NOT FOUND A CAUSE TO THESE PROBLEMS AT THIS TIME. | 1
3036815 | IT WAS REPORTED THAT THE CODE BLUE RESUSCITATOR DID NOT ALLOW OXYGEN TO FLOW THROUGH DURING USE. THE DEVICE WAS ATTACHED TO A SECOND OXYGEN LINE AND DID NOT FLOW. THE DEVICE WAS REPLACED AND PROCEDURE RESUMED. THE CUSTOMER REPORTED OXYGEN TUBING WAS KINKED AN PREVENTED OXYGEN FROM FLOWING TO THE PT. THE RESUSCITATOR REMAINED FUNCTIONAL. THERE WAS NO INJURY REPORTED. THE EXACT DATE OF EVENT WAS UNK. AT THIS TIME, THE LOT NUMBER IS UNK. THEREFORE, THE DEVICE MANUFACTURE DATE CANNOT BE DETERMINED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A F/U REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
2904495 | THE CUSTOMER REPORTED AN INCIDENT INVOLVING THE CRITICAL CARE MANAGER (ELECTRONIC HEALTH RECORD) SOFTWARE APPLICATION, WHEREBY AN ERROR IN THE ELECTRONIC DOCUMENTATION FOR EXTENDING A MEDICATION ORDER RESULTED WITH A PT MISSING SIX DAYS OF AN ELECTROLYTE MINERAL - NAC1 50MG (SALT). THE INCIDENT DID NOT RESULT IN ANY INJURY TO THE PT. OUR INVESTIGATION INTO THE REPORTED INCIDENT HAS CONCLUDED THAT THE ROOT CAUSE RESULTED FROM A LIMITATION IN SOFTWARE PERFORMANCE WITHIN THE PICIS PERIOP AND CRITICAL CARE SUITE (ELECTRONIC HEALTH RECORD) OF SOFTWARE APPLICATIONS. OUR FINDINGS HAVE REVEALED THAT CUSTOMERS MAY EXPERIENCE AN ERROR IN PT CASE RECORD DOCUMENTATION FOR MEDICATION ORDER. WHEREBY, AN ORDER MAY FAIL TO BE EXTENDED OR DISCONTINUED IN DOCUMENTATION AT WORKSTATIONS OTHER THAN THE ONE WHERE THIS ACTION WAS TAKEN. THIS ERROR MAY OCCUR DURING A SPECIFIC TIMING OF SOFTWARE SYSTEM EVENTS AND COMBINATION OF DOCUMENTED ACTIONS, AND PRECLUDED BY AN INTERRUPTION IN OUR APPLICATION SYSTEM UPDATE SERVICES THAT IS USUALLY CAUSED BY A NETWORK DISRUPTION. THE ERROR IS ONLY SEEN IN WORKFLOWS WHERE USERS DOCUMENT WITHIN THE SAME PT RECORD AT DIFFERENT WORKSTATIONS. AN INDICATION OF THE ERROR CONDITION WAS OBSERVED ON THE APPLICATION FLOWSHEET DISPLAY BY THE PRESENCE OF A RED STATUS ICON, WHICH IS NORMALLY USED TO INDICATE THE STATUS OF PENDING MEDICATION ORDER TASKS. HOWEVER, WHEN OPENING THE ORDER STATUS WINDOW THE PAGE WAS BLANK DUE TO THE SOFTWARE ERROR. REPORTS OF THIS ERROR CONDITION FROM CLIENTS ARE RARE, AS MULTIPLE SEQUENTIAL CONDITIONS MUST TAKE PLACE. A RESOLUTION TO ADDRESS THE POTENTIAL FOR THIS ERROR IS AVAILABLE IN A CURRENT SOFTWARE SERVICE PACK. THE CUSTOMER WHO REPORTED THE INCIDENT HAS BEEN RECOMMENDED TO UPGRADE TO THE LATEST AVAILABLE SERVICE PACK. PICIS IS PRESENTLY EVALUATING THE NEED TO COMMUNICATE WITH CUSTOMER ABOUT THE POTENTIAL FOR SEEING THIS ERROR. | 1
3422151 | RECALL OF INFINI NUCLEAR MEDICINE SYSTEM. WE ONLY HAVE THE OPTIMA MR450W 1.5TR MRI FOR INTELIRAD PACS. OUR IT OPERATIONS DIRECTOR WITH THE RADIOLOGY GROUP CONFIRMED THAT WE DO NOT USE THIS VERSION EFFECTING THE RECALL. | 0
3041516 | A PATIENT IN THE (B)(6) WAS UNDERGOING THERAPEUTIC PLASMA EXCHANGE (TPE) TREATMENT. APPROXIMATELY 10 MINUTES INTO THE TREATMENT THE PRISMA MACHINE GENERATED SEVERAL "AIR IN SET" ALARMS AND AIR WAS OBSERVED INSIDE THE PLASMA FILTER. THE TREATMENT WAS DISCONTINUED AND THE EXTRACORPOREAL BLOOD WAS NOT RETURNED TO THE PATIENT. THE ESTIMATED BLOOD LOSS TO THE PATIENT WAS 200 ML. THE PATIENT WAS NOT SYMPTOMATIC AS A RESULT OF THE BLOOD LOSS HOWEVER, IT WAS REPORTED THAT THE PATIENT WAS TRANSFUSED WITH A UNIT OF BLOOD THAT EVENING. THERE WAS NO FURTHER MEDICAL OR LABORATORY DATA PROVIDED. THE PRISMA MACHINE GENERATED SEVERAL "AIR IN SET" ALARMS, WHICH IS IN ACCORDANCE WITH SPECIFICATIONS. NO MACHINE MALFUNCTION WAS REPORTED. THE FILTER SET WAS DISCARDED AND THEREFORE, NOT AVAILABLE FOR INSPECTION. THE REVIEW OF THE DEVICE HISTORY RECORD OF THIS LOT AND THE COMPLAINT HISTORY FILES HAS SHOWN THAT THERE WERE NO NONCONFORMITIES AND NO OTHER COMPLAINT REPORTED ON THIS LOT. | 0
3359413 | ABOUT AN HOUR INTO THE PROCEDURE (APHERESIS FOR WHITE CELL COLLECTION), THE KIT FAILED AND PROCEDURE HAD TO BE ABORTED AND RESTARTED. THE BLOOD IN THE SYSTEM WAS NOT RETURNED TO PATIENT. THE KIT WAS SENT TO COMPANY FOR ANALYSIS. BLOOD LOSS TO THE PATIENT WAS APPROXIMATELY 150 CC.======================MANUFACTURER RESPONSE FOR APHERESIS COLLECTION KIT, SPECTRA OPTIA (PER SITE REPORTER).======================DEVICE RETURNED TO MANUFACTURER IMMEDIATELY FOLLOWING OCCURRENCE. STILL AWAITING RESPONSE FROM COMPANY ON THEIR ANALYSIS OF KIT. | 0
3354105 | THE EVENT OCCURRED DURING AN APHERESIS COLLECTION PROCEDURE. ABOUT 3.5 HOURS INTO THE PROCEDURE, A BLOOD LEAK INTO CENTRIFUGE WAS NOTED. THE PROCEDURE HAD TO BE ABORTED AND RESTARTED. RINSE BACK WAS UNABLE TO BE DONE, SO SOME BLOOD WAS LOST (~< 150CC). A PIN POINT LEAK WAS THEN NOTED JUST LEFT OF THE COLLECT LINE WHICH IS UNLIKELY TO HAVE OCCURRED FROM LOADING THE KIT INCORRECTLY. ======================MANUFACTURER RESPONSE FOR APHERESIS COLLECTION KIT, SPECTRA OPTIA (PER SITE REPORTER).======================KIT WAS RETURNED TO MANUFACTURER IMMEDIATELY. STILL AWAITING RESPONSE ON THEIR EVALUATION OF THIS KIT. | 0
3485960 | PATIENT HERE FOR PLASMAPHERESIS TREATMENT. UPON HITTING THE START RUN BUTTON, AN ALARM APPEARED WHICH STATED, "AIR DETECTED IN RETURN LINE." ATTEMPTS WERE MADE TO CLEAR THE ALARM BY FOLLOWING THE PROMPTS ON THE MACHINE. MACHINE COULD NOT CLEAR THE ALARM. CARE WAS DELAYED UNTIL THE ANOTHER MACHINE WAS AVAILABLE.====================== MANUFACTURER RESPONSE FOR PLASMAPHERESIS TUBING, TERUMO TUBING SET (PER SITE REPORTER) ====================== TERUMO BCT WAS CONTACTED AT THE TIME. OPERATOR WAS INSTRUCTED TO CLEAN AND PRESSURE THE SENSOR. TUBING SET WAS SENT BACK TO THE COMPANY FOR QUALITY CONTROL. THE SERVICE REP CAME TO DO PREVENTATIVE MAINTENANCE AND WAS NOT ABLE TO REPRODUCE THE ALARM. | 0
3428922 | PLASMAPHERESIS STARTED. RBC DETECTOR ALARMED AND HEMOLYSIS NOTED IN THE COLLECTION TUBING. TERUMO WAS CONTACTED AND AFTER MANY ATTEMPTS TO CORRECT BY INCREASING THE HEMATOCRIT VALUE AND REDUCING THE INLET FLOW RATE WITHOUT RESOLVING THE PROBLEM, THE STAFF PERFORMED RINSE-BACK, THEN PRIMED AND LOADED NEW TUBING. THIS RESOLVED THE ISSUE.THE PATIENT DID NOT SUFFER ANY ILL EFFECTS.====================== MANUFACTURER RESPONSE FOR PLASMAPHERESIS MACHINE AND TUBING, SPECTRA OPTIA (PER SITE REPORTER).====================== THE TUBING WAS RETURNED TO THE MANUFACTURER.THE SET WAS INSPECTED FOR MISSING PARTS, MIS-ASSEMBLIES, KINKS AND OCCLUSIONS. NONE WERE FOUND. BASED ON THE DESCRIPTION OF THE ISSUE AND THE EVALUATION OF THE SET, THE ROOT CAUSE FOR THIS ALARM WAS AN RBC SPILL-OVER, LIKELY CAUSED BY THE USE OF THE PREVIOUS DAYS HEMATOCRIT. ENSURING THE MOST CURRENT AND ACCURATE PATIENT INFORMATION AVAILABLE IS ENTERED INTO THE SPECTRA OPTIA MACHINE AS EARLY AS POSSIBLE IN THE PROCEDURE IS KEY TO THE SYSTEM PERFORMANCE. NO CORRECTION WS PERFORMED FOR THIS INCIDENT BY TERUMO BCT. TRENDS ARE REGULARLY MONITORED TO DETERMINE APPROPRIATE CORRECTIVE ACTIONS BY MANUFACTURING OR ENGINEERING. | 0
2837253 | ON (B)(6) 2012, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT SERIAL DOWNLOADER RECHARGER S/N (B)(4) BECAME HOT TO TOUCH AND BEGAN TO SMOKE WHILE IN USE. BASED ON THE INFO AVAILABLE AT THIS TIME, THERE WERE NO INJURIES ASSOCIATED WITH THIS EVENT. (B)(4). THE INVESTIGATION WAS COMPLETED ON (B)(4) 2013. THE FAILURE WAS DUE TO A DEFECTIVE TRANSISTOR.><APOC INCIDENT #(B)(4). | 0
3252407 | PT ADMITTED WITH SEVERE BRADYCARDIA AND SOME HEART BLOCK. WHILE UNDERGOING PACEMAKER PLACEMENT, INADVERTENTLY, PART OF THE SHEATH THAT WAS A PEEL-BACK SHEATH TO INSTALL THE WIRES BROKE OFF AND COULD NOT BE RETRIEVED. A VENOGRAM CONFIRMED THE PRESSURE OF AN 8-10CM SECTION OF SHEATH THAT WAS IN THE LOWER PORTION OF THE SUBCLAVIAN VEIN. AFTER REVIEWING THIS WITH THE INTERVENTIONAL RADIOLOGIST AND SPEAKING WITH CARDIOVASCULAR THORACIC SURGEONS, IT WAS FELT IT WAS NOT RETRIEVABLE DUE TO THE DEPTH AND ANGLE. A RECOMMENDATION WAS MADE THAT THE PT BE ANTICOAGULATED UNTIL HE COULD ENDOTHELIALIZE THIS AREA NATURALLY. | 0
2926794 | CUSTOMER REPORTED TO BECKMAN COULTER, INC. (BEC) THAT PINK SMOKE CAME OUT OF THE COULTER LH 750 SLIDESTAINER (LH 750 SLIDESTAINER) WHEN THEY POWERED UP THE UNIT WHICH HAD BEEN IDLE FOR SOME TIME. CUSTOMER REPORTED THEY DID NOT SEE SPARKS OR FLAMES. CUSTOMER REPORTED THEY DID NOT CALL THE FIRE DEPARTMENT. CUSTOMER REPORTED THEY POWERED OFF AND UNPLUGGED THE LH 750 SLIDESTAINER UNIT. THERE WAS NO REPORT PATIENT RESULTS WERE AFFECTED. THERE WAS NO REPORT OF ANY ADVERSE EVENT OR INJURY REQUIRING MEDICAL INTERVENTION OR PATIENT TREATMENT. A BEC FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S SITE TO EVALUATE THE INSTRUMENT. CUSTOMER REPORTED TO THE FSE THAT THERE HAD BEEN AN ACCIDENTAL SPILL ON THE NIGHT WHEN THE INCIDENT OCCURRED AND THE INSTRUMENT WAS GIVING AGITATION ERRORS. THE CUSTOMER DID NOT PROVIDE ANY SPECIFICS ON THE ACCIDENTAL SPILL. THE FSE DID NOT OBSERVE ANY VISIBLE FLUID WHILE ON SITE. THE FSE FOUND DAMAGE TO THE BASE BOARD THAT HAD SHORTED OUT CAUSING THE SMOKE. THE FSE FOUND THE AGITATION MODULE AND MOTOR, AS WELL AS OTHER COMPONENTS, HAD NO VISIBLE SIGNS OF DAMAGE. THE FSE REPLACED THE BASEBOARD TO REPAIR THE INSTRUMENT. (B)(4). | 0
3339296 | CULTURES WERE OBTAINED FROM A DEEP SKIN INFECTION INVOLVING AN IMPLANTED MEDICAL DEVICE. MULTIPLE CULTURES GREW SERRATIA MARCESCENS. THE ANTIBIOTIC SENSITIVITIES WERE LOST FROM THE MDDS SECTION OF THE EHR, OR THEY WERE NEVER POSTED DUE TO AN INTERFACE FAILURE. A WORK AROUND WAS REQUIRED TO FIND THE RESULTS, BUT THEY REMAIN ABSENT FROM THE SILO OF THE EHR WHERE THEY SHOULD APPEAR. THIS DEFECT CAUSES DELAYS IN CARE AND ADVERSITY DUE TO DELAYS IN PINPOINTING THE CORRECT ANTIBIOTIC TO USE IN THIS CRITICAL SITUATION. THIS GENRE OF FLAW RAISES DOUBT IN THE HEALTH CARE PROFESSIONALS AS TO WHETHER THE PRESENTATION OF RESULTS ON ANY PT ARE ACCURATE. | 1
3339298 | THE MDDS DISPLAY OF DATA ON THE EHR SHOW RESULTS OF MANY TESTS DONE ON DIFFERENT DATES IN THE SAME COLUMN, WITH EACH TEST SEPARATED BY A HORIZONTAL LINE. HOWEVER, WHEN THE USER EYE SCROLLS DOWN THE SCREEN LOOKING AT RESULTS, THE DATE AT THE TOP OF THE COLUMN STAYS IN MIND, RENDERING MANY OF THE RESULTS ERRONEOUS AS THE USER THINKS ABOUT IT, FOR THE DATE IN QUESTION. THAT RESULTS IN MISTREATMENT. FOR INSTANCE, A LOW POTASSIUM RESULT DONE A WEEK AGO WILL SHOW UP IN THE SAME VERTICAL COLUMN AS A GLUCOSE LEVEL DONE TODAY IF THE PT HAS NOT HAD ANY POTASSIUM LEVELS OBTAINED. SINCE THE ONE OF A WEEK AGO. THE USER SEES THE LOW POTASSIUM AND ORDERS POTASSIUM SUPPLEMENTS HAVING BEEN FOOLED AS TO THE DATE OF THAT LOW POTASSIUM. THIS RESULTS IN INNUMERABLE MISTREATMENTS, THOUGH I AM NOT AWARE OF ANY DEATHS FROM THIS, YET. THIS DEVICE SHOULD BE RECALLED SO THAT THE COLUMNS WITH DIFFERENT DATES ARE NOT COMMINGLED. | 1
2845222 | IN INSTANCES WHERE AN EMR TRANSMITS MULTIPLE ORDER MESSAGES ON DIFFERENT PATIENTS WITH IDENTICAL ORDER IDS TO THE VALIDATION QUEUE OF CERNER WEBCONNECT, CERNER WEBCONNECT WILL: DISPLAY TO THE USER OF THE VALIDATION QUEUE THE PATIENT ASSOCIATED WITH THE LAST ORDER RECEIVED AND WILL FAIL TO DISPLAY THE PATIENT(S) ASSOCIATED WITH THE FIRST ORDERS RECEIVED HAVING THE SAME ORDER ID WHERE THE SAME TEST IS ORDERED ON ALL PATIENTS HAVING THE DUPLICATE ID ORDER. THIS PREVENTS A USER FROM BEING ABLE TO RELEASE THE ORDER FROM THE VALIDATION QUEUE TO THE RECEIVING SYSTEM FOR ANY ORDERS RECEIVED PRIOR TO THE LAST ORDER HAVING THE DUPLICATE ORDER IDS; DISPLAY TO THE USER OF THE VALIDATION QUEUE THE PATIENT ASSOCIATED WITH THE FIRST ORDER RECEIVED AND WILL FAIL TO DISPLAY THE PATIENT(S) ASSOCIATED WITH SUBSEQUENT ORDERS HAVING THE SAME ORDER ID WHERE DISTINCTLY DIFFERENT TESTS ARE ORDERED ON EACH PATIENT HAVING ORDERS WITH THE SAME ORDER ID. THIS PREVENTS A USER FROM BEING ABLE TO RELEASE THE ORDER FROM THE VALIDATION QUEUE TO THE RECEIVING SYSTEM FOR ANY ORDERS RECEIVED AFTER THE FIRST ORDER HAVING DUPLICATE ORDER IDS. WHEN THE ORDER IS RELEASED FROM THE VALIDATION QUEUE FOR THE ORDER THAT IS DISPLAYED IN THE VALIDATION QUEUE, THE OUTBOUND HL7 MESSAGE TRANSMITTED FROM CERNER WEBCONNECT MAY CONTAIN THE ADT, PID, PV1, ORC, OBR, AND IN MESSAGE SEGMENTS FOR ANOTHER PATIENT WITH THE SAME ORDER ID, NOT THE SELECTED PATIENT RECORD RELEASED BY THE VALIDATION QUEUE USER. CERNER HAS NOT RECEIVED COMMUNICATION ON ANY ADVERSE PATIENT EVENTS AS A RESULT OF THIS ISSUE. CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON (B)(4) 2012 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDED A DESCRIPTION OF THE ISSUE. CERNER CORPORATION WILL PROVIDE FOLLOW-UP REPORTS WHEN THE SOFTWARE MODIFICATION IS AVAILABLE. | 1
2788896 | IT WAS REPORTED THAT THE REMOTE WEBSITE TRANSMISSION DISPLAYS INTERVAL VALUES THAT ARE INCORRECT. THE APPROPRIATE SAFETY VALUE WAS CONFIRMED BY TECHNICAL SUPPORT. NO PATIENT COMPLICATIONS WERE REPORTED AS A RESULT OF THIS EVENT. THE INFORMATION SUBMITTED REFLECTS ALL RELEVANT DATA RECEIVED. IF ADDITIONAL RELEVANT INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 1
3289145 | PATIENT WAS IN CYSTO ROOM FOR A PROCEDURE. PATIENT WAS LYING ON THE TABLE. STAFF BELIEVE PATIENT'S ARMS WERE ON HER CHEST HOWEVER, STERILE DRAPE WAS COVERING PATIENT'S ARMS. AS THE TABLE WAS TRANSFERRING FROM HEAD TO FOOT, THE PATIENT'S THUMB WAS CAUGHT BETWEEN THE FOOT EDGE OF THE TABLE AND A 1/4" THICK RAIL. WHEN STAFF NOTICE THE PATIENT'S THUMB WAS CAUGHT BETWEEN THIS SPACE, TABLE POSITION WAS REVERSED TO FREE THE PATIENT'S THUMB. ONCE PATIENT RECOVERED, A XRAY WAS TAKEN AND "NO FRACTURE" WAS INDICATED. PATIENT SAID THUMB WAS SORE, BUT WAS NOT IN ANY SIGNIFICANT PAIN. A FOLLOW-UP WAS MADE TWO DAYS LATER AND THE PATIENT MADE NO MENTION OF DISCOMFORT OR PAIN FROM THE INCIDENT.======================MANUFACTURER RESPONSE FOR RAD FLUORO UNIT, CYSTO, DIRECT DIGITAL, UROSKOP OMNIA (PER SITE REPORTER).======================MANUFACTURER SENT IN A FIELD SERVICE ENGINEER TO INSPECT. WE SHOULD BE SEEING A REPORT INDICATING ANY FINDING. NOTHING AT THE TIME OF THE INSPECTION INDICATED ANY EQUIPMENT MALFUNCTIONS.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?CYSTOSCOPY; LEFT URETEROSCOPY; LEFT URETERAL STENT CHANGE UNDER FLOUOROSCOPY; AND EXAM UNDER ANESTHESIA.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3065463 | PATIENT WAS IN THE PROCESS OF UNDERGOING A LEFT CARDIAC CATHETERIZATION WHEN THE TOSHIBA IMAGING EQUIPMENT MALFUNCTIONED. THE C-ARM HYPERHANDLE WHICH CONTROLS ANGULATION FOR IMAGING STOPPED FUNCTIONING. THE STAFF ATTEMPTED TO REBOOT THE SYSTEM TO GET THE HYPERHANDLE TO WORK. REBOOTING THE SYSTEM WAS NOT EFFECTIVE. THE PATIENT WAS STABILIZED AND TRANSFERRED TO ANOTHER ROOM TO FINISH THE CASE. THERE WAS AN APPROXIMATE 30-MINUTE DELAY IN THE PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT INCIDENT AND WITHOUT ANY HARM TO THE PATIENT. A TOSHIBA REPAIRMAN CAME TO THE HOSPITAL ON (B)(6) 2013 TO ASSESS AND REPAIR THE SYSTEM. IN HIS REPORT HE FOUND THAT THE BLUE POSITIONING KNOB ON THE HYPERHANDLE WAS MECHANICALLY JAMMED. HE REPLACED THE HANDLE, TESTED IT FOR ACCURACY AND CALIBRATED THE SYSTEM AND FOUND ALL IN WORKING ORDER.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?LEFT CARDIAC CATHETERIZATION.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #2IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3185675 | IT WAS REPORTED TO HEALTHCARE-ID (HID) BY (B)(4) THAT ON 10 OCCASIONS SINCE (B)(6) 2012, FOR DONORS USING THE INTERNET COMPUTER ASSISTED SELF INTERVIEW (ICASI) OF DONOR-ID (B)(4), IT APPEARS THAT THE RESPONSES TO THE FIRST ONE OR TWO QUESTIONS HAVE BEEN OMITTED AND THE RESPONSES HAVE SHIFTED TO BE APPLIED TO THE INCORRECT QUESTIONS. ON AT LEAST ONE OCCASION A DONOR THAT SHOULD HAVE BEEN DEFERRED WAS COLLECTED. THE OCCURRENCES WERE RESTRICTED TO A SINGLE DONOR CENTER. THE METHOD USED BY (B)(4) FOR ICASI PUTS THREE BARCODES ON THE DONOR RECEIPT THAT IS PRINTED AND BROUGHT BY THE DOCTOR TO THE COLLECTION SITE. BARCODE SCANNERS AT (B)(4) ARE PROGRAMMED TO INSERT AND AUTOMATIC "ENTER" AT THE END OF THE BARCODE READ SO THAT STAFF USERS DO NOT HAVE TO CLICK THE ENTER KEY ON THE SCREEN AFTER THE SCAN. INFO RECEIVED FROM (B)(4) WAS THAT THESE DONORS ALL CLAIMED TO HAVE COMPLETED THE ENTIRE QUESTIONNAIRE DURING THE ICASI PROCESS. STAFF USERS HAVE ROUTINELY BEEN OBSERVED SCANNING THE THREE BARCODES IN RAPID SUCCESSION. ALTHOUGH THIS IS DUE TO USER ERROR, WE HAVE INITIATED INVESTIGATION INTO A CHECK AGAINST THE USER MAKING THIS ERROR TO BE IMPLEMENTED IN A FUTURE VERSION OF THE DONOR-ID WEB SOFTWARE. CONCLUSION: WHEN THE SOFTWARE IS USED AS INTENDED, IT WORKS PROPERLY. IN EVERY INSTANCE WHERE THE USER IS KNOWN TO WAIT FOR THE PROMPT TO SCAN THE BARCODE THE ENTIRE BARCODE IS REGISTERED IN THE DATABASE AND THERE ARE NO MISMATCHED RESPONSES. BLOOD CENTER USERS MUST FOLLOW THE ON-SCREEN PROMPTS FOR TIMING OF BARCODE SCANS FROM THE ICASI RECEIPT. IN THOSE CASES WHEN THE BARCODE IS SCANNED TOO SOON INCOMPLETE DATA MAY BE RECEIVED AND THERE ARE AMPLE INDICATIONS THAT SOMETHING IS AMISS WITH THE RESPONSE RECORDED. AS PREVIOUSLY STATED IN THE CONTINUED INVESTIGATION SECTION, WHEN THE RESPONSES ARE SHIFTED IN ORDER, THE USER MUST TAKE ACTION ON FOUR SEPARATE LEVELS FOR THE DONOR'S RESPONSES. ANY ONE OF THESE SHOULD BE A CAUSE FOR THE USER TO QUESTION WHAT HAS HAPPENED AND GIVE THOUGHT TO REJECTING THE ENTIRE SET OF RESPONSES. FOR THE CASES SENT TO HID FOR INVESTIGATION THE USER WOULD HAVE HAD ALL FOUR OF THESE CIRCUMSTANCES TO DEAL WITH. ADDITIONALLY, THE DISCREPANCIES ARE ALSO NOT BEING CAUGHT DURING REVIEW WHICH CAN OCCUR ON-SCREEN OR FROM THE PRINTED DONOR RECORD. ENVIRONMENTAL ISSUES ALSO CONTRIBUTED TO THE OCCURRENCE OF THIS ISSUE. WHEN THE NETWORK OR SERVER RUN SLOWLY, THE PROCESSING OF THE BARCODES TAKES LONGER WHICH INCREASES THE CHANCE THAT THE USER WILL ATTEMPT TO SCAN THE BARCODE TOO SOON (PRIOR TO RECEIVING THE PROMPT THAT THE NEXT SCAN CAN BE PERFORMED). THE BLOOD CENTER CUSTOMER HAS THE RESPONSIBILITY TO PROVIDE THE ENVIRONMENT IN WHICH PROCESSING TAKES PLACE IN A TIMELY MANNER. | 1
3424394 | WHILE IN USE, DURING A ROBOTIC COLON SURGERY PROCEDURE, THE BIPOLAR FENESTRATED FORCEPS INSTRUMENT BEGAN TO MALFUNCTION. IT WAS NOTED THAT THERE WERE BROKEN WIRES AT THE TIP OF THE INSTRUMENT. | 0
3439843 | WHILE DOING THE PROCEDURE, THE PHYSICIAN NOTICED SMALL WIRES PROTRUDING FROM END OF INSTRUMENT, WHICH WERE REMOVED FROM THE PATIENT. THERE WAS NO HARM TO THE PATIENT. | 0
3166177 | THE SYNGO LAB DATA MANAGER SOFTWARE APPLICATION DOES NOT TRANSFER TEST RESULTS FROM THE ADVIA CENTAUR INSTRUMENT TO THE LABORATORY INFORMATION SYSTEM (LIS). THE CUSTOMER MUST MANUALLY SEARCH FOR THE RESULTS AND RELEASE THEM TO THE LIS. THERE ARE NO REPORTS OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCES DUE TO RESULTS NOT BEING TRANSFERRED FROM THE ADVIA CENTAUR INSTRUMENT TO THE LIS. AN URGENT MEDICAL DEVICE CORRECTION (UMDC) ENTITLED "SYNGO LAB DATA MANAGER VERSION VA11B AND VA12A SYSTEMS: LIMITATIONS AND SOFTWARE ANOMALIES" WAS SENT TO CUSTOMERS IN MAY 2013 TO NOTIFY CUSTOMERS THAT RESULTS MAY NOT BE RECEIVED FROM THE INSTRUMENT OR TRANSFERRED TO THE LABORATORY INFORMATION SYSTEM. THE UMDC PROVIDES CUSTOMERS WITH ACTIONS TO IMPLEMENT WHILE A NEW SOFTWARE VERSION IS DEVELOPED. | 1
3299532 | PATIENT WAS TO RECEIVE FLOURO IMAGE FROM C-ARM TO FURTHER IDENTIFY SURGICAL SITE. C-ARM WOULD NOT FLURO EVEN AFTER REBOOT. C-ARM WAS SWITCHED POTENTIALLY CAUSING HARM IN DELAY OF SURGERY. NO PATIENT HARM.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?NOT KNOWN.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2992376 | A VIDEO CAPSULE STUDY WAS DONE AND DOWNLOADED. THE PHYSICIAN REVIEWED THE STUDY TWO DAYS LATER AND NOTED ONLY 15 MINUTES OF THE 8 HOUR STUDY WAS RECORDED. PLEASE NOTE THAT THIS PROCEDURE WAS DONE IN THE MAIN OPERATING ROOM. THE PATIENT HAS AN AUTOMATED IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (AICD) AND ON A MONITORED BED, ALL OF WHICH MAY INTERFERE WITH DATA TRANSMISSION AND DEACTIVATE THE CAPSULE. THE PHYSICIAN IS AWARE OF THESE RISKS. ONE DAY AFTER THE ORIGINAL VIDEO CAPSULE STUDY, BOTH THE DATA RECEIVER AND SENSOR BELT WERE USED SUCCESSFULLY TO DO ANOTHER CAPSULE STUDY. OF COURSE A NEW VIDEO CAPSULE WAS USED. THIS IS THE ONLY SINGLE USE ITEM OF THE CAPSULE STUDY. THE MANUFACTURER WAS CONTACTED AND A DE-IDENTIFIED DVR COPY OF THE STUDY WAS SENT FOR EVALUATION. | 1
3231124 | THE PATIENT WAS UNDERGOING AN MRI OF THE C-SPINE WITH AND WITHOUT CONTRAST. NEAR THE END OF THE EXAM THE THE MRI UNIT GAVE AN "X,Y,Z GRADIENT ERROR". THE UNIT WAS SHUT DOWN AND RESTARTED WITHOUT SUCCESS IN CLEARING THE ERROR. NO PATIENT HARM. THE VENDOR WAS NOTIFIED. GRADIENT COIL HAD TO BE REPLACED.======================MANUFACTURER RESPONSE FOR GE MRI HD XT 1.5 TESLA, GA MRI HD XT 1.5 TESLA (PER SITE REPORTER).======================FIELD REPRESENTATIVE ON SITE ON THREE SEPARATE DAYS WITHIN A WEEK TIME FRAME. SERVICE REPAIR WITH GRADIENT COIL REPLACED AND MRI PM COMPLETED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?MRI OF C-SPINE WITH AND WITHOUT CONTRAST.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2909289 | PATIENT HAD AN MRI OF THE BREAST. AFTER THE PROCEDURE THE PATIENT STATED THAT THERE WAS BILATERAL TINGLING OF THUMBS DURING THE PROCEDURE. THE SCAN LASTED APPROXIMATELY 45 MINUTES. THE PATIENT STATED THE SQUEEZE BALL WAS SQUEEZED TWICE TO ALERT THE STAFF REGARDING FEELING SOMETHING UNUSUAL, BUT THE STAFF NEVER HEARD AN ALARM. THE BALL WAS THEN CHECKED TO VERIFY THAT IT WAS WORKING PROPERLY AND IT WAS. THE PATIENT'S THUMBS WERE EXAMINED IMMEDIATELY AFTER THE MRI PROCEDURE AND APPEARED NORMAL. THE FOLLOWING DAY, THE PATIENT CALLED TO SAY THAT THERE NOW WAS A BLISTER ON THE LEFT THUMB. NO BLISTER ON THE RIGHT THUMB. THE PATIENT WAS SEEN IN THE ED AND A 1 CM RUPTURED BLISTER WAS NOTED ON THE LEFT THUMB. NO ADDITIONAL MEDICAL TREATMENT REQUIRED AND NO FURTHER MEDICAL TREATMENT WAS DOCUMENTED IN THE EMERGENCY DEPARTMENT RECORD. PATIENT DID HAVE WEDDING RING ON THE WEDDING RING FINGER (LEFT HAND) DURING PROCEDURE, BUT THIS IS NOT CONTRA-INDICATED.THE STAFF HAVE NEVER EXPERIENCED THIS TYPE OF PROBLEM BEFORE AND THEIR EXPERIENCE IS THAT A PATIENT WITH AN MRI BURN SHOWS EVIDENCE OF THE BURN IMMEDIATELY. IN THIS CASE, THE BLISTER APPEARED THE NEXT DAY AND WAS ONLY ON ONE THUMB.======================MANUFACTURER RESPONSE FOR MRI, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================THE MANUFACTURER CAME ON-SITE AND PERFORMED THE FOLLOWING SYSTEMS CHECKS:GRADIENT COOLANT CIRCULATOR, GRADIENT CHILLER, FLUID LEVELS OF THE HEAT EXCHANGER, WATER TEMPERATURE, PT AIR FLOW/ BODY COIL AIR BLOWER SYSTEM, CHECKED SYSTEM ERROR LOGS, BREAST COIL INSPECTED, NO VISUAL IRREGULARITIES OR ARTIFACTS. RAN FULL SPT HEAD/BODY, CHECKED RF OUTPUT PEAK POWER, RAN FORWARD/REVERSE QUADRATURE CHECK, UPM FUNCTIONAL CHECK, CHECKED PT ALERT SYSTEM, AND CHECKED PT INTERCOM. ALL SYSTEMS PASSED AND NO ERRORS FOUND. CUSTOM ER REPLACED RF AMP BECAUSE WE COULD NOT ACHIEVE RF PEAK BODY POWER. SYSTEM PASSED ALL REQUIRED TESTS AND WAS TURNED OVER TO THE CUSTOMER FOR CLINICAL USE. | 0
2938981 | THE ROBOTIC GRASPER WOULD NOT LET GO OF THE TISSUE. WE USED THE EMERGENCY ALLEN WRENCH TO RELEASE IT, BUT IT STILL WOULD NOT RELEASE. WE HAD TO DO A TOTAL SYSTEM SHUTDOWN TO GET THE GRASPER TO OPEN ITS JAWS. THERE WAS NO PATIENT INJURY.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTICALLY ASSISTED COLECTOMY AND RECTOPEXY. | 0
3488814 | VESSEL SEALER WOULD NOT CUT. | 0
3500947 | WHEN USING A FENESTRATED BIPOLAR FOR THE DAVINCI ROBOT ARM #2, STAFF NOTICED AN EXPOSED WIRE AND THE INSTRUMENT NO LONGER RESPONDED "AS NORMAL". STAFF WAS UNABLE TO STRAIGHTEN OUT THE ARTICULATION ON THE FENESTRATED BIPOLAR, SO WHEN IT WAS "PULLED OUT" OF THE HOLDING A WIRE SNAPPED OFF AND DROPPED INTO THE PATIENT'S ABDOMINAL CAVITY. WIRE WAS IMMEDIATELY SEEN AND RETRIEVED WITH A LAPAROSCOPIC INSTRUMENT. | 0
2894066 | MEGA SUTURECUT NEEDLE DRIVER BROKE WHEN MD WAS USING ON IT ON PATIENT. NO PIECES LEFT IN PATIENT. PIANO-TYPE WIRE BROKE INSIDE THE SHAFT OF THE INSTRUMENT MAKING IT INOPERABLE. | 0
3496512 | DURING THE PROCEDURE, THE SURGEON REPORTED THAT THE VESSEL SEALING DEVICE WAS NOT SEALING PROPERLY. THE DEVICE WAS REMOVED FROM THE FIELD, AND A NEW DEVICE WAS UTILIZED. IT IS UNKNOWN IF THE NEW DEVICE WAS OF THE SAME OR DIFFERENT LOT. | 0
3535280 | EXPOSED WIRE ON THE TIP OF THE FENESTRATED BIPOLAR OPERATED BY THE DOCTOR NICKED THE PATIENT. THE CONDITION APPEARED TO BE REPAIRED UPON THE END OF SURGERY. THE PATIENT WAS TRANSFERRED TO THE POST ANESTHESIA UNIT IN STABLE CONDITION. | 0
3283765 | BROKEN BIPOLAR/SPRING. ALL PARTS WERE RETRIEVED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED MYOMECTOMY/MORCELLATION.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3490871 | THE DAVINCI ROBOTIC INSTRUMENTS MALFUNCTIONED WHILE IN USE. THE BIPOLAR MARYLAND FORCEP WIRES LOCATED AT THE WRIST WERE CUT. THE PROGRASP INSTRUMENT WAS NOT OPENING PROPERLY. | 0
3490869 | DURING ROBOTIC MYOMECTOMY, WHILE USING THE DAVINCI ENDOWRIST INSTRUMENT 8MM TENACULUM FORCEPS, THE CABLE ON THE FORCEP SNAPPED. THE DOCTOR STATES THERE WAS NO HARM TO THE PATIENT AND NO PIECE OF THE CABLE WAS LEFT BEHIND. IT IS VISUALIZED ATTACHED TO THE INSTRUMENT. | 0
3503167 | DURING ROBOTIC MYOMECTOMY WHILE USING THE DAVINCI ENDOWRIST INSTRUMENT 8MM TENACULUM FORCEPS, THE CABLE ON THE FORCEPS SNAPPED. MD STATES NO HARM TO THE PATIENT NO PIECE OF CABLE WAS LEFT BEHIND. IT IS VISUALIZED ATTACHED TO THE INSTRUMENT. | 0
3321511 | PROGRASP DIDN'T OPEN OR ARTICULATE AS THEY SHOULD. REPLACED TIMES 2, AND THE THIRD PROGRASP WORKED PROPERLY. THERE WAS NO HARM TO THE PATIENT.HARMONIC HANDPIECE WOULDN'T SCREW INTO INSTRUMENT. THERE WAS NO HARM TO THE PATIENT. A SECOND ONE WAS GIVEN TO THE TECH AND USED. FIRST NOT CHARGED TO PATIENT AND SAVED FOR THE REP TO INVESTIGATE THE PROBLEM. | 0
3494527 | DURING A ROBOTIC LYMPH NODE DISSECTION, THE DAVINCI PROGRASP FORCEPS WERE NOT OPERATING CORRECTLY. THE INSTRUMENT WAS REMOVED AND REPLACED BY A NEW BACK UP INSTRUMENT. | 0
3291298 | DURING THE PROCEDURE THE VESSEL SEALER BLADE WOULD NOT RETRACT INTO THE INSTRUMENT, CAUSING A REVOCABLE FAULT. THE INSTRUMENT WAS REMOVED AND A NEW ONE WAS OPENED. SINCE THE BLADE WOULD NOT RETRACT, IT LEFT THE BLADE EXPOSED. NO HARM TO PATIENT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?OPERATIVE PROCEDURE:1. DA VINCI ROBOTIC LAPAROSCOPIC HYSTERECTOMY.2. BILATERAL SALPINGECTOMY.3. MORCELLATION OF UTERUS WITH FIBROIDS.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3294441 | THE WHITE TIP PORTION OF THE SHEARS BECAME LOOSE. DEVICE WAS REPLACED WITH A NEW SHEAR INSERT. THE DEFECTIVE ITEM WAS REMOVED OFF THE SURGICAL FIELD.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?HIATAL HERNIA.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3299572 | THE ROBOTIC INSTRUMENT, THE ENDOWRIST ONE DA VINCI SI VESSEL SEALER WAS NOT RECOGNIZD AS A MEDICAL DEVICE AND THEREFORE UNABLE TO USE DURING PROCEDURE.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?1. DA VINCI ROBOTIC LAPAROSCOPIC HYSTERECTOMY WITH BILATERALSALPINGO-OOPHORECTOMY.2. ABDOMINOPELVIC WASHINGS.3. BILATERAL PELVIC LYMPHADENECTOMY.4. PARAAORTIC LYMPHADENECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3164067 | BEFORE USING THE MARYLAND BIPOLAR FORCEPS (REF 420172) FOR THE ROBOTIC CHOLECYSTECTOMY WITH THE DAVINCI, THE SCRUB TECH NOTICED THAT THE STRINGS WERE BROKEN. THERE ARE USES LEFT ON THE INSTRUMENT TO RECEIVE REIMBURSEMENT.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3164183 | DURING ROBOTIC CASE, THE MEGA SUTURECUT NEEDLE DRIVER WAS NOT FUNCTIONING PROPERLY. THE INSTRUMENT WAS IMMEDIATELY REMOVED FROM THE SURGICAL FIELD AND REPLACED.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3172013 | WHILE IN USE THE BLADE WOULD NOT RETRACT. THE DEVICE WAS REMOVED FROM PT. NO PATIENT HARM.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3439656 | WHILE USING THE DA VINCI MEGA SUTURE NEEDLE DRIVER, THE CABLE BROKE AT THE WRIST JOINT. INSTRUMENT HAD 2 LIVES LEFT. | 0
3229280 | ROBOTIC CARDIERE FORCEP REPORTED BY SURGEON AS NOT PROPERLY CLOSING. INSTRUMENT PASSED OFF FIELD AND NEW FORCEPS OPENED. | 0
3424406 | WHILE USING THE FENESTRATED BIPOLAR FORCEPS DURING SURGERY, THE SURGEON NOTED THAT THE TIPS OF THE FORCEPS WERE MISALIGNED. | 0
3113453 | CADIERE FORCEPS USED IN ROBOTIC SURGERY WAS NOT WORKING PROPERLY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3113346 | HARMONIC TIP BROKE OFF, BUT WAS RETRIEVED.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3113448 | DURING LAPAROSCOPIC ROBOTIC ASSISTED TAH, SURGEON NOTED THAT THE ENDOWRIST INSTRUMENT, MARYLAND BIPOLAR FORCEPS 8MM HAD SPARKED WHILE IN USE AND THE TIPS WERE NOT CLOSING PROPERLY. HE STATED THAT THERE WAS NO HARM TO THE PATIENT. THE INSTRUMENT WAS REMOVED FROM THE FIELD AND REPLACED WITH ANOTHER BIPOLAR.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3176943 | ROBOTIC TOTAL ABDOMINAL HYSTERECTOMY IN PROGRESS WHEN SURGICAL TEAM NOTICED A SMALL BLACK OBJECT IN ABDOMINAL CAVITY, VISIBLE ON MONITOR. OBJECT WAS REMOVED AND INSPECTED; BLACK IN COLOR AND HARD TO THE TOUCH. TECH REMOVED INSTRUMENTS FROM TROCARS. 5MM MONOPOLAR CAUTERY DEVICE REMOVED FROM 8MM DAVINCI TROCAR AND EXAMINED. OUTER, BLACK COVERING OF DEVICE HAD AN APPROXIMATELY 1.5 MM AREA ON SHEATH THAT WAS FRAYED AND DAMAGED. DEVICE REMOVED FROM STERILE FIELD. ABDOMEN WAS IRRIGATED AND SUCTIONED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TOTAL ABDOMINAL HYSTERECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3268250 | DA VINCI PROGRASP INSTRUMENT STOPPED WORKING WHILE MD WAS USING INSTRUMENT. JAWS WOULD NOT OPEN. OR TECH STARTED TO REMOVE THE INSTRUMENT FROM THE STERILE FIELD AND THE SILVER METAL TAB BUTTON BROKE OFF ON THE OUTSIDE OF THE PT. BUTTON RETRIEVED AND INSTRUMENT REMOVED FROM THE FIELD AND REPLACED WITH ANOTHER DA VINCI INSTRUMENT. ORIGINAL INSTRUMENT HAD 4 LIVES LEFT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TOTAL LAPAROSCOPIC HYSTERECTOMY | 0
3171948 | DURING THE ROBOTIC CASE, THE HARMONIC INSERT JAW TIP BROKE OFF. THE PIECE WAS RETRIEVED WITH NO APPARENT INJURY TO THE PATIENT.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3229243 | DURING PROCEDURE, PK DISSECTING FORCEPS CABLE BROKE AFTER ONLY THREE USES. THE DEVICE WAS REMOVED FROM THE FIELD, AND ANOTHER DEVICE WAS UTILIZED. IT IS UNKNOWN IF THE NEW DEVICE WAS OF THE SAME OR DIFFERENT LOT.======================MANUFACTURER RESPONSE FOR DISSECTING FORCEPS, PK DISSECTING FORCEPS (PER SITE REPORTER).======================NONE AT THIS TIME.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?LAPAROSCOPIC BILATERAL SALPINGO-OOPHORECTOMY AND EXTENSIVE ADHESIOLYSIS.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3207155 | INSTRUMENT PLACED INTO ROBOTIC. NOTICED THE WIRES WERE BROKEN, THE DEVICE WAS REMOVED AND REPLACED. ALL PIECES INTACT AND THE EQUIPMENT SAVED. NO HARM TO THE PATIENT.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3417852 | DURING A ROBOTIC ENDOSCOPIC PROCEDURE, THE RF DISSECTING FORCEPS FAILED TO GRASP TISSUE CORRECTLY. THE FORCEPS WOULD NOT CLOSE COMPLETELY. THE HANDPIECE WAS REMOVED AND REPLACED, AND THE PROCEDURE CONTINUED WITHOUT FURTHER ACTION. NO VISILE DAMAGE WAS NOTED AT THE TIME OF THE EVENT. ====================== MANUFACTURER RESPONSE FOR DISSECTING FORCEPS, ROBOTIC, PK DISSECTING FORCEPS (PER SITE REPORTER).====================== MANUFACTURER PROVIDED RMA (PACKAGING) FOR RETURN AND EVALUATION OF DEVICE. | 0
3417849 | DURING A ROBOTIC ENDOSCOPIC PROCEDURE, WHEN HANDPIECE WAS INSTALLED AND INSERTED THE TIP WOULD NOT MOVE OR ARTICULATE. THE HANDPIECE REPLACED AND SURGERY COMPLETED WITHOUT INCIDENT. NO VISIBLE DAMAGE WAS NOTED AT THE TIME OF THE EVENT.====================== MANUFACTURER RESPONSE FOR MEGA SUTURE NEEDLE DRIVER, ROBOTIC, MEGA SUTURE NEEDLE DRIVER (PER SITE REPORTER).====================== MFG PROVIDED RGA (PACKAGING) TO RETURN AND EVALUATE HANDPIECE. | 0
3417848 | DURING AN ENDOSCOPIC ROBOTIC PROCEDURE THE SURGEON STATED THAT NEEDLE DRIVER HANDPIECE WOULD NOT MOVE OR ARTICULATE. THE HANDPIECE WAS REPLACED AND PROCEDURE CONTINUED WITHOUT INCIDENT. NO DAMAGE WAS NOTED ON THE DEVICE AT THE TIME OF THE EVENT.====================== MANUFACTURER RESPONSE FOR MEGA NEEDLE DRIVER, MEGA NEEDLE DRIVER (PER SITE REPORTER).====================== MFG PROVIDED RGA (PACKAGING) FOR RETURN AND EVALUATION OF DEVICE. | 0
3417641 | WHILE IN USE, THE CABLE THAT CONTROLS THE BITING ACTION OF THE INSTRUMENT TIP SNAPPED. NO PIECES CAME OFF, AND THE CABLE DID NOT COME INTO CONTACT WITH THE PATIENT. THE TIP WAS NOT BEING USED FOR BITING ACTION AT THE TIME THE CABLE BROKE. THE DEVICE WAS BEING OPERATED BY THE SURGEON PRIOR TO CONTACT WITH THE TISSUE. | 0
2913010 | TIP OF THE GRASPING RETRACTOR FOR DAVINCI ROBOT WAS FLOPPY AND SURGEON WAS UNABLE TO CONTROL INSTRUMENT WITH ROBOTIC ARM.======================MANUFACTURER RESPONSE FOR GRASPING RETRACTOR, DAVINCI (PER SITE REPORTER).======================MANUFACTURER TO EVALUATE DEVICE. | 0
2913025 | PRIOR TO A CASE A GRASPING RETRACTOR WAS FOUND TO HAVE A BROKEN CABLE. | 0
2913037 | PRIOR TO A PROCEDURE, DAVINCI FENESTRATED BIPOLAR WAS FOUND TO HAVE BROKEN CABLES IN THE TIP.======================MANUFACTURER RESPONSE FOR FENESTRATED BIPOLAR, DAVINCI (PER SITE REPORTER).======================MFR TO EVALUATE DAMAGE TO INSTRUMENT. | 0
2913052 | PRIOR TO A PROCEDURE A CABLE ON THE GRASPING RETRACTOR WAS FOUND TO BE BROKEN. | 0
2973586 | PATIENT WAS HAVING A ROBOTIC PROCTOSIGMOIDECTOMY WITH TAKE DOWN OF COLOVAGINAL FISTULA. DURING THE PROCEDURE THE SURGICAL PA NOTICED THAT THE #1 ROBOTIC ARM WAS NOT ACTING RIGHT. SHE STATED THAT WHILE VISUALLY EXAMINING THE AREA AROUND THE ENTRANCE SITE TO SEE IF THERE WAS A PROBLEM, THE ROBOTIC ARM ALL OF A SUDDEN JERKED FORWARD MOVING FURTHER INTO THE PATIENT AND THEN JERKED BACK. IT WAS DESCRIBES AS A STABBING MOTION. THIS CAUSED AN INJURY TO THE AORTIC ARTERY. THE ABDOMEN WAS IMMEDIATELY OPENED AND VASCULAR SURGEON CALLED TO THE ROOM TO ASSIST. AFTER THE REPAIR OF THE VESSEL THE SURGERY CONTINUED. PATIENT WAS TRANSFERRED POST OP TO ICU. AS DOCUMENTED. AFTER THE OCCURRENCE I TALKED WITH THE STAFF IN THE ROOM AND THEN CALLED THE COMPANY SPOKE WITH SOMEONE WITH TECHNICAL SERVICE, OUR SERVICE ENGINEER, AND OUR SALES REPRESENTATIVE. ROBOT SEQUESTERED UNTIL FOLLOW UP WITH EVERYONE IS COMPLETED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC PROCTOSIGMOIDECTOMY WITH TAKE DOWN OF COLOVAGINAL FISTULA. | 0
3148690 | DURING THE OR PROCEDURE, THE DAVINCI SCISSOR TIP WAS BURNED DURING USE. BURN WENT THROUGH SCISSOR TIP TO INSTRUMENT AND BURNED A COATING ON END OF SCISSOR. THE PACKAGING WAS NOT RETAINED THUS CANNOT SUBMIT ALL THE NUMBERS ON PACKAGE, JUST THE DEVICE ITSELF.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?THE PATIENT WAS HAVING ROBOTIC SURGERY FOR REMOVAL OF UTERINE FIBROIDS. THERE WAS NO HARM TO THE PATIENT.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3543803 | THE PLASTIC INNER CANNULA CAME OUT OF THE DA VINCI MONOPLOAR SCISSOR DEVICE. THE DEVICE WAS BEING IRRIGATED WHEN THE BLOCKAGE WAS NOTED, AND THE PLASTIC SLEEVE FELL OUT OF THE DEVICE. THIS CAUSED A SLIGHT DELAY IN THE CASE. THERE WAS NO HARM TO THE PATIENT. THE DEVICE WAS NOT USED ON THE PATIENT. | 0
3049359 | WHILE TAKING THE DA VINCI S PERMANENT CAUTERY HOOK OUT OF THE TROCAR DURING A ROBOTIC CHOLECYSTECTOMY, THE HOOK WOULD NOT COME OUT OF THE TROCAR. THE HOOK CAME OUT WITH THE TROCAR REMOVAL WITH NO HARM TO THE PATIENT. THE HOOK HAS A LITTLE NUB ON THE SIDE OF THE INSTRUMENT THAT DIDN'T ALLOW IT TO COME OUT. THE HOSPITAL NEEDS REIMBURSEMENT FROM COMPANY FOR UNUSED USAGES.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2913001 | SCISSOR IS SEPARATING NEAR THE TIP.======================MANUFACTURER RESPONSE FOR MONOPOLAR CURVED SCISSOR, DAVINCI (PER SITE REPORTER).======================MANUFACTURER TO EVALUATE DAMAGE AND REPORT BACK. | 0
3034433 | DA VINCI FENESTRATED BIPOLAR CAUTERY STOPPED WORKING DURING THE PROCEDURE. DESPITE HAVING 8 USES LEFT. CAUTERY CORD AND INSTRUMENT WERE CHANGED AND CAUTERIZATION WORKED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TOTAL HYSTERECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2973634 | DURING THE ROBOTIC HYSTERECTOMY AFTER THE PORT WAS INSERTED AND THE GAS WAS INSTILLED INTO THE ABDOMEN, THE CAMERA FAILED TO ALLOW 3-D IMAGING. THE STAFF WAS UNABLE TO FIX THE ISSUE. THE ROBOTIC APPROACH WAS ABORTED. THE SURGEON ELECTED TO PERFORM AN OPEN ABDOMINAL HYSTERECTOMY. | 0
3349207 | WHILE BEING IMAGED BY A GE INFINIA II HAWKEYE 4, THE PATIENT'S FOOT GOT WEDGED BETWEEN A ROLLER AND THE OUTER SHELL OF THE CAMERA. THERE WERE NO FRACTURES OR OTHER INJURIES TO THE PATIENT. | 0
3498929 | THE DEVICE IS REPEATEDLY FREEZING AND NOT TRANSMITTING IMAGES PROPERLY RESULTING IN REPEAT IMAGES NEEDED FROM PATIENTS AND UNNECESSARY EXPOSURE TO ADDITIONAL X RAYS. | 1
3476760 | GE PREVENTIVE MAINTENANCE CHECKS DISCOVERED THERE MAY HAVE BEEN INSTALLATION ERROR BY OMITTING 4 SCREWS IN THE MOUNTING KIT. OUR BIOMEDICAL TECHNICIAN CHECKED AND THE SCREWS WERE INSTALLED. THE MANUFACTURER PROMPTED THE VISUAL EXAMINATION. NO PATIENTS OR STAFF WERE HARMED. | 0
3476788 | DURING GE PREVENTIVE MAINTENANCE, IT WAS DISCOVERED THAT 4 MOUNTING SCREWS MAY BE MISSING GIVING THE POTENTIAL OF THE VIDEO MONITOR TO FALL. OUR BIOMEDICAL TECHNICIAN CHECKED THE MONITORS TO ENSURE THE SCREWS WERE THERE, AND THEY WERE THERE. THE MONITOR WAS INSTALLED CORRECTLY. | 0
3476786 | GE PREVENTIVE MAINTENANCE CHECKS DISCOVERED IMPROPER INSTALLATION MAY HAVE OCCURRED AND 4 SCREWS MAY HAVE BEEN OMITTED GIVING THE POTENTIAL FOR THE MONITOR TO FALL. OUR BIOMEDICAL TECHNICIAN CHECKED THE MOUNTING AND THE SCREWS WERE THERE, SO THERE WAS NOT THE POTENTIAL TO FALL. NO PATIENTS OR STAFF WERE HARMED. | 0
2901133 | IT WAS REPORTED THAT DURING AN EXAM A PATIENT SUSTAINED A LACERATION FROM THE BENT EDGE OF THE MYLAR WINDOW FRAME AND REQUIRED STITCHES. PATIENT DATA NOT PROVIDED. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. | 0
3439645 | WE WERE DOING A TRIPHASIC LIVER AND THE SCANNER DURING THE CONTRAST PORTION OF THE TEST HAD AN ERROR AND SAID SCANNER WAS DIS-ENABLED. PATIENT HAD FULL CONTRAST INJECTION AND HAD TO RETURN FOR TEST AGAIN. I CONTACTED CLINICAL ENGINEERING TO CHECK CT SCANNER ERROR LOG; SCANNER IS NOW WORKING FINE. | 0
2943840 | DURING DAVINCI HYSTERECTOMY NEAR START OF PROCEDURE MARYLAND BIPOLAR FORCEP IN ARM 2 WAS NOT MOVING CORRECTLY. PHYSICIAN COULD NOT GET IT TO MOVE WHERE HE WANTED. ASSISTANT TRIED REMOVING INSTRUMENT AND RESETTING THE CASSETTE. THAT DID NOT WORK. INSTRUMENT REMOVED AND A SMALL PIECE OF WIRE NEAR TIP OF INSTRUMENT WAS NOTED TO BE EXPOSED. INSTRUMENT HAD 3 LIVES LEFT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?HYSTERECTOMY. | 0
3254025 | WHEN SMALL CLIP APPLIER ENDOWRIST WAS BEING LOADED WITH THE SMALL CLIP, ONE OF THE TIPS ON THE ENDOWRIST BROKE. INSTRUMENT WAS NOT INSIDE OF THE PATIENT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC THYMECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2989481 | 8MM CARDIERE FORCEPS BROKE DURING USE. WIRE AT PULLEY SYSTEM FRAYED AND EXPOSED.MANUFACTURER RESPONSE FOR CARDIERE FORCEPS, INTUITIVE ENDOWRIST (PER SITE REPORTER).COMPLAINT WAS FILED WITH COMPANY RMA. DEVICE TO BE RETURNED TO COMPANY ONCE COMPLAINT IS VERIFIED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED THREE FIELD ESPHAGECTOMY. | 0
2989501 | DURING OPERATION MEGA SUTURECUT NEEDLE DRIVER WIRE BROKE AT TIP OF INSTRUMENT. NEEDLE DRIVER REMOVED FROM STERILE FIELD.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED HYSTERECTOMY, BSO, SACROCOLPOPEXY AND TVT. | 0
2989506 | DURING A ROBOTIC ASSISTED HYSTERECTOMY IT WAS NOTED BY THE SURGEON THAT THE WIRE WAS FRAYED AT THE TIP OF THE INSTRUMENT AT THE PULLEY SYSTEM. THE MEGA SUTURECUT NEEDLE DRIVER WAS REMOVED FROM THE PATIENT AND USE.MANUFACTURER RESPONSE FOR MEGA SUTURECUT NEEDLE DRIVER, INTUITIVE ENDOWRIST (PER SITE REPORTER).COMPANY RECEIVED COMPLAINT RMA AND DEVICE WAS RETURNED TO COMPANY FOR THEIR REVIEW.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED TOTAL LAPAROSCOPIC HYSTERECTOMY AND SACROCOLPOPEXY. | 0
2989528 | DURING A ROBOTIC ASSISTED SACROCOLPOPEXY THE METAL CORD/WIRE AT THE TIP OF THE INSTRUMENT AT THE PULLEY SYSTEM BROKE. THE MEGA SUTURE CUT NEEDLE DRIVER WAS IMMEDIATELY REMOVED FROM SERVICE AND REPLACED.MANUFACTURER RESPONSE FOR MEGA SUTURE CUT NEEDLE DRIVER, INTUITIVE SURGICAL ENDOWRIST (PER SITE REPORTER).COMPLAINT WAS FILED RMA AND DEVICE WAS RETURNED FOR THEIR REVIEW.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED SACROCOLPOPEXY. | 0
2989510 | DURING ROBOTIC ASSISTED HYSTERECTOMY IT WAS NOTED THAT THE PULLEY CABLE AT THE TIP OF THE PK ENDOWRIST DISSECTOR WAS FRAYED. THE DISSECTOR WAS IMMEDIATELY REMOVED FROM THE PATIENT AND REPLACED.MANUFACTURER RESPONSE FOR PK ENDOWRIST DISSECTING FORCEPS, PK ENDOWRIST DISSECTING FORCEPS (PER SITE REPORTER).COMPLAINT FILED RMA AND DEVICE RETURNED TO THE COMPANY FOR THEIR REVIEW.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED HYSTERECTOMY AND LSO. | 0
2982170 | DURING A ROBOTIC HYSTERECTOMY IT WAS NOTED THAT A WIRE THE TIP OF THE INSTRUMENT AT THE PULLEY SYSTEM WAS BROKEN. SURGEON DID NOT NOTE ANY UNUSUAL MOVEMENTS OR ISSUES HANDLING THE INSTRUMENT. THERE WERE 4 USES LEFT ON THIS INSTRUMENT.MANUFACTURER RESPONSE FOR PK DISSECTING FORCEP, ENDOWRIST PK DISSECTING FORCEP (PER SITE REPORTER).COMPLAINT WAS REGISTERED RMA AND DEVICE WAS RETURNED FOR THEIR REVIEW.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED HYSTERECTOMY. | 0
3095497 | NOTICED WIRE WAS COMING THROUGH NEEDLE DRIVER. NO HARM TO PATIENT. NEEDLE DRIVER BROKEN AND WIRE EXPOSED DURING USE. REPLACED WITH NEW NEEDLE DRIVER. NO PT HARM.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED MYOMECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3501150 | SURGEON USED DISPOSABLE BLADE (ATTACHED TO NON-DISPOSABLE HANDLE) FOR LAPAROSCOPIC REMOVAL OF UTERUS. HE WAS CUTTING UTERUS FOR 15-20 MINUTES WHEN NO TIP CAME OFF HANDLE. BLADE RETRIEVED BY SURGEON AND SEQUESTERED OFF THE TABLE. SCRUB TECH ASSEMBLED BLADE TO HANDLE DURING PROCEDURE AND NOT FORMALLY TRAINED ON HOW TO ASSEMBLE. REP NOTIFIED. IN-SERVICE FOR STAFF SCHEDULED.NOTE - TAILS OF BLADE APPEARED TO BE BENT WHEN BLADE FELL OFF. QUESTIONABLE DEFECTIVE PRODUCT VERSUS POOR ASSEMBLY BY STAFF.======================MANUFACTURER RESPONSE FOR SNAP-FIT PADDLE BLADE, SNAP-FIT PADDLE BLADE (PER SITE REPORTER).======================MANUFACTURER NOTIFIED AND WILL PICK UP DEVICE. | 0
3088940 | DURING ROBOTIC PROCEDURE, SURGEON SAW SPARK WHILE USING A MARYLAND ROBOTIC INSTRUMENT. THE MARYLAND WAS IMMEDIATELY REMOVED. NO PATIENT HARM. SPARKING WAS AT THE INSTRUMENT PIECE BUT NOT THE ROBOTIC EQUIPMENT.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3248803 | DURING A PROCEDURE, THE ROBOT MALFUNCTIONED WITH A NON'-RECOVERABLE" FAULT. INTUITIVE BIOMED 800# CALLED. THE ROBOT HAND WAS CLAMPED ONTO THE PATIENTS BLADDER AND WOULD NOT RELEASE. THEY SUGGESTED TO RETRIEVE A PARTICULAR WRENCH TO OPEN INSTRUMENT JAWS. IT TOOK A FEW MINUTES TO FIND THIS "SPECIAL" STERILE ALLEN WRENCH. WHEN LOCATED, STAFF WERE ABLE TO USE THE TOOL TO RELEASE THE CLAMP. NO PATIENT HARM.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?PROSTATECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3420110 | THE CAUTERY ON THE DAVINCI ROBOT, BOTH MONOPOLAR AND BIPOLAR, STOPPED FUNCTIONING DURING SURGERY WHICH RESULTED IN NO MEANS TO CONTROL BLEEDING IF NEEDED. THERE WAS NO LOSS OF POWER TO THE CAUTERY DEVICES DURING THIS TIME AND OTHER MEASURES WERE TAKEN TO MAINTAIN CAUTERY DURING SURGERY. INCIDENT WAS REPORTED TO DAVINCI COMPANY REPRESENTATIVE AND ROBOTIC EQUIPMENT WAS CHECKED OUT THAT DAY. THE OUTLET AND EXTENSION CORDS WERE CHECKED OUT BY MAINTENANCE. THE PATIENT SUFFERED NO HARM DURING THIS SITUATION. | 0
3420119 | WHILE PERFORMING THE LAPAROSCOPIC ROBOTIC ASSISTED SACROCOLPOPEXY, THE SURGICAL TEAM NOTICED TWO HOLES IN THE BLADDER. IMMEDIATE REPAIR WAS PERFORMED PER PHYSICIAN. | 0
3420104 | PHYSICIAN WAS PERFORMING A ROBOTIC HYSTERECTOMY. SHE REQUESTED TO PREPARE FOR AN OPEN PROCEDURE. THE STAFF OBSERVED ON THE MONITOR THAT THERE WAS SOME BLEEDING. WE BEGAN GATHERING ITEMS TO OPEN. PHYSICIAN REQUESTED A SUTURE AND THE BLEEDING SUBSIDED. SHE PUT A STITCH WHERE THE HEMATOMA WAS TO SEE IF IT WAS GETTING ANY BIGGER. IT APPEARED NOT TO. WE DID NOT GO OPEN. SHE THEN FINISHED THE PROCEDURE AND CHECKED AGAIN FOR BLEEDING AND CHECKED THE HEMATOMA AREA, IT SEEMED THE SAME AND THEN SHE CLOSED. PATIENT WAS TRANSPORTED TO RECOVERY. HOWEVER IT WAS REPORTED TO US LAST WEEK THAT THIS PATIENT HAD OTHER COMPLICATIONS AND WENT INTO HEMORRHAGIC SHOCK THE MORNING AFTER HER INITIAL SURGERY. AND SHE WAS TAKEN EMERGENTLY FOR AN EXPLORATORY LAPAROTOMY AND WAS FOUND TO HAVE AN INJURY TO HER RIGHT EXTERNAL ILIAC ARTERY. | 0
3273426 | DURING THE CASE THE INSTRUMENT FAILED TO DELIVER BIPOLAR ENERGY. THE CORD WAS REPLACED BUT THE PROBLEM CONTINUED. A NEW INSTRUMENT WAS OBTAINED AND THE PROCEDURE CONTINUED WITHOUT FURTHER INCIDENT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED HYSTERECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2915371 | DURING PROCEDURE, MONOPOLAR CURVED SCISSOR IN ARM 1 (DAVINCI CASE) PHYSICIAN SAID HE COULD NOT OPEN OR MOVE SCISSOR IN ANY DIRECTION. ENTIRE TROCAR REMOVED WITH ARM ATTACHED AND INSTRUMENT IN TROCAR. INSTRUMENT REMOVED FROM TROCAR AND AT END OF SCISSOR NEAR TIP, BLACK PLASTIC UNDER THE ORANGE PAPER WAS BENT OUTWARD AND TORN DOWNWARD. INSTRUMENT HAD 6 LIVES LEFT. NEW SCISSOR OPENED. NO HARM TO PATIENT. DEVICE SENT BACK TO COMPANY. | 0
3088893 | INSERTS FOR THE HARMONIC INSTRUMENT FELL APART DURING ATTEMPTED USE (X2). BOTH INSERTS BROKE IN THE SAME MANNER WHILE BEING USED BY A SURGEON VERY FAMILIAR WITH THE DEVICE. PHYSICIAN WAS ABLE TO RETRIEVE BROKEN PIECE FROM PATIENT WITH NO RESIDUAL HARM TO THE PATIENT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?MYOMECTOMY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3195395 | DURING ROBOTIC SURGICAL PROCEDURE THE SURGEON NOTED THAT GRASPER FORCEPS HANDPIECE WOULD NOT RETRACT FROM THE CURVED CANNULA ASSEMBLY. AFTER SEVERAL ATTEMPTS THE COMPLETE GRASPING FORCEPS AND CANNULA ASSEMBLY WAS REMOVED FROM THE ROBOTIC ARM AND EXCHANGED WITH A NEW GRASPER/FORCEPS SET. THE SURGERY WAS COMPLETED WITHOUT FURTHER INCIDENT. THE SURGEON NOTED THAT THE CANNULA HAD CUT THROUGH THE GRASPER FORCEPS INSULATION COATING AND PREVENTED THE FORCEPS FROM BEING REMOVED FROM CANNULA. SURGEON NOTED THAT ALL OF THE FORCEP MATERIAL REMAINED INTACT AND DID NOT ENTER PATIENT SURGICAL FIELD. BOTH CANNULA AND ATTACHED FORCEPS SENT TO BIOMEDICAL FOR EVALUATION AND REPORTING TO MANUFACTURER.======================MANUFACTURER RESPONSE FOR ROBOTIC GRASPING FORCEPS, CROCODILE GRASPER (PER SITE REPORTER).======================MANUFACTURER ISSUED RMA FOR RETURN AND EVALUATION.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC SURGICAL PROCEDURE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO.DEVICE #2IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
2918782 | DURING ROBOTIC PROCEDURE, SURGEON WAS USING A MEGA SUTURECUT NEEDLE DRIVER AND THE CABLE IN THE INSTRUMENT SNAPPED. THE CABLE WAS INTACT. INSTRUMENT WAS TAKEN OFF THE STERILE FIELD AND TAKEN OUT OF SERVICE. THERE WAS NO HARM TO THE PATIENT. | 0
2979934 | DURING ROBOTIC MYOMECTOMY PROCEDURE PHYSICIAN STATED THAT THE PK DISSECTING FORCEP TIPS WERE NOT APPROXIMATING CORRECTLY. ON FURTHER INSPECTION OF THE PK DISSECTING FORCEP, ONE OF THE WIRES ON THE TIP OF THE PULLEY SYSTEM THAT CONTROLS THE DISSECTORS TIPS WAS BROKEN. THE INSTRUMENT HAD 5 USES LEFT WHEN IT BROKE.MANUFACTURER RESPONSE FOR PK DISSECTING FORCEP, INTUITIVE SURGICAL (PER SITE REPORTER).COMPLAINT WAS REGISTERED RMA AND DEVICE WAS RETURNED FOR THEIR REVIEW.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC MYOMECTOMY. | 0
3439824 | THE WIRES ON THE LARGE SUTURECUT NEEDLE DRIVER INSTRUMENT WERE NOTED TO BE FRAYED ON INSPECTION PRIOR TO USE. THERE WERE LIVES/USES STILL LEFT ON THE INSTRUMENT. A NEW DEVICE WAS USED WITHOUT INCIDENT.====================== MANUFACTURER RESPONSE FOR INTUITIVE LARGE SUTURECUT NEEDLEDRIVER, LARGE SUTURECUT NEEDLEDRIVER (PER SITE REPORTER).====================== NONE AT THIS TIME. | 0
3026842 | DURING A ROBOTIC HYSTERECTOMY CASE, THE SURGEON FOUND THE PK DISSECTING FORCEPS ENDOWRIST INSTRUMENT TO BE FAULTY. ON FURTHER INSPECTION, THE WIRES THAT CONTROL THE INSTRUMENT BROKE. THE ABDOMEN WAS SURVEYED BY SURGEON AND PA; IT WAS CONFIRMED NO FOREIGN BODIES WERE LEFT IN PATIENT ABDOMEN. IT WAS ALSO NOTED THAT OUT OF THE 10 USES OF THE PK DISSECTING FORCEPS, THERE WERE 2 USES LEFT.MANUFACTURER RESPONSE FOR PK ENDOWRIST DISSECTING FORCEPS, PK ENDOWRIST DISSECTING FORCEPS (PER SITE REPORTER).DEVICE WAS RETURNED FOR THEIR QUALITY REVIEW. THEY WILL RESPOND WITH THEIR RESULTS.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?ROBOTIC ASSISTED HYSTERECTOMY. | 0
3417214 | THE PATIENT HAD A ROBOTIC COLPOSUSPENSION WITH MESH AND A BILATERAL SALPINGO-OOPHORECTOMY (S&O). SEVERAL DAYS POST OP SHE DEVELOPED POST OPERATIVE PAIN, NAUSEA AND VOMITING. GENERAL SURGERY WAS CONSULTED. AFTER SOME TESTS WERE DONE THE PATIENT WAS TAKEN BACK TO THE OR FOR AN EXPLORATORY LAPAROTOMY. A SMALL BOWEL IATROGENIC INJURY WAS FOUND AND REPAIRED. THE PATIENT TOLERATED THE PROCEDURE WELL. | 0
3417211 | DURING A ROBOTIC ASSISTED HYSTERECTOMY, A HOLE WAS NOTED IN THE BLADDER. THE BLADDER WAS REPAIRED AND A CYSTOSCOPY PERFORMED BY THE PHYSICIAN. | 0
3527349 | SINGLE SITE HOOK CAUTERY SNAPPED ON INSERTION AND INSTRUMENT WAS UNUSABLE. BRAND NEW INSTRUMENT. REF (B)(4). SURGEON INITIATED THE DISSECTION; HOWEVER, ONE OF THE INSTRUMENTS WAS NOT FUNCTIONAL AND WE HAD TO PLACE AN ADDITIONAL PORT IN THE LEFT ABDOMEN TO USE A THIRD ARM FROM THE ROBOT. THIS WAS THE FIRST SINGLE PORT DAVINCI CASE THAT HAS BEEN DONE AT THIS HOSPITAL. BECAUSE OF THIS THERE WAS ONLY ONE SET OF INSTRUMENTS AVAILABLE WITHOUT BACKUPS. CASE WOULD HAVE BEEN OPTIMIZED IF WE HAD A REPLACEMENT SINGLE HOOK CAUTERY. WITH NO BACK UP INSTRUMENT WE HAD TO IMPROVISE AND MAKE ANOTHER INCISION. | 0
3203952 | DURING DAVINCI, AN INSTRUMENT GOT STUCK IN THE DAVINCI ARM AND WOULDN'T COME OUT. AFTER A FEW MINUTES OF MANIPULATION THE PK FINALLY CAME OUT. UPON INSPECTION OF INSTRUMENT A WIRE HAD SNAPPED WHICH DISPLACED IT AND CAUSED THE BROKEN END TO STICK OUT. NO HARM TO PATIENT. A NEW PK WAS OPENED. BROKEN INSTRUMENT WAS SENT BACK TO COMPANY.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3420062 | THIS PATIENT UNDERWENT A ROBOTIC HYSTERECTOMY ABOUT THREE YEARS AGO. SHE DEVELOPED A VAGINAL LEAKAGE AND WAS NOTED TO HAVE ON AN INTRAVENOUS PYELOGRAM (IVP) AND A CT SCAN, A HIGH-GRADE OBSTRUCTION IN THE DISTAL LEFT URETER. SHE UNDERWENT A CYSTOSCOPIC EVALUATION, WHICH WAS CONSISTENT WITH A LOW LEVEL INJURY. A NEPHROSTOMY TUBE WAS PLACED. SHE HAD DEVELOPMENT OF A DISTAL FISTULA FORMATION BETWEEN THE URETER AND THE VAGINAL WALL ASSOCIATED WITH INCONTINENCE. | 0
3420063 | PATIENT CAME IN ELECTIVELY FOR ROBOTIC COLPOSUSPENSION WITH MESH, TENSION-FREE VAGINAL TAPE (TVT), ANTERIOR POSTERIOR (AP) REPAIR. DURING THE PROCEDURE VASCULAR SURGERY HAD TO BE CALLED EMERGENTLY TO THE OPERATING ROOM TO ASSESS PATIENT WHO HAD MASSIVE BLEEDING DURING THE ROBOTIC PROCEDURE. THE ABDOMEN HAD BEEN OPENED BY THE PRIMARY SURGEON AND PRESSURE WAS BEING USED TO CONTROL THE BLEEDING. THE PATIENT HAD TO HAVE HER ABDOMEN OPENED WITH A MIDLINE INCISION FOR REPAIR OF LEFT ILIAC VEIN. AFTER EXPOSURE OF THE AREA MULTIPLE SMALL VENOTOMIES IN THE LEFT ILIAC VEIN WERE NOTED ADJACENT TO AN AREA CONTAINING GORTEX SUTURE AND MESH. THESE AREAS WERE REPAIRED. | 0
3021230 | DURING THE ROBOTIC PROCEDURE, THE SURGEON WANTED TO USE THE VESSEL SEALER. THE TECH LOADED THE VESSEL SEALER INTO ARM 1 OF THE PATIENT CART OF THE ROBOT. IT WAS WITNESSED THAT THE VESSEL SEALER WILL BURN OUT AND CUT ON THE FIRST TRY, AFTERWARDS THE SURGEON TRIED TO USE IT AGAIN AND IT WAS WITNESSED THAT THE VESSEL SEALER WOULD BURN ONLY. ANOTHER DEVICE WAS NEEDED. THERE WAS NO HARM TO THE PATIENT. | 0
3190310 | SNAPFIT SCALPEL WHICH ATTACHES TO THE INSTRUMENT CAME OFF OF INSTRUMENT WHEN SURGEON WAS ROBOTICALLY BIFURCATING THE UTERUS. ABDOMEN WAS SEARCHED ROBOTICALLY AND THE SCALPEL WAS RECOVERED AND NOTED TO BE BENT BUT INTACT. PER SURGEON, MAY HAVE BEEN BENT FROM STRESS AND LONG USE WITH A VERY LARGE FIBROID IN UTERUS.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?PATIENT WITH HISTORY OF FIBOIDS AND ENDOMETRIOSIS. UNDERWENT ROBOTIC LAPAROSCOPIC HYSTERECTOMY REMOVED WITH MORCELLATION, BILATERAL SALPINGECTOMY, AND APPENDECTOMY. PATIENT HAD A 16-WEEK SIZE MULTILOBULAR FIBROID UTERUS WITH SAPROPHYTIC CONNECTIONS OF HER UTERUS AND FIBROIDS TO THE BOWEL.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3308203 | ROBOT VESSEL SEALER WAS USED BY SURGEON BUT IT WAS NOT CUTTING WELL. ITEM WAS REMOVED FROM USE AND A NEW ONE WAS OPENED. NO PROBLEMS WITH THE NEW ONE. THE DEVICE AND ITS ORIGINAL PACKAGING WERE GIVEN TO THE HOSPITAL'S MATERIALS MANAGEMENT DEPARTMENT TO RETURN TO THE MANUFACTURER. | 0
3014770 | LARGE FIBROID REMOVED FROM UTERUS - FIBROID BEING HELD WITH DAVINCI TENACULUM. MONOPOLAR CURVED SCISSORS USED TO CUT FIBROID IN HALF. TENACULUM AND SCISSOR NEVER TOUCHED AND KEPT A GOOD DISTANCE APART WHILE CUTTING. HAVING TROUBLE GETTING SCISSOR TO WORK AND BURN NOTED TO PERITONEUM. NO DAMAGE TO BOWEL OR BLADDER - NO OTHER PROCEDURE NEEDED. DURING TIME MONOPOLAR SCISSOR APPEARED TO BE NOT WORKING IT APPARENTLY ARCED AND ENERGY TRANSFERRED TO TENACULUM. TENACULUM NEAR THE JOINT OF THE INSTRUMENT IS WHAT CAUSED THE BURN ON PERITONEUM. NO ONE SAW IT ARC - IT WAS ASSUMED BY STAFF AND DOCTOR THAT IT ARCED. POSSIBLY FROM DISSECTING ADHESIONS. NO VISIBLE DAMAGE OR IMPERFECTIONS NOTED. INSULATION INTACT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?DAVINCI MYOMECTOMY. | 0
3234347 | A RADIOLOGIC TECHNOLOGIST FROM A DENTAL IMAGING CENTER REPORTED TO A MIDMARK SALES REPRESENTATIVE THAT DURING THE PROCESS OF POSITIONING A PATIENT, A VANTAGE PANORAMIC X-RAY UNIT, SERIAL NUMBER (B)(4), WOULD CONTINUE MOVING DOWN AFTER THE "DOWN" BUTTON ON THE PATIENT TABLE WAS RELEASED. THE RADIOLOGIC TECHNOLOGIST REMOVED THE PATIENT FROM THE MACHINE. THERE WAS ONLY ONE OCCURRENCE AND THE MACHINE PERFORMED NORMALLY AFTERWARDS. THE MIDMARK SALES REPRESENTATIVE SPOKE TO THE RADIOLOGIC TECHNOLOGIST AND CONFIRMED THERE WAS NO INJURY. THE UNIT IS EQUIPPED WITH AN EMERGENCY STOP SWITCH. THE USER MANUAL EXPLAINS THAT "DEPRESSING THE BUTTON WILL IMMEDIATELY HALT ALL MOTOR MOVEMENT". A MIDMARK SALES REPRESENTATIVE PERFORMED AN ON-SITE INVESTIGATION AND CONFIRMED THAT THE VERTICAL MOVEMENT SWITCHES ALLOWED THE COLUMN TO MOVE UP, DOWN AND STOP UPON COMMAND. THE EMERGENCY STOP SWITCH WAS CONFIRMED TO FUNCTION AS INTENDED. THE VERTICAL MOVEMENT SWITCH AND THE ASSOCIATED CABLE WERE REPLACED AND RETURNED TO MIDMARK FOR FURTHER ANALYSIS. THE ALLEGED FAILURE COULD NOT BE REPRODUCED. | 0
3397073 | THE TELEMETRY UNITS/BOXES HAVE A DESIGN FAILURE IN THAT THE SLIDE-TOP CLOSURE WHICH ENCASES THE BATTERIES MUST BE MANIPULATED TO ENABLE A CONNECTION. THE BATTERY CONNECTION IS REQUIRED FOR THE PT'S CARDIAC RHYTHM TO SHOW ON THE TELEMETRY MONITORS. THE TELEMETRY TECHNICIANS AND THE NURSES MOVE THE SLIDE-LID BACK AND FORTH UNTIL THE POSITION IS DISCOVERED WHICH ALLOWS THE RHYTHM TO SHOW ON THE MONITOR. OF COURSE, IF THERE IS SUBSEQUENT MOVEMENT OF THE SLIDE-LID, THE CONNECTION CAN BE LOST (WHICH WAS THE CASE IN THIS PARTICULAR INCIDENT). THIS WAS DISCOVERED WHEN THE SITUATION WAS REPORTED UP TO A NURSING SUPERVISOR, WHO THEN NOTIFIED THE RISK MANAGER. AN INVESTIGATION WAS STARTED. SINCE THE DATE OF THE EVENT WITH THIS PT, 17 ADDITIONAL TELEMETRY UNIT/BOXES HAVE BEEN SENT BACK TO GE. I (WE) ARE CONCERNED BECAUSE THIS ISSUE GOES BEYOND THIS SINGLE PT. IT COULD AFFECT OTHER PTS AS WELL, BOTH WITHIN OUR HOSPITAL AND OTHERS. GE HAD PREVIOUSLY ON ((B)(6) 2013) TESTED 100%OF OUR UNITS/BOXES WITH NO ISSUES FOUND. AND THE RE-TESTING OF THE UNIT/BOX ATTACHED TO THIS PT WAS RETESTED WITH NO FINDINGS. HOWEVER, THEY ARE TESTED ON SIMULATOR AND OF COURSE THE CONDITIONS ARE STATIC WITH A SIMULATOR, WHILE THEY ARE DYNAMIC WITH A PT WHO MOVES, SWEATS, SCRATCHES, ETC (ALL OF WHICH MAY CAUSE MOVEMENT OF THE PARTIALLY OPEN SLIDE-LID). HOWEVER, THE PURPOSE OF THIS REPORT IS TO MAKE THE FDA AWARE OF A PRODUCT MALFUNCTION THAT MUST BE ADDRESSED IMMEDIATELY. THIS PT HAD HER UNIT/BOX ATTACHED, BUT THE SIGNAL WAS LOST. AND, NOW WE HAVE LEARNED THIS IS A COMMON PROBLEM WITH THESE UNITS/BOXES. (B)(4). | 0
2785160 | THE CUSTOMER REPORTED THAT THEY OBSERVED A "NO COMM" TECHNICAL MESSAGE ON THE SCREEN OF THE MONITORING DEVICE, INDICATING MONITORING WAS NOT OCCURRING. SIXTEEN TELEMETRY ONLY PTS WERE UNMONITORED FOR APPROX 20 MINUTES UNTIL THEY COULD BE RE-ADMITTED IN ORDER TO RESUME MONITORING. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
2806726 | THE CUSTOMER REPORTED THE APEXPRO TELEMETRY SERVER DID NOT BOOT BACK UP AFTER A POWER ISSUE. THE CUSTOMER LOST USE OF THEIR MONITOR FOR 45 MINUTES. THE CLINICAL INFO CENTERS (CIC'S) WERE STILL WORKING AND DISPLAYING NO COMM MESSAGES. WHEN THE SYSTEM WAS TURNED BACK ON, THE CIC'S DISPLAYED WAVEFORMS APPROPRIATELY. THE CUSTOMER HAD BEDSIDE MONITORS AVAILABLE BUT NOT ENOUGH TO COVER THE NUMBER OF PTS THAT EXPERIENCED THE LOSS OF MONITORING DURING THIS INCIDENT. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. A F/U REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. SERIAL NUMBER AND DEVICE MANUFACTURE DATE - UNK AT THIS TIME. | 1
2844407 | THE CUSTOMER REPORTED THAT THE CLINICAL INFO CENTER SHUT DOWN DUE TO AN ISSUE WITH THE APEXPRO TELEMETRY SERVER (ATS) FAN FAILURE. THIS CAUSED LOSS OF MONITORING FOR THIRTY MINUTES AFFECTING FIVE PTS UNTIL ALTERNATE ATS MONITORING WAS ESTABLISHED. THERE WAS NO REPORTED PT INJURY ASSOCIATED WITH THIS EVENT. THE SCOPE OF THIS INVESTIGATION IS BASED UPON THE INFO OBTAINED FROM THE FIELD ENGINEER (FE). AFTER ARRIVING ONSITE, THE FE FOUND THAT THE APEXPRO TELEMETRY SERVER (ATS) HAD SHUT DOWN. AFTER POWERING THE UNIT BACK UP, THE FE RAN THE WEBMIN DIAGNOSTICS AND NOTED THAT THE CENTRAL PROCESSING UNIT (CPU) FAN WAS NOT OPERATING. A REPLACEMENT CPU FAN WAS INSTALLED IN THE UNIT ON (B)(4) 2012. THE FE CONFIRMED THAT THE ATS IS NOW BACK ONLINE AND FUNCTIONING NORMALLY AND THE USER IS ABLE TO ADMIT PTS TO THE SYSTEM. REVIEW OF THE SERVICE HISTORY OF THE ATS SHOWED THAT THE CPU FAN HAS NOT BEEN REPLACED SINCE THE UNIT WAS INSTALLED APPROX FOUR YEARS AGO. | 1
3193451 | TELEMETRY PAKS NOT TRANSFERRING THE PT WAVE FORM. PROBLEM IS OCCURRING WITH MULTIPLE PACKS OF THE SAME SYSTEM ON THE SAME FLOOR BUT NOT AT THE SAME TIME. | 1
3232783 | PATIENT WAS ON TELEMETRY MONITORING. MONITOR WAS REMOVED AND SYSTEM WAS PUT INTO STANDBY MODE WHILE BEING SENT FOR A CT EXAM. UPON RETURN TO DEPT PATIENT WAS RECONNECTED TO TELEMETRY MONITOR BUT MONITOR NEVER RECONNECTED TO THE CENTRAL STATION. DURING ROUNDS PATIENT WAS FOUND UNRESPONSIVE. | 1
2929963 | TELEMETRY PACK MONITOR #77--BATTERIES STARTED TO SIZZLE AND POPPED. BATTERY ACID THEN STARTED TO DRIZZLE OUT OF THE BATTERY, CAUSING FUMES. PATIENT'S BLANKET WAS BURNED BY THE HOT BATTERIES TAKEN OUT OF THE MONITOR BY THE PATIENT CARE TECH.======================MANUFACTURER RESPONSE FOR TELEMETRY MONITOR, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================NEED TO SEND TELEMETRY BOX TO MANUFACTURER.><TELEMETRY PACK MONITOR #77--BATTERIES STARTED TO SIZZLE AND POPPED. BATTERY ACID THEN STARTED TO DRIZZLE OUT OF THE BATTERY, CAUSING FUMES. PATIENT'S BLANKET WAS BURNED BY THE HOT BATTERIES TAKEN OUT OF THE MONITOR BY THE PATIENT CARE TECH.======================MANUFACTURER RESPONSE FOR TELEMETRY MONITOR, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================NEED TO SEND TELEMETRY BOX TO MANUFACTURE. | 1
3515411 | EVENT: THIS EVENT OCCURRED IN A (B)(6) HOSPITAL ON (B)(6) 2013 AND REPORTED TO US BY THE HOSPITAL ON (B)(4) 2013. A MEDICAL STAFF NOTICED THE PT ECG WAVEFORM ON THE CENTRAL MONITOR STATION, (B)(4), SCREEN TO BE A SQUARE WAVEFORM INDICATING THE PT BEING MONITORED ON A TELEMETRY TRANSMITTER MONITOR, LX-5230, WAS OFF-LINE, OUT OF RANGE, OR NOT CONNECTED. THE MEDICAL STAFF PERSON RUSHED INTO PT ROOM AND FOUND THE PT HAD EXPIRED. UPON EXAMINATION OF THE TELEMETRY TRANSMITTER MONITOR, IT WAS DETERMINED THAT THE BATTERY POWER OF TELEMETER LX-5230 WAS EXHAUSTED. THIS EVENT WAS DUE TO USER ERROR AND NOT DUE TO ANY MALFUNCTION OF DEVICE, WHICH WAS ALSO ACKNOWLEDGED BY THE FACILITY. THE TELEMETRY TRANSMITTER MONITOR INVOLVED IN THIS EVENT IS THE FUKUDA DENSHI MODEL LX-5230 WHICH IS SOLD IN THE (B)(4) MARKET BUT WHICH IS SIMILAR TO THE FUKUDA DENSHI TELEMETRY TRANSMITTER MONITOR LX-5630 WHICH IS DISTRIBUTED AND SOLD IN THE US MARKET. THIS EVENT OCCURRED SOLELY AS A RESULT OF USER ERROR AND NOT DUE TO ANY MALFUNCTION OF THE DEVICE. AND AS PER FUKUDA DENSHI CO. LTD. QA SYSTEM SOP, IS WAS DETERMINED TO BE A REPORTABLE EVENT AND THEREFORE AN FDA FORM 3500A WAS ORDERED TO BE FILED. IN ADDITION, AS A RESULT OF THIS EVENT, THE HOSPITAL STAFF WERE PROVIDED WITH ADDITIONAL INSTRUCTION AND THOROUGH EXPLANATION OF DEVICE OPERATION AND THE IMPORTANCE OF ROUTINELY CHECKING THE DEVICE POWER SOURCE AND MONITORING OF THE CENTRAL STATION. | 1
3267253 | NURSE CHANGED BATTERIES IN TELEMETRY UNIT #46 AND LEFT ROOM. SHORTLY THEREAFTER, THE PATIENT CALLED FOR ASSISTANCE, WHEN NURSE ENTERED THE ROOM THE TELEMETRY UNIT WAS OFF THE PATIENT AND ON THE MATTRESS. THE PATIENT STATED "ITS HOT". THE NURSE TOUCHED THE UNIT AND IT WAS VERY HOT TO TOUCH. THE NURSE OPENED THE TELEMETRY UNIT AND THE BATTERIES WERE MAKING A SIZZLING SOUND AND APPEARED MELTED. | 1
3244633 | WHILE THE PHYSICIAN WAS HANDLING THE DEVICE DURING THE PROCEDURE, THE DEVICE FAILED TO AUTO FLEX TO 90 DEGREES. THE PROCEDURE WAS DELAYED WHILE THE PROBE WAS CHANGED.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?TEE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3520553 | DURING MONITORING, THE PATIENT COMPLAINED OF PAIN IN THE LOCATION OF THE TELEMETRY MONITOR. THE NURSE (RN) NOTICED 2 REDDISH-BLUE AREAS ON THE PATIENT'S UPPER FOREARM. AS THE RN PICKED UP THE TELEMETRY MONITOR TO MOVE IT TO FURTHER ASSESS THE PATIENT'S FOREARM, SHE NOTICED THAT THE TELEMETRY MONITOR WAS VERY HOT. THE RN IMMEDIATELY DISCONNECTED THE MONITOR FROM THE PATIENT AND NOTIFIED THE MONITOR TECH THAT THE PATIENT WOULD BE OFF THE MONITOR FOR A LITTLE BIT. THE RN STEPPED OUTSIDE OF THE PATIENT'S ROOM WITH THE TELEMETRY MONITOR AND ATTEMPTED TO REMOVE THE BATTERY; THE BATTERY WAS TOO HOT TO REMOVE WITH BARE HANDS. THE PATIENT WAS THEREAFTER CONNECTED TO ANOTHER TELEMETRY MONITOR. THE PATIENT'S WOUND WAS ASSESSED BY A PHYSICIAN, A WOUND CONSULT WAS ORDERED, AND SILVER SULFADIAZINE WAS APPLIED. THERE WAS NO BLISTERING. BIOMED CHECKED THE TELEPACK AND WAS ABLE TO DUPLICATE THE ISSUE WITH THE PACK HEATING UP. BIOMED WILL BE SENDING TO THE DEVICE TO THE MANUFACTURER FOR REPAIR. BIOMED TRIED A NEW SET OF BATTERIES WITH THE PACK AND THEY ALSO HEATED UP. THE BATTERY TABS ARE NOT BROKEN OFF AND THE BOTTOM BATTERY TERMINAL APPEARS TO BE SLIGHTLY BENT. THERE ARE NO OTHER NOTICEABLE DAMAGES TO THE PACK. THE BATTERY TERMINAL IS SOMEWHAT DIRTY AND NEEDS TO BE REPLACED. | 0
3272943 | THE CUSTOMER REPORTED A PLATELET CONTAMINATION WITH RED BLOOD CELLS. IT IS NOT KNOWN AT THIS TIME IF WHITE BLOOD CELLS (WBCS) WERE ALSO FOUND IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR INJURY OCCURRED. INVESTIGATION: PER THE CUSTOMER, THE PLATELET CONTAMINATION WAS NOTED FOLLOWING AN UPGRADE OF THE DEVICE. THE SERVICE TECHNICIAN NOTED A DEFECTIVE TEST CASSETTE UPON RETURNING TO CALIBRATE THE DEVICE. THERE HAVE BEEN NO FURTHER COMPLAINTS ABOUT THE DEVICE REGARDING THE REPORTED CONDITION SINCE THE RE-CALIBRATION. ROOT CAUSE: THE LIKELY ROOT CAUSE OF THE PLATELET CONTAMINATION IS A CALIBRATION ISSUE WITH THE RED BLOOD CELL DETECTOR DUE TO A WORN TEST CASSETTE USED FOR THE ORIGINAL CALIBRATION. CORRECTION: A NEW TEST CASSETTE WAS OBTAINED AND THE DEVICE WAS CALIBRATED AS APPROPRIATE.><. | 0
3428260 | THE CUSTOMER REPORTED THAT WHEN THE DONOR WAS DETACHED, A LEAK WAS NOTICED AT THE LUER. THE PRODUCT WAS DESTROYED. NO MEDICAL INTERVENTION WAS REQUIRED FOR THIS EVENT. THE DONOR WAS CALLED THE NEXT DAY TO DETERMINE IF HE HAD ANY SIDE EFFECTS. HE DID NOT HAVE ANY SIDE EFFECTS TO REPORT. TERUMO BCT IS AWAITING THE RETURN OF THE DISPOSABLE SET. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE DISPOSABLE SET WAS RETURNED FOR INVESTIGATION. A LEAK WAS CONFIRMED AT THE JOINT BETWEEN THE MANIFOLD AND THE NEEDLE-LESS ACCESS (NLA) PORT. THE MANIFOLD WAS FOUND TO HAVE A STRESS CRACK. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE STRESS MARKS ON THE COMPONENT AND THE LEAK OBSERVED IN THE RETURNED PART INDICATE THAT THE ROOT CAUSE IS RELATED TO A DEFECTIVE PART FROM THE VENDOR THAT OCCURRED DURING ASSEMBLY. TESTING HAS SHOWN THAT NOT ALL COMPONENTS WITH STRESS MARKS RESULT IN A LEAK. CORRECTIVE ACTIONS: TERUMO BCT'S STAFF WERE MADE AWARE OF THIS INCIDENT. A 100% INSPECTION AND SORT OF THE NEEDLE-LESS ACCESS COMPONENTS HAS BEEN IMPLEMENTED TO REDUCE THE OCCURRENCE OF THESE LEAKS. THE VENDOR WAS ALSO MADE AWARE OF THIS INCIDENT AND ACTIONS ARE IN PROGRESS TO IMPROVE THE MOLDING PROCESS. | 0
2847083 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE FOLLOW-UP #1 REPORT FOR THIS EVENT.><(B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2944874 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED><THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
2994868 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO PLASMA LINE OCCLUSION. CORRECTIVE ACTION: AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED RWBCS RELATED TO PLASMA LINE OCCLUSIONS. | 0
2786377 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
3236694 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE RUN DATA FILES (RDF) WERE IDENTIFIED AND ANALYZED AS A GROUP. A REVIEW OF THESE FILES DID NOT INDICATE AN OBVIOUS CAUSE FOR THE REPORTED ELEVATED WBC LEVELS. EVIDENCE OF THE FOLLOWING TYPES OF FAILURES WERE FOUND:IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. BASED ON THE AVAILABLE INFORMATION, IT IS ALSO POSSIBLE THAT SEVERAL OF THESE LEUKO REDUCTION FAILURES COULD BE DONOR-RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. REVIEW OF THE RUN DATA FILES DID NOT INDICATE AN OBVIOUS ROOT CAUSE FOR THE REPORTED ELEVATED WBC CONTENT. POSSIBLE ROOT CAUSES WERE PROVIDED IN FOLLOW-UP #1 FOR THIS REPORT. CORRECTIVE/PREVENTIVE ACTION: AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED RWBCS RELATED TO PLASMA LINE OCCLUSIONS.><THIS RECORD IS BEING FILED TO CORRECT DATA THAT WAS DUPLICATED IN THE INITIAL REPORT DUE TO A TECHNICAL ERROR. THE TECHNICAL ERROR HAS NOW BEEN RESOLVED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED><INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2774792 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><(B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RDF DID NOT INDICATE A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR MAY HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2873815 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE TRIPLE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><(B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING THE PROCEDURE. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED DURING A PORTION OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
3075363 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. IT IS POSSIBLE THAT THE HIGH NUMBER OF ACCESS ALERTS AND ADJUSTMENTS THAT OCCURRED THROUGHOUT THE PROCEDURE DISRUPTED THE STEADY-STATE OF THE SYSTEM AND CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
3236727 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THEELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT.THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBCTESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT.DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT.THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT ID: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE RUN DATA FILES (RDF) WERE IDENTIFIED AND ANALYZED AS A GROUP. A REVIEW OF THESE FILES DID NOT INDICATE AN OBVIOUS CAUSE FOR THE REPORTED ELEVATED WBC LEVELS. EVIDENCE OF THE FOLLOWING TYPES OF FAILURES WERE FOUND;IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. BASED ON THE AVAILABLE INFORMATION, IT IS ALSO POSSIBLE THAT SEVERAL OF THESE LEUKOREDUCTION FAILURES COULD BE DONOR-RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THIS RECORD IS BEING FILED TO CORRECT DATA THAT WAS DUPLICATED IN THE INITIAL REPORT DUE TO A TECHNICAL ERROR. THE TECHNICAL ERROR HAS NOW BEEN RESOLVED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED><ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. REVIEW OF THE RUN DATA FILES DID NOT INDICATE AN OBVIOUS ROOT CAUSE FOR THE REPORTED ELEVATED WBC CONTENT. POSSIBLE ROOT CAUSES WERE PROVIDED IN FOLLOW-UP #2 FOR THIS REPORT. CORRECTIVE/PREVENTIVE ACTION: AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED RWBCS RELATED TO PLASMA LINE OCCLUSIONS. | 0
2950848 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE DOUBLE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION:THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED,SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD ALSO BE DONOR-RELATED.><INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2748980 | THE CUSTOMER REPORTED ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RDF INDICATE THE POSSIBILITY THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED DURING A PORTION OF THE RUN. THE OCCLUSION OF THE PLASMA LINE CAUSES THE PLASMA LINE TO NO LONGER CONTRIBUTE TO THE PLATELET PUMP DURING THE PCA SUBSTATE, WHICH IN TURN CAUSES THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS. THIS CAN CAUSE WBCS TO ESCAPE THE LEUKOREDUCTION SYSTEM CHAMBER AND END UP IN THE PRODUCT BAG. THE ORIENTATION OF THE TUBING AS LOADED IN THE CENTRIFUGE HEX HOLDER UNDER CERTAIN FLOW CONDITIONS, MAY CAUSE THE PLASMA LINE TO PINCH OFF. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. ALGORITHMS HAVE BEEN INCORPORATED INTO THE NEXT VERSION OF SOFTWARE FOR THE TRIMA ACCEL SYSTEM. THESE ALGORITHMS RECOGNIZE A POTENTIAL ELEVATED WBC COUNT DUE TO THE FAILURE MODE WITNESSED DURING THIS INVESTIGATION. THE SOFTWARE WILL FLAG THE PROCEDURE TO MEASURE OR "VERIFY WBC'S IN THE PLATELET PRODUCT". THE NEW ALGORITHMS ARE ADDED IN THE 6.1.2 TRIMA EQUIPMENT SOFTWARE UPGRADE. INTERNAL CAPAS HAVE BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS. | 0
3019366 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THIS FAILURE. | 0
2859094 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE DOUBLE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><(B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
3089758 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE DOUBLE PLATELET PRODUCT AT THEIR (B)(6) SITE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. THE CUSTOMER PROVIDED ELEVATED WBC DATA FOR MULTIPLE PROCEDURES. THE DATA FOR THESE PROCEDURAL EVENTS WAS ANALYZED AS A GROUPING. INVESTIGATION OF AVAILABLE RUN DATA FILES (RDFS)IDENTIFIED SEVERAL PROCEDURAL EVENTS, AS LISTED BELOW, THAT MAY HAVE LED TO ELEVATED WBC LEVELS IN COLLECTED PRODUCT. ROOT CAUSE: RDFS FOR THE PROCEDURES WITH HIGH RWBC WERE RECEIVED, INVESTIGATED AND ANALYZED. THE DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCTS WERE IDENTIFIED AND ARE AS FOLLOWS: PRE-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT THE LRS CHAMBER HAS REACHED ITS WBC HOLDING CAPACITY BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. POST-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT AFTER AN LRS PURGE, THE LRS CHAMBER HAS RE-FILLED WITH WBCS BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. PLASMA LINE OCCLUSIONS - SIGNALS IN THE RDF SUGGEST THAT THE PLASMA LINE WAS PARTIALLY OCCLUDED DURING A PORTION OF THE PROCEDURE. THIS CAN CAUSE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN EXPECTED, DECREASING THE HOLDING CAPACITY OF THE LRS CHAMBER. OTHER EVENTS - THESE CAN INCLUDE; CENTRIFUGE STOPS, MULTIPLE ACCESS ALARMS & FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO A NON STEADY STATE CHAMBER CONDITIONS, AND RBC SPILLOVER SWITCH CAN POSSIBLY LEAD TO A HIGHER THAN EXPECTED WBC CONTENT. ALSO INCLUDED ARE DONOR RELATED CHARACTERISTICS. CORRECTION: A TRIMA SOFTWARE VERSION UPGRADE HAS BEEN COMPLETED AT ALL OF THE CUSTOMER¿S CENTERS. THE CUSTOMER PROVIDED DATA GATHERED FROM A TERUMO BCT RECOMMENDED TRIAL OF AVERAGING THE LAST 3 PLATELET COUNTS OF DONORS. THIS DATA SUGGESTS THAT USING THE AVERAGE PLATELET COUNT SHOWED A MORE CONSISTENT PROCESS, WHICH WILL REDUCE VARIATION SHOWN TO CONTRIBUTE TOWARDS RWBC FAILURES. | 0
2807031 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR CATHEDRAL COURT SITE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. THIS ELEVATED WBC INCIDENT WAS ANALYZED AS A GROUPING OF COMPLAINTS SINCE (B)(6) 2012 FROM THIS CUSTOMER. THE TRIMA SYSTEM WILL IDENTIFY CERTAIN EVENTS THAT OCCUR DURING THE PROCEDURE THAT CAN POSSIBLY AFFECT THE WBC CONTENT OF THE COLLECTED PRODUCTS, THE OPERATOR WILL BE NOTIFIED AND THE REASON DISPLAYED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED RUNS. THE ANALYSIS IDENTIFIED 4 DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT, AND THE FREQUENCY THAT THESE EVENTS OCCURRED WITH THE CUSTOMER. THESE ARE: PRE-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER FILLS UP BEFORE SYSTEM EXPECTS, POST-PURGE SATURATIONS ((B)(4)) - LRS CHAMBER RE-FILLS BACK UP BEFORE EXPECTED, PLASMA LINE OCCLUSIONS ((B)(4)) - CURRENTLY UNDER INVESTIGATION FOR MITIGATION AND DETECTION OF THESE EVENTS, EVENT BASED - INCLUDES CENTRIFUGE STOP, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO NON STEADY STATE CHAMBER CONDITIONS, DONOR RELATED. | 0
3089754 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE DOUBLE PLATELET PRODUCT AT THEIR (B)(6) SITE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. THE CUSTOMER PROVIDED ELEVATED WBC DATA FOR MULTIPLE PROCEDURES. THE DATA FOR THESE PROCEDURAL EVENTS WAS ANALYZED AS A GROUPING. INVESTIGATION OF AVAILABLE RUN DATA FILES (RDFS) IDENTIFIED SEVERAL PROCEDURAL EVENTS, AS LISTED BELOW, THAT MAY HAVE LED TO ELEVATED WBC LEVELS IN COLLECTED PRODUCT. ROOT CAUSE: RDFS FOR THE PROCEDURES WITH HIGH RWBC WERE RECEIVED, INVESTIGATED AND ANALYZED. THE DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCTS WERE IDENTIFIED AND ARE AS FOLLOWS: PRE-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT THE LRS CHAMBER HAS REACHED ITS WBC HOLDING CAPACITY BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. POST-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT AFTER AN LRS PURGE, THE LRS CHAMBER HAS RE-FILLED WITH WBCS BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. PLASMA LINE OCCLUSIONS - SIGNALS IN THE RDF SUGGEST THAT THE PLASMA LINE WAS PARTIALLY OCCLUDED DURING A PORTION OF THE PROCEDURE. THIS CAN CAUSE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN EXPECTED, DECREASING THE HOLDING CAPACITY OF THE LRS CHAMBER. OTHER EVENTS - THESE CAN INCLUDE; CENTRIFUGE STOPS, MULTIPLE ACCESS ALARMS AND FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO A NON STEADY STATE CHAMBER CONDITIONS, AND RBC SPILL OVERS WHICH CAN POSSIBLY LEAD TO A HIGHER THAN EXPECTED WBC CONTENT. ALSO INCLUDED ARE DONOR RELATED CHARACTERISTICS. CORRECTION: A TRIMA SOFTWARE VERSION UPGRADE HAS BEEN COMPLETED AT ALL OF THE CUSTOMER¿S CENTERS. THE CUSTOMER PROVIDED DATA GATHERED FROM A TERUMO BCT RECOMMENDED TRIAL OF AVERAGING THE LAST 3 PLATELET COUNTS OF DONORS. THIS DATA SUGGESTS THAT USING THE AVERAGE PLATELET COUNT SHOWED A MORE CONSISTENT PROCESS, WHICH WILL REDUCE VARIATION SHOWN TO CONTRIBUTE TOWARDS RWBC FAILURES. | 0
3090390 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE DOUBLE PLATELET PRODUCT AT THEIR (B)(6) SITE. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. THE CUSTOMER PROVIDED ELEVATED WBC DATA FOR MULTIPLE PROCEDURES. THE DATA FOR THESE PROCEDURAL EVENTS WAS ANALYZED AS A GROUPING. INVESTIGATION OF AVAILABLE RUN DATA FILES (RDFS)IDENTIFIED SEVERAL PROCEDURAL EVENTS, AS LISTED BELOW, THAT MAY HAVE LED TO ELEVATED WBC LEVELS IN COLLECTED PRODUCT. ROOT CAUSE: RDFS FOR THE PROCEDURES WITH HIGH RWBC WERE RECEIVED, INVESTIGATED AND ANALYZED. THE DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCTS WERE IDENTIFIED AND ARE AS FOLLOWS: PRE-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT THE LRS CHAMBER HAS REACHED ITS WBC HOLDING CAPACITY BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. POST-PURGE SATURATIONS - SIGNALS IN THE RDF SUGGEST THAT AFTER AN LRS PURGE, THE LRS CHAMBER HAS RE-FILLED WITH WBCS BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. PLASMA LINE OCCLUSIONS - SIGNALS IN THE RDF SUGGEST THAT THE PLASMA LINE WAS PARTIALLY OCCLUDED DURING A PORTION OF THE PROCEDURE. THIS CAN CAUSE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN EXPECTED, DECREASING THE HOLDING CAPACITY OF THE LRS CHAMBER. OTHER EVENTS - THESE CAN INCLUDE; CENTRIFUGE STOPS, MULTIPLE ACCESS ALARMS & FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO A NON STEADY STATE CHAMBER CONDITIONS, AND RBC SPILLOVER SWITCH CAN POSSIBLY LEAD TO A HIGHER THAN EXPECTED WBC CONTENT. ALSO INCLUDED ARE DONOR RELATED CHARACTERISTICS. CORRECTION: A TRIMA SOFTWARE VERSION UPGRADE HAS BEEN COMPLETED AT ALL OF THE CUSTOMER¿S CENTERS. THE CUSTOMER PROVIDED DATA GATHERED FROM A TERUMO BCT RECOMMENDED TRIAL OF AVERAGING THE LAST 3 PLATELET COUNTS OF DONORS. THIS DATA SUGGESTS THAT USING THE AVERAGE PLATELET COUNT SHOWED A MORE CONSISTENT PROCESS, WHICH WILL REDUCE VARIATION SHOWN TO CONTRIBUTE TOWARDS RWBC FAILURES.><INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2957491 | THE CUSTOMER REPORTED AN ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT AT THEIR (B)(6). THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. THE CUSTOMER PROVIDED ELEVATED WBC DATA FOR MULTIPLE PROCEDURES. THE DATA FOR THESE PROCEDURAL EVENTS WAS ANALYZED AS A GROUPING. INVESTIGATION OF AVAILABLE RUN DATA FILES (RDFS) IDENTIFIED SEVERAL PROCEDURAL EVENTS, AS LISTED BELOW, THAT MAY HAVE LED TO ELEVATED WBC LEVELS IN COLLECTED PRODUCT. ROOT CAUSE: RDFS FOR THE PROCEDURES WITH HIGH RWBC WERE RECEIVED, INVESTIGATED AND ANALYZED. THE ANALYSIS INCLUDED BOTH FLAGGED AND UNFLAGGED PROCEDURES. THE DIFFERENT TYPES OF EVENTS THAT LED TO ELEVATED WBC CONTENT IN PLATELET PRODUCT WERE IDENTIFIED AND ARE AS FOLLOWS: PRE-PURGE SATURATIONS¿ SIGNALS IN THE RDF SUGGEST THAT THE LRS CHAMBER HAS REACHED ITS WBC HOLDING CAPACITY BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. POST-PURGE SATURATIONS¿ SIGNALS IN THE RDF SUGGEST THAT AFTER AN LRS PURGE, THE LRS CHAMBER HAS RE-FILLED WITH WBCS BEFORE EXPECTED BY THE TRIMA ACCEL SYSTEM. PLASMA LINE OCCLUSIONS¿ SIGNALS IN THE RDF SUGGEST THAT THE PLASMA LINE WAS PARTIALLY OCCLUDED DURING A PORTION OF THE PROCEDURE. THIS CAN CAUSE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN EXPECTED, DECREASING THE HOLDING CAPACITY OF THE LRS CHAMBER. OTHER EVENTS ¿ THESE CAN INCLUDE; CENTRIFUGE STOPS, MULTIPLE ACCESS ALARMS & FLOW ADJUSTMENTS THAT CAN CONTRIBUTE TO A NON STEADY STATE CHAMBER CONDITIONS, AND RBC SPILLOVERS WHICH CAN POSSIBLY LEAD TO A HIGHER THAN EXPECTED WBC CONTENT. ALSO INCLUDED ARE DONOR RELATED CHARACTERISTICS. CORRECTION: A TRIMA SOFTWARE VERSION UPGRADE HAS BEEN COMPLETED AT ALL OF THE CUSTOMER¿S CENTERS. THE CUSTOMER PROVIDED DATA GATHERED FROM A TERUMO BCT RECOMMENDED TRIAL OF AVERAGING THE LAST 3 PLATELET COUNTS OF DONORS. THIS DATA SUGGESTS THAT USING THE AVERAGE PLATELET COUNT SHOWED A MORE CONSISTENT PROCESS, WHICH WILL REDUCE VARIATION SHOWN TO CONTRIBUTE TOWARDS RWBC FAILURES.><INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED | 0
2773595 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBC COUNTS.><TERUMO BCT INTERNAL REFERENCE: (B)(4). INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RDF INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER TOWARD THE END OF THE PROCEDURE WHEN THE ACCESS ALERTS WERE OCCURRING. BASED ON THE AVAILABLE INFORMATION, IT IS ALSO POSSIBLE THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. IN ORDER TO EFFICIENTLY COLLECT PLATELETS, THE SYSTEM NEEDS TO OPERATE AT A STEADY STATE. ACCESS ISSUES DURING THE PROCEDURE SHOULD BE MINIMIZED TO THE EXTENT POSSIBLE PER THE OPERATOR'S MANUAL. NUMEROUS ALERTS CAN SOMETIMES DELAY THE SYSTEM FROM REACHING STEADY STATE AND RESULT IN NON-OPTIMAL COLLECTIONS. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2861768 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
3092469 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED RWBCS RELATED TO PLASMA LINE OCCLUSIONS.><INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, "THOUNOT" CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED | 0
3031129 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, BUT NOT CONCLUSIVE, THAT THE PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED><THIS REPORT IS BEING FILED TO PROVIDE ADDITIONAL INFORMATION. INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED RWBCS RELATED TO PLASMA LINE OCCLUSION. | 0
2957661 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDE IN THE INITIAL REPORT FOR THIS EVENT. AN INTERNAL CAPA HAS BEEN INITIATED TO EVALUATE REPORTS OF ELEVATED WBCS RELATED TO PLASMA LINE OCCLUSIONS.><INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. THE SIGNALS IN THE RUN DATA FILE INDICATE IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT TH PLASMA LINE MAY HAVE BEEN PARTIALLY OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP, IT COULD CAUSE THE FLOW THROUGH THE LRS CHAMBER TO BE HIGHER THAN THE SYSTEM EXPECTS, POSSIBLY ALLOWING SOME WBCS TO ESCAPE. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
2870515 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. (B)(4). INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND ANY CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. IT IS POSSIBLE THAT THE NUMEROUS ACCESS PRESSURE ALERTS COULD HAVE DISRUPTED THE STEADY STATE OF THE SYSTEM AND COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. ADDITIONALLY, BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><(B)(4). THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
3435541 | THE CUSTOMER REPORTED A LEAK AT THE NEEDLESS ACCESS LOCATED ON THE INLET/RETURN TUBING. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. INVESTIGATION: THE CUSTOMER CLEANED THE DISPOSABLE SET AND CONFIRMED THE LEAK UNDER WATER PRESSURE THE CUSTOMER RETURNED THE INJECTION PORT ASSEMBLY, SAMPLE BAG AND THE AC/DRAW/RETURN MANIFOLD WERE RETURNED FOR INVESTIGATION. LEAK TESTING DID NOT REVEAL A LEAK UNTIL THE INJECTION PORT WAS MANIPULATED. PLASTIC STRESS MARKS WERE NOTED AT THE LEAK LOCATION. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE STRESS MARKS ON THE COMPONENT AND THE LEAK OBSERVED IN THE RETURNED PART INDICATE THAT THE ROOT CAUSE IS RELATED TO A DEFECTIVE PART FROM THE VENDOR THAT OCCURRED DURING ASSEMBLY. TESTING HAS SHOWN THAT NOT ALL COMPONENTS WITH STRESS MARKS RESULT IN A LEAK. CORRECTIVE ACTION: TERUMO BCT'S STAFF WERE MADE AWARE OF THIS INCIDENT. A 100% INSPECTION AND SORT OF THE NEEDLE-LESS ACCESS COMPONENTS HAS BEEN IMPLEMENTED TO REDUCE THE OCCURRENCE OF THESE LEAKS. THE VENDOR WAS ALSO MADE AWARE OF THIS INCIDENT AND ACTIONS ARE IN PROGRESS TO IMPROVE THE MOLDING PROCESS. | 0
3424157 | PATIENT WAS GIVEN A TRUE RESULT METER WHICH HAS BEEN GIVING THE PATIENT FALSE READINGS, THEREFORE, PUTTING HIS LIFE AT RISK DUE TO UNDER DOSING AND OVER DOSING HIMSELF ON HIS INSULIN FOR HIS BLOOD SUGAR READINGS. PATIENT WAS USING THIS METER BECAUSE IT WAS THE METER PROVIDED BY HIS INSURANCE GROUP; (B)(6). PERSON STOPPED USING THE PRODUCT ON (B)(6) 2013. THE PROBLEM STOPPED AFTER THE PERSON STOPPED USING THE PRODUCT. THE PROBLEM WAS REPORTED TO THE DOCTOR'S OFFICE, THE PHARMACY AND THE INSURANCE COMPANY. | 0
3367874 | IT WAS REPORTED THAT A PT WAS INJURED DUE TO UNEXPECTED MISPLACEMENT OF THE BIOPSY NEEDLE. THE PT REPORTEDLY UNDERWENT A COELIOSCOPY ON THE SAME DAY FOR DRAINAGE AND OVARIAN HEMOSTASIS DUE TO AN HEMOPERITONEUM. OOCYTE RETRIEVAL RESCHEDULED. PER COMMUNICATION FROM THE HOSPITAL, THERE WERE A TOTAL OF 5 PTS THAT WERE SIMILARLY AFFECTED. THIS IS PT 3 OF 5. THE FOLLOWING MDRS ARE ASSOCIATED WITH THE 5 PTS: 9612283-2013-00003; 9612283-2013-00004; 9612283-2013-00005; 9612283-2013-00006; 9612283-2013-00007. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED. UNDER EUROPEAN LAW, PT INFORMATION IS CONSIDERED CONFIDENTIAL AND WILL NOT BE RELEASED BY THE HOSPITAL. | 0
2800503 | SYSTEM WAS PUT INTO STANDBY AND WHEELED OVER TO ER FOR A SCAN, THE SYSTEM WAS PLUGGED IN AND WOULD NOT BOOT UP. CUSTOMERS HAVE BEEN NOTIFIED THAT THE SYSTEM DOES NOT START UP FROM THE BATTERY POWERED STANDBY MODE AND THAT IT WILL BE DISABLED TEMPORARILY BY A FIELD MODIFICATION INSTRUCTION. A DESIGN CHANGE IS CURRENTLY BEING CONDUCTED TO CORRECT THIS PROBLEM. WHEN IT IS COMPLETED ALL UNITS WILL BE RETROFITTED WITH THE CHANGE.><ISSUE -SYSTEM DOES NOT START FROM STANDBY MODE. IT WAS FOUND THAT THE SYSTEM HAS A LIMITATION THAT IT IS NOT POSSIBLE TO START THE POWER UNTIL APPROXIMATELY 5 SECONDS AFTER STANDBY LED LIGHTS (UNTIL THE INTERNAL CIRCUITS BECOME STABLE) AFTER THE AC PLUG IS INSERTED. IF THE POWER SWITCH IS PRESSED WITHIN THIS TIME (5 SECONDS),THE ATX UNIT JUDGES THAT THE POWER SUPPLY DOES NOT START UP AND IS ABNORMAL, AND THEN MAY FORCIBLY STOP THE POWER UNIT. THIS DOES NOT OCCUR WHEN THE POWER SWITCH IS PRESSED 15 SECONDS AFTER STANDBY LED LIGHTS PER THE OPERATION MANUAL. WE ARE IN THE PROCESS OF DETERMINING A COURSE OF ACTION TO TAKE TO FIX THESE ISSUES. | 0
2915414 | DURING AN INTERNAL INVESTIGATION BY THE MFR, IT WAS DISCOVERED THAT WHEN 'FLOW MODEL' IS SELECTED IN COLOR FLOW OR PULSED DOPPLER MODE WITH CARDIAC APPLICATION, THE ACOUSTIC OUTPUT INCREASES ABOVE ITS DESIGN SPECIFICATION AND THE PROBE SURFACE BECOMES HOT. THERE WAS NO REPORT OF INJURY. IN ADDITION TO RAPIDLY INCREASED SURFACE TEMPERATURE OF THE PROBE, THIS EVENT CAUSED THE PROBE ACOUSTIC OUTPUT TO INCREASE WELL ABOVE THE FDA LIMIT OF 720 MW/CM^2, THUS CAUSING THE DEVICE TO FAIL ITS DESIGN SPECIFICATIONS RELATING TO ULTRASOUND EMISSIONS. THIS ISSUE WAS PREVIOUSLY REPORTED FOR THE LOGIQ S8 ON MFR REPORT NO. 9710090-2012-00004. THE PRIMARY DIFFERENCE BETWEEN THE TWO SYSTEMS IS THE CHANNEL COUNT. THE LOGIQ S7 HAS 128/64 CHANNELS, AND THE LOGIQ S8 HAS 192 CHANNELS. ALL OTHER FEATURES ARE VIRTUALLY THE SAME. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED. THE SERIAL NUMBER OF THE DEVICE IS NOT AVAILABLE. THE DEVICE MFR DATE IS NOT AVAILABLE. | 0
3050567 | A CUSTOMER REPORTED THAT THE SYSTEM PRESENTED DIFFERENT MEASUREMENTS DURING THE DIAGNOSTIC TESTING. THE TESTING WAS COMPLETED WITH NO HARM TO THE PT. THE COMPANY REP EXAMINED THE SYSTEM AND COULD NOT REPLICATE THE REPORTED EVENT. THE SYSTEM WAS THEN TESTED AND MET ALL PRODUCT SPECIFICATIONS. THERE WAS NO SAMPLE RETURNED FOR EVAL AND NO ADDITIONAL INFORMATION PROVIDED RELATED TO THIS EVENT. FOR THIS REASON, STEPS COULD NOT BE TAKEN TO REPLICATE OR CONFIRM THE REPORTED EVENT. A REVIEW OF COMPLAINTS FOR THE LAST (B)(4) DID NOT INDICATE ANY ADDITIONAL SIMILAR REPORTS FOR THIS SYSTEM. THE ROOT CAUSE COULD NOT BE DETERMINED. (B)(4). | 1
3057900 | A CUSTOMER REPORTS THE SYSTEM DID NOT PERFORM THE BIOMETRY CORRECTLY AND THE MEASUREMENTS WERE NOT ACCURATE. THE EXAM WAS CANCELLED WITH NO PT HARM. THE CUSTOMER CALLED THE COMPANY REPRESENTATIVE TO ASSIST WITH TROUBLESHOOTING. THE COMPANY REPRESENTATIVE INSTRUCTED THE CUSTOMER TO RECONFIGURE THEIR TECHNIQUE OVER THE PHONE AND STATED THAT THE SYSTEM WAS FUNCTIONING PROPERLY. THE COMPANY REPRESENTATIVE INSTRUCTED THE CUSTOMER ON THE PROPER TECHNIQUE SETTINGS. THE CUSTOMER REPORTED EVENT OF A DIFFERENCE IN MEASUREMENTS BETWEEN CONTACT AND IMMERSION TECHNIQUES COULD NOT BE REPLICATED OR CONFIRMED. THE CUSTOMER DID NOT REQUEST SERVICE. THERE WAS NO SAMPLE RETURNED FOR EVALUATION AND NO ADDITIONAL INFORMATION PROVIDED RELATED TO THIS EVENT. FOR THIS REASON, STEPS COULD NOT BE TAKEN TO REPLICATE OR CONFIRM THE REPORTED EVENT. A REVIEW OF COMPLAINTS FOR THE LAST (B)(4) DID NOT INDICATE ANY ADDITIONAL SIMILAR REPORTS FOR THIS SYSTEM. THE ROOT CAUSE OF THE CUSTOMER REPORTED EVENT IS MOSTLY LIKELY DUE IMPROPER SETTINGS USED BY THE CUSTOMER. (B)(4). | 1
3349965 | A DOCTOR REPORTED THE READINGS FROM A BIOMETRY UNIT WERE INCORRECT. THERE WAS NO PATIENT HARM REPORTED. ADD'L INFO HAS BEEN REQUESTED BUT NONE HAS BEEN RECEIVED. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4). | 1
3183634 | SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR FAILED TO ALARM FOR V-RUN. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.SPACELABS IS EVALUATING THIS EVENT AND WILL FILE A FOLLOW UP REPORT WHEN OUR EVALUATION IS CONCLUDED. (B)(4): PLACEHOLDER. | 1
2916240 | CUSTOMER REPORTED THAT A PATIENT CODED BUT THE MONITOR/MODULE DID NOT ALARM.><THE PATIENT 'CODED' BUT THE MONITOR/MODULE DID NOT ALARM. SPACELABS HAS CONCLUDED OUR EVALUATION OF THIS EVENT. THE HOSPITAL DID NOT PROVIDE ANY PRINTED STRIPS SHOWING PATIENT WAVEFORMS AT THE TIME OF THE EVENT. PATIENT DATA FROM THE ICS RETROSPECTIVE REVIEW SYSTEM WAS NOT AVAILABLE FOR THE TIME OF THE EVENT. A SPACELABS FIELD SERVICE ENGINEER RESPONDED TO THE COMPLAINT. THE DEVICES WERE TESTED AT THE CUSTOMER SITE AND FOUND TO BE OPERATING IN ACCORDANCE WITH SPECIFICATIONS. SPACELABS REQUESTED THAT THE MONITOR, MODULE AND HOUSING BE RETURNED TO (B)(6) FOR EVALUATION AND ALL WERE FOUND TO BE OPERATING IN ACCORDANCE WITH SPECIFICATIONS. SPACELABS HAS CONCLUDED THAT THE REPORTED ISSUE COULD NOT BE DUPLICATED, AND THE HOSPITAL SUPPLIED NO EVIDENCE OF PRODUCT FAILURE. SPACELABS CONSIDERS THIS ISSUE IS CLOSED.><THE CUSTOMER REPORTED THAT THE MONITOR/MODULE DID NOT ALARM WHEN A PATIENT "CODED" AND THE PATIENT DIED. THE HOSPITAL DID NOT PROVIDE ANY PRINTED STRIPS SHOWING PATIENT WAVEFORMS AT THE TIME OF THE INCIDENT. DATA FROM THE ICS RETROSPECTIVE REVIEW SYSTEM WAS NOT AVAILABLE FOR THE TIME OF THE EVENT SO THERE WAS NO HISTORICAL EVENT DATA THAT COULD BE RETRIEVED AT THE CUSTOMER SITE. A SPACELABS FIELD SERVICE ENGINEER RESPONDED TO THE COMPLAINT. THE DEVICES WERE TESTED AT THE CUSTOMER SITE AND FOUND TO BE OPERATING IN ACCORDANCE WITH SPECIFICATIONS. SPACELABS REQUESTED THAT THE MONITOR, MODULE AND HOUSING ALL BE RETURNED TO (B)(4) FOR EVALUATION. SPACELABS WILL FILE A FOLLOW UP REPORT WHEN OUR EVALUATION HAS CONCLUDED. | 1
2973565 | THIS EVENT OCCURRED DURING CRYOABLATION OF A RENAL MASS. CT-HAS A "PRE-TEST" PROCESS FOR ALL PROBES AND BOTH GASES ARGON AND HELIUM WHICH IS CONDUCTED PRIOR TO EVERY PROCEDURE. ALL PROBES WERE PRE TESTED AND "PASSED". THE PROCESS CONSISTS OF PARTIALLY CREATING AN "ICE BALL" AND THEN THAWING WHICH IS DUPLICATED DURING AND POST PROCEDURE. THREE OF THE FOUR PROBES FUNCTIONED PROPERLY DURING AND POST PROCEDURE. THE FOURTH FAILED TO THAW AND THE END OF PROCEDURE. AT THAT POINT THE COMPANY WAS CONTACTED TWICE FOR TECHNICAL SUPPORT AND THE CT TEAM FOLLOWED INSTRUCTIONS GIVEN. FIRST, IT WAS REQUESTED THAT THE CONNECTION OF THE PROBE INTO THE MAIN PORTION BE SWITCHED. IT WAS DONE WITHOUT CHANGING THE ICE BALL TEST RESULTS. THE COMPANY IMMEDIATELY WAS CONTACTED ONCE AGAIN. IT WAS REQUESTED THAT THE EQUIPMENT BE TURNED OFF FOR A 'RESET' AND TURNED BACK ON AGAIN.THE ICE BALL WAS SUFFICIENT TO TREAT THE TUMOR. SINCE THE MALFUNCTIONING PROBE WAS IN THE CENTER OF THE ICE BALL WE HAD TO WAIT FOR NATURAL THAW. THERE WAS A SMALL PERINEPHRIC HEM, BUT THIS CAN BE EXPECTED. THE PT WAS ADMITTED FOR PAIN CONTROL, NOT BECAUSE OF DEVICE MALFUNCTION. THE DEVICE EVENT DID NOT HARM THE PATIENT.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?CRYOABLATION OF A RENAL MASS. | 0
3294478 | THE MACHINE WAS TESTED BEFORE THE PATIENT WAS INDUCED. THE MACHINE STOPPED WORKING AFTER IT WAS RUNNING FOR AWHILE. THE PATIENT WAS ALREADY ASLEEP WHEN EQUIPMENT FAILED. IT HAPPENED BEFORE THE SURGICAL PREP OF THE AREA. CASE ABORTED BY MD. PATIENT WAS TRANSFERRED TO PACU.WHAT WAS THE ORIGINAL INTENDED PROCEDURE?CYSTOSCOPY CRYOABLATION OF PROSTATE.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3286805 | THE VACUUM CHAMBER APPEARS TO HAVE FAILED WITHIN THE PROBE. THE ICE FROZE ALL THE WAY UP THE SHAFT. A STERILE GLOVE WITH WARM SALINE WAS USED TO AVOID FREEZE DAMAGE. NO INURY. (B)(4). EVALUATION CONFIRMED THE REPORTED ISSUE OF SHAFT FROSTING. FROSTING DUE TO A VACUUM SLEEVE FAILURE.><SHAFT FROSTING. PER RADIOLOGIST AT HOSPITAL, "USED A STERILE GLOVE WITH WARM SALINE TO AVOID FREEZE DAMAGE". CASE CONTINUED, NO INJURY TO PATIENT AND THE PATIENT IS DOING FINE. | 0
3000442 | PER PERSON MAKING THE REPORT, WHO WAS NOT PRESENT DURING THE PROCEDURE, ERROR MESSAGE - WOULD NOT WORK IN ANY PORT. PHYSICIAN ALSO HAD A QUESTION ABOUT THE "TROCAR" AND ASKED IF THE PROBE HAD BEEN CHANGED.ADDITIONAL INFORMATION RECEIVED ON (B)(4) 2013 - 2ND PHYSICIAN HELPED TO TROUBLESHOOT AND GET CASE BACK ON TRACK AND COMPLETED. 2ND PHYSICIAN FELT IT WAS A FAULTY OR DEFECTIVE PROBE.FURTHER INFORMATION RECEIVED ON (B)(4) 2013 - NO INJURY TO PATIENT, 1 HOUR DELAY, 2ND PHYSICIAN HAD THEM OPEN A FRESH PROBE TO FINISH THE CASE. NO PATIENT INFORMATION PROVIDED. MDR REPORTABILITY DETERMINATION CHANGED FROM NO MDR REQUIRED TO MDR REQUIRED BASED ON INFORMATION RECEIVED (B)(4) 2013><CORRECTED DATA - EVALUATION CODES CHANGED. ORIGINAL MDR REPORTED FAILURE WAS NOT CONFIRMED. FAILURE WAS CONFIRMED AND WAS RELATED TO THE DEVICE.INITIAL SUBMISSION CONCLUSION WAS INCORRECT. EVALUATION ON (B)(4), 2013 FOUND THE PROBE WAS CLOGGED. | 0
3329463 | WE ARE FOLLOWING UP ON A MEDICAL DEVICE CORRECTION FROM GE ON THE TEC 6 AND TEC 6 PLUS VAPORIZERS THAT WAS SENT OUT FROM GE ON THE 6 AUGUST 2013... WE HAVE THE TEC PLUS 6 IN OUR FACILITY - THE CLINICAL ENGINEERING DIRECTOR TESTED THEM AT OUR FACILITY AND 63% OR 12 OUT OF 19 HAD A LEAK (LOW PRESSURE LEAK TEST). WE FOLLOWED THE INSTRUCTIONS FROM GE HEALTHCARE: "RE: THE LOW PRESSURE LEAK TEST MAY NOT DETECT THE FULL RANGE OF LEAKS FROM THE SEAL IN THE TEC-6 AND TEC-6 PLUS VAPORIZERS. A LEAK MAY CAUSE A REDUCTION IN THE FRESH GAS VOLUME DELIVERED TO THE BREATHING SYSTEM... PERFORM THE USER MANUAL RECOMMENDED PRE-OPERATIVE CHECK WHICH INCLUDES THE LOW PRESSURE LEAK TEST ON EVERY VAPORIZER WITH THE DIAL TURNED TO THE 12% SETTING INSTEAD OF THE SPECIFIED 1% SETTING AS PER THE ATTACHED ADDENDUM...DISCONTINUE USE OF THE TEC 6 OR TEC 6 PLUS VAPORIZER IF IT DOES NOT PASS THE LOW PRESSURE LEAK AT 12%, REMOVE IT FROM SERVICE AND CONTACT A GE HEALTHCARE REPRESENTATIVE FOR A REPAIR/REPLACEMENT." | 0
3354111 | PER THE RE-CALL, LOW PRESSURE LEAK TEST WAS ASKED TO BE VALIDATED. 20/20 OF OUR PIECES FAILED. WE ARE CURRENTLY WORKING WITH GE ON SERVICE AND REPLACEMENT. ====================== MANUFACTURER RESPONSE FOR TEC 6 PLUS VAPORIZER, TEC 6 VAPORIZER (PER SITE REPORTER).====================== WE ARE WAITING FOR GE TO REPAIR/REPLACE. | 0
3315279 | IN A LETTER ISSUED (B)(4) 2013 (B)(4), GE HEALTHCARE ISSUED A DEVICE CORRECTION LETTER ON THEIR TEC 6 AND TEC 6 PLUS VAPORIZERS. IN THE LETTER IT STATED "A LEAK MAY CAUSE A REDUCTION IN THE FRESH GAS VOLUME DELIVERED TO THE BREATHING SYSTEM." AS SUGGESTED IN THE LETTER, THE TEST WAS PERFORMED ON OUR VAPORIZERS. OF THE 21 VAPORIZERS IN OUR TECH 6 INVENTORY 13 WERE FOUND TO BE BAD. THIS HAS PLACED A STRAIN ON OUR ANESTHESIA DEPARTMENT AS WE HAVE PULLED THEM FROM SERVICE. OUR RISK DEPARTMENT ASKED THAT THIS REPORT BE FILED TO DOCUMENT THIS ISSUE. | 0
3145484 | IT WAS REPORTED THAT A WIRELESS MODULE IS GENERATING A RADIO DISABLED ALERT AND RESULTED IN A NON-CURRENT MASTER DRUG LIBRARY. THIS WAS DISCOVERED IN THE MED/SURG DEPARTMENT. IT WAS ALSO REPORTED THERE WAS NO PATIENT INJURY. (B)(4). THE DEVICE WAS RECEIVED BY BAXTER AND THE BAXTER EVALUATION IS STILL IN PROGRESS. WHEN THE DEVICE EVALUATION IS COMPLETE A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
3096006 | IT WAS REPORTED A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMER'S NETWORK. THERE WAS NO PT INJURY REPORTED. (B)(4). BAXTER'S DEVICE EVALUATION IS IN PROGRESS. WHEN COMPLETE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
3257092 | IT WAS REPORTED THAT A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMERS NETWORK. IT WAS ALSO REPORTED THAT THERE WAS NO PT INVOLVEMENT. (B)(4). THE WIRELESS MODULE WAS RECEIVED AND EVALUATED BY BAXTER. THE MODULE WAS RECEIVED INOPERABLE DUE TO A "WILL NOT CONNECT" CONDITION. THE MODULE FAILED DUE TO THE RADIO PCB (SPECIFIC COMPONENT NOT IDENTIFIED) RENDERING THE UNIT UN-REPAIRABLE. THE UN-REPAIRABLE MODULE WAS REMOVED FROM SERVICE. | 0
3165851 | IT WAS REPORTED THAT A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMER'S NETWORK. IT WAS ALSO REPORTED THERE WAS NO PATIENT INJURY. (B)(4). MODULE RECEIVED INOPERABLE WITH A "POWER UP TIMEOUT" CONDITION DUE TO CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PCB AS A RESULT OF FLUID INTRUSION, WHICH CAUSED THE COMPONENTS TO FAIL, RENDERING THE UNIT UN-REPAIRABLE. THE UN-REPAIRABLE MODULE WAS REMOVED FROM SERVICE. | 0
3149169 | IT WAS REPORTED THAT A WIRELESS MODULE IS "NOT TRANSMITTING." IT WAS ALSO REPORTED THAT THERE WAS NO PT INJURY. (B)(4). THE WIRELESS MODULE WAS RECEIVED BY BAXTER. THE BAXTER DEVICE EVAL IS STILL IN PROGRESS AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON THE COMPLETION OF THE EVAL. | 0
3222084 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE WOULD NOT CONNECT TO THE CUSTOMER'S NETWORKS. IT WAS ALSO REPORTED THERE WAS NO PATIENT INVOLVEMENT. (B)(4). THE DEVICE WAS RETURNED AND EVALUATED BY BAXTER. EVALUATION FOUND UNIT INOPERABLE CAUSED BY A FAILURE ON THE RADIO PCB RENDERING THE UNIT UN-REPAIRABLE. THE UN-REPAIRABLE MODULE WILL BE REMOVED FROM SERVICE. | 0
3179065 | IT WAS REPORTED THAT A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMER'S NETWORK. IT WAS ALSO REPORTED THERE WAS NO PT INJURY. (B)(4). BAXTER'S DEVICE EVAL IS IN PROGRESS. WHEN COMPLETE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | 0
3149144 | IT WAS REPORTED THAT A WIRELESS MODULE "TURNS ON AND OFF ON ITS OWN" IT WAS ALSO REPORTED THAT THERE WAS NO PT INJURY. MFR REF NO: (B)(4). THE WIRELESS MODULE WAS RECEIVED AND EVALUATED BY BAXTER. THE REPORTED PROBLEM COULD NOT BE CONFIRMED OR REPRODUCED AND THE UNIT WAS TESTED WITH THE UNIT PASSING. ADDITIONALLY, THE MODULE WAS COUPLED WITH A KNOWN GOOD PUMP, WITH THE BAXTER MEDINA GATEWAY CONFIGURATION INSTALLED, MAKING SURE THE BATTERY AND PUMP ARE ABLE TO COMMUNICATE WITH THE BAXTER MEDINA GATEWAY AND RECEIVE THE IP ADDRESS AND GATEWAY INFO FROM THE BAXTER MEDINA NETWORK. | 0
3168179 | IT WAS REPORTED THAT A DEVICE WOULD NOT CONNECT TO THE CUSTOMER'S WIRELESS NETWORK. NO PT INJURY WAS REPORTED. MANUFACTURER REF NO.:(B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE IN QUESTION. THE DEVICE WAS CONFIRMED TO BE INOPERABLE DUE TO A "CHECK BATTERY" CONDITION. THE DEVICE TESTING DUE TO A FAILURE ON THE RADIO PRINTED CIRCUIT BOARD. THE DEVICE COULD NOT BE REPAIRED AND WAS REMOVED FROM SERVICE. | 0
3101836 | IT WAS REPORTED THAT A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMER'S WIRELESS NETWORK. IT WAS ALSO REPORTED THERE WAS NO PT INVOLVEMENT. MFR REF NO.: (B)(4). THE BATTERY MODULE WAS REC'D INOPERABLE BY BAXTER. THE DEVICE ALARMED FOR "CHECK BATTERY" DUE TO CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PCB AS A RESULT OF FLUID INTRUSION. THIS CAUSED THE COMPONENTS TO FAIL, RENDERING THE UNIT UN-REPAIRABLE. THE UN-REPAIRABLE WIRELESS MODULE WAS REMOVED FROM SERVICE. | 0
3135177 | IT WAS REPORTED THAT WIRELESS MODULE FOR A SPECTRUM INFUSION PUMP WILL NOT CONNECT TO THE NETWORK. IT WAS FURTHER REPORTED THAT THE CUSTOMER "WAS UNSURE IF AN INFUSION WAS PROGRAMMED WITHOUT THE MOST RECENT MDL AS A RESULT OF THE NETWORK CONNECTIVITY PROBLEM." IT WAS ALSO REPORTED THAT THERE WAS NO PT INJURY. (B)(4). THE DEVICE WAS RECEIVED BY BAXTER AND THE REPORTED SYMPTOM WAS CONFIRMED. THE WIRELESS MODULE WAS NOT WITHIN SPECIFICATION. THE MODULE FAILED TO PASS TESTING DUE TO A FAILURE ON THE RADIO PCB, RENDERING THE UNIT UN-REPAIRABLE. THIS UNIT WILL BE REMOVED FROM SERVICE. | 0
2977818 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE WAS "SMOKING AND IS NOW BURNT." IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY. (B)(4). BAXTER EVALUATED THE DEVICE IN QUESTION AND OBSERVED FLUID RESIDUE INSIDE OF THE DEVICE. A SHORT WAS ALSO OBSERVED ON THE WIRELESS MODULE FLEX PRINTED CIRCUIT BOARD AND FOUND TO BE CAUSED BY FLUID INTRUSION. AT THIS TIME, THE DATE OF EVENT IS UNKNOWN. | 0
3098642 | IT WAS REPORTED A WIRELESS MODULE WOULD NOT CONNECT TO THE CUSTOMER'S NETWORK. IT WAS ALSO REPORTED THAT THERE WAS NO PT INJURY. (B)(4). COULD NOT REPRODUCE, THE DEVICE MEETS SPECIFICATION FOR THE REPORTED SYMPTOM OF "WILL NOT CONNECT". THE UNIT WAS TESTED WITH UNIT PASSING. ADDITIONALLY, THE MODULE WAS COUPLED WITH A KNOWN GOOD PUMP, WITH THE BAXTER MEDINA GATEWAY CONFIGURATION INSTALLED, MAKING SURE THE BATTERY AND PUMP ARE ABLE TO COMMUNICATE WITH THE BAXTER MEDINA GATEWAY AND RECEIVED THE IP ADDRESS AND GATEWAY INFORMATION FROM THE BAXTER MEDINA NETWORK. THE DEVICE WILL BE RETURNED TO THE CUSTOMER. | 0
3486463 | DEVICE READS CONSISTENTLY HIGH BUT PT HAS NO SYMPTOMS OF HYPERGLYCEMIA. THIS HAPPENS MOSTLY WHEN THE DEVICE IS COLD. PT HAS CALLED BAYER AND A REPLACEMENT DEVICE WILL BE SENT. | 1
3032178 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE RESTARTS A PUMP. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INJURY. (B)(4). THE DEVICE WAS RECEIVED AND EVALUATED BY BAXTER. THE EVALUATION CONFIRMED THAT THE BATTERY MODULE WILL RESTART THE DEVICE DUE TO A FAILED WIRELESS FLEX ASSEMBLY. FURTHER EVALUATION DETERMINED THAT THIS CAUSED BY FLUID INTRUSION. AN INSUFFICIENT AMOUNT OF SEALANT WAS APPLIED TO THE BATTERY CASE WHICH ALLOWED THE FLUID INTRUSION. THE WIRELESS MODULE WAS REPLACED. | 0
3237733 | THE OLYMPUS WM-NP1 SERIES MOBILE WORKSTATION WITH SEPARATION TRANSFORMER IS USED IN MEDICAL FACILITIES TO ACCOMMODATE A RANGE OF OLYMPUS EQUIPMENT TO FACILITATE GI ENDOSCOPY, ENDOSCOPIC ULTRASOUND, RESPIRATORY AND SURGICAL ENDOSCOPIC PROCEDURES. THE WORKSTATION WHICH IS THE SUBJECT OF THIS REPORT WAS BEING USED TO ACCOMMODATE RECORDING EQUIPMENT DURING A LAPAROTOMY PROCEDURE. THE PLUG OF THE POWER CABLE OF THE WORKSTATION WAS CONNECTED TO A POWERSTRIP. AT THE BEGINNING OF THE PROCEDURE THE PLUG OF THE POWER CABLE OF THE WORKSTATION SMOKED. A NURSE RECONNECTED THE PLUG TO THE SOCKET ON THE WALL AND IT CAUGHT FIRE. THERE IS NO REPORT OF INJURY TO PT OR USER AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. THE CABLE HAS BEEN RETURNED TO KEYMED LTD AND A FULL EVAL IS TAKING PLACE TO DETERMINE THE ROOT CAUSE OF THE EVENT / DEVICE MALFUNCTION. A F/U REPORT WILL BE SUBMITTED ONCE THE ROOT CAUSE HAS BEEN DETERMINED. THERE IS NO REPORT OF INJURY TO PT OR USER AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. | 1
3299570 | NURSE WAS SETTING UP ROOM IN PREPARATION TO RECEIVE A PATIENT. BROUGHT COMPUTER DOWN TO HER LEVEL AND MONITOR MOUNT SNAPPED OFF. MONITOR FELL, HITTING NURSE IN FACE. NURSE SUFFERED ABRASION TO SIDE OF NOSE AND IMMEDIATE DIZZINESS. PER OTHER STAFF IN ROOM AT TIME, NO EXCESSIVE FORCE APPLIED. PART BROKEN WAS ALUMINUM, APPROXIMATELY AN INCH IN DIAMETER.DEVICE #1IS THIS A LABORATORY DEVICE OR LABORATORY TEST?NO. | 0
3480537 | THE RADIOLOGIC TECHNOLOGIST WAS POSITIONING THE DEVICE TO PERFORM AN X-RAY EXAM WHEN HE FELT SOMETHING POKE INTO HIS HAND. HE DISCOVERED THAT THE X-RAY GRID HAD SPLINTERED AND PENETRATED HIS HAND.====================== MANUFACTURER RESPONSE FOR X-RAY GRID FOR USE WITH GE FLASH PAD DETECTOR, GE SNAP - ON GRID FOR FLASH PAD DETECTOR (PER SITE REPORTER).====================== GE HAS STATED IN THE PAST THAT THEY ARE WORKING ON A RE-DESIGN WHICH WILL BE MORE DURABLE WITH LESS CHANCE OF CARBON FIBER SPLINTERING. WE HAVE REPLACED 20 + GRIDS OF THIS TYPE IN THE PAST 18 MONTHS WITH NO NEW DESIGN OFFERED TO US. STAFF HAVE EXPRESSED CONCERN OF CARBON FIBER SPLINTERS. SEVERAL EMPLOYEES HAVE RECEIVED SPLINTERS. | 0
2982730 | CERNER'S DISPLAY OF INFORMATION ON ORDER SETS IS UNNECESSARILY CONFUSING, CLUTTERED AND PRONE TO ERROR. IT IS DIFFICULT TO DETERMINE WHICH LINES OF TEXT APPLY TO WHICH LINES OF ORDER. THIS IS A SYSTEMIC PROBLEM WITH ALL CERNER ORDER SETS. I NOTICED A NEW ORDER SET FOR VIT K REVERSAL OF AN ELEVATED INR AND COULD NOT EASILY DETERMINE WHAT TEXT WENT WITH WHAT ORDER. I ASKED OUR HEAD PHARMACIST AND LEAD HOSPITALIST, AND NEITHER COULD READILY NAVIGATE THE INFORMATION ON THIS SCREEN. THIS SAME NEEDLESS CONFUSION IS PRESENT ON ORDER SETS FOR INSULIN AND FOR ANTIBIOTICS FOR PNEUMONIA, INCREASING THE PROBABILITY THAT THE PHYSICIAN WILL MISTAKENLY ORDER AN UNTENDED THERAPY AND NOT ORDER AN INTENDED THERAPY. | 1
3453136 | A SURGEON REPORTED THAT SINUS IMAGES WERE INCORRECTLY MARKED RESULTING IN THE WRONG SIDE OF THE PT BEING OPERATED ON. SOFTWARE VERSION UNK THUS MANUFACTURE DATE IS ALSO UNK. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3229519 | DURING USE OF THE MCKESSON ANESTHESIA-RX DEVICE (MEDICATION DISPENSING CART), AN "UNHANDLED EXCEPTION ERROR" WAS RECEIVED BY THE USER. WHEN THE ERROR OCCURRED, THE SOFTWARE APPLICATION ON THE ANESTHESIA-RX DEVICE BECAME TEMPORARILY UNAVAILABLE WHICH COULD RESULT IN THE DELAY OF ACCESS TO MEDICATIONS FOR A PT, AS REPORTED BY THE USER, AND IF RECOMMENDED BACK-UP PROTOCOLS ARE NOT FOLLOWED OR CAUSE AN EXCESSIVE DELAY. NO PT INJURIES WERE REPORTED AS A RESULT OF THIS MALFUNCTION. THE MANUFACTURER INVESTIGATION DETERMINED THAT THE CONFIGURATION STANDARD DEFAULT VALUE OF 100 SQL CONNECTIONS WAS EXCEEDED BY THE USER AS WELL AS A SERIAL THREADING ISSUE ASSOCIATED WITH SCANNING AND PRINTING FUNCTIONS WHICH RESULTED IN THE ASSOCIATED ERRORS REPORTED. THE ISSUES CAUSING THE ERRORS WERE RESOLVED BY INCREASING THE SQL CONNECTIONS TO 500 AND THEN BY INSTALLING A SOFTWARE UPDATE (UPGRADE 1.5). THE VERSIONS OF SOFTWARE DETERMINED TO BE POTENTIALLY AFFECTED BY ONE OR BOTH OF THE ROOT CAUSES IDENTIFIED ARE 8.0.1 AND 8.0.2. ALTHOUGH THE POTENTIAL FOR PT HARM HAS BEEN IDENTIFIED TO BE REMOTE, MCKESSON AUTOMATION WILL BE PROVIDING A SOLUTION FOR ALL IDENTIFIED POTENTIALLY AFFECTED CUSTOMERS TO PREVENT A RECURRENCE. | 1
3527990 | THE PACS ADMINISTRATOR FROM THE REPORTING FACILITY ALLEGED THAT A TECHNOLOGIST INADVERTENTLY MARKED A STUDY AS "REPORTED", AND THAT THE RADIOLOGIST DID NOT RECEIVE THE STUDY CORRECTLY MARKED FOR REQUIRING A RADIOLOGY REVIEW UNTIL 3 DAYS LATER. UPON FURTHER REVIEW, THE RADIOLOGIST SUBSEQUENTLY LEARNED THAT THE PT HAD EXPIRED. MCKESSON HAS CONTACTED THE REPORTING FACILITY TO OBTAIN FURTHER INFORMATION IN ITS INVESTIGATION. HORIZON MEDICAL IMAGING ALLOWS THE STATUS OF A STUDY TO BE MARKED AS "PERFORMED", "REVIEWED", NEEDS OVER-READ", "DICTATED", OR "REPORTED". USERS AT THE REPORTING FACILITY ARE GRANTED PERMISSIONS TO CHANGE THE STATUS OF A STUDY BASED ON ESTABLISHED CONFIGURATION SETTINGS. UNDER NORMAL OPERATING CONDITIONS A TECHNOLOGIST AT THE REPORTING FACILITY HAS PERMISSIONS TO CHANGE THE STATUS OF STUDIES FROM "PERFORMED" TO "REVIEWED". THE REPORTING FACILITY INDICATED THAT TECHNOLOGISTS ARE TRAINED ONLY TO MARK STUDIES AS "REVIEWED". WHEN THE (B)(6) MEDICAL IMAGING AUTHORIZATION SERVICE USED TO MANAGE USER PERMISSIONS IS INOPERATIVE OR INACCESSIBLE, THE SYSTEM IS DESIGNED TO ENSURE HIGH AVAILABILITY OF CRITICAL FUNCTIONS TO ENABLE CARE PROVIDERS TO ACCESS STUDIES. UNDER THESE CONDITIONS, THE SYSTEM GRANTS FULL PERMISSIONS TO ALL USERS TO ENSURE THAT CRITICAL FEATURES REMAIN AVAILABLE. WHEN OPERATING IN THIS MODE, THE USERS ARE PROVIDED WITH MENU OPTIONS TO MARK THE STUDY AS "REVIEWED", "NEEDS OVER-READ", "DICTATED", OR "REPORTED". BASED ON OUR INVESTIGATION, WHICH INCLUDED A REVIEW OF THE SYSTEM ACTIVITY LOGS FROM THE REPORTING FACILITY, THE TECHNOLOGIST WAS GRANTED FULL PERMISSIONS, AS THE AUTHORIZATION PROCESS WAS INACCESSIBLE DUE TO A SYSTEM COMMUNICATION ISSUE. THE SYSTEM GRANTED THOSE FULL PERMISSIONS TO THE TECHNOLOGIST, AS DESIGNED. THE TECHNOLOGIST INADVERTENTLY SELECTED THE INCORRECT STATUS FOR THE PT STUDY, MARKING IT AS "REPORTED" INSTEAD OF "REVIEWED", RESULTING IN AN ALLEGED DELAY SENDING THE STUDY TO THE RADIOLOGIST TO READ. FURTHER INVESTIGATION IS ONGOING. | 1
3055572 | IT WAS LEARNED FROM ONE CUSTOMER SITE OF AN OCCURRENCE WHERE THE DATA DISPLAYED DID NOT MATCH THE ACTUAL PATIENT DATA WHEN THE CASE WAS RECALLED IN THE ANESTHESIA CARE RECORD (ACR) AS THE DISPLAYED DATA INCLUDED DATA FROM ANOTHER CASE. THERE WERE NO ADVERSE EVENTS OR PATIENT INJURIES RESULTANT FROM THE REPORTED ALLEGED MALFUNCTION. ALTHOUGH THE COMPANY WAS NOT ABLE TO DUPLICATE THE ALLEGED PROBLEM, FURTHER INVESTIGATION AND ANALYSIS THROUGH SIMULATED TESTING DETERMINED THAT THE MALFUNCTION IS LIKELY THE RESULT OF ABNORMAL LATENCY DELAYS EXPERIENCED DUE TO THE CUSTOMER'S DELAYED WORKSTATION AND NETWORK CONNECTIVELY DURING PATIENT CASE DATA UPDATES. IT WAS DETERMINED THAT THE USER OPENED A NEW PATIENT CASE RECORD ON THE SAME WORKSTATION BEFORE THE PRIOR RECORD UPDATE WAS COMPLETED, THEREBY RESULTING IN THE NEWLY OPENED CASE DISPLAYING INFORMATION FROM THE PREVIOUS PATIENT CASE BEING UPDATED. THE ANALYSIS AND TESTING DETERMINED THAT THE ALLEGED EVENT IS ISOLATED AND UNLIKELY TO OCCUR WITH ANY POTENTIAL PATIENT INJURY TO BE REMOTE, IF NOT IMPROBABLE. AS A PRECAUTIONARY MEASURE, THE COMPANY IS ADVISING ALL CUSTOMERS POTENTIALLY AFFECTED BY THE MALFUNCTION AND WILL PROVIDE A MITIGATIVE SOFTWARE UPDATE. | 1

## 2014
4131209 | CONSUMER REPORTED THAT DUE TO HOME BLOOD PRESSURE MONITOR READING LOW HER HUSBAND HAD A STROKE. CONSUMER REPORTED HUSBAND WAS SITTING ON CHAIR AND ARM WAS RESTING ON CHEST. CONSUMER STATED HER HUSBAND IS TAKING DIFFERENT MEDICATIONS: HYDROCHLOROTHIAZIDE, INVOKANA, SYNTHROID, COZAAR, ZOCOR. CONSUMER CLAIMED BATTERIES WERE REPLACED 2 WEEKS AGO. CONSUMER CLAIMED HER HUSBAND SPENT 4 DAYS IN THE HOSPITAL. PER CONSUMER, THERE WAS A COMPARISON DONE AT THE HOSPITAL WITH HOME UNIT AND HOME UNIT WAS READING LOW. THEY DID NOT KNOW WHAT NORMAL BLOOD PRESSURE (BP) READINGS IS. CONSUMER REPORTED HAS THREE SHEETS OF READINGS HER HUSBAND TOOK WITH HOME BP MONITOR. CONSUMER HAS PURCHASED AUTOMATIC ARM BP MONITOR (BP710) SINCE CANNOT BE WITHOUT THE BP MONITOR. CONSUMER REFUSED TO SEND THE UNIT BACK TO FURTHER INVESTIGATION SINCE THIS IS HER PROOF THIS HAPPENED TO HER HUSBAND. CUSTOMER SERVICE REP EXPLAINED THE PROPER PROCEDURE TO FOLLOW WHEN TAKING BLOOD PRESSURE MEASUREMENT. DURING FOLLOW UP WITH QUALITY REP, END USER'S WIFE CLAIMED CUSTOMER CLAIMING LOW READINGS FOR 2 MONTHS WHILE USING THE UNIT. APPROXIMATE READINGS AROUND 148/XX AND AT THE HOSPITAL HIS READINGS WERE 207/60. PT WAS ADVISED BY DOCTOR TO PURCHASE UNIT TO MONITOR BP FROM (B)(6) 2014. CUSTOMER STARTED ACTING WEIRD AND INCOHERENT WHICH PROMPTED TAKING HIM TO THE HOSPITAL. WHEN CUSTOMER GOT TO HOSPITAL THE DIAGNOSIS WAS A BROKEN BLOOD VESSEL. CUSTOMER CLAIMING 20-50 POINT DIFFERENCE IN READINGS BETWEEN OMRON UNIT AND HOSPITAL BP UNIT READINGS. AT THE HOSPITAL, HIS BLOOD PRESSURE WAS 160 AND DOCTOR HAD TROUBLE BRINGING IT DOWN. AFTER CHANGING MEDICATION IT WAS DOWN TO AROUND 130 AT THE TIME HE WAS RELEASED FROM THE HOSPITAL. CUSTOMER WRIST SIZE IS 7 3/4. CUSTOMER BOUGHT NEW UNIT (BP710) AND CURRENTLY GETTING READINGS AROUND 144/XX. QUALITY ANALYSIS REMINDED CONSUMER THAT THE HOME MONITOR IS NOT A DIAGNOSTIC DEVICE AND TAKING MEDICATION AS PRESCRIBED BY THE DOCTOR. CUSTOMER CLAIMED MEDICATION WAS NOT ADJUSTED BASED ON HOME UNIT BP READINGS. CUSTOMER AGREED TO SEND THE UNIT BACK FOR INSPECTION. A POSTAGE PAID LABEL HAS BEEN SENT FOR RETRIEVAL OF THE HOME UNIT FOR INSPECTION. QUALITY REP CALLED CUSTOMER AGAIN ON (B)(6) 2014 TO GET AN UPDATE ON SHIPPING UNIT BACK FOR INVESTIGATION. UNIT HAS NOT BEEN RECEIVED BACK AS OF (B)(6) 2014.><(B)(4). | 0
4116853 | MULTIPLE REPORTS OF POWER SONIC BATTERY PACK ON RUBBERMAID MED SOLUTIONS WORK STATION ON WHEELS OVERHEATING, RESULTING IN NOXIOUS SMELL ON UNIT. BATTERIES ARE BEING REPLACED ON UNITS AS THIS SMELL HAS OCCURRED SEVERAL TIMES IN FACILITY.====================== MANUFACTURER RESPONSE FOR WORKSTATION ON WHEELS, RUBBERMAID (PER SITE REPORTER)======================EFFORTS ARE UNDERWAY TO REPLACE BATTERIES ON ALL UNITS. | 0
3782754 | (B)(4). THE DEVICE WAS EVALUATED BY BAXTER. THE EVALUATION FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE DUE TO FLUID ON THE INTERNAL COMPONENTS. THE WIRELESS BATTERY MODULE FLEX ASSEMBLY WAS REPLACED WITH A NEW WIRELESS BATTERY MODULE FLEX ASSEMBLY WHICH ALLEVIATED THE SYMPTOM.><THE CUSTOMER REPORTED THAT THE SPECTRUM WIRELESS MODULE HAS FLUID DAMAGE". NO PATIENT INJURY WAS REPORTED. NO FURTHER INFORMATION WAS AVAILABLE." | 0
3561847 | MFR REF NO.: (B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVAL WAS PERFORMED. THE EVAL FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE. THE WIRELESS BATTERY MODULE WAS PAIRED WITH BOTH A KNOWN GOOD SPECTRUM PUMP WITH SOFTWARE VERSION 6.05.11 INSTALLED AND A KNOWN GOOD SPECTRUM PUMP WITH SOFTWARE VERSION 6.02.06 INSTALLED WHICH FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE DUE TO A CHECK BATTERY" MESSAGE. FURTHER INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION WHICH CAUSED THE COMPONENTS TO FAIL | 0
4336389 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE G SUFFERED FLUID INTRUSION. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INVOLVEMENT.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED AND REPRODUCED THE REPORTED FLUID INTRUSION. CORROSION WAS FOUND ON THE RADIO PRINTED CIRCUIT BOARD, RENDERING THE DEVICE UN-REPAIRABLE. THE DEVICE WAS RETIRED FROM SERVICE. | 0
3779454 | IT WAS REPORTED THAT A WIRELESS MODULE HAS SUFFERED FLUID INTRUSION. IT WAS ALSO REPORTED THERE WAS NO PATIENT INJURY.><(B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION FOUND THE MODULE'S WIRELESS CONNECTION CAPABILITIES INOPERABLE DUE TO CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PCB. AS A RESULT OF THE FLUID INTRUSION, THE COMPONENTS HAVE FAILED, RENDERING THE UNIT UNSERVICEABLE DUE TO THE MAC ADDRESS ASSIGNED AS THE PRODUCT SERIAL NUMBER. | 0
3934095 | IT WAS REPORTED THAT A WIRELESS BATTERY MODULE EXPERIENCED FLUID INTRUSION. THERE WAS NO REPORTED PATIENT INCIDENT.><MANUFACTURER REF NO.: (B)(4). THE DEVICE WAS RETURNED TO BAXTER AN AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED AND REPRODUCED THE REPORTED SYMPTOM FLUID INTRUSION. INVESTIGATION REVEALED CORROSION ON THE WIRELESS MODULE FLEX AS A RESULT OF FLUID INTRUSION. THE WIRELESS BATTERY MODULE FLEX ASSEMBLY WAS REPLACED. | 0
3922347 | MANUFACTURER REF NO.: (B)(4). THE DEVICE WAS RETURNED TO BAXTER AND AN EVALUATION WAS PERFORMED. THE EVALUATION CONFIRMED AND REPRODUCED THE REPORTED SYMPTOM OF WATER DAMAGE. EVALUATION FOUND CORROSION ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION, WHICH CAUSED THE COMPONENTS TO FAIL. THE DEVICE HAS BEEN RETIRED FROM SERVICE.><IT WAS REPORTED THAT A WIRELESS MODULE HAS WATER DAMAGE. ANY PATIENT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNKNOWN. | 0
3700481 | MANUFACTURER REF NO. : (B)(4). THE DEVICE WAS NOT IDENTIFIED OR RETURNED TO BAXTER FOR EVALUATION. THEREFORE, AN EVALUATION COULD NOT BE COMPLETED. IF THE DEVICE IS IDENTIFIED AND RETURNED, AN EVALUATION WILL BE COMPLETED AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><IT WAS REPORTED THAT AN UNSPECIFIED NUMBER OF SPECTRUM WIRELESS MODULES HAVE FLUID INTRUSION. ANY PATIENT INVOLVEMENT, INJURY OR MEDICAL INTERVENTION IS UNKNOWN. | 0
3577247 | (B)(4). BAXTER RECEIVED AND EVALUATED THE DEVICE. THE REPORTED SYMPTOM FLUID INTRUSION WAS CONFIRMED. CORROSION WAS FOUND ON THE WIRELESS MODULE FLEX AND RADIO PRINTED CIRCUIT BOARD AS A RESULT OF FLUID INTRUSION. IT CAUSED THE COMPONENTS TO FAIL RENDERING THE UNIT NOT REPAIRABLE.><IT WAS REPORTED THAT A WIRELESS BATTERY MODULE HAD FLUID INTRUSION. THERE WAS NO PT INVOLVEMENT. | 0
3560546 | (B)(4) BAXTER RECEIVED AND EVALUATED THE DEVICE. IT WAS FOUND OUT OF SPECIFICATION WITH RESPECT TO THE REPORTED SYMPTOM OF FLUID INTRUSION. THIS CAUSED THE BATTERY TO BE UNABLE TO CONNECT TO THE WIRELESS NETWORK. THE DEVICE WAS FOUND TO NOT BE REPAIRABLE AND WILL BE RETIRED FROM SERVICE.><IT WAS REPORTED THAT A SPECTRUM WIRELESS BATTERY MODULE HAD FLUID INTRUSION. IT WAS ALSO REPORTED THAT THERE WAS NO PATIENT INVOLVEMENT. | 0
3779293 | IT WAS REPORTED THAT THE DEVICE HAD FLUID INTRUSION. THE CUSTOMER HAS STATED THAT THERE WAS NO PT INJURY.><(B)(4). THE DEVICE WAS RETURNED AND EVALUATED BY BAXTER. BAXTER EVAL COULD NOT REPRODUCE THE REPORTED SYMPTOM OF FLUID INTRUSION" IN THE WIRELESS DEVICE. EVAL FOUND THAT THE WIRELESS MODULE CONNECTION CAPABILITIES ARE INOPERABLE DUE TO A "CHECK BATTERY (1)" CONDITION CAUSED BY CORROSION ON THE WIRELESS MODULE FLEX AND FAILURE OF THE RADIO PRINTER CIRCUIT BOARD RENDERING THE UNIT UN-REPAIRABLE. THE DEVICE IS UN-REPAIRABLE AND WAS REMOVED FROM SERVICES." | 0
4308887 | ONSITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THE EQUIPMENT PERFORMED TO SPECIFICATIONS. TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. DURING TESTING, IT WAS NOTED THAT THE TELEMETRY RECEIVER MODULE WAS CONFIGURED WITH THE MULTIVIEW I OPTION (OPTION ¿B¿). PER THE PRODUCT¿S DATA SHEET, THE MULTIVIEW I OPTION PROVIDES ALARMS FOR ABNORMAL BEATS PER MINUTE, HIGH AND LOW HEART RATE, ASYSTOLE, VENTRICULAR FIBRILLATION, VENTRICULAR RUNS, AND COUPLETS. IT DOES NOT PROVIDE AN ALARM FOR SUPRAVENTRICULAR TACHYCARDIA (SVT). THERE WAS NO EQUIPMENT MALFUNCTION. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED. PLACEHOLDER.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITORED WITH TRANSMITTER MODEL 91343, RECEIVER MODULE MODEL 90478 AND CENTRAL MONITOR MODEL 91387-38 DID NOT GENERATE ALARMS FOR SUPRAVENTRICULAR TACHYCARDIA (SVT) ON (B)(6) 2014 FROM 7:30 TO 10:00 A.M. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITOR FAILED TO ALARM ON SUPRAVENTRICULAR TACHYCARDIA(SVT). NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><ONSITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THE EQUIPMENT WORKED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER.SPACELABS IS EVALUATING THIS REPORTED EVENT AND WILL FILE A FOLLOW UP REPORT WHEN OUR EVALUATION IS CONCLUDED. PLACEHOLDER. | 0
4020024 | ><INFANT ON BILISOFT PHOTOTHERAPY BLANKET, LARGE PAD. PAD COVER HAS TIES ON THE BACK TO HELP KEEP BLANKET AROUND INFANT. THIS IS A SMALL, WIGGLY INFANT. AFTER A LAB DRAW, RN WENT TO REWRAP INFANT AND NOTICED THE TIES OF THE BILIPAD COVER WERE AROUND THE INFANT'S NECK. RN QUICKLY UNTIED AND TUCKED TIES BEHIND THE PAD. INFANT THEN SWADDLED IN BILIBLANKET WITH JUST SWADDLING BLANKETS. PARENTS CAME INTO NURSERY SHORTLY THEREAFTER TO PICK UP INFANT AND WERE NOTIFIED OF THE EVENT AND TO NOT USE THE TIE STRAPS. PARENTS UNDERSTANDING. INFANT HAD BEEN WIGGLING AND CRYING, BUT WAS NEVER COMPROMISED OXYGEN WISE. COLOR WAS PINK AND INFANT ACTIVE. THIS COULD HAVE BEEN A WORSE SITUATION. RECOMMEND TO NOT USE STRAPS AND TO CUT THEM OFF OF BILIBLANKET PAD COVERS FOR FUTURE USE. | 0
3673309 | ><CRYO CS UNIT TURNED ON AND BOOTED UP WITHOUT ISSUE. GASES HOOKED UP AND TURNED ON. PATIENT BROUGHT INTO ROOM AND PLACED UNDER ANESTHESIA. CRYO PROBES OPENED AND HOOKED UP TO THE CS. PRETEST FUNCTION USED, NO GAS FLOW. TURNED OFF GASES AND CS, REBOOTED MACHINE, AND TRIED PRETEST ONCE AGAIN WITHOUT WORKING. INFORMED DOCTOR OF ISSUE AND CALLED SERVICE HOTLINE. WAS ADVISED OF STEPS TO FOLLOW. ONCE COMPLETED, REBOOTED CS, TURNED GASES ON AND RETRIED TEST CYCLE. THIS TIME IT WORKED AND THE GASES FLOWED. THE CASE WAS CONTINUED FOLLOWING THE 40 MINUTE DELAY FOR TROUBLE SHOOTING. | 0
3824019 | ><AFTER PERFORMING A RENAL ULTRASOUND STUDY IT WAS REPORTED THAT THE LOGIQ E9 SYSTEM MISCALCULATED THE RESISTIVE INDEX (RI) TO BE LOWER THAN THE ACTUAL VALUE. THIS OCCURRED ON SEPARATE OCCASIONS WITH TWO PT STUDIES. IN BOTH CASES, THE USER WAS STORING IMAGES AND THE SYSTEM INTERMITTENTLY AND RANDOMLY DISPLAYED THE MISCALCULATED RI WITHIN THE IMAGE. NO SERIOUS INJURY WAS REPORTED. | 0
4272081 | INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4).><A CUSTOMER REPORTED THAT THE CALIPER RULE PRESENTED MEASURES THAT LEFT THE DOCTOR SUSPICIOUS OF THE BIOMETRY RESULTS. THE EXAM WAS COMPLETED USING THE SAME EQUIPMENT. ADD'L INFO HAS BEEN REQUESTED. | 0
3815630 | THE SERVICE RECORD STATES THAT THE CUSTOMER CANCELLED THE REQUEST. THERE WAS NO SAMPLE RETURNED FOR EVALUATION AND NO ADDITIONAL INFORMATION PROVIDED RELATED TO THIS EVENT. FOR THIS REASON, STEPS COULD NOT BE TAKEN TO REPLICATE OR CONFIRM THE REPORTED EVENT. THE ROOT CAUSE CANNOT BE DETERMINED CONCLUSIVELY. (B)(4).><A CUSTOMER REPORTED MEASURES ARE PRESENTING VARIATION. ADDITIONAL INFORMATION HAS BEEN REQUESTED. | 0
4172065 | A SERVICE VISIT WAS PERFORMED. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT RECEIVED TO DATE. (B)(4).><A SURGEON REPORTED THAT A FEW MINUTES AFTER THE MEASUREMENTS, THE SCREEN OF THE SYSTEM STARTED TO FLICKER FOR 4 PATIENTS. AFTER THE INTRAOCULAR LENS (IOL) IMPLANT SURGERY THESE PATIENTS HAD 1 TO 3 DIOPTERS OF RESIDUAL MYOPIA. HE ALSO REPORTED THAT THE EVENT WILL BE PERMANENT FOR THIS PATIENT SINCE NO ADDITIONAL SURGERY WILL BE PERFORMED. THE AFFECTED EYE IS THE FOR THE RIGHT EYE (OD). IN SURGEON'S OPINION THE DEVICE CONTRIBUTED TO THE EVENT BECAUSE THE CALCULATED LENS STRENGTH WAS USED AS PROVIDED BY THE SYSTEM. THE SURGEON ALSO REPORTED THAT THE EVENT WAS NOTED BY THE FIRST TIME 1 DAY AFTER SURGERY. ADDITIONAL INFORMATION HAS BEEN REQUESTED BUT NOT RECEIVED TO DATE. THIS IS ONE OF FIVE MEDICAL DEVICE REPORTS BEING FILLED FOR THIS EVENT. THIS REPORT IS FOR THE THIRD PATIENT. | 0
3997795 | INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADD'L REPORTABLE INFO BECOMES AVAILABLE. (B)(4).><A CUSTOMER REPORTED THE SYSTEM WAS NOT TAKING READINGS CORRECTLY. THERE WAS NO PT HARM REPORTED. | 0
4359782 | ><THE HOSPITAL REPORTED THAT THEY INTERNALLY CONFIRMED ON (B)(6) 2014 THAT THE VELOCITIES ARE DIFFERENT BETWEEN THE 4C-D AND RM6C PROBES WHEN USED ON VOLUSON E8 K8E10807 WITH ANOTHER PATIENT, AND THEY CONFIRMED THAT PERFORMING THE SAME MEASUREMENTS USING THE SAME PROBES ON ANOTHER VOLUSON E8 ULTRASOUND SYSTEM DID NOT HAVE THE SAME QUANTITY OF DIFFERENCE. IN OTHER WORDS ,THE SAME MEASUREMENT USING THE TWO DIFFERENT PROBES ON ANOTHER ULTRASOUND SYSTEM PRODUCED COMPARABLE RESULTS, WHILE ON SYSTEM K8E10807 THE MEASUREMENT DIFFERENCE WAS BEYOND WHAT WAS TYPICALLY EXPECTED. THEREFORE, THEY SUSPECTED THE VOLUSON E8 K8E10807 OF PRODUCING INCORRECT MCA DOPPLER VELOCITY MEASUREMENTS WHEN USING THE RM6C PROBE. RETROSPECTIVELY, THE MCA VELOCITY MEASUREMENT (PERFORMED ON (B)(6)2014) WAS SUSPECTED TO BE FALSE, IN WHICH CASE THE C-SECTION MAY HAVE BEEN UNNECESSARY. | 0
4200514 | INVESTIGATION: TWO TRIMA ACCEL SETS WERE RECEIVED FOR INVESTIGATION. EACH KIT WAS LEAK TESTED AND A LEAK WAS FOUND AT THE CONTROL Y CONNECTOR OF EACH SET. SPECIFICALLY, THE LEAK SOURCE WAS AT THE KNIT LINE OF THE RBC PORT OF THE CONTROL Y. THE CUSTOMER ALSO SENT PHOTOGRAPHIC EVIDENCE OF THE BLOOD SPRAY CONFIRMING THE SOURCE OF THE LEAK. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THE LEAK IN THE CONTROL Y PARTS IS DUE TO A LEAK PATH AT THE KNIT LINE AND IS LOCATED ON THE FIRST SHOULDER OF THE CONTROL Y NEAR THE RBC PORT. THE KNIT LINE IS CREATED AS A NORMAL PART OF THE MOLDING PROCESS. IN THIS CASE, THE INJECTION PROCESS WAS SHUT OFF PREMATURELY AND THE MATERIAL DID NOT COMPLETELY FILL ALONG THE KNIT LINE AND A VOID OR GAP RESULTED. DURING CENTRIFUGATE THE CONTROL Y IS SUBJECT TO HIGH INTERNAL PRESSURE AND AN OPENING CAN FORM ALONG THE KNIT LINE WHICH CAN ALLOW FOR A LEAK PATH. THIS TYPE OF LEAK IS A VERY FINE SPRAY WHICH MAY NOT SIGNAL THE CENTRIFUGE'S LEAK DETECTOR TO ALARM. CORRECTIVE ACTION: THE ROOT CAUSE HAS BEEN IDENTIFIED AND AN IMMEDIATE CORRECTIVE ACTION HAS BEEN IMPLEMENTED AT THE SUPPLIER. THE CORRECTIVE ACTION IS TO INCREASE THE MAXIMUM TIME ALLOWED FOR FILLING OF THE MOLD TO ENSURE AN ADEQUATE RESIN INJECTION TIME. THIS ALLOWS FOR A COMPLETE KNIT LINE OF THE PART.><THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDRS TO IDENTIFY RECORDS IN WHICH AN EVENT OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED.><THE CUSTOMER REPORTED THAT THEY EXPERIENCED A LEAK ON 2 DISPOSABLE SETS. THE BLOOD LEAK WAS FOUND INSIDE THE CENTRIFUGE. THEY DID NOT RECEIVE AN ALARM. NO MEDICAL INTERVENTION WAS NECESSARY. DUE TO EU PERSONAL DATA PROTECTION LAWS, THE PATIENT INFORMATION IS NOT AVAILABLE FROM THE CUSTOMER. THIS REPORT IS BEING FILED IN RESPONSE TO THE CUSTOMER FILING A SAE REPORT WITH THEIR LOCAL AUTHORITIES. | 0
4189351 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WITHOUT A LEUKOREDUCTION (LR) FILTER WAS RECEIVED FOR INVESTIGATION. NO PRODUCT HAD BEEN COLLECTED IN THE COLLECTION BAGS WHICH WERE ALL PRESENT ON THE SET. UPON VISUAL INSPECTION NO MISASSEMBLIES WERE NOTED WITH THE SET. NO OCCLUSIONS WERE NOTED AND FLUID MOVED EASILY WITHIN THE SET. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER- THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. HOWEVER, IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT THE NUMEROUS APS PRESSURE ALARMS THROUGHOUT THE PROCEDURE, WHICH WILL CAUSE THE PUMPS TO PAUSE, MAY HAVE CONTRIBUTED TO ADDITIONAL CELLS BEING RELEASED FROM THE CHAMBER DURING THE COLLECTION. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE MAY BE DONOR RELATED. | 0
4225031 | INVESTIGATION: THE DISPOSABLE SET WAS RETURNED FOR INVESTIGATION. UPON INSPECTION, RED CELLS WERE NOTED IN THE PLATELET LINE. PAS WAS ADDED. RED BLOOD CELL CENTRIFUGE CLAMP WAS FOUND TO BE NOT CORRECTLY OCCLUDING LINE. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. REVIEW OF THE RUN DATA DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER THAN EXPECT WHITE CELL COUNT OF THE PLATELET PRODUCT. HOWEVER, IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT WBCS MAY HAVE ESCAPED THE LRS CHAMBER LATE IN THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE MAY BE DONOR RELATED. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT. | 0
4074486 | INVESTIGATION: THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE GREATER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE RUN DATA FILE ANALYSIS DID NOT SHOW A CONCLUSIVE ROOT CAUSE FOR THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. HOWEVER, IT IS POSSIBLE, THOUGH NOT CONCLUSIVE, THAT WBCS MAY HAVE ESCAPED THE LRS CHAMBER LATE IN THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT IS POSSIBLE THAT THIS LEUKOREDUCTION FAILURE MAY BE DONOR RELATED. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4301813 | INVESTIGATION: THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. SIGNALS IN THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE REPORTED RESULTS. INVESTIGATION IS IN-PROGRESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. SIGNALS IN THE RUN DATA FILE DO NOT INDICATE A CONCLUSIVE CAUSE FOR THE GREATER THAN EXPECTED RWBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION, THAT HAS THE POTENTIAL FOR INJURY. | 0
3734165 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4251453 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WHITE BLOOD CELL (RWBC) TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE REPORTED RESULTS. | 0
3526469 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: THIS DISPOSABLE WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS, OR CORRECTIVE AND PREVENTIVE ACTIONS BY TERUMO BCT QUALITY ASSURANCE. THE ANALYSIS OF THE RUN DATA FILE INDICATES ITIS POSSIBLE THAT THE PLASMA LINE MAY HAVE BEEN OCCLUDED NEAR THE END OF THE PROCEDURE. IF THE PLASMA LINE DOES NOT CONTRIBUTE PROPERLY TO THE PLATELET PUMP IT IS POSSIBLE THE FLOW THROUGH THE LRS CHAMBER COULD BE HIGHER THAN EXPECTED BY THE SYSTEM. THIS COULD ALLOW SOME WBCS TO ESCAPE AND CONTRIBUTE TO THE HIGHER THAN EXPECTED WBC CONTENT REPORTED FOR THIS COLLECTION. ORIENTATION OF THE HEX IN THE HEX HOLDER MAY CONTRIBUTE TO THE ABOVE.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4358576 | THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDR'S TO IDENTIFY RECORDS IN WHICH AN EVENT OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED APPROPRIATELY IN THE MDR FORM. THIS SUPPLEMENT IS BEING FILED TO MODIFY INFORMATION IN AND TO ALIGN WITH THE REPORTED EVENT.><ADDITIONAL INFORMATION: THE APHERESIS MANAGER WAS CONTACTED TO DETERMINE IF THE FACILITY WANTED RE-TRAINING, REGARDING TECHNIQUES TO OPTIMIZE THE VENIPUNCTURE PROCEDURE. THE CUSTOMER DECLINED THE RE-TRAINING.><INVESTIGATION: PER THE CUSTOMER, THEY HAD DIFFICULTIES WITH THE VENIPUNCTURE AND NOTICED SLOW BLOOD FLOW DURING THE PROCEDURE. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER REPORTED THAT DURING THE FIRST RETURN CYCLE OF A TRIPLE PLATELET COLLECTION PROCEDURE, THEY RECEIVED MULTIPLE ALARMS INCLUDING 'SYSTEM TEST FAILURE' ALARM. WHILE THE OPERATOR DISCONNECTED AND REMOVED THE NEEDLE FROM THE DONOR'S ARM, THEY NOTED A 'LONG' BLOOD CLOT REMOVED WITH THE NEEDLE. NO MEDICAL INTERVENTION WAS NECESSARY FOR THIS EVENT AND NO ADDITIONAL FOLLOW-UP VISIT WAS REQUIRED. PATIENT (DONOR) FULL IDENTIFIER: (B)(6). THE DISPOSABLE KIT IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO THE POTENTIAL FOR INJURY IF THE SAME FAILURE WERE TO REOCCUR.><INVESTIGATION: A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THERE WERE NO INDICATIONS IN THE FILE FOR WHY THERE WOULD HAVE BEEN CLOTTING IN THE RETURN LINE. THE TRIMA SYSTEM OPERATED AS INTENDED. ROOT CAUSE: THE SET WAS NOT AVAILABLE FOR RETURN AND ANALYSIS. THE RDF ANALYSIS SHOWS MULTIPLE LOW APS PRESSURE ALERTS DURING THE PROCEDURE. ADDITIONALLY, CUSTOMER STATEMENTS INDICATE THAT THE CAUSE FOR THE BLOOD CLOT, ALERTS, AND ALARM ARE RELATED TO SLOW BLOOD FLOW FROM POOR VENIPUNCTURE. | 0
3449913 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED. | 0
3853367 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(4). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
3632761 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER TOWARD THE END OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKOREDUCTION FAILURE COULD BE DONOR-RELATED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
3645693 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. IT IS POSSIBLE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT THE HIGH NUMBER OF ACCESS ALERTS AND ADJUSTMENTS THAT OCCURRED THROUGHOUT THE PROCEDURE DISRUPTED THE STEADY-STATE OF THE SYSTEM AND CONTRIBUTED TO THE ELEVATED WBC CONTENT MEASURED IN THE PLATELET PRODUCT. IT IS ALSO POSSIBLE THAT A SAMPLING, CALCULATION OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
3410428 | INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. IT IS POSSIBLE THAT THE HIGH NUMBER OF ACCESS ALERTS AND ADJUSTMENTS THAT OCCURRED THROUGHOUT THE PROCEDURE DISRUPTED THE STEADY-STATE OF THE SYSTEM AND CONTRIBUTED TO THE ELEVATED WBC CONTENT MEASURED IN THE PLATELET PRODUCT. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4233772 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
3833710 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS EVENT. THERE WERE NO EVENTS NOTED IN THE DHR THAT WOULD HAVE CONTRIBUTED TO THE ELEVATED WBC COUNT EXPERIENCED BY THE CUSTOMER. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4068039 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: . A REVIEW OF THE DEVICE HISTORY RECORD (DHR) FOR THIS UNIT SHOWED NO IRREGULARITIES DURING MANUFACTURING THAT WERE RELEVANT TO THIS ISSUE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. POSSIBLE ROOT CAUSES WERE PROVIDED IN THE INITIAL REPORT FOR THIS EVENT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE, NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
3458414 | INVESTIGATION: THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT WAS LIKELY A RESULT OF AN ESCAPE OF WBCS FROM THE LRS CHAMBER DURING A PORTION OF THE PROCEDURE. THERE ARE NO EVENTS (ADJUSTMENTS, CHANGES IN PUMP SPEED, SUBSTATE CHANGES, ETC.) IN THE PROCEDURE THAT CORRESPONDS WITH THE ONSET OF THE OVERLOADING OF THE LRS CHAMBER. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THIS LEUKO REDUCTION FAILURE COULD BE DONOR-RELATED. INVESTIGATION EVALUATION AND CORRECTIVE ACTIONS ARE IN-PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
4013408 | INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE (RDF) WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. AFTER REVIEWING THE SIGNALS IN THE RUN DATA FILE, IT CANNOT BE RULED OUT THAT THE HIGH NUMBER OF ACCESS ALERTS THAT OCCURRED THROUGHOUT THE PROCEDURE DISRUPTED THE STEADY-STATE OF THE SYSTEM AND CONTRIBUTED TO THE ELEVATED WBC CONTENT MEASURED IN THE PLATELET PRODUCT. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER-THAN-EXPECTED WBC CONTENT IN THE PLATELET PRODUCT.><THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6). THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY. | 0
3484281 | THE CUSTOMER WOULD LIKE THE RUN DATA FILE INVESTIGATED TO DETERMINE A POSSIBLE CAUSE FOR THE ELEVATED WHITE BLOOD CELL (WBC) CONTENT IN THE PLATELET PRODUCT. THERE WAS NOT A TRANSFUSION RECIPIENT OR PATIENT INVOLVED AT THE TIME OF THE RESIDUAL WBC TESTING, THEREFORE NO PATIENT INFORMATION IS REASONABLY KNOWN AT THE TIME OF THE EVENT. DONOR UNIT #: (B)(6) THE DISPOSABLE SET IS UNAVAILABLE FOR RETURN BECAUSE THE CUSTOMER DISCARDED IT. THIS REPORT IS BEING FILED DUE TO A DEVICE MALFUNCTION THAT HAS THE POTENTIAL FOR INJURY.><INVESTIGATION: THE DISPOSABLE SET WAS UNAVAILABLE FOR RETURN AND INVESTIGATION. THE RUN DATA FILE WAS ANALYZED FOR THIS EVENT. THE DEVICE HISTORY RECORDS (DHR) WERE REVIEWED FOR THIS LOT. THERE WERE NO ISSUES NOTED IN THE DHR THAT WOULD CONTRIBUTE TO THE ELEVATED WBC COUNT THAT THE CUSTOMER EXPERIENCED. ROOT CAUSE: A DEFINITIVE ROOT CAUSE FOR THE OBSERVED LEUKOREDUCTION FAILURE REMAINS UNDETERMINED AT THIS TIME. THE ANALYSIS OF THE RUN DATA FILE DID NOT FIND A CONCLUSIVE CAUSE FOR THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT REPORTED FOR THIS COLLECTION. NO UNUSUAL PROCESS VARIABLE WAS IDENTIFIED AND THE SIGNALS IN THE RUN DATA FILE INDICATE THAT THE TRIMA ACCEL SYSTEM OPERATED AS INTENDED. BASED ON THE AVAILABLE INFORMATION, IT CANNOT BE RULED OUT THAT THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT COULD BE DONOR-RELATED. IT ALSO CANNOT BE RULED OUT THAT A SAMPLING, CALCULATION, OR OTHER PROCESS ERROR COULD HAVE CONTRIBUTED TO THE HIGHER THAN EXPECTED WBC CONTENT IN THE PLATELET PRODUCT. | 0
4302161 | THE CUSTOMER REPORTED A DONOR REACTION POST DONATION. PER THE CUSTOMER, THE DONOR PASSED OUT IN THE BATHROOM AFTER THE DONATION. THE DONOR WAS TAKEN TO THE HOSPITAL. DONOR'S STATUS IS NOT KNOWN AT THIS TIME. PER THE CUSTOMER, THE DONOR COMPLETED THE DONATION PROCEDURE WITH NO ADVERSE SYMPTOMS. THE CUSTOMER HAS DECLINED TO PROVIDE ANY FURTHER INFORMATION ON THE EVENT. THE CUSTOMER DECLINED TO PROVIDE THE PATIENT IDENTIFIER, AGE, AND WEIGHT OF THE PATIENT. THE DISPOSABLE SET IS NOT AVAILABLE FOR RETURN BECAUSE IT WAS DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO MEDICAL INTERVENTION IN THE FORM OF THE DONOR BEING TAKEN TO THE HOSPITAL.><THE CUSTOMER DID NOT PROVIDE THE LOT NUMBER PERTAINING TO THIS EVENT,THEREFORE A DEVICE HISTORY RECORD (DHR) SEARCH COULD NOT BE CONDUCTED FOR THIS SPECIFIC INCIDENT. ALL LOTS MUST MEET ACCEPTANCE CRITERIA BEFORE RELEASE. ROOT CAUSE: THIS DISPOSABLE SET WAS UNAVAILABLE FOR SPECIFIC ROOT CAUSE ANALYSIS. BASED ON INFORMATION PROVIDED BY THE CUSTOMER, IT IS POSSIBLE THAT THE PATIENT EXHIBITED AN ALLERGIC REACTION TO ACD-A AND/OR ETO, DUE TO THEIR CURRENT PHYSIOLOGICAL CONDITION. PER THE THERAPEUTIC APHERESIS HANDBOOK: A PHYSICIAN'S REFERENCE, 2ND EDITION, FIRST TIME PROCEDURE PATIENTS REACTIONS MAY OCCUR IN APPROXIMATELY 4.8% OF PROCEDURES, AND 1.2% OF PROCEDURE PATIENTS EXPERIENCE REACTIONS TO CITRATE, BASED ON THE RESULTS OF A SURVEY OF THERAPEUTIC PROCEDURES. OF THESE REACTIONS, MOST WERE MILD AND WELL TOLERATED. THE TRIMA ACCEL SYSTEM HAS MANY SAFETY FEATURES. HOWEVER A DONOR REACTION CAN OCCUR RAPIDLY. THEREFORE, IT IS IMPERATIVE THAT THE TRIMA ACCEL SYSTEM AND THE DONOR BE MONITORED THROUGHOUT THE PROCEDURE.><INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDR'S, PER FDA REQUEST, TO IDENTIFY RECORDS IN WHICH A SERIOUS INJURY OR MEDICAL INTERVENTION OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED AS A SERIOUS INJURY ON THE MDR FORM. THIS SUPPLEMENT IS BEING FILED TO MODIFY INFORMATION PER FDA REQUEST. | 0
4181907 | THE CUSTOMER REPORTED THAT OVER THE LAST WEEKS SINCE THE INSTALLATION OF TRIMA SOFTWARE VERSION 6.06 AND THE ACTIVATION OF THE PLASMA RINSEBACK FEATURE, THEY HAVE SEEN SEVERAL DONORS EXPERIENCE 'SEVERE' AC REACTIONS DURING THE PLASMA RINSEBACK. PER THE CUSTOMER, SOME DONORS HAD PROBLEMS BREATHING CORRECTLY. THOSE DONORS DID NOT HAVE REACTIONS WHEN THEY WERE DONATING PLATELETS ON TRIMA SOFTWARE VERSION 5.21. SPECIFIC DATES OF THE EVENTS AND SPECIFIC PATIENT INFORMATION IS NOT AVAILABLE AT THIS TIME. IT IS NOT KNOWN AT THIS TIME IF MEDICAL INTERVENTION WAS NECESSARY FOR ANY OF THE EVENTS. THE DISPOSABLE KITS ARE NOT AVAILABLE FOR RETURN BECAUSE THEY WERE DISCARDED BY THE CUSTOMER. THIS REPORT IS BEING FILED DUE TO INSUFFICIENT INFORMATION PROVIDED AT THIS TIME TO DETERMINE IF A MALFUNCTION WITH THE POTENTIAL FOR INJURY HAS OCCURRED.><PER THE CUSTOMER, MEDICAL INTERVENTION WAS NOT NECESSARY FOR THESE EVENTS. IN 2 CASES, THEY INCREASED THE PUMP RATE OF THE CA INFUSION PUMPS ON THE TRIMA MACHINE AND DID NOT CONSIDER IT AS MEDICAL INTERVENTION. THE PATIENT'S SYMPTOMS WAS NAUSEA AND HE FELT TINGLING OF LIPS AND HANDS. THE PATIENT'S AGE IS NOT AVAILABLE AT THIS TIME. THE CUSTOMER REPORTED 5 CASES OF DONOR REACTION, BUT SENT PATIENT (DONOR) INFORMATION FOR 6 INCIDENTS. SPECIFIC INCIDENT DATES WERE NOT PROVIDED BY THE CUSTOMER.><INVESTIGATION: THE DEVICE WAS TESTED FOR FUNCTIONALITY, WITH NO PROBLEMS IDENTIFIED. NO ADDITIONAL REPORTS HAVE BEEN RECEIVED FOR THIS DEVICE REGARDING DONOR REACTIONS. A REVIEW OF THE LAST YEAR OF SERVICE HISTORY FOR THIS DEVICE, INDICATED NO OTHER REPORTS RELATED TO THIS ISSUE. FOUR RUN DATA FILES (RDFS) WERE ANALYZED FOR THESE EVENTS. IN ALL FOUR PROCEDURES, THE MAXIMUM CONCURRENT PLASMA COLLECTION ENDED APPROXIMATELY HALFWAY THROUGH THE PROCEDURE. DURING THE PLASMA COLLECTION, MUCH OF THE AC IS GOING TO THE PLASMA PRODUCT, AND THEREFORE THE AC INFUSION RATE TO THE DONOR IS LESS. WHEN PLASMA COLLECTION HAS COMPLETED, THE AC INFUSION RATE TO THE DONOR INCREASES, AS THE DONOR¿S PLASMA IS NOW BEING RETURNED. SINCE PLASMA COLLECTION IS ENDING ABOUT HALFWAY THROUGH THE PROCEDURES, DONORS HAVE HIGHER INFUSION RATES FOR THE SECOND HALF OF THEIR PROCEDURES. FROM THE INFORMATION PROVIDED, IT SOUNDS LIKE THE PATIENTS ARE NOT REACTING TO THE AC AT ALL UNTIL THE PLASMA RINSEBACK PROTOCOL IS TURNED ON. THE RDF REVIEW SHOWED THAT THE AC INFUSION RATE DURING RINSEBACK IS ABOUT EQUAL TO OR LESS THAN (NEVER GREATER THAN) THE AC INFUSION RATE DURING PLATELET ONLY COLLECTION. THE AC INFUSION RATE DECREASES OVER THE COURSE OF PLASMA RINSEBACK. IT IS NOT CLEAR WHY DONORS ARE REACTING DURING PLASMA RINSEBACK, BUT NOT PRIOR TO RINSEBACK. SOME DONORS HAVE NORMAL AC REACTIONS (NOT SEVERE) AT THE END OF PLASMA COLLECTION DUE TO THE INCREASE IN INFUSION RATE. THIS EVENT HAS BEEN REFERENCED IN A SOFTWARE ISSUE TRACKING SYSTEM FOR FURTHER ANALYSIS. ROOT CAUSE: SOFTWARE ISSUE.><THIS RECORD WAS IDENTIFIED DURING A RETROSPECTIVE REVIEW OF MDRS TO IDENTIFY RECORDS IN WHICH AN EVENT OCCURRED, BUT THE TYPE OF REPORTABLE EVENT WAS NOT INDICATED APPROPRIATELY.><INVESTIGATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE PROVIDED.><INVESTIGATION: TERUMO BCT'S FIELD PERFORMANCE ENGINEER PERFORMED A SITE ASSESSMENT TO INVESTIGATE WHAT CAN BE DONE TO ALLEVIATE THE CUSTOMER'S ISSUE. NO SPECIFIC REMARKS WERE MADE DURING THE ASSESSMENT REGARDING ANY MORE AC REACTIONS DURING PLASMA RINSEBACK. PREVENTATIVE MAINTENANCE WAS PERFORMED ON THE DEVICE. NO ISSUES WERE FOUND DURING PM. AC MANAGEMENT IS COVERED IN THE TRIMA OPERATOR'S MANUAL. THE MANUAL CAUTIONS THE OPERATOR TO MONITOR THE DONOR THROUGHOUT THE PROCEDURE IN ORDER TO RESPOND QUICKLY TO DONOR REACTIONS. DONOR REACTIONS TO AC ARE A KNOWN POSSIBLE SIDE EFFECT. ROOT CAUSE: NO DEFECT WAS FOUND WITH THE DEVICE. FURTHER INVESTIGATION INDICATES THERE WAS NO SOFTWARE ISSUE, AS INITIALLY SUSPECTED, AND THAT THE SOFTWARE PERFORMED AS INTENDED. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED AT THIS TIME. THE KNOWN POSSIBLE SIDE EFFECTS FOR AC REACTIONS ARE: MILD TINGLING OF THE LIPS, LIGHTHEADEDNESS AND NAUSEA. IN RARE CASES TREMBLING, TWITCHING OR INCREASE IN HEART RATE MAY OCCUR.><DUE TO EU PERSONAL DATA PROTECTION LAWS, THE PATIENT'S AGE IS NOT AVAILABLE FROM THE CUSTOMER. | 0
4135220 | FLUID INTRUSION TO THE WIRELESS/BATTERY MODULE. CAUSING EXTENSIVE ELECTRONIC DAMAGE/CORROSION/OXIDATION TO THE MAIN PCBA/WIRELESS ANTENNA. PICTURES SENT TO MFR SHOWING DAMAGE WITH NO FEED BACK. THIS WAS A THIRD ATTEMPT TO REPORT DAMAGE/FAILURE OF DEVICE, DO NOT BELIEVE HOSPITALS SHOULD BE BILLED FOR THE REPAIR OF A DESIGN ISSUE. THE DESIGN OF THE PUMP AS IT SITS ON THE IV POLE ALLOWS FOR DIRECT FLOW OF ANY FLUID OF SPIKED BAG LOCATED ABOVE UNIT AND/OR COMING FROM ANY OTHER SOURCE TO FLOW DIRECTLY DOWN TO BATTERY LATCH AND INTO THE ELECTRONICS. THESE DEVICES WHICH WE ARE REPORTING DO NOT NEED TO BE WATER PROOF BUT SHOULD BE ENCOURAGED TO BE FLUID RESISTANT. IN TALKING TO TECH SUPPORT THE TECH STATED THAT THE PROBLEM HAD BEEN SEEN INSIDE THE PUMP ITSELF. THE STRONGEST CONCERN IS WITH PT IMPACT. WHAT WOULD BE THE WORST SITUATION FLUID ENTERING DEVICES COULD PROPOSE OTHER THAN STRAIGHT FAILURE. BIOMEDICAL ENGINEERING OPENED AND PHOTOGRAPHED DEVICES AND SENT PHOTOS TO THE MFR RECORD # (B)(4). (B)(4). | 0
3959945 | A SPACELABS FIELD SERVICE ENGINEER (FSE) VISITED THE HOSPITAL AND TESTED THE INCIDENT DEVICES ON-SITE WITH NO PROBLEM. THE DEVICES PASSED ALL TESTS AND GENERATED ALARMS FOR V-TAC. WE ARE COLLECTING MORE INFORMATION FROM THE FSE. THE INVESTIGATION IS UNDERWAY. NO ONE HAS BEEN INJURED AS A RESULT OF THIS ALLEGED MALFUNCTION. WE WILL FILE A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS CONCLUDED.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY MONITOR FAILED TO ALARM ON A RUN OF VENTRICULAR TACHYCARDIA (V-TAC).><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITORED WITH TRANSMITTER MODEL 90347-05, RECEIVER MODULE MODEL 90478 AND CENTRAL MONITOR MODEL 91387-38 FAILED TO GENERATE AN ALARM FOR VENTRICULAR TACHYCARDIA (VTACH) AT 8:05 A.M. ON (B)(6) 2014. NO ONE WAS INJURED AS THE RESULT OF THIS EVENT.><THIS EVENT WAS REPORTED TO SPACELABS 28 DAYS AFTER THE DATE OF OCCURRENCE. ON SITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THAT ALL DEVICES PERFORMED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. THE CUSTOMER ONLY PROVIDED A SHORT RHYTHM STRIP OF THE EVENT THAT WAS NOT LONG ENOUGH FOR ANALYSIS. THE CUSTOMER DID NOT PROVIDE A RETROSPECTIVE DATABASE WHICH IS NEEDED FOR A ROOT CAUSE DETERMINATION. AN INVESTIGATION TO ROOT CAUSE COULD NOT BE PERFORMED DUE TO THE LACK OF DATA PROVIDED BY THE CUSTOMER. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED. | 0
3676484 | A NEW PATIENT CAME IN TO THE PROGRESSIVE CARE UNIT (PCU). WE PLACED A NEW BATTERY IN THE CARDIAC MONITOR FOR THE PCU. AFTER FEW MINUTES THE MONITOR TECHNICIAN CALLED THE NURSE (RN) AND TOLD HER TO REPLACE BATTERY, MONITOR TECHNICIAN WAS TOLD THAT THE RN JUST PUT A BATTERY IN THE CARDIAC MONITOR FOR THIS NEW PATIENT. THE RN CHECKED THE MONITOR AND IT WAS HOT WHERE THE BATTERY IS IN, AND WHEN SHE OPENED IT, THE BATTERY WAS VERY HOT AND STARTING TO MELT. SHE WAS ABLE TO REMOVE ONE BATTERY BUT SHE BURNED HER RIGHT THUMB IN DOING SO. SHE LEFT THE OTHER BATTERY IN AND SHOWED IT TO THE MONITOR TECHNICIAN. THERE WAS A LIQUID NOTED COMING OUT FROM THE BATTERY. THE DEVICE WAS TAKEN OUT OF SERVICE. | 0
3947576 | NURSE STATED THAT SHE DID NOT ADMIT THE TELEMETRY 11 PATIENT TO THE SYSTEM. REVIEW OF SYSTEM LOGS CONFIRMS THAT THE PATIENT WAS NOT ADMITTED TO THE SYSTEM.><REFERENCE IMPORTER # (B)(4). | 0
4080818 | TELEMETRY MONITOR MALFUNCTION AND MONITOR TECH UNABLE TO VIEW PATIENT RHYTHM FOR LTAC PATIENTS AT 15:45. MALFUNCTION INVOLVED A TOTAL OF 6 PATIENTS WHERE PATIENT TELEMETRY WAS NOT VIEWABLE BY TECH. MINDRAY REPRESENTATIVE TROUBLESHOOTING AND UNABLE TO IDENTIFY PROBLEM. ON (B)(6) 2014, CEO REQUESTING ON-SITE VISIT FOR TELEMETRY EVALUATION AND REPAIR. MINDRAY ON SITE NAD TELEMETRY REPAIRED AT 1830. PATIENTS REMAIN SAFE AND FREE FROM HARM DURING THE TIME PERIOD. | 0
4041561 | THE PATIENT ON TELEMETRY WAS BEING MONITORED AT THE CENTRAL STATION WHEN THE ECG TRACE DISAPPEARED WITHOUT THE STAFF NOTICING OR WITHOUT ANY ALARMS BEING TRIGGERED.><SPACELABS IS EVALUATING THIS REPORTED EVENT AND WILL FILE A FOLLOW UP REPORT WHEN OUR EVALUATION IS CONCLUDED.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITORED WITH TRANSMITTER MODEL 90341-50, TELEMETRY RECEIVER MODULE MODEL 90478, AND CENTRAL MONITOR MODEL 91387-38 EXPERIENCED AN ECG TRACE DROP FROM THE CENTRAL MONITOR FROM 6:12 A.M. TO 1:45 P.M. ON (B)(6) 2013. THE DIGITAL HEART RATE DISPLAY REMAINED AT 81 BEAT PER MINUTE THROUGHOUT THIS TIME. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><THERE WERE NO ALARMS GENERATED FOR THE ECG TRACE DROP AND THE EVENT WENT UNNOTICED BY CLINICAL STAFF. THIS EVENT WAS DOCUMENTED IN THE PATIENT RETROSPECTIVE DATABASE PROVIDED BY THE CUSTOMER. THE INVOLVED DEVICES WERE REMOVED FROM SERVICE BY THE CUSTOMER. ON SITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THAT ALL DEVICES PERFORMED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. THE TELEMETRY TRANSMITTER AND RECEIVER MODULE WERE SHIPPED TO SPACELABS FOR FURTHER INVESTIGATION. A ROOT CAUSE ASSESSMENT WAS PERFORMED. ADDITIONAL FUNCTIONAL TESTS WERE ALSO PERFORMED. THE REPORTED PROBLEM WAS NOT REPRODUCIBLE; THEREFORE, NO ROOT CAUSE WAS DETERMINED. IN SUMMARY, WE WERE UNABLE TO REPLICATE THE ISSUE AND CANNOT CONCLUSIVELY DETERMINE ROOT CAUSE OF THE REPORTED FAILURE. WE ARE CLOSING THE INVESTIGATION INTO THIS PARTICULAR INCIDENT. WE WILL CONTINUE TO MONITOR SUCH ISSUES AND REOPEN THIS INCIDENT¿S COMPLAINT INVESTIGATION IF APPROPRIATE. | 0
3944932 | THE CUSTOMER PROVIDED THE RETROSPECTIVE PATIENT DATABASE, WHICH WAS REVIEWED BY A SPACELABS LEAD SOFTWARE ENGINEER. ECG WAVEFORMS WERE EXTRACTED FROM THE DATA AND RE-PLAYED THROUGH AN ANALOG SYSTEM FOR ANALYSIS TO CONFIRM INTERPRETATION FROM THE ECG ALGORITHM CONTAINED WITHIN THE TELEMETRY RECEIVER MODULE. THE ECG ALGORITHM AND ANALOG ANALYSIS DID NOT GENERATE AN ALARM AT THE EVENT TIME BECAUSE THE THRESHOLD CLASSIFICATION FOR RECOGNITION OF VFIB BY THE PRODUCT WAS NOT MET. THIS EVENT HAD LOW WAVEFORM AMPLITUDE AND ABERRANT MORPHOLOGY WITH SUBTLE DIFFERENCES IN TIMING, SLOPE AND BEAT INTERVALS FROM THE TWO ECG LEADS USED FOR ANALYSIS. LOW WAVEFORM AMPLITUDE AND OTHER ECG SIGNAL ISSUES ARE DISCUSSED IN THE SPACELABS CLINICAL PARAMETERS MANUALS PROVIDED TO THE CUSTOMER. IN CONCLUSION, THERE WAS NO EQUIPMENT MALFUNCTION. SPACELABS VFIB ALARM THRESHOLDS WERE NOT MET THEREFORE NO ALARM WAS GENERATED. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED. PLACEHOLDER.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY CENTRAL MONITOR FAILED TO ALARM ON VENTRICULAR FIBRILLATION (V-FIB). PATIENT WAS SHOCKED WITH IMPLANTABLE CARDIOVERTER DEFIBRILLATOR (ICD) AND CONVERTED TO A PACED RHYTHM.><A SPACELABS FIELD SERVICE ENGINEER (FSE) INVESTIGATED THE INCIDENT DEVICE ONSITE. SPACELABS IS FOLLOWING UP WITH THE FSE TO GET MORE INFORMATION. FULL DISCLOSURE DATA WAS PROVIDED BY THE CUSTOMER. INVESTIGATION IS UNDERWAY. WE WILL FILE A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS CONCLUDED. PLACEHOLDER.><.><SPACELABS RECEIVED A REPORT THAT A PATIENT MONITORED WITH TELEMETRY TRANSMITTER MODEL 96281, RECEIVER MODULE MODEL 90478 AND CENTRAL MONITOR MODEL 91387-38 FAILED TO GENERATE AN ALARM FOR VENTRICULAR FIBRILLATION WITH TORSADE (VFIB) ON (B)(6) 2014 AT 11:55 P.M. THE PATIENT¿S INTERNAL CARDIAC DEFIBRILLATOR DETECTED AND TREATED THE VFIB RETURNING THE PATIENT TO A NORMAL CARDIAC RHYTHM. NO ONE WAS INJURED AS A RESULT OF THIS EVENT. | 0
3936332 | SPACELABS IS EVALUATING THIS REPORTED EVENT AND WILL FILE A SUPPLEMENTAL REPORT WHEN OUR EVALUATION IS CONCLUDED. PLACEHOLDER.><ONSITE TESTING OF THE INVOLVED DEVICES WAS SCHEDULED BY A SPACELABS FIELD SERVICE ENGINEER (FSE); HOWEVER, THE FACILITY WITNESS TO THE EVENT WAS NOT AVAILABLE TO VERIFY WHICH EQUIPMENT WAS IN USE OR ANY EVENT DETAILS. NO OTHER FACILITY STAFF MEMBER PROVIDED EVENT DETAILS WHEN ASKED. THE FACILITY BIOMED SPECIFICALLY STATED HE WAS UNAWARE OF ANY ISSUES AND THE TELEMETRY SYSTEM WAS PERFORMING AS EXPECTED. THE FSE CONTACTED THE CUSTOMER TWO ADDITIONAL TIMES TO OBTAIN EVENT DETAILS, BUT NO ONE FROM THE FACILITY PROVIDED ANY INFORMATION OR REPORTED ANY ISSUES. THE CUSTOMER DID NOT PROVIDE EQUIPMENT LOGS, SERIAL NUMBERS, RECORDINGS OR A PATIENT RETROSPECTIVE DATABASE NEEDED FOR AN EVENT INVESTIGATION. ROOT CAUSE COULD NOT BE DETERMINED DUE TO INSUFFICIENT DATA FROM THE CUSTOMER. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY MONITOR FAILED TO ALARM FOR ARRHYTHMIA LASTED 6 SECONDS. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><SPACELABS RECEIVED A REPORT THAT A TELEMETRY PATIENT MONITORED WITH RECEIVER MODULE MODEL 90478 AND CENTRAL MONITOR MODEL 91387-38 FAILED TO GENERATE AN AUDIBLE ALARM OR PRINTOUT FOR A WITNESSED SIX SECONDS ARRHYTHMIA ON (B)(6) 2014. THE ALARM EVENT DID DISPLAY VISUALLY AT THE CENTRAL MONITOR AND SENT AN ALARM EVENT TO ICS-G2 (THE RETROSPECTIVE DATABASE). NO ONE WAS INJURED AS A RESULT OF THIS EVENT. | 0
3557828 | TELEMETRY SYSTEM ON MED SURG FAILED AT 4:30 AM. RN SUPERVISOR ATTEMPTED TO TROUBLESHOOT. CALLED MINDRAY TO TROUBLESHOOT. MATERIALS MANAGER AND MINDRAY ATTEMPTED TROUBLESHOOT WITHOUT SUCCESS. MINDRAY TECH CAME TO REPAIR, HAD TO RETURN TO (B)(4) TO RETRIEVE PART. SYSTEM REPAIRED AND PTS PLACED BACK ON TELEMETRY AT 12:10 PM. ACCORDING TO TECH, THE POWER SUPPLY IN A RECEIVER MALFUNCTIONED, LIKELY DUE TO POWER FLUCTUATION. ALSO NOTED THAT BATTERIES IN UNINTERRUPTED POWER SUPPLY WERE IN NEED OF REPLACEMENT AND MAY HAVE CONTRIBUTED. | 0
3977516 | TWO TELEMETRY TECHNICIANS WITNESSED A 15 BEAT RUN OF VTACH IN THE CENTRALIZED TELEMETRY MONITORING STATION WHICH DID NOT PRODUCE AN ALARM. THE TELEMETRY TECHNICIANS WERE ABLE TO ALERT CLINICIANS AND THERE WAS NO PATIENT HARM ASSOCIATED WITH THIS EVENT, HOWEVER SIGNIFICANT POTENTIAL TO CAUSE HARM IF VTACH WERE MISSED (I.E. IN FACILITIES THAT DO NOT HAVE CENTRALIZED TELEMETRY MONITORING). THE MANUFACTURER ACKNOWLEDGE THAT VTACH HAD OCCURRED AND THAT IT WAS MISSED BY THEIR ARRHYTHMIA DETECTION SOFTWARE. | 0
3606526 | PT HAD TELEMETRY UNIT IN PLACE. HE HAD AN ICD THAT WAS NOTED TO FIRE ON WHILE ON MONITOR, BUT NO HEART RHYTHM WAS NOTED. THE TELEMETRY MONITOR DID NOT ALARM OR PICK UP THAT NO HEART RHYTHM WAS CAPTURED. | 0
3709515 | PATIENT HAD TELEMETRY UNIT IN PLACE. TELEMETRY MONITOR DID NOT ALARM WHEN PATIENT WENT INTO A JUNCTIONAL RHYTHM. WHEN PATIENT WENT FROM JUNCTIONAL RHYTHM. WHEN PATIENT WENT FROM JUNCTIONAL RHYTHM INTO A LETHAL RHYTHM, ONLY A VISUAL ALARM WAS NOTED AND NO AUDITORY ALARM SOUNDED. | 0
3770884 | ON (B)(6) 2013, THE PT HAD BEEN AMBULATED TO THE BATHROOM AND UPON RETURN TO THE BED AT 0705, THE PT BECAME UNRESPONSIVE. CAROTID PULSE CHECKED BY THE STAFF. NO PULSE WAS FELT. CODE WAS CALLED. THE PT WAS SUCCESSFULLY RESUSCITATED AND TRANSFERRED TO THE CRITICAL CARE UNIT. THE PT HAD BEEN IN ATRIAL FIBRILLATION PRIOR TO THIS EVENT. THE PT'S CARDIAC MONITORING SYSTEM FAILED TO RECOGNIZE A BRADYCARDIA RHYTHM OR ASYSTOLE. PT HAD PAUSES OVER SIX (6) SECONDS. SYSTEM BPM (BEATS PER MINUTE) WERE BLANK. NO ALARM SOUNDED. | 0
4121852 | TELE EQUIPMENT DOWN FOR 45 MINUTES. (B)(6) CONTACTED; PTS TRIAGED APPROPRIATELY; RETURNED SPONTANEOUSLY. NO HARM TO PT. THIS DOWNTIME INCLUDED A TOTAL OF 6 PTS. NO HARM TO ANY OF THE 6 PTS OCCURRED. | 0
4121849 | HALF OF TELE SYSTEM WENT DOWN FOR 30 MINUTES; PT TRIAGED OFF TO OTHERSIDE; FUNCTION RETURNED; UHS AND PHILIPS IN TO MANAGE SYSTEM. TOTAL OF TWO (2) PTS INCLUDED IN DOWNTIME. NO HARM TO EITHER PT. | 0
3976234 | THE ELECTRODE HAS BEEN REQUESTED SENT BACK FOR INVESTIGATION TO DETERMINE IF THIS PROBLEM HAS BEEN CAUSED BY A PRODUCT MALFUNCTION OR A USER ERROR. BASED ON THE EVENT DESCRIPTION THIS CASE COULD BE CAUSED BY A USER ERROR SINCE THE ELECTRODE TEMPERATURE WAS SET HIGHER THAN RECOMMENDED AND FOR 9 HOURS INSTEAD OF ONLY 4 HOURS. IT WILL BE DETERMINED IF THIS EVENT WAS CAUSED BY A PRODUCT MALFUNCTION OR A USER ERROR WHEN THE ELECTRODE HAS BEEN INVESTIGATED. BASED ON THE PICTURES RECEIVED THE BLISTERS ARE EVALUATED TO BE SECOND DEGREE BURN MARKS. NO INFORMATION OF TREATMENT APPLIED HAS BEEN RECEIVED. THIS INJURY IS EVALUATED AS REPORTABLE, BUT THE INJURY IS ONLY EVALUATED AS MINOR AS IT DOES NOT MEET THE CRITERIA OF A SERIOUS INJURY, AS DESCRIBED IN THE INSTRUCTIONS FOR THIS FORM.><ACCORDING TO THE COMPLAINT SKIN BLISTERS WERE RECEIVED AFTER USING THE TCM4 SENSOR FOR 9 HOURS AT THE SAME LOCATION WITH THE TEMPERATURE SET AT 45 DEGREES. 10 PATIENTS WERE AFFECTED. CURRENTLY, THE PATIENT INFORMATION RECEIVED FOR ALL PATIENTS IS IDENTICAL. MORE INFORMATION HAS BEEN REQUESTED. THIS REPORT REGARDS PATIENT NR 3. MDR REPORTS FOR THE AFFECTED PATIENTS WILL BE SUBMITTED WITH MDR REFERENCE 3002807968-2014-00023 TO 3002807968-2014-00032. | 0
3976217 | ACCORDING TO THE COMPLAINT SKIN BLISTERS WERE RECEIVED AFTER USING THE TCM4 SENSOR FOR 9 HOURS AT THE SAME LOCATION WITH THE TEMPERATURE SET AT 45 DEGREES. 10 PATIENTS WERE AFFECTED. CURRENTLY, THE PATIENT INFORMATION RECEIVED FOR ALL PATIENTS IS IDENTICAL. MORE INFORMATION HAS BEEN REQUESTED. THIS REPORT REGARDS PATIENT NR 4. MDR REPORTS FOR THE AFFECTED PATIENTS WILL BE SUBMITTED WITH MDR REFERENCE 3002807968-2014-00023 TO 3002807968-2014-00032.><THE ELECTRODE HAS BEEN REQUESTED SENT BACK FOR INVESTIGATION TO DETERMINE IF THIS PROBLEM HAS BEEN CAUSED BY A PRODUCT MALFUNCTION OR A USER ERROR. BASED ON THE EVENT DESCRIPTION THIS CASE COULD BE CAUSED BY A USER ERROR SINCE THE ELECTRODE TEMPERATURE WAS SET HIGHER THAN RECOMMENDED AND FOR 9 HOURS INSTEAD OF ONLY 4 HOURS. IT WILL BE DETERMINED IF THIS EVENT WAS CAUSED BY A PRODUCT MALFUNCTION OR A USER ERROR WHEN THE ELECTRODE HAS BEEN INVESTIGATED. BASED ON THE PICTURES RECEIVED THE BLISTERS ARE EVALUATED TO BE SECOND DEGREE BURN MARKS. NO INFORMATION OF TREATMENT APPLIED HAS BEEN RECEIVED. THIS INJURY IS EVALUATED AS REPORTABLE, BUT THE INJURY IS ONLY EVALUATED AS MINOR AS IT DOES NOT MEET THE CRITERIA OF A SERIOUS INJURY, AS DESCRIBED IN THE INSTRUCTIONS FOR THIS FORM. . | 0
4356208 | ACCORDING TO THE COMPLAINT TWO PATIENTS RECEIVED SKIN BLISTERS AFTER USING THE TCM4 FOR A SLEEP STUDY. TWO TCM4 MONITORS WAS SENT TO RADIOMETER FOR EVALUATION BUT THE CUSTOMER DID NOT KNOW WHICH OF THE TWO MONITORS WERE USED FOR THE STUDY. THIS REPORT CONCERNS PATIENT (B)(4). MDR REPORTS FOR THE OTHER AFFECTED PATIENT WILL BE SUBMITTED WITH MDR REFERENCE 3002807968-2014-00054. | 0
4064347 | THE T:CONNECT SOFTWARE USED TO DOWNLOAD INSULIN PUMP SETTINGS DOES NOT DOWNLOAD THE T-SLIM PUMP SETTINGS ACCURATELY. THIS WILL LEAD A PROVIDER TO POTENTIALLY MAKE INACCURATE INSULIN PRESCRIPTIONS. THE SPECIFIC ISSUE IS THAT FOR A CARBOHYDRATE RATIO LESS THAN 10 THE DECIMAL POINTS ARE NOT DISPLAYED AND ARE INSTEAD ROUNDED TO THE NEAREST WHOLE VALUE. FOR PTS UPLOADING AT HOME, LEADING TO RECOMMENDATIONS BETWEEN VISITS, THIS PUTS THE PROVIDER AT SIGNIFICANT RISK FOR GIVING THE WRONG INSTRUCTIONS AND PUTS THE PT AT RISK FOR POTENTIAL INSULIN OVERDOSE. FOR AN IMPORTANT MINORITY OF PTS WHICH INCLUDE THIRD TRIMESTER PREGNANT WOMEN, AN EXAMPLE WOULD BE A CARBOHYDRATE RATIO OF 2.4. THIS DISPLAYS AS 2. THE PROVIDER WANTS TO MAKE A 10% CHANGE IN THE DOSE AND TELLS THE PT TO CHANGE TO 1.8 INSTEAD OF 2.2. WE DEPEND ON THE PT BEING PARTICULARLY SAVVY TO CATCH THIS MISTAKE. THAT IS AN UNSAFE ASSUMPTION. TO MY KNOWLEDGE THE COMPANY HAS NOT MADE THIS PUBLIC EVEN THOUGH THEY HAVE BEEN AWARE OF IT FOR SOME TIME. (B)(4). | 0
4076598 | DAVINCI HYSTERECTOMY; BLADE STAYS EXTENDED, DID NOT RETRACT. | 0
3743661 | PATIENT WAS HAVING A ROBOT-ASSISTED PROCEDURE. WHEN PUTTING IN VERESS NEEDLE, GOT BLOOD BACK. PATIENT DEVELOPED A HEMATOMA FROM VERESS NEEDLE PLACEMENT. | 0
4264267 | LAP SHEAR HANDLE DID NOT ROTATE DURING MANIPULATION DURING SURGERY. ETHICON REP WAS PRESENT WHILE THE DEVICE WAS BEING USED. WE SEQUESTERED THE HANDPIECE INVOLVED AND THERE WAS NO PROBLEM WITH THE HARMONIC GENERATOR ITSELF.======================MANUFACTURER RESPONSE FOR LAP SHEARS, HARMONIC ACE+7 LAPAROSCOPIC SHEARS (PER SITE REPORTER).======================UNKNOWN. | 0
3750508 | VESSEL SEALER USED FOR LESS THAN 5 MINUTES WHEN GREEN ALERT BOX POPPED UP INSTRUMENT NOT ATTACHED" PK USED FOR REST OF CASE. NO INJURY TO PATIENT." | 0
4095415 | BLADE EXPOSED AND FAILED SELF-CHECK. ANOTHER DEVICE OBTAINED AND PROCEDURE CONTINUED. NO PATIENT HARM.THIS IS THE SECOND EVENT AT THIS FACILITY WITH THIS DEVICE IN THE LAST MONTH. IN BOTH CASES, THE DEVICE WAS THE SAME CATALOG AND LOT NUMBER. | 0
4109351 | THE PATIENT CAME IN FOR A TOTAL LAPAROSCOPIC ROBOTIC HYSTERECTOMY AND BILATERAL SALPINGO-OOPHORECTOMY. DURING THE CONDUCT OF THE OPERATION LAPAROSCOPICALLY AND MOBILIZATION OF DENSE ADHESIONS BETWEEN THE SIGMOID COLON AND THE RIGHT PELVIC SIDEWALL, THERE WAS AN INJURY WITH THE MONOPOLAR SCISSORS TO THE RIGHT EXTERNAL ILIAC ARTERY. ROBOTIC ASSISTED HYSTERECTOMY CONVERTED TO LAPAROTOMY AFTER HEMORRHAGE EVENT. CONVERSION TO LAPAROTOMY AND OVERSEWING OF THIS ARTERIAL INJURY WAS ACCOMPLISHED WITH SUTURES. THE PATIENT WAS TAKEN TO RECOVERY ROOM IN STABLE CONDITION. | 0
4095420 | BLADE EXPOSED AND FAILED SELF-CHECK. ANOTHER DEVICE OBTAINED AND PROCEDURE CONTINUED. NO PATIENT HARM. | 0
4066408 | THE CAUTERY HOOK SPARKED WHILE IT WAS IN USE. THE INSTRUMENT WAS REMOVED FROM THE STERILE FIELD. THERE WAS NO INJURY TO THE PATIENT. | 0
4137720 | 5MM DAVINCI THORACIC GRASPER IN USE. SURGEON NOTICED A TEAR IN THE BLACK INSULATION AND ASKED THE INSTRUMENT TO BE REMOVED. AS THE INSTRUMENT WAS REMOVED APPROXIMATELY ONE INCH SECTION OF THE INSULATION CAME OFF. NO VISIBLE INSULATION PARTICLES IN THE PATIENT; SURGEON SEARCHED CAVITY. DAVINCI WAS NOTIFIED OF INSTRUMENT PROBLEM. | 0
3605081 | WHILE TAKING DOWN ADHESIONS, MADE INCIDENTAL HOLE IN SMALL BOWEL, SMALL BOWEL REPAIR DONE. | 0
3831945 | DOCTOR NOTICED THE VESSEL SEALER WAS NOT REGISTERED ON THE ROBOT AND SAW THAT THE LIGHT ABOVE THE CORD WHICH SHOULD BE BLUE WAS RED. HE CALLED INTUITIVE AND THEY ADVISED HIM TO GET A NEW CORD AND RESTART THE ROBOT. WHEN HE DID THIS, IT REGISTERED THE NEW VESSEL SEALER. THE COMPANY WANTED HIM TO SEND THE CORD AND INSTRUMENT VESSEL SEALER IN THAT DID NOT REGISTER TO CHECK FOR MALFUNCTION. | 0
3904281 | THE DAVINCI CAUTERY HOOK THAT WE USE BROKE DURING USE. THE INSTRUMENT WAS IMMEDIATELY REMOVED IN ITS ENTIRETY. | 0
3994218 | THE AUDIO COMMUNICATION SYSTEM FOR THE INTUITIVE DAVINCI SI ROBOT IS POORLY DESIGNED. THE MICROPHONE THAT ENABLES THE CLINICIAN AT THE ROBOT CONTROL CONSOLE TO HEAR CLINICIAN AT BEDSIDE REQUIRES A DISTANCE OF LESS THAN 3 FEET BETWEEN PERSON SPEAKING AND THE MICROPHONE AND ALSO REQUIRES LINE OF SITE. THE MICROPHONE IS ATTACHED TO A TOWER WITH A LARGE FOOTPRINT, WHICH IS DIFFICULT TO PLACE IN THE IDEAL LOCATION. IN OUR SETUP, THE MICROPHONE IS LOCATED BEHIND THE CLINICIAN AT THE BEDSIDE. THE VENDOR WILL NOT SUPPORT A SOLUTION FOR PLACING THE MICROPHONE IN A LOCATION WHERE IT CAN PICK UP THE CLINICIAN AT THE BEDSIDE. MOREOVER, THE MICROPHONE IS OPEN AT ALL TIMES SUCH THAT SOUND COMING OUT OF THE AMPLIFIED SPEAKER IS PICKED UP BY THE MICROPHONE, POTENTIALLY CREATING AUDIO FEEDBACK. AN ADDITIONAL PROBLEM IS THAT THE SECOND MICROPHONE INSIDE THE ROBOT CONTROL CONSOLE IS TRIGGERED BY THE AUDIO LEVEL OF THE CLINICIAN'S VOICE. THIS DESIGN SCHEME CAUSES CHOPPING OF THE AUDIO STREAM. THERE IS NO MEANS FOR THE USER TO REMOVE THE TRIGGERING. FINALLY, WHEN THE AUDIO IS ROUTED THROUGH THE PROCESSOR FOR THE ROBOT, A DELAY IS NOTED. THE SOUND DIRECT FROM THE CLINICIAN AT THE CONSOLE TO THE LISTENER'S EAR ARRIVES BEFORE THE SOUND COMING OUT OF THE AMPLIFIED SPEAKER.MANUFACTURER RESPONSE: THE MANUFACTURER HAS CHANGED THE HARDWARE ASSOCIATED WITH THE MICROPHONE AND CONFIRMED THAT THE AUDIO COMMUNICATION IS PERFORMING ACCORDING TO THEIR SPECIFICATIONS. HOWEVER, DUE TO THE CONSTRAINTS OF THE FDA REGULATORY CLEARANCE OF THE DEVICE AND THE LACK OF CATALOG ITEMS, THE VENDOR IS NOT WILLING OR ABLE TO SOLVE THE AUDIO COMMUNICATION PROBLEM. | 0
4033892 | DA VINCI LARGE SUTURE CUT NEEDLE DRIVER WAS CONNECTED TO ROBOTIC ARM DURING PROCEDURE. IT WAS NOTED ON SCREEN (WHILE INSIDE PATIENT) THAT THE WIRE SNAPPED. UPON REMOVAL FROM ROBOT ARM AND INSPECTION, NOTED THAT WIRE HAD INDEED SNAPPED YET IT WAS INTACT STILL ON THE DEVICE. NO HARM WAS NOTED TO THE PATIENT. | 0
3588766 | WHILE UNDOCKING ROBOT AT THE CONCLUSION OF A ROBOTIC ASSISTED RADICAL NEPHRECTOMY, SCRUB NURSE NOTICED A BULGE UNDER THE SCISSOR TIP COVER OF THE ROBOTIC ENDOWRIST. SCISSOR TIP WAS DIFFICULT TO REMOVE. INSTRUMENT WAS PASSED OFF THE STERILE FIELD, WHEN REMOVED; MULTIPLE PIECES OF THE ENDOWRIST SHAFT WERE FOUND TO BE LOOSE UNDER THE SCISSOR TIP. SURGEON NOTIFIED. LAPAROSCOPE USED TO VIEW THE SURGICAL FIELD AND EVALUATED FOR LOOSE PIECES REMAINING. NONE NOTED, AREA THOROUGHLY IRRIGATED AND SUCTIONED OUT. REMAINING PIECES OF INSTRUMENT WERE COLLECTED. | 0
3983011 | THE PATIENT UNDERWENT A ROBOTIC ASSISTED LAPAROSCOPIC RADICAL NEPHRECTOMY, LEFT ADRENALECTOMY, AND REPAIR OF SPLENIC CAPSULAR TEAR. THE PATIENT RETURNED TO THE OPERATING ROOM (OR) ONE WEEK LATER FOR EXPLORATORY LAPAROTOMY AND HAS A COLON INJURY. THE OPERATING SURGEON BELIEVES THIS MAY BE A THERMAL BURN FROM THE PRIOR SURGERY; LIKELY FAILURE OF INSULATION ON THE CANNULA. | 0
4239378 | SURGEON WAS USING THE ENDOWRIST ONE VESSEL SEALER, WHEN HE WAS DONE THE JAWS WERE OPENED BUT THE BLADE WAS STILL SHOWING. | 0
3735551 | THE DEVICE WAS WORKING PROPERLY, AND THEN ALL OF THE SUDDEN IT ALARMED DEVICE MALFUNCTION. REMOVED THE VESSEL SEALER." VESSEL SEALER WAS REMOVED AND PLACED BACK ON AND STILL NOT WORKING | 0
4088666 | DURING DAVINCI PROCEDURE, BLADE OF VESSEL SEALER BECAME EXPOSED. BLADE WAS STRAIGHT OUT BETWEEN JAWS OF VESSEL SEALER. INSTRUMENT REMOVED AND REPLACED WITH NEW INSTRUMENT. NO HARM TO PATIENT. | 0
3605075 | PATIENT HAD ROBOTIC HYSTERECTOMY, THERE WERE TWO LARGE TEARS IN BOTH RIGHT AND LEFT LATERAL WALLS NEAR THE CUFF. | 0
4033954 | DURING SURGERY, THE TECHNICIAN WAS PLACING INSTRUMENT TO ARM #1, UPON INSERTION, AN ALARM WAS NOTED STATING THAT DEVICE CANNOT BE IDENTIFIED". ATTEMPTED TO PLACE 3 TIMES AND STILL NOTED SAME THING. TECHNICIAN USED A NEW SMALL CLIP APPLIER INSTRUMENT | 0
3644098 | PATIENT WAS HAVING A LAPAROSCOPIC ROBOTIC ASSISTED NEPHROURETERECTOMY. WHILE SURGEON WAS USING A ROBOTIC 8MM ENDOWRIST MEGA NEEDLE DRIVER, CABLE ON INSTRUMENT BROKE. INSTRUMENT WAS REMOVED FROM FIELD AND NEW INSTRUMENT USED. | 0
3750615 | VESSEL SEALER MALFUNCTION. BLADE EXPOSED. NOT CUTTING. OPENED 2 OF THE VESSEL SEALER AND DID NOT USE EITHER ONE BECAUSE OF THE SAME PROBLEM. BOTH HAD SAME LOT # AND EXPIRATION DATE. NO HARM TO PATIENT. | 0
3743766 | DURING ROBOTIC LYMPH NODE DISSECTION, THE PHYSICIAN CUT THE LEFT URETER. THE UROLOGIST ON CALL FIXED THE URETER FOLLOWING ORGINIAL PROCEDURE ROBOTICALLY. | 0
3743764 | DURING THE COURSE OF DISSECTION, THERE DID APPEAR TO BE A SMALL RENT WITHIN THE MUCOSA OF THE SIGMOID COLON ABOUT 3 CM IN DIAMETER. GIVEN THIS RENT AND THE CONTINUED DIFFICULTY IN CLEARLY IDENTIFYING THE URETERS AS NECESSARY TO FREE THE OVARIES, THE PHYSICIAN WAS CONSULTED AND CAME TO THE OR AND ASSISTED IN COMPLETING THE BSO, LYSIS OF ADHESIONS, AND URETEROLYSIS AND REPAIRED THE RENT IN THE SEROSA OF THE SIGMOID COLON. | 0
4176509 | UPON USING THE VESSEL SEALER DEVICE ON PATIENT TISSUE, IT DID NOT SEAL AND THERE WAS AN ERROR CODE ON THE INTUITIVE SURGICAL SCREEN WE WERE LOOKING AT. WE THEN TOOK THE ITEM OUT AND RESET IT INTO THE ROBOTIC ARM AND IT WOULD NOT TAKE/READ THE INSTRUMENT. THE DEVICE WAS SEQUESTERED AND THE INSTRUMENTATION WAS PROCESSED PROCEEDED TO OPEN A NEW DEVICE, THIS ONE FORTUNATELY WORKED OUT JUST FINE. NO HARM TO THE PATIENT IN THIS CASE. | 0
4146579 | THE PATIENT WAS UNDERGOING A SINGLE SITE ROBOTIC TOTAL ABDOMINAL HYSTERECTOMY WITH BILATERAL SALPINGO-OOPHERECTOMY (TAH-BSO) AND CYSTOSCOPY. THE SURGEON WAS USING A SINGLE SITE CURVED NEEDLE DRIVER TO COMPLETE THE SUTURING AT THE END OF THE CASE, WHEN THE HEAD OF THE NEEDLE DRIVER BROKE OFF INTO THE PATIENT. THE SURGEON HAD TO LOCATE AND REMOVE SEVERAL LOOSE PIECES FROM THE PATIENT, PROLONGING TIME UNDER ANESTHESIA. THE SIDE OF THE DRIVER TIP, WHICH IS MADE OF HARD PLASTIC, APPEARS TO HAVE BROKEN OFF. IT IS ALSO UNCLEAR WHETHER A SMALL PIECE OF GRAY MATERIAL THAT HOLDS THE DRIVER HINGE IN PLACE WAS MISSING ON THIS NEEDLE DRIVER. IT WAS NOT FOUND IN THE PATIENT OR IN THE MATERIAL RETURNED TO US FROM THE OPERATING ROOM. WE THINK IT MAY NOT HAVE BEEN PLACED ON THE HINGE DURING MANUFACTURING, THUS ALLOWING THE DRIVER TO FALL APART AND BREAK UNDER TENSION. WE HAVE PHOTOS FOR REVIEW OF THE DRIVER WE RECOVERED AND A NEW DRIVER WE PULLED APART. | 0
4079937 | TIP OF DA VINCI HARMONIC ACE CURVED SHEARS INSERT BROKE OFF DURING SURGERY. TIP WAS RECOVERED BY SURGEON AND NO HARM TO PATIENT WAS CAUSED. | 0
3842637 | THE SURGEON ASKED FOR A UROLOGY CONSULT DUE TO A POSSIBLE THERMAL INJURY TO A URETER. A SECOND PHYSICIAN CAME IN AND CONFIRMED THERE WAS A URETERAL INJURY. | 0
4134477 | THE ROBOTIC ARM/STAPLER CLAMPED DOWN ON THE PATIENT'S COLON; IT WENT THROUGH THE INITIAL ASSESSMENT AND MADE AN ERROR CODE. THE PA TRIED TO RELEASE THE ARM FROM THE ROBOT AND IT WOULD NOT RELEASE THE CLAMP/HOLD IT HAD ON THE PATIENT'S COLON. THEY HAD TO USE THE RELEASE WRENCH TO RELEASE THE COLON FORM THE STAPLER. THERE WAS NO HARM TO THE PATIENT. STAPLER WAS REMOVED FROM THE ROBOTIC ARM AND TAKEN OFF THE FIELD.======================MANUFACTURER RESPONSE FOR ROBOTIC STAPLER, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================MANUFACTURER WAS NOTIFIED BY ROBOTICS COORDINATOR, REGA# ISSUED. | 0
3828908 | IN THE OPERATING ROOM, A PATIENT WAS UNDERGOING A ROBOTICALLY ASSISTED LEFT LOWER LOBECTOMY AND LYMPHNODE RESECTION. THE DA VINCI MARYLAND FORCEPS WERE IN USE WHEN THE OR STAFF REPORTED THAT THEY SAW SMOKE AND SMELLED A PLASTIC OR METAL ODOR. THEY ALSO REPORTED SEEING AN ENERGY SPARK COMING FROM THE DEVICE. THIS DEVICE WAS REMOVED IMMEDIATELY AND REPLACED WITH A NEW ONE. THERE WAS NO HARM TO THE PATIENT. | 0
4043825 | DURING ROBOTIC HYSTERECTOMY PATIENT SUSTAINED A PERI-URETHRAL LACERATION. THE PHYSICIAN DID A REPAIR OF THE PERI-URETHRAL LACERATION. | 0
4195430 | THE DEVICE MESSAGE DURING THE TEST PHASE STATED THAT: THE HANDPIECE WAS NOT TIGHTLY SECURED TO THE NONDISPOSABLE HANDPIECE". THE HANDPIECE AND THE DISPOSABLE DEVICE WERE CHANGED OUT AND ANOTHER NEW HANDPIECE AND DISPOSABLE WAS PLACED ON STERILE FILED | 0
3738112 | WHILE THE MARYLAND DISSECTOR INSTRUMENT WAS BEING USED INSIDE THE PATIENT, THE STAFF NOTICED A BURNING SMELL AND LATER, SMOKE. UPON INVESTIGATION, THEY NOTED THAT THE INSULATION WAS FRIED AT THE END AND APPEARED BURNED. THE IMMEDIATELY STOPPED USE AND REMOVED THE DEVICE. THERE WAS NO INJURY OR HARM TO THE PATIENT. | 0
3738115 | AN ENDOWRIST VESSEL SEALER WAS BEING USED DURING THIS CASE FROM THE BEGINNING. TWO HOURS INTO THE SURGERY, THE VESSEL SEALER DISPLAYED A MALFUNCTION MESSAGE ON THE SCREEN: BLADE MAY BE EXPOSED". DIFFICULTY WAS ALSO NOTED IN REMOVING THE INSTRUMENT FROM THE PORT; THERE WAS NO HARM TO THE PATIENT." | 0
3998300 | SURGEON WAS SUTURING THE UTERUS AT THE END OF SURGERY USING THE DEVICE IN QUESTION. THE DEVICE BROKE AND A CABLE FROM THE INSTRUMENT WAS VISIBLE TO THE SURGEON AND STAFF. THERE WAS NO INJURY TO THE PATIENT AND NO DAMAGE TO THE SURGICAL FIELD. THE DEVICE WAS REMOVED FROM THE SURGICAL FIELD AND NO DEBRIS WAS SEEN AS A RESULT OF THE DEVICE FAILURE. THE PROCEDURE WAS COMPLETED WITH A REPLACEMENT DEVICE OF THE SAME TYPE. | 0
4076758 | DURING DAVINCI HYSTERECTOMY THE VESSEL SEALER MADE THE ROBOT ALARM. ROBOT SCREEN SAID RECOVERABLE FAULT HAD OCCURRED AND SAID BLADE OF VESSEL SEALER MAY BE EXPOSED. BLADE WAS EXPOSED IN CENTER OF SEALER. INSTRUMENT REMOVED FROM PATIENT AND FROM FIELD. NO HARM TO PATIENT. | 0
3703000 | A LARGE SUTURECUT ROBOTIC INSTRUMENT MALFUNCTIONED DURING A GYNECOLOGICAL PROCEDURE. THE ROBOTIC INSTRUMENT WAS ONLY USED FOUR OUT OF TEN TIMES WHEN EXAMINED BY THE SURGICAL TECHNICIAN AND THEY NOTED THAT THE TIPS HAD WIRES STICKING OUT OF THEM. THIS DEVICE WAS TAKEN FROM THE FIELD AND REPLACED WITH A FUNCTIONING UNIT; THERE WAS NO HARM TO THE PATIENT. | 0
3975686 | WHILE PERFORMING A ROBOTIC HYSTERECTOMY, SACRAL COLPOPEXY, THE MEGA SUTURE CUT NEEDLE DRIVER BROKE WHILE IN THE PATIENT. ALL PIECES WERE RETRIEVED AND NO HARM WAS CAUSED TO THE PATIENT. | 0
4268334 | WHILE PERFORMING A ROBOTIC HYSTERECTOMY IT WAS NOTED THAT THERE WAS A HOLE IN THE BLADDER. IT WAS REPAIRED ROBOTICALLY WITH 3-0 VICRYL SUTURE. | 0
3605127 | UNKNOWN IF RELATED TO SURGERY." SMALL BOWEL OBSTRUCTION STATUS POST GASTRIC BYPASS." | 0
3811962 | IN THE OPERATING ROOM, UPON FIRST CONNECTING THE DA VINCI ROBOT BOVIE, WHEN CAUTERY WAS CONNECTED TO SCISSORS AND FORCEPS, THE CURRENT ACTIVATED WITHOUT THE SURGEON TOUCHING THE FOOT PEDAL. THE IMPROPER ACTIVATION OF THE CAUTERY INSTRUMENT RESULTED IN A SUPERFICIAL BURN TO THE PATIENT'S SIGMOID COLON APPROXIMATELY 2 CM IN LENGTH. THE BURN WAS OVER SEWN BY THE SURGEON IN ORDER TO STRENGTHEN THE AREA. THE BIPOLAR FORCEPS WERE THEN REPLACED, AND THERE WERE NO FURTHER COMPLICATIONS WITH THE DA VINCI ROBOT. THE INJURY HAPPENED WHEN THE CAUTERY CORDS WERE CONNECTED TO THE ROBOTIC FENESTRATED BIPOLAR FORCEPS AND THE MONOPOLAR SCISSORS. THE CURRENT WAS ACTIVATED ALTHOUGH THE SURGEON WAS NOT PRESSING THE CONTROL. THE CAUTERY CORDS WERE UNPLUGGED FROM THE CAUTERY UNIT. THE BURN APPEARED TO HAVE BEEN FROM THE FORCEPS; HOWEVER BOTH INSTRUMENTS WERE REMOVED FROM THE SURGICAL FIELD AND SENT THROUGH THE SUPPLY CHAIN PROCESS FOR FAULTY EQUIPMENT. THE INTUITIVE/DA VINCI REPRESENTATIVE WAS AVAILABLE AND CALLED INTO THE ROOM TO HELP WITH THE EVALUATION. THE SURGEON ALSO CONFIRMED THAT NO OTHER TRIGGERING SOURCE ON THE CAUTERY UNIT WAS CONNECTED. THE REPRESENTATIVE WAS INFORMED AT THE TIME OF THE EVENT AND WAS CALLED IN TO THE OPERATING ROOM TO ASSIST WITH EVALUATION. | 0
3912507 | THE DAVINCI ROBOT WAS IN THE DOCKED POSITION WITH THE PROGRASP FORCEPS ALREADY INSIDE AND SET. THE PHYSICIAN WAS TRYING TO CONTROL THE INSTRUMENT BUT THE INSTRUMENT WOULD NOT WORK/ARTICULATE. THE PHYSICIAN ASSISTANT WAS ON THE STERILE FIELD ASSISTING WHEN SHE NOTICED ON THE ROBOTIC SYSTEM SCREEN THAT THE WIRE ON THE INSTRUMENT WAS LOOSE. THIS INSTRUMENT WAS REMOVED FROM THE ROBOT ARM AND REPLACED WITH NEW PROGRASP FORCEPS. THE PROCEDURE WAS COMPLETED WITHOUT HARM TO PATIENT. UPON INSPECTION, NOTICED THAT THE WIRE IS INDEED LOOSE ALREADY. | 0
4212515 | DURING A ROBOTIC LEFT LOWER LOBECTOMY, A MONOPOLAR CURVED SCISSOR ENDOWRIST WAS PLACED THROUGH A TROCAR AND WHEN THE TIP WAS VIEWED IN THE CHEST CAVITY, THERE APPEARED TO BE A 1CM SECTION MISSING OF THE ORANGE/BLACK SECTION AND THE TIP WAS AT A 90 DEGREE ANGLE. AFTER THE LOBE WAS REMOVED, THE PIECE WAS LOCATED BY THE SURGEON AND APPEARED AS A COMPLETE MATCH. | 0
3792083 | PATIENT WAS UNDERGOING ROBOTIC ASSISTED LAPAROSCOPIC MYOMECTOMY. DURING THE PROCEDURE, AN ARM #3 THE SCISSORS WERE USED TO CREATE A LONGITUDINAL VERTICAL INCISION AND CONTINUED DISSECTION THROUGH THE MYOMETRIUM UNTIL THE FIBROID WAS REACHED. DURING THIS PROCESS, THE WIRE BROKE ON THE MEGA SUTURE CUT. MEGA SUTURE CUT REMOVED FROM FIELD. SURGEON LOOKED IN ABDOMEN FOR ANY BROKEN PARTS. NONE FOUND. EVENT WAS REPORTED TO INTUITIVE. PATIENT DID NOT SUSTAIN ANY INJURY AND CASE COMPLETED SUCCESSFULLY. THIS WAS A NEAR MISS. THERE WAS ONE LIFE LEFT ON THE MEGA SUTURE CUT DEVICE PRIOR TO START OF CASE.MANUFACTURER RESPONSE FOR INTUITIVE MEGA SUTURE CUT, MEGA SUTURE CUT (PER SITE REPORTER).====================== NO RESPONSE AT THIS TIME. | 0
3933947 | DURING A ROBOTIC SURGERY, THE INSTRUMENT WIRE WOULD NOT PROPERLY SET. WHEN IN USE, THE FORCEP DID NOT GRASP TISSUE. THE SURGEON TRIED GRASPING WITH THE DEVICE, HOWEVER, THE MOTION WAS NOT ACCEPTED, WHICH IS THE INTENDED USE. THE INSTRUMENT HAS 7 OUT OF 10 LIVES LEFT. | 0
4071811 | DURING DA VINCI CASE, MARYLAND DISSECTOR DID NOT WORK PROPERLY. THE TIPS WERE NOT APPROXIMATING, A DIFFERENT EQUIPMENT WAS UTILIZED. | 0
3975595 | DURING ROBOTIC SINGLE-INCISION CHOLECYSTECTOMY, THE SURGEON PLACED THE SINGLE SITE PORT AND IT BROKE IN HALF. THE SURGEON WAS ABLE TO RETRIEVE BOTH PIECES FROM THE PATIENT. | 0
3933960 | INCIDENT HAPPENED DURING SURGERY. DOCTOR WAS USING THE VESSEL SEALER AT THAT TIME WHEN AN ALARM SOUND WAS HEARD AND ON THE SCREEN. IT READ: BLADE EXPOSES VESSEL SEALER". INSTRUMENT TURNED OFF AND REMOVED RIGHT AWAY FROM PATIENT. WHEN THE ALARM SOUNDED THE VESSEL SEALER WASN'T TOUCHING ANY TISSUE. NO HARM WAS DONE TO THE PATIENT. UPON INSPECTION OF DEVICE | 0
3735600 | THE SURGEON WAS USING THE VESSEL SEALER AND THE UNIT FAULTED". INSTRUMENT WAS RE-SEATED | 0
4109348 | TWO (2) SETS OF DA VINCI VESSEL SEALERS MALFUNCTIONED DURING THIS CASE. THE ALARM STATES BLADE EXPOSED". THE BLADE WOULD NOT RETRACT BACK. ONE WAS OPENED BEFORE THE START OF THE CASE AND MALFUNCTIONED DURING THE MIDDLE OF THE CASE. THE RN SEQUESTERED THE MALFUNCTIONING EQUIPMENT AND GAVE THE- NEW VESSEL SEALER AS REQUESTED BY THE DOCTOR. THE NEW VESSEL SEALER WAS PLACED TO THE ARM AND AFTER LESS THAN 10 MINS THE ALARM WENT OFF WITH THE SAME EXACT REASON. THE SECOND VESSEL SEALER WAS THEN AGAIN SEQUESTERED BY THE RN. INCIDENT OF MALFUNCTIONED EQUIPMENT WAS REPORTED TO SUPPLY CHAIN. CHARGE NURSE MADE AWARE. AFTER THE SECOND TIME | 0
3687298 | VESSEL SEALER USED DURING ROBOTIC CYSTECTOMY. VESSEL SEALER BLADE DID NOT RETRACT BACK; INSTRUMENT FAILED TO SEAL COMPLETELY. OPENED ANOTHER VESSEL SEALER AND IT WORKED FINE, NO PATIENT INJURY. MALFUNCTIONING PRODUCT TO BE SENT BACK TO VENDOR FOR QUALITY ANALYSIS AND CREDIT/REIMBURSEMENT GIVEN. | 0
3586786 | PATIENT WAS UNDERGOING ROBOTIC ASSISTED MYOMECTOMY. SWITCHED ROBOTIC MONOPOLAR SCISSOR WITH TIP COVER ON RIGHT SIDE OF PATIENT ABDOMEN INTO ROBOTIC LARGE NEEDLE DRIVER. ROBOTIC MONOPOLAR SCISSOR WAS STUCK IN THE ROBOTIC 8MM TROCAR. REMOVED TROCAR WITH THE SCISSOR AS A UNIT FROM THE PATIENT. AREA BETWEEN SCISSOR AND SHAFT WAS LOOSE. NO OBVIOUS HARM TO PATIENT OTHER THAN SMALL EXTENSION OF OR TIME. | 0
4192392 | A SPRING BROKE INSIDE OF THE INSTRUMENT, AN 8 MM MARYLAND BIPOLAR FORCEP. THE DEVICE HAD TO BE REMOVED BECAUSE IT MALFUNCTIONED AND IN ORDER TO REMOVE THE INSTRUMENT, THE ENTIRE TROCAR IT WAS PASSED THROUGH NEEDED TO ALSO BE REMOVED DUE TO THE BROKEN SPRING. A NEW CANNULA AND INSTRUMENT WERE USED AFTER THE MALFUNCTIONED SET FAILED.THERE WAS NO HARM TO THE PATIENT. THE INSTRUMENT WAS SENT TO OUR CENTRAL STERILE PROCESSING FOR DECONTAMINATION AND FOR NOTIFICATION TO THE VENDOR. | 0
4272527 | DURING LAPAROSCOPIC SUPRACERVICAL HYSTERECTOMY DAVINCI INSTRUMENT PIECE BROKE AND A SMALLER PIECE BROKE OFF AND FELL. | 0
3641087 | BLADE TIP (ATTACHED BY SURGICAL TECHNICIAN BEFORE PROCEDURE) CAME OFF OF HANDLE DURING USE INSIDE PATIENT WHILE PERFORMING LAPAROSCOPIC HYSTERECTOMY WITH DAVINCI ROBOT. SURGEON HAD TO RETRIEVE DEVICE LAPAROSCOPICALLY TO PREVENT MAKING A LARGER INCISION. RETRIEVAL TIME WAS APPROXIMATELY 15 MINUTES RESULTING IN LONGER THAN ANTICIPATED SURGICAL AND ANESTHESIA TIME. THE DEVICE WAS RECOVERED WITHOUT DAMAGE TO INTERNAL ORGANS. | 0
3904729 | ><A PATIENT WAS BEING POSITIONED FOR AN X-RAY OF THE LUMBAR SPINE ON THE X-RAY TABLE, WHEN SUDDENLY SPARKS CAME OUT OF THE CABLE WIRES OF THE X-RAY TUBE. THE X-RAY TECH IMMEDIATELY SHUT THE MACHINE OFF. THERE WERE NO INITIAL PROBLEMS WITH THE DEVICE, TO OUR KNOWLEDGE, OTHER THAN THE FACT THE EQUIPMENT IS OLD (1993). THERE WERE NO INJURIES TO THE PATIENT OR THE X-RAY TECH. | 0
3979411 | ><RADIOLOGY TECH WAS OPERATING X-RAY MACHINE DURING CONTRAST EXAM. THE TECH ATTEMPTED TO TAKE AN IMAGE AT THE DIRECTION OF THE RADIOLOGIST, BUT THE MACHINE MISFIRED. ON THE SECOND ATTEMPT THE MACHINE TOOK AN IMAGE AS DIRECTED. | 0
4084944 | A DENTIST REPORTED TO DEALER WHO THEN REPORTED TO MIDMARK THAT THE COLUMN OF A VANTAGE X-RAY UNIT, SERIAL NUMBER (B)(4), CONTINUED MOVING UP OR DOWN AFTER THE MOVEMENT CONTROL BUTTON WAS RELEASED. THE DENTAL OFFICE REPORTED TO THE MIDMARD REPRESENTATIVE THAT THE UNIT HAS BEEN HAVING THIS ISSUE FOR A WHILE AND HAS GOTTEN WORSE LATELY. THE MIDMARK REPRESENTATIVE CALLED THE DENTAL OFFICE AND CONFIRMED THAT THERE WAS NO INJURY. | 0
3549755 | ORGANISMS TESTED: 10 GRAM NEGATIVE ORGANISMS SENT IN THE BREAKPOINT IMPLEMENTATION TOOL PACK FROM CAP. DOCUMENTED TO BE RESISTANT TO MEROPENEM. 3/10 ORGANISMS (E. COLI, MORGANELLA, CITROBACTER) GAVE FALSE SUSCEPTIBILITY RESULTS ON PHOENIX. TESTED TWICE ON PHOENIX (BD) INSTRUMENT. PHOENIX REPORTING AS SENSITIVE BOTH TIMES. THE ORGANISMS WERE TESTED ON THREE OTHER SYSTEMS (ETEST, DISK DIFFUSION AND TREK MICRO DILUTION PANEL) IN PARALLEL. ALL THESE THREE METHODS PRODUCED CORRECT RESULTS OF MEROPENEM RESISTANT.CONSEQUENCE OF THIS FINDING: WE CAN NO LONGER REPORT SUSCEPTIBLE MEROPENEM RESULTS FROM PHOENIX FOR ENTEROBACTERIACEAE. WE ARE SETTING UP KB DISK DIFFUSION IN PARALLEL. EXTRA WORK FOR TECHS. EXTRA COST FOR THE LAB (REAGENTS, PLATES ETC). WE ARE OK TO ACCEPT MEROPENEM INTERMEDIATE OR RESISTANT RESULTS FROM PHOENIX FOR THE ENTEROBACTERIACEAE. | 0
3661995 | THE CUSTOMER REPORTED APPROXIMATELY 20 LITERS OF DILUTED DETERGENT LEAKED FROM THE BECKMAN COULTER PK7300 AUTOMATED MICROPLATE SYSTEM. THE LEAK WAS NOT CONTAINED WITHIN THE INSTRUMENT AND DILUTED DETERGENT LEAKED ONTO THE FLOOR OF THE LABORATORY. THE CUSTOMER WAS WEARING PERSONAL PROTECTIVE EQUIPMENT CONSISTING OF GLOVES AND A LABORATORY COAT AT THE TIME OF THE EVENT AND NO INJURY OR DIRECT EXPOSURE TO THE LEAK WAS REPORTED. NO ERRONEOUS PATIENT RESULTS WERE GENERATED AND THERE WAS NO CHANGE OR AFFECT TO PATIENT TREATMENT IN CONNECTION WITH THIS EVENT.><ADDITIONAL INITIAL REPORTER INFORMATION : (B)(6). A FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S SITE. THE FSE EVALUATED THE INSTRUMENT AND DETERMINED THAT THE CAUSE OF THE LEAK WAS ATTRIBUTED TO A DEFECTIVE DETERGENT PUMP. THE FSE REPLACED THE DETERGENT PUMP TO RESOLVE THE ISSUE AND NO FURTHER LEAKS WERE NOTED. (B)(4). | 0
3588765 | ENDOWRIST CLIP APPLIERS WITH WECK HEM-O-LOK CLIPS WAS BEING USED WHEN THE CLIPS WOULD CLOSE BUT NOT LOCK. MULTIPLE ATTEMPTS WERE MADE WITH DIFFERENT CLIPS FROM DIFFERENT PACKS. A TOTAL OF 5 PACKS AND DIFFERENT APPLIERS WERE USED TO RULE OUT CALIBRATION ISSUES. NO BLOOD LOSS WAS NOTED AND NO INCISION WAS MADE UNTIL THE CLIPS WORKED APPROPRIATELY. | 0
3712151 | DURING A ROBOTIC EXPLORATION, THE BLADE OF THE DEVICE BROKE. | 0
3842787 | ROBOTIC CAMERA ARM LOST FUNCTIONALITY OF THE SCOPE AND CAMERA HEAD TWO HOURS INTO THE PROCEDURE. SCOPE WAS REMOVED, THE CASE WAS ABORTED AND COMPLETED LATER THE SAME DAY WHEN A REPLACEMENT ARM WAS OBTAINED BY THE VENDOR. | 0
4272267 | DURING A ROBOTIC ASSISTED LAPAROSCOPIC HYSTERECTOMY, WHILE THE SURGEON WAS OPERATING THE VESSEL SEALING DEVICE, IT FAILED TO ENGAGE IN ORDER TO SEAL THE VESSEL. ANOTHER DEVICE WAS OPENED. | 0
4373531 | THE HARMONIC SCALPEL STOPPED WORKING APPROXIMATELY 30 MINUTES INTO A LAPAROSCOPIC GASTRIC BYPASS CASE. OPERATING ROOM (OR) STAFF ATTEMPTED TO RESET THE MACHINE AND TIGHTEN THE HANDPIECE. THE INSTRUMENT WOULD NOT WORK. A NEW HARMONIC SCALPEL WAS OPENED AND USED FOR THE REMAINDER OF THE CASE. | 0
3861218 | THE INTUITIVE SURGICAL ENDOWRIST ONE VESSEL SEALER BLADE WOULD NOT RETRACT. | 0
4302665 | THIS NEW DISPOSABLE DEVICE MALFUNCTIONED. THE BLADE DID NOT FULLY RETRACT. THE DEVICE WAS REMOVED FROM SERVICE. IT WAS SENT TO OPERATING ROOM EDUCATOR AND THEN ON TO THE OPERATING ROOM MATERIALS MANAGEMENT. | 0
4134439 | THE ROBOTIC INSTRUMENT NEEDLE DRIVER (WIRING) BROKE DURING PROCEDURE. | 0
3816570 | WHILE MD WAS SUTURING DURING A ROBOTIC TOTAL LAPAROSCOPIC HYSTERECTOMY CASE, THE MEGA NEEDLE DRIVER'S CABLE SNAPPED AT THE TIP. | 0
3590834 | WHEN LOADING VESSEL SEALER INSTRUMENT INTO THE DA VINCI INSTRUMENT ARM TO USE, IT TRIGGERED A RECOVERABLE FAULT. STAFF REMOVED IT FROM THE ARM AND ATTEMPTED TO RELOAD IT A FEW TIMES TO NO AVAIL. OPENED ANOTHER ONE ONTO THE STERILE FIELD AND IT DID THE SAME THING. STAFF TRIED TO LEAD IT IN A DIFFERENT ARM BUT HAD THE SAME RESULTS. THE CIRCULATOR THEN CALLED THE INTUITIVE SURGICAL SUPPORT NUMBER. IT WAS DISCOVERED THAT THE ERBE ENERGY SOURCE HAD A RED LIGHT ON. STAFF HAD TO REBOOT THE SYSTEM AND THEN THE SECOND INSTRUMENT WORKED FINE. THE FIRST INSTRUMENT HAD BEEN HANDED OFF OF THE FIELD BY THEN. | 0
4212535 | WHILE PERFORMING A ROBOTIC COLON PROCEDURE, THE SURGEON WAS USING THE THE ROBOTIC GRASPING RETRACTOR INSTRUMENT AND THE CABLE BROKE. | 0
4095349 | DURING PROCEDURE ON THE FIRST USE THE SCISSOR TIP WAS NOT CLOSING COMPLETELY.THE DEVICE WAS GIVEN TO THE SALES REP. NO REPORT BACK AS OF THIS REPORT. | 0
3682740 | 1. ITEM DEFECTIVE - WIRE HANGING OUT.2. SEALER STUCK INSIDE TROCAR.OR STAFF MADE MANUFACTURER AWARE. | 0
4252965 | THE CABLES ON THE JAW OF THE ENDOWRIST INSTRUMENT (MEGASUTURECUT NEEDLE DRIVER) BROKE WHEN THE DOCTOR WAS SUTURING. | 0
3763165 | DURING THE SINGLE SITE ROBOTIC CHOLECYSTECTOMY, THE MONOPOLAR ROBOTIC HOOK WAS DIFFICULT TO INSERT INTO THE ROBOTIC CANNULA. A SECOND HOOK AND NEW CANNULA WERE USED AND THE SAME DIFFICULTY WAS ENCOUNTERED. | 0
4252967 | DR. WAS USING THE PROGRASP FOR HIS ROBOTIC INGUINAL HERNIA REPAIR WHEN ALL OF A SUDDEN THE ENDOWRIST INSTRUMENT STOPPED ARTICULATING AS IT SHOULD. | 0
3678599 | DURING CASE, DA VINCI 8 MM MEGA SUTURECUT NEEDLE DRIVER BROKEN WITH CABLE PROTRUDING ON THE TIP AND BOTH TIPS STILL ATTACHED. DEVICE PRE-CHECKED PRIOR TO USE, NO INDICATION OF ANY PROBLEM. WORKING FOR FEW HOURS DURING CASE WHEN SURGEON NOTIFIED SURGICAL TECHNICIAL TO REMOVE DEVICE FROM SITE. DISCOVERED BY ATTENDING SURGEON WHEN DEVICE FAILED TO FUNCTION. X-RAYS ORDERED AND TAKEN. NO VISIBLE PARTS LEFT PER X-RAY. BROKEN INSTRUMENT COMPARED TO NON-BROKEN ONE. ALL PARTS APPEARED TO BE PRESENT. THIS WAS USE 8 OF 10 FOR THIS INSTRUMENT. | 0
4088677 | DAVINCI SUTURE CUT NEEDLE HOLDER CABLE BROKE DURING OPERATION. | 0
4033899 | DURING ROBOTIC SCARAL COLPOPEXY PROCEDURE, PHYSICIAN WAS SUTURING WITH THE ROBOTIC INSTRUMENTATION AND NOTICED WHITE SPECS APPEARING IN THE OPERATIVE FILED. UPON LOOKING AROUND THE SURGICAL SITE, IT WAS NOTICED THAT THE CADIERE FORCEPS AND THE GRASPING RETRACTOR SHAFTS HAD BEEN CROSSED OVER EACH OTHER. THE INSTRUMENT SHAFTS WERE RUBBING TOGETHER AND CAUSING FLAKING TO OCCUR. THE INSTRUMENTS WERE REMOVED AND THE PATIENT'S ABDOMEN WAS THOROUGHLY IRRIGATED TO FLUSH OUT THE FLAKES. | 0
4004959 | DURING ROBOTIC ASSISTED LAPAROSCOPIC HYSTERECTOMY THE ENDOWRIST ONE VESSEL SEALER STOPPED CUTTING WHILE IN USE. | 0
3865533 | DA VINCI FORCEP JAWS ON DEVICE CLOSED AND WOULD NOT OPEN. NO FURTHER INFORMATION. | 0
4302708 | WHILE MD WAS USING THE MEGA SUTURECUT NEEDLE DRIVER, THE CABLE BROKE AT THE TIP OF THE INSTRUMENT. | 0
4117109 | THE VESSEL SEALER MALFUNCTIONED AND DID NOT RETRACT AS EXPECTED. STEPPED ON PEDAL AND RECEIVED AN ERROR MESSAGE AND LOCKED IT. | 0
4026979 | DURING A ROBOTIC INGUINAL HERNIA REPAIR CASE, THE SURGEON WAS USING THE MEGA SUTURECUT NEEDLE DRIVER AND THE CABLE AT THE TIP OF THE INSTRUMENT BROKE. | 0
3716258 | THE DAVINCI ROBOTIC MARYLAND BIPOLAR FORCEP BROKE. | 0
3772195 | THE INSTRUMENT PROGRASP DOES NOT OPEN/GRASP. | 0
3631466 | DURING THE PROCEDURE, THE SURGEON NOTED THE ROBOTIC TENACULUM BROKE AND SPRANG INTO ARIGHT ANGLE POSITION. HE WAS UNABLE TO MOVE IT, AND PIECES OF THE CABLE WERE REMOVED IN AN ENDOCATCH AND NO FURTHER CABLE WAS SEEN. THE TENACULUM WAS REMOVED. THERE WAS NO SIGN OF OTHER INSTRUMENT PARTS OR TENACULUM OR CABLE, AND AFTER THE SURGEON INSPECTED AND IRRIGATED, A CYSTOSCOPY WAS PERFORMED DUE TO THE GENERAL DIFFICULTY REMOVING THE UTERUS. THE URETERAL ORIFICES HAD GOOD STRONG EFFLUX OF BLUE-COLORED URINE, AND THERE WAS NO INJURY TO THE ACTUAL BLADDER. | 0
3763211 | DURING ROBOTIC HYSTERECTOMY THE BLADE FAILED. | 0
3915255 | UPON PREPARING FOR A ROBOTIC ASSISTED CASE, THE TECHNICIAN WAS CHECKING THE INSTRUMENTATION IN THE ROBOTIC INSTRUMENT ARM AND THE FENESTRATED BIPOLAR FORCEPS WAS NOT BEING RECOGNIZED BY THE INSTRUMENT ARM. | 0
4302615 | THE INTUITIVE SURGICAL VESSEL SEALER WAS ATTEMPTED TO FIRE AND IT FAILED. IT FIRED ON THE SECOND ATTEMPT, BUT THEN FAILED ON THE NEXT TWO CONSECUTIVE ATTEMPTS TO GET THE SEALER TO FIRE.======================MANUFACTURER RESPONSE FOR INTUITIVE SURGICAL VESSEL SEALER, DISPOSABLE VESSEL SEALER (PER SITE REPORTER).======================CURRENTLY AWAITING RESPONSE FROM MANUFACTURER. | 0
3587988 | INSTRUMENT WAS MISSING A PIECE OF BLACK INSULATION NEAR TAN TIP OF INSTRUMENT - HOOK CAUTERY.======================MANUFACTURER RESPONSE FOR HOOK CAUTERY, SINGLE SITE DAVINCI PERMANENT HOOK CAUTERY (PER SITE REPORTER).======================NO RESPONSE YET. | 0
3628377 | THE VESSEL SEALER BLADE ON THE VESSEL SEALER WASN'T WORKING PROPERLY. | 0
3951906 | INSTRUMENT TIP UNEVEN AND DOES NOT CLOSE WELL. SURGEON STATED THE INSTRUMENT NOT WORKING PROPERLY. | 0
4043835 | FACILITY HAVING CONTINUED ISSUES WITH THE DAVINCI STERILE CAMERA ARM DRAPES-UPON PREPARING THE DRAPE FOR USE ON THE STERILE FIELD, THE TECHNICIAN HAS DISCOVERED HOLES IN THE PAPER HEAT SHIELD PORTION LOCATED JUST ABOVE THE PLASTIC SCOPE HOLDER PIECE. THIS HAS OCCURRED IN EVERY TYPE OF PACKAGING: 3-ARM ACCESSARY KIT, 4-ARM ACCESSARY KIT AND THE SINGLE CAMERA ARM DRAPE. | 0
3800016 | DURING ROBOTIC ASSISTED SURGERY, THE CABLE ON THE WORKING END BROKE. THIS RESULTED IN A SLIGHT DELAY TO OBTAIN ANOTHER DEVICE. THERE WAS NO HARM TO THE PATIENT OR BLOOD LOSS. | 0
4251486 | DURING ROBOTIC SURGERY CASE, A FENESTRATED BIPOLAR FORCEP AND A MONOPOLAR CURVED SCIZZORS ROBOTIC INSTURMENTS WERE FOUND TO NOT HAVE THE RELEASE BUTTONS ON THE INSTRUMENTS. CHARGE RN WAS BALE TO GIVE VERBAL INSTRUCTIONS TO STAFF ON HOW TO MANUALLY RELEASE THE INSTRUMENT OFF THE INSTRUMENT ARM. | 0
3915222 | BIPOLAR NOT CAUTERIZING WHEN ACTIVATED. | 0
3994288 | INTUITIVE SURGICAL SINGLE SITE M-L CLIP APPLIER NOT WORKING. | 0
4029795 | DURING ROBOTIC SURGERY, DOCTOR ATTEMPTED TO TAKE CONTROL OF THE ROBOTIC ARMS BUT WAS UNABLE TO. HE TRANSFERRED THE CONTROL OVER TO THE SECOND SURGEON CONSOLE TO SEE IF THIS WAS A FAILURE ON THE PRIMARY CONSOLE BUT WAS UNABLE TO TAKE CONTROL OF THE ROBOTIC ARMS WITH THE 2ND CONSOLE. USING THE TROUBLE SHOOTING TECHNIQUE WHICH WAS EXPLAINED DURING THE TRAINING, HE PRESSED THE EMERGENCY STOP BUTTON AND THE ROBOT SHUT DOWN AND HE THEN REBOOTED THE ROBOT. HE DID NOT NEED TO REMOVE THE WRIST INSTRUMENTS FROM THE ROBOT ARMS AND DID NOT NEED TO REMOVE THE TROCARS OR WRIST INSTRUMENT TO PERFORM THIS FUNCTION. AFTER THE ROBOT WAS REBOOTED THE ROBOT PERFORMED NORMALLY DOCTOR WAS ABLE TO CONTROL THE ARMS. THE DAVINCI SERVICE ENGINEER WAS NOTIFIED OF THIS EVENT AND CHECKED THE ROBOT ERROR LOG AND FOUND NO ERROR MESSAGES. THIS EXACT SCENARIO OCCURRED THE WEEK BEFORE. AT THAT TIME THE ENGINEER WAS NOTIFIED AND FOUND NO ERROR IN THE LOG, HE CAME TO THE HOSPITAL TO EVALUATE THE ROBOT AND FOUND NO MALFUNCTION. THE SECOND EVENT SHOWED NO MALFUNCTION AND THE ENGINEER INDICATED THAT HE BELIEVES IT WAS THE MOTHER BOARD THAT NEEDED REPLACEMENT. HE FELT THE COMMUNICATION FROM THE MOTHER BOARD TO MOVE THE ARMS WAS NOT WORKING PROPERLY. | 0
4116745 | JAWS ON INSTRUMENT ARE MISALIGNED. | 0
3792074 | PATIENT WAS UNDERGOING A ROBOTIC NISSEN FUNDOPLICATION AND HIATAL HERNIA REPAIR. DURING THE PROCEDURE, WHILE IN USE, THE TIP OF THE HARMONIC SCALPEL BROKE OFF INTO THE PATIENT. THE TIP WAS RETRIEVED WITHOUT ANY HARM TO THE PATIENT. | 0
3628331 | THE ROBOTIC VESSEL SEALER DOES NOT CUT. | 0
3666467 | ROBOTIC VESSEL SEALER PART NUMBER 410322 BY INTUITIVE SURGICAL FAULTED DURING CASE. THIS IS THE THIRD ONE THAT HAS FAILED OUT OF A BOX OF SIX. THEY ALL HAVE THE SAME LOT NUMBERS AND EACH ONE FAILED WITH EXPOSED BLADE MESSAGE. IT WAS 2 DIFFERENT SURGEONS ON 2 DIFFERENT TYPES OF CASES. THERE WAS NO PATIENT INJURY DUE TO FAILURE. | 0
3682726 | THE BLADE ON VESSEL SEALER STAYS EXPOSED AND WILL NOT RETRACT. | 0
3644122 | THE DAVINCI ROBOT INSTRUMENT WOULDN'T CLOSE OR STRAIGHTEN COMPLETELY WHILE IN PATIENT DURING THE REMOVAL OF INSTRUMENT. THE INSTRUMENT WAS CLOSED MANUALLY, BUT STILL WOULDN'T STRAIGHTEN COMPLETELY. THE INSTRUMENT WAS REMOVED THROUGH METAL PORT WITH NO HARM TO PATIENT. THE SHEATH WAS REMOVED FROM INSTRUMENT. THE INSTRUMENT WAS FOUND TO BE BROKEN. | 0
4008150 | THE BLADE WAS EXPOSED UPON USE. SURGEON STATED THE INSTRUMENT NOT WORKING PROPERLY. | 0
3911850 | DURING A ROBOTIC SURGERY, THE FIRST ASSIST PLACED THE DAVINCI ROBOTIC PROGRASP FORCEPS INTO THE INSTRUMENT ARM PER THE SURGEON'S REQUEST. UPON VIEWING THE TIP OF THE INSTRUMENT HE IMMEDIATELY NOTICED THAT THE CABLE WAS FRAYED AT THE WRISTED JOINT. | 0
4156059 | THE SYSTEM WAS NOT RETURNED FOR EVAL, HOWEVER, A TREND DOWNLOAD WAS PERFORMED AND THE DATA WAS REVIEWED. PER THE SYSTEM CONFIGURATION INFO THE SYSTEM WAS RUNNING IN CENTRAL MONITORING MODE. PER THE TREND DATA, THE SYSTEM WAS MONITORING AND GENERATING ALARMS ON THE SAFETY NET VIEW DURING THE PERIOD IN QUESTION AS DESIGNED.><IT WAS REPORTED BY THE CUSTOMER THAT A PT DIED WHILE BEING MONITORED ON THE PT SAFETY NET SYSTEM DURING THE NIGHT. THE PT WAS ADMITTED TINT THE SYSTEM AND THE ALARMS WERE NOTED AS ACTIVE AT THE NURSING STATION AS WELL AS THE BEDSIDE. INFO WAS MADE AVAILABLE TO MASIMO ON (B)(6) 2014 THAT THE PT SUFFERED AN EMBOLISM THAT CAUSED ALARMS ON BOTH THE VIEW AND THE BEDSIDE MONITOR. | 0
3760716 | ><MOTHER WAS IN LABOR, WITH NORMAL PROGRESSIONS FOR BOTH THE MOM AND NEONATE. THE CERNER FETALINS SYSTEM WENT DOWN AFTER POWER SURGE WITH A < 1 SECOND OUTAGE. AFTER THE SYSTEM WENT DOWN, A REGISTERED NURSE (RN) WAS ASSIGNED FULL-TIME TO THE ROOM. NEWBORN ENTERED A DISTRESSED STATE, THE RN SHOUTED FOR ASSISTANCE AND A DOCTOR WAS CALLED. STAFF SPEAKING WITH DOCTOR WAS NOT ABLE TO GIVE FULL INFORMATION FROM THE NURSES STATION AS THE CENTRAL MONITOR WAS DOWN. THE DOCTOR RESPONDED AND STARTED THE DELIVERY. FULL INFO WAS NOT AVAILABLE DURING THE DELIVERY AS THE VIEW OF THE FETAL MONITOR WAS BLOCKED BY STAFF WORKING WITH THE PATIENT. DELIVERY WAS ACCOMPLISHED SUCCESSFULLY, WITH AN APGAR OF 8/9 BECAUSE THE INFANT WAS NOTED TO HAVE A BROKEN CLAVICAL INCIDENTAL TO THE DELIVERY. THIS IS SUBMITTED AS A POTENTIAL NEAR MISS IF STAFFING DID NOT ALLOW A RN FULL TIME IN THE ROOM AFTER THE FAILURE OF THE MONITORING SYSTEM, AN ADVERSE EVENT MAY HAVE OCCURRED. THE PROBLEM WITH THE SYSTEM WAS THE POWER FAILURE CHANGED THE TIME ON ONE OF THE CERNER DEVICES TO OUTSIDE THE +/- 30 SECOND WINDOW PERMITTED. IT RESOLVED THE ISSUE IN APPROXIMATELY 2 HOURS (EASY FIX, TROUBLESHOOTING TOOK SOME TIME) USING A COLLABORATIVE TEAM, CERNER, CORPORATE IT, INFORMATICS, BIOMEDICAL STAFF. MANUFACTURER IS INVOLVED IN RESOLUTION OF THE PROBLEM. AS STATED ABOVE THIS IS A HEADS-UP REGARDING TECHNOLOGY FAILURE AND THE POTENTIAL FOR SIGNIFICANT PATIENT HARM. OUTCOME ON THIS PROCEDURE WAS NOT SIGNIFICANTLY CHANGED BY THE FAILURE OF THE EQUIPMENT. | 0
3631473 | GI CAPSULE PROCEDURE FAILURE AND THE NEED TO RESCHEDULE PATIENT. THIS PATIENT WAS SCHEDULED FOR AN UPPER GI ENDOSCOPY WITH THE PLACEMENT OF A GI CAPSULE (PILL CAMERA). THE PLACEMENT OF THE PILL, RATHER THAN SWALLOWING THE PILL, WAS NEEDED IN HIS CASE DUE TO AN ABNORMALITY WITH HIS ANATOMY. HE WILL BE HAVING SURGERY TO REPAIR THIS DEFECT AFTER THE GI CAPSULE STUDY IS COMPLETED. UNFORTUNATELY, WHEN THE PILL WAS PLACED THE RECEIVER BELT WAS NOT DETECTING AND RECORDING. THE NURSE GOT ANOTHER BELT AND REPLACED THE ORIGINAL BELT. AT THIS POINT THE RECORDING BEGAN, HOWEVER, THERE WAS APPROXIMATELY 20 MINUTES THAT CONTAINED NO IMAGE. IN ADDITION, ONCE THE VIDEO WAS DOWNLOADED TO A CD AND VIEWED THE PILL HAD TRAVERSED RAPIDLY INTO THE COLON SO THAT THE AREA OF CONCERN, THE SMALL BOWEL, WAS NOT CAPTURED. | 0
3908387 | CAPSULE AND DEVICE SUCCESSFULLY ADMINISTERED AND APPLIED. UPON RETRIEVING DATA RECORDER, NO INFORMATION RECORDED. | 0
4176236 | THE PHYSICIAN INFORMED THE NURSE THAT A GASTROINTESTINAL (GI) PILL STUDY DID NOT FULLY DOWNLOAD AND THAT ALMOST ALL OF THE STUDY WAS NOT RECORDED AND DID NOT DOWNLOAD. GIVEN TECHNICAL SUPPORT WAS CALLED AND A TICKET NUMBER WAS CREATED. THE STUDY WAS DOWNLOADED ONTO A CD AND IS BEING SENT TO GIVEN TECHNICAL SUPPORT SO THEY CAN RESEARCH THE PROBLEM. | 0
3596520 | WENT TO ADMINISTER GI PILL TO PT FOR ANEMIA. PHYSICIAN NOTIFIED THAT PT WAS ON A TELEMETRY UNIT AND IT IS KNOWN THAT TELEMETRY MAY INTERFERE WITH PILL CAPTURING/RECORDING. PHYSICIAN SAID TO GO AHEAD WITH PROCEDURE. PT ATTACHED TO SENSOR BELT AND GI PILL MONITOR/RECORDER. RECORDER DEVICE WOULD BLINK RED, YELLOW, AND BLUE AT DIFFERENT TIMES. OTHER ENDOSCOPY NURSE CALLED AND CAME TO LOOK AT DEVICE. MD CALLED AND NOTIFIED OF PILL NOT RECORDING CORRECTLY. MD ORDERED TO D/C TELEMETRY ON PATIENT AND CONTINUE WITH PROCEDURE. GI RECORDING DEVICE WAS BLINKING BLUE CORRECTLY AGAIN AND THE GI PILL WAS SWALLOWED BY THE PT. DEVICE WAS WORKING PROPERLY WHEN WE LEFT. ENDOSCOPY UNIT WAS CALLED BY PT AND MD TO INFORM THAT THE MONITOR/RECORDING DEVICE STOPPED WORKING AND WAS HAD NO LIGHTS BLINKING ANYWHERE ON IT. RN CALLED GIVEN REP AND REP SAID TO TRY SWITCHING OUT SENSOR BELTS. RN WENT TO PT'S ROOM, REINITIALIZED/TURNED BACK ON MONITOR/RECORDING DEVICE AND SWITCHED SENSOR BELTS. THE MONITORING DEVICE WAS NOT SHOWING ANY LIGHTS/SIGNALS IN THE UPPER RIGHT HAND AREA WHERE THE PILL SENSOR IS. RN CALLED REP AGAIN AND WAS TOLD TO CALL GIVEN TECH SUPPORT. CALLED AND SPOKE WITH TECH SUPPORT AND WAS TOLD TO TAKE ANOTHER SENSOR BELT UP TO THE PT AND CALLED BACK WHEN WITH THE PT. WENT BACK TO PT'S ROOM WITH ANOTHER SENSOR BELT AND CALLED TECH SUPPORT AGAIN AND WAS TOLD TO DETACH PT FROM MONITOR AND TAKE MONITOR BACK TO THE ENDOSCOPY CENTER WHERE THE BASE/CRADLE IS AND TO CALL TECH SUPPORT BACK AGAIN. ONCE BACK IN THE ENDOSCOPY CENTER, CALLED TECH SUPPORT AGAIN AND FOLLOWED THE DIRECTIONS TO TROUBLE SHOOT. WAS INFORMED THAT MONITOR WAS NOT CAPTURING CORRECTLY AND PT WOULD NOT BE ABLE TO COMPLETE STUDY DUE TO MALFUNCTION OF THE RECORDING DEVICE. MD AND PT NOTIFIED THAT PROCEDURE CANNOT BE FINISHED AND MAY NOT HAVE BEEN RECORDED DUE TO MALFUNCTION, PT TOLD TO CALL MD'S OFFICE TOMORROW. FOLLOWING TECH SUPPORT INSTRUCTIONS, CD WAS BURNED OF THE GI PILL STUDY THAT DID CAPTURE AND SENT TO GIVEN TECH SUPPORT. ALL SENSOR BELTS WERE TESTED PER TECH SUPPORT INSTRUCTIONS AND WORKING PROPERLY. RECORDER/MONITORING DEVICE WAS REMOVED, TAGGED, AND GIVEN TO CUC. GIVEN TECH SUPPORT TICKET # 300283272. GI CAPSULE ID: SW8-BSM-R/LOT #2013-27/22428S. | 0
3959233 | ><PATIENT WAS HAVING A CHEST X-RAY AS FOLLOW UP TO PNEUMONIA. TWO EXPOSURES WERE TAKEN, BOTH SHOWED NO IMAGES POST PROCESSING. NO ERROR CODE WAS GIVEN. THE SYSTEM ACTED AS NORMAL. THE PATIENT WAS MOVED INTO ANOTHER ROOM AND RE-EXPOSED FOR ANOTHER CHEST X-RAY. | 0
3962875 | ><PT WAS STANDING FOR UPRIGHT X-RAY. THE TECHNOLOGIST TOOK IMAGE BUT CONTROL PANEL MADE AN UNUSUAL SOUND DISPLAYING A VERY RAPID EXPOSURE. TECHICIAN PROCESSED IMAGE, BUT NO IMAGE SHOWED (MEANING EQUIPMENT FAILURE). TECHNICIAN REPOSITIONED PATIENT AND TOOK ANOTHER X-RAY. THE CONTROL PANEL MADE THE SAME UNUSUAL SOUND AND DISPLAYED A VERY RAPID EXPOSURE. THE IMAGE WAS PROCESSED, BUT AGAIN, THERE WAS NO IMAGE. THE PATIENT WAS TAKEN TO ANOTHER X-RAY ROOM TO FINISH THE EXAM. | 0
4275961 | ><MANUFACTURE DATE UNKNOWN. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
4195512 | CARTRIDGES FOR QUALITY CONTROL CHECKS ARE INDICATING NEED FOR REPLACEMENT WHEN THEY DO NOT NEED TO BE REPLACED. WHEN ONE OF THE 3 IS REPLACED, THE OTHER ONES NO LONGER WORK. ACCORDING TO SIEMENS, WHEN ONE IS REMOVED IT CREATES A MICROGAP SO THAT THE MACHINE THINKS THE OTHER CARTRIDGE IS NOT WORKING. SIEMENS HAS RECOMMENDED THAT STAFF HOLD THE OTHER TWO CARTRIDGES IN PLACE WITH THEIR HANDS WHEN THEY REMOVE ONE SO NO MICROGAP IS CREATED, WHICH HAS SEEMED TO WORK. SIEMENS HAS BEEN CREDITING THE FACILITY WITH THE AMOUNT LEFT IN THE CARTRIDGES. FIELD SERVICE HAS BEEN OUT TO LOOK AT THE MACHINES AND INDICATES THEY ARE WORKING CORRECTLY. RESPIRATORY THERAPISTS GO TO THE OTHER ABG MACHINE IN THE FACILITY WHEN ONE DOES NOT WORK SO NO KNOWN PATIENT INJURY HAS RESULTED.======================MANUFACTURER RESPONSE FOR ABG MACHINE, RAPIDPOINT 405 (PER SITE REPORTER).======================SIEMENS HAS RECOMMENDED THAT STAFF HOLD THE OTHER TWO CARTRIDGES IN PLACE WITH THEIR HANDS WHEN THEY REMOVE ONE SO NO MICROGAP IS CREATED, WHICH HAS SEEMED TO WORK. | 0
4195509 | CARTRIDGES FOR QUALITY CONTROL CHECKS ARE INDICATING NEED FOR REPLACEMENT WHEN THEY DO NOT NEED TO BE REPLACED. WHEN ONE OF THE 3 IS REPLACED, THE OTHER ONES NO LONGER WORK. ACCORDING TO SIEMENS, WHEN ONE IS REMOVED IT CREATES A MICROGAP SO THAT THE MACHINE THINKS THE OTHER CARTRIDGE IS NOT WORKING. SIEMENS HAS RECOMMENDED THAT STAFF HOLD THE OTHER TWO CARTRIDGES IN PLACE WITH THEIR HANDS WHEN THEY REMOVE ONE SO NO MICROGAP IS CREATED, WHICH HAS SEEMED TO WORK. SIEMENS HAS BEEN CREDITING THE FACILITY WITH THE AMOUNT LEFT IN THE CARTRIDGES. FIELD SERVICE HAS BEEN OUT TO LOOK AT THE MACHINES AND INDICATES THEY ARE WORKING CORRECTLY. RESPIRATORY THERAPISTS GO TO THE OTHER ABG MACHINE IN THE FACILITY WHEN ONE DOES NOT WORK SO NO KNOWN PATIENT INJURY HAS RESULTED. | 0
4172768 | THE PFA INSTRUMENT USED FOR CLOSURE TIME TESTING IS NOT OPERATIONAL. SPECIMENS ARE BEING SENT TO A SISTER LABORATORY FOR TESTING. A COUPLE HOUR DELAY IN RESULTS MAY BE DUE TO TRANSPORT TIME. EVERY EFFORT IS BEING MADE TO SECURE A LOANER INSTRUMENT BUT ARE UNABLE TO AT THIS TIME.======================MANUFACTURER RESPONSE FOR CLOSURE TIME DEVICE, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================LOANER BECAME AVAILABLE FROM MFR. OBTAINED LOANER AND PLACED IN SERVICE. | 0
3817599 | CABLE BROKE ON DA VINCI INSTRUMENT WHILE IN USE. NO PATIENT HARM. | 0
4120797 | ><A 3 CC LUER LOCK SYRINGE WITH 25 GAUGE 1 INCH NEEDLE CRACKED WHEN THE MEDICAL ASSISTANT BEGAN PUSHING THE PLUNGER TO INJECT A SUBCUTANEOUS MEDICATION. THE MEDICATION SPRAYED ONTO THE PATIENT'S UPPER BODY AND ON THE MEDICAL ASSISTANT. | 0
3803885 | THE SIEMENS CUSTOMER CARE CENTER (CCC) WAS CONTACTED BY THE CUSTOMER. AFTER DATA EVALUATION, THE CCC DETERMINED THAT THE CUSTOMER IS UNABLE TO SET THE CHLORIDE ASSAY TO ENABLED ON ONE INSTRUMENT PORTFOLIO ON THE SYNGO LAB DATA MANAGER. THE CCC ALSO DETERMINED THAT IF THE CUSTOMER IS UNABLE TO RUN THE CHLORIDE ASSAY, THE SAMPLE WOULD NEED TO BE MANUALLY RUN ON ANOTHER INSTRUMENT. SIEMENS IS CURRENTLY INVESTIGATING THIS ISSUE.><THE CUSTOMER STATED THAT THEY WERE UNABLE TO HAVE CERTAIN ASSAYS ENABLED ON ONE INSTRUMENT WITH THE SYNGO LAB DATA MANAGER. THE CUSTOMER STATED THERE IS A POTENTIAL FOR THE OPERATOR TO BE UNAWARE OF AN ASSAY NOT PROCESSED BECAUSE IT WAS QUERIED BY AN INSTRUMENT THAT COULD NOT PERFORM THE ASSAY. THERE ARE NO REPORTS OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCES DUE TO CERTAIN ASSAYS NOT BEING ENABLED ON THE INSTRUMENT. | 0
3893468 | ><PATIENT WENT TO MRI AT 1300. PRIOR TO LEAVING, TEMPERATURE WAS DOCUMENTED AT 1200 AS 36.9. UPON RETURN FROM MRI AT 1500, TEMPERATURE WAS DOCUMENTED AS 39.0. PATIENT WAS ON THE PREVALON TURN AND POSITION SYSTEM DURING MRI. NO APPARENT INJURY. PATIENT MONITORED.DEVICE INSTRUCTIONS FOR USE DO NOT INDICATE THAT THE DEVICE CANNOT BE USED IN THE MRI SUITE. NO INDICATIONS REGARDING MRI COMPATABILITY ARE LISTED ON THE DEVICE INFORMATION SHEETS. | 0
3880655 | ><REPORT RECEIVED OF A USER ERROR RESULTING IN A SUCTION SWAB DISENGAGEMENT. REPORTER STATED THAT ON (B)(6) 2014, PT HAD A SCHEDULED ESOPHAGOGASTRODUODENOSCOPY (EGD) THE REASON FOR PROCEDURE WAS UNK BY REPORTER. DURING THE ESG, AN OBJECT DESCRIBED AS A PORTION OF GREEN FOAM WITH A SMALL WHITE PLASTIC TUBE IN THE CENTER WAS NOTED IN THE PT'S STOMACH. REPORTER IDENTIFIED THE OBJECT AS THE TIP OF A SAGE SUCTION SWAB. REPORTER STATED THE OBJECT WAS REMOVED SO THAT IT WOULD NOT OBSTRUCT THE DRAINAGE FLOW OF PT'S NASOGASTRIC TUBE. REPORTER STATED UPON VIEWING OBJECT, IT WAS EVIDENT THE PT HAD BITTEN OFF THE TIP OF THE SUCTION SWAB. REPORTER DESCRIBED INVOLVED PT AS CONFUSED, DISORIENTED, UNCOOPERATIVE WITH A HISTORY OF BITING. REPORTER INDICATED THAT WHILE HOSPITALIZED, THERE WAS NO REPORT OF THE PT BITING OFF A SUCTION SWAB. REPORTER STATED SHE WAS CERTAIN INCIDENT OCCURRED IN PT'S HOME PRIOR TO HOSPITALIZATION ALTHOUGH THE FAMILY NEVER MENTIONED THAT FACT, WHEN ASKED ABOUT THE LOT NUMBER, EXPIRATION DATE, AND SCU NUMBER, AS REPORTED ON FORM 3500, REPORTER STATED SHE DID NOT KNOW THE DETAILS FROM WHAT KIT THE SWAB CAME. REPORTER EXPLAINED SHE SELECTED A SAGE BRAND #(B)(4) Q4 COMPLETE CARE KIT FROM THE STOREROOM, AS IT CONTAINED SUCTION SWABS, AND COMPLETED SECTION E OF THE FORM 3500 USING THAT KIT'S LOT NUMBER, EXPIRATION DATE, AND SCU. SHE DID NOT INQUIRE OF THE FAMILY ABOUT THE BRAND OF PRODUCT USED AT HOME. THROUGH REQUESTED, NO ADDITIONAL INFO WAS AVAILABLE. | 0
3792992 | ><THE GREEN PORTION OF THE SUCTION SWAB HAD TO BE EXTRACTED DURING AN EGD PROCEDURE AS IT WAS OBSTRUCTING THE NASOGASTRIC TUBE PLACEMENT. THE FOREIGN BODY WAS BROKEN OFF FROM THE STICK. | 0
4247010 | ATTACHED PHOTOS SHOW THE CRACKS IN THE HOOD. BECAUSE THE REPORTED INCIDENT OCCURRED WITH TWO VP2000 INSTRUMENTS, A SEPARATE MDR 3005248192-214-00073 WAS ALSO SUBMITTED ON (B)(4) 2014 FOR THE FIRST DEVICE. FOLLOW-UP MDR REPORTS WILL BE SUBMITTED AFTER ELEVATED COMPLAINT INVESTIGATION 4265-ECINV CONCLUDES.><THE ABBOTT VP 2000 PROCESSOR IS A DEVICE DESIGNED TO AUTOMATE AND STANDARDIZE SLIDE SPECIMEN PROCESSING AND ROUTINE SLIDE STAINING FOR THE LABORATORY. A CUSTOMER REPORTED THAT BOTH THE FRONT AND BACK HOODS ON TWO VP2000 INSTRUMENTS HAVE CRACKED. THE CRACKS FIRST APPEARED APPROXIMATELY TWO WEEKS AFTER THE INSTRUMENTS WERE INSTALLED AND PROGRESSED WITH TIME. AN ABBOTT FIELD SERVICE ENGINEER (FSE) REPLACED THE FRONT HOOD AND REAR HOOD OF EACH INSTRUMENT AND VERIFIED THAT THE HOODS OPENED AND CLOSED PROPERLY. NO DEATH OR INJURY WAS REPORTED. ALTHOUGH A PRODUCT MALFUNCTION HAS NOT BEEN CONFIRMED, THE CRACKS IN THE HOOD HAVE THE POTENTIAL TO EXPOSE THE USER TO SHARP EDGES AND POTENTIAL USER INJURY SUCH AS CUTS OR ABRASIONS. BECAUSE THE HOODS ARE USED TO COVER SLIDES CONTAINING BIOHAZARDOUS MATERIALS AND BECAUSE THE HOOD MATERIAL ITSELF MAY HAVE REMNANTS OF BIOHAZARDOUS MATERIALS ON ITS SURFACE, THERE IS POTENTIAL FOR THE USER TO SUSTAIN A SERIOUS INJURY FROM A SHARP EDGE OF THE CRACKED HOOD AND TO REQUIRE FURTHER TREATMENT BECAUSE OF EXPOSURE TO BIOHAZARDOUS MATERIALS. THIS INCIDENT IS REPORTED BASED ON THE POTENTIAL FOR SERIOUS INJURY. ELEVATED COMPLAINT INVESTIGATION (B)(4) IS IN PROCESS.><SUMMARY OF (B)(4) FOR MDR 3005248192-2014-00073 FOLLOW-UP REPORT 1: NOTE: A SEPARATE FORM 3500A (MDR 3005248192-2014-00073 FOLLOW-UP REPORT 1) WAS COMPLETED FOR A SECOND VP2000 DEVICE (SERIAL NUMBER (B)(4)) REPORTED AT THE SAME CUSTOMER SITE. THE FOLLOWING ELEVATED COMPLAINT INVESTIGATION SUMMARY INCLUDES BOTH DEVICES. INVESTIGATION INTO THIS COMPLAINT INCLUDED: QUALITY DATA REVIEW, RETURN EVALUATION AND COMPLAINT HISTORY REVIEW. EXISTING DATA REVIEW: PICTURES ATTACHED IN EACH COMPLAINT WERE REVIEWED TO IDENTIFY CORRELATION TO THE INFORMATION PROVIDED WITHIN THE COMPLAINTS ACTIVITY" TAB. MULTIPLE CRACKS COULD BE SEEN ON THE HOODS OF THESE INSTRUMENTS IN THE ATTACHED PICTURES. THE VP2000 PROCESSOR OPERATIONS MANUAL (30-608384 | 0
4101754 | (B)(4). IT HAS BEEN COMMUNICATED THAT THE DEVICE IS NOT AVAILABLE FOR EVALUATION. WITHOUT THE DEVICE, IT IS NOT POSSIBLE FOR CODMAN TO CONDUCT A PROPER INVESTIGATION. SINCE A LOT NUMBER HAS BEEN PROVIDED, A REVIEW OF THE MANUFACTURING RECORDS WILL BE REVIEWED. WE ANTICIPATE THAT THE EVALUATION WILL REVEAL THAT THE DEVICE CONFORMED TO SPECIFICATIONS PRIOR TO RELEASE. IF ANYTHING OTHERWISE IS FOUND THEN A FOLLOW UP REPORT WILL BE FILED. IF AT SOME POINT THE DEVICE DOES BECOME AVAILABLE, THIS COMPLAINT WILL BE RE-OPENED, EVALUATED AND A FOLLOW UP REPORT WILL BE FILED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR COMPLAINTS. AT THE PRESENT TIME, THIS COMPLAINT IS CONSIDERED CLOSED. DEVICE NOT RETURNED.><THE AFFILIATE REPORTED: PATIENT RETURNED TO THEATRE FOR ICP MONITOR TO BE CHANGED (LE3382), AS THE ICP WAS RANGING IN MEASUREMENT OVERNIGHT. NEW ICP MONITOR (LE10302) AND MICRO SENSOR WERE USED. THE SENSOR WAS ZEROED CORRECTLY, SUTURED AND CLOSED. THE ICP RANGED FROM 10 ? 22 MMGH. SURGEON WAS CONCERNED ABOUT THE RANGE OF THE ICP READING AND CHANGED OVER TO ANOTHER ICP (LE3751) AND NEW MICRO SENSOR. THE SENSOR ZEROED, THE ICP READING INITIALLY FLUCTUATED THEN READ -99. DUE TO THIS, THE SURGEON WAS STILL CONCERNED AND PLACED IN AN EVD. (DURING USE ON PATIENT).><THE FOLLOWING PRODUCTS WERE REPORTED UNDER THIS COMPLAINT: P/N 82-6631; S/N (B)(4) - RETURNED TO CODMAN FOR EVALUATION. P/N 82-6631; S/N (B)(4) - RETURNED TO CODMAN FOR EVALUATION. P/N 82-6631; S/N UNK. - NOT RETURNED TO CODMAN FOR EVALUATION. P/N 82-6635; S/N (B)(4) - NOT RETURNED TO CODMAN FOR EVALUATION. P/N 82-6635; S/N (B)(4) - NOT RETURNED TO CODMAN FOR EVALUATION. FOR THE ICP EXPRESS UNIT P/N 82-6635; S/N (B)(4), THE PRODUCT WAS NOT RETURNED TO CODMAN FOR EVALUATION. IN THE ABSENCE OF THE COMPLAINT SAMPLE, THE MANUFACTURING RECORDS WERE REVIEW AND IT WAS VERIFIED THAT THIS UNIT WAS CONFORMING TO THE REQUIRED SPECIFICATIONS WHEN RELEASED FROM THE MANUFACTURING ON JANUARY 21, 2000. COMPLAINT IS CONSIDERED CLOSED AT THIS TIME. IF THE COMPLAINT SAMPLES BECOME AVAILABLE, THIS COMPLAINT WILL BE REOPENED AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS OR SIMILAR COMPLAINTS. DEVICE NOT RETURNED. | 0
3828984 | THE MANUFACTURER'S INVESTIGATION IS ON-GOING. ELEKTA WILL PROVIDE FURTHER INFORMATION WHEN THE INVESTIGATION IS COMPLETE.><CHANNEL ARTEFACT SPREAD BY FILTERING TO RESEMBLE EPILEPTIC ACTIVITY. THE CUSTOMER REPORTED THAT ONE CHANNEL HAS LOW-AMPLITUDE (ORDER OF NEURONAL SOURCE) ARTEFACT OCCURING EVERY NOW AND THEN. TSSS PROCESSING USING MAXFILTER-2.2 DID NOT REMOVE THIS ARTEFACT, BUT SPREAD IT TO CLOSE-BY CHANNELS. ADDITIONAL TEMPORAL FILTER IN GRAPH (4-30 HZ) SMOOTHED THE TIME COURSE OF THE ARTEFACTS TO RESEMBLE EPILEPTIC INTER-ICTAL SPIKES. THE ARTEFACTS LOCALIZED TO THE FRONTAL CORTEX, TO THE LOCATION THAT WELL MATCHED WITH THE CLINICAL PROFILE OF THE PATIENT. OTHER EPILEPTIC ACTIVITY WAS NOT FOUND. WITHOUT LOOKING AT THE RAW DATA, THIS ARTEFACTUAL SOURCE COULD HAVE BEEN ASSUMED TO BE EPILEPTIC SOURCE. THE CASE DISPLAYS KNOWN BUT RARE BEHAVIOUR OF MAXFILTER PREPROCESSING: IF BAD CHANNEL IS NOT MARKED/DETECTED AND THE ARTEFACT IS LOW IN AMPLITUDE, SO THAT IT DOES NOT HAVE ENOUGH POWER TO BE DETECTED BY THE TEMPORAL EXTENSION OF TSSS, THE SPATIAL FILTER (I.E. SSS) WILL SMOOTH THE ARTEFACT TO RESEMBLE REAL MAGNETIC FIELD. TYPICALLY THE ARTEFACT IS SUPPRESSED BY EITHER DETECTING THE CHANNEL AS BAD OR BY TERMPORAL PROJECTION (IN TSSS).><THE MANUFACTURER'S INVESTIGATION CONCLUDED THAT THE PROBLEM WAS THAT BOTH AUTOMATIC BAD CHANNEL DETECTION AND THE TEMPORAL SSS (TSSS) FAILED TO RECOGNIZE SHORT-DURATION JUMPING ARTEFACTS ON A SINGLE MEG CHANNEL. THEREFORE, MAXFILTER TSSS PROCESSING CAUSED SPREADING OF SUCH SUDDEN JUMPS ON SINGLE CHANNEL AS ARTEFACTS TO SEVERAL MEG CHANNELS. THE CHARACTERISTICS OF THE BAND-PASS FILTERED ARTIFACT RESEMBLED AN EPILEPTIC SPIKE. THE USER PERFORMED TYPICAL ANALYSIS OF LOCALIZATION OF EPILEPTIC ACTIVITY ON THE DATA. THE MEG RESULTS WERE IN ACCORDANCE WITH THE CLINICAL PICTURE OF THE PATIENT, AND NO OTHER (REAL) EPILEPTIC ACTIVITY WAS FOUND IN THE DATA. HOWEVER, USER SUSPECTED THE VALIDITY OF THE RESULTS BY USING NORMAL CLINICAL PRACTICES AND BY FOLLOWING INSTRUCTIONS, E.G. LACK OF CORRESPONDING ACTIVITY IN THE EEG DATA MEASURED SIMULTANEOUSLY WITH MEG DATA. USER CONSULTED ELEKTA PERSONNEL, AND THE ACTIVITY WAS CONFIRMED TO BE AN ARTIFACT BY REVIEWING THE RAW DATA TO WHICH MAXFILTER PROCESSING WAS APPLIED, AS DESCRIBED IN MAXFILTER USER MANUAL AND ACNS AND ACMEGS GUIDELINES. AFTER MANUALLY SELECTING THE ILL-BEHAVING CHANNEL, AND OMITTING IT FROM MAXFILTER PROCESSING, NO ARTIFACT REMAINED IN THE PROCESSED DATA. NO HARM WAS REALIZED TO THE PATIENT. EVEN IF THE ANALYSIS HAD BEEN ACCEPTED BY THE CLINICAL USER, THE RESULT WOULD NOT HAVE BEEN TYPICALLY USED ALONE FOR THE PRESURGICAL PLANNING, BUT ONLY IN AGREEMENT WITH OTHER IMAGING MODALITIES AND IN REFERRING PHYSICIAN'S CAREFUL CONSIDERATION. THE ROOT CAUSE WAS THAT MAXFILTER DID NOT RECOGNIZE THE PROBLEMATIC CHANNEL MEG1023 AS A BAD ONE, BECAUSE: 1. THE AUTOMATIC BAD CHANNEL SCANNING DID NOT FIND THE ARTIFACT DUE TO LARGE DISTURBANCES CREATED BY THE VNS DEVICE. 2. THE JUMPS HAVE SHORT DURATION (<100 MS) AND OCCURRED INFREQUENTLY. THEREFORE, TSSS-PROCESSING IN MAXFILTER DID NOT INCLUDE THE ARTIFACT SHAPE IN THE DISTURBANCE WAVEFORMS WHICH WERE PROJECTED OUT WHEN USING DEFAULT PARAMETER SETTINGS. THUS, MAXFILTER PROCESSING DID NOT REMOVE THE ARTIFACT. FIELD CORRECTIVE ACTION (FCO 145 036 01 113) WAS RELEASED TO CUSTOMERS JUNE 2014. USER NOTICE RELEASED TO CUSTOMERS CLARIFIES THE GUIDANCE PROVIDED BY THE INSTRUCTIONS FOR USE, PROVIDING SPECIFIC GUIDANCE ON HOW TO INSPECT THE OUTPUT DATA IN ORDER TO DETECT SIMILAR ARTIFACTS. THIS IS A VERY RARE COMBINATION OF INDIVIDUAL EVENTS THAT TOGETHER CAUSED THE ARTEFACT TO RESEMBLE POTENTIALLY PATHOLOGICAL ACTIVITY. NO HARM OR RISK OF HARM WAS REALIZED TO THE PATIENT. | 0
4203992 | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT A RED ALERT WAS DETECTED FOR THIS DEVICE VIA REMOTE MONITORING. NO ADVERSE PATIENT EFFECTS WERE REPORTED.><EFFORTS TO OBTAIN ADDITIONAL PRODUCT ISSUE DETAILS WERE UNSUCCESSFUL. NO OTHER ADVERSE EVENTS HAVE BEEN REPORTED. THIS PRODUCT ISSUE WILL BE UPDATED IF ADDITIONAL INFORMATION IS RECEIVED. | 0
3657273 | THIS ISSUE AFFECTS THE DOSE RANGE CHECKING (DRC) FUNCTIONALITY IN CERNER MILLENNIUM APPLICATIONS. DRC IS AN AUTOMATED METHOD TO COMPARE A MEDICATION ORDER DOSE AGAINST A PRE-ESTABLISHED SAFE RANGE TO ASSIST THE CLINICIAN IN PROVIDING THE CORRECT DOSAGE TO A PT. THE ISSUE OCCURS WHEN A MEDICATION ORDER IS ENTERED IN CERNER MILLENNIUM APPLICATIONS WHERE MULTUM DRUG CONTENT IS USED; A DOSE RANGE CHECKING ALERT MAY NOT BE DISPLAYED. CERNER HAD NOT RECEIVED COMMUNICATION ON ANY ADVERSE PT EVENTS AS A RESULT OF THIS ISSUE.><CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION (B)(4) 2013 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE FLASH NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE, INCLUDES STEPS TO HELP CLIENTS IDENTIFY IF THEY ARE AFFECTED, AND NOTIFIES CLIENTS OF THE CONTENT UPDATE TO RESOLVE THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CONSIDERS THIS ISSUE TO BE RESOLVED AND NO FURTHER NARRATIVE IS REQUIRED FOR FOLLOW-UP. | 0
3637725 | ><THIS ISSUE AFFECTS THE DOSE RANGE CHECKING (DRC) FUNCTIONALITY IN SURGINET, INET, AND FIRSTNET. DRC IS AN AUTOMATED METHOD TO COMPARE MEDICATION ORDER DOSE AGAINST A PRE-ESTABLISHED SAFE RANGE TO ASSIST THE CLINICIAN TO PROVIDED THE CORRECT DOSAGE TO PT. THE ISSUE OCCURS WHEN A MEDICATION ORDER IS ENTERED IN FIRSTNET, INET, OR SURGICAL POWERCHART, A DOSE RANGE CHECKING ALERT MAY NOT BE DISPLAYED. CERNER HAS NOT RECEIVED COMMUNICATION ON ANY ADVERSE PT EVENTS AS A RESULT OF THIS ISSUE. | 0
3724173 | THE MEDICAL ADMINISTRATION RECORD AND RECONCILIATION RECORD RANDOMLY OMITS DOSES OF DISEASE CRITICAL MEDICATIONS, INCREASING THE LIKELIHOOD OF MISTAKES AND OTHER ERRORS IN DOSING AND ADMINISTRATION OF DISEASE CRITICAL MEDICATIONS. IT PLACES PTS AT RISK FOR EITHER UNDER DOSING OR OVERDOSING OF THESE DISEASE. | 0
3621876 | COBE SPECTRA DEVELOPED POWER FAILURE. OPERATOR FOLLOWED INSTRUCTION TO RESUME TREATMENT BUT MACHINE CONTINUED TO FAIL, INDICATING DIFFERENT CODE NUMBER EACH TIME. UNABLE TO RETURN BLOOD MANUALLY AND REMOVE SETS AFTER TOOK OFF PATIENT WITHOUT RETURNING BLOOD. CHANGED OUTLETS AND TURNED MACHINE ON/OFF WITH NO RESULTS. THERAPY DELAYED. | 0
3678819 | TUBING KIT FOR OPTIA TPE (THERAPEUTIC PLASMA EXCHANGE) MALFUNCTIONED DURING PROCEDURE; LOUD BANG HEARD, WHITE UPPER PLASTIC BRACE HOUSING NOTED TO BE SHREDDED. PROCEDURE ABORTED AND NEW TUBING KIT SET UP. PATIENT DID NOT RECEIVE RINSEBACK; APPROXIMATELY 195 CCS OF WB NOT RETURNED TO PATIENT. | 0
3779248 | ON THREE OCCASIONS IN THE PAST THREE MONTHS LEAKS HAVE OCCURRED. ON PLASMAPHERESIS: ON TWO DIFFERENT DATES THIS WAS DETECTED; LEAK DETECTED IN CENTRIFUGE APPROXIMATELY 2 MINUTES INTO THE COLLECTION.ON ONE DATE: DURING SALINE PRIME MULTIPLE RETURN SALINE LINE FAILED INTEGRITY TEST" ALARMS OCCURRED | 0
3959521 | DOING A PLASMA EXCHANGE WITH A COBE SPECTRA. THE MACHINE MALFUNCTIONEDAND GAVE THE PATIENT A BOLUS OF ACDA. IT WAS A FAILURE OF THE ACDA PUMP THAT CAUSED THE PATIENT TO RECEIVE TOO MUCH ACDA.MANUFACTURER RESPONSE FOR DIALYSIS, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================AFTER COMPLETING THE INTERROGATION ON THE SYSTEM, BELOW ARE OUR FINDINGS:THE FAILURE OF THE ENCODER IN THE ANTICOAGULANT (AC) PUMP MOTOR CAUSING THE MOTOR TO GO INTO RUN AWAY MODE GIVING A F20 ERROR CODE WAS THE ONLY REASON FOR THE FAILURE. AS PRECAUTIONS THE AC, PLASMA PUMP MOTORS, PUMP DRIVER BOARD, HALL EFFECTS SWITCH, R.A.M BATTERIES WERE REPLACED AND FULL PM WAS PERFORMED TO INSURE PROPER PERFORMANCE. UNIT RAN FOR ONE HOUR WITH ALL FOUR MOTORS AT 200RPMS NO ERRORS OCCURRED. | 0
3750564 | PERFORMING TPE WITH PRISMAFLEX. BLOOD NOTED OUTSIDE OF BLOOD FILTERS- BLOOD LEAK DETECTOR ALARMED APPROPRIATELY - TREATMENT IMMEDIATELY DISCONTINUED. VISUAL INSPECTION APPEARS TO BE DEFECT IN MEMBRANE SEPARATING TWO COMPARTMENTS. | 0
4049320 | SPACELABS IS EVALUATING THIS REPORTED EVENT AND WILL FILE A SUPPLEMENTAL REPORT WHEN OUR EVALUATION IS CONCLUDED. PLACEHOLDER.><SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR MODEL 91393 DISPLAYED THE RED FLASHING LIGHTS OF A VISUAL ALARM, BUT DID NOT PROVIDE AN AUDIBLE ALARM TONE ON (B)(6) 2014. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><THE FACILITY BIOMED MET WITH THE CLINICAL STAFF TO CONFIRM THE COMPLAINT DESCRIPTION. THE FACILITY BIOMED TESTED THE MONITOR AND FOUND THAT IT PERFORMED TO SPECIFICATIONS, INCLUDING PROVIDING AN OPERATIONAL AUDIBLE ALARM. THE FACILITY BIOMED COULD NOT REPRODUCE THE PROBLEM. THE CUSTOMER HAS NOT REPORTED ANY ADDITIONAL ISSUES WITH THIS MONITOR. SPACELABS IS INVESTIGATING THE ISSUE AND WILL FILE A SUPPLEMENTAL REPORT.><SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR MODEL 91393 DISPLAYED THE RED FLASHING LIGHTS OF A VISUAL ALARM, BUT DID NOT PROVIDE AN AUDIBLE ALARM TONE ON (B)(6) 2014. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR, MODEL 91393, HAD RED FLASHING LIGHTS OF VISUAL ALARM, BUT WAS NOT MAKING AN AUDIBLE ALARM TONE. NO ONE WAS INJURED AS A RESULT OF THIS EVENT.><THE FACILITY BIOMED MET WITH THE CLINICAL STAFF TO CONFIRM THE COMPLAINT DESCRIPTION. THE FACILITY BIOMED TESTED THE MONITOR AND FOUND THAT IT PERFORMED TO SPECIFICATIONS, INCLUDING PROVIDING AN OPERATIONAL AUDIBLE ALARM. THE FACILITY BIOMED COULD NOT REPRODUCE THE PROBLEM. THE CUSTOMER HAS NOT REPORTED ANY ADDITIONAL ISSUES WITH THIS MONITOR TO SPACELABS. A SPACELABS PRODUCT SUPPORT SPECIALIST MADE MULTIPLE ATTEMPTS TO CONTACT THE FACILITY BIOMED VIA PHONE MESSAGES AND EMAIL TO OBTAIN ADDITIONAL INFORMATION. HOWEVER, THE FACILITY BIOMED DID NOT RETURN ANY OF THE MESSAGES. DUE TO INSUFFICIENT INFORMATION FROM THE CUSTOMER A ROOT CAUSE COULD NOT BE DETERMINED. THIS REPORT IS CONSIDERED FINAL AND THE ISSUE CLOSED. PLACEHOLDER. | 0
3801407 | THREE PATIENT SAMPLE TUBES WERE DROPPED BY A VERSACELL SAMPLE MANAGEMENT SYSTEM WHILE BEING TRANSFERRED TO AN IMMULITE INSTRUMENT FOR TESTING. THE SAMPLES SPILLED AND WERE CLEANED BY THE LABORATORY STAFF, WHO WORE PERSONAL PROTECTIVE EQUIPMENT. THE SAMPLES COULD NOT BE PROCESSED BY THE INSTRUMENT AND NEW SAMPLES HAD TO BE OBTAINED FROM THE PATIENTS. THE SPILLED CONTENTS FROM THE DROPPED SAMPLE TUBES DID NOT CONTAMINATE ANY OTHER SAMPLES SCHEDULED FOR TESTING. THERE ARE NO KNOWN REPORTS OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCES DUE TO THE SAMPLE TUBES BEING DROPPED BY THE VERSACELL SAMPLE MANAGEMENT SYSTEM PRIOR TO BEING SAMPLED BY THE INSTRUMENT.><THE CUSTOMER CONTACTED THE SIEMENS CUSTOMER CARE CENTER (CCC). DURING TROUBLESHOOTING WITH THE CUSTOMER, IT WAS DETERMINED THAT THE VERSACELL SAMPLE MANAGEMENT SYSTEM GRIPPERS APPEARED TO BE WORKING AS EXPECTED. THERE WERE NO OBSTRUCTIONS AND ALL O-RINGS WERE IN PLACE. A SIEMENS CUSTOMER SERVICE ENGINEER (CSE) WAS DISPATCHED TO THE CUSTOMER SITE. THE CSE REPLACED THE X-DIRECTION SERVER MOTOR PRINTED CIRCUIT BOARD AND ADJUSTED THE ARM ALIGNMENT. THE CAUSE OF THE DROPPED SAMPLE TUBES IS UNKNOWN. THE VERSACELL SAMPLE MANAGEMENT SYSTEM IS PERFORMING ACCORDING TO SPECIFICATIONS. NO FURTHER EVALUATION OF THIS DEVICE IS REQUIRED. | 0
3869053 | CUSTOMER REPORT: WE GAVE A PATIENT A DRUG THAT THEY WERE ALLERGIC TO TODAY BECAUSE THE PCM DELETED THE DRUG. PLEASE SEE THE AUDIT TRAIL." AFTER INVESTIGATION | 0
4046743 | ><EVALUATION SUMMARY: (B)(4). CLINICAL EVALUATION SUMMARY AND CONCLUSION A HYPOTHETICAL IMPLANT SCENARIO WAS CREATED AND MODELLED ON THE VARISEED TREATMENT PLANNING SYSTEM TO SIMULATE THE IMPACT OF POTENTIAL ERRORS RELATED TO SEED ACTIVITY. FIVE CASES WERE PRODUCED ON A TREATMENT PLANNING SYSTEM WITH FIVE DIFFERENT TREATMENT PLANS. THE PLANS WERE MODIFIED TO IMITATE THE ERROR REPORTED IN (B)(4) AND THE IMPACT ON DOSE DISTRIBUTION WAS THEN ANALYSED. GEC-ESTRO1 RECOMMENDATIONS WERE USED TO EVALUATE THE QUALITY OF PLANNING AND THE POTENTIAL FOR CLINICAL IMPACT. THE LIMITATION OF THIS DOSIMETRIC MODELLING IS THAT THE FIVE CASES ARE SAMPLES AND CANNOT EXACTLY REPRODUCE THE SIZE, SHAPE AND THE DOSIMETRIC TREATMENT PLAN FOR THE SPECIFIC PATIENT WHICH IS UNK. CLINICAL RISK: THE CUSTOMER COULD HAVE MISTAKENLY USED THE INCORRECT SEED ACTIVITY AS INDICATED ON THE LABEL. DOSIMETRY MODELLING: THE DISCREPANCY BETWEEN ACTUAL SEED ACTIVITY AND THE LABEL IS APPROX 8%. TO ASSUME WORST CASE SCENARIO, 5 CLINICALLY ACCEPTABLE TREATMENT PLANS WERE MODIFIED, SO THE SEED ACTIVITY IS INCREASED BY +8% AND THEN DECREASED BY -8% TO INVESTIGATE THE IMPACT OF DOSE DISTRIBUTION. RESULT: INCREASING THE SEED STRENGTH BY 8% COULD LEAD TO OVER DOSAGE OF THE URETHRA BY 11.3%. THIS IS LIKELY TO INCREASE THE RISK OF RADIATION TOXICITY TO THE URETHRA AND CLINICAL COMPLICATIONS INCLUDING URINARY INCONTINENCE. PROSTATE GLAND IS ALSO MORE LIKELY TO BE OVERDOSED BY 18.6% AND 14.4GY WHICH POTENTIALLY COULD BE ASSOCIATED WITH SIDE EFFECTS. ON THE OTHER HAND, REDUCTION OF SEED STRENGTH BY 8% WOULD SIGNIFICANTLY UNDER DOSE THE PROSTATE BY 15% AND 14GY. THIS MAY LEAD TO REDUCTION TO TARGET VOLUME COVERAGE WHICH CAN BE ASSOCIATED WITH TREATMENT FAILURE AND INCREASED RISK OF RELAPSE.4. CONCLUSION: THE EFFECT OF 8% CHANGE IN SEED ACTIVITY COULD LEAD TO SIGNIFICANT CHANGES TO THE DOES DISTRIBUTION WITHIN THE TREATMENT PLAN. THIS MAY BE ASSOCIATED WITH SUB OPTIMAL CLINICAL OUTCOMES. MANUFACTURER EVALUATION: WITH THE FOLLOW-UP INFORMATION RECEIVED, THIS REPORT IS NOW CONSIDERED SERIOUS, BASED ON HYPOTHETICAL IMPLANT SCENARIO. ALTHOUGH IN THIS INSTANCE, THE IMPLANT WAS NOT AFFECTED, DUE TO THE NATURE OF THE ERROR THIS MALFUNCTION WAS ASSESSED AS HAVING THE POTENTIAL TO CAUSE DEATH OR SERIOUS INJURY SHOULD IT RECUR. ADDITIONAL LOT #130237A. | 0
3692357 | ON (B)(6) 2014, CARESTREAM HEALTH REC'D AN INQUIRY REGARDING AVAILABILITY OF PT ALLERGY INFO IN VUE RIS. CARESTREAM WAS INFORMED THAT THE (B)(6) YR OLD ONCOLOGY PT HAD PASSED AWAY.><CARESTREAM HEALTH SERVICE PERSONNEL CONDUCTED AN INVESTIGATION OF THIS ISSUE TO DETERMINE HOW THE VUE RIS WAS WORKING DURING THE EVENT. THE INVESTIGATION INCLUDED COLLECTION OF LOGS SHOWING ALL OF THE MESSAGES REC'D BY THE HOSPITAL'S VUE RIS FROM THEIR SENDING SYSTEM (3RD PARTY VENDOR). THE VUE RIS CONTAINS THE PT INFO SENT BY THE USER THROUGH THEIR 3RD PARY SYSTEM. THE LOG SHOW TWO (2) TYPES OF MESSAGES REC'D BY THE VUE RIS; ORM (ORDERS) AND ADT. THE ORM MESSAGES FROM THE HOSPITAL CONTAIN PT INFO, INCLUDING ALLERGY INFO. THE HOSPITAL DOES NOT ENTER ALLERGY INFO IN ADT MESSAGES. THE ORDER IN WHICH VUE RIS RECEIVES THE MESSAGES IS IMPORTANT; THE LAST MESSAGE REC'D IS CONSIDERED THE MASTER AND LATEST UPDATE ON THE PT. LOG SHOW THE ALLERGY INFO TRANSMITTED FROM THE HOSPITAL SYSTEM TO THE VUE RIS IN THE ORM MESSAGE FOR THIS PT ON (B)(6) 2014. THE LOGS SHOW A USER AT THE HOSPITAL SENT AN ADT MESSAGE DISCHARGING THE PT; THIS ADT MESSAGE DID NOT INCLUDE ALLERGY INFO BECAUSE THAT INFO IS NOT INCLUDED IN ADT MESSAGES. ALLERGY INFO WAS CURRENT IN VUE RIS FOR THE PT AS OF THE LAST STUDY PERFORMED ((B)(6) 2014:16:30). THE VUE RIS REC'D A PT DISCHARGE NOTIFICATION (ADT_A03) FROM THE HOSPITAL'S ORDERING SYSTEM AT ((B)(6) 2014:14:17) THAT DID NOT CONTAIN THE ALLERGY INFO. AFTER (B)(6) 2014, THE VUE RIS REC'D A SERIES OF PT UPDATE (ADT_A08) MESSAGES; AGAIN THESE MESSAGES DID NOT CONTAIN ANY ALLERGY INFO. | 0
4116928 | ><GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE PROVIDED AFTER THE INVESTIGATION HAS BEEN COMPLETED. | 0
4058344 | ><(B)(4). CONCLUSION: AN ASSIGNABLE CAUSE COULD NOT BE DETERMINED AS THE PRODUCTS MET ALL APPROVED RELEASE SPECIFICATIONS AT THE TIME OF MANUFACTURE BEFORE THEY WERE RELEASED. BASED ON REVIEW OF ALL THE INFORMATION GATHERED DURING THIS INVESTIGATION-ASSOCIATED SAGE DOCUMENTATION, MANUFACTURING PROCESS, COMPLAINT HISTORY, AND INSPECTION OF CURRENT INVENTORY/RETENTION SAMPLE , SEVERAL EXTERNAL SOURCES MAY HAVE CONTRIBUTED TO THE REPORTED COMPLAINT CONDITION. POTENTIAL SOURCES OF PRODUCT CONTAMINATION INCLUDE PRE-STATING ACTIVITIES PRIOR TO PROCEDURE, THE CATHETER USED , OR IMPROPER WASHING OF THE SPERM SAMPLES (SEE FISH BONE DIAGRAM ABOVE). THE RETENTION SAMPLES OF THE SUBJECT LOTS TESTED SHOW THAT THE PRODUCTS MEETS ALL APPROVED RELEASE SPECIFICATIONS. THE CUSTOMER DID NOT RETURN THE ACTUAL SAMPLES INVOLVED. THE PRODUCT PERFORMS AS INTENDED. NO FURTHER ACTIONS ARE NEEDED AT THIS TIME. THE INVESTIGATION MAY BE REEVALUATED WITH ANY NEW ADDITIONAL INFORMATION RECEIVED. | 0
3918112 | COMP REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S POWER SUPPLY AND SPO2 INTERFACE BOARD. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><THE CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR PLUS, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 0
4322942 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S POWER SUPPLY AND RESEATING THE CABLES. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED THE PASSPORT 2 MONITOR SHUT DOWN, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 0
4217053 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPO2 BOARD. UNIT WAS SAFETY TESTED TO FACTORY'S SPECS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 0
4037863 | COMPANY REP EVALUATED SYS. CORRECTIONS INCLUDED REPLACEMENT OF THE SPO2 CABLE ASSEMBLY. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR PLUS MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 0
4190162 | ISSUE CURRENTLY UNDER INVESTIGATION.><CUSTOMER REPORTED THAT THE PASSPORT V MONITOR FELL OFF FROM THE ROLLING STAND AND THE NURSE CAUGHT IT. NO PT INJURY REPORTED. | 0
3418385 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE POWER MODULE ASSEMBLY. | 0
3918035 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF UNIT'S ECG MODULE.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 7 MONITOR, WHICH MAY HAVE AFFECTED ECG MONITORING. NO PT INJURY WAS REPORTED. | 0
4067428 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2007. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
3903071 | IT WAS REPORTED THAT A HOSPITAL HANDHELD WAS EXPEIRENCING SCREEN FREEZES AND A NEW PROGRAMMING COMPUTER WAS REQUESTED. IT IS UNKNOWN IF ANY TROUBLESHOOTING WAS PERFORMED. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><THE SUSPECTED PROGRAMMING COMPUTER HAS BEEN RETURNED ON 01/11/2016 FOR THE ANALYSIS. THE FLASHCARD SOFTWARE HAS BEEN NOT RETURNED TO DATE. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE REPORTED FROZEN SCREEN" WAS NOT VERIFIED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT A KNOWN GOOD FLASHCARD COULD NOT BE FULLY INSERTED INTO THE RETURNED HANDHELD. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A BENT FLASHCARD SLOT PIN. ONCE THE PIN WAS STRAIGHTENED | 0
4059021 | ON (B)(6) 2014 DURING A REVIEW OF PROGRAMMING HISTORY, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2011 WHICH RESULTED IN A SETTINGS CHANGE. WHILE ATTEMPTING TO CORRECT THE SETTINGS, A PARTIAL PROGRAMMING EVENT OCCURRED WHICH WAS CORRECTED BUT THE SETTINGS WERE STILL NOT FULLY CORRECTED WHEN THEY WERE CHANGED AS A RESULT OF THE FAULTED SYSTEM DIAGNOSTIC TEST. THE MAGNET OUTPUT WAS NOT PROGRAMMED BACK OFF AS INTENDED.><ANALYSIS OF PROGRAMMING HISTORY. | 0
4067439 | DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2011. THE PHYSICIAN ONLY CORRECTED THE NORMAL MODE AND MAGNET MODE OUTPUT CURRENTS. NO PATIENT ADVERSE EVENTS WERE REPORTED.><REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY. | 0
3610519 | REVIEW OF THE PATIENT'S PROGRAMMING HISTORY WAS PERFORMED IN FILE (B)(4) ON (B)(4) 2014. IT WAS NOTICED THAT UPON INITIAL INTERROGATION ON (B)(6) 2008, THE OUTPUT CURRENTS WERE AT 0MA. ON THE PREVIOUS VISIT ON (B)(6) 2007, THERE WAS NO FINAL INTERROGATION PERFORMED PRIOR TO THE PATIENT LEAVING THE CLINIC. TWO PROGRAMMING EVENTS WERE PERFORMED BUT A FINAL INTERROGATION DID NOT SHOW IF THE PROGRAMMING WAS SUCCESSFUL. THEREFORE, IT CANNOT BE RULED OUT THAT AN INCOMPLETE SYSTEM DIAGNOSTIC TEST INADVERTENTLY CHANGED THE SETTINGS. THIS FILE WAS CREATED TO HOUSE THE PROGRAMMING ANOMALY.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3629794 | PRODUCT ANALYSIS OF THE GENERATOR FOUND THAT THE DEVICE WAS AT FAULTED SETTINGS. ATTEMPTS ARE UNDERWAY FOR ADDITIONAL INFORMATION; HOWEVER, NO OTHER INFORMATION HAS BEEN PROVIDED.><.><ATTEMPTS TO OBTAIN ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. | 0
3408389 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT DURING INTERROGATION ON OFFICE VISIT ON (B)(6) 2011 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE SAME OFFICE VISIT. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2011. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE MAGNET ON TIME WAS NOT CORRECTED. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
4067762 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2010 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2010. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2010. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2010 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE CORRECTED ON (B)(6) 2010. NO PATIENT ADVERSE EVENTS WERE REPORTED.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3563154 | ANALYSIS OF PROGRAMMING HISTORY.><A REVIEW OF THE VNS PROGRAMMING HISTORY FOUND THAT A FAULTED SYSTEM DIAGNOSTIC TEST WAS PERFORMED ON (B)(6) 2012, WHICH CHANGED THE VNS SETTINGS. ALTHOUGH SOME OF THE SETTINGS WERE CHANGED, THE OFF TIME WAS NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE VISIT. | 0
3558506 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT DURING INTERROGATION ON OFFICE VISIT ON (B)(6) 2012 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE SAME OFFICE VISIT. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST WHICH OCCURRED ON (B)(6) 2012. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE DEVICE OFF TIME AND MAGNET ON TIME WERE NOT CORRECTED. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2012 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
3813070 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT AN INTERRUPTED GENERATOR DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2014. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE MAGNET OUTPUT CURRENT AND MAGNET ON-TIME WERE NOT CORRECTED. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
4147761 | .><IT WAS REPORTED THAT THE PHYSICIAN WAS HAVING DIFFICULTY INTERROGATING THE VNS PATIENT¿S DEVICE. THE PHYSICIAN WAS FINALLY ABLE TO INTERROGATE THE DEVICE BUT NOTED THAT THE DEVICE WAS DISABLED. THE PATIENT¿S DEVICE OUTPUT CURRENT WAS PREVIOUSLY PROGRAMMED TO 1.75MA. THE PHYSICIAN PROGRAMMED THE PATIENT¿S DEVICE BACK ON TO AN OUTPUT CURRENT OF 0.5MA. THE PATIENT STATED THAT HE WAS UNABLE TO PERCEIVE STIMULATION FOLLOWING HIS PREVIOUS OFFICE VISIT. CLINIC NOTES WERE RECEIVED INDICATING THAT THE PATIENT WAS LAST SEEN ON (B)(6) 2014. THE NOTES INDICATE THAT THE PHYSICIAN ALSO HAD DIFFICULTY INTERROGATING THE PATIENT¿S DEVICE DURING THE OFFICE VISIT. THE PHYSICIAN STATED THAT THE PATIENT¿S DEVICE SETTINGS WERE CHANGED UNINTENTIONALLY DURING THIS VISIT. | 0
3392320 | REVIEW OF DEVICE PROGRAMMING HISTORY.><TROUBLESHOOTING WAS REQUESTED BY A NURSE PRACTITIONER WHO WAS WITH A PATIENT IMPLANTED WITH VNS. THE NP WAS ABLE TO INTERROGATE THE PATIENT¿S DEVICE AND PERFORM SYSTEM DIAGNOSTIC TESTING WITHOUT ISSUE. THE SYSTEM DIAGNOSTICS WERE ALL WITHIN NORMAL LIMITS AND THERE WERE NO DELAYS OR COMMUNICATION ERRORS WITH THE DEVICE. HOWEVER, WHEN THE NP ATTEMPTED TO RUN MAGNET MODE DIAGNOSTICS, AN ERROR MESSAGE WAS REPEATEDLY OBSERVED STATING THE PULSE GENERATOR DID NOT SENSE A MAGNET SWIPE DURING THE EXECUTION OF THE MAGNET MODE DIAGNOSTICS. PLEASE EXECUTE THE DIAGNOSTIC AGAIN AND VERIFY THAT THE PULSE GENERATOR IS SWIPED WITH THE MAGNET BEFORE STARTING THE DIAGNOSTICS. IF YOU HAVE ANY QUESTIONS | 0
4006510 | .><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2011 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2010. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2010 DID NOT IDENTIFY ANY DIAGNOSTIC TESTS. THE DEVICE WAS INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2010 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE FAULTED DIAGNOSTIC TEST LIKELY OCCURRED BETWEEN (B)(6) 2010 AND (B)(6) 2011 ON A DIFFERENT PROGRAMMING SYSTEM. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
3546742 | DURING REVIEW OF PROGRAMMING HISTORY, IT WAS NOTED THAT AN INTERRUPTED SYSTEM DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS. THE SETTINGS WERE NOT CORRECTED AT THE TIME OF THE EVENT. NO ADVERSE EVENTS HAVE BEEN REPORTED AS A RESULT OF THIS.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3849521 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><CLINIC NOTES WERE RECEIVED INDICATING THAT THE VNS PATIENT WENT TO THE ER ON (B)(6) 2011 DUE TO AN INCREASE IN SEIZURES. THE PATIENT¿S DEVICE SETTINGS WERE ADJUSTED AND THE PATIENT FELT BACK TO NORMAL. THE PATIENT¿S DEVICE WAS LATER INTERROGATED AND HER DEVICE SETTINGS WERE REPORTEDLY RESET. REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY SHOWED NO ANOMALIES IN PROGRAMMING THROUGH THE FINAL INTERROGATION ON (B)(6) 2011. ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. THE INCREASE IN SEIZURES WAS REPORTED IN MANUFACTURER REPORT #1644487-2012-00170. | 0
3497219 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD FREEZES INTERMITTENTLY DURING INTERROGATION. IT WAS REPORTED THAT THIS HAS BEEN AN ON GOING ISSUE FOR A WHILE. NO PATIENTS WERE AFFECTED BY THE FREEZE BECAUSE THE PHYSICIAN WAS ABLE TO USE ANOTHER PROGRAMMING SYSTEM. TROUBLESHOOTING WAS PERFORMED AND THE FREEZE WAS ABLE TO BE DUPLICATED. A HARD RESET WAS PERFORMED. A NEW PROGRAMMING TABLET WAS REQUESTED TO BE SENT TO THE PHYSICIAN. THE HANDHELD IS EXPECTED TO BE RETURNED, BUT HAS NOT BEEN RECEIVED TO DATE.><(B)(4).><THE HANDHELD AND FLASHCARD WERE RECEIVED BY THE MANUFACTURER. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED FROZEN SCREEN ALLEGATION WAS NOT VERIFIED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE SOFTWARE WOULD PAUSE FOLLOWING AN INTERROGATION. THIS IS EXPECTED BEHAVIOR FOR THE SOFTWARE CONSIDERING THE DATABASE SIZE. DURING THE ANALYSIS, NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
3403563 | .><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WOULD NOT GO PAST THE INITIAL SCREEN AFTER A HARD RESET. IT WAS REPORTED THAT THE SCREEN INDICATES TO TAP SCREEN TO CONTINUE, BUT THE HANDHELD WAS FROZEN. TROUBLESHOOTING WAS PERFORMED BY CHECKING TO ENSURE THE LOCK BUTTON WAS NOT ENGAGED AND BY PERFORMING A HARD RESET. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET AND THE HANDHELD WAS RETURNED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
3406687 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT ON OFFICE VISIT ON (B)(6) 2010 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE PREVIOUS OFFICE VISIT ON (B)(6) 2010. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2010 WAS WITHIN NORMAL LIMITS. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2010 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; THEREFORE, THE SETTINGS WERE NOT CORRECTED PRIOR TO THE PATIENT LEAVING THE OFFICE. THE SETTINGS WERE NOT CORRECTED ON 10/10/2010. NO PATIENT ADVERSE EVENTS WERE REPORTED.><PREVIOUSLY SUBMITTED MDR INCORRECTLY IDENTIFIED THE SOFTWARE VERSION AND OMITTED THE LOT NUMBER OF THE SOFTWARE PRODUCT. THIS REPORT IS BEING SUBMITTED TO CORRECT THIS INFORMATION. PREVIOUSLY SUBMITTED MDR OMITTED THE MANUFACTURE DATE OF THE SUSPECT MEDICAL DEVICE. THIS REPORT IS BEING SUBMITTED TO CORRECT THIS INFORMATION. | 0
3854325 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD CONSTANTLY FREEZES, EVEN AFTER PERFORMING A HARD RESET. TROUBLESHOOTING WAS PERFORMED WHICH DID NOT RESOLVE THE ISSUE. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING COMPUTER. THE HANDHELD AND FLASHCARD HAVE NOT BEEN RECEIVED FOR ANALYSIS TO DATE. ATTEMPTS TO OBTAIN ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><.><THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON 11/20/2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE CRADLE AND NO ANOMALIES WERE IDENTIFIED IN ITS ABILITY TO PROVIDE POWER TO THE HANDHELD. THE CRADLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON 11/20/2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
3793654 | PRODUCT INFORMATION WAS RECEIVED. THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD FREEZES AFTER INTERROGATION. IT WAS REPORTED THAT THE HANDHELD WILL REMAIL FROZEN ON THE SAME SCREEN FOR AT LEAST TWENTY MINUTES AS THE INFORMATION IS BEING PERFORMED. TROUBLESHOOTING WAS PERFORMED BY A HARD RESET AND REINSERTION AND CLEANING OF THE FLASHCARD. A NEW PROGRAMMING COMPUTER WAS PROVIDED TO THE PHYSICIAN. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><. | 0
4064602 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT A GENERATOR DIAGNOSTIC TEST OCCURRED THAT CAUSED AN UNINTENDED CHANGE IN DEVICE SETTINGS DURING AN OFFICE VISIT ON (B)(6) 2012. THE PATIENT¿S DEVICE WAS NOT SUBSEQUENTLY INTERROGATED DURING THE VISIT. ON (B)(6) 2013, THE PATIENT¿S DEVICE SETTINGS WERE INDICATIVE OF A GENERATOR DIAGNOSTIC TEST. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
3589558 | .><PRODUCT INFORMATION WAS OBTAINED.><LOT #, OTHER, CORRECTED DATA: PREVIOUSLY SUBMITTED MDR INDICATED THIS INFORMATION WAS UNKNOWN. IT IS NOW KNOWN. THIS REPORT IS BEING SUBMITTED TO CORRECT THIS DATA. MANUFACTURE DATE, CORRECTED DATA: PREVIOUSLY SUBMITTED MDR INDICATED THIS INFORMATION WAS UNKNOWN. IT IS NOW KNOWN. THIS REPORT IS BEING SUBMITTED TO CORRECT THIS DATA.><A COMPANY REPRESENTATIVE REPORTED THAT HIS PROGRAMMING SYSTEM WAS NOT FUNCTIONING. IT WAS REPORTED THAT THE HANDHELD IS LOCKING UP AND THE SCREEN IS NOT RESPONDING. THE COMPANY REPRESENTATIVE REPORTED THAT A HARD RESET WAS PERFORMED AND THAT ALL TROUBLESHOOTING SOLUTIONS HAVE BEEN ATTEMPTED AND THAT NOTHING IS WORKING. THE REPRESENTATIVE REQUESTED A NEW PROGRAMMING HANDHELD BE SENT TO HIM. THE HANDHELD WAS RETURNED FOR ANALYSIS ON (B)(6) 2014. ANALYSIS WAS COMPLETED ON (B)(6) 2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE FLASHCARD WAS NOT RETURNED FOR ANALYSIS. | 0
3709338 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY. DATE OF EVENT; CORRECTED DATA: ADDITIONAL INFORMATION INDICATES THAT THE CROSS PROGRAMMING EVENT OCCURRED ON (B)(6) 2014.><IT WAS REPORTED THAT A CROSS PROGRAMMING EVENT IS BELIEVED TO HAVE OCCURRED WITH THE PATIENT. THE NURSE PRACTITIONER REPORTED THAT THE PATIENT WAS LAST SEEN IN (B)(6) 2013 AND THAT THE SETTINGS PROGRAMMED AT THIS VISIT ARE NOT THE SAME AS THE PATIENT WAS FOUND AT CURRENTLY. IT WAS REPORTED THAT THE DEVICE SERIAL NUMBER AND OTHER PARAMETERS WERE CORRECT, BUT THE OUTPUT CURRENT WAS DIFFERENT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE VNS PATIENT¿S DEVICE SETTINGS WERE FOUND TO BE LOWER THAN WHAT THEY WERE PREVIOUSLY. THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EVENTS. REVIEW OF ADDITIONAL PROGRAMMING HISTORY SHOWS THAT THE LAST KNOWN SETTINGS PRIOR TO THE EVENT WERE FROM (B)(6) 2012 IN WHICH THE OUTPUT CURRENT WAS PROGRAMMED TO 1.75MA. THE CROSS PROGRAMMING EVENT APPEARS TO HAVE OCCURRED ON (B)(6) 2014.><. | 0
3552423 | ANALYSIS OF PROGRAMMING HISTORY.><A REVIEW OF PROGRAMMING AND DIAGNOSTIC DATA WAS PERFORMED, ON (B)(6) 2013 FAILED DIAGNOSTICS OCCURRED AND CAUSED THE SETTINGS TO CHANGE. ALL SETTINGS, EXCEPT FOR THE MAGNET ON TIME WERE PROGRAMMED BACK TO THE ORIGINAL SETTINGS. IT IS UNCLEAR IF THE MAGNET ON TIME WAS INTENTIONALLY LEFT AT 30 SECONDS OR NOT. | 0
3746464 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT DURING INTERROGATION ON OFFICE VISIT ON (B)(6) 2007 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE SAME OFFICE VISIT. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2007 WAS WITHIN NORMAL LIMITS. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE DEVICE FREQUENCY AND MAGNET ON TIME WERE NOT CORRECTED. THE DEVICE WAS INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2007 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; HOWEVER, THE PHYSICIAN DID NOT CORRECT THE SETTINGS. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
4193105 | IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS SLOW TO RESPOND. A HARD RESET WAS PERFORMED WHICH RESOLVED THE PROBLEM. IT WAS LATER REPORTED THAT THE EVENT WAS HAPPENING MORE FREQUENTLY WITH THE SCREEN BECOMING NON-RESPONSIVE ON START-UP AND IN BETWEEN PATIENTS. IT WAS REPORTED THAT THIS HAS OCCURRED TWICE WHEN THE HANDHELD WAS POWERED ON. A HARD RESET WAS ABLE TO RESOLVE THE ISSUE, BUT THE SITE WANTED A NEW PROGRAMMING COMPUTER. A NEW PROGRAMMING COMPUTER WAS PROVIDED TO THE NURSE PRACTITIONER. IT WAS REPORTED THAT THE NEW PROGRAMMING COMPUTER HAS NOT HAD ANY PROBLEMS. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><.><THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
4059957 | ANALYSIS OF PROGRAMMING HISTORY.><DURING REVIEW OF PROGRAMMING HISTORY IT WAS NOTED THAT AN INCOMPLETE DIAGNOSTICS OCCURRED WHICH CHANGE THE PATIENT TO UNINTENTIONAL SETTINGS. ALTHOUGH THERE WAS A FINAL INTERROGATION TO CHECK THE SETTING THE SETTING WERE NO CORRECTED PRIOR TO THE PATIENT LEAVING. SETTINGS WERE CORRECTED AT A LATER APPOINTMENT. | 0
4051940 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2012. UPON INITIAL INTERROGATION ON THE NEXT VISIT ON (B)(6) 2012, THE SETTINGS WERE AT UNINTENDED PARAMETERS. THE DEVICE WAS NOT INTERROGATED PRIOR TO THE PATIENT LEAVING THE HOSPITAL ON (B)(6) 2012 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS. NO PATIENT ADVERSE EVENTS WERE REPORTED.><REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY. | 0
3430487 | DURING REVIEW OF THE PATIENT'S PROGRAMMING HISTORY, IT WAS OBSERVED THAT THE DEVICE WAS TURNED ON THE DATE OF IMPLANT, BUT A SYSTEM DIAGNOSTIC TEST WAS PERFORMED AFTERWARDS AND A FINAL INTERROGATION WAS NOT PERFORMED. THE DEVICE WAS THEREFORE TURNED TO 0MA INADVERTENTLY DUE TO A PROGRAMMING ANOMALY AS EVIDENT UPON INITIAL INTERROGATION ON (B)(6) 2010. NO PATIENT ADVERSE EVENTS WERE REPORTED.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3902313 | REVIEW OF THE AVAILABLE PROGRAMMING AND DIAGNOSTIC HISTORY.><DURING REVIEW OF PROGRAMMING AND DIAGNOSTIC HISTORY, IT WAS OBSERVED THAT THE VNS PATIENT¿S DEVICE WAS PROGRAMMED TO SETTINGS INDICATIVE OF AN INTERRUPTED SYSTEM DIAGNOSTIC TEST ON (B)(6) 2009. NO PRIOR PROGRAMMING AND DIAGNOSTIC HISTORY IS AVAILABLE SO THE DATE OF THE INTERRUPTED SYSTEM DIAGNOSTIC TEST IS UNKNOWN. THE PATIENT WAS IMPLANTED ON (B)(6) 2009. NO PATIENT ADVERSE EVENTS WERE REPORTED. | 0
3571935 | DURING REVIEW OF THE IN-HOUSE PROGRAMMING/DIAGNOSTIC HISTORY DATABASE, IT WAS OBSERVED THAT DURING INTERROGATION ON OFFICE VISIT ON (B)(6) 2011 THE PATIENT'S SETTINGS WERE DIFFERENT THAN WHAT WERE PROGRAMMED AT THE SAME OFFICE VISIT. THE SETTINGS FOUND WERE INDICATIVE OF A FAULTED DIAGNOSTIC TEST; HOWEVER, REVIEW OF SYSTEM DIAGNOSTIC TESTING FROM OFFICE VISIT ON (B)(6) 2011 WAS WITHIN NORMAL LIMITS. THE PHYSICIAN CORRECTED THE SETTINGS; HOWEVER, THE DEVICE OFF TIME WAS NOT CORRECTED. THE DEVICE WAS INTERROGATED PRIOR TO THE PATIENT LEAVING THE OFFICE ON (B)(6) 2011 AS RECOMMENDED BY DEVICE MANUFACTURER TO ENSURE THE DEVICE IS AT THE CORRECT SETTINGS; HOWEVER, THE PHYSICIAN DID NOT CORRECT THE OFF TIME BACK TO EFFICACIOUS SETTING. THE SETTING WAS CORRECTED ON (B)(6) 2012. NO PATIENT ADVERSE EVENTS WERE REPORTED.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3908447 | IT WAS REPORTED THAT THE NEUROLOGIST¿S HANDHELD DEVICE HAD A FROZEN SCREEN. THE HANDHELD DEVICE WAS CHARGED AND A HARD RESET WAS PERFORMED, BUT THE ISSUE DID NOT RESOLVE. THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 0
4157513 | THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE PHYSICIAN¿S HANDHELD DEVICE WOULD REPEATEDLY FREEZE WHEN ATTEMPTING TO PROGRAM. THE PHYSICIAN REPORTEDLY EXPERIENCED THESE ISSUES FOR MANY YEARS WHICH ONLY RESOLVED AFTER THE SOFTWARE FLASHCARD WAS RESEATED. THE HANDHELD DEVICE AND SOFTWARE FLASHCARD HAVE NOT BEEN RETURNED TO DATE.><. | 0
4061588 | ANALYSIS OF PROGRAMMING HISTORY.><(SERIAL #); CORRECTED DATA: THIS INFORMATION WAS INADVERTENTLY LEFT OFF OF INITIAL MFR. REPORT.><ON (B)(6) 2014 DURING A REVIEW OF PROGRAMMING HISTORY, IT WAS OBSERVED THAT A FAULTED SYSTEM DIAGNOSTIC TEST OCCURRED ON (B)(6) 2005 WHICH RESULTED IN A SETTINGS CHANGE. THE SETTINGS CHANGE WAS NOT NOTED DURING THAT APPOINTMENT AS NO FINAL INTERROGATION WAS PERFORMED. THE SETTINGS CHANGE WAS NOT FOUND UNTIL THE NEXT APPOINTMENT AND NOT CORRECTED UNTIL (B)(6) 2006.><REVIEW OF DEVICE PROGRAMMING HISTORY IDENTIFIED THAT ON (B)(6) 2006, ONLY THE MAGNET OUTPUT CURRENT WAS ADJUSTED. ON (B)(6) 2006, ALL SETTINGS WERE ADJUSTED. | 0
3556076 | A REVIEW OF PROGRAMMING AND DIAGNOSTIC DATA WAS PERFORMED, WHICH IDENTIFIED THAT ON (B)(6) 2012, THE PATIENT WAS FOUND TO BE AT FAULTED SYSTEM DIAGNOSTICS SETTINGS UPON INITIAL INTERROGATION. THE MAGNET OUTPUT CURRENT WAS CORRECTED ON THAT DATE AND THE NORMAL OUTPUT CURRENT WAS CORRECTED ON (B)(6) 2012; HOWEVER, THE REST OF THE SETTINGS WERE NOT CHANGED BACK TO PREVIOUS SETTINGS.><ANALYSIS OF PROGRAMMING HISTORY. | 0
3430144 | .><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD SCREEN IS DARK WITH THE BATTERY CHARGED. IT WAS REPORTED THAT TROUBLESHOOTING WAS PERFORMED WHICH DID NOT RESOLVE THE ISSUE. A HARD RESET WAS PERFORMED AND THE ON/OFF BUTTON WAS PRESSED FOR A FEW SECONDS. A NEW PROGRAMMING SYSTEM WAS PROVIDED TO THE PHYSICIAN. THE HANDHELD WAS RETURNED TO MANUFACTURER FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
3846113 | TYPE OF DEVICE, NAME; CORRECTED DATA: THE INITIAL MDR INCORRECTLY REPORTED THE SUSPECT DEVICE AS THE PROGRAMMING SOFTWARE. THE SUSPECT DEVICE IS THE PROGRAMMING COMPUTER. DEVICE MANUFACTURE DATE; CORRECTED DATA: THE INITIAL MDR INCORRECTLY REPORTED THE SUSPECT DEVICE AS THE PROGRAMMING SOFTWARE. THE SUSPECT DEVICE IS THE PROGRAMMING COMPUTER.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD IS BLOCKED AND ASKING FOR SCREEN ALIGNMENT. A HARD RESET WAS PERFORMED; HOWEVER, IT WAS REPORTED THAT THE HANDHELD WAS COMPLETELY BLOCKED AT THE SCREEN ALIGNMENT". THE PHYSICIAN REQUESTED A REPLACEMENT PROGRAMMING SYSTEM. A NEW PROGRAMMING DEVICE WAS PROVIDED TO THE PHYSICIAN.><THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><." | 0
4274103 | IT WAS REPORTED THAT THE TABLET DEVICE WAS SHATTERED DUE TO UNKNOWN REASONS. THE TABLET DEVICE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><THE USER OF THE TABLET REPORTED THAT IT WAS NOT BELIEVED THAT THE DEVICE HAD BEEN DROPPED, AND THE REASON FOR THE CRACKED SCREEN WAS UNCLEAR. AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET, AND THE REPORTED ALLEGATION WAS VERIFIED. A VISUAL INSPECTION OF THE TABLET VERIFIED THAT THE DISPLAY WAS CRACKED. ALSO DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE TABLET WOULD NOT POWER ON. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH THE POWER BUTTON BEING INSTALLED UPSIDE-DOWN. AS A RESULT, WHEN THE POWER BUTTON WAS PRESSED, THE POST ON THE POWER BUTTON WAS STRIKING THE POWER BOARD LED AND NOT THE ON/OFF SWITCH, PREVENTING THE TABLET FROM POWERING ON. BASED ON THIS INFORMATION, IT APPEARS THAT THE BUTTON WAS LIKELY ACCIDENTLY REMOVED WHILE IN THE FIELD, WHICH MAY HAVE BEEN A RESULT OF POTENTIALLY DROPPING THE TABLET OR PULLING UP ON THE POWER BUTTON. ONCE THE POWER BUTTON WAS REMOVED AND REINSTALLED, NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET WERE IDENTIFIED DURING THE ANALYSIS.><THE INITIAL REPORT INADVERTENTLY DID NOT REPORT THE FOLLOWING: IT WAS REPORTED THAT THE USER DID NOT BELIEVE THAT THE DEVICE HAD BEEN DROPPED, AND THE REASON FOR THE CRACKED SCREEN WAS UNCLEAR. DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
4334607 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN¿S TABLET DEVICE SHOWED AN ¿UNABLE TO OPEN PORT¿ MESSAGE. A NEW SERIAL CABLE WAS PROVIDED BUT DID NOT RESOLVE THE ISSUE. IT WAS NOTED THAT THE USB PORT ON THE TABLET WAS LOOSE AND THE SOURCE OF THE PROBLEM. THE PHYSICIAN HAD PREVIOUSLY STEPPED ON THE SERIAL CABLE WHICH MAY HAVE CAUSED DAMAGE TO THE USB PORT. THE TABLET DEVICE WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS. A BROKEN WIRE CONNECTION IN THE SERIAL CABLE DB9 HOOD ASSEMBLY WAS FOUND. ONCE THE WIRE WAS SOLDERED ON TO THE PCB, NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. NO ANOMALIES ASSOCIATED WITH THE USB PORT WERE IDENTIFIED DURING THE ANALYSIS. | 0
3918611 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD HAD A MESSAGE TO CLEAN ALL DATA IN THE MEMORY | 0
3632939 | IT WAS REPORTED THAT A PROGRAMMING SYSTEM WAS NOT WORKING DUE TO COMMUNICATION ISSUES. THE PROGRAMMING SYSTEM WOULD WORK INTERMITTENTLY AND THE COMMUNICATION DIFFICULTIES WERE CONFIRMED ON A DEMO DEVICE. THE 9V BATTERY WAS NOT THE CAUSE OF THE COMMUNICATION ISSUES. THE PROGRAMMING SYSTEM WAS RETURNED TO THE MANUFACTURER WHERE THE FAILURE TO PROGRAM WAS CONFIRMED AND ATTRIBUTED TO AN INTERMITTENT CONDUCTOR IN THE SERIAL DATA CABLE. ONCE THE CABLE WAS SUBSTITUTED, ALL COMMUNICATION ERRORS CLEARED. NO ISSUES WITH THE RETURNED FLASHCARD WERE NOTED. DURING ANALYSIS OF THE RETURNED HANDHELD, IT WAS IDENTIFIED THAT SOME OF THE SYNC CABLE CONNECTOR LEADS WERE DAMAGED AND THAT THE SOLDER CONNECTIONS BETWEEN THE SYNC CABLE CONNECTOR LEADS AND THE PCB WERE CRACKED. BECAUSE OF THE DAMAGE, THE HANDHELD WAS UNABLE TO RECEIVE POWER FROM THE AC ADAPTER. THE CAUSE FOR THE DAMAGE TO THE CONNECTOR AND SOLDER CONNECTIONS IS MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE AS IT APPEARS THE CONNECTOR DETENTS ARE NOT BEING COMPRESSED WHEN REMOVING OR MANIPULATING THE SERIAL DATA CABLE, THUS PLACING AN EXTENSIVE AMOUNT OF FORCE ON THE PCB AND ATTACHED CABLE RECEPTACLE. IT WAS ALSO IDENTIFIED THAT THE SERIAL CABLE WAS UNABLE TO ESTABLISH COMMUNICATION USING A KNOWN GOOD WAND AND THE RETURNED HANDHELD. THE CAUSE FOR THE COMMUNICATION DIFFICULTIES IS ASSOCIATED WITH A BROKEN WIRE CONNECTION ON THE AXIM CONNECTOR PLUG PCB. ONCE THE WIRE WAS SOLDERED BACK ONTO THE DELL AXIM CONNECTOR PLUG PCB, THE SERIAL CABLE WAS ABLE TO ESTABLISH COMMUNICATION BETWEEN THE HANDHELD AND WAND. THE MOST LIKELY CAUSE FOR THE WIRE BREAKS CAN BE ASSOCIATED WITH MISHANDLING OF THE SERIAL CABLE. THE COMMUNICATION ERRORS ASSOCIATED WITH THE INTERMITTENT CONDUCTOR IN THE WAND WILL BE REPORTED VIA ALTERNATE SUMMARY REPORTING.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3473282 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><WHEN THE MANUFACTURER'S REPRESENTATIVE WENT TO VISIT THE PHYSICIAN, IT WAS REPORTED THAT THE PROGRAMMING SYSTEM (WAND AND HANDHELD) WERE NOT WORKING. AFTER CHECKING THE EQUIPMENT, IT APPEARED THAT THE COMMUNICATION ISSUES WERE REPETITIVE. SOME TIMES THE PROGRAMMING SYSTEM COULD BE USED WITHOUT ANY PROBLEMS, AND OTHER TIMES IT COULD NOT UNDER THE SAME CONDITIONS. THE SYSTEM WAS TESTED SEVERAL TIMES IN THE SAME ROOM WITH KNOWN WORKING GENERATORS (THE GENERATORS WERE INTERROGATED BY AN ALTERNATE PROGRAMMING SYSTEM AND THEREFORE WERE KNOWN TO BE WORKING). ATTEMPTS TO ISOLATE THE MALFUNCTIONING UNIT, BETWEEN THE WAND AND HANDHELD, WERE UNSUCCESSFUL. AS THE SITE HAS AN ALTERNATE PROGRAMMING SYSTEM, PRODUCT ANALYSIS OF THE WAND (REPORTED IN MFR #: 1644487-2012-02849) INDICATED THAT THERE WERE NO ANOMALIES AND IT WAS WORKING CORRECTLY. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE REPORTED ALLEGATION WAS VERIFIED. DURING THE ANALYSIS, 2 ANOMALIES WERE IDENTIFIED THAT HAD A NEGATIVE IMPACT ON THE DEVICE PERFORMANCE. THE FIRST ANOMALY IS ASSOCIATED WITH THE MISSING BATTERY COVER. WITHOUT THE BATTERY COVER, THE HANDHELD WILL NOT POWER ON. THE SECOND ANOMALY IS ASSOCIATED WITH A DAMAGED 120V POWER CORD. AN ANALYSIS IDENTIFIED THAT THE 120V POWER CORD THAT PROVIDED POWER TO THE AC ADAPTER WAS DAMAGED AND NOT MAKING ELECTRICAL CONTACT. THE CAUSE FOR THE DAMAGE TO THE CORD IS ASSOCIATED WITH MISHANDLING. ONCE THE BATTERY DOOR AND 120V CORD WERE REPLACED, NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. | 0
4325910 | .><IT WAS REPORTED THAT THE PHYSICIAN'S PROGRAMMING TABLET WAS GIVING AN ERROR MESSAGE UNABLE TO OPEN PORT". A DIFFERENT USB CABLE WAS USED WITH THE TABLET AND THE ISSUE RESOLVED. THE PHYSICIAN WAS PROVIDED A NEW USB CABLE. THE NONFUNCTIONAL USB CABLE IS EXPECTED TO BE RETURNED FOR ANALYSIS | 0
3505870 | .><IT WAS REPORTED THAT THE PHYSICIAN WAS UNABLE TO GET THE PROGRAMMING SYSTEM TO TURN ON TO ALLOW HIM TO GET THE PATIENT'S VNS PARAMETERS. REVIEW OF THE DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION. FOLLOW UP WITH THE PHYSICIAN CONFIRMED THAT EVERYTHING WAS ALRIGHT WITH THE PATIENT AND THE ISSUE WAS WITH THE HANDHELD. THE HANDHELD WAS PLUGGED INTO THE OUTLET FROM FRIDAY UNTIL MONDAY AND WAS ON THE FLOOR; HOWEVER, THE HANDHELD DID NOT TURN ON THAT MONDAY. THE PHYSICIAN TRIED TO POWER UP THE DEVICE AGAIN ON (B)(6) 2013 AFTER HAVING IT UNPLUGGED. THE PHYSICIAN PLUGGED THE HANDHELD INTO THE OUTLET AND THE SCREEN CAME ON RIGHT AWAY. THE FOLLOWING DAY, THE PATIENT CAME INTO THE OFFICE AND THE PHYSICIAN WAS ABLE TO USE THE HANDHELD WITHOUT ANY PROBLEMS. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED. | 0
3821442 | DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><THE HANDHELD SERIAL CABLE WAS RECEIVED FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><.><AN ANALYSIS WAS PERFORMED ON THE RETURNED SERIAL CABLE AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE SERIAL CABLE WERE NOTED DURING TESTING. THE SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE COMPANY REPRESENTATIVE WAS EXPERIENCING DIFFICULTY INTERROGATING A GENERATOR INTRAOPERATIVELY. IT WAS REPORTED THAT THE HOSPITAL'S HANDHELD WAS ABLE TO COMMUNICATE WITH THE DEVICE. TROUBLESHOOTING WAS PERFORMED WHICH NARROWED DOWN THE HANDHELD SERIAL CABLE TO THE REASON THE DEVICE WAS UNABLE TO BE INTERROGATED. THE HANDHELD SERIAL CABLE IS EXPECTED TO BE RETURNED, BUT HAS NOT BEEN RECEIVED TO DATE. | 0
3469202 | ON (B)(6) 2013, IT WAS REPORTED THAT THE PHYSICIAN TRIED TO DISABLE THE PATIENT'S DEVICE TO 0 MA; HOWEVER, WHEN HE PRESSED THE ZERO VALUE, THE HANDHELD DID NOT RESPOND TO THAT CHOICE. FOLLOW UP WITH THE PHYSICIAN FOUND THAT AFTER TROUBLESHOOTING WITH THE MANUFACTURER'S REPRESENTATIVE, THE PHYSICIAN WAS ABLE TO DISABLE THE DEVICE SO THE PATIENT COULD GET ECT THERAPY. THE PROGRAMMING SYSTEM WAS THEREFORE FUNCTIONING OKAY AND THE PATIENT'S DEVICE WAS ABLE TO BE PROGRAMMED. REVIEW OF THE HANDHELD DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><. | 0
4220701 | .><IT WAS INITIALLY REPORTED THAT THE PHYSICIAN WAS RECEIVING AN ERROR ON HIS TABLET. TROUBLESHOOTING WAS PERFORMED AND THE ISSUE WAS ISOLATED TO A FAULTY USB SERIAL ADAPTOR. THE USB SERIAL ADAPTOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. PRODUCT ANALYSIS IS PLANNED BUT HAS NOT BEEN COMPLETED.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><ADDITIONAL INFORMATION WAS RECEIVED THAT PRODUCT ANALYSIS WAS COMPLETED ON THE USB CABLE. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NOT FURTHER ANOMALIES WERE IDENTIFIED. | 0
4084021 | AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD SERIAL CABLE, AND THE REPORTED ALLEGATION WAS VERIFIED. VISUAL ANALYSIS OF THE SERIAL CABLE VERIFIED THAT THE CABLE WAS PULLED OUT OF THE TRANSCEIVER CASE STRAIN RELIEF. ALSO DURING TESTING, IT WAS IDENTIFIED THAT THE SERIAL CABLE WAS UNABLE TO ESTABLISH COMMUNICATION USING A KNOWN GOOD WAND AND HANDHELD. THE CAUSE FOR THE COMMUNICATION DIFFICULTIES IS ASSOCIATED WITH THREE BROKEN WIRE CONNECTIONS TO THE TRANSCEIVER PCB. ONCE THE WIRES WERE SOLDERED BACK ONTO THE TRANSCEIVER PCB, THE SERIAL CABLE WAS ABLE TO ESTABLISH COMMUNICATION BETWEEN THE HANDHELD AND WAND. THE CAUSE FOR THE REPORTED ALLEGATIONS ARE MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH.><IT WAS REPORTED THAT THE EPILEPTOLOGIST HAD A FRAYED SERIAL CABLE FOR HIS HANDHELD DEVICE AND THAT HE WAS UNABLE TO USE HIS PROGRAMMING SYSTEM IN MULTIPLE PATIENTS. THE PROGRAMMING WAND WAS TESTED WITH A PROGRAMMING TABLET AND CONFIRMED TO BE FUNCTIONING PROPERLY. THE FRAYED SERIAL CABLE HAS NOT BEEN RETURNED TO DATE.><IT WAS REPORTED THAT THE HANDHELD SERIAL CABLE WAS RETURNED. THE PRODUCT WAS RECEIVED FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE. GOOD FAITH ATTEMPTS TO OBTAIN PRODUCT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><.><. | 0
3812355 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><THE HANDHELD ANALYSIS WAS COMPLETED ON (B(4) 2014. NO ANOMALIES ASSOCIATED WITH THE BATTERY LATCH WERE IDENTIFIED. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE MAIN BATTERY WAS SWOLLEN. THE SWOLLEN BATTERY BENT THE BATTERY COVER CAUSING THE BATTERY LATCH SWITCH TO REGISTER THAT THE LATCH WAS UNLOCKED, THUS CAUSING THE HANDHELD TO POWER OFF. NO FURTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS HARD TO POWER ON AFTER BEING CHARGED. ONCE THE HANDHELD POWERED ON AN INTERROGATION WAS ATTEMPTED AT WHICH TIME THE HANDHELD POWERED OFF. UPON RESTARTING THE HANDHELD AN ERROR MESSAGE WAS OBTAINED INDICATING THAT THE BATTERY DOOR LATCH WAS OPEN. TROUBLESHOOTING CONFIRMED THAT THE LATCH WAS CLOSED; HOWEVER, IT WAS NOTED THAT THE DOOR APPEARED TO BE BOWED. THE BATTERY WAS CONFIRMED TO BE NORMAL IN SIZE AND THE HANDHELD WAS STORED AT NORMAL RECOMMENDED TEMPERATURES. THE BATTERY CONFIRMED TO BE AT 100% AND AGAIN SHUT DOWN WITH THE SAME ERROR MESSAGE BEING RECEIVED UPON RESTARTING A SECOND TIME. A NEW PROGRAMMING COMPUTER WAS PROVIDED TO THE PHYSICIAN AND THE HANDHELD AND FLASHCARD WERE RETURNED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
4027185 | THIS EVENT IS A KNOWN ISSUE WITH SOFTWARE VERSION 8.0 AND RIOR MODEL 250 SOFTWARE, AND V10.0 SOFTWARE IS THE SAME SOFTWARE AS V8.0 ON A TABLET INTERFACE. THIS ISSUE HAS BEEN CORRECTED WITH V8.1 SOFTWARE. IN THESE EVENTS, MULTIPLE MAGNET ACTIVATION TIMES BEING DISPLAYED ON THE SAME LINE WHEN VIEWING PATIENT MAGNET HISTORY WITH VNS PROGRAMMING SOFTWARE. THE ISSUE IS RELATED TO A LOCAL ARRAY BEING MIS-DECLARED AS A STATIC VARIABLE WITHIN THE PROGRAMMING SOFTWARE WHICH IS TRIGGERED WHEN VIEWING PATIENT MAGNET ACTIVATION HISTORY WHICH SPANS THE TIME PERIOD BEFORE AND AFTER A GENERATOR TOTAL OPERATING TIME ROLLOVER.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN ASSISTANT'S HANDHELD HAD A DARK SCREEN. THE POWER BUTTON WAS PRESSED WHICH ILLUMINATED THE SCREEN BRIEFLY, BUT THE HANDHELD SCREEN RETURNED DARK. DURING TROUBLESHOOTING THE COMPANY REPRESENTATIVE IDENTIFIED THAT THE HANDHELD DID NOT SHOW SMALL ICONS ON THE BOTTOM OF THE DATABASE SCREEN. THE PHYSICIAN ASSISTANT WAS PROVIDED A NEW PROGRAMMING TABLET. THE HANDHELD IS EXPECTED TO BE RETURNED, BUT HAS NOT BEEN RECEIVED TO DATE. | 0
3546419 | ON (B)(6) 2013, IT WAS REPORTED THAT THIS PHYSICIAN¿S HANDHELD DEVICE WOULD NOT RESPONDING SCREEN TOUCHES AROUND THE ¿INTERROGATE DEVICE¿ SCREEN. A HARD RESET WAS PERFORMED, AND THE DEVICE PROCEEDED PAS THE SCREEN ALIGNMENT WITH NO ERROR. ALL THE OTHER MAIN MENU OPTIONS WORKED (I.E. USER PREFERENCES AND VIEW DATABASE) AS INTENDED. THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED AND ARE PENDING ANALYSIS.><. | 0
3938100 | ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD AND THE REPORTED ALLEGATION WAS VERIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. THE HANDHELD SERIAL CABLE WAS RETURNED FOR ANALYSIS. ANALYSIS OF THE SERIAL CABLE WAS COMPLETED ON 06/24/2014. DURING THE ANALYSIS IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD AN OPEN ELECTRICAL CONNECTION IN THE POWER PLUG WIRE. THE CAUSE FOR THE OPEN CONNECTION IS UNKNOWN AND COULD NOT BE IDENTIFIED DURING THE ANALYSIS. THIS WAS CAPTURED ON MANUFACTURER REPORT # 1644487-2014-01349.><DEVICE MANUFACTURE DATE: THE INCORRECT PRODUCT INFORMATION WAS INADVERTENTLY PUT ON THE INITIAL REPORT.><IT WAS REPORTED THAT THE SURGEON¿S HANDHELD DEVICE WOULD NOT POWER ON WHEN THE SURGEON WAS IN THE OR. WHEN NOT IN USE, THE HANDHELD DEVICE REMAINED PLUGGED IN TO THE POWER ADAPTOR TO CHARGE THE DEVICE. THE LOCK SWITCH AND BATTERY HATCH WERE CONFIRMED NOT TO BE THE SOURCE OF THE ISSUE. THE HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE.><. | 0
3548194 | .><THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><ON (B)(6) 2013, IT WAS REPORTED THAT THIS PHYSICIAN HAD AN ISSUE WITH HIS PROGRAMMING SYSTEM. THE PHYSICIAN REPORTED THAT THE HANDHELD DEVICE WILL NOT HOLD A CHARGE: THE HHD DIES UPON BEING REMOVED FROM THE CHARGER. THE SUSPECT DEVICE HAS NOT BEEN RETURNED TO DATE.><AN ANALYSIS WAS PERFORMED ON THE RETURNED FLASHCARD, AND NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD, AND THE REPORTED FAILURE TO HOLD A CHARGE ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
4102901 | .><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><ANALYSIS OF THE RETURNED HANDHELD DEVICE AND SOFTWARE FLASHCARD WAS COMPLETED. IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE NURSE PRACTITIONER'S HANDHELD WAS NOT HOLDING A CHARGE, NOR WOULD THE SCREEN RESPOND AFTER A HARD RESET. TROUBLESHOOTING WAS PERFORMED BY CONFIRMING THAT THE HARD RESET WAS PERFORMED CORRECTLY AND THE SCREEN WAS CLEANED, THE BATTERY WAS REMOVED AND REINSERTED AND THE SCREEN LOCK WAS CONFIRMED TO NOT BE ENGAGED. TROUBLESHOOTING WAS UNABLE TO RESOLVE THE ISSUES. THE NURSE PRACTITIONER WAS PROVIDED A NEW PROGRAMMING TABLET AND THE HANDHELD WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
3479375 | (B)(4).><THE PHYSICIAN'S OFFICE REPORTED NO ISSUES WITH THE NEW REPLACEMENT TABLET AND EXISTING WAND. GOOD FAITH ATTEMPTS FOR PRODUCT RETURN OF THE HANDHELD DEVICE HAVE BEEN UNSUCCESSFUL TO DATE. THE HANDHELD DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS NOT ALWAYS RESPONDING TO TAPPING ON THE SCREEN WHEN TRYING TO INTERROGATE. IT WAS REPORTED THAT THIS HAS BEEN OCCURRING FOR APPROXIMATELY TWO WEEKS. THE PHYSICIAN WAS GIVEN A NEW PROGRAMMING TABLET. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><THE HANDHELD AND FLASHCARD WAS RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2014. THE SERIAL CABLE WAS NOT RETURNED FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
3976506 | .><ANALYSIS OF THE RETURNED HANDHELD DEVICE WAS COMPLETED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE PSYCHIATRIST¿S HANDHELD DEVICE WAS NO LONGER HOLDING A CHARGE. THE PSYCHIATRIST POWER CYCLED THE DEVICE AND ATTEMPTED TO CHARGE THE HANDHELD WITH THE POWER ADAPTOR; HOWEVER, THE ISSUE DID NOT RESOLVE. THE HANDHELD DEVICE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY. | 0
4084091 | THE HANDHELD AND FLASHCARD WERE RECEIVED FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON 11/18/2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON 11/18/2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WILL ONLY INTERROGATE PATIENT'S DEVICES WHEN IT IS PLUGGED INTO THE CHARGER. IT WAS REPORTED THAT WHENT HE HANDHELD IS REMOVED FROM THE CHARGER THE CHARGED LIGHT IS GREEN, BUT WILL NOT COMMUNICATE WITH THE GENERATORS. A NEW PROGRAMMING COMPUTER WAS REQUESTED. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE. | 0
4101663 | .><IT WAS REPORTED THAT PLUGGING IN THE HANDHELD TO CHARGE RESOLVED THE ISSUE.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS NOT HOLDING A CHARGE. A NEW PROGRAMMING TABLET WAS REQUESTED. TROUBLESHOOTING PERFORMED BY A COMPANY REPRESENTATIVE IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON. A HARD RESET WAS PERFORMED AFTER THE HANDHELD WAS PLUGGED INTO THE WALL AND THE UNIT POWERED ON. THE INITIAL START UP PROCESS FOUND NO ISSUES. THE HANDHELD WAS ABLE TO INTERROGATE A DEMO GENERATOR WITHOUT DIFFICULTY. THE HANDHELD WAS RECEIVING CHARGE WHILE PLUGGED INTO THE WALL, BUT IMMEDIATELY GAVE A WARNING THAT THE BATTERY WAS DEPLETED. THE COMPANY REPRESENTATIVE REPORTED THAT AFTER 20 MINUTES OF CHARGE THE HANDHELD CHARGING BAR HAD PROGRESSED SINCE INITIALLY PLUGGING THE UNIT INTO THE ELECTRICAL OUTLET. IT IS UNKNOWN IF THE HANDHELD CONTINUED TO CHARGE AND WHETHER OR NOT THE CHARGE WAS RETAINED BY THE UNIT. NO ADDITIONAL RELEVANT INFORMATION HAS BEEN RECEIVED TO DATE. | 0
4178457 | ANALYSIS OF THE RETURNED HANDHELD DEVICE AND SOFTWARE FLASHCARD WAS COMPLETED. IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ESTABLISH COMMUNICATION WITH A KNOWN GOOD WAND AND GENERATOR. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A BROKEN WIRE CONNECTION IN THE SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED. THE HANDHELD DID NOT SHUT-DOWN OR POWER-OFF BY ITSELF DURING THE ANALYSIS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS NOT WORKING PROPERLY. IT WAS REPORTED THAT THE HANDHELD WAS POWERING DOWN WITHOUT BEING PROMPTED. TROUBLESHOOTING WAS PERFORMED, BUT DID NOT RESOLVE THE ISSUE. A NEW PROGRAMMING TABLET WAS PROVIDED TO THE PHYSICIAN. THE HANDHELD WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
4077155 | ANALYSIS OF THE RETURNED HANDHELD DEVICE AND SOFTWARE FLASHCARD WAS COMPLETED. IT WAS IDENTIFIED THAT THE HANDHELD WAS RECEIVED WITH THE LOCK BUTTON IN THE LOCKED POSITION. SINCE THE LOCK BUTTON WAS LOCKED, THE HANDHELD BUTTONS AND TOUCHSCREEN WERE UNRESPONSIVE. ONCE THE LOCK BUTTON WAS MOVED TO THE UNLOCKED POSITION NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE HANDHELD DEVICE WAS NOT WORKING PROPERLY. THE HANDHELD WAS RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><NAME AND ADDRESS, TITLE, FIRST NAME, LAST NAME, ADDRESS, ZIP/POSTAL, PHONE #; CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR PROVIDED THE WRONG INITIAL REPORTER. OCCUPATION; CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR PROVIDED THE WRONG INITIAL REPORTER. INITIAL REPORTER ALSO SENT REPORT TO FDA; CORRECTED DATA: THE PREVIOUSLY SUBMITTED MDR PROVIDED THE WRONG INITIAL REPORTER. IT IS UNKNOWN IF THE REPORTER ALSO REPORTED THE EVENT TO FDA.><. | 0
3926575 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><ANALYSIS OF THE RETURNED HANDHELD SERIAL CABLE WAS COMPLETED. NO ANOMALIES ASSOCIATED WITH THE SERIAL CABLE COMMUNICATION WERE IDENTIFIED DURING THE ANALYSIS. HOWEVER, IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD 2 OPEN ELECTRICAL CONNECTIONS IN THE POWER PLUG WIRE. BECAUSE OF THE OPEN WIRE CONNECTIONS, THE AC ADAPTER WAS UNABLE TO SUPPLY POWER TO THE HANDHELD. NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE SERIAL CABLE FOR THE HANDHELD DEVICE WAS CAUSING INTERMITTENT COMMUNICATION ISSUES DUE TO A LOOSE CONNECTION. TROUBLESHOOTING WAS PERFORMED AND DETERMINED THE ROOT CAUSE OF THE ISSUES WAS THE SERIAL CABLE. THE SERIAL CABLE HAS NOT BEEN RETURNED TO DATE.><THE HANDHELD SERIAL CABLE WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><DEVICE MANUFACTURING RECORDS WERE REVIEWED.REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE HANDHELD DEVICE PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. | 0
4274319 | .><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WAS UNRESPONSIVE. IT WAS REPORTED THAT THE HANDHELD WOULD NOT ADVANCE PAST THE ALIGNMENT SCREEN. TROUBLESHOOTING WAS PERFORMED BY A COMPANY REPRESENTATIVE. THE HANDHELD WAS POWERED OFF AND THEN POWERED BACK ON, A HARD RESET WAS PERFORMED, AND THE BATTERY WAS REMOVED AND REINSERTED. THE HANDHELD WAS PLUGGED INTO AN ELECTRICAL OUTLET AND THE UNPLUGGED AND THE SCREEN WAS CLEANED AROUND THE EDGES. THE SCREEN LOCK WAS NOT ENGAGED. THE HANDHELD WAS NOT ABLE TO ADVANCE PAST THE ALIGNMENT SCREEN AFTER TROUBLESHOOTING. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET. THE HANDHELD IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><THE HANDHELD AND SOFTWARE WERE RECEIVED BY THE MANUFACTURER. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD. NO SOFTWARE ERRORS ASSOCIATED WITH THE HANDHELD WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
4318602 | (B)(4). CORRECTED DATA: INADVERTENTLY DID NOT INCLUDE THE UDI ON THE INITIAL REPORT.><.><IT WAS REPORTED ON (B)(6) 2014 THAT DURING THE PATIENT¿S GENERATOR REPLACEMENT, PRIOR TO OPENING THE NEW GENERATOR THEY ATTEMPTED TO INTERROGATE IT AND WAS RECEIVING A ¿FAILURE TO OPEN PORT¿ MESSAGE. THE MESSAGE DISPLAYED INSTANTLY WHEN SHE SELECTED THE ¿INTERROGATE¿¿ OPTION ON THE FIRST SCREEN. IT DIDN¿T ADVANCE TO THE SCREEN WITH THE OPTIONS TO ¿START INTERROGATION¿ OR ¿CANCEL¿. THE WAND WAS UNPLUGGED AND RE-PLUGGED AND STILL RECEIVED THE SAME MESSAGE. THE TABLET WAS TURNED OFF AND BACK ON AND STILL RECEIVED THE SAME MESSAGE. IT WAS NOTED THAT THE PROGRAMMING SYSTEM WAS FINALLY ABLE TO INTERROGATE THE PATIENT AND THE SURGERY WAS COMPLETED WITH NO FURTHER ISSUES. IT WAS NOTED THAT THE PHYSICIAN IS GOING TO KEEP THE TABLET AND WILL SEND IT BACK FOR PRODUCT ANALYSIS IF HE EXPERIENCES THE SAME ISSUES AGAIN. | 0
3782117 | A COMPANY REPRESENTATIVE INDICATED THAT THE HANDHELD WAS NOT HOLDING A CHARGE. IT WAS REPORTED THAT THE HANDHELD WAS LEFT PLUGGED INTO THE OUTLET ALL NIGHT AND WHEN UNPLUGGED THE HANDHELD IMMEDIATELY LOST THE CHARGE. THE HANDHELD WAS PLUGGED INTO ANOTHER OUTLET WHICH DID NOT RESOLVE THE ISSUE. A NEW HANDHELD WAS PROVIDED TO THE COMPANY REPRESENTATIVE AND THE HANDHELD AND FLASHCARD THAT WOULD NOT CHARGE WERE RETURNED TO MANUFACTURER FOR ANALYSIS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2014. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT SOME OF THE SYNC CABLE CONNECTOR LEADS WERE DAMAGED AND THAT SOME OF THE SOLDER CONNECTIONS BETWEEN THE SYNC CABLE CONNECTOR LEADS AND THE HANDHELD PCB WERE CRACKED. BECAUSE OF THE DAMAGE, THE HANDHELD WAS UNABLE TO RECEIVE POWER FROM THE AC ADAPTER. THE CAUSE FOR THE DAMAGE TO THE CONNECTOR AND SOLDER CONNECTIONS IS MOST LIKELY ASSOCIATED WITH MISHANDLING OF THE DEVICE AS IT APPEARS THE CONNECTOR DETENTS ARE NOT BEING COMPRESSED WHEN REMOVING OR MANIPULATING THE SERIAL DATA CABLE, THUS PLACING AN EXTENSIVE AMOUNT OF FORCE ON THE PCB AND ATTACHED CABLE RECEPTACLE. NO OTHER ANOMALIES WERE IDENTIFIED. ANALYSIS OF THE FLASHCARD WAS RECEIVED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
3736396 | .><THE PHYSICIAN INDICATED THAT HE WANTED REPLACEMENT OF HIS HANDHELD. IT WAS REPORTED THAT TROUBLESHOOTING WAS PERFORMED AND IDENTIFIED THAT THE HANDHELD WAS THE ISSUE. THE PHYSICIAN REPORTED THAT HE WAS ABLE TO USED HIS PROGRAMMING WAND AND ANOTHER HANDHELD WITHOUT ANY ISSUES. A REPLACEMENT TABLET WAS PROVIDED TO THE PHYSICIAN AND THE HANDHELD WAS RETURNED FOR ANALYSIS. THE HANDHELD WAS RECEIVED FOR ANALYSIS WITHOUT AN AC ADAPTER. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(6) 2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING A KNOWN GOOD AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(6) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
4295121 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN HANDHELD DEVICE WAS NOT FUNCTIONING PROPERLY. DETAILS REGARDING THE ISSUES WERE NOT PROVIDED. TROUBLESHOOTING WAS PERFORMED, BUT THESE UNKNOWN ISSUES DID NOT RESOLVE. THE HANDHELD DEVICE AND SOFTWARE FLASHCARD HAVE BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY. ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><ANALYSIS OF THE RETURNED HANDHELD DEVICE AND SOFTWARE FLASHCARD WAS COMPLETED. IT WAS IDENTIFIED THAT THE HANDHELD WOULD NOT POWER ON. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH A SWOLLEN MAIN BATTERY. THE SWOLLEN BATTERY BENT THE BATTERY COVER CAUSING THE BATTERY LATCH SWITCH TO REGISTER THAT THE LATCH WAS UNLOCKED, THUS PREVENTING THE HANDHELD FROM POWERING ON. ADDITIONALLY, IT WAS IDENTIFIED THAT THE MAIN BATTERY COULD NOT POWER THE HANDHELD FOR MORE THAN AN HOUR GIVING THE INDICATION THAT THE MAIN BATTERY IS DEFECTIVE. NO OTHER ANOMALIES WERE IDENTIFIED. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 0
3970466 | THE HANDHELD ANALYSIS WAS COMPLETED ON 08/07/2014. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON 08/07/2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE SURGEON¿S PROGRAMMING SYSTEM WAS NOT WORKING WHILE IN THE OR. THE HANDHELD DEVICE AND PROGRAMMING WAND WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. THE PROGRAMMING WAND PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE HANDHELD DEVICE IS CURRENTLY UNDERWAY.><. | 0
4251804 | IT WAS REPORTED ON (B)(6), 2014 THAT THE HANDHELD HAD A CRACKED SCREEN. IT WAS NOTICED UPON ITS ARRIVAL WHEN IT WAS SHIPPED TO THE DISTRIBUTOR. THE HANDHELD AND FLASHCARD WERE RETURNED FOR PRODUCT ANALYSIS ON (B)(6) 2014. PRODUCT ANALYSIS IS STILL UNDERWAY AND HAS NOT BEEN COMPLETED TO DATE.><.><PRODUCT ANALYSIS WAS COMPLETED ON THE HANDHELD AND FLASHCARD ON 11/21/2014. A VISUAL ANALYSIS OF THE HANDHELD WAS ABLE TO VERIFY THAT THE DISPLAY WAS CRACKED. THE CAUSE FOR THE CRACKED SCREEN IS ASSOCIATED WITH MISHANDLING OF THE DEVICE. ONCE THE SCREEN WAS REPLACED WITH A KNOWN GOOD SCREEN, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
3790460 | .><IT WAS REPORTED THAT THE NURSE¿S HANDHELD DEVICE HAD RESET ITSELF A COUPLE OF TIMES. ATTEMPTS FOR ADDITIONAL RELEVANT INFORMATION WERE MADE BUT HAVE BEEN UNSUCCESSFUL TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT THE HANDHELD ISSUE DID NOT OCCUR FREQUENTLY AND HAD NOT REOCCURRED SINCE THE NURSE WAS INSTRUCTED TO REPLACE THE HANDHELD BATTERY. | 0
4118368 | .><MANUFACTURER DEVICE HISTORY RECORDS WERE REVIEWED. REVIEW OF THE HANDHELD MANUFACTURER DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY SPECIFICATIONS WERE PASSED PRIOR TO DISTRIBUTION.><THE HANDHELD AND RELATED SOFTWARE PRODUCT INFORMATION WAS OBTAINED. THE HANDHELD SERIAL CABLE ADAPTOR HAS NOT BEEN RECEIVED BY THE MANUFACTURER TO DATE.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><THE HANDHELD SERIAL ADAPTER CABLE WAS RECEIVED BY THE MANUFACTURER FOR ANALYSIS. HOWEVER, ANALYSIS HAS NOT BEEN COMPLETED TO DATE.><ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE HANDHELD WAS FUNCTIONING PROPERLY AND THE CHARGING ISSUES WERE ATTRIBUTED TO A BROKEN SERIAL CABLE ADAPTOR. THE BROKEN SERIAL CABLE ADAPTOR HAS NOT BEEN RETURNED TO DATE.><IT WAS REPORTED THAT THE PHYSICIAN¿S HANDHELD DEVICE WAS NOT HOLDING A CHARGE AND WAS NOT POWERING ON. THE PHYSICIAN USED DIFFERENT OUTLETS THROUGHOUT HIS OFFICE BUT WAS UNABLE TO RESOLVE THE CHARGING ISSUE. IT WAS NOTED THAT THE ORANGE CHARGING INDICATOR WAS NOT DISPLAYED. THE HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE.><AN ANALYSIS WAS PERFORMED ON THE RETURNED SERIAL CABLE AND A CAUSE FOR THE REPORTED ALLEGATION WAS IDENTIFIED. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD AN OPEN ELECTRICAL CONNECTION IN THE POWER PLUG WIRE. THE CAUSE FOR THE OPEN CONNECTION IS UNKNOWN AND COULD NOT BE IDENTIFIED DURING THE ANALYSIS. | 0
3687767 | ANALYSIS OF A RETURNED HANDHELD WAS COMPLETED ON (B)(6) 2014. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD AN OPEN ELECTRICAL CONNECTION IN THE POWER CABLE. AS A RESULT OF THE OPEN WIRE CONNECTION, THE HANDHELD WOULD RECEIVE POWER FROM THE AC ADAPTER INTERMITTENTLY. NO OTHER ANOMALIES WERE IDENTIFIED.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. | 0
4331845 | ON 11/(B)(6) 2014, IT WAS REPORTED THAT THE PHYSICIAN¿S PROGRAMMING HANDHELD ¿DIES A LOT¿. THE HANDHELD HAS NOT BEEN RETURNED FOR PRODUCT ANALYSIS TO DATE.><.><THE HANDHELD AND FLASHCARD WERE RECEIVED FOR PRODUCT ANALYSIS ON 02/03/2015. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. | 0
3384974 | .><ON (B)(6) 2013, THE PHYSICIAN REPORTED THAT THE HANDHELD WAS STUCK AT THE TAP SCREEN AT THE INITIAL SCREEN. THE PHYSICIAN SAID THAT THE HANDHELD WAS NOT RESPONDING TO THE SCREEN TAPS. AFTER PERFORMING A HEARD RESET, THE PHYSICIAN SAID THAT THE HANDHELD WAS STUCK ON THE FIRST SCREEN THAT SAYS TO TAP THE SCREEN (THE ALIGNMENT SCREEN"). DESPITE MULTIPLE HARD RESETS | 0
3630159 | IT WAS REPORTED THAT THE HANDHELD CHARGER WAS NOT WORKING. PER THE NURSE, THE CHARGER WAS DEFINITELY BROKEN AND WOULD NOT CHARGE THE HANDHELD. THE CHARGER OF ANOTHER PROGRAMMING SYSTEM WAS BEING USED TO CHARGE THE HANDHELD AS A SOLUTION TO THE ISSUE. A REPLACEMENT CHARGER WAS REQUESTED. ATTEMPTS HAVE BEEN MADE FOR ADDITIONAL INFORMATION AND PRODUCT RETURN; HOWEVER, THEY WERE UNSUCCESSFUL. NO OTHER INFORMATION HAS BEEN PROVIDED.><THE HANDHELD AND FLASHCARD WERE RECEIVED ON (B)(4) 2014. ANALYSIS OF THE FLASHCARD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. ANALYSIS OF THE HANDHELD WAS COMPLETED ON (B)(4) 2014. NO ANOMALIES ASSOCIATED WITH THE POWER ADAPTER WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS IT WAS IDENTIFIED THAT HANDHELD WAS UNABLE TO CHARGE THE MAIN BATTERY. THE CAUSE FOR THE ANOMALY IS ASSOCIATED WITH BROKEN SOLDER CONNECTIONS ON THE HANDHELD MAIN BOARD.><DEVICE FAILURE OCCURRED BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
4155234 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><MANUFACTURER REVIEWED DEVICE HISTORY RECORDS.REVIEW OF THE TABLET DEVICE HISTORY RECORDS CONFIRMED ALL QUALITY SPECIFICATIONS WERE PASSED PRIOR TO DISTRIBUTION.><AN ANALYSIS WAS PERFORMED ON THE RETURNED TABLET USB TO DB9 CABLE, AND THE REPORTED ALLEGATION WAS VERIFIED. THE CAUSE FOR THE REPORTED ALLEGATION IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION RELATED TO A STRESS-INDUCATED FRACTURE IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NO FURTHER ANOMALIES WERE IDENTIFIED.><IT WAS REPORTED THAT THE MEDICAL PROFESSIONAL¿S PROGRAMMING TABLET SERIAL CABLE FAILED. TROUBLESHOOTING WAS PERFORMED BY THE COMPANY REPRESENTATIVE WHICH CONFIRMED THAT THE SERIAL CABLE WAS ISSUE AS IT WAS TRIED WITH DIFFERENT WANDS AND TABLETS. A REPLACEMENT SERIAL CABLE WAS PROVIDED TO THE SITE, AND THE FAULTY SERIAL CABLE WAS RETURNED TO THE MANUFACTURER FOR ANALYSIS. HOWEVER, PRODUCT ANALYSIS HAS NOT BEEN COMPLETED TO DATE. | 0
3926496 | DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><IT WAS REPORTED THAT THE PHYSICIAN'S PROGRAMMING TABLET WAS GIVING THE MESSAGE UNABLE TO OPEN THE PORT. THE SYSTEM CANNOT FIND THE FILE SPECIFIED." THE SERIAL ADAPTER CORD WAS DISCONNECTED AND INTERROGATION WAS ATTEMPTED AGAIN ON THE PATIENT'S DEVICE AND THE SAME MESSAGE WAS RECEIVED. ANOTHER SERIAL ADAPTER CORD WAS NOT AVAILABLE; HOWEVER | 0
3472350 | ANALYSIS OF THE POWER CORD WAS COMPLETED ON (B)(4) 2013. NO FUNCTIONAL ANOMALIES ASSOCIATED WITH THE AC ADAPTER¿S PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. VISUAL ANALYSIS WAS ABLE TO VERIFY THAT THE OUTER INSULATION ON THE LOW VOLTAGE SIDE WAS ABRADED, EXPOSING THE SHIELD WIRE. SINCE THE SHIELD WIRE IS CONNECTED TO GROUND, THERE WAS NO RISK OF ELECTRICAL SHOCK. NO FUNCTIONAL ANOMALIES ASSOCIATED WITH THE AC ADAPTER WERE NOTED DURING TESTING USING A KNOWN GOOD HANDHELD AND SERIAL CABLE. THE AC ADAPTER PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS OF THE HANDHELD CONFIRMED ALL QUALITY TESTS WERE PASSED PRIOR TO DISTRIBUTION.><IT WAS REPORTED THA THE PHYSICIAN'S HANDHELD WAS NOT POWERING ON. TROUBLESHOOTING WAS PERFORMED AND THE ISSUE WAS NARROWED DOWN TO THE POWER CORD. A NEW POWER CORD WAS PLACED AND THE HANDHELD CHARGE AND POWERED ON. THE PHYSICIAN WAS PROVIDED A NEW POWER CORD AND THE NON FUNCTIONAL POWER CORD WAS RETURNED FOR ANALYSIS. ANALYSIS OF THE POWER CORD IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. | 0
3942164 | ON (B)(6) 2014 THE HANDHELD AND FLASHCARD WERE RETURNED FOR PRODUCT ANALYSIS. AN ANALYSIS WAS PERFORMED ON THE RETURNED HANDHELD AND THE UNRESPONSIVE SCREEN EVENT WAS CONFIRMED. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE HANDHELD WAS UNABLE TO ADVANCE PAST THE WELCOME SCREEN. THE CAUSE FOR THE DISPLAY ANOMALY IS ASSOCIATED WITH RESISTANCE VALUES BEING HIGHER THAN EXPECTED IN THE TOUCH SCREEN CIRCUITRY. ONCE THE DISPLAY WAS REPLACED WITH A KNOWN GOOD DISPLAY, NO FURTHER ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE IDENTIFIED DURING THE ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE NEUROLOGIST¿S HANDHELD SCREEN WAS UNRESPONSIVE. HARD RESETS WERE PERFORMED BUT THE ISSUE DID NOT RESOLVE. THE HANDHELD DEVICE HAS BEEN RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><. | 0
4116009 | DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><.><THE PROGRAMMING TABLET WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS OF THE TABLET WAS COMPLETED ON 12/03/2014. NO ANOMALIES ASSOCIATED WITH THE MAIN BATTERY WERE IDENTIFIED DURING THE ANALYSIS. DURING THE ANALYSIS, IT WAS IDENTIFIED THAT THE TABLET COULD NOT CHARGE THE MAIN BATTERY, AND AS A RESULT WAS ALSO NOT ABLE TO BE POWERED USING THE MAIN BATTERY. THE CAUSE FOR THE IDENTIFIED ANOMALY IS ASSOCIATED WITH A LOOSE BATTERY CABLE TO THE MOTHERBOARD. ONCE THE CABLE WAS RESEATED, NO FURTHER ANOMALIES WERE IDENTIFIED DURING THE ANALYSIS. VISUAL ANALYSIS ALSO IDENTIFIED DAMAGE TO THE BACK COVER OF THE TABLET. THE CAUSE OF THE DAMAGE COULD NOT BE IDENTIFIED DURING THE ANALYSIS.><A COMPANY REPRESENTATIVE'S PROGRAMMING TABLET HAS BEEN SHUTTING OFF AND WILL ONLY STAY OPERATIONAL IF THE DEVICE IS PLUGGED DIRECTLY INTO THE ELECTRICAL OUTLET. THE BATTERY STATUS INDICATOR SHOWS 100% CHARGED AND THE BATTERY STATUS BUTTON ILLUMINATES RED LIGHTS WHEN CHARGING AS EXPECTED. WHEN THE TABLET IS DISCONNECTED THE TABLET POWERS DOWN. THE COMPANY REPRESENTATIVE WAS PROVIDED A NEW PROGRAMMING TABLET. THE NON-FUNCTIONAL TABLET IS EXPECTED TO BE RETURNED FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE. | 0
3985299 | DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><PRODUCT INFORMATION WAS OBTAINED. ANALYSIS OF THE RETURNED SERIAL CABLE WAS COMPLETED. THE CAUSE FOR THE ERROR MESSAGE IS ASSOCIATED WITH A DISCONNECTED WIRE CONNECTION IN THE RETURNED SERIAL CABLE. ONCE THE WIRE WAS SOLDERED ONTO THE SERIAL CABLE PCB, NOT FURTHER ANOMALIES WERE IDENTIFIED.><.><IT WAS REPORTED THAT THE NEUROLOGIST¿S TABLET DEVICE HAD A FAULTY USB ADAPTOR. THE CONNECTION BETWEEN THE USB ADAPTOR AND PROGRAMMING WAND WAS REPORTED AS ¿BAD.¿ THE NEUROLOGIST WAS UNABLE TO PERFORM INTERROGATIONS AND THE TABLET DISPLAYED AN ¿UNABLE TO OPEN PORT¿ ERROR MESSAGE. THE FAULTY USB ADAPTOR WAS RETURNED TO THE MANUFACTURER WHERE ANALYSIS IS CURRENTLY UNDERWAY.><MANUFACTURE DATE, CORRECTED DATA: PREVIOUSLY SUBMITTED MDR INADVERTENTLY DID NOT INCLUDE THE MANUFACTURE DATE. | 0
3665098 | DEVICE FAILURE IS SUSPECTED, BUT DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><THE PHYSICIAN WAS PROVIDED A NEW HANDHELD SERIAL CABLE AND THE ORIGINAL HANDHELD SERIAL CABLE WAS RETURNED FOR ANALYSIS. THE HANDHELD SERIAL CABLE WAS RECEIVED ON (B)(4) 2014. ANALYSIS OF THE HANDHELD SERIAL CABLE IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><IT WAS REPORTED BY THE NEUROLOGIST THAT HIS HANDHELD WORKS ONLY WHILE HOLDING THE SERIAL CABLE IN A SPECIFIC POSITION. COMMUNICATION ISSUES OCCUR WHEN THE CABLE IS NOT HELD.><ANALYSIS OF THE HANDHELD SERIAL CABLE WAS COMPLETED ON (B)(4) 2014. AN ANALYSIS WAS PERFORMED ON THE RETURNED SERIAL CABLE AND THE REPORTED ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE SERIAL CABLE WERE NOTED DURING TESTING. THE SERIAL CABLE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><THE PHYSICIAN WAS PROVIDED A NEW HANDHELD SERIAL CABLE. THIS INFORMATION WAS INADVERTENTLY LEFT OFF OF INITIAL MFR. REPORT. | 0
3824942 | THE PHYSICIAN RETURNED THE REPLACEMENT HANDHELD AND SOFTWARE; HOWEVER, THE NON-FUNCTIONING POWER CORD WAS NOT RETURNED. THE PHYSICIAN KEPT THE REPLACEMENT POWER CORD AND THE NON-FUNCTIONING POWER CORD. THE REPLACEMENT HANDHELD AND FLASHCARD PERFORMED ACCORDING TO SPECIFICATIONS.><IT WAS REPORTED THAT PHYSICIAN WAS HAVING ISSUES WITH HIS HANDHELD DEVICE CABLES. THE PHYSICIAN UNPLUGGED AND RECONNECTED THE POWER ADAPTOR AND CHARGING CABLE BUT THE ISSUE WAS NOT RESOLVED. MISHANDLING AND STORAGE OF THE DEVICE IS NOT SUSPECTED TO HAVE CAUSED THE FAILURE TO CHARGE. THE SUSPECT DEVICE HAS NOT BEEN RETURNED TO DATE.><. | 0
3892039 | IT WAS REPORTED THAT THE PHYSICIAN HANDHELD WAS PRESENTING A FAUX CONTACT". IT WAS REPORTED THAT A REPLACEMENT WOULD BE PROVIDED AND THE HANDHELD WILL BE RETURNED FOR ANALYSIS. THE HANDHELD WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY | 0
3884370 | THE HANDHELD DEVICE AND SOFTWARE FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR ANALYSIS. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><.><IT WAS REPORTED THAT THE NURSE WAS HAVING INTERMITTENT COMMUNICATION ISSUES WITH HER PROGRAMMING SYSTEM. THE WAND SHOWED DATA TRANSMISSION AND THE HANDHELD SHOWED THE INTERROGATION PROGRESS BAR, BUT THEN AN ERROR MESSAGE APPEARED AND THE INTERROGATION FAILED. THE WAND BATTERY WAS REPLACED, THE HANDHELD DEVICE WAS FULLY CHARGED, AND A HARD RESET WAS PERFORMED; HOWEVER, THE ISSUES DID NOT RESOLVE. ANOTHER PROGRAMMING SYSTEM WAS ABLE TO SUCCESSFUL COMMUNICATE WITH THE PATIENTS¿ DEVICES. IT WAS DETERMINED THAT THERE WAS NO ISSUE WITH THE WAND. FURTHER FOLLOW-UP REVEALED THAT THE HANDHELD¿S SCREEN WORKED INTERMITTENTLY. THE SCREEN WOULD FREEZE AND AT ONE POINT PIXILATED OVER AND WENT BLACK. A HARD RESET WAS PERFORMED AND THE SOFTWARE FLASHCARD WAS RESEATED; HOWEVER, THE SCREEN ISSUE CONTINUED TO OCCUR. IT WAS NOTED THAT THE SERIAL CABLE CONNECTION AT THE BASE OF THE HANDHELD WAS LOOSE. THE SERIAL CABLE NEEDED TO BE PLACED IN A SPECIFIC ORIENTATION IN ORDER TO COMMUNICATE. THE HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE. | 0
3562829 | .><THE PHYSICIAN REPORTED THAT THE BATTERY IN THE HANDHELD DIED COMPLETELY AND THAT THE BATTERY IS PROBABLY BAD. THE PHYSICIAN REPORTED THAT HE HAS KEPT THE UNIT PLUGGED INTO THE OUTLET. THE PHYSICIAN HAS ANOTHER PROGRAMMING SYSTEM SO NO PATIENTS WERE AFFECTED. NO ADDITIONAL INFORMATION HAS BEEN OBTAINED TO DATE. | 0
3821329 | IT WAS REPORTED THAT THE NEUROSURGEON¿S TABLET DEVICE COULD NOT BE CONNECTED TO THE PROGRAMMING WAND BECAUSE THE CONNECTION HAD BEEN STRIPPED AND WAS NO LONGER TIGHT. THE TABLET DEVICE IS EXPECTED TO BE RETURNED TO THE MANUFACTURER BUT HAS NOT BEEN RECEIVED TO DATE.><ON (B)(6) 2014, IT WAS REPORTED THAT AN EXHAUSTIVE SEARCH OF THE REPORTING FACILITY WAS UNSUCCESSFUL IN LOCATING THE SUSPECT MEDICAL DEVICE. THE DEVICE IS NOT EXPECTED FOR RETURN AT THIS TIME.><THE PROGRAMMING TABLET IS EXPECTED TO BE RETURNED TO MANUFACTURER FOR ANALYSIS, BUT HAS NOT BEEN RECEIVED TO DATE.><. | 0
3941364 | IT WAS REPORTED THAT THE PHYSICIAN KEPT THE CHARGER AND THAT THE CHARGER WORKS WITH WITH OTHER HANDHELDS.><ANALYSIS OF THE RETURNED HANDHELD DEVICE WAS COMPLETED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD WILL NOT HOLD A CHARGE. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING COMPUTER AND THE HANDHELD WAS RETURNED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE. ATTEMPTS TO OBTAIN ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><. | 0
4095318 | .><IT WAS REPORTED THAT THE PHYSICIAN'S HANDHELD SERIAL CABLE HAS TO BE MANIPULATED AT THE BASE OF THE HANDHELD IN ORDER TO WORK. IT WAS REPORTED THAT IT IS UNCLEAR WHEN THE ISSUE WAS FIRST OBSERVED. A REPLACEMENT SERIAL CABLE WAS PROVIDED TO THE PHYSICIAN AND THE NONFUNCTIONAL SERIAL CABLE WAS RETURNED FOR ANALYSIS. ANALYSIS OF THE SERIAL CABLE IS UNDERWAY, BUT HAS NOT BEEN RECEIVED TO DATE.><DEVICE FAILURE OCCURRED, BUT DID NOT CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY.><ANALYSIS OF THE RETURNED HANDHELD SERIAL CABLE WAS COMPLETED. IT WAS IDENTIFIED THAT THE SERIAL CABLE HAD AN OPEN ELECTRICAL CONNECTION IN THE POWER PLUG WIRE. THE CAUSE FOR THE OPEN CONNECTION IS UNKNOWN AND COULD NOT BE IDENTIFIED DURING THE ANALYSIS. | 0
3889259 | ADDITIONAL INFORMATION WAS RECEIVED THAT THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. PRODUCT ANALYSIS IS PLANNED BUT HAS NOT BEEN COMPLETED.><ADDITIONAL INFORMATION WAS RECEIVED THAT PRODUCT ANALYSIS WAS COMPLETED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD PERFORMANCE WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS. NO ANOMALIES ASSOCIATED WITH FLASHCARD SOFTWARE OR DATABASES WERE IDENTIFIED DURING THE FLASHCARD ANALYSIS. THE FLASHCARD AND SOFTWARE PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><IT WAS REPORTED THAT THE HANDHELD DEVICE WAS NOT PROGRAMMING ON AND DID NOT LOAD THE PROGRAMMING SOFTWARE. A HARD RESET WAS PERFORMED AND THE ISSUE WAS RESOLVED. FURTHER FOLLOW-UP REVEALED THAT THE HANDHELD DEVICE WOULD NOT POWER ON DESPITE CHARGING. THE POWER ADAPTOR AND SERIAL CABLE CONNECTIONS WERE CONFIRMED TO BE SECURE. IT WAS NOTED THAT THE HANDHELD POWER INDICATOR WOULD NOT TURN ON. THE HANDHELD DEVICE HAS NOT BEEN RETURNED TO DATE.><DEVICE MANUFACTURING RECORDS WERE REVIEWED. REVIEW OF MANUFACTURING RECORDS CONFIRMED THAT THE HANDHELD DEVICE PASSED ALL FUNCTIONAL TESTS PRIOR TO DISTRIBUTION. | 0
3850600 | ><PATIENT DATA NOT CURRENTLY AVAILABLE. EMAIL ADDRESS OF REPORTER UNKNOWN. DATE OF MANUFACTURE UNKNOWN AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 0
4319214 | EVALUATION- CONCLUSIONS, CORRECTED DATA: (B)(4).><IT WAS REPORTED THAT THE HANDHELD WOULD NOT CHARGE WHEN PLUGGED INTO THE ELECTRICAL OUTLET. IT WAS REPORTED THAT WHILE PLUGGED IN, THE POWER INDICATOR TURNS GREEN; HOWEVER, THE HANDHELD WOULD NOT POWER ON. IT WAS LATER REPORTED THAT THE HANDHELD WAS FINE; HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. ATTEMPTS TO OBTAIN ADDITIONAL RELEVANT INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE.><. | 1
3488661 | IT WAS INITIALLY REPORTED THAT THE PHYSICIAN¿S HANDHELD WOULD NOT HOLD A CHARGE. THE COMPANY REPRESENTATIVE WENT TO THE SITE BUT WAS UNABLE TO TROUBLESHOOT THE HANDHELD. NO PATIENT WAS ADVERSELY AFFECTED BY THESE ISSUES. THE HANDHELD AND FLASHCARD WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. PRODUCT ANALYSIS IS PLANNED BUT HAS NOT BEEN COMPLETED.><PRODUCT ANALYSIS WAS APPROVED ON (B)(4) 2013. AN ANALYSIS WAS PERFORMED ON THE HANDHELD AND THE REPORTED ¿FAILURE TO HOLD A CHARGE¿ ALLEGATION WAS NOT VERIFIED. NO ANOMALIES ASSOCIATED WITH THE HANDHELD WERE NOTED DURING TESTING USING THE AC ADAPTER OR THE MAIN BATTERY WITH A FULL CHARGE. THE HANDHELD PERFORMED ACCORDING TO FUNCTIONAL SPECIFICATIONS.><. | 1
4259346 | .><IT WAS REPORTED THAT THE PHYSICIAN'S PROGRAMMING TABLET SCREEN WAS CRACKED AND IT WAS UNKNOWN HOW THIS OCCURRED. THE PHYSICIAN WAS PROVIDED A NEW PROGRAMMING TABLET. THE TABLET WITH THE CRACKED SCREEN IS EXPECTED TO BE RETURNED, BUT HAS NOT BEEN RECEIVED TO DATE.><THE PROGRAMMING TABLET WAS RECEIVED FOR ANALYSIS. ANALYSIS IS UNDERWAY, BUT HAS NOT BEEN COMPLETED TO DATE.><ANALYSIS OF THE TABLET WAS COMPLETED ON 03/06/2015. A VISUAL INSPECTION OF THE TABLET VERIFIED THAT THE DISPLAY WAS CRACKED. THE CAUSE FOR THE CRACKED SCREEN IS ASSOCIATED WITH MISHANDLING OF THE TABLET. NO FURTHER ANOMALIES ASSOCIATED WITH THE TABLET WERE IDENTIFIED DURING THE ANALYSIS. | 1
4036722 | IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO® 3 SYSTEM, AND UNWANTED PACING MALFUNCTION OCCURRED. THE PROCEDURE WAS COMPLETED WITHOUT PATIENT CONSEQUENCE. PER 21 CFR, PART 803, THIS COMPLAINT IS REPORTABLE BECAUSE DEVICE MALFUNCTION OF THIS TYPE COULD POTENTIALLY CONTRIBUTE TO A DEATH OR SERIOUS INJURY.><THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. CONCOMITANT PRODUCTS USED: PRODUCT NAME:SMART TOUCH NAV: US CATALOG #: UNKNOWN, LOT #: UNKNOWN. PRODUCT NAME:BIOSENSE WEBSTER CS: US CATALOG #: UNKNOWN, LOT #: UNKNOWN. (B)(4).><MANUFACTURER¿S REFERENCE #: (B)(4) IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO® 3 SYSTEM, AND UNWANTED PACING MALFUNCTION OCCURRED. THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED. BACKPLANE WAS REPLACED AND ISSUE WAS RESOLVED. BACKPLANE WAS SENT TO HTC FOR INVESTIGATION AND THE ISSUE WAS NOT DUPLICATED CUSTOMER COMPLAINT WAS NOT CONFIRMED. THE CARD TESTED AT AQUARIUM STATION AND PASSED ECG TESTS AT SERVICE JIG UTILITY .THE CARD PASSED ISOLATION TEST. REPORTED PROBLEM WAS NOT CONFIRMED. THE HISTORY OF CUSTOMER COMPLAINTS ASSOCIATED WITH CARTO 3 SYSTEM # 11222 WAS REVIEWED. 4 OUT OF 28 ADDITIONAL REPORTED COMPLAINTS MAY BE RELATED TO THE REPORTED ISSUE. A DHR REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. | 1
4226110 | OPPORTUNITY FOR ERROR IDENTIFIED. LOOKING FOR SAFEGUARD TO ADD TO MACHINE. MACHINE ALARMS WHEN CONCENTRATION OF SOLUTION NEEDS TO BE CORRECTED (XYLENE, ALCOHOL, ETC). MACHINE ASKS TECHNICIAN, DID YOU CHANGE SOLUTION?" THERE IS THE POTENTIAL FOR TECHNICIAN TO RESPOND "YES" | 1
3931729 | IT WAS REPORTED THAT ARCHIVED DATA WAS REMOVED FROM THE LONG TERM STORAGE (LTS) DEVICE MOUNTED ON THE SYNGO IMAGING SYSTEM. AFTER LTS ARCHIVING PROCEDURE WAS CHANGED AT THIS CUSTOMER'S SITE, IT WAS DETERMINED THAT APPROXIMATELY 24 MILLION OBJECTS OUT OF MORE THAN 100 MILLION THAT CONTAINED PATIENT IMAGES WERE REMOVED FROM THE LTS. THE POSSIBILITY TO RECOVER THIS DATA ON THE FILE SYSTEM LEVEL IS CURRENTLY UNDER INVESTIGATION.><ACCORDING TO OUR FACTORY EXPERTS THERE ARE TWO PROCEDURES AVAILABLE TO PERFORM LTS ARCHIVING IN SYNGO IMAGING: IMPLICIT AND EXPLICIT ARCHIVING. WITH IMPLICIT ARCHIVING" THE DATA TO BE ARCHIVED IS DIRECTLY WRITTEN INTO A FOLDER ON THE ARCHIVE. THIS FOLDER IS USUALLY MOUNTED ON THE SYNGO IMAGING SYSTEM. ONCE THIS WRITING OPERATION IS SUCCESSFULLY PERFORMED | 1
3697254 | WHEN DICOM IMAGES ARE VISUALIZED WITH AYCAN OSIRIX WHERE ONE DICOM FILE CONTAINS IMAGE DATA WITH AREAS OF DIFFERENT RESOLUTION VALUES (E.G. TWO AREAS OF DIFFERENT RESOLUTION VALUES - WHICH IS DIFFERENT TO THE 'TYPICAL' DICOM DATASET), WHILE DOING E.G. LENGTH DETERMINATION/CALCULATION THE SOFTWARE DOES NOT RECOGNIZE MORE THAN 1 RESOLUTION VALUES PER IMAGE, WHICH MIGHT RESULT IN INCORRECT LENGTH ETC. DETERMINATION/CALCULATION. 'TYPICAL' DICOM DATASETS USE ONLY ONE RESOLUTION PER IMAGE/FILE AND THE RESOLUTION IS CONSISTENT THROUGHOUT THE WHOLE SERIES. IMAGES WITH DIFFERENT RESOLUTIONS IN ONE FILE HAPPEN TO BE ULTRASOUND IMAGES WHERE, DURING TYPICAL WORKFLOW, MEASUREMENTS ARE USUALLY DONE DIRECTLY AT THE MODALITY DURING ACQUISITION.><CORRECTIVE ACTIONS: AN UPDATE FOR THE SOFTWARE AYCAN OSIRIX WAS CREATED WHICH COVERS THE PROBLEM. ALL USERS ARE INFORMED ABOUT THE PROBLEM AND THE UPDATE. (B)(4). | 1
4113544 | AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 7TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THE CUSTOMER, THE INITIAL REPORTER, IN THIS NEW EVENT CONTACTED AGFA ON FEBRUARY 14, 2014 AND REPORTED A READING PHYSICIAN HAD DISCOVERED AN INCORRECT STATEMENT IN THE PULMONARY ARTERY FINDINGS WITHIN THEIR RMAT REPORTING. THE CUSTOMER CONFIRMED THERE WERE NO REPORTS OF PATIENT HARM DURING THIS EVENT. THE CUSTOMER STATED NO RMAT CUSTOMIZATION REQUESTS HAD BEEN MADE TO AGFA, AND THE SITE'S SUPER USER HAD NOT MADE ANY CUSTOM CHANGES TO THE PULMONARY ARTERY FINDINGS WITHIN THE REPORTING TOOL, ABIDING BY AGFA'S RMAT TERMS OF USE. THE CUSTOMER DESCRIBED WHEN REPORTS WERE CREATED FOR TRANS ESOPHAGEAL ECHOCARDIOGRAM (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND STRESS (STR) STUDIES, AND WHEN THE END-USER SELECTED THE MILD TO MODERATE" RADIO BUTTON IN THE PULMONARY ARTERY PATHOLOGY SECTION | 1
4339739 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 16TH OCCURRENCE ARE 26 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 28 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINDING AND FINDING SENTENCE ARE NOT ANATOMICALLY CORRECT. POTENTIAL IMPACT: WHEN THE USER ATTEMPTS TO SELECT THAT THERE IS AKINESIS OF THE MID ANTERIOR WALL THE SENTENCE PRODUCED WILL DESCRIBE THE AKINESIS AS BEING ANTERIOR WALL, WHICH COULD GROSSLY OVERESTIMATE THE DISEASE STATE AND POSSIBLY LEAD TO INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | 1
4279592 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL MEASUREMENT LABEL WAS POST TRAB MUSCULAR TWO PEAK GRAD" - THIS WAS CHANGED TO "PDA PK GRAD". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
3958963 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SVC AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING ID=100033 THE SELECTION TO CREATE THE FINDING SENTENCE IN IMPAX CV REPORTING READS APPENDAGE VISUALISED." THE FINDING SENTENCE THAT IS CREATED READS "THE LEFT ATRIAL APPENDAGE IS FREE OF CLOTS IN MULTIPLE VIEWS AND NORMAL CONTRACTIONS." BECAUSE THE SELECTION PRESENTED TO THE USER DOES NOT INDICATE "FREE OF CLOTS" IT IS POSSIBLE THE CLINICAL USER CAN CREATE A FINDING SENTENCE THAT IS INACCURATE. (B)(4). NO PT HARM HAS BEEN REPORTED DUE TO THIS EVENT. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REF #: Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING.><AGFA SUBMITTED MDR REPORT #: 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLANTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN | 1
4112953 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# ISZ-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
3957828 | AGFA SUBJECTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 10TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING ID = 249 HAS BEEN CHANGED FROM THE AORTIC VALVE LEAFLETS ARE SEVERELY THICKENED WITH REDUCED SYSTOLIC EXCURSION." TO "THE AORTIC VALVE LEAFLETS ARE THICKENED WITH REDUCED SYSTOLIC EXCURSION." HISTORICALLY | 1
4285955 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL FINDING/SENTENCE/SHORT DESCRIPTION IN THE STUDY INDICATION WAS A FOLLOW-UP STUDY FOR SVT" - THIS WAS CHANGED TO A REGULAR INDICATION OF "BRADYCARDIA". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4113192 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. STUDY ID/S AND STUDY DATE FOR THIS MFR REPORT #: (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
3966457 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINING SENTENCE FOR FINDING ID = (B)(4). THIS ISSUE IS RELATED TO THE POTENTIAL FOR HISTORICAL REPORT CONTENT TO BE CHANGED IF RESIGNED OR PULLED FORWARD AS A TEMPLATE TO CREATE A NEW REPORT. THIS CONFIGURATION IN IMPAX CV REPORTING FOR THIS FINDING USED TO USE A MEAS_ID1025 WHICH IS A VELOCITY VALUE NOT A GRADIENT AS INDICATED IN THE FINDING SENTENCE. TODAY THE SENTENCE APPEARS WITHOUT A VALUE (THIS IS EVIDENT BECAUSE OF THE PARENTHESES ARE MISSING IN THE SENTENCE) MEANING THAT THIS ISSUE WOULD NOT BE CREATED IN A NEW REPORTED CREATED BY A CUSTOMER TODAY. HOWEVER, OLDER REPORTS WHICH WERE CREATED PRIOR TO THE REMOVAL OF THE MEAS_ID VALUE WOULD CONTAIN A MEASURE THAT IS NOT CONSISTENT WITH THE FINDING SENTENCE. (B)(4). NO PATIENT HARM HAS BEEN REPORTED DUE TO THIS EVENT. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA, FDA REFERENCE # IS Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING.><AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTIONS PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS (B)(6) OCCURRENCE ARE 15 DIFFERENT STUDY DATES, IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). | 1
4293149 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: GUI TYPE CHANGED WHICH COULD LEAD TO CONTRADICTORY STATEMENTS. BIO PROSTHETIC FUNCTION FINDING GROUP CHANGED FROM RADIO TO BOX, ALLOWING FOR CONTRADICTORY STATEMENTS TO APPEAR ON THE REPORT. REPORTS WITH INAPPROPRIATE USE OF GUI TYPE MITRAL VALVE PROSTHETIC MECHANICAL PROSTHETIC FUNCTION. POTENTIAL IMPACT: CONTRADICTORY STATEMENTS CAN NOW BE MADE BASED ON FINDINGS IN THE GROUP. REPORT PRESENTED DOWNSTREAM IS DIFFERENT FROM WHAT THE READING PHYSICIAN INTENDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6). | 1
4254381 | AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 12TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 12TH OCCURRENCE ARE 9 DIFFERENT STUDY DATES FOR MAY 2010, IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 48 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGED HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL SENTENCE/CONCLUSION STATEMENT REFERRED TO NORMAL VENTRICULAR OUTFLOW TRACTS" - THIS WAS CHANGED TO REFER TO "NORMAL PULMONARY VALVE | 1
4113560 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
3957811 | AGFA SUBJECTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 10TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING ID = (B)(4) HAS BEEN CHANGED FROM 'THE AORTIC VALVE LEAFLETS ARE SEVERELY THICKENED WITH REDUCED SYSTOLIC EXCURSION.' TO 'THE AORTIC VALVE LEAFLETS ARE THICKENED WITH REDUCED SYSTOLIC EXCURSION.' HISTORICALLY, THIS WILL CHANGE ANY AMENDED REPORT THAT PREVIOUSLY INCLUDED THE PHRASE 'SEVERELY THICKENED.' NO PATIENT HARM HAS BEEN REPORTED. CURRENTLY 59 AMENDED REPORTS INCLUDE THIS STATEMENT. THERE ARE 55 UNIQUE STUDY DATES WITHIN THE 59 REPORTS. A MDR WILL BE REPORTED FOR EACH OF THE STUDY DATES. THE MEDICAL RECORD NUMBER AND STUDY DATE FOR THIS MDR IS: (B)(4); STUDY DATE: (B)(4) 2012. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO FDA. FDA REFERENCE # IS Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. | 1
4279601 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE LABEL FOR THE FINDING READS PARTIAL ACUTE/CHRONIC DVT" | 1
4293400 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: THE ORIGINAL SENTENCE FINDING OF A MASS IS VISUALIZED AT THE APEX OF THE RIGHT VENTRICLE" HAS BEEN CHANGED TO: "THE PRESENCE OF A MASS IN THE RIGHT VENTRICLE CANNOT BE RULED OUT FROM THIS EXAM. IF CLINICALLY APPROPRIATE AN MRI COULD MORE THOROUGHLY EVALUATE THE TRANSTHORACIC ECHO FINDINGS. IN THIS ISSUE | 1
4293367 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE ORIGINAL SENTENCE FINDING OF THE RIGHT VENTRICULAR SYSTOLIC PRESSURE IS ESTIMATED TO BE 45-50MMHG" HAS BEEN CHANGED TO:" THE RIGHT VENTRICULAR SYSTOLIC PRESSURE IS ESTIMATED TO BE BETWEEN 90 AND 95 MMHG". IN THIS ISSUE | 1
4254342 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL SENTENCE/CONCLUSION STATEMENT REFERRED TO NORMAL INTACT AORTIC ARCH WITH NORMAL MAIN AND BRANCH PULMONARY ARTERIES.¿ - THIS WAS CHANGED TO REFER TO "NORMAL AORTIC ARCH WITH NO COARCTION OR DILATION.¿ POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4113613 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4293192 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: GUI TYPE CHANGED WHICH COULD LEAD TO CONTRADICTORY STATEMENTS. BIO PROSTHETIC FUNCTION FINDING GROUP CHANGED FROM RADIO TO BOX, ALLOWING FOR CONTRADICTORY STATEMENTS TO APPEAR ON THE REPORT. REPORTS WITH INAPPROPRIATE USE OF GUI TYPE SEPTAL NOTCH. POTENTIAL IMPACT: CONTRADICTORY STATEMENTS CAN NOW BE MADE BASED ON FINDINGS IN THE GROUP. REPORT PRESENTED DOWNSTREAM IS DIFFERENT FROM WHAT THE READING PHYSICIAN INTENDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | 1
4293184 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: GUI TYPE CHANGED WHICH COULD LEAD TO CONTRADICTORY STATEMENTS. BIO PROSTHETIC FUNCTION FINDING GROUP CHANGED FROM RADIO TO BOX, ALLOWING FOR CONTRADICTORY STATEMENTS TO APPEAR ON THE REPORT. REPORTS WITH INAPPROPRIATE USE OF GUI TYPE SEPTAL NOTCH. POTENTIAL IMPACT: CONTRADICTORY STATEMENTS CAN NOW BE MADE BASED ON FINDINGS IN THE GROUP. REPORT PRESENTED DOWNSTREAM IS DIFFERENT FROM WHAT THE READING PHYSICIAN INTENDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6). | 1
4113188 | AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 7TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THE CUSTOMER, THE INITIAL REPORTER, IN THIS NEW EVENT CONTACTED AGFA ON FEBRUARY 14, 2014 AND REPORTED A READING PHYSICIAN HAD DISCOVERED AN INCORRECT STATEMENT IN THE PULMONARY ARTERY FINDINGS WITHIN THEIR RMAT REPORTING. THE CUSTOMER CONFIRMED THERE WERE NO REPORTS OF PATIENT HARM DURING THIS EVENT. THE CUSTOMER STATED NO RMAT CUSTOMIZATION REQUESTS HAD BEEN MADE TO AGFA, AND THE SITE'S SUPER USER HAD NOT MADE ANY CUSTOM CHANGES TO THE PULMONARY ARTERY FINDINGS WITHIN THE REPORTING TOOL, ABIDING BY AGFA'S RMAT TERMS OF USE. THE CUSTOMER DESCRIBED WHEN REPORTS WERE CREATED FOR TRANS ESOPHAGEAL ECHOCARDIOGRAM (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND STRESS (STR) STUDIES, AND WHEN THE END-USER SELECTED THE MILD TO MODERATE" RADIO BUTTON IN THE PULMONARY ARTERY PATHOLOGY SECTION | 1
4285897 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL FINDING/SENTENCE/SHORT DESCRIPTION IN THE STUDY INDICATION WAS A FOLLOW-UP STUDY FOR AOV | 1
4113427 | AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 7TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THE CUSTOMER, THE INITIAL REPORTER, IN THIS NEW EVENT CONTACTED AGFA ON FEBRUARY 14, 2014 AND REPORTED A READING PHYSICIAN HAD DISCOVERED AN INCORRECT STATEMENT IN THE PULMONARY ARTERY FINDINGS WITHIN THEIR RMAT REPORTING. THE CUSTOMER CONFIRMED THERE WERE NO REPORTS OF PATIENT HARM DURING THIS EVENT. THE CUSTOMER STATED NO RMAT CUSTOMIZATION REQUESTS HAD BEEN MADE TO AGFA, AND THE SITE'S SUPER USER HAD NOT MADE ANY CUSTOM CHANGES TO THE PULMONARY ARTERY FINDINGS WITHIN THE REPORTING TOOL, ABIDING BY AGFA'S RMAT TERMS OF USE. THE CUSTOMER DESCRIBED WHEN REPORTS WERE CREATED FOR TRANS ESOPHAGEAL ECHOCARDIOGRAM (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND STRESS (STR) STUDIES, AND WHEN THE END-USER SELECTED THE MILD TO MODERATE" RADIO BUTTON IN THE PULMONARY ARTERY PATHOLOGY SECTION | 1
4279702 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE LABEL FOR THE FINDING READS COMPLETE ACUTE/CHRONIC DVT" | 1
4293456 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SVC AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: THE ORIGINAL SENTENCE FINDING OF POSTERIOR WALL HYPERTROPHY IS OBSERVED" HAS BEEN CHANGED TO: "THERE IS AN ASYMMETRIC INFERIOR WALL HYPERTROPHY OF THE LEFT VENTRICLE CONSISTENT WITH HYPERTROPHIC CARDIOMYOPATHY." IN THIS ISSUE | 1
4279440 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL MEASUREMENT LABEL WAS POST TRAB MUSCULAR TWO PEAK GRAD" - THIS WAS CHANGED TO "PDA PK GRAD". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4293369 | AN AGFA CLINICA ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE ORIGINAL SENTENCE FINDING OF THE RIGHT VENTRICULAR SYSTOLIC PRESSURE IS ESTIMATED TO BE 20-25 MMHG" HAS BENE CHANGED TO "THE RIGHT VENTRICULAR SYSTOLIC PRESSURE IS ESTIMATED TO BE BETWEEN 65 AND 70 MMHG." IN THIS ISSUE | 1
4113526 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4308488 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: INCORRECT MEASUREMENT METHOD BEING POPULATED IN FINDING SENTENCE: THE FINDING SENTENCE REFERRING TO AVA BY PLANIMETRY HAS BEEN CHANGED TO IMPORT THE AVA MEASURED BY VTI. POTENTIAL IMPACT: AN INCORRECT METHOD FOR OBTAINING TO AVA WILL BE IMPORTED, LEADING TO A POTENTIAL INCORRECT AREA BY VTI. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 13TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 13TH OCCURRENCE ARE 28 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 30 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. | 1
4254369 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 12TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 12TH OCCURRENCE ARE 3 DIFFERENT STUDY DATES FOR MAY 2010, IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 14 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL SENTENCE/CONCLUSION STATEMENT REFERRED TO NORMAL INTACT AORTIC ARCH WITH NORMAL MAIN AND BRANCH PULMONARY ARTERIES.¿ - THIS WAS CHANGED TO REFER TO "NORMAL AORTIC ARCH WITH NO COARCTION OR DILATION.¿ POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4339719 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINDING AND FINDING SENTENCE ARE NOT ANATOMICALLY CORRECT. POTENTIAL IMPACT: WHEN THE USER ATTEMPTS TO SELECT THAT THERE IS AKINESIS OF THE MID ANTERIOR WALL THE SENTENCE PRODUCED WILL DESCRIBE THE AKINESIS AS BEING ANTERIOR WALL, WHICH COULD GROSSLY OVERESTIMATE THE DISEASE STATE AND POSSIBLY LEAD TO INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 16TH OCCURRENCE ARE 26 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 28 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. | 1
4113128 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
3854036 | .><AGFA'S AWARENESS DATE IS MAY 14, 2014 FOR THIS EVENT. AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A CUSTOMER IN THE US. A 9TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). WITHIN THIS 9TH OCCURRENCE ARE 7 DIFFERENT STUDY DATES FROM YEARS 2011, 2012, AND 2013, IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: MEASUREMENT LABEL IS CHANGED ON FINDING SENTENCE THE PRESSURE HALF TIME OF THE MITRAL VALVES IS ( ) MSEC." TO "THE PRESSURE HALF TIME OF THE MITRAL VALVES IS ( ) MMHG." (ONE LABEL INDICATES TIME | 1
4279676 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE LABEL FOR THE FINDING READS COMPLETE ACUTE/CHRONIC DVT" | 1
4308579 | THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE U.S. A 13TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 13TH OCCURRENCE ARE 28 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 30 REPORTS CONTAIN THIS INCORRECT FINDING INFO. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MGR. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: INCORRECT MEASUREMENT METHOD BEING POPULATED IN FINDING SENTENCE: THE FINDING SENTENCE REFERRING TO AVA BY PLANIMETRY HAS BEEN CHANGED TO IMPORT THE AVA MEASURED BY VTI. POTENTIAL IMPACT: AN INCORRECT METHOD FOR OBTAINING THE AVA WILL BE IMPORTED, LEADING TO A POTENTIAL INCORRECT AREA BY VTI. | 1
4285939 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL FINDING/SENTENCE/SHORT DESCRIPTION IN THE STUDY INDICATION WAS A FOLLOW-UP STUDY FOR SVT" - THIS WAS CHANGED TO A REGULAR INDICATION OF "BRADYCARDIA". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4293439 | AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE UNITED STATES. A 14TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MGMT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AFGA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION ITO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: THE ORIGINAL SENTENCE FINDING OF A LEFT VENTRICULAR PSEUDOANEURYSM OF THE BASAL WALL IS OBSERVED". THIS WAS CHANGED TO: "THERE ARE ECHO FINDINGS TO SUGGEST THE PRESENCE OF THROMBUS IN THE PSEUDOANEURYSM". IN THIS ISSUE | 1
4113597 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4113552 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4308559 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE U.S. A 13TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 13TH OCCURRENCE ARE 28 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 30 REPORTS CONTAIN THIS INCORRECT FINDING INFO. AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MGR. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: INCORRECT MEASUREMENT METHOD BEING POPULATED IN FINDING SENTENCE: THE FINDING SENTENCE REFERRING TO AVA BY PLANIMETRY HAS BEEN CHANGED TO IMPORT THE AVA MEASURED BY VTI. POTENTIAL IMPACT: AN INCORRECT METHOD FOR OBTAINING THE AVA WILL BE IMPORTED, LEADING TO A POTENTIAL INCORRECT AREA BY VTI.><THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | 1
4113627 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4)><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4293175 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: GUI TYPE CHANGED WHICH COULD LEAD TO CONTRADICTORY STATEMENTS. BIO PROSTHETIC FUNCTION FINDING GROUP CHANGED FROM RADIO TO BOX, ALLOWING FOR CONTRADICTORY STATEMENTS TO APPEAR ON THE REPORT. REPORTS WITH INAPPROPRIATE USE OF GUI TYPE MITRAL VALVE PROSTHETIC BIOPROSTHETIC FUNCTION. POTENTIAL IMPACT: CONTRADICTORY STATEMENTS CAN NOW BE MADE BASED ON FINDINGS IN THE GROUP. REPORT PRESENTED DOWNSTREAM IS DIFFERENT FROM WHAT THE READING PHYSICIAN INTENDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | 1
4113591 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
3957834 | AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US, A (B)(4) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING (B)(4) HAS BEEN CHANGED FROM THE AORTIC VALVE LEAFLETS ARE SEVERELY THICKENED WITH REDUCED SYSTOLIC EXCURSION." TO "THE AORTIC VALVE LEAFLETS ARE THICKENED WITH REDUCED SYSTOLIC EXCURSION." HISTORICALLY | 1
4113159 | (B)(4). EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING.><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4279447 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL MEASUREMENT LABEL WAS POST TRAB MUSCULAR TWO PEAK GRAD" - THIS WAS CHANGED TO "PDA PK GRAD". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4339691 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 16TH OCCURRENCE ARE 26 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 28 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINDING AND FINDING SENTENCE ARE NOT ANATOMICALLY CORRECT. POTENTIAL IMPACT: WHEN THE USER ATTEMPTS TO SELECT THAT THERE IS AKINESIS OF THE MID ANTERIOR WALL THE SENTENCE PRODUCED WILL DESCRIBE THE AKINESIS AS BEING ANTERIOR WALL, WHICH COULD GROSSLY OVERESTIMATE THE DISEASE STATE AND POSSIBLY LEAD TO INCORRECT TREATMENT. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | 1
3957836 | AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA FOR A SITE IN THE US, A 10TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN (B)(4).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING ID (B)(6) HAS BEEN CHANGED FROM THE AORTIC VALVE LEAFLETS ARE SEVERELY THICKENED WITH REDUCED SYSTOLIC EXCURSION." TO "THE AORTIC VALVE LEAFLETS ARE THICKENED WITH REDUCED SYSTOLIC EXCURSION." HISTORICALLY | 1
4113539 | AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 7TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THE CUSTOMER, THE INITIAL REPORTER, IN THIS NEW EVENT CONTACTED AGFA ON FEBRUARY 14, 2014 AND REPORTED A READING PHYSICIAN HAD DISCOVERED AN INCORRECT STATEMENT IN THE PULMONARY ARTERY FINDINGS WITHIN THEIR RMAT REPORTING. THE CUSTOMER CONFIRMED THERE WERE NO REPORTS OF PATIENT HARM DURING THIS EVENT. THE CUSTOMER STATED NO RMAT CUSTOMIZATION REQUESTS HAD BEEN MADE TO AGFA, AND THE SITE'S SUPER USER HAD NOT MADE ANY CUSTOM CHANGES TO THE PULMONARY ARTERY FINDINGS WITHIN THE REPORTING TOOL, ABIDING BY AGFA'S RMAT TERMS OF USE. THE CUSTOMER DESCRIBED WHEN REPORTS WERE CREATED FOR TRANS ESOPHAGEAL ECHOCARDIOGRAM (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND STRESS (STR) STUDIES, AND WHEN THE END-USER SELECTED THE MILD TO MODERATE" RADIO BUTTON IN THE PULMONARY ARTERY PATHOLOGY SECTION | 1
4278920 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL MEASUREMENT LABEL WAS POST TRAB MUSCULAR TWO PEAK GRAD" - THIS WAS CHANGED TO "PDA PK GRAD". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4279698 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: THE LABEL FOR THE FINDING READS PARTIAL ACUTE/CHRONIC DVT" | 1
4308528 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 13TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 13TH OCCURRENCE ARE 28 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 30 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: INCORRECT MEASUREMENT METHOD BEING POPULATED IN FINDING SENTENCE: THE FINDING SENTENCE REFERRING TO AVA BY PLANIMETRY HAS BEEN CHANGED TO IMPORT THE AVA MEASURED BY VTI. POTENTIAL IMPACT: AN INCORRECT METHOD FOR OBTAINING TO AVA WILL BE IMPORTED, LEADING TO A POTENTIAL INCORRECT AREA BY VTI. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6). | 1
3957777 | AGFA SUBMITTED MDR REPORT #: 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US. A (B)(4) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLANTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10.><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SVC AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PT REPORTS. SPECIFICALLY: FINDING SENTENCE FOR FINDING (B)(4) HAS BEEN CHANGED FROM THE AORTIC VALVE LEAFLETS ARE SEVERELY THICKENED WITH REDUCED SYSTOLIC EXCURSION." TO "THE AORTIC VALVE LEAFLETS ARE THICKENED WITH REDUCED SYSTOLIC EXCURSION." HISTORICALLY | 1
4293172 | AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. (B)(6).><AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: GUI TYPE CHANGED WHICH COULD LEAD TO CONTRADICTORY STATEMENTS. BIO PROSTHETIC FUNCTION FINDING GROUP CHANGED FROM RADIO TO BOX, ALLOWING FOR CONTRADICTORY STATEMENTS TO APPEAR ON THE REPORT. REPORTS WITH INAPPROPRIATE USE OF GUI TYPE SEPTAL NOTCH. POTENTIAL IMPACT: CONTRADICTORY STATEMENTS CAN NOW BE MADE BASED ON FINDINGS IN THE GROUP. REPORT PRESENTED DOWNSTREAM IS DIFFERENT FROM WHAT THE READING PHYSICIAN INTENDED. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4). | 1
4308526 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: INCORRECT MEASUREMENT METHOD BEING POPULATED IN FINDING SENTENCE: THE FINDING SENTENCE REFERRING TO AVA BY PLANIMETRY HAS BEEN CHANGED TO IMPORT THE AVA MEASURED BY VTI. POTENTIAL IMPACT: AN INCORRECT METHOD FOR OBTAINING TO AVA WILL BE IMPORTED, LEADING TO A POTENTIAL INCORRECT AREA BY VTI. THERE HAS BEEN NO REPORTED PATIENT HARM FOR THIS OCCURRENCE. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA VIA REFERENCE #Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(6).><AGFA SUBMITTED MDR REPORT #1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 13TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTED MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTION PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS 13TH OCCURRENCE ARE 28 DIFFERENT STUDY DATES IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). CURRENTLY, 30 REPORTS CONTAIN THIS INCORRECT FINDING INFORMATION. | 1
4285935 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL FINDING/SENTENCE/SHORT DESCRIPTION IN THE STUDY INDICATION WAS RULE OUT CONGENITAL HEART DISEASE" - THIS WAS CHANGE TO "PULMONARY HYPERTENSION". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
4279403 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATION HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: ORIGINAL MEASUREMENT LABEL WAS POST TRAB MUSCULAR TWO PEAK GRAD" - THIS WAS CHANGED TO "PDA PK GRAD". POTENTIAL IMPACT: IF THE CHANGE WAS MADE AFTER REPORTS HAD BEEN GENERATED WITH THE ORIGINAL SENTENCE SELECTED | 1
3966473 | AN AGFA CLINICAL ANALYST PERFORMED A RETRO-ANALYSIS AND REPORTED THE FINDINGS TO AGFA SERVICE AND AGFA PRODUCT QUALITY MANAGER. AGFA'S INVESTIGATION INTO THIS OCCURRENCE OF RMAT CUSTOMIZATIONS HAS REVEALED THAT THIS SPECIFIC CHANGE HAD THE POTENTIAL TO INTRODUCE CLINICAL INACCURACIES IN PATIENT REPORTS. SPECIFICALLY: FINING SENTENCE FOR FINDING ID = (B)(4). THIS ISSUE IS RELATED TO THE POTENTIAL FOR HISTORICAL REPORT CONTENT TO BE CHANGED IF RESIGNED OR PULLED FORWARD AS A TEMPLATE TO CREATE A NEW REPORT. THIS CONFIGURATION IN IMPAX CV REPORTING FOR THIS FINDING USED TO USE A MEAS_ID1025 WHICH IS A VELOCITY VALUE NOT A GRADIENT AS INDICATED IN THE FINDING SENTENCE. TODAY THE SENTENCE APPEARS WITHOUT A VALUE (THIS IS EVIDENT BECAUSE OF THE PARENTHESES ARE MISSING IN THE SENTENCE) MEANING THAT THIS ISSUE WOULD NOT BE CREATED IN A NEW REPORTED CREATED BY A CUSTOMER TODAY. HOWEVER, OLDER REPORTS WHICH WERE CREATED PRIOR TO THE REMOVAL OF THE MEAS_ID VALUE WOULD CONTAIN A MEASURE THAT IS NOT CONSISTENT WITH THE FINDING SENTENCE. (B)(4). NO PATIENT HARM HAS BEN REPORTED DUE TO THIS EVENT. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA, FDA REFERENCE # IS Z-2112-10. AGFA WILL FOLLOW THE RMAT POST MARKET VERIFICATION WORK INSTRUCTIONS TO CORRECT THE SENTENCE FINDING. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING.><AGFA SUBMITTED MDR REPORT # 1225058-2010-00001 TO THE FDA ON 06/07/2010 FOR A SITE IN THE US. A (B)(6) OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THIS IS AN INTERNAL DISCOVERY DETERMINED DURING THE IMPLEMENTATION OF THE ASSOCIATED PROBLEM CORRECTIONS PLAN, RMAT VERIFICATION, AS REPORTED IN FDA Z-2112-10. WITHIN THIS (B)(6) OCCURRENCE ARE 15 DIFFERENT STUDY DATES, IN WHICH AN INDIVIDUAL MDR REPORT WILL BE SUBMITTED FOR EACH ASSOCIATED STUDY DATE AND MEDICAL RECORD NUMBER (MRN). | 1
4113545 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA IS REPORTING A MEDICAL DEVICE REPORT FOR EACH OF THE 226 STUDY DATES. AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4113085 | EVAL SUMMARY: AGFA IS CURRENTLY REVIEWING THE RMAT LOGS TO IDENTIFY WHEN THE CUSTOMIZATION CHANGE WAS INTRODUCED INTO THE PRODUCTION ENVIRONMENT. (B)(4). AGFA'S CLINICAL ANALYST ALSO DETERMINED FROM A STUDY SELECTED FROM THE POTENTIALLY AFFECTED RMAT LOG, THAT THE INCORRECT SENTENCE FINDING LINE WAS LOCATED DIRECTLY ABOVE THE ACTUAL MEASUREMENTS (WHICH ARE CORRECT) AND THAT A USER WOULD BE MORE LIKELY TO LOOK AT THE ACTUAL MEASUREMENTS IN A REPORT TO FORMULATE CONCLUSIONS AND TREATMENT DECISIONS; AN INCLINATION AS TO WHY THE READING PHYSICIANS MISSED THE INCORRECT SENTENCE FINDING OF HYPOTENSION" VERSUS "HYPERTENSION " IN THE MANY PREVIOUS PULMONARY ARTERY PATHOLOGY REPORTS. A REPORTABLE CORRECTION IS UNDERWAY FOR THIS ISSUE AND HAS BEEN REPORTED TO THE FDA. FDA REFERENCE# IS Z-2112-10. ANY FURTHER INVESTIGATION FOR THE SITE DESCRIBED IN THIS REPORT WILL BE DOCUMENTED IN THE ONGOING CFR PART 806 REPORTING. (B)(4).><AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7 | 1
4113057 | AGFA SUBMITTED MDR REPORT# 1225058-2010-00001 TO THE FDA ON JUNE 7, 2010 FOR A SITE IN THE US. A 7TH OCCURRENCE IS BEING REPORTED FOR THE SAME ISSUE/SAME DEVICE: IMPAX CV RESULTS MANAGEMENT ADMINISTRATION TOOL (RMAT). THE CUSTOMER, THE INITIAL REPORTER, IN THIS NEW EVENT CONTACTED AGFA ON FEBRUARY 14, 2014 AND REPORTED A READING PHYSICIAN HAD DISCOVERED AN INCORRECT STATEMENT IN THE PULMONARY ARTERY FINDINGS WITHIN THEIR RMAT REPORTING. THE CUSTOMER CONFIRMED THERE WERE NO REPORTS OF PATIENT HARM DURING THIS EVENT. THE CUSTOMER STATED NO RMAT CUSTOMIZATION REQUESTS HAD BEEN MADE TO AGFA, AND THE SITE'S SUPER USER HAD NOT MADE ANY CUSTOM CHANGES TO THE PULMONARY ARTERY FINDINGS WITHIN THE REPORTING TOOL, ABIDING BY AGFA'S RMAT TERMS OF USE. THE CUSTOMER DESCRIBED WHEN REPORTS WERE CREATED FOR TRANS ESOPHAGEAL ECHOCARDIOGRAM (TEE), TRANSTHORACIC ECHOCARDIOGRAM (TTE) AND STRESS (STR) STUDIES, AND WHEN THE END-USER SELECTED THE MILD TO MODERATE" RADIO BUTTON IN THE PULMONARY ARTERY PATHOLOGY SECTION | 1
4155289 | MCKESSON INVESTIGATED THE ISSUE AND CONCLUDED THAT (B)(4) DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT HARM. WHEN IMAGES ARE INITIALLY ACQUIRED FROM A MODALITY AND RECEIVED BY (B)(4), THE STUDY STATUS APPEARING IN THE WORK LIST IS SET TO PERFORMED". A TECHNOLOGIST REVIEWS THE STUDY AND CHANGES THE STATUS TO "REVIEWED". REVIEWED STUDIES WILL APPEAR ON THE RADIOLOGIST STUDY LIST AND IS AVAILABLE FOR INTERPRETATION. IN THE REPORTED CASE | 1
3815398 | THE REPORTING FACILITY INDICATED THAT THE HIGH PRIORITY ALERT ICON DID NOT DISPLAY ON THE MCKESSON RADIOLOGY STUDY LIST, AS EXPECTED. NO HARM TO THE PATIENT WAS REPORTED.><THE REPORTING FACILITY INDICATED THAT THEY USE THE PRIORITY ALERT ICON FOR 'HOLD AND CALL' SITUATIONS, WHEREBY A PATIENT MUST REMAIN AT THE HEALTHCARE FACILITY FOR IMMEDIATE TREATMENT, OR UNTIL A PHYSICIAN HAS REVIEWED THE STUDY. MCKESSON MEDICAL IMAGING COMPANY COMPLETED AN INVESTIGATION AND CONCLUDED THAT THE SYSTEM MALFUNCTIONED. THE HIGH PRIORITY ALERT ICON IS ONLY DISPLAYED WHEN THERE IS ALSO A STAT STUDY PRESENT IN THE MCKESSON RADIOLOGY STUDY LIST. WHEN THE HIGH PRIORITY ALERT ICON IS NOT DISPLAYED AS INTENDED, THERE IS POTENTIAL FOR A DELAY IN CARE. MCKESSON MEDICAL IMAGING COMPANY CUSTOMER SUPPORT WILL WORK WITH POTENTIALLY AFFECTED CUSTOMERS TO APPLY A SOFTWARE UPDATE TO MCKESSON RADIOLOGY 12.0 FOR CORRECTING THE ERROR IDENTIFIED. | 1
4152308 | A SERVICE CALL WAS PERFORMED AND THE CYNOSURE SERVICE ENGINEER FOUND THAT THE SHUTTER ON THE HANDPIECE HAD A LOOSE SCREW. THE SCREW WAS TIGHTENED AND THE LASER PERFORMED TO SPECIFICATION. THIS IS AN ISOLATED INCIDENT.><LASER FIRING ON ITS OWN. NO PATIENTS WERE INJURED. | 1
4046808 | A DATABASE SEARCH WAS PERFORMED, REVIEWED BY A CLINICAL HEALTHCARE PROFESSIONAL AND ENGINEERS AT PHILIPS. THE DATABASE SEARCH CONFIRMED THAT ON (B)(6) 2014, RN CALLED THE PHILIPS CUSTOMER SERVICE PERSONNEL WITH THE ISSUE OF WHEN THEY USED SEQUENCE FOR THEIR PULLBACK | 1
4173115 | ><CLINICIANS HAVE NOTED THAT THEY HAVE WITNESSED CONTINUOUS REBOOTING OF THE DASH MONITORS. DURING OUR INVESTIGATION IT WAS NOTED THAT THE NURSES HAVE REPORTED THIS PROBLEM DURING A PATIENT'S ADMISSION OR DISCHARGE OR WHEN THEY ARE RESTARTING THE MONITOR. THERE HAVE BEEN SEVERAL REPORTS IN THE FALL OF 2013. | 1
3817772 | E-MAIL ADDRESS AND OCCUPATION OF REPORTER IS UNK. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><THE CUSTOMER REPORTS THERE WAS WHITE SMOKE COMING FROM THE PROCARE MONITOR B20. THERE WAS NO PT INVOLVEMENT AS THE SMOKE WAS NOTICED PRIOR TO THE PROCEDURE. | 1
3908688 | THE HOSPITAL REPORTED A NURSE TURNED ON THE MONITOR IN PREPARATION OF PT ARRIVAL, OBSERVED SMOKE COMING FROM THE BACK OF THE UNIT, IMMEDIATELY TURNED UNIT OFF AND UNPLUGGED IT. A PT WAS NOT CONNECTED TO THE MONITOR AT THE TIME.><THE PRODUCT WAS SHIPPED FROM THE MANUFACTURING SITE 02/04/2009. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
4351348 | PATIENT INFORMATION IS NOT AVAILABLE. REPORTER'S OCCUPATION IS UNKNOWN. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE.><INVESTIGATION OF THE ECG WAVEFORMS FROM SINGLE LEAD (12-LEAD WAVEFORMS NOT AVAILABLE) WAS CONDUCTED. CONCLUSION WAS THAT THERE ARE HEART RATE (HR) INCREASES DUE TO ARTIFACTS, HOWEVER, THE HR INCREASES ARE INTERMITTENT AND SOME OF THE HR INCREASES WERE SEEN ALSO IN THE PLETH WAVEFORM SUGGESTING PERFUSION. THE B850 LOGS FROM (B)(6) 2014 19:01:18 TO (B)(6) 2014 18:32:32 SHOW EIGHT (8) HR HIGH ( TACHY) ALARMS WITH LONGEST DURATION OF 6 SECONDS WITH THE HR HIGH ALARM LIMIT OF 200 /MIN. THIS INDICATES THAT THE HR WAS BELOW ALARM LIMIT MOST OF THE TIME. THE INVOLVED PDM MODULE WAS TESTED FOR LINEARITY, DYNAMIC RANGE, AND NOISE PERFORMANCE, PASSED ALL TESTS AND NO ISSUES WERE FOUND. THE INVOLVED ECG CABLE WAS REQUESTED FOR INVESTIGATION BUT WAS NOT AVAILABLE. GE ENGINEERS VISITED THE REPORTING HOSPITAL. THEY WERE ABLE TO DEMONSTRATE THAT THE ECG ELECTRODES WERE CONTRIBUTING TO THE ARTIFACT SEEN IN THE ECG WAVEFORM. AFTER THE CUSTOMER CHANGED TO A DIFFERENT MANUFACTURED ECG ELECTRODE, THE NOISE ARTIFACT NO LONGER EXISTED AND FALSE ARRHYTHMIA CALLS STOPPED. NO FURTHER COMPLAINTS HAVE BEEN REPORTED OF THIS ISSUE FROM THE CUSTOMER SINCE THE CHANGE OF THE ECG ELECTRODES. THE INVESTIGATION CONCLUDED THAT THE ISSUE WAS CAUSED BY THE COMBINATION OF HIGH SKIN-TO-ELECTRODE IMPEDANCE IN CONJUNCTION WITH AN EXTERNAL SOURCE OF ELECTRICAL INTERFERENCE. THE ECG MEASUREMENT CIRCUITRY WILL BE MORE SENSITIVE TO EXTERNAL ENVIRONMENTAL NOISE IF THE ECG ELECTRODE IMPEDANCE IS HIGH. THE B850 USER'S MANUAL COVERS SKIN PREPARATION TECHNIQUES: SHAVE ANY HAIR FROM THE ELECTRODE SITE; GENTLY RUB THE SURFACE OF THE SKIN TO INCREASE CAPILLARY BLOOD FLOW; CLEAN THE SKIN WITH ALCOHOL OR MILD SOAP AND WATER SOLUTION TO REMOVE SKIN OIL AND DEAD OR ABRADED SKIN CELLS; DRY THE SKIN COMPLETELY BEFORE APPLYING THE ELECTRODES. IT WAS OBSERVED THAT THE SECTION ON SKIN PREPARATION WAS NOT FOLLOWED BY THE CUSTOMER. THE B850 USER'S MANUAL (2062971-001) COVERS TROUBLESHOOTING TECHNIQUES. FOR NOISY ECG SIGNAL THE ECG TROUBLESHOOTING GIVES THE FOLLOWING SOLUTION: CHECK THE ELECTRODE QUALITY AND POSITIONING. DO NOT PLACE ELECTRODES ON BODY HAIR | 1
3492254 | THE RECORD REPORTS A LOSS OF THE DEFIC SYNC FUNCTION ON THE PT DATA MODULE. THE ISSUE WAS FOUND DURING SET-UP SO THERE WAS NO PT INVOLVEMENT.><PT DATA NOT CURRENTLY AVAILABLE. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3624217 | LIMITED INFORMATION PROVIDED. THE CUSTOMER REPORTS A DINAMAP PRO 100 SPONTANEOUSLY BURNED UP FROM THE BOTTOM. THERE WAS NO REPORTED PATIENT INJURY ASSOCIATED WITH THIS EVENT.><A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3994192 | PHILIPS FOOTBALL" NON INVASIVE BLOODPRESSURE MONITOR MALFUNCTIONED AND WOULD NOT OPERATE. EQUIPMENT PULLED FROM SERVICE AND REPLACED." | 1
3588735 | ><THE SAGE COMFORT ESSENTIAL BATH WIPES LEAK. THE BAG OPENS WHEN IN THE WARMERS AND THE SOAP DRAINS ONTO THE WARMING UNITS. | 1
3894327 | IT WAS REPORTED THAT, APPROXIMATELY 2 HOURS AFTER THE LAST CONTROL OF THE DEVICE FUNCTION, THE STAFF NOTICED THE DEATH OF THE PT. THE MONITOR WAS IN OPERATION WITH ALL CABLES CONNECTED BUT NO ALARMS WERE PRESENT AT THE CENTRAL STATION IN THE PERIOD BEFORE.><THE INVESTIGATION HAS JUST STARTED, RESULTS WILL BE PROVIDED IN A FOLLOW-UP/FINAL REPORT. | 1
3537676 | THE REPORTED EVENT HAPPENED DURING A LAB TEST W/O PT INVOLVEMENT. DRAGER WAS ABLE TO REPRODUCE THE DESCRIBED BEHAVIOR OF THE COCKPIT MONITOR. BUT THE PT NEAR M540 MONITOR GENERATES THE RED ALARM. FINAL INVESTIGATION RESULTS WILL BE REPORTED IN A F/U REPORT.><IT WAS REPORTED THAT: A YELLOW SPO2-LOW WAS DISPLAYED AND THE CORRESPONDING ALARM TONE IS GENERATED ON THE M540 AND THE COCKPIT. THE ALARM WAS SILENCED FOR 2 MINS VIA THE AUDIO PAUSE. IF THEREAFTER SPO2 DROPS BELOW THE SECOND ALARM LIMIT FOR CRITICAL SPO2 ALARM, THE LIFE THREATENING ALARM WAS NOT SHOWN ON THE COCKPIT AND NO ALARM TONE WAS GENERATED. | 1
4180239 | STAFF REPORT: PT IN BED 13 IN AFIB, NO ALARM OR MESSAGE ON MONITOR. I DEFAULTED THE MONITOR ALARMS, CHECKED AND DOUBLE CHECKED THE ALARMS WITH NO SUCCESS. I CALLED BIOMED, I DON'T KNOW IF ANYTHING WAS DONE BUT THE ALARM MESSAGE" WAS BACK LATER THAT DAY.THE NEXT DAY MY PT (BED 13) WENT FROM ST TO AF | 1
3961445 | DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED.><IT WAS REPORTED THAT 2 M300S APPEARED TO DISPLAY INACCURATE PT WAVEFORMS AT THE INFINITY CENTRAL STATION (ICS) AND AT THE M300A. IT WAS REPORTED THAT AT APPROX 7:00AM ON (B)(6) 2014, THE M300 FROM BED 61 DISPLAYED WHAT APPEARED TO BE A VTACH RHYTHM AND AT APPROX 10:00AM ON THE SAME DAY, THE M300 FROM BED 62 ALSO BEGAN TO DISPLAY WHAT APPEARED TO BE A VTACH RHYTHM. IT WAS REPORTED THAT A NURSE RESPONDED TO THE PTS AND A CARDIZEM DRIP WAS STARTED ALTHOUGH THE NURSE REPORTED THAT IN EACH CASE, THE PT DID NOT SHOW SYMPTOMS OF IRREGULAR RHYTHM AND THE NURSE DID NOT DETECT THE DISPLAYED RHYTHM WHEN LISTENING TO THE PT'S HEART. THERE WERE NO INJURIES REPORTED. (B)(4). | 1
3853647 | IT WAS REPORTED THAT DOCTOR (B)(6) FROM (B)(6) HOSPITAL-MEDINAH NICU DEPARTMENT, WHO OBSERVED THIS CASE, SAID THAT HE HAD PATIENT CONNECTED WITH INFINITY VISTA XL MONITOR (SPO2 SENSOR ONLY). THE PATIENT WAS DEAD, BUT THE O2 SATURATION STILL APPEARED ON THE MONITOR. DRAEGER REFERENCE NUMBER: (B)(4).><DRAEGER IS STILL INVESTIGATION THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
3600599 | DRAEGER IS STILL INVESTIGATION THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED.><IT WAS REPORTED THAT ON (B)(6) 2014, AT 7:55 AM, PATIENT WAS FOUND IN CARDIAC ARREST USING THE M300. THE NURSE IN THE PATIENT ROOM NOTICED THE M300 V LEAD WAS BETWEEN TWO FINGERS. THE PATIENT'S HEART RATE ON THE M300 WAS NOTED TO BE AT 60/MINUTE BUT IN CARDIAC ARREST. THEY COMPLAIN THE INTENSIVE CARE UNIT NEVER CALLED THEM TO GET NOTIFIED OF THE LEAD DISCONNECT. THERE WAS A PATIENT DEATH REPORTED. (B)(4). | 1
3770400 | DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED.><IT WAS REPORTED THAT DURING MONITORING OF A PATIENT AT BED SITE 1 AT THE INTENSIVE CARE UNIT (ICU) OF THE HOSPITAL, THE FOLLOWING SITUATION OCCURRED ON (B)(6) 2014 AT 9 PM. THE MONITOR MEASURE THE LAST SPO2 VALUE AT ABOUT 8:59 PM. THIS ALSO WAS SAVED AT THE MVWS (MULTIVIEW WORKSTATION). ALSO, THE CENTRAL STATION SAVED THAT A TECHNICAL PROBLEM OCCURRED AND MARKED THIS WITH A GREY NUMERIC FIELD AND 3 STARS (***). AT 9PM AT NIBP MEASUREMENT WAS STARTED. DURING THIS THE VALUE OF SPO2 WAS NOT DISPLAYED. UP FROM THIS TIME THE MVWS DID NOT SAVE ANY SPO2 VALUES ANYMORE. ONLY 3 STARS WERE DISPLAYED AND SAVED. ONE CAN PRINT OUT FROM THE LONG-TERM RECORD THAT THE SPO2 VALUE IS CONSTANTLY DECREASING. IT WAS FURTHER REPORTED THAT NEITHER THE MONITOR NOR THE MVWS GENERATED AN ALARM; THE HEART FREQUENCY STAYED THE SAME AT 48 PACES. IT WAS REPORTED THAT THE PATIENT HAS A PACEMAKER ACCORDING TO STAFF. IT IW NOT KNOWN IF THE PACER DETECTION WAS NO AT THIS TIME. AT 9:15 PM THE PATIENT WAS RESUSCITATED. DRAEGER REFERENCE NUMBER: (B)(4). | 1
3538982 | IT WAS REPORTED THAT SEVERAL GEN 1 ICS (INFINITY CENTRAL STATION) THAT HAD BEEN IN STORAGE AT THE HOSPITAL WERE RECENTLY INSTALLED TO REPLACE MVWS (MULTIVIEW WORK STATION) ON SEVERAL FLOORS INCLUDING SEVERAL TELEMETRY UNITS. THE ICS WERE LOADED WITH VF8.11.1 SOFTWARE AND SEVERAL ARE USED TO MONITOR TRUST TELEMETRY PATIENTS. THE USERS HAVE REPORTED THAT THE ICS THAT ARE MONITORING TRUST TELEMETRY PATIENTS WILL INTERMITTENTLY RESET. THE USER REPORTED THAT AFTER A RESET, IT TAKES SEVERAL MINUTES FOR CENTRAL STATION MONITORING TO BE RESTORED. THERE HAVE BEEN NO PT INJURIES REPORTED ASSOCIATED WITH THESE EVENTS. (B)(4).><DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A FOLLOW-UP REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
3418415 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF MAIN CPU BOARD AND PERFORMANCE OF OPERATIONAL AND SAFETY INSPECTIONS. | 1
4090075 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENTS OF THE UNIT'S MOTHER BOARD. UNIT WAS SAFETY TESTED TO FACTORY'S SPECS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. | 1
3418440 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF POWER SUPPLY PC BOARD. THE UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR PLUS MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
4304397 | CUSTOMER REPORTED AN ISSUE WITH THE V-SERIES MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND COULD NOT DUPLICATE THE REPORTED PROBLEM. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4234381 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S ENCODER. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4234275 | DEVICE WAS SHIPPED TO (B)(4) CENTER AND FOUND TO HAVE A BURNT CHIP ON THE DISPLAY BOARD. CHIP WAS REPLACED AND FUNCTIONALLY TESTED.><CUSTOMER REPORTED PASSPORT PATIENT MONITOR SMOKED. NO PATIENT INJURY WAS REPORTED. | 1
4006042 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNITS O2 MODULE. UNIT WAS CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PM9000 EXPRESS MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3418383 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND REPLACED THE CPU. UNIT WAS TESTED AND CALIBRATED. | 1
4219005 | UNIT HAS BEEN RETURNED TO THE COMPANY'S REPAIR CENTER FOR EVALUATION.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 1
4090074 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED.><UNIT HAS BEEN RETURNED TO THE COMPANY'S REPAIR CENTER FOR FURTHER EVAL. | 1
3515229 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 5 PATIENT MONITOR'S GAS MODULE WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><UNIT WILL BE RETURNED FOR EVALUATION. | 1
4219043 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD COMPONENTS. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4218977 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR,WHICH MAY HAVE AFFECTED SPO2 AND CO2 MONITORING. NO PATIENT INJURY WAS REPORT.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND DUPLICATED THE PROBLEM. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S SPO2 BOARD, CO2 MODULE AND COOLING FAN. UNIT WS TESTED TO FACTORY'S SPECIFICATIONS. | 1
3917925 | UNIT IS BEING RETURNED TO THE COMPANY'S REPAIR CENTER FOR FURTHER EVALUATION.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3418388 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF POWER PIN BOARD.><CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 1
3515241 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE POWER PIN BOARD KIT.><CUSTOMER REPORTED AN ISSUE WITH THE V SERIES MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 1
3418270 | CUSTOMER REPORTED AN ISSUE WITH THE ACCUTORR V MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT AND REPAIRED CONNECTOR ON SPO2 BOARD. THE UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
4144720 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND VERIFIED THE PROBLEM. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S DISPLAY ASSEMBLY. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4219026 | CUSTOMER REPORTED THE DPM 6 MONITOR HAD A MISSED VENTRICULAR FIBRILLATION ALARM, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE SYSTEM AND IDENTIFIED AN ISSUE WITH THE SYSTEM'S ALARM SETTING THAT WAS TURNED OFF. THE ALARM SETTING WAS TURNED ON AND THE PROBLEM WAS RESOLVED. | 1
3418259 | DOCUMENTATION FROM THE REPORTED EVENT IS CURRENTLY UNDER REVIEW.><CUSTOMER REPORTED THAT THE PANORAMA CENTRAL STATION DID NOT FOR A CARDIAC ARRHYTHMIA. NO PT INJURY WAS REPORTED. | 1
3795403 | COMPANY REPRESENTATIVE EVALUATED MONITOR AND WAS UNABLE TO REPRODUCE THE REPORT ISSUE.><CUSTOMER REPORTED ISSUE WITH THE PASSPORT 2 MONITOR WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3832035 | COMPANY REP EVALUATED THE UNIT AND REPLACED THE CPU BOARD. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PM-8000 EXPRESS MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
4144737 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND VERIFIED THE PROBLEM. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
3418433 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR'S DISPLAY, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENTS OF THE UNIT'S DISPLAY AND POWER SUPPLY. THE UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS. | 1
4219035 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR SHUT DOWN, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD. | 1
4006018 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6/7 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE SYSTEM. CORRECTIONS INCLUDED REPAIRING THE COLD SOLDER ON ALL FRONT PANEL CONNECTORS. UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTOR'S SPECS. | 1
3515262 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S MAIN CONTROL BOARD. UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3745046 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE SPO2 BOARD. UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
3418443 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE REPAIRED COLD SOLDER AT ALL CONNECTORS ON FRONT PANEL BOARD. THE UNIT WAS CALIBRATED, PERFORMANCE AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
3795407 | THE UNIT AND BATTERY WERE REMOVED FROM THE FACILITY AND IS CURRENTLY BEING ANALYZED BY THE MANUFACTURER.><THE CUSTOMER REPORTED THE PASSPORT 2 MONITOR'S BATTERY OVERHEATED CAUSING SMOKE AND TRIGGERED FIRE ALARMS. THE UNIT WAS NOT IN USE ON A PATIENT DURING THE TIME THE EVENT OCCURRED. NO INJURY WAS REPORTED. | 1
3786620 | COMPANY REP EVALUATED THE UNIT AND REPLACED THE UNIT'S SPO2 BOARD. UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE SPECTRUM MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
4006002 | CUSTOMER REPORTED AN ISSUE WITH THE BENEVIEW MONITOR WHICH MAY HAVE AFFECTED O2 MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S GAS BENCH. UNIT WAS CALIBRATE AND SAFETY TESTED TO FACTORY'S SPECS. | 1
4122248 | CUSTOMER REPORTED AN ISSUE WITH THE DPM 4 MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD. | 1
3418441 | COMPANY REPRESENTATIVE REPAIRED COLD SOLDER AT ALL CONNECTORS ON FRONT PANEL BOARD. THE UNIT WAS CALIBRATED, PERFORMANCE AND SAFETY TESTED TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE DPM 6 MONITOR, WHICH MAY HAVE AFFECTED SPO2 MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3574003 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S MAIN BOARD. UNIT WAS TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED THE PASSPORT 2 MONITOR SHUT DOWN, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
3418364 | COMPANY REPRESENTATIVE EVALUATED THE UNIT, REPLACED THE CO2 AND PERFORMED CALIBRATION.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT V MONITOR, WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PT INJURY WAS REPORTED. | 1
4122236 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR SHUT DOWN, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD COMPONENTS. UNIT WAS CALIBRATED AND TESTED TO FACTORY'S SPECS. | 1
3573999 | CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT AND FOUND A FAULTY CABLE. CORRECTIONS INCLUDED REPLACEMENT OF THE SPO2 CABLE. | 1
3832068 | UNITS ARE BEING RETURNED TO MINDRAY FOR FURTHER EVAL.><CUSTOMER REPORTED ISSUES RELATED TO THEIR V SERIES MONITORS, WHICH MAY HAVE AFFECTED SPO2, ECG, AND O2 MONITORING. NO PT INJURY WAS REPORTED. SERIAL NUMBERS: (B)(4). | 1
4322982 | CUSTOMER REPORTED AN ISSUE WITH THE V-SERIES MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S DOCKING STATION BOARD. UNIT WAS SAFETY TESTED TO FACTORY'S SPECIFICATIONS. | 1
4293969 | CUSTOMER REPORTED AN ISSUE WITH THE SPECTRUM MONITORS, WHICH MAY HAVE AFFECTED CO2 MONITORING. NO PATIENT INJURY WAS REPORTED.><COMPANY REPRESENTATIVE EVALUATED UNITS AND DUPLICATED THE REPORTED PROBLEM. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CO2 MODULES. PROBLEM WAS RESOLVED. | 1
3457954 | COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE CO2 MODULE. THE UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED GAS MONITORING. NO PT INJURY WAS REPORTED. | 1
4006052 | COMPANY REPRESENTATIVE EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S DISPLAY. UNIT WAS CALIBRATED AND SAFETY TESTED UNIT TO FACTORY SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE PASSPORT 2 MONITOR, WHICH MAY HAVE AFFECTED PATIENT MONITORING. NO PATIENT INJURY WAS REPORTED. | 1
4321441 | CUSTOMER REPORTED THAT THE DPM 7 MONITOR DID NOT AUDIBLY ALARM WHEN A VENTRICULAR FIBRILLATION AND VENTRICULAR TACHYCARDIA EVENT ON (B)(6) 2014. NO PT INJURY WAS REPORTED.><DOCUMENTATION FROM THE REPORTED EVENT WAS REVIEWED BY MNDRAY'S ENGINEERING TEAM. (PLEASE NOTE THE INVESTIGATION WAS PERFORMED BEING EVENT REPORTS PROVIDED FROM THE PANORAMA CENTRAL STATION. ARRHYTHMIA CALLS ARE INITIATED AND PROCESSED BY THE DPM 7 AND THE PANORAMA REPORTS ARE REFLECTIVE OF EVENTS OCCURRING ON THE DPM 7.) THE REPORT CONFIRM THAT THE REPORTED ARRHYTHMIA EVENTS WERE INITIATED BY THE DPM 7. THESE EVENT REPORTS DO NOT, HOWEVER, PROVIDE SPECIFIC INFO ABOUT THE MONITORS AUDIBLE AND VISUAL ALARM INDICATORS. MINDRAY SERVICE REP ALSO EVALUATED THE UNIT IN QUESTION AND CONFIRMED THE MONITOR IS PERFORMING AS SPECIFIED, THAT IS, THE UNIT ALARM BOTH AUDIBLY AND VISUALLY WHEN ARRHYTHMIA EVENTS WERE INITIATED USING A SIMULATOR. | 1
3418427 | COMPANY REPRESENTATIVE EVALUATED THE UNIT AND COULD NOT DUPLICATE THE REPORTED ERROR MESSAGES. AS A PREVENTATIVE MAINTENANCE, THE CPU BOARD WAS REPLACED. THE UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECIFICATIONS.><CUSTOMER REPORTED AN ISSUE WITH THE SPECTRUM MONITOR, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED. | 1
3457929 | CUSTOMER REPORTED THE PASSPORT 2 MONITOR INTERMITTENTLY SHUTS DOWN, WHICH MAY HAVE AFFECTED PT MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE UNIT. CORRECTIONS INCLUDED REPLACEMENT OF THE UNIT'S CPU BOARD. THE UNIT WAS CALIBRATED AND SAFETY TESTED TO FACTORY SPECIFICATIONS. | 1
4177540 | (B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.><.><THE CUSTOMER REPORTED A SPEAKER MALFUNCTION AND CONFIRMED THAT THAT SPEAKER WAS COMPLETELY OUT OF SOUND. NO PATIENT HARM WAS REPORTED. | 1
3997680 | CUSTOMER REPORTED THAT WHEN THE PM-8000 WAS MONITORING A PATIENT IN THE HOSPITAL, THERE'S INACCURATE MEASUREMENT OF HEART RATE APPEARED, THE HR WAS DISPLAYED TO BE OVER 200BPM. THE DOCTOR REBOOTED THE MONITOR, THEN IT CAN WORK FINE. NO PATIENT WAS DEAD OR INJURED.><BECAUSE THE PM-8000 DOESN'T HAVE THE LOG FILE FUNCTION, THE DEVICE WAS RETURNED FOR ANALYSIS. THE RETURNED MONITOR'S ECG FUNCTION WAS INSPECTED AND TESTED AND THE TEST RESULT SHOWED THAT IT'S PROPER FUNCTIONING AND MEET THE PRODUCT SPECIFICATION. THE CUSTOMER REPORTED PROBLEM COULD NOT BE DUPLICATED. | 1
3836242 | ><WE HAVE A NEW COMPLAINT FROM (B)(6) ABOUT A LEAK OF THE HMEF750S. | 1
3777371 | ON (B)(4) 2014, MEDCOM HAS LEARNT ABOUT A SOFTWARE DEFECT RELATED TO VERISUITE 1.8 VERSION B600.4. THE DEFECT CAUSES MALFUNCTIONING OF THE SOFTWARE WHEN A SLICE OF THE PLANNING CT IS MISSING OR DUPLICATED. VERISUITE WILL DISPLAY A MESSAGE IN CASE OF OMITTED OR DUPLICATE SLICES (A DUPLICATE SLICE WILL BE OMITTED). HOWEVER, WHEN THE USER CONTINUES PATIENT POSITION VERIFICATION BASED ON CT DATASETS WITH DUPLICATE OR MISSING SLICES, THE SOFTWARE DEFECT MAY LEAD TO A SHIFT OF THE ISOCENTER POSITION AND THUS TO A WRONG CORRECTION VECTOR AND INCORRECT PATIENT POSITIONING.><ALL AFFECTED SITES HAVE BEEN CONTACTED ON (B)(4) 2014 VIA A FIELD SAFETY NOTICE (SEE ATTACHED). THE FIELD SAFETY NOTICE RECOMMENDS THE USER NOT TO CONTINUE TREATMENT IF VERISUITE INFORMS ABOUT MISSING SLICES VIA AN ONSCREEN MESSAGE, I.E. IF THE CT DATASET WAS NOT COMPLETELY TRANSFERRED TO VERISUITE. THE USER SHALL RESTART VERISUITE AND RELOAD THE COMPLETE DATASET IN CASE OF MISSING SLICES. THE AFFECTED SITES ARE: (B)(4). THE SITES WERE ALSO INFORMED IN THE FIELD SAFETY NOTICE TO UPGRADE AS SOON AS POSSIBLE TO THE NEWEST VERISUITE VERSION 1.8 B641.3/.4, WHICH IS AVAILABLE AND INCLUDES A BUG FIX FOR THE REPORTED PROBLEM. | 1
3929747 | AMBU ACTION BLOCK PAIN PUMP DISPOSABLE PAIN MANAGEMENT SYSTEM (AABPDPM) PUMP (SOMETIMES CALLED A BALLOON) WAS FILLED WITH BUPIVACINE 1625MH/650ML FOR POST OP FEMORAL BLOCK FOLLOWING TOTAL KNEE ARTHROPLASTY. CATHETER WAS INSERTED BY ANESTHESIA FOR FEMORAL NERVE BLOCK IN PACU. PUMP CONTAINING BUPIVACAINE AND ATTACHED TUBING WERE PRIMED, CONNECTED TO CATHETER PORT AND UNCLAMPED FOR ADMINISTRATION BY SURGICAL UNIT RN. NURSE WAS ACCUSTOMED TO A DIFFERENT PUMP WITH A BUILT-IN FLOW METER THAT DELIVERED 5ML/HR. THE AABPDPM HAS A SEPARATELY-PACKAGED BLOCK-PUMP THAT IS MEANT TO ATTACH BETWEEN THE PUMP TUBING AND THE CATHETER PORT. THE BLOCK-PUMP KIT REGULATES FLOW RATE. RN DID NOT NOTICE THE ABSENCE OF A FLOW METER. THERE IS NO BARRIER IN PLACE TO PREVENT PUMP TUBING FROM CONNECTING DIRECTLY TO CATHETER PORT. PATIENT RECEIVED BUPIVACAINE VIA FEMORAL NERVE CATHETER AT AN UNREGULATED ATE. MEDICATION WAS STARTED AT 1456. AT APPROXIMATELY 1700, PATIENT COMPLAINED OF NUMBNESS IN ARMS AND TONGUE, ROOM SPINNING, AND SLIGHT NAUSEA. PATIENT DEVELOPED DIFFICULTY FOCUSING AND ANSWERING QUESTIONS, SEIZURE-LIKE ACTIVITY, AND DIFFICULTY BREATHING. CPR STARTED AT 1720. PATIENT PRONOUNCED DEAD AT 1929. DEVICE SENT TO OFFICE OF MEDICAL EXAMINER. | 1
3927217 | OVER SEVERAL MONTHS, WHEN RADIOGRAPHS ARE ACCESSED VIA THE PACS MDDS DEVICE LINKED TO THE EHR RESULTS SILOS, THE RADIOGRAPH VIEWED PREVIOUSLY REMAINS ON THE SCREEN DESPITE A NEW NAME BEING ENTERED. WHEN THE NEW PT NAME IS ENTERED, THE PACS MDDS DEVICE DOES ALL SORTS OF BLINKING AND SEARCHING, BUT WHAT COMES UP IS THE PRIOR STUDY, AND THEN, AFTER SEVERAL SECONDS, THE PT'S RADIOGRAPH THAT WAS SOUGHT APPEARS. THIS HAS RESULTED IN ASSESSMENTS AND MANAGEMENT FOR THE WRONG PT AND AN INCORRECT DIAGNOSIS. THIS LOOMS PARTICULARLY OMINOUS SINCE THE FDA IS NOW STATING THAT IT DOES NOT THINK MDDS DEVICE NEED TO UNDERGO SURVEILLANCE. | 1
3553997 | DELAY IN CARE DUE TO BUG IN SCIMAGE TELERADIOLOGY SOFTWARE, WHICH PREVENTED TIMELY TRANSMISSION OF MULTIPLE STAT RADIOLOGY EXAMS. APPLICATION IS ERRONEOUSLY REJECTING ASSOCIATIONS FROM THE FACILITY PACS DUE TO A LICENSE FILE ISSUE. BUG HAS BEEN REPORTED TO MANUFACTURER MULTIPLE TIMES WITHOUT RESOLUTION. | 1
4218913 | DURING DEVELOPMENT OF NORDICBRAINEX V2.1, DATASETS THAT CAN GIVE RISE TO THE DESCRIBED SCENARIO ARE BLOCKED UPON LOADING. FOR NORDICTUMOREX, ALL CUSTOMERS HAVE BEEN INFORMED AND WHERE POSSIBLE ENCOURAGED TO CHANGE TO NORDICBRAINEX. OTHERWISE WE HAVE RECOMMENDED TO EVALUATE THE DICOM DATA USED FOR CLINICAL PURPOSE TO MAKE SURE THAT THEY ARE NOT CONFIGURED IN A WAY THAT TRIGGERS THIS BUG.><A BUG WAS DISCOVERED DURING TESTING INTERNALLY IN THE MFR COMPANY. THE BUG WAS EVALUATED TO HAVE A PT RISK. | 1
3677334 | ><THIS RECORD REFLECTS PATIENT C OF 4 PATIENTS. PATIENT DATA NOT CURRENTLY AVAILABLE. CHILDREN WERE 12 TO 18 MONTHS OLD. DATE OF EVENT: NO PRECISE DATE OBTAINED. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | 1
3695836 | THE FOLLOWING ISSUES WERE REPORTED: IN MO TREATED WINDOW, THE SOLUTION AMOUNT DOES NOT PROPAGATE CORRECTLY EVERY TIME IT IS CHANGED. CHANGE IS DISPENSE DETAIL DOES NOT SHOW THE CHANGED VOLUME, ONLY IF CHANGE FROM CHANGE ORDER WINDOW. WHEN THE USER ENTERS A SOLUTION WITH A LONG NAME (40 CHARACTERS MAX), WHEN THEY GET TO THE MAR, IT EITHER DOES NOT DISPLAY THE SPECIFIED VOLUME AT ALL, OR (DEPENDING ON THE LENGTH OF THE SOLUTION NAME), IT MAY CUT OFF THE VOLUME. FOR EXAMPLE, A 100ML SOLUTION IS DISPLAYED AS 100ML SOLUTION. THERE WAS NO REPORT OF MISTREATMENT. NO FURTHER INFO WAS REPORTED.><(B)(4) | 1
3612604 | I WRITE TO REPORT PROBLEMS I HAVE BEEN EXPERIENCING WITH MY ABOVE-REFERENCED INSULIN INFUSION SYSTEM. I HAVE BEEN USING THIS DEVICE FOR ABOUT TWO YEARS. ABOUT FIVE MONTHS AGO, THE MFR NOTIFIED ME THE FDA HAD APPROVED ITS NEW NEXTGEN" INSULIN DELIVERY SYSTEM | 1
3652813 | I AM A TYPE 1 DIABETIC (INSULIN DEPENDENT). I USE THE OMNIPOD INSULIN PUMP MADE BY INSULET CORP. IN THE PAST 11 MONTHS, MY PUMP HAS FAILED (TURNED ITSELF OFF AND STOPS GIVING ME INSULIN) OVER 11 TIMES (EACH INCIDENT HAS BEEN REPORTED DIRECTLY TO INSULET AND IS A MATTER OF RECORD). IT HAS FAILED 5 TIMES JUST SINCE (B)(6) 2014). IT CAUSES BLOOD SUGAR ISSUES BECAUSE I AM WITHOUT INSULIN UNTIL I AM ABLE TO GET ACCESS TO MORE INSULIN AND ANOTHER INFUSION SET. IT HAS EVEN HAPPENED WHILE I WAS DRIVING. WHEN I AM WITHOUT INSULIN, MY BLOOD SUGAR GOES HIGH. I CONTINUE TO CALL INSULET'S PUMP HELP LINE. THE ONLY THING THEY SAY THEY CAN DO IS REPLACE THE DISPOSABLE POD PART." THEY CONFIRMED ON THE PHONE TODAY THAT I AM USING THE DEVICE CORRECTLY AND THEY ARE UNSURE WHY THE PRODUCT CONTINUES TO FAIL. THE MANAGER | 1
4364140 | (B)(4). WHEN THE INVESTIGATION IS COMPLETED, A F/U REPORT WILL BE SENT TO THE FDA.><PHILIPS RECEIVED A REPORT FROM A CUSTOMER THAT A FEMALE PT WAS POSITIONED FEET FIRST SUPINE, WITH THE TORSO XL COIL AROUND THE BELLY FOR AN EXAMINATION OF THE CERVIX. AFTER THE EXAMINATION A BLISTER OF 2CM IN DIAMETER WAS OBSERVED AROUND THE BELLY BUTTON. | 1
3603207 | THE PATIENT'S DAUGHTER PHONE HOMEDICS ON (B)(6) 2013 TO REPORT HER MOTHER'S UNIT (B)(4) WAS READING LOWER THAN HER DOCTOR'S. AT THE TIME OF THAT CALL, SHE REPORTED THAT SHE HAD TAKEN HER (B)(4) TO A DOCTOR APPOINTMENT, BUT THERE WAS NO CONFIRMATION THIS (B)(4) GAVE A FAULTY READING IN THE PRESENCE OF THE DOCTOR. HOWEVER, HOMEDICS RECEIVED A LETTER FROM THE PATIENT ON (B)(6) 2013 IN WHICH SHE WRITES, AFTER A CHECK-UP WITH MY PRIMARY DOCTOR AFTER A HOSPITAL STAY FOR A STROKE | 1
3602993 | (B)(4).><AN EVALUATION OF THIS FAILURE MODE FROM THE DESIGN FAILURE MODES AND EFFECTS ANALYSIS OF THIS PRODUCT, THAT THE USE OF THE PRODUCT IS A MANNER LIKELY TO CAUSE ADVERSE HEALTH CONSEQUENCE IS IMPROBABLE. THE ROOT CAUSE OF THIS COMPLAINT CANNOT BE IDENTIFIED, SINCE THE PERFORMANCE OF THE RETURNED SUBJECT DEVICE WAS VERIFIED AND THE RESULTS WERE WITHIN THE ACCEPTANCE CRITERIA. HEREWITH THE INVESTIGATION REPORT AS ATTACHMENT. | 1
3960755 | ><WE HAD A RECENT ERROR WHERE U100 INSULIN WAS AUTOMATICALLY CONVERTED TO U500 ON PATIENT'S DEPART MEDS REC IN CERNER HAS BUILT ALL INSULIN'S UNDER THE SAME SYNONYM. WE ASKED CERNER 2 YEARS AGO TO SEPARATE U500 DUE TO CONCERNS WITH CONVERSION ERRORS, BUT THEY HAVE DECLINED TO DO THIS. PT COUNSELING PROVIDED: UNK. RELEVANT MATERIALS PROVIDED: NONE. (B)(6). | 1
3695649 | ><RE: DEFECTIVE MEDICAL PRODUCT COMPROMISING PEDIATRIC PATIENT CARE AT (B)(6). DEAR (B)(4), AS A PEDIATRICIAN, I FEEL THE OBLIGATION TO SUGGEST" TO YOU THAT YOUR COMPANY HAS SOLD A DEFECTIVE MEDICAL SYSTEM TO (B)(6). IT COMPROMISES PEDIATRIC PATIENT CARE. NEXTGEN LISTS GESTATIONAL AGE (HOW LONG A BABY WAS IN HIS MOTHER'S BELLY) UNDER VITAL SIGNS (A LIST | 1
3959919 | SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR MODEL 91393 WITH COMMAND MODULE MODEL 91496 FAILED TO ALARM FOR 24 BEATS OF VENTRICULAR TACHYCARDIA (VTACH) AT 12:13 A.M. (B)(6) 2014.><.><TESTING OF THE SUBJECT MONITOR ONSITE BY A SPACELABS FIELD SERVICE ENGINEER (FSE) CONFIRMED THE DEVICE WAS FUNCTIONING AS DESIGNED. THE MONITOR PASSED ALL FUNCTIONAL TESTS AND GENERATED BOTH VISUAL AND AUDIO ALARMS. SPACELABS IS FURTHER EVALUATING THE PATIENT'S RETROSPECTIVE DATA. THE INVESTIGATION IS UNDERWAY. NO ONE HAS BEEN INJURED AS A RESULT OF THIS ALLEGED MALFUNCTION. WE WILL FILE A SUPPLEMENTAL REPORT WHEN OUR INVESTIGATION IS CONCLUDED.><ONSITE TESTING OF THE INVOLVED DEVICES BY A SPACELABS FIELD SERVICE ENGINEER CONFIRMED THAT THE EQUIPMENT WORKED TO SPECIFICATIONS. THE TESTING WAS WITNESSED BY A FACILITY STAFF MEMBER. PATIENT RETROSPECTIVE DATA FOR THE EVENT IN QUESTION WAS EVALUATED BY A SPACELABS SOFTWARE LEAD ENGINEER. THE ECG ALGORITHM, WHICH PROCESSES THE ECG SIGNAL IN THE COMMAND MODULE, DETECTED NOISY SIGNALS DURING THIS EPISODE OF VENTRICULAR TACHYCARDIA (VTACH). A ¿NOISY SIGNAL¿ STATE SUSPENDS ECG PROCESSING TO PREVENT FALSE ALARMS. THE HIGH FREQUENCY NOISE DETECTOR COMPONENT OF THE ECG ALGORITHM TRACKS LARGE AND SMALL PEAKS IN THE ECG WAVEFORM AND APPLIES A TIERED THRESHOLD TO THE TWO PEAK VALUES TO DETERMINE WHETHER OR NOT ARTIFACT IS PRESENT. THE VERY UNUSUAL EXTRA NOTCHES PRESENT IN THE ECG COMPLEXES FOR THIS EVENT (WHICH ARE UNUSUAL FOR ANY PATIENT AND WERE NOT PRESENT IN 40 OTHER ECG COMPLEXES PRESENTED FOR THIS PATIENT) WERE ENOUGH TO CAUSE THE HIGH FREQUENCY NOISE DETECTOR TO INTERPRET THE ECG WAVEFORM AS ARTIFACT, WHICH RESULTED IN FAILURE TO DETECT THIS VTACH EVENT. A REVIEW OF SIMILAR COMPLAINTS FOUND NO OTHER REPORTS WITH THIS TYPE OF FAILURE TO ALARM ON VTACH. THIS SUPPLEMENTAL REPORT IS CONSIDERED FINAL AND THE ISSUE IS CLOSED.><SPACELABS RECEIVED A REPORT THAT A PATIENT MONITOR, MODEL 91393, FAILED TO ALARM FOR A RUN OF 24 BEATS OF VTACH. THE MONITOR DID GENERATE AN AUDIO ALARM DUE TO TN ARTERIAL PRESSURE EVENT. | 1
4364218 | A CUSTOMER COMPLAINT WAS REPORTED THROUGH THE NATUS NEUROLOGY (B)(4) DISTRIBUTOR. AN (B)(6) CUSTOMER WAS USING A NATUS ENDEAVOR CR DESKTOP SYSTEM ON A PATIENT TO MONITOR EEG/EMG FUNCTIONS. THE CUSTOMER WAS PERFORMING PATIENT STIMULATION IN A NON-RECURRENT MODE AND REPORTED THAT A DOUBLE STIMULATION OCCURRED ON THE PATIENT USING A COMPUTER MOUSE AND BY CLICKING ON THE ASSOCIATED SYSTEM COMPUTER ICON. AFTER PURPOSELY PERFORMING THE FIRST STIMULATION, A SECOND STIMULATION PULSE OCCURRED SUPPOSEDLY WITHOUT FURTHER INPUT FROM THE DOCTOR. FOLLOWING THE SECOND STIMULATION, THE PATIENT EXPERIENCED A SEIZURE WHICH REQUIRED DRUG INTERVENTION TO CONTROL.><REGISTERED INTERNALLY AS A COMPLAINT (REFERENCE (B)(4)) FOR FURTHER INVESTIGATION BY MANUFACTURER. THE SYSTEM MANUFACTURING RECORDS WERE REVIEWED. ALL SYSTEM COMPONENT INSPECTION TESTS WERE IN SPECIFICATION AND NO ISSUES ASSOCIATED WITH THE PRODUCT MANUFACTURER WERE IDENTIFIED THAT MAY HAVE CONTRIBUTED TO THE COMPLAINTS CLAIM. EVALUATION PERFORMED BY THE MANUFACTURER USING DUPLICATE EQUIPMENT COULD NOT REPRODUCE THE REPORTED MULTIPLE STIMULATION FAULT EXCEPT THROUGH REPEATED ACTIVATION OF THE MOUSE/ICON INTERFACE, AN EXPECTED FUNCTION. THE SYSTEM SOFTWARE CODE WAS ALSO REVIEWED CONFIRMING THAT A DOUBLE STIMULATION CANNOT OCCUR ON ITS OWN. EVEN IF THIS PURPOSELY OCCURS THROUGH REPEATED ACTIVATION, STIMULATION OF THIS KIND USING SUCH EQUIPMENT IS NOT CONSIDERED HAZARDOUS, IS NOT CONSIDERED ADEQUATE TO INITIATE A PATIENT SEIZURE, NOR IS SUCH A RESPONSE REPORTED IN LITERATURE AS A POTENTIAL RISK IN ELECTRO-DIAGNOSTIC MEDICINE. | 1
3853464 | (B)(4)><IT WAS REPORTED THAT WHEN A CHANGE IS MADE TO A DRUG NAME, ONLY THE NAME OF THE DRUG IS CHANGED. THE FOLLOWING ASSOCIATED FIELDS WITH THE DRUG ARE NOT CHANGED: - ADMIN INSTRUCTIONS, ERX INSTRUCTIONS, QUANTITY AND UNITS, REFILLS, PRN. THERE WAS NO ACTUAL MISTREATMENT REPORTED AS A RESULT OF THIS ISSUE BASED ON THE AVAILABLE INFORMATION. | 1
4188084 | THE DOOR ON THE BATTERY COMPARTMENT OFTEN WEARS OR BREAKS OFF WHEN BEING OPENED TO CHANGE THE BATTERIES.MANUFACTURER RESPONSE FOR TELE TRANSMITTER/RECEIVER, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================BATTERY DOOR REPLACEMENTS ARE AVAILABLE FOR PURCHASE.MANUFACTURER RESPONSE FOR TELEMETRY TRANSMITTER/RECEIVER, (BRAND NOT PROVIDED) (PER SITE REPORTER).======================BATTERY COVERS ARE AVAILABLE FOR PURCHASE AS A REPLACEMENT PART. | 1
3918231 | ><GENERAL ELECTRIC (GE) PATIENT DATA MODULE (PDM) SPORADICALLY DISCHARGED THE PATIENT FROM THE GE MEDICAL B650 PATIENT MONITOR. THIS ALLOWS FOR POTENTIAL MISSED ALARMS AND LOSS OF DATA FROM THE PDM. THE PATIENT MONITOR EXHIBITED WHAT APPEARED TO BE A COMMUNICATION ERROR. THERE WAS A LOSS OF ALL PARAMETERS FROM THE SCREEN. NO ADVERSE PATIENT OUTCOME. | 1
3918227 | ><BATTERY FAILURE MESSAGE APPEARED ON THE B650 PATIENT MONITOR. THE BATTERY WAS MOVED TO ANOTHER DEVICE AND THE BATTERY FAILURE FOLLOWED WITH THE BATTERY TO THAT MONITOR. TO TROUBLESHOOT FURTHER, THE UNIT WAS UNPLUGGED AND RAN ON BATTERY FOR 60 SECONDS AND PLUGGED BACK IN. THE BATTERY FAILURE WENT AWAY UPON RE-POWERING THE UNIT. | 1
3965945 | ><TRANSFERRING PATIENT TO COMMODE IN ROOM. PATIENT BECAME UNRESPONSIVE, APNEIC, BRADYCARDIC AND HAD A SIX SECOND PAUSE- ASYSTOLE. PATIENT WAS BAGGED X 2-3 BREATHS AND RETURNED TO BASELINE LEVEL OF CONSCIOUSNESS (LOC) AND ASKED WHAT IS WRONG?"MONITOR DID NOT ALARM DURING THIS TIME!!UPON INVESTIGATING MONITOR HISTORY | 1
3569529 | ><THE REGISTERED NURSE CHECKED ICU PATIENTS RHYTHM BY PULLING UP THE PATIENT IN THE CENTRAL MONITOR AT THE NURSES' STATION. HE SAW THAT THE SPO2 AND RESPIRATORY RATE LINES WERE FLAT AND HE WAS CONCERNED ABOUT THE RHYTHM HE SAW SO HE WENT TO THE ROOM TO CHECK THE PATIENT. UPON ENTERING THE ROOM | 1
4222463 | ><PATIENT STARTED ALARMING ASYSTOLE DESPITE HAVING A REAL RHYTHM AT A RATE IN THE 60'S. BIOMEDICAL ENGINEERING CONTACTED WHO CHANGED LEAD PLACEMENTS WITHOUT CHANGE. REMOVED PATIENT FROM THE SECONDARY ALARM NOTIFICATION SYSTEM (SANS) TO PREVENT ALARM FATIGUE AND IMPLEMENTED TELE TECH TO MONITOR CENTRAL MONITORING STATION.BIOMED FOLLOW-UP: THE ISSUE WAS THE AMPLITUDE FROM THE QRS COMPLEX WAS NOT ENOUGH TO GIVE A RATE, WHICH CAUSED A FALSE ALARM. GE HAS NO DEFINITIVE FIX FOR THE ISSUE. THE BEST FIX SO FAR IS TO USE THE GE PACER LEAD CONFIGURATION, TO SCRUB THE SKIN WITH A 4X4, USE SIGNA GEL AND THE PACER LEAD CONFIG, AND CHECK PATIENTS EVERY DAY ON ALARMS. | 1
4092067 | ><NO COMMUNICATION" MESSAGE FROM MONITOR | 1
3979478 | ><PATIENT WAS IN RECOVERY BAY 7 CONNECTED TO A GE CARESCAPE B650. ACCORDING TO THE NURSE, THE MONITOR TURNED ITSELF OFF AND THEN BACK ON. NO INJURY WAS REPORTED FROM THE INCIDENT.====================== MANUFACTURER RESPONSE FOR MONITOR, MULTI PARAMETER, CARESCAPE B650 (PER SITE REPORTER)======================WE HAD CONTACTED OUR GE REPRESENTATIVE AND WERE INFORMED GE IS WORKING ON PROVIDING AN SOLUTION TO THIS ISSUE. | 1
4240903 | ><INVESTIGATION REVEALED THE LOSS OF THE ECG DEFIB SYNC OUTPUT FUNCTION FROM THE PDM DEFIB SYNC CONNECTOR WAS CAUSED BY FAILURE OF THE PDM DEFIB SYNC OUTPUT BOARD. | 1
3777343 | ><ON THE APEXPRO TELEMETRY BOX TO CLOSE THE BATTERY DOOR, A RUBBER GRIP STRIP GIVES THE NURSE BETTER ACCESS TO CLOSE THE DOOR. THE GRIP STRIPS ARE CONSTANTLY FALLING OFF, MAKING IT ALMOST IMPOSSIBLE TO CLOSE OR OPEN THE BATTERY DOOR. I HAVE TO ORDER THE ENTIRE OUTSIDE CASE, LABELS, AND THE KEYPAD TO FIX THE ISSUE. | 1
3594426 | ><MONITOR REPORTED ARRHYTHMIA HISTORY AT 1530 AS A PVC EVENT AT 1336. NO PRINT OUT WAS RECORDED. NO PATIENT HARM. | 1
3415182 | AN OLYMPUS FIELD SERVICE ENGINEER ATTENDED THE (B)(4) MEDICAL CENTER TO CARRY OUT AN INITIAL INVESTIGATION INTO THE EVENT. THE ENGINEER FOUND THAT THE PROBLEM ORIGINATED FROM THE MAIN PLUG OF THE WORKSTATION NOT BEING FULLY INSERTED INTO THE POWER SOCKET WHICH IS LOCATED IN THE CEILING. IT WAS NOT NECESSARY TO RESET THE CONTROL AT THE REAR OF THE TRANSFORMER. THE ENGINEER FULLY INSERTED THE MAIN PLUG INTO THE POWER POCKET IN THE CEILING AND THE POWER SUPPLY TO THE WORKSTATION WAS RE-INSTATED. THERE WAS NO REPORT OF INJURY TO THE PT AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. DEVICE EVAL: THE DEVICE IS NOT BEING RETURNED FOR EVAL AS THERE IS NO MALFUNCTION OF THE DEVICE.><THE OLYMPUS WM-NP2 MOBILE WORKSTATION IS INTENDED FOR USE IN MEDICAL FACILITIES UNDER THE DIRECTION OF A TRAINED PHYSICIAN AND HAS BEEN DESIGNED TO BE USED WITH A RAGE OF OLYMPUS EQUIPMENT TO FACILITATE GI ENDOSCOPY, ENDOSCOPIC ULTRASOUND, RESPIRATORY AND SURGICAL ENDOSCOPIC PROCEDURES. AT THE END OF AN ENDOSCOPIC SUBMUCOSAL DISSECTION OF THE STOMACH, AFTER THE SEPARATION WAS COMPLETED AND THE BLEEDING HAD STOPPED, THE WM-NP2 WHICH IS THE SUBJECT OF THIS REPORT, LOST POWER RESULTING IN THE LOSS OF ENDOSCOPIC IMAGE. THE PT WAS MOVED TO ANOTHER ROOM AND THE PROCEDURE WAS COMPLETED WITH ANOTHER WORKSTATION. THERE IS NO REPORT OF INJURY TO PT AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. | 1
4171143 | OLYMPUS WORKSTATIONS ARE DESIGNED FOR USE WITH A RANGE OF OLYMPUS EQUIPMENT TO FACILITATE ENDOSCOPY PROCEDURES. THE WORKSTATION HAS TRANSFORMER WHICH IS LOCATED UNDERNEATH AND TO THE REAR OF THE LOWER SHELF. THE TRANSFORMER IS CONNECTED TO THE MAINS POWER SUPPLY BY A POWER CORD AND PLUG. OLYMPUS EQUIPMENT IS CONNECTED TO POWER CORD AND PLUG. OLYMPUS EQUIPMENT IS CONNECTED TO THE TRANSFORMER TO FACILITY ENDOSCOPY PROCEDURES. (B)(4). HAS BEEN MADE AWARE THAT DURING AN UNKNOWN PROCEDURE THE POWER OF THE WM-NP2 WAS SHUT DOWN AND A SPARK WAS GENERATED FROM THE PLUG OF THE POWER CABLE. THERE IS NO REPORT OF PATIENT INJURY AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION.><THE SUBJECT DEVICE HAS BEEN RETURNED TO OLYMPUS (B)(4). A BURN ON THE PLUG OF POWER CABLE HAS BEEN OBSERVED. (B)(4) HAS REQUESTED THE RETURN OF THE POWER CORD AND PLUG TO ENABLE US TO CONDUCT A FULL INVESTIGATION. THERE IS NO REPORT OF PATIENT INJURY AND THIS REPORT IS SUBMITTED IN AN ABUNDANCE OF CAUTION. | 1
3492174 | DEVICE INVESTIGATION: INVESTIGATION HAS DETERMINED THAT THIS ISSUE IS THE RESULT OF THE DONOR USING THE BROWSER BACK BUTTON OR APPROPRIATE KEY STROKE TO NAVIGATE TO A PREVIOUS PAGE FROM AN ADD'L INFO PAGE. THE BROWSER WILL LOAD THE BROWSER EXPIRED PAGE. THEN THE DONOR UTILIZES THE REFRESH OPTION TO RE-LOAD THE PAGE WHICH CAN CAUSE DATA TO BE OVERWRITTEN IN THE DATABASE. THE APPLICATION IS EXPECTING THE DONOR TO USE THE CONTROL BUTTONS (NEXT, PREVIOUS, SAVE, CANCEL, ETC) PROVIDED WITH THE APPLICATION FOR NAVIGATION. THE INVALID RESPONSE IS ONLY SEEN IN THE DATABASE, NOT THE QUESTION SUMMARY PAGE, AND ONCE THE QUESTIONNAIRE IS COMPLETED THE ENTRIES IN THE DATABASE ARE SAVE CORRECTLY. FURTHER INVESTIGATION REPORTED THAT A CLIENT MAY USE (B)(4)'S KIOSK MODE FOR THEIR QUESTIONNAIRES. IN KIOSK MODE, THE NAVIGATION BUTTONS ARE NOT DISPLAYED TO THE DONOR, HOWEVER, THE DONOR CAN STILL USE SHORTCUTS (I.E. ALT-LEFT ARROW FOR BACK, F5 FOR REFRESH, ETC) TO GO BACK AND REFRESH THE PAGE. ON (B)(6) 2013, HAEMONETICS SOFTWARE SOLUTIONS SUBMITTED THE FOLLOWING: PART 806 REPORT IN CONNECTION WITH THE FIELD ACTION OF (B)(4). THE DATES BELOW ARE FROM THE 806 REPORT (B)(6) 2013 - HEALTH RISK ASSESSMENT (HRA) DATE OF INCIDENT REPORTED/DISCOVERED. FIRST EMAIL COMPLAINT FROM CUSTOMER. THE INFO RECEIVED FROM THE CUSTOMER WAS NOT ENOUGH TO EVALUATE RISK. ON (B)(4) 2013, COMPLAINT ASSESSED IN PRODUCT COMPLAINT REVIEW COMMITTEE (PCRC). DETERMINED IT WAS A SOFTWARE BUG NOT A COMPLAINT. ON (B)(4) 2013 - EMAIL COMPLAINT FROM SAME CUSTOMER THAT A SERVER ERROR OCCURRED. WORKED WITH THE CUSTOMER TO OBTAIN FURTHER INFO AND DID FURTHER INVESTIGATION TO TRY AND RECREATE THE ISSUE. ON (B)(4) 2013 - COMPLAINT ASSESSED IN PCRC. DETERMINED IT NEEDED A RISK ASSESSMENT. ON (B)(4) 2013 - HRA WAS COMPLETED (B)(4) 2013 - HRA SIGNATURE FOR COMPLETE BY. ON (B)(4) 2013 - DEVELOPED RECALL STRATEGY AND SAFETY NOTIFICATION. ON (B)(4) 2013 - SAFETY NOTIFICATION AND CORRECTED SOFTWARE VERSION SENT OUT TO ALL CUSTOMERS. ON (B)(4) 2013 - OBTAINED REMAINING HRA SIGNATURES. THE 806 E-FILED AND HARDCOPY SENT VIA (B)(4). FDA RESPONSE RECEIVED ON (B)(4) 2013. RECALL #B-0008-2014.><WHEN A DONOR IS TAKING A SELF-ADMINISTERED QUESTIONNAIRE AND IS PROMPTED FOR ADD'L INFO, IF THE DONOR NAVIGATES THROUGH THE BROWSER'S HISTORY, AND THEN USES THE BROWSER'S REFRESH" BUTTON | 1
4179939 | AT APPROXIMATELY 11:00 PM, A SURGICAL ATTENDANT, WHILE ON ROUTINE ROUNDS, SMELLED SMOKE IN A NON-ATTENDED OPERATING ROOM. THE ATTENDANT IMMEDIATELY UNPLUGGED THE WORKSTATION ON WHEELS (WOW) CART, PULLED IT AWAY FROM THE ANESTHESIA MACHINE WHERE IT WAS IN VERY CLOSE PROXIMITY (ANESTHESIA MACHINE INCLUDES OXYGEN), AND THEN THREW WATER ON IT. THE BATTERY OF THE WOW LOCATED AT THE BASE OF THE CART, APPEARED AS IF IT HAD EXPLODED AS EVIDENCED BY MELTED PARTS OF THE EXTERIOR FLOOR BASE AS WELL AS SOME BURNT DAMAGE TO THE FLOOR. THE WOW WAS REMOVED AND SEQUESTERED ON THE LOADING DOCK. THE OPERATING ROOM 1B WAS CLOSED AT THE TIME OF DISCOVERY AND REMAINS OUT-OF-USE WHILE THE INVESTIGATION ENSUES. NO STAFF OR PATIENTS WERE INJURED.THE DEVICE IS A HIGH EFFICIENCY AC POWERED COMPUTER CART WITH 40 AMP LITHIUM IRON (LIFEPO4) BATTERY (16 INCH HEIGHT ADJUSTMENT, INTERNAL STORAGE FOR SMALL FORM CPUS, FULLY ADJUSTABLE KEYBOARD TRAY, KEYBOARD LIGHT). | 1
3729625 | MEDICATION ERROR. THIS EVENT WAS RELATED TO HEALTH INFORMATION TECHNOLOGY. SPECIFICALLY, THE MANNER IN WHICH THE SYSTEM PROCESSED AN ORDER FOR FREE WATER REPLACEMENT. ORDER WAS INTENDED TO BE 20ML/HR FOR 6 HOURS FOR A TOTAL OF 120ML. IT WAS A 1000 ML BAG SO FIRST THEY PUT 20ML/HR WHICH WOULD HAVE DEFAULTED THE INFUSE OVER TIME TO 50 HRS. THEY TRIED TO CHANGE THE INFUSE OVER TO 6 HOURS WHICH THEN CHANGES THE ML/HR RATE TO 167 ML/HR. THEY DID NOT NOTICE THIS HAD CHANGED. THEY DID NOT REALIZE THAT THEY NEEDED TO GO INTO DETAILS TAB TO SHOW THE TIME FRAME SO THE PATIENT GOT 167ML/HR INSTEAD OF 20ML/HR FOR 6 HOURS. ORDER WAS VERIFIED BY PHARMACY AND ADMINISTERED AT THAT RATE VIA PERIPHERAL IV. AFTER APPROXIMATELY 1L OF FLUID HAD INFUSED, PATIENT SHOWED SEIZURE ACTIVITY. PICU TEAM CALLED TO BEDSIDE. TREATMENT PROVIDED FOR SEIZURES AND CRITICAL SODIUM AND POTASSIUM.THE CERNER POWERCHART SOFTWARE SYSTEM HAS FUNCTIONALITY FOR CONTINUOUS MEDICATIONS/FLUID ORDERING THAT DOES NOT PREVENT USER ERROR." WHEN THE PROVIDER ORDERS THE FLUID | 1
3831895 | USERS WERE UNABLE TO LOGIN TO A SELECT NUMBER OF CLINICAL WORKSTATIONS AFTER AN UPDATE/APPLICATION WAS INSTALLED WHICH ALLOWS USERS TO RESET THEIR OWN PASSWORDS. MOST NOTABLY, A CARDIOLOGY COMPUTER PROVIDING CLINICAL DATA TO THE CATH LAB WAS ALSO IMPACTED BY THE FIX DEPLOYED TO THE AFOREMENTIONED. PRIMARILY CLASS 1 TYPE MACHINES (WHICH HAVE A SINGLE SIGN ON SOFTWARE) WERE IMPACTED. IN AUDITION, A CARDIOLOGY COMPUTER WHICH ACTS LIKE A SERVER TO PROVIDE DATA TO THE LAB, AND IS RUNNING WINDOWS XP WAS MISIDENTIFIED AS A CLASS 1 MACHINE AND WAS ALSO IMPACTED. IT SHOULD BE NOTED THAT THIS SOFTWARE UPDATE HAS BEEN ROLLING OUT IN BATCHES OVER THE LAST THREE WEEKS WITHOUT INCIDENT.THERE WAS NO ACTUAL SOFTWARE" THAT CAUSED THE PROBLEM. RATHER | 1
4098118 | DESPITE AN INCREASING AND MARKEDLY ELEVATED POTASSIUM, THE POTASSIUM PILL THAT WAS ON THE COMPUTERIZED MEDICATION ADMINISTRATION NURSING TASK LIST WAS ADMINISTERED TO THE PATIENT, CAUSING FURTHER INCREASE IN SERUM POTASSIUM LEVELS. THE CLINICAL DECISION SUPPORT IS DUMB. THE LAB DATA IS UNKNOWN BY THE NURSING STAFF BECAUSE IT DOES NOT COME TO THEM, BUT RATHER, GOES DIRECTLY IN TO THE MDDS LAB SECTION. THE NURSES, LIKEWISE, DO NOT KNOW THAT A POTASSIUM LEVEL WAS ORDERED AND OBTAINED AND ANALYZED, AS THE INFORMATION AVAILABLE TO THEM IS IN SEPARATE NON-CONTIGUOUS SILOS. THE WORK FLOW OF CARE DOES NOT COMPENSATE FOR THE IMPEDIMENTS OF THE EHR AND MDDS. | 1
4098127 | THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH A WOBBLY GAIT, ATAXIA. A VITAMIN B12 LEVEL WAS ORDERED BY THE NEUROLOGY CONSULTANT, IN ADDITION TO OTHER TESTS. IT TAKES SEVERAL DAYS FOR THE RESULT TO APPEAR, WHICH ULTIMATELY GETS SENT TO THE EHR AND STORED AS OF THE DATE THE SAMPLE WAS OBTAINED. THE LEVEL WAS ABNORMAL, LOW, AND CONSISTENT WITH THE GAIT PROBLEM; BUT NO ONE SAW THE RESULT, INCLUDING THE NEUROLOGIST. THE NEUROLOGIST TREATED THE PATIENT FOR PARKINSON'S DISEASE INSTEAD, ERRONEOUSLY. RESULTS ARE COMMONLY MISSED WHEN STORED IN COMPLEX EHR SYSTEMS WITH POOR PRESENTATION ON SCREENS WITH MEANINGFULLY USELESS EXCESS DESCRIPTION, POSTED ON A DAY THAT EXCEEDS (FURTHER IN THE PAST) THE DEFAULT DISPLAY OF RESULTS. | 1
4114532 | THE MDDS DISPLAY OF LABORATORY DATA IS MISLEADING BECAUSE THE DATE AT THE TOP OF THE COLUMN DOES NOT APPLY TO ALL LAB DATA BELOW IN THE SAME GENERAL COLUMN. THUS, THERE MAY BE DATA, FOR INSTANCE A LOW HEMOGLOBIN LEVEL, THAT WAS OBTAINED ONE YEAR AGO, THAT APPEARS IN THE SAME COLUMN AS YESTERDAY'S RESULT. THIS LEADS TO ERRORS AND DANGEROUS DECISIONS, BECAUSE THEY ARE BASED ON DATA THAT IS NOT CURRENT BUT MAY BE THOUGHT TO BE CURRENT. DEVICES WITH SUCH MISLEADING FORMATS SHOULD BE RECALLED. | 1
3906231 | PATIENT WITH FEVER AND RESPIRATORY FAILURE. CHEST X-RAY ORDERED AND WAS PERFORMED. RADIOGRAPHS ON THIS PATIENT WERE NOT ACCESSIBLE ON THE MDDS LINKED TO THE PROTOUCH EHR DEVICE. THIS RESULTED IN DELAY OF DIAGNOSIS AND CARE. THERE WERE NOT ANY ALTERNATIVE WORK AROUNDS TO ACCESS THIS CHEST X-RAY, INCLUDING SEARCH MECHANISMS. NO STAFF MEMBER COULD FIND IT EITHER. | 1
3913670 | A CRITICALLY ILL PT UNDER WENT RADIOGRAPHIC EVAL OF CHEST AND ABDOMEN. THE LAST NAME OF THE PT CONTAINED ONE APOSTROPHE. THE RADIOGRAPH IMAGES COULD NOT BE ACCESSED ON THE EHR RESULTS MDDS. IT WAS DETERMINED THAT THE ONE ENTERING THE PT'S NAME AT THE IMAGING VENDOR ENTERED A DOUBLE APOSTROPHE, RATHER THAN ONE. IT COULD NOT BE CORRECTED FOR DAYS, ONCE THE IMAGES WERE FOUND, 5 DAYS AFTER THEY WERE DONE. IT TOOK ANOTHER 3 DAYS FOR PACS VENDOR TO CORRECT THIS MISIDENTIFICATION ISSUE. THE VENDOR'S DEVICE IS DEFECTIVE BECAUSE IT ALLOWED ABSURDITY (THERE IS NEVER A NAME WITH CONSECUTIVE APOSTROPHES) AND IT FAILED TO WARN OF THE ERROR. THESE MDDS DEVICES NEED TIGHTER REGULATION, SURVEILLANCE, AND SAFETY. | 1
4046963 | ><THE REPORTING FACILITY ALLEGED THAT THE HEMOGLOBIN VALUE DISPLAYED IN ONE OF THE (B)(4) CARDIOLOGY HEMO APPLICATION SCREENS WAS NOT CORRECTLY CONVERTED IN CALCULATED RESULTS WHICH ARE USED TO SUPPORT IN THE TREATMENT DECISION FOR CARDIAC PATIENTS. THE REPORTING FACILITY REVIEWED AFFECTED PATIENT RECORDS AND DETERMINED THAT NO CHANGES TO TREATMENT DECISIONS WERE REQUIRED. NO HARM TO PATENTS HAS BEEN REPORTED AS A RESULT OF THIS ISSUE. | 1
3645369 | ><AN INVESTIGATION WAS PERFORMED USING A SIMULATED TESTING ENVIRONMENT TO DETERMINE THE ROOT CAUSE OF THE REPORTED EVENT. THE RESULTS OF THE INVESTIGATION CONCLUDED THAT INCORRECT INFORMATION WAS DISPLAYED ON THE PAE SCREENS OF THE REPORTING FACILITY. ALTHOUGH THE INFORMATION WAS INCORRECTLY DISPLAYED ON THE PAE SCREEN, THE DATABASE STORAGE AND ASSOCIATED REPORTS REFLECTED THE CORRECT PATIENT MEDICAL CONDITION DATA. THE ACTUAL DEFECT DID NOT OCCUR EVERY TIME AND WAS SPORADIC IN NATURE. FURTHER ANALYSIS DETERMINED THAT THE MALFUNCTION WAS LIKELY CAUSED BY A DEFECT IN THE DEVICE SOFTWARE. THERE HAVE BEEN NO ADVERSE EVENTS OR PATIENT INJURIES RESULTANT FROM THE REPORTED ALLEGED MALFUNCTION. | 1
3987001 | ><OXYGEN TUBING SELF-DISCONNECTS AT NRB MASK CONNECTION. DATES OF USE: (B)(6) 2014 - (B)(6) 2014. DIAGNOSIS FOR USE: SUPPLEMENTAL OXYGENATION TO SEDATED PATIENTS. | 1
4182408 | PATIENT RECEIVED RF BURN FROM SKIN TO SKIN CONTACT BETWEEN BREAST AND CHEST DURING ABDOMINAL MRI STUDY.><THE BURN WAS ALLEGEDLY CAUSED BY A SKIN TO SKIN LOOP FROM THE BREAST TOUCHING THE CHEST. THE PATIENT WAS WEARING A HOSPITAL GOWN AND THERE WAS PAD BETWEEN COIL AND PATIENT'S CHEST.. THE CABLE WAS ROUTED STRAIGHT WITH NO LOOPS AND PATIENT WAS NOT TOUCHING THE COIL OR THE BORE WALL. ALLEGEDLY THE BURN IS SLIGHTLY SMALLER THAN A QUARTER, BETWEEN FIRST AND SECOND DEGREE PER THE REVIEWING PHYSICIAN. COILS WERE CHECKED AND THERE WAS NO APPARENT DEFECT AND NO DAMAGE WAS NOTED. A QUALITY ASSURANCE CHECK WAS CONDUCTED FOR THE COILS (I.E., SIGNAL TO NOISE AND ALL COIL SET UP CHECKS) AND THEY PASSED. | 1
3815974 | ><DURING THE END OF THE PROCEDURE (2300 SHOCKS) SYSTEM STOPPED SHOCKING AND AN ODOR/BURNING SMELL WAS NOTED. TRIED TO TROUBLESHOOT AND RESTART MACHINE BUT BURNING SMELL CONTINUED. STONE MOSTLY BROKEN UP SO CASE STOPPED.PATIENT'S CURRENT CONDITION: SIGNIFICANT PORTION OF STONE DESTROYED. UNKNOWN IF REMAINING PORTION OF STONE WILL PASS. | 1
4251571 | ><PATIENT UNDER GENERAL ANESTHESIA. LITHROTRIPSY MACHINE STOPPED WORKING MID-CYCLE. COMPANY EMPLOYEE THAT WAS HERE OPERATING THE MACHINE CALLED FOR SERVICE. UNABLE TO RESTART AFTER SEVERAL ATTEMPTS TO REBOOT. DR. ABORTED CASE. PATIENT WAS BROUGHT OUT OF ANESTHESIA AND CASE WAS RESCHEDULED. NO HARM TO PATIENT, ALTHOUGH PATIENT REQUIRED ANOTHER SURGICAL PROCEDURE FOR LITHOTRIPSY WITH ANESTHESIA.====================== MANUFACTURER RESPONSE FOR LITHOTRIPSY, LITHOTRON (PER SITE REPORTER)======================THE MAIN GENERATOR FAILED. THEY HAVE REPLACED THE MAIN GENERATOR ON THE MACHINE. PERFORMANCE TESTING SHOWED NO ADDITIONAL PROBLEMS. PATIENT'S REPEAT PROCEDURE WAS PERFORMED WITHOUT INCIDENT. | 1
3610055 | ><LITHOTRIPSY MACHINE FAILED AFTER LESS THAN 100 IMPULSES. UNABLE TO REBOOT THE MACHINE. REMAINING CASES FOR THE DAY REQUIRING THE LITHOTRIPSY MACHINE WERE CANCELLED. THE EQUIPMENT WAS REMOVED BY THE VENDOR. THE MALFUNCTION WAS WITHIN THE UNIT, WHICH WAS JUST SHORT OF IMPULSES NEEDED FOR PREVENTATIVE MAINTENANCE. THE LITHOTRIPTOR WAS REPAIRED AND IS NOW AVAILABLE. | 1
4114586 | ><CHILD SEEN IN URGENT CARE FOR A LACERATION ON RIGHT EYE, JUST ABOVE RIGHT EYE EYEBROW. MD PLACE LIQUIBAND FLEX ON LACERATION, AND A SMALL AMOUNT DRIBBLED DOWN CHEEK AND A TINY BIT RAN INTO THE OUTER CORNER OF HIS RIGHT EYE CAUSING THE EYELASHES AND LIDS TO ADHERE TO EACH OTHER. MD NOTED THAT THIS PRODUCT IS TOO LOOSE" MORE TENDS TO BE "RUNNY". NOT PLEASED WITH PRODUCT PERFORMANCE." | 1
3944058 | THE PANEL PROCESSING CUSTOMIZATION WINDOW IN (B)(4) ALLOWS THE OPERATOR TO CUSTOMIZE THE 24 HOUR HOLD OPTIONS FOR (B)(6) USING DRIED OVERNIGHT PANELS. TIS WAS DISCOVERED DURING CURRENT SOFTWARE DEVELOPMENT ACTIVITIES. PER THE DRIED GRAM POSITIVE PROCEDURAL MANUAL. ACCURATE DETECTION OF RESISTANCE REQUIRES AN EXTENDED INCUBATION TIME OF 24 HOURS FOR (B)(6) ORGANISM AND (B)(6) ANTIMICROBICS.><THIS IS AN INTERNAL COMPLAINT. FURTHER INVESTIGATION REVEALED A GAP IN THE PROCEDURAL REQUIREMENTS FOR ASSAY AND SOFTWARE DESIGN CONTROLS. THE PROCEDURE DID NOT MAKE PROVISIONS TO DOCUMENT SOFTWARE CONSIDERATIONS, IN ADDITION TO THE ALGORITHM. THE PROCEDURE WAS REVISED TO INCLUDE THIS REQUIREMENT, AS WELL AS OTHER REQUIREMENTS FOR ADDITIONAL APPLICABLE DETAILED LABELING INFORMATION. FOR EXAMPLE, INSTRUCTIONS WERE ADDED FOR THE LABPRO OPERATOR'S GUIDE IN ORDER TO DOCUMENT REQUIREMENTS THAT AFFECT THE SOFTWARE. (B)(6). | 1
3839607 | THE CUSTOMER CONTACTED THE SIEMENS TECHNICAL SOLUTIONS CENTER (TSC) AND STATED THAT PATIENT RESULTS POSTED BY THE SYNGO LAB DATA MANAGER APPLICATION SOFTWARE SHOULD INDICATE RESULT COMMENTS AND THE MEASUREMENT DATE AND TIME ON PATIENT REPORTS. THERE ARE NO KNOWN REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO THE SYNGO LAB DATA MANAGER APPLICATION SOFTWARE NOT INDICATING RESULT COMMENTS AND THE MEASUREMENT DATE AND TIME ON PATIENT REPORTS.><SIEMENS HEALTHCARE DIAGNOSTICS HAS INVESTIGATED THE INSTANCES OF RESULTS BEING RELEASED TO THE LABORATORY INFORMATION SYSTEM WHEN THEY SHOULD HAVE BEEN HELD FOR REVIEW. SIEMENS HEALTHCARE DIAGNOSTICS HAS CONFIRMED THAT CERTAIN CONFIGURATION PARAMETERS SELECTED DURING CONFIGURATION OF REFERENCE RANGES FOR ASSAYS MAY CAUSE THE REFERENCE RANGES TO NOT HOLD RESULTS AS EXPECTED. URGENT MEDICAL DEVICE CORRECTION (UMDC) 14-30 ENTITLED SYNGO LAB DATA MANAGER: CUSTOM REFERENCE RANGE AND SAMPLE REPORT ISSUES" WAS SENT TO CUSTOMERS IN THE UNITED STATES IN MAY 2014. URGENT FIELD SAFETY NOTICE (UFSN) 14-29 ENTITLED "SYNGO LAB DATA MANAGER: CUSTOM REFERENCE RANGE AND SAMPLE REPORT ISSUES" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN MAY 2014. THE UMDC/UFSN INSTRUCTS CUSTOMERS TO REVIEW THE CUSTOM REFERENCE RANGES THEY CURRENTLY HAVE IN PLACE | 1
3694609 | INSTRUMENT MANAGER IS ALSO LISTED AS MMH SOFTWARE, BLOOD BANK STAND ALONE PRODUCT. INSTRUMENT MANAGER IS CREATED USING INTERSYSTEMS CACHE, WHICH IS AN ADVANCED OBJECT DATABASE. ON (B)(6) 2013, DATA INNOVATIONS DETERMINED THAT THE CAUSE OF THE RULES NOT PROCESSING AFTER THE DIVIDE BY ZERO ISSUE OCCURRED WAS DUE TO A CACHE ERROR TRAPPING ISSUE. THIS ISSUE WAS CONFIRMED BY INTERSYSTEMS TO BE A PROBLEM WITH CACHE VERSIONS: 2011.1 THROUGH 2012.0, THESE VERSIONS WERE RELEASED FOR USE IN 2011. THIS WAS CORRECTED IN INTERSYSTEMS CACHE VERSION: 2012.1.0 IN 2012 WITH CHANGE CDS2047. ALL INSTRUMENT MANAGER 8.12 VERSIONS (V8.12.01, V8.12.02, V8.12.10, V8.12.11, V8.12.20, AND V8.12.21) WERE BUILT USING AFFECTED VERSIONS OF CACHE. THE FOLLOWING CACHE VERSIONS WERE USED TO CREATE THE FOLLOWING INSTRUMENT MANAGER (IM) VERSIONS: 2011.1.2.701.0.11458 - FOR IM V8.12.01, V8.12.02, V8.12.10, V8.12.11, V8.12.20, 2011.1.2.701.0.12368 - FOR IM V8.12.21. NOTIFICATION OF BUGS IN CACHE IS MADE VIA EMAIL ANNOUNCEMENTS FROM INTERSYSTEMS. INTERSYSTEMS DID NOT NOTIFY DATA INNOVATIONS OF THE BUG THAT CAUSED THIS ERROR TRAP NOR DID INTERSYSTEMS NOTIFY DATA INNOVATIONS WHEN THE ISSUE WAS CORRECTED. DATA INNOVATIONS SUBMITTED RECALL # 1225673-12/17/2013-002-C TO THE FDA.><CUSTOMER SUPPORT INCIDENT # (B)(4) WAS SUBMITTED TO DATA INNOVATIONS (B)(4) ON (B)(6) 2013 BY ROCHE DIAGNOSTICS FOR (B)(6) FOR A RULES ISSUE THAT CAUSED THE RULES TO NOT PROCESS WHEN THE RULE HAD A DIVISION OF ZERO. THE LAB IS A HIGH VOLUME LAB AND WAS REVISING RULES TO FIT THEIR NEEDS WHEN THE LAB DETERMINED THAT RULES WERE NOT PROCESSED AFTER A DIVIDE BY ZERO ERROR OCCURS. ROCHE DIAGNOSTICS' TECHNICIAN CORRECTED THE PROBLEM FOR THE CUSTOMER BY REVISING THE RULE TO CHECK FOR A ZERO VALUE BEFORE DOING DIVISION. THERE WAS NO PT HARM REPORTED IN THIS INCIDENT BECAUSE THE LABORATORY CAUGHT AND CORRECTED THE RULE DURING TESTING OF THOSE RULES. IF A RULESET CONTAINS ENOUGH RULES TO CREATE A ROUTINE LARGER THAN 64KG IN SIZE AND AN ERROR OCCURS OR IF A RULE ATTEMPTS TO DIVIDE BY 0, SUBSEQUENT RULES MAY NOT BE EVAL. DURING THE SAVING OF RULES, THE RULES ARE CONVERTED FROM THEIR DISPLAY FORMAT INTO GENERATED ROUTINES THAT GET COMPILED AND EXECUTED AT RUNTIME. IF THE SIZE OF THE ROUTINE IS LARGER THAN 64KG, THIS ISSUE MAY OCCUR WHEN AN ERROR OCCURS. FOR THIS ISSUE TO OCCUR THE FOLLOWING CRITERIA MUST BE MET: THE GENERATED ROUTINE MUST BE GREATER THAN 64KB; IF THE PARENT OF THE CHILD RULE DOES NOT HAVE AN IF STATEMENT, THE ERROR MUST OCCUR IN THE SECOND CHILD OR SUBSEQUENT CHILD RULE. HOWEVER, IF THE PARENT OF THE CHILD RULE HAS AN IF STATEMENT, THE ERROR COULD OCCUR IN ANY CHILD RULE. THE RULE SET CRITERIA ARE MET, THE SUBSEQUENT SIBLING RULES IN THAT GROUP OF THE CHILD RULE THAT ERRED WILL NOT BE EVALUATED. | 1
3677423 | I'M LOOKING TO FIND OUT IF AN ACTIVELY LOCKED/ UNLOCKED, REMAINDERED PULL DISPENSING SYSTEM SHOULD BE FDA CONTROLLED. I HAVE CONCERNS THAT FLAWS IN THE DESIGN, IF IT WERE NOT APPROVED BY THE FDA, COULD RESULT IN PT HARM (NO PILLS, TOO MANY PILLS, ETC). | 1
4174486 | USER FACILITY REPORTS LOWER THAN REFERENCE WITH A PROTIME MICROAGULATION SYSTEM. PROTIME GENERATED A 1.1 INR RESULT. PATIENT WAS TESTED WITHIN 24 HOURS AT REFERENCE LABORATORY AND THE RESULT WAS 4.7 INR. PATIENT'S THERAPEUTIC RANGE: 2.0 - 3.0 INR. NO ADVERSE EVENT(S) WERE REPORTED.><(B)(4). CUSTOMER TESTED WITH CUVETTE LOT#D4K3C019. METHOD: ACTUAL DEVICE NOT EVALUATED. PROCESS EVALUATION PERFORMED. NO NCMRS OR ANOMALIES IDENTIFIED FOR THIS INSTRUMENT. CUVETTE DEVICE HISTORY RECORDS REVIEWED AND NO RELATED ISSUES WERE IDENTIFIED. RESULT CODE: NO RESULTS AVAILABLE SINCE NO EVALUATION PERFORMED. CONCLUSION: DEVICE FAILURE RELATED TO PATIENT CONDITION. IT WAS REPORTED THAT THE PATIENT WAS ON LOVENOX. IT IS COMMON FOR PATIENT'S ON LOVENOX TO REGISTER LOWER THAN EXPECTED INR RESULTS USING THE PROTIME MICROAGULATION SYSTEM. ITC HAS REQUESTED ALL DATA REQUIRED FOR FORM 3500A. | 1
3599047 | ><I USED AN INSTANT COLD COMPRESS (HENRY SCHEIN INSTANT COLD COMPRESS PART # (B)(4)). THE COMPRESS WAS PLACED ON MY RIGHT WRIST AREA. WITHIN 10 MINUTES I BEGAN TO EXPERIENCE BURNING SENSATION. COMPRESS WAS IMMEDIATELY REMOVED. THE AREA OF MY RIGHT ARM WAS A REDDISH COLOR, HARD TO TOUCH, STINGING AND BURNING. I IMMEDIATELY CONTACTED MY PCP AND WAS SEEN WITHIN 1.5 HOUR. DURING MY MEDICAL VISIT THE PHYSICIAN EXAMINED MY ARM AND DETERMINED I HAD SUSTAINED A CHEMICAL BURN FROM THE COMPRESS. THE COMPRESS PACK DID NOT HAVE ANY LEAKS OR PUNCTURES AND NO EXPIRATION DATE. I AM RIGHT HAND DOMINANT. DIFFICULTY TO PERFORM WORK, HOME, AND NORMAL HYGIENE. I HAVE BEEN DIAGNOSED WITH 2ND, 3RD & 4TH DEGREE BURNS. I AM EXPERIENCING PAIN IN MY ENTIRE RIGHT ARM FROM WRIST TO ELBOW. HARD TO GRIP OBJECTS, USE COMPUTER, COOK, ETC. VERY PAINFUL, UNEXPECTED EVENT. | 1
3732002 | THE PHYSICIAN PERFORMS A CARDIAC CATHETERIZATION AND DURING THE PROCEDURE THE PATIENT IS PLACED IN DIFFERENT CONDITIONS" (FOR EXAMPLE: WITH OXYGEN AT 30% AND THEN WITH OXYGEN AT 100%). THE SYSTEM RECORDS INFORMATION DURING EACH OF THESE CONDITIONS AND RECORDS IT ELECTRONICALLY. THE ELECTRONIC VERSION OF THE CATHETERIZATION HAD THE CORRECT INFORMATION. AT THE TIME THE REPORT WAS PRINTED | 1
3651217 | ><GE GIRAFFE OMNIBED SET TO 35 DEGREES CELSIUS, 80% HUMIDITY. OMNIBED DISPLAY READ 81% HUMIDITY. EXTERNAL HUMIDITY TEST DEVICE READ 100%+. | 1
4184301 | THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION, AS THE DEVICE IS BEING RETAINED BY THE REPORTING FACILITY. A LOT HISTORY REVIEW (LHR) REVIEW IS NOT POSSIBLE, AS NO MANUFACTURING LOT NUMBER HAS BEEN PROVIDED BY THE COMPLAINANT.><NURSE WHO USES THE SITE RITE VISION II SYSTEM COMPLAINED THAT THE SYSTEM IS DIFFICULT TO PUSH AND MANEUVER ESPECIALLY WHEN THE SYSTEM GAINS MOMENTUM. THE NURSE HAS COMPLAINED THAT PUSHING THE SYSTEM HAS CAUSED PAIN IN THEIR HIPS, KNEE AND OTHER JOINTS. THE RN NOW GOES TO A CHIROPRACTOR ONCE A WEEK DUE TO SHOULDER PAIN. | 1
3971515 | (B)(4). THE INVESTIGATION DETERMINED THAT ICT CALIBRATOR LOTS HAVE DECREASED TOTAL PROTEIN CONTENT. THE DECREASED TOTAL PROTEIN CONTENT IS CORRELATED WITH INCREASED AMMONIA LEVELS IN THE CALIBRATORS AMMONIA, A KNOWN INTERFERENT, CAN IMPACT THE PERFORMANCE OF THE ICT MODULES POTASSIUM MEASUREMENT. A PRODUCT CORRECTION LETTER HAS BEEN ISSUED TO THE ICT SERUM CALIBRATOR CUSTOMERS WHO RECEIVED SHIPMENTS OF LOTS 21399UN13, 67873UN13, 74710UN13, 82008UN13, AND 88884UN13. THE LETTER IS INSTRUCTING THE CUSTOMERS THAT THEY CAN CONTINUE USE OF THE LOT IF THE CALIBRATION CURVE FOR POTASSIUM IS VERIFIED WITH AT LEAST TWO LEVELS OF CONTROLS ACCORDING TO ESTABLISHED QUALITY CONTROL REQUIREMENTS FOR THE CUSTOMER LABORATORY. THE LETTER IS ALSO INSTRUCTING THE CUSTOMERS TO RECALIBRATE USING A DIFFERENT CALIBRATOR LOT IF QUALITY CONTROLS FOR THE POTASSIUM ASSAY ARE NOT WITHIN THE ESTABLISHED QUALITY CONTROL REQUIREMENTS OF THE CUSTOMER LABORATORY (NEW LOTS ARE AVAILABLE) STABILITY MONITORING OF IN-DATE LOTS IS TO BE CONDUCTED><THE DATE ENTERED IN SECTION DATE RECVD BY MFR:" IN THE INITIAL REPORT WAS INCORRECT. THE CORRECT DATE IS 07/10/2014.><THE CUSTOMER STATED THAT THE POTASSIUM QUALITY CONTROLS SHIFTED LOW WHILE USING THE ARCHITECT ICT SERUM CALIBRATOR LOT 74710UN13. NO DISCREPANT PATIENT RESULTS OR IMPACT TO PATIENT MANAGEMENT WAS REPORTED." | 1
3992054 | (B)(4). THE INVESTIGATION DETERMINED THAT ICT CALIBRATOR LOTS HAVE DECREASED TOTAL PROTEIN CONTENT. THE DECREASED TOTAL PROTEIN CONTENT IS CORRELATED WITH INCREASED AMMONIA LEVELS IN THE CALIBRATORS AMMONIA, A KNOWN INTERFERENT, CAN IMPACT THE PERFORMANCE OF THE ICT MODULES POTASSIUM MEASUREMENT. A PRODUCT CORRECTION LETTER HAS BEEN ISSUED TO THE ICT SERUM CALIBRATOR CUSTOMERS WHO RECEIVED SHIPMENTS OF LOTS 21399UN13, 67873UN13, 74710UN13, 82008UN13, AND 88884UN13. THE LETTER IS INSTRUCTING THE CUSTOMERS THAT THEY CAN CONTINUE USE OF THE LOT IF THE CALIBRATION CURVE FOR POTASSIUM IS VERIFIED WITH AT LEAST TWO LEVELS OF CONTROLS ACCORDING TO ESTABLISHED QUALITY CONTROL REQUIREMENTS FOR THE CUSTOMER LABORATORY. THE LETTER IS ALSO INSTRUCTING THE CUSTOMERS TO RECALIBRATE USING A DIFFERENT CALIBRATOR LOT IF QUALITY CONTROLS FOR THE POTASSIUM ASSAY ARE NOT WITHIN THE ESTABLISHED QUALITY CONTROL REQUIREMENTS OF THE CUSTOMER LABORATORY (NEW LOTS ARE AVAILABLE) STABILITY MONITORING OF IN-DATE LOTS IS TO BE CONDUCTED><THE CUSTOMER STATED THAT POTASSIUM QUALITY CONTROLS ARE RUNNING LOW WHILE TESTING WITH THE ARCHITECT ICT SERUM CALIBRATOR LOT 74710UN13. NO DISCREPANT PATIENT RESULTS HAVE BEEN REPORTED AND NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 1
4163139 | HELIUM OUTFLOW / FILL PORT ON A MAQUET INTRA AORTIC BALLOON PUMP DEVICE WSA FOUND TO BE OCCLUDED BY A DEAD END MALE/FEMALE LUER LOCK PLUG AND AS A RESULT, THE IAB WOULD NOT INFLATE AFTER BEING DEPLOYED. THIS PLUG EASILY FITS INTO THE OUTFLOW PORT. IT WAS NOT DETECTED UNTIL AFTER THE BALLOON WAS DISCONTINUED FROM THE PT AND DISCONNECTED FROM THE PUMP ITSELF. IT WAS VERY DIFFICULT TO SEE THE PLUG AND IT FIT SNUGLY ONTO THE END PORTS AS IF IT BELONGED THERE. NOTE - THE PT EXPIRED BEFORE HER DEATH WAS NOT CAUSED BY THIS EVENT. WHEN A DECISION WAS MADE TO INSERT AN IABP, CPR WAS ALREADY IN PROGRESS AND THE PT HAD BEEN WITHOUT VITAL SIGNS FOR SEVERAL MINUTES. | 1
3647140 | ><BROWNISH PARTICLES AND DISCOLORATION ON THE INSIDE OF A NEWLY OPENED STERILE 10CC HYPODERMIC SYRINGE. ENTIRE CASE OF SYRINGES WAS DISCARDED. | 1
4016723 | THE FIELD SERVICE REPRESENTATIVE (FSR) COULD NOT CONFIRM THE REPORTED ISSUE, AS THE HEATER COOLER UNIT COOLED TO THREE DEGREES CELSIUS. THE FSR CHECKED THE UNIT'S CALIBRATION AND OPERATION INCLUDING CHECKING FOR FLOW BLOCKAGES AND MIXING VALVE OPERATION. THE UNIT OPERATED TO MANUFACTURER SPECIFICATIONS AND WAS RETURNED TO CLINICAL USE. IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS AND/OR CONCLUSION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.><IT WAS REPORTED THAT DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS (CPB) PROCEDURE, THE HEATER COOLER UNIT WAS NOT COOLING DOWN TO DESIRED TEMPERATURE. THE DEVICE WAS NOT CHANGED OUT. THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO DELAYS, NO BLOOD LOSS, OR NO ADVERSE CONSEQUENCES TO THE PATIENT. PER THE CLINICAL REVIEW ON (B)(6) 2014: DURING CARDIOPLEGIA DELIVERY, THE PERFUSIONIST (CCP) HAD ONE CHANNEL OF THE HEATER COOLER UNIT SET TO A WATER SET POINT (MAINTAIN MODE) OF THREE DEGREES CELSIUS. THE CCP HEARD A SINGLE AUDIBLE TONE THAT LET HIM BECOME AWARE OF AN ISSUE, BUT NO ERROR CODE WAS OBSERVED. ACCORDING TO THE CCP, THE WATER PUMP FOR THIS CHANNEL SHUT DOWN. THE CCP RAISED THE SET POINT TO FIVE DEGREES CELSIUS AND THE WATER PUMP RESUMED. HE FINISHED THE REST OF THE CASE WITH THE CARDIOPLEGIA WATER SET POINT OF FIVE DEGREES CELSIUS AND HE HAD NO FURTHER ISSUES THE REMAINDER OF THE PROCEDURE. AFTER THE PROCEDURE, THE CCP CONTACTED THE FIELD SERVICE REPRESENTATIVE (FSR) AND TOOK THE UNIT OUT OF SERVICE. (B)(4)/MDR #1828100-2014-00665 CLINICAL REVIEW CONTINUED: THE CCP STATED THE FIELD SERVICE REPRESENTATIVE (FSR) REPAIRED THE UNIT AT THE HOSPITAL AND THERE HAVE BEEN NO FURTHER ISSUES. THE CASE WAS COMPLETED SUCCESSFULLY WITHOUT DELAY AND WITHOUT ASSOCIATED BLOOD LOSS. THERE WAS NO HARM OBSERVED OR REPORTED. | 1
4110121 | IT WAS REPORTED THAT DURING SET-UP OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THERE WERE WATER LEVEL ALARMS ON BOTH CHANNELS OF THE HEATER COOLER UNIT. THE FRONT PANEL LIGHT CAME ON (RED LIGHT WITH WATER BUCKET ICON). AS A RESULT, AN ALTERNATE DEVICE WAS EMPLOYED. THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WERE NO DELAYS, NO BLOOD LOSS, OR NO ADVERSE CONSEQUENCES TO THE PATIENT.><THE FIELD SERVICE REPRESENTATIVE (FSR) FOUND THE USER DID NOT HAVE ENOUGH CHLORINE IN THE WATER WHEN HE TESTED THE UNIT. THE FSR SPOKE WITH THE PERFUSIONIST (CCP) ABOUT THE CHLORINE LEVELS. NO PARTS BEING RETURNED FOR EVALUATION. IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS AND/OR CONCLUSION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY. | 1
4349506 | THIS COMPLAINT IS RELATED TO MDR 1828100-2014-01061. PER THE FIELD SERVICE REPRESENTATIVE, THIS ERROR IS CAUSED BY A FAILING MIX VALVE, WHICH IS CURRENTLY BACK ORDERED.><IT WAS REPORTED THAT DURING USE OF THE DEVICE FOR A CARDIOPULMONARY BYPASS (CPB) PROCEDURE, AN E0E" ERROR CODE WAS OBSERVED ON THE HEATER COOLER UNIT. THE DEVICE WAS NOT CHANGED OUT | 1
3688196 | EXCESSIVE FLUID RETENTION, CELLULITIS, PAIN, INFECTION, LABORED BREATHING, FLUID IN LUNGS, BECAME WHEELCHAIR BOUND, COULD NOT DO ANYTHING FOR SELF. HAD TO LEAVE JOB AND WAS HOSPITALIZED AFTER SUSPECTING PD CATHETER WAS MALFUNCTIONING; IT WAS FOUND TO BE WORKING PROPERLY. HAD TO HAVE ANOTHER PROCEDURE TO PLACE A CATHETER FOR HOME ANTIBIOTICS. FAMILY WAS WATCHING ME DROWN. EMOTIONAL ROLLER COASTER AS I BEGAN TO PREPARE FOR MY DEATH. HOPELESS AND LOT OF ANGUISH. PHYSICAL HEALTH WAS POOR. VITALS WERE UNSTABLE. I WAS DOING HOME PD, THEN HAVING TO GO INTO CENTER TO HAVE HD TO HELP PULL OFF FLUID THAT MACHINE WAS DUMPING IN AND NOT TAKING OFF. AFTER SEVERAL REEVALUATIONS AND ASSESSMENTS, THE DECISION WAS MADE TO CEASE PD. ALMOST 3 MONTHS LATER, I AM STILL HAVING EXTRA HD TREATMENTS TO PULL EXCESS PD FLUID OFF SUCH AS TODAY (B)(6)2014. | 1
3406111 | (B)(4). THE DEVICE WAS RETURNED AND THE EVALUATION WAS COMPLETED. THIS IS AN ANCILLARY SERVICE EVENT. THE INCREASED INTRA PERITONEAL VOLUME (IIPV) EVENT WAS CONFIRMED THROUGH AN EVENT HISTORY LOG REVIEW. THE CAUSE WAS ONE OR MORE CYCLE ADVANCES TO THE NEXT FILL WHEN A SLOW / NO FLOW OCCURRED, ABOVE THE MINIMUM DRAIN VOLUME THRESHOLD. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.><DURING EVALUATION OF A RETURNED HOMECHOICE DEVICE, ONE INCREASED INTRA-PERITONEAL VOLUME (IIPV) EVENT WAS IDENTIFIED. THIS OCCURRED IN THE THERAPY INITIATED ON (B)(6) 2013, DURING NIGHT DRAIN CYCLE SIX. THE PATIENT'S ULTRAFILTRATION READING WAS 1835ML, INDICATING THE HOME PATIENT (HP) DRAINED 1835ML MORE THAN THEIR MAXIMUM PROGRAMMED FILL VOLUME OF 2500ML. NO ADDITIONAL INFORMATION IS AVAILABLE. | 1
3994432 | THE SYSTEM IS INTERMITTENTLY SHUTTING DOWN WHICH IS CAUSING INCONVENIENCE. AWAITING THE DEVICE FOR FURTHER INVESTIGATION. NO HARM TO PT REPORTED AND NO ADD'L TREATMENT NEEDED.><THE SYSTEM IS INTERMITTENTLY SHUTTING DOWN WHICH IS CAUSING INCONVENIENCE. | 1
3949764 | INCONSISTENT FUNCTIONALITY WITH THE HEMODYNAMIC MONITORING SYSTEM CAUSED 3 X MESSAGES OF ASTERISKS INSTEAD OF NUMBERS. THE POTENTIAL FOR PT HARM IS OF CONCERN, AS THIS OCCURS DURING CARDIAC CATHETERIZATION. THE MFR WAS NOTIFIED RIGHT AWAY. THEIR ENGINEERS ARE WORKING ON IT. THIS OCCURRED AFTER A SOFTWARE UPDATE. REBOOTING THE SYSTEM SEEMS TO CORRECT THE PROBLEM, HOWEVER, THE PT INFO IS NOT AVAILABLE DURING THAT TIME. | 1
4100919 | ><FROM 2011 TO DATE WE HAVE HAD TO REPLACE 12 OF THE 45 FRONT CONNECTOR PANELS ON THE GE S/5 PRETN HEMODYNAMIC MODULES REPORTS OF INOPERABLE/INACCURATE BP MONITORING FROM THE NIBP CONNECTION. OVER APPROXIMATELY 2 MONTH PERIOD FROM (B)(6) 2014 7 OF THE MODULES HAVE NEEDED REPLACEMENT. THIS HAS BEEN ATTRIBUTED TO A LOOSE" NIBP CONNECTION RESULTING IN INACCURATE OUTPUT OR COMPLETE LOSS OF SIGNAL DUE TO NIBP CABLE FALLING OUT OF THE MODULE." | 1
4152609 | UROLOGY CASE USING GYRUS WORKSTATION GENERATOR. GENERATOR STARTED TO OVERHEAT. AN ELECTRICAL SMELL AND SMOKE WAS NOTED. UNIT WAS UNPLUGGED AND TAKEN TO CLINICAL ENGINEERING. INTERNAL WIRING HAD BURNED. NO INJURY TO PT. POSSIBLE USER ERROR. | 1
3716251 | ><THE HOSPITAL RECENTLY PLACED A NEW AVANCE ANESTHESIA UNIT INTO USE IN ROOM 6. ONE OF THE FLOW SENSOR TRANSDUCER FLAPPERS GOT STUCK IN THE UP POSITION (PART # 1503-3858-000). THIS PROBLEM HAS BEEN OCCURRING FREQUENTLY ON THE NEW AVANCE MODELS AT THE HOSPITAL. THE GE FIELD SERVICE REP AND GE SALES REP ARE BOTH AWARE OF THIS ISSUE. | 1
4124586 | ><DURING OR CASE ANESTHESIA MACHINE DEVELOPED CIRCUIT LEAK.FROM BIOMED INVESTIGATION: CHECKING THE LOGS, THERE IS ONLY ONE ENTRY FOR THE MIXER FAILURE. ERROR = MIXER O2 SELECTION VLV FAIL ISSUE/CAUSE = ALTERNATE O2 SCREEN MIXER STATUS BIT: STS_SELV_VOXY_NOTIFY_FAIL . REASON = THE STATUS OF THE O2 SELECTOR VALVE DOES NOT MATCH THE COMMANDED STATE. MEDIUM AC, MIXERRESOLUTION = REBOOT SYSTEM. IF PROBLEM CONTINUES, REPLACE THE MIXER. PN = MIXER ASSEMBLY - COMPLETE (B)(6). | 1
4302664 | ><AT CONCLUSION OF CABG, THE SURGEON ASKED THE ANESTHESIOLOGIST TO SWITCH THE PATIENT OFF HEART/LUNG BYPASS MACHINE BACK TO VENTILATOR. VENTILATOR ON ANESTHESIA MACHINE WOULD NOT FUNCTION PROPERLY (ERROR MESSAGE: NOT ENOUGH PRESSURE). VENTILATOR WAS CONNECTED TO O2 COLUMN AND BACKUP O2 TANK. THE MACHINE WOULD NOT SWITCH OVER TO BACKUP O2 TANK. COLUMN AND TANK TESTED AND OPERATING NORMALLY. DELAY IN REMOVING PATIENT FROM BYPASS AS PORTABLE VENTILATOR CONNECTED. EVENT LOG CHECKED POST FAILURE. STATED THAT ALTERNATE O2 TANK ACTIVATED DUE TO MIXER FAILURE. STAFF REPORTED THAT O2 TANK DID NOT ACTIVATE EITHER. MIXER REPLACED. | 1
3852965 | ><ONCE THE PATIENT INTUBATED, THE POSITIVE END-EXPIRATORY PRESSURE (PEEP) WAS MEASURING 10 CM. THE ANESTHESIA MACHINE PEEP WAS SET AT 0CM(ZERO), WITH PEEP MEASURING 10 CM. THE MACHINE WAS THEN TURNED OFF, AND THE PATIENT WAS VENTILATED BY HAND WITH AN AMBU BAG WITH 100% OXYGEN. THE LUNGS WERE DETERMINED TO BE CLEAR BUT THE PATIENT'S HEART RATE (SVT) WAS 140BPM, THEY WERE ALSO HYPOTENSIVE. | 1
4029800 | ><UPON TURNING THE PATIENT TO THE PRONE POSITION, THE ENDOTRACHEAL (ET) TUBE BECAME BLOCKED. CLINICIANS DID NOT OBSERVE ANY ALARMS FROM THE ANESTHESIA MACHINE FOR WELL OVER 1 MINUTE, DURING WHICH TIME THE PATIENT WAS NOT BEING VENTILATED. STAFF COULD TELL FROM THE DISPLAY THERE WAS A PROBLEM, AND CORRECTED THE OBSTRUCTION BEFORE ANY INJURY OCCURRED.THIS APPEARS TO BE AN ISSUE WITH THE SOFTWARE. WHEN USING THE PCV-VG VENTILATOR MODE, IF THE ET TUBE BECOMES KINKED OR OBSTRUCTED, THE ALARM SEQUENCE IS DIFFERENT THAN THE OTHER MODES. ALL OUR AISYS MACHINES BEHAVE THE SAME WAY.IN THE OTHER MODES SUCH AS VCV OR SIMV/PSV MODE, IF THE ET TUBE IS OBSTRUCTED, THE ALARM GOES OFF IN ABOUT 5 SECONDS WITH FLASHING RED ON THE SCREEN AND A RAPID 5 BEEP ALARM EVERY 10 SECONDS. THIS IS HARD TO MISS.IN PCV-VG MODE, IF THE ET TUBE IS OBSTRUCTED, THE FIRST ALARM DOES NOT APPEAR FOR 45 SECONDS, AND IS A SINGLE BEEP WITH 'TV NOT ACHIEVED' MESSAGE, NO FLASHING. AFTER ANOTHER 45 SECONDS, A SECOND ALARM 'MV LOW' AND 3 BEEPS EVERY 30 SECONDS APPEARS. IT IS VERY EASY TO MISS THE SINGLE BEEP OF THE FIRST ALARM. THE SECOND ALARM IS NOT FLASHING RED AND WOULD BE EASY ENOUGH TO MISS THE FIRST TIME.CLINICIANS REPORT THERE HAVE BEEN SEVERAL NEAR-MISS EVENTS CAUSED BY THIS PROBLEM.====================== MANUFACTURER RESPONSE FOR ANESTHESIA MACHINE, (BRAND NOT PROVIDED) (PER SITE REPORTER). ====================== ISSUE IS UNDER INVESTIGATION. | 1
3775725 | ><ON (B)(6) 2014, A PT WAS LIFTED OFF OF THE GE INFINIA HAWKEYE TABLE ONTO A GURNEY AND WHEELED OUT OF THE ROOM. THE TECH PUSHED THE UNLOAD" BUTTON TO SET THE TABLE UP FOR THE NEXT PT AND WALKED OUT OF THE ROOM. WHEN THE TABLE WAS FULLY LOWERED | 1
3599331 | ABBOTT FREESTYLE LITE BLOOD GLUCOSE TEST STRIPS LOT NUMBER: 1378940, DATE: 04/2015. THIS CONTAINER OF STRIPS GAVE READING BETWEEN 58 AND 132 WITHIN A 5 MINUTE PERIOD. THIS COULD NOT BE THE CASE THAT MY HUSBAND'S BLOOD SUGAR WOULD CHANGE LIKE THIS. HE DOES NOT USE INSULIN. THIS IS NOT A LOT THAT HAS BEEN RECALLED. IT HAS BEEN PRETTY CONSTANT OVER THE PAST FEW MONTHS THAT HE HAS TO TAKE A READING FOUR TIMES IN THE MORNING TO GET SOME READINGS THAT ARE CLOSE. WHEN I HAVE CALLED THE COMPANY IN THE PAST, WE ARE TOLD IT IS NORMAL TO HAVE READINGS VARY BY 20%. THIS IS NOT REASONABLE. PLEASE HELP THE COMPANY TO TAKE THE STEPS TO HAVE MORE ACCURATE TEST STRIPS. THEY ARE EXPENSIVE. THE STRIPS FROM THIS CONTAINER HAVE ALL BEEN USED. | 1
3584096 | THE FOLLOWING ISSUE WAS DISCOVERED VIA A PUBLISHED INDIVIDUAL CASE REPORT BY THE MITEK SAFETY OFFICER: INDIVIDUAL CASE REPORTS (COMPLICATIONS). AUTHOR: CAVAIGNAC (2013) FRANCE. REFERENCE: CAVAIGNAC E, PAILHÉ R, REINA N, CHIRON P, AND LAFFOSSE JM. MASSIVE PROXIMAL EXTRAVASATION AS A COMPLICATION DURING ARTHROSCOPIC ANTERIOR CRUCIATE LIGAMENT RECONSTRUCTION. KNEE SURG RELAT RES 2013;25(2):84-87. ADDRESS: ETIENNE CAVAIGNAC, MD. INSTITUT LOCOMOTEUR - SERVICE DE CHIRURGIE ORTHOPÉDIQUE ET TRAUMATOLOGIQUE, CHU RANGUEIL, 1, AVENUE JEAN POULHÈS TSA 50032, 31059 TOULOUSE CEDEX 9, FRANCE. TEL: +33-6007342595, FAX: +33-561322232. E-MAIL: CAVAIGNAC.ETIENNE@GMAIL.COM. PATIENTS GROUP: N/A. PATIENT CHARACTERISTICS: AGE: (B)(6). GENDER: MALE. INDEX OPERATION: ACL REPAIR (OLD ACL TEAR + POSTERIOR HORN AND MEDIAL MENISCUS TEAR > FEMORAL SHAFT FRACTURE). DEVICES: FMS DUO+ TIME OF COMPLICATION: DURING OPERATION (OPERATION ENDED AFTER 105 MIN. AND 15 LITERS OF IRRIGATION FLUID.) TYPE OF COMPLICATION: EXTRAVASATION (THIGH, RIGHT ILIAC FOSSA, PERINEUM, AND SCROTUM). - - JOINT: KNEE. DEVICE: FMS DUO+ (PRESSURE VALUE ERROR). SUTURE: -- OTHER DEVICES: - SIGNS AND SYMPTOMS: SWELLING OF THE THIGH, RIGHT ILIAC FOSSA, PERINEUM AND SCOTUM. TISSUE DAMAGE: NO COMPARTMENT SYNDROME, PULSES WERE NORMAL, NO SENSORY OR MOTOR DISTURBANCES, BLOOD PRESSURE NORMAL, NO SIGN OF BLOOD LOSS. IMAGING STUDIES: ABDOMINAL AND PELVIC ULTRASOUND: EFFUSION OF THE SUBPERITONEAL SPACE AND SCROTUM. REOPERATION (RELATED/UNRELATED): ACL REPAIR USING THE TUNNELS FROM THE FIRST SURGERY (RELATED). MEDICAL TREATMENT: FUROSEMIDE. F/U TIME: 5 YRS. IMAGING STUDIES: - OBJECTIVE OUTCOMES: NO SEQUELAE. FUNCTIONAL OUTCOMES: - SUBJECTIVE OUTCOMES:><OUR PATIENT SAFETY DEPARTMENT DISCOVERED THIS PUBLISHED WHITE PAPER DETAILING A HISTORICAL EVENT IN WHICH SOME MITEK DEVICES WERE IMPLICATED. WE CANNOT CONFIRM THAT THIS ISSUE HAD BEEN PREVIOUSLY REPORTED TO MITEK, SO WE ARE FILING AN ADVERSE EVENT REPORT TO DOCUMENT THE EXPERIENCE DESCRIBED IN THE ARTICLE. AT THIS POINT IN TIME, NO FURTHER ACTION IS WARRANTED. HOWEVER, THIS FILE WILL REMAIN RECEPTIVE TO ANY POTENTIAL FORTHCOMING INFORMATION RECEIVED THAT IS PERTINENT AND GERMANE TO THIS ISSUE. ALSO, MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. | 1
4005040 | PATIENT WITH DIAGNOSIS OF MENORRHAGIA WAS UNDERGOING A HYSTEROSCOPIC RESECTION OF FIBROID. THE PATIENT WAS BROUGHT INTO THE OPERATING ROOM. HER CERVIX WAS VISUALIZED AND SEQUENTIALLY DILATED TO ALLOW THE OPERATIVE HYSTEROSCOPE. THE HYSTEROSCOPE WAS INTRODUCED INTO THE CAVITY. DEFICIT IN THE KARL STORZ FLUID MANAGEMENT SYSTEM WAS INCREASING THROUGHOUT THE CASE. WHEN CHECKING ALL THE CONNECTIONS, THE POLE FELL DOWN THAT WAS HOLDING THE CANNISTERS. THIS WAS FELT TO BE THE REASON THEY WERE NOT GETTING ACCURATE MEASUREMENTS. USING THE MORCELLATOR DEVICE, THE FIBROID WAS REMOVED. DURING THE CASE, ONCE IT WAS DETERMINED THAT THE KARL STORZ FLUID MANAGEMENT SYSTEM WASN'T FUNCTIONING PROPERLY, FLUID WAS MANUALLY MEASURED. WHEN THE DEFICIT HAD REACHED 800 MILLITERS OF SODIUM CHLORIDE, THE SURGEON COMPLETED RESECTING. THE INSTRUMENTS WERE REMOVED. GOOD HEMOSTATSIS WAS NOTED. THE PATIENT WAS THEN TRANSFERRED TO RECOVERY ROOM IN GOOD CONDITION. | 1
3766044 | THIS REPORT WAS FILED LATE BECAUSE THE INCIDENT WAS INCORRECTLY CLASSIFIED AS NOT MDR REPORTABLE. A RETROSPECTIVE REVIEW OF COMPLAINTS; HOWEVER, IDENTIFIED THIS INCIDENT AS MDR REPORTABLE. AS A RESULT AN UPDATED VERSION OF 6.1.1. OF THE FITTING SOFTWARE HAS BEEN PROVIDED TOGETHER WITH A CUSTOMER LETTER FOR THE ACOUSTICIANS, INFORMING THEM ABOUT THE ERROR ON (B)(6) 2009. THE LOW LIKELIHOOD OF AN OCCURRENCE OF THE PROBLEM IN COMBINATION WITH THE RISK LEVEL DID NOT NECESSITATE FURTHER ACTION. IT WAS DETERMINED AFTER DISCUSSIONS WITH SHI AUDIOLOGY THAT CUSTOMER NOTIFICATION IS NOT REQUIRED SINCE THE ERROR ONLY OCCURS UNDER VERY SPECIFIC CONDITIONS, AND IN THOSE INSTANCES AN AUDIOLOGY FOLLOWING PROPER FITTING PROTOCOLS WOULD DETECT AND REMEDIATE THE PROBLEM THROUGH 2CM3 COUPLER OR PROBE-MICROPHONE MEASUREMENTS. SECTION 6 OF 2003 AAA PROTOCOL REQUIRES THAT THE AUDIOLOGIST MATCH THE ELECTROACOUSTICAL PERFORMANCE OF THE INSTRUMENT TO THE PRESCRIBED 2CM3 COUPLER TARGET VALUES FOR GAIN AND OUTPUT LIMITING, OR PERFORM PROBE-MICROPHONE MEASUREMENTS. EITHER OF THESE METHODS PRECLUDES AN AUDIOLOGIST FROM FITTING AN INSTRUMENT AFFECTED BY THE INCORRECT CLINICAL FIT ALGORITHM. AS OF TODAY ACOUSTICIANS ARE USING LATER VERSIONS OF THE SOFTWARE 6.4 AND HIGHER IN ORDER TO BE ABLE TO FIT HE CURRENT HEARING AIDS ON THE MARKET.><A BUG HAS BEEN DETECTED IN FITTING SOFTWARE CONNEXX 6.1 DURING AN INTERNAL SOFTWARE CHECK. THIS BUG ONLY OCCURS IN A RARE COMBINATION OF SOFTWARE SETTINGS. NORMALLY THE FAILURE WOULD BE DETECTED BY VERIFICATION OF THE TARGETS WITH AN INDEPENDENT MEASUREMENT SYSTEM TO CORRECT THE SETTING IF NECESSARY. AS A RESULT THE LIKELIHOOD OF OCCURRENCE OF THIS FAILURE IS NEARLY ZERO. POTENTIAL IMPACT: THE INCORRECT AUDIOGRAM TRANSFORMATION WOULD CAUSE A HIGHER GAIN THAN DESIRED TO COMPENSATE FOR THE INDIVIDUAL HEARING LOSS. THIS ERROR MAINLY IMPACTS LOW FREQUENCIES AND EXCEEDS UP TO 19BD ONLY. IT WAS DETERMINED AFTER DISCUSSIONS WITH SHI AUDIOLOGY THAT CUSTOMER NOTIFICATION IS NOT REQUIRED SINCE THE ERROR ONLY OCCURS UNDER VERY SPECIFIC CONDITIONS, AND IN THOSE INSTANCES AN AUDIOLOGIST FOLLOWING PROPER FITTING PROTOCOLS WOULD DETECT AND REMEDIATE THE PROBLEM THROUGH 2CM3 COUPLER OR PROBE-MICROPHONE MEASUREMENTS. | 1
4021145 | ON (B)(6) 2014, AT 13:35, A 40 WEEKS AND 5 DAYS PREGNANT FEMALE PRESENTED TO HOSPITAL FOR INDUCTION OF LABOR SECONDARY TO POST TERM PREGNANCY AND NEWLY DIAGNOSED GESTATIONAL HYPERTENSION. FETAL MONITOR WAS CONNECTED. ON (B)(6) 2014, AT 21:27, A FEMALE INFANT WAS BORN WITH TIGHT DOUBLE NUCHAL CORD NOTED AND HANDED IMMEDIATELY TO PEDIATRICIAN. CHEST COMPRESSIONS AND RESPIRATORY ASSISTANCE WERE STARTED AND CODE PINK CALLED AT LESS THAN ONE MINUTE OF LIFE WITH APGARS 0/0/0/2. THE INFANT WAS RESUSCITATED AND TRANSPORTED VIA CRITICAL CARE TEAM TO OUTSIDE FACILITY NICU. THE INFANT EXPIRED ON (B)(6) 2014, AT 06:03. | 1
3546655 | THE DEVICE WAS NOT RETURNED TO OLYMPUS FOR EVALUATION. HOWEVER, THE USER FACILITY EXAMINED THE COLONOSCOPE AND STATED IT WAS FUNCTIONING NORMALLY. THE DEVICE INSTRUMENT HISTORY WAS REVIEWED AND IT WAS NOTED THAT THE DEVICE WAS LAST SERVICED IN (B)(4) 2012. THE EXACT CAUSE OF THE PATIENT'S OUTCOME CANNOT BE CONCLUSIVELY DETERMINED.><OLYMPUS WAS INFORMED THAT DURING DIAGNOSTIC COLONOSCOPY, A TEAR OCCURRED IN THE SIGMOID COLON. THE MILD BLEEDING WAS CONTROLLED WITH CLIPS. NO ADDITIONAL INFORMATION WAS PROVIDED. | 1
3652379 | THE DEVICE WAS STERILIZED AT THE USER FACILITY PRIOR TO BEING RETURNED TO OLYMPUS FOR EVAL. THE BIOPSY CHANNEL AND THE SUCTION CHANNEL WERE CHECKED WITH A BOROSCOPE, AND WHITE RESIDUE AND DEBRIS WERE FOUND THROUGHOUT. IN ADDITION, THE DISTAL END COVER WAS CRACKED WHICH CAUSED THE DEVICE TO FAIL THE INSULATION TEST. THE DEVICE WAS SERVICED AND RETURNED TO THE USER FACILITY. THE EXACT CAUSE OF THE USER'S REPORT COULD NOT BE CONCLUSIVELY DETERMINED. THIS REPORT WILL BE SUPPLEMENTED IF ADD'L AND RELEVANT INFO BECOMES AVAILABLE AT A LATER TIME. AN OLYMPUS ENDOSCOPY SUPPORT SPECIALIST (ESS) WAS DISPATCHED TO THE USER FACILITY TO OBSERVE THE USER FACILITY'S REPROCESSING PRACTICES. NO REPROCESSING DEVIATIONS WERE NOTED. PLEASE CROSS-REFERENCE THE FOLLOWING REPORTS FOR THE OTHER FOUR COLONOSCOPIES: 2951238-2014-00034, 2951238-2014-00035, 2951238-2014-00036 AND 2951238-2014-00037.><OLYMPUS WAS INFORMED THAT FIVE COLONOSCOPIES TESTED POSITIVE FOR KLEBSIELLA, CITROBACTER, PSEUDOMONAS, STENOTROPHOMONAS, E.COLI, ENTEROCOCCUS AND OTHER NON-SPECIFIED GRAM NEGATIVE BACTERIA. THE USER FACILITY CULTURED THE DEVICE AS PART OF THEIR MICROBIOLOGICAL SURVEILLANCE TESTING. THE DEVICE WAS REPROCESSED IN A NON-OLYMPUS AUTOMATIC ENDOSCOPE REPROCESSOR PRIOR TO CONDUCTING THE TEST. THERE WAS NO PT INFECTION ASSOCIATED WITH THIS REPORT. | 1
4323028 | WE DISCOVERED TWO SIGNIFICANT PT SAFETY ISSUES THAT COULD CAUSE A PROVIDER TO ORDER THE WRONG TYPE OF INSULIN. THE FIRST CONCERNS A DESIGN FLAW THAT ALLOWS ONE TO ERRONEOUSLY ORDER U-500 INSULIN IN OUR AMBULATORY CARE ELECTRONIC MEDICAL RECORD SYSTEM. CURRENTLY, WHEN ONE GOES TO ORDER REGULAR INSULIN IN OUR AMBULATORY CARE SYSTEM (ECLINICALWORKS OR ECW), BOTH INSULIN R U-100 AND INSULIN R U-500 POP UP. UPON TYPING IN INSULIN R" ONE GETS TWO OPTIONS: "INSULIN REGULAR HUMAN" AND "INSULIN REGULAR HUMAN CON". PRESUMABLY THE PART THAT IS CUT OFF REFERS TO CONCENTRATED TO INDICATE IT IS THE U-500 INSULIN. HOWEVER "CON" COULD REFER TO ANY NUMBER OF THINGS AND IN ANY CASE THAT TERMINOLOGY IS NON-STANDARD AND NON-SPECIFIC AS TO WHAT THAT OPTION IS. THUS NOT ONLY IS IT NOT CLEARLY LABELED OR CLEAR WHAT THE SECOND OPTION IS | 1
3987623 | ><THIS REPORT WAS INCORRECTLY FILED AS MDR MFR. REPORT #2124823-2013-00022. THE CORRECT MDR MFR. REPORT # IS 9617277-2014-00003. PATIENT DATA NOT AVAILABLE. THE COPY" BUTTON ON THE MAC5500 IS INTENDED ONLY TO PRINT A COPY OF THE ECG LAST TAKEN | 1
4046855 | THE VITAL SIGNS ON THIS PATIENT WERE ELECTRONICALLY SENT TO THE EHR. THE HEART RATE AT REST WAS BETWEEN 100 AND 120 BPM WITH END STAGE HEART FUNCTION AND SOME TYPE OF RESPIRATORY SYMPTOMS. ASSOCIATED WITH THE RUSH TO BEAT THE 23 HOUR OBSERVATION DEADLINE, THE PATIENT WAS DISCHARGED FROM THE HOSPITAL WITHOUT CONSIDERATION OF THE RESTING TACHYCARDIA, BECAUSE THE NUMBER WAS SILENT IN THE VITAL SIGNS SILO OUT OF SIGHT AND OUT OF MIND. THE PATIENT WENT ON TO SUFFER A MYOCARDIAL INFARCTION, ASSOCIATED WITH THE FAILURE TO MANAGE THE SILENT" TACHYCARDIA. JUST WHO RECEIVES AND ADDRESSES THE VITAL SIGNS IN ELECTRONIC DATA SYSTEMS? THE DEVICE | 1
3955738 | ONE DAY AFTER ADMISSION TO AN ACUTE CARE HOSPITAL FOR TREATMENT OF SEPSIS AND ENDOCARDITIS, THE PT INFORMED STAFF THAT SHE WAS BEING STALKED AND WANTED TO BE IDENTIFIED BY AN ALIAS. A NEW NAME WAS ENTERED, REGISTERED, AND ESTABLISHED FOR HER EHR AND MDDS. HOWEVER, THE DIGITAL RECORDS OF THE FIRST DAY IN THE HOSPITAL UNDER HER REAL NAME WERE NO LONGER ACCESSIBLE. WHEN THE NAME ENTERED ON THE EHR ON ADMISSION WAS SEARCHED, THE NAME ENTERED TO THE ALIAS, AND THE RECORDS AND DATA IN THE MDDS APPEARED BUT STARTED ON DAY TWO. INACCESSIBLE WERE THE RECORDS AND DATA IN THE EHR AND MDDS FROM DAY ONE, COMPROMISING CARE AND UNDERSTANDING OF THE CASE. EHR SYSTEMS HAVE FAILED THIS PT, AND FAIL TO ENABLE FLEXIBILITY IN IDENTIFYING PTS IN UNIQUE CIRCUMSTANCES. | 1
3999615 | THE PT WITH CONGESTIVE HEART FAILURE WITH ARRHYTHMIAS WENT FOR EVAL AT ER AFTER INGESTING EXCESS DIURETICS. THE DOCTOR IN THE ER WAS OF THE IMPRESSION THAT THE POTASSIUM WAS NORMAL. THE EHR MDDS DID NOT ADEQUATELY DISCLOSE THAT THE SAMPLE OF BLOOD WAS HEMOLYZED, FALSELY ELEVATING THE POTASSIUM, PROVIDING THE ILLUSION IT WAS NORMAL, WHEN IT WAS SEVERELY LOW. THERE WAS A TINY ASTERISK IN THE CORNER OF THE MDDS GRID, REQUIRING A CLICK TO FIND OUT THAT THE RESULT WAS SPURIOUS. | 1
4006278 | DATA IS OFTEN BEING DEPOSITED IN THE EHR IN SILENCE. NO ONE BECOMES AWARE OF NEW DATA ON PTS. COMMUNICATION IS STUNTED. THERE IS A COGNITIVE FOCUS ON THE POORLY USABLE EHRS ASSOCIATED WITH LAPSES IN ATTENTION TO THE VITAL ISSUES SUCH AS VITAL SIGNS WHEN THE EHR SAPS ATTENTION OF THE USER. IN THIS CASE, THE HEART RATE WAS 135 BEATS PER MINUTE FOR MORE THAN 12 HOURS WITHOUT COMMUNICATION TO DOCTOR STAFF. AIDES OBTAIN THE VITAL SIGNS BY MACHINE THAT TRANSMIT DIRECTLY TO THE EHR. WHO KNEW. NO ONE CALLED THE DOCTOR, NO ONE KNEW OR PROCESSED THE CRITICAL SIGNS. THE PT HAD DELIRIUM FROM HYPOTENSION REQUIRING FLUID RESUSCITATION. THE POINT OF THIS IS THAT THESE SYSTEMS OF DEVICES DISTRACT AND DISRUPT COGNITIVE FUNCTION OF THE USER BECAUSE OF HUMAN FACTOR DEFICIENCY IN DESIGN, CREATING, IN TURN, FERTILE FIELD FOR LIFE THREATENING COMPLICATIONS. | 1
4317208 | RECENTLY, IT CAME OT OUR ATTENTION THE BPOC SYSTEM HAS A GAPING HOLE RELATIVE TO THE SAFE ADMINISTRATION OF IV MEDICATIONS. WE BELIEVE THE REPORTED ISSUE MAY HAVE HAPPENED WITH AN UPGRADE BUT CAN NOT BE CERTAIN. WHEN PREPARING COMPOUNDED IV MEDICATIONS, THE COMPONENTS OF THE IV ARE BARCODE VERIFIED AGAINST THE CPOE ORDER TO ENSURE CORRECT PRODUCT SELECTION PRIOR TO MIXING. THIS SCAN INCLUDES THE CARRIER SOLUTION AND ANY ADDITIVES. THIS SYSTEM HAS BEEN RELIABLE SINCE IMPLEMENTATION 7 YEARS AGO. AT THE BEDSIDE, THIS NURSE IS EXPECTED TO SCAN THE BARCODE ON THE PATIENT LABEL WHICH REFERENCES THE SPECIFIC RX # OF THE ORDER. THIS MATCHES THE CORRECT IV PRODUCT DIRECTLY TO THE PATIENT. HOWEVER, WE WERE RECENTLY MADE AWARE WHEN A NURSE SCANS THE CARRIER SOLUTION BARCODE AT THE BEDSIDE HE/SHE RECEIVES POSITIVE CONFIRMATION OF THE CORRECT PRODUCT. WHAT THIS MEANS IS EVERY PRODUCT MIXED IN 100 ML OF 0.9% NORMAL SALINE, FOR EXAMPLE, WOULD GIVE THE SAME FEEDBACK IF THE CARRIER SOLUTION BARCODE IS SCANNED. AS A QUICK FIX, WE HAVE STARTED TO OBSCURE THE CARRIER SOLUTION BARCODE SO THE NURSES ONLY HAVE ONE BARCODE ON THE PRODUCT, THAT BEING THE RX # BARCODE. WE ARE SUBMITTING A PATIENT SAFETY TASK TO OUR EHR VENDOR AS WELL. THIS BEGS THE QUESTION, WHAT ARE WE MISSING DURING TESTING FOR UPGRADES AND HOW MANY OTHER LATENT ERRORS ARE IN THE EHR SYSTEMS. (B)(6). | 1
3949787 | THE PT WAS EXPERIENCING A METABOLIC DIURESIS. THE DEGREE TO WHICH THE PT WITH HEART DYSFUNCTION NEEDED INTRAVENOUS FLUIDS WAS DEPENDENT ON THE DEGREE OF ORTHOSTATIC HYPOTENSION, IF ANY. THE ORDER TO OBTAIN THESE BLOOD PRESSURES WAS ENTERED. THE TECH USING A BP MACHINE HAD THEM ENTERED IN TO THE MDDS EHR, BUT THERE WAS NOT ANY ALERT THAT THEY WERE ENTERED OR OBTAINED. THE NURSE DID NOTHING DESPITE WHAT WAS A GREATER THAN 50 MMHG DROP IN SYSTOLIC BLOOD PRESSURE. THE SILENT SILO REARS ITS UGLY HEAD AGAIN. THERE IS AN INTRINSIC DEFECT IN THE DEVICE IN THAT IT FAILS TO WARN THAT NEW DATA WAS ENTERED OR ARRIVED. | 1
4270836 | ><PT IS A PHYSICIAN ATTEMPTING TO USE E-CLINICAL WORKS TO ACHIEVE MEANINGFUL USE" BY USING MANDATED "PATIENT PORTAL". THE PT PORTAL DOES NOT WORK | 1
4253525 | DELAY OF CASE BECAUSE VOLCANO EAGLE EYE CATHETER FAILED DURING RECOGNITION TO PIM PHASE. CATHETER BLACKED OUT DURING AN IVUS RUN. CATHETER WAS NOT RECOGNIZED BY PIM. NO HARM TO PATIENT. | 1
3690255 | THIS MEDIAL DEVICE REPORT IS BEING SUBMITTED DUE TO A RECENT FDA INSPECTION HELD AT AGFA'S (B)(6) LOCATION. IT WAS DETERMINED THAT 17 ADDITIONAL OCCURRENCES IDENTIFIED WITH THE SAME EVENT ISSUE AS DESCRIBED IN MDR REPORT , FDA # 9616389-2013-00003 REQUIRED REPORTING TO DOCUMENT THE ADDITIONAL OCCURRENCES. THIS REPORT IS FOR A 13TH EVENT FOR THE OVERALL ISSUE, AND THE 2ND OCCURRENCE FROM THE SAME SITE REPORTED IN 9616389-2014-00011. FOR THIS NEW OCCURRENCE, AN AGFA ENGINEER REPORTED ON (B)(6), 2013, THAT THE SITE HAD EXPERIENCED THEIR DX-D100 UNIT EXHIBITING INTERMITTENT MOVEMENT AGAIN WHEN DRIVING THE DX-D100 UNIT. THE AGFA ENGINEER REPLACED STRAIN GAUGES AND ADJUSTED VOLTAGES AND THE UNIT IS WORKING SMOOTHLY. NO HARM WAS REPORTED FOR THIS EVENT. AGFA INVESTIGATION WAS ALREADY UNDERWAY FOR A REPORTABLE CORRECTION TO THE FDA ON MAY 15, 2013: FDA REFERENCE # Z-1487-13. FOR THE CORRECTIONS, AGFA IMPLEMENTED MANDATORY SERVICE BULLETINS IN WHICH ALL AFFECTED UNITS WERE REPLACED WITH NEW FIRMWARE AND NEW DIGITAL MOTION CONTROL BOARDS. ALL OTHER DOCUMENTATION FOR THE DX-D100 REPORTABLE CORRECTION WILL BE REPORTED VIA FDA Z-1487-13.><. | 1
3895939 | THIS MEDIAL DEVICE REPORT IS BEING SUBMITTED DUE TO AN ADDITIONAL EVENT THAT WAS ORIGINALLY REPORTED IN MDR REPORT, FDA # 9616389-2013-00003. THIS EVENT OCCURRED AT THE SAME SITE AS FDA MDR 9616389-2014-00024. THIS REPORT IS FOR A 20TH EVENT FOR THE ORIGINAL REPORTED MDR EVENT, IN WHICH THE CUSTOMER INFORMED AGFA ON (B)(6) 2014, THAT THE SITE HAD EXPERIENCED THEIR DX-D100 UNIT EXHIBITING ERRATIC MOVEMENT WHEN DRIVING THE UNIT. THE DX-D100 UNIT WAS NEWLY INSTALLED AT THE CUSTOMER SITE AND THIS EVENT OCCURRED ON THE FIRST DAY OF USE WITH THE NEW UNIT. THE CUSTOMER EXPERIENCED INTERMITTENT MOVEMENT WITH THE UNIT MOVING IN THE OPPOSITE DIRECTION. WHEN THE TECHNICIAN WAS BACKING THE UNIT UP AND LET GO OF THE HANDLE THE UNIT WOULD SLIGHTLY MOVE FORWARD. WHEN THE TECHNICIAN WAS MOVING THE UNIT FORWARD AND LET GO OF THE HANDLE THE UNIT WOULD SLIGHTLY MOVE BACKWARDS. NO HARM HAS BEEN REPORTED FOR THIS EVENT AND THE CUSTOMER, AT THIS TIME, HAS ELECTED TO CONTINUE TO USE THE NEWLY INSTALLED UNIT. AGFA INVESTIGATION WAS ALREADY UNDERWAY FOR A REPORTABLE CORRECTION TO THE FDA ON MAY 15, 2013: FDA REFERENCE # Z-1487-13. FOR ANY POTENTIAL CORRECTIONS FOR THIS DX-D 100 UNIT, DOCUMENTATION WILL BE PROVIDED VIA FDA Z-1487-13.><. | 1
3690270 | THIS MEDIAL DEVICE REPORT IS BEING SUBMITTED DUE TO A RECENT FDA INSPECTION HELD AT AGFA'S (B)(4) LOCATION. IT WAS DETERMINED THAT 17 ADDITIONAL OCCURRENCES IDENTIFIED WITH THE SAME EVENT ISSUE AS DESCRIBED IN MDR REPORT, FDA # 9616389-2013-00003, BUT AT DIFFERENT SITES, REQUIRED REPORTING TO DOCUMENT THE ADDITIONAL OCCURRENCES. THIS REPORT IS FOR A 10TH EVENT IN WHICH A CUSTOMER ENGINEER INFORMED AGFA ON (B)(4) 2013, THAT HE HAD PERFORMED THE DX-D100 LATEST VERSION UPGRADE OF THE FIRMWARE, PER A MANDATORY SERVICE BULLETIN BY AGFA, AND CHECKED THE GAUGES, VOLTAGE AND ENCODER CONNECTIONS. THE UNIT WORKED PERFECTLY THE FIRST DAY, BUT WITHIN A FEW DAYS THE UNIT EXHIBITED ERRATIC MOVEMENT. AS THIS UNIT WAS INSTALLED NEARLY TWO MONTHS PRIOR TO THIS EVENT, THE UNIT WAS PULLED FROM USE AND A BACKUP SYSTEM USED BY THE SITE. AGFA DETERMINED THERE WAS A DEFECTIVE PART FOR THE DIGITAL MOTION CONTROL BOARD; THEREFORE, AFTER WORKING WITH SEDECAL, THE VENDOR SUPPLYING THE REPLACEMENT PARTS, AGFA REPLACED THE DIGITAL MOTION CONTROL BOARD, GAUGES AND ADJUSTED VOLTAGES WITH GOOD PARTS. THE DX-D100 UNIT FOR THIS SITE IS NOW WORKING AS EXPECTED. NO HARM HAS BEEN REPORTED FOR THIS EVENT. AGFA INVESTIGATION WAS ALREADY UNDERWAY FOR A REPORTABLE CORRECTION TO THE FDA ON (B)(4). FOR THE CORRECTIONS, AGFA IMPLEMENTED MANDATORY SERVICE BULLETINS IN WHICH ALL AFFECTED UNITS WERE TO BE REPLACED WITH NEW FIRMWARE AND NEW DIGITAL MOTION CONTROL BOARDS. ALL OTHER DOCUMENTATION FOR THE DX-D100 REPORTABLE CORRECTION WILL BE REPORTED VIA (B)(4).><. | 1
3690227 | THIS MEDIAL DEVICE REPORT IS BEING SUBMITTED DUE TO A RECENT FDA INSPECTION HELD AT AGFA'S (B)(6) LOCATION. IT WAS DETERMINED THAT 17 ADDITIONAL OCCURRENCES IDENTIFIED WITH THE SAME EVENT ISSUE AS DESCRIBED IN MDR REPORT , FDA # 9616389-2013-00003, REQUIRED REPORTING TO DOCUMENT THE ADDITIONAL OCCURRENCES. THIS REPORT IS A 2ND OCCURRENCE FOR THE SAME SITE AS REPORTED IN 9616389-2013-00003, AND THE 7TH EVENT OVERALL, REPORTED ON (B)(6), 2013, IN WHICH A DIFFERENT DX-D100 UNIT EXHIBITED SPORADIC AND UNINTENDED MOVEMENT WHEN DRIVING THE UNIT. THE AGFA CSE REPLACED THE DIGITAL MOTION CONTROL BOARD, GAUGES AND ADJUSTED VOLTAGES . THE DX-D100 UNIT FOR THIS SITE IS NOW WORKING AS EXPECTED. NO HARM HAS BEEN REPORTED FOR THIS EVENT. AGFA INVESTIGATION WAS ALREADY UNDERWAY FOR A REPORTABLE CORRECTION TO THE FDA ON MAY 15, 2013: FDA REFERENCE # Z-1487-13. FOR THE CORRECTIONS, AGFA IMPLEMENTED MANDATORY SERVICE BULLETINS IN WHICH ALL AFFECTED UNITS WERE TO BE REPLACED WITH NEW FIRMWARE AND NEW DIGITAL MOTION CONTROL BOARDS. ALL OTHER DOCUMENTATION FOR THE DX-D100 REPORTABLE CORRECTION WILL BE REPORTED VIA FDA Z-1487-13.><. | 1
3446316 | (B)(4). THE DEFECTIVE DOWNLOADER WAS INADVERTENTLY DISPOSED OF AT THE CUSTOMER SITE. THEREFORE NO PRODUCT IS AVAILABLE FOR FAILURE ANALYSIS AND AN INVESTIGATION WILL BE DEEMED INCONCLUSIVE.><ON (B)(6) 2013, ABBOTT POINT OF CARE WAS CONTACTED BY A CUSTOMER WHO REPORTED THAT A SERIAL DOWNLOADER SMOKED WHEN A POWER CORD WAS PLUGGED INTO THE DEVICE. THE CUSTOMER STATES THAT THEY NOTICED THE POWER LIGHT BLINKING AND WHEN THEY REMOVED THE ANALYZER FROM DOWNLOADER AND REPLACED IT THE DOWNLOADER MADE CLICKING NOISE AND THEN MADE SMOKE. NO OTHER PRODUCTS WERE AFFECTED AND THE DOWNLOADER WAS REPLACED BY THE APOC DISTRIBUTOR. THE CUSTOMER STATES THAT THE DEFECTIVE DOWNLOADER WAS DISPOSED OF AND THE SERIAL NUMBER IS UNKNOWN. THE PRODUCT IS NOT AVAILABLE FOR INVESTIGATION AND THE CAUSE IS UNDETERMINED. BASED ON THE INFORMATION AVAILABLE AT THE TIME, THERE WERE NO INJURIES ASSOCIATED WITH THIS EVENT. | 1
3823100 | OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) WAS INFORMED THAT ONCE A SNARE OF THE SUBJECT EMR KIT OPENED AND IT DID NOT OPEN AGAIN. THE DOCTOR ENDED UP USING 7 EMR KIT. THERE WAS NO PATIENT INJURY REGARDING THIS REPORT.><THE SUBJECT DEVICE WAS NOT RETURNED TO OMSC FOR INVESTIGATION SINCE THE USER FACILITY DISCARDED THE DEVICE. THE EXACT CAUSE OF THE USER'S EXPERIENCE COULD NOT BE CONCLUSIVELY DETERMINED. THIS REPORT IS BEING SUBMITTED AS A MEDICAL DEVICE REPORT IN AN ABUNDANCE OF CAUTION. | 1
3554195 | THE DISTRIBUTOR REPORTED THAT A PIECE OF PACKAGING WAS IDENTIFIED IN THE SEAL OF THE POUCH DURING THEIR INITIAL INSPECTION OF REC'D PRODUCT. THE DEFECT CREATED A BREACH IN THE SEAL. THE DEVICE WAS NOT SENT TO A USER FACILITY.><DEVICE EVALUATION: ONE UNUSED SUSPECT DEVICE WAS RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD DID NOT REVEAL ANY EXCEPTION DOCUMENTS. THE COMPLAINT DATABASE WAS REVIEWED AND FOUND NO SIMILAR COMPLAINTS FOR THIS LOT NUMBER. A VISUAL INSPECTION OF THE RETURNED DEVICE FOUND THAT THE POUCH HAD FRAGMENTS OF MATERIAL CAUGHT IN THE VENDOR SEALED AREA CREATING AN AIR PASSAGE. THE FRAGMENTS APPEAR TO BE TYVEK MATERIAL NOTCHES FROM THE VENDOR SEALING PROCESS. MERIT BELIEVES THIS TO BE AN ISOLATED EVENT AND HAS INFORMED THE SUPPLIER. | 1
4222497 | A PRE-TEEN FEMALE WITH A HISTORY OF HYPOPLASTIC LEFT HEART SYNDROME WHO UNDERWENT STAGED PALLIATION WITH A NORWOOD PROCEDURE FOLLOWED BY A MODIFIED BLALOCK-TAUSSIG SHUNT AND THEN FONTAN PALLIATION WITH A 20 MM EXTRACARDIAC FONTAN. PATIENT HAD PERSISTENT CYANOSIS FOR WHICH SHE HAD CARDIAC CATHETERIZATION AND EXTENSIVE IMMOBILIZATION OF VENOATRIAL COLLATERAL VESSELS. PATIENT AGAIN HAD DECREASED OXYGEN SATURATIONS WITH ACTIVITY AND SHE PRESENTED A FEW WEEKS AGO FOR CARDIAC CATHETERIZATION AND POSSIBLE COIL EMBOLIZATION OF COLLATERAL VESSELS. DURING THE PROCEDURE, ANGIOGRAM WAS ATTEMPTED USING A POWER INJECTOR. HIGH PRESSURE TUBING WAS USED WITH A RATING OF 1200 PSI TO DELIVER A DYE SOLUTION FOR THE ANGIOGRAM. DURING THE FIRST ATTEMPT, THE PRESSURE GOT ONLY TO 718 PSI WHEN THE TUBING'S LUER LOCK ASSEMBLY FRACTURED APART AT ITS BASE CAUSING DYE SOLUTION TO SPRAY OUT OF THE TUBING UNDER PRESSURE. A SECOND TUBING SET WAS QUICKLY SET UP AND REACHED ONLY 612 PSI BEFORE THE TUBING'S LUER LOCK ASSEMBLY ALSO FRACTURED APART AT THE BASE. A THIRD SET OF TUBING WAS OBTAINED WHICH REACHED 748 PSI WITH THE SAME RESULTS.HIGH PRESSURE TUBING THAT COMES STOCKED THE CATH LAB TRAYS (PRE-PACKAGED WITH DRAPES, SUCTION TUBING, BOWLS AND OTHER ITEMS), BUT WHICH IS OFTEN UTILIZED INITIALLY OR IS UTILIZED IN A DIFFERENT CAPACITY WAS FOUND TO WORK APPROPRIATELY. IN THIS CASE, ANOTHER TRAY WAS OPENED TO REMOVE THAT TUBING, WHICH WAS SUCCESSFULLY USED FOR THE ANGIOGRAMS.THE CCL TEAM IDENTIFIED THAT THEY WERE HAVING PROBLEMS WITH THE BRAND OF HIGH PRESSURE TUBING THAT THEY ORDERED SEPARATELY FOR USE AS EXTRA OR BACK-UP TUBING. EACH PACKAGE CONTAINS INDIVIDUALLY PACKAGED TUBING. THE TEAM REMOVED ALL OF THIS TUBING FROM THEIR SUPPLY AND REPORTED THE INCIDENTS TO MATERIALS MANAGEMENT AND THE MANUFACTURER WAS NOTIFIED. THEY WERE RESEARCHING OTHER PRODUCTS TO USE AS BACK-UP TUBING FOR THE FUTURE. | 1
4151070 | MALE PRE-SCHOOLER WITH HISTORY OF HYPOPLASTIC LEFT HEART SYNDROME DIAGNOSED IN UTERO, WITH CORONARY ARTERY TO LEFT VENTRICULAR FISTULAE. HE HAD UNDERGONE MULTIPLE CARDIAC SURGERIES INCLUDING COIL CLOSURE OF THE FISTULAE, NORWOOD/SANO PROCEDURES, BIDIRECTIONAL GLENN PROCEDURE WITH BANDING OF THE RIGHT VENTRICLE TO PULMONARY ARTERY CONDUIT, ATRIAL SEPTECTOMY AND TRICUSPID VALVE REPAIR FOR TRICUSPID REGURGITATION. THE PATIENT WAS UNDERGOING A RIGHT AND LEFT HEART CATHETERIZATION PROCEDURE WITH ANGIOGRAPHY AND POSSIBLE COIL EMBOLIZATION OF SIGNIFICANT VENOVENOUS OR AORTOPULMONARY COLLATERAL VESSELS. THE HIGH-PRESSURE IV TUBING WAS HOOKED UP TO THE POWER-INJECTOR PUMP SYSTEM TO DELIVER THE DYE SOLUTION FOR ANGIOGRAPHY; THE PUMP REACHED 852 PSI WHEN THE TUBING FRACTURED AT THE LUER-LOCK CONNECTION TO THE PATIENT'S IV LINE. THE LUER LOCK ASSEMBLY CAME APART FROM THE BASE. A SECOND SET OF TUBING WAS PREPARED, AND WHEN INFUSION WAS STARTED, IT REACHED 677 PSI WHEN THE LUER CONNECTOR CAME APART AT THE SAME PLACE. TUBING IS LISTED AS CAPABLE OF ADMINISTERING UP TO 1200 PSI. | 1
3670319 | ORDER GIVEN TO INSTILL MEDICATION VIA THE RECTUM. STAFF CHOSE TO USE THIS DEVICE TO INSTILL THE MEDICATION. THE FIRST DOSE WAS ADMINISTERED THROUGH THE PORT LABELED IRRIG" WITH NO ISSUE. STAFF THEN ADMINISTERED SECOND DOSE THROUGH GREEN PORT LABELED "INF 45 ML". THIS IS THE BALLOON INFLATION PORT. STAFF MISINTERPRETED "INF" TO MEAN INFUSION. PATIENT WAS BROUGHT TO THE OPERATING ROOM FOR FREE AIR IN THE ABDOMEN AND WAS FOUND TO HAVE A RECTAL TEAR CAUSED BY THE OVERINFLATED BALLOON WHICH CONTAINED APPROX 530 ML OF FLUID." | 1
3535999 | A CUSTOMER REPORTED THAT ONE OF THE ANNOTATION TOOLS CONTAINED WITHIN THE DPS 1.1 SOFTWARE APPEARED TO MISCALCULATE THE AREA OF A REGION OF INTEREST ON A SCANNED SLIDE. INTERNAL INVESTIGATION CONFIRMED THAT THE ANNOTATION TOOLS USED WITHIN THE DPS VERSION 1.1 IMAGE VIEWER WILL DISPLAY INCORRECT MEASUREMENTS AND AREA CALCULATIONS FOR A REGION OF INTEREST ON A SCANNED SLIDE, UNDER CERTAIN USE CONDITIONS.><EVAL SUMMARY: ISSUE SUMMARY - A CUSTOMER REPORTED THAT ONE OF THE ANNOTATION TOOLS CONTAINED WITHIN THE DPS 1.1 SOFTWARE APPEARED TO MISCALCULATE THE AREA OF A REGION OF INTEREST ON A SCANNED SLIDE. EVAL - INVESTIGATION AT OUT COMPANY CONFIRMED THAT THE ANNOTATION TOOLS USED WITHIN THE DPS VERSION 1.1 IMAGE VIEWER WILL DISPLAY INCORRECT MEASUREMENTS AND AREA CALCULATIONS FOR A REGION OF INTEREST ON A SCANNED SLIDE, UNDER CERTAIN USE CONDITIONS. SOURCE CODE REVIEW IDENTIFIED WHERE THIS ERROR WAS INTRODUCED INTO DPS 1.1. CUSTOMER WAS INSTRUCTED NOT TO USE ANY AREA CALCULATIONS DISPLAYED BY THE VIEWER ANNOTATION TOOLS INCLUDING ALL SHAPES AND THE LINEAR MEASUREMENT TOOL UNTIL A SOLUTION COULD BE IDENTIFIED. CORRECTIVE ACTION - CUSTOMER WERE NOTIFIED OF A FIELD CORRECTIVE ACTION ON OR ABOUT (B)(4) 2013 AND WERE INSTRUCTED TO 1) DISCONTINUE USE OF THE MEASUREMENT AND AREA CALCULATION FEATURES OF THE ANNOTATION TOOLS (LINE, RECTANGLE/SQUARE, ELLIPSE/CIRCLE, AND FREEFORM) IN THE DPS IMAGE VIEWER; 2) VERIFY MEASUREMENTS AND CALCULATIONS ON GLASS SLIDES, IF ANY OF THE TOOLS HAVE BEEN USED FOR MEASUREMENTS OR AREA CALCULATIONS ON CLINICAL CASES; AND 3) NOTIFY ANY OTHER USERS OF THE SOFTWARE AT THEIR SITE OF THIS ISSUE. A NEW VERSION (1.2) OF DPS SOFTWARE IS CURRENTLY UNDER DEVELOPMENT. ENGINEERING IS REVIEWING AND REVISING TEST CASES TO ENSURE APPROPRIATE TEST CASE COVERAGE PRIOR TO EXECUTING FORMAL VERIFICATION TESTING. CUSTOMERS WILL BE PROVIDED WITH AN UPDATED VERSION OF THE SOFTWARE WHEN IT BECOMES AVAILABLE. THIS CORRECTION WAS REPORTED TO THE FDA (B)(4) DISTRICT OFFICE ON (B)(4) 2013 AS MEDICAL DEVICE CORRECTION AND REMOVAL REPORT NUMBER 3008491142-10/15/13-001-C. | 1
3528372 | TECHNOLOGIST WAS STANDING BEHIND THE ACQUISITION WORK STATION (AWS) FACING THE X-RAY SHIELD AND THE PT WAS STANDING FACING THE WALL TO THE RIGHT OF THE GANTRY. THE TECHNOLOGIST TURNED TO HER RIGHT SLIGHTLY WITH HER HEAD DOWN TO HANG AN APRON AND THE PT WAS STILL IN THE SAME POSITION WHEN THE X-RAY SHIELD SUDDENLY SHATTERED. THE TECHNOLOGIST AND PT STATED THEY DID NOT TOUCH OR BUMP THE GLASS IN ANY WAY PRIOR TO THIS EVENT. THE TECHNOLOGIST REC'D SOME MINOR LACERATIONS ON HER ARM, HANDS, AND NECK AT THE SHOULDER. THE TECHNOLOGIST WAS NOT TREATED BY A PHYSICIAN, BUT WAS LOOKED AT BY THE NURSE THAT IS ON STAFF AT THE SITE. THE PT DID NOT RECEIVE ANY INJURIES.><AN HOLOGIC FIELD ENGINEER WAS DISPATCHED TO THE SITE TO INSPECT THE ACQUISITION WORK STATION (AWS) AND ENVIRONMENT. THE FIELD ENGINEER FOUND NO INDICATION OF WHAT CAUSED THE EVENT. THE SYSTEM HAS BEEN INSTALLED AT THE SITE SINCE (B)(4) 2011. THE FIELD ENGINEER REPLACED THE X-RAY SHIELD ON THE SYSTEM IN QUESTION AND INSPECTED THE X-RAY SHIELDS ON THE FIFTEEN REMAINING SELENIA DIMENSIONS SYSTEMS AT THIS SITE FOR STRESS, CHIPS OR CRACKS. NONE WERE FOUND. THERE MOUNTING HARDWARE WAS ALSO INSPECTED AND IT WAS DETERMINED THEY ALL HAD THE CORRECT COMPONENTS AND WERE PROPERLY MOUNTED TO THE AWS. THE X-RAY SHIELDS ARE MADE OF TEMPERED GLASS (SAFETY GLASS). TEMPERED GLASS BREAKS INTO SMALLER IRREGULAR SHAPED PIECES INSTEAD OF SHARP SHARDS AS WITH NON-TEMPERED GLASS. SAMPLES OF THE BROKEN X-RAY SHIELD WERE SENT TO THE GLASS MFR FOR EVAL. IT WAS CONFIRMED THE PIECES WERE TEMPERED GLASS, BUT THE CAUSE OF THE FAILURE COULD NOT BE DETERMINED. AT THIS TIME, HOLOGIC HAS NOT IDENTIFIED A CAUSE FOR THIS INCIDENT, BUT CONTINUES TO WORK WITH THE GLASS MFR TO ENSURE THE X-RAY SHIELDS CONTINUE TO MEET SAFETY REQUIREMENTS. | 1
4105448 | COMPLAINT WAS RECEIVED FOR A SCANX DUO ((B)(4), DATE OF MFR, 01/21/2011). THE IMPLICATED SCANX DUO WAS MANUFACTURED 01/21/2011 AND HAS BEEN IN SERVICE FOR TWO YEARS. OBSERVATIONS: THE UNIT WAS TAMPERED WITH BECAUSE WE OBSERVED THAT THE BOTTOM FUSE TAB WAS MISSING. THE REAR COVER SCREWS WERE LOOSE. THE MAIN PCB WAS INSPECTED AND THERE WERE NO VISIBLE BURN MARKINGS OR BURNT COMPONENTS. THE INTERNAL POWER SUPPLY WAS REMOVED AND THERE WAS A BURN MARK ON THE MAIN CAVITY NEXT TO IT. THE IMPLICATED UNIT WAS IN USE FOR NEARLY TWO YEARS WHEN THE INTERNAL POWER SUPPLY FAILED, RENDERING THE UNIT INOPERABLE. THE FAILURE OF THE POWER SUPPLY IS DEEMED AN ISOLATED INCIDENT. | 1
4117910 | IT WAS UNKNOWN IF THE INITIAL REPORTER SENT A REPORT TO THE FDA. THE PRODUCT INVOLVED IN THE COMPLAINT WAS NOT RETURNED.><IT WAS UNKNOWN IF THE INITIAL REPORTER SENT A REPORT TO THE FDA.><PATIENT QUESTIONED THE BOLUS ADVICE ON HER PALM CENTRO BOLUS CALCULATOR. PATIENT REPORTED SHE CALCULATED A CARBOHYDRATE OF 15 GRAMS AND THE BOLUS CALCULATOR ADVISED A BOLUS OF 0.8 UNITS. PATIENT STATED HER CARBOHYDRATE RATIO IS 1 UNIT FOR 15 GRAMS OF CARBOHYDRATES. NO ADVERSE EVENT REPORTED. PRODUCT WAS REPLACED AND REQUESTED RETURN OF THE BOLUS CALCULATOR AND PUMP. | 1
4168993 | VIDEO SCANNER DID NOT PRINT IMAGES OF CASE IT WAS SUPPOSED TO. UNABLE TO PRINT AT ALL AND UNKNOWN WHAT CAUSED DEVICE NOT TO PRINT.======================MANUFACTURER RESPONSE FOR VIDEO SCANNER - PRINTER, MEDXCHANGER (PER SITE REPORTER).======================SALES REPRESENTATIVE TOLD US HOW TO RETRIEVE THE PICTURES. | 1
3979913 | PHYSICIAN WAS ORDERING INSULIN FOR PT AND WAS USING THE OUTPATIENT MEDICATION REVIEW AS A SOURCE FOR CONTINUING HER HOME REGIMEN. THE OUTPATIENT MEDICATION REVIEW POPULATES DRUGS FROM MULTUM DATABASE ((B)(4)) AND THE FORMATTING CANNOT BE CHANGED. THE PHYSICIAN USED THE CONCENTRATION OF INSULIN GLARGINE AND INSULIN LISPRO AS THE DOSE AND ORDERED INSULIN GLARGINE 100 UNITS TWICE DAILY AND INSULIN LISPRO 100 UNITS THREE TIMES DAILY BECAUSE OF THE WAY IT WAS FORMATTED AND DISPLAYED. THE ORDER WAS SUBSEQUENTLY VERIFIED, BUT THE NURSE NOTICED THE ERROR BEFORE THE PT RECEIVED AN OVERDOSE OF INSULIN. PRIOR TO USING THE OUTPATIENT MEDICATION REVIEW, OUR INSTITUTION HAD A DIFFERENT APPLICATION FOR HOME MEDICATIONS WHERE THE DESCRIPTION OF THE DRUGS WERE ABLE TO BE CHANGED. WE HAD TAKEN OUT THE CONCENTRATIONS OF THE INSULIN FROM THEIR DESCRIPTIONS BECAUSE IT BECAME A RECURRING ISSUE FOR THE PHYSICIANS TO ORDER THE CONCENTRATION RATHER THAN THE DOSE. NOW, WE ARE NOT ABLE TO EDIT THE DESCRIPTION OF THE DRUG IN THE MULTUM DATABASE AND HAVE NO CHOICE BUT TO USE IT. THIS HAS LEAD TO MORE ERRORS IN INSULIN PRESCRIBING. | 1
4064095 | SIEMENS HEALTHCARE DIAGNOSTICS HAS INVESTIGATED THIS ISSUE AND IDENTIFIED A TASK SCHEDULER ISSUE WITH THE CENTRALINK¿ DATA MANAGEMENT SYSTEM, SOFTWARE VERSION 14. THE ISSUE MAY CAUSE THE SOFTWARE TO STOP EXECUTING COMMANDS, INCLUDING UPLOADING VALIDATED RESULTS TO THE LIS. THE ISSUE IS RELATED TO AN INTERNAL SOFTWARE TIMER THAT OVERFLOWS AFTER 24 DAYS. AUTOMATED RULES TRIGGERED BY TASK SCHEDULERS OR BY ACTIONS IN THE CENTRALINK USER INTERFACE MAY ALSO BE AFFECTED. AS EACH CUSTOMER CONFIGURATION IS UNIQUE, IN ADDITION TO THE DELAY IN UPLOADING OF VALIDATED RESULTS, THE SPECIFIC CONSEQUENCES OF DISCONTINUED COMMAND EXECUTION WILL VARY. URGENT MEDICAL DEVICE CORRECTION (UMDC) 10819364 ENTITLED CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS IN THE UNITED STATES AND URGENT FIELD SAFETY NOTICE (UFSN) 10819365 ENTITLED "CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN AUGUST 2014. THE UMDC AND UFSN DESCRIBES THE ISSUE AND PROVIDES GUIDANCE ON HOW TO PREVENT THE ISSUE FROM OCCURRING.><ALL TASKS ON A CENTRALINK DATA MANAGEMENT SYSTEM STOPPED UNEXPECTEDLY. THERE ARE NO REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO ALL TASKS UNEXPECTEDLY STOPPING ON THE CENTRALINK DATA MANAGEMENT SYSTEM." | 1
4193325 | A SIEMENS TECHNICAL APPLICATION SPECIALIST (TAS) WAS DISPATCHED TO THE CUSTOMER SITE. THE TAS SPECIALIST DETERMINED THAT THE ISSUE WAS CAUSED BY A CHANGE IN THE REPORTING RANGES THAT THE TAS HAD PREVIOUSLY MADE. THE TAS CHANGED THE ANION GAP CALCULATION SO THAT IT WOULD NOT CALCULATE IF A RESULT HAD A COMPARATOR. THE TAS ALSO MADE CHANGES TO THE ADVIA 1800 LINEARITY AND CONFIGURED A LOW LIMIT FOR CO2. THE CAUSE OF THE ANION GAP BEING CALCULATED WHEN IT SHOULD NOT HAVE WAS DUE TO A CONFIGURATION ISSUE. THE CENTRALINK DATA MANAGEMENT SYSTEM IS PERFORMING ACCORDING TO SPECIFICATIONS. NO FURTHER EVALUATION IS REQUIRED.><THE OPERATOR OF A CENTRALINK DATA MANAGEMENT SYSTEM CONTACTED THE SIEMENS CUSTOMER CARE CENTER (CCC) BECAUSE AN ANION GAP HAD BEEN CALCULATED FOR A PATIENT SAMPLE WHEN IT SHOULD NOT HAVE BEEN CALCULATED. THE ADVIA 1800 INSTRUMENT HAD PRODUCED A RESULT OF 7.2 FOR CARBON DIOXIDE (CO2), WHICH WAS BELOW THE CO2 DETECTION LIMIT OF 10 ON THE CENTRALINK. THE CO2 RESULT WAS TRANSMITTED TO THE CENTRALINK AS <10 AND AN ANION GAP OF 25 WAS CALCULATED. BASED ON THE PATIENT'S SODIUM, CHLORIDE, AND CO2 VALUES, THE CUSTOMER KNEW THAT THE RESULT OF 25 WAS INCORRECT. THE SAMPLE WAS REPEATED FOR CO2 ON THE INSTRUMENT AND RESULTED AS 6.2, WHICH WAS ALSO TRANSMITTED TO THE CENTRALINK AS <10. THE CENTRALINK PRODUCED AN INCORRECT ANION GAP VALUE OF 25 A SECOND TIME, WHICH WAS REPORTED TO THE PHYSICIAN(S). THERE ARE NO KNOWN REPORTS OF PATIENT INTERVENTION OR ADVERSE HEALTH CONSEQUENCES DUE TO THE CENTRALINK DATA MANAGEMENT SYSTEM CALCULATING AN ANION GAP RESULT WHEN IT SHOULD NOT HAVE. | 1
4064790 | SIEMENS HEALTHCARE DIAGNOSTICS HAS INVESTIGATED THIS ISSUE AND IDENTIFIED A TASK SCHEDULER ISSUE WITH THE CENTRALINK¿ DATA MANAGEMENT SYSTEM, SOFTWARE VERSION 14. THE ISSUE MAY CAUSE THE SOFTWARE TO STOP EXECUTING COMMANDS, INCLUDING UPLOADING VALIDATED RESULTS TO THE LIS. THE ISSUE IS RELATED TO AN INTERNAL SOFTWARE TIMER THAT OVERFLOWS AFTER 24 DAYS. AUTOMATED RULES TRIGGERED BY TASK SCHEDULERS OR BY ACTIONS IN THE CENTRALINK USER INTERFACE MAY ALSO BE AFFECTED. AS EACH CUSTOMER CONFIGURATION IS UNIQUE, IN ADDITION TO THE DELAY IN UPLOADING OF VALIDATED RESULTS, THE SPECIFIC CONSEQUENCES OF DISCONTINUED COMMAND EXECUTION WILL VARY. URGENT MEDICAL DEVICE CORRECTION (UMDC) 10819364 ENTITLED CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS IN THE UNITED STATES AND URGENT FIELD SAFETY NOTICE (UFSN) 10819365 ENTITLED "CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN AUGUST 2014. THE UMDC AND UFSN DESCRIBES THE ISSUE AND PROVIDES GUIDANCE ON HOW TO PREVENT THE ISSUE FROM OCCURRING.><A SCHEDULED TASK ON A CENTRALINK DATA MANAGEMENT SYSTEM STOPPED. THERE ARE NO REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO A SCHEDULED TASK STOPPING ON THE CENTRALINK DATA MANAGEMENT SYSTEM." | 1
4064826 | A SCHEDULED TASK ON A CENTRALINK DATA MANAGEMENT SYSTEM STOPPED UNEXPECTEDLY. THE OPERATOR REBOOTED THE CENTRALINK SERVER. THERE ARE NO REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO A SCHEDULED TASK STOPPING UNEXPECTEDLY ON THE CENTRALINK DATA MANAGEMENT SYSTEM.><SIEMENS HEALTHCARE DIAGNOSTICS HAS INVESTIGATED THIS ISSUE AND IDENTIFIED A TASK SCHEDULER ISSUE WITH THE CENTRALINK¿ DATA MANAGEMENT SYSTEM, SOFTWARE VERSION 14. THE ISSUE MAY CAUSE THE SOFTWARE TO STOP EXECUTING COMMANDS, INCLUDING UPLOADING VALIDATED RESULTS TO THE LIS. THE ISSUE IS RELATED TO AN INTERNAL SOFTWARE TIMER THAT OVERFLOWS AFTER 24 DAYS. AUTOMATED RULES TRIGGERED BY TASK SCHEDULERS OR BY ACTIONS IN THE CENTRALINK USER INTERFACE MAY ALSO BE AFFECTED. AS EACH CUSTOMER CONFIGURATION IS UNIQUE, IN ADDITION TO THE DELAY IN UPLOADING OF VALIDATED RESULTS, THE SPECIFIC CONSEQUENCES OF DISCONTINUED COMMAND EXECUTION WILL VARY. URGENT MEDICAL DEVICE (B)(4) ENTITLED CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS IN THE UNITED STATES AND URGENT FIELD SAFETY NOTICE (UFSN) 10819365 ENTITLED "CENTRALINK DATA MANAGEMENT SYSTEM TASK SCHEDULER ISSUE" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN (B)(4) 2014. THE UMDC AND UFSN DESCRIBES THE ISSUE AND PROVIDES GUIDANCE ON HOW TO PREVENT THE ISSUE FROM OCCURRING." | 1
4188039 | IT WAS REPORTED THAT THE T:CONNECT LOGBOOK WAS DUPLICATING BOLUS AND CARBOHYDRATE DATA. THERE WAS NO IMPACT TO THE PATIENT.><IF NEW INFORMATION BECOMES AVAILABLE, A FOLLOW UP REPORT WILL BE SUBMITTED. | 1
3905124 | IT WAS REPORTED THAT THE DAILY INSULIN USE WAS A DIFFERENT VALUE ON THE DASHBOARD PAGE OF THE T:CONNECT DATA MANAGEMENT SYSTEM AS COMPARED WITH THE DAILY INSULIN USE VALUE ON THE ACTIVITY SUMMARY PAGE OF THE T:CONNECT DATA MANAGEMENT SYSTEM. | 1
3610683 | THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.><A CUSTOMER REPORTED THEIR ADC BLOOD GLUCOSE METER WAS NOT RETAINING THE DATE AND TIME SETTINGS. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT PROPERLY SET, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT. | 1
3962150 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THEIR METER WERE NOT DISPLAYING PROPERLY, AND THEY REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. NOTE: THE DEVICE MANUFACTURER DATE FOR THE SUSPECT DEVICE IS UNKNOWN. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.><THIS FOLLOW-UP SERVES AS ADDITIONAL INFORMATION. METER SERIAL NUMBER (B)(4) WAS RETURNED ON 01-AUG-2014. | 1
3565594 | A CUSTOMER REPORTED INCORRECT TIME/DATE ON THE DISPLAY OF THEIR ADC BLOOD GLUCOSE METER. THE CUSTOMER ADDITIONALLY REPORTED BEING A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 1
4130529 | A CUSTOMER REPORTED RECEIVING AN E6 MESSAGE ON THE DISPLAY OF HIS ADC BLOOD GLUCOSE METER. IT WAS THEN ADDITIONALLY IDENTIFIED BY ADC CUSTOMER SERVICE THAT THE DATE AND TIME SETTINGS IN THE METER WERE NOT PROPERLY SET, AND HE REPORTED TO BE A USER OF THE PRECISION LINK DATA MANAGEMENT SYSTEM. THERE WAS NO REPORT OF DEATH, SERIOUS INJURY OR MISTREATMENT ASSOCIATED WITH THIS EVENT.><THERE IS A KNOWN MALFUNCTION WITH THE PRECISION LINK SOFTWARE THAT CAN LEAD TO INCORRECT TRENDING OF RESULTS. THIS OCCURS WHEN RESULTS, OBTAINED ON A METER WITH INCORRECT DATE AND TIME, ARE UPLOADED TO A COMPUTER WITH PRECISION LINK SOFTWARE. CUSTOMERS AND RETAILERS HAVE BEEN NOTIFIED THROUGH THE ADC FA21DEC2006 LETTER. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED. | 1
3652806 | THE PALL DATA 5 SOFTWARE PRESENTED AN O2 LEVEL READING OF PASS" FOR A SAMPLE POUCH TEST RESULT. THE TEST RESULT VALUES WERE INDICATIVE OF AN INVALID TEST RESULT ACCORDING TO CBC STANDARD OPERATING PROCEDURE. UPON COMPLETION OF INVESTIGATIONS | 1
3806906 | AN ECHOCARDIOGRAM WAS ORDERED VIA CPOE TO EVALUATE A LOW CARDIAC OUTPUT STATE. THE UPDATE ON THE RESULTS SCREEN FOR IMAGING STATED THAT THE TEST WAS COMPLETED, AND THE TIME IT WAS COMPLETED. THERE IS AN OBLIGATE APPROXIMATE 24 HOURS DELAY FOR THE REPORT AND IMAGES TO APPEAR. AFTER 36 HOURS, THERE WAS NOT ANY RESULT AND AFTER 48 HOURS, A CALL WAS MADE TO THE ANCILLARY DEPARTMENT. THE IMAGING WAS NEVER DONE, EVEN THOUGH THE UPDATE STATED COMPLETE. THIS DELAY IN KNOWING CARDIAC FUNCTION STATUS WAS A CRITICAL ADVERSITY AND IMPEDIMENT TO THE CARE OF THIS PT WHO DIED A FEW DAYS LATER FROM AN INFECTION, THOUGH THIS DELAY WAS NOT THE CAUSE OF THE INFECTION. THE FAILURE OF ANCILLARY DEPARTMENTS TO GET ACCURATE ORDERS IS AN ONGOING PROBLEM WITH ALL CPOE SYSTEMS DUE TO INTERFACE AND INTEROPERABILITY ISSUES. THEY CAN NOT BE TRUSTED TO ENABLE TIMELY TESTING WITHOUT INTENSE VIGILANCE. | 1
3840183 | ORDERING IV FLUIDS IS A COMPLEX TASK ON THE CPE DEVICE. THE ORDERING SCREENS ARE NOISY AND SEEK MUCH IRRELEVANT INFORMATION WHICH SERVES TO DISTRACT AND CONFUSE THE USE AND THE TASK AT HAND. THE DIFFICULTY IN NAVIGATING THE ORDERING SCREENS RESULTS IN CLINICALLY COUNTER INTUITIVE INFUSIONS. THIS PATIENT WITH SEPSIS, HYPOTENSION, AND TACHYCARDIA WAS ORDERED FLUIDS AT 125 CC PER HOUR IN THE EMERGENCY WARD. LATER IN THE DAY, THE PATIENT BECAME HYPOTENSIVE AND TACHYCARDIC WITH A SPIKE OF FEVER. WHAT WAS NOT KNOWN BY THE CLINICAL TEAM WAS THAT IN THE TINY FONT OF THAT IV FLUID ORDER WAS THAT THE IV FLUIDS WERE TO STOP AUTOMATICALLY AFTER 3 LITRES. THE NURSES STOPPED THE FLUID DESPITE THE SEPSIS, BECAUSE THE CPOE DEVICE REQUIRED THAT, AND THERE WAS NOT ANY NOTIFICATION OR CLINICAL DECISION SUPPORT WARNING OF SAME. THE IV FLUIDS WERE RESTARTED AFTER THE DETERIORATION THAT REQUIRED TRANSFER TO ICU. THE STEALTH CESSATION OF DISEASE CRITICAL FLUID WAS LIFE THREATENING. THE COMPLEXITIES AND POOR USABILITY OF THE IV FLUID ORDER MERIT INVESTIGATION, SURVEILLANCE AND POSSIBLE RECALL. | 1
3805735 | AN ECHOCARDIOGRAM WAS ORDERED VIA CPOE TO EVALUATE A LOW CARDIAC OUTPUT STATE. THE UPDATE ON THE RESULTS SCREEN FOR IMAGING STATED THAT THE TEST WAS COMPLETED, AND THE TIME IT WAS COMPLETED. THERE IS AN OBLIGATE APPROXIMATE 24 HOUR DELAY FOR THE REPORT AND IMAGES TO APPEAR. AFTER 36 HOURS, THERE WAS NOT ANY RESULT AND AFTER 48 HOURS, A CALL WAS MADE TO THE ANCILLARY DEPARTMENT. THE IMAGING WAS NEVER DONE, EVEN THOUGH THE UPDATE STATED COMPLETE. THIS DELAY IN KNOWING CARDIAC FUNCTION STATUS WAS A CRITICAL ADVERSITY AND IMPEDIMENT TO THE CARE OF THIS PT WHO DIED A FEW DAYS LATER FROM AN INFECTION, THOUGH THIS DELAY WAS NOT THE CAUSE OF THE INFECTION. THE FAILURE OF ANCILLARY DEPARTMENTS TO GET ACCURATE ORDERS IS AN ONGOING PROBLEM WITH ALL CPOE SYSTEMS DUE TO INTERFACE AND INTEROPERABILITY ISSUES. THEY CAN NOT BE TRUSTED TO ENABLE TIMELY TESTING WITHOUT INTENSE VIGILANCE. | 1
4114511 | URINE TESTING AND CULTURE WAS ORDERED BUT NEITHER DID REACH THE RECIPIENT TO EXECUTE THE ORDER. THERE IS NOT A USER FRIENDLY PROGRAM TO RECONCILE THE ORDERS TO DETERMINE EXACTLY WHAT WAS DONE OR NOT DONE. DISEASE CRITICAL TESTS ARE DELAYED RESULTING IN DELAYS OF TREATMENT, THREATENING LIFE AND LIMB. | 1
3807187 | THE COMMUNICATIONS IN THIS HOSPITAL THAT HAS DEPLOYED CPOE AND EHR SYSTEMS HAVE BEEN HAMPERED EVER SINCE FOR INNUMERABLE REASONS, MOST SIGNIFICANTLY THE USER UNFRIENDLINESS OF THE CPOE AND EHR DEVICES. THE PATIENT WHO SUSTAINED INJURY WAS AT RISK FOR RESPIRATORY ARREST DUE TO A NARROWING IN THE TRACHEA. THERE IS NO PLACE ON THE EHR TO LIST SUCH A LIFE THREATENING CONDITION THAT WOULD BE VISIBLE TO EACH AND EVERY CARE TEAM MEMBER WHO OPENED THE EHR FOR THIS PATIENT. THE LACK OF COMMUNICATION BREEDS UNFAMILIARITY WITH THE SPEED AT WHICH THE PATIENT COULD DETERIORATE, IN ADDITION TO THE INTRINSIC DELAY IN GETTING TREATMENTS ORDERED BY THE CPOE SYSTEM. IT HAS BEEN REPORTED THAT PATIENTS ARE NEGLECTED WHEN THE NURSE AND CARE TEAM MEMBERS FOCUS ON THE COMPUTERS RATHER THAN THE PATIENTS, SOMETHING THAT IS COMMONPLACE CONSIDERING THE POOR USABILITY OF THE EHRS. THE PT IN QUESTION DEVELOPED ACUTE RESPIRATORY DISTRESS. CARE WAS DELAYED DUE TO THE ABOVE MENTIONED REASONS. THE PT SUSTAINED A COMPLETE RESPIRATORY ARREST THAT LED TO A CARDIAC ARREST AND ANOXIC BRAIN INJURY. THE EHR DEVICE IS DEFECTIVE IN THAT IT FAILS TO ENABLE EXPEDIENT COMMUNICATION OF THE CRITICAL LIFE THREATENING PROBLEMS OF EVERY PATIENT WHICH CAUSES NEGLECT AND DELAYS IN CARE WHEN LIFE IS ON THE LINE. | 1
3724172 | CPOE SYSTEM BY THIS VENDOR IS POORLY USABLE AND PROMOTES ERRORS. INTENDED RECIPIENT OF ORDERS, ESPECIALLY ANCILLARY SERVICE(S), IS CONFUSED BY THE EXCESS SUPERFLUOUS DESCRIPTION CONTAINED IN THE ORDER DESCRIPTION RESULTING IN WRONG TESTS OR WRONG TIMING OF TESTS. AN ARTERIAL BLOOD GAS WAS ORDERED TO BE PERFORMED PRIOR TO APPLICATION OF THE RESPIRATORY SUPPORT CALLED BIPAP, BUT THE ORDER WAS CONFUSING RESULTING IN THE BIPAP BEING APPLIED PRIOR TO DRAWING THE BLOOD GAS SAMPLE. THE RESULT WAS CLINICALLY MEANINGLESS AS IT DID NOT ACCURATELY DEPICT BASELINE CLINICAL CONDITION. DIAGNOSES AND CLINICAL STRATEGIES ARE DISRUPTED, TO THE DETERMINANT OF PT SAFETY, DIAGNOSES, AND OUTCOME, BECAUSE OF THE FAILURE OF THE CPOE TO CLEARLY DESCRIBED WHAT TESTS ARE TO BE DONE AND WHEN. THIS PROBLEM IS WIDESPREAD AND COMPROMISES CARE AND SAFETY ACROSS MOST CLINICAL SETTINGS. | 1
4045010 | PT HAD LOVENOX, ASPIRING, AND WARFARIN ORDERED BY (B)(4), WITH NOTES STATING THAT THE LOVENOX IS TO BE DISCONTINUED WHEN THE INR MEASURING WARFARIN ANTICOAGULATION REACHED ALEVE OF 2.0. THE INR HAD BEEN MORE THAN 3.0 FOR 4 DAYS YET PT CONTINUED TO RECEIVE ALL THREE POTENT DRUGS. THE PT DEVELOPED GASTROINTESTINAL BLEEDING FROM THE EXCESSIVE ANTICOAGULATION. DESPITE THE (B)(4) AND DECISION SUPPORT, THE DOCTORS, NURSES, AND PHARMACY CONTINUED TO PROVIDE THE MEDICATIONS. THE SYSTEM IS A FAILURE IN THAT IT DID NOT WARN OFF THE EXCESSIVE ANTICOAGULATION. | 1
4031429 | THE CLINICAL DECISION SUPPORT FAILED TO WARN OF AN ALLERGY LISTED UNDER THE NAME OF A COMBINATION DRUG, LOTREL". THE PT WAS ALLERGIC TO ENALAPRIL (AND ALL ACE INHIBITORS) DUE TO A COUGH. BUT WHO KNEW WHICH COMPONENT OF THE COMBINATION DRUG CAUSED THE ALLERGY? NO ONE. THE NEXT DAY AFTER RECEIVING THE AMLODIPINE COMPONENT OF LOTREL | 1
3955739 | IN A CASE WITH DETERIORATING KIDNEY FUNCTION, RESPIRATORY FAILURE, AND ALIMENTARY TRACT CANDIDIASIS, EVAL AND MANAGEMENT ORDERS WERE ENTERED VIA THE CPOE MACHINE FOR ANALGESIA, GENITO-URINARY FUNCTION, AND PULMONARY FUNCTION. THESE ORDERS WERE SPORADICALLY EXECUTED, COMPROMISING THE CARE OF THE PT AND DELAYING DIAGNOSIS. THERE IS CONFUSION AND BEFUDDLEMENT IN NURSES WHO MAY BE INADEQUATELY TRAINED ON THESE DEVICES OR OVERWHELMED BY THE EXCESS VERBIAGE ACCOMPANYING THE ORDERS ON THEIR SCREEN, LIMITING THEIR UNDERSTANDING OF WHAT NEEDS TO BE DONE AND WHEN. THE PT WAS NEGLECTED, SUSTAINING SEVERE HYPOXIA FURTHER COMPROMISING THE PT FUNCTION AND PROGNOSIS. | 1
3770927 | PT HAD HYPOTENSION, CARDIAC ENZYMES WERE ORDERED STAT USING THE CPOE MACHINE. THE ORDER LAY IN THE RECIPIENT'S SILO, FOR AN HOUR BEFORE BEING CARRIED OUT DELAYING DIAGNOSIS. THERE IS NOT ANY WARNING THAT NEW ORDERS HAVE ARRIVED, UBIQUITOUSLY DELAYING CARE. | 1
3999638 | PT WITH DEHYDRATION IN SETTING OF SEVERE NON ISCHEMIC CARDIOMYOPATHY WAS RECEIVING IV FLUIDS AT OVER 100 CC PER HOUR ORDERED BY CPOE. BECAUSE THE BNP WAS ELEVATED, THE IV FLUIDS WERE DISCONTINUED USING CPOE. THREE HOURS LATER, THE PT WAS STILL RECEIVING THE IV FLUIDS, CAUSING A FLUID OVERLOADED STATE. THE NURSE HAD NOT SEEN THE ORDER BECAUSE THE NURSE DID NOT KNOW THERE WAS A NEW ORDER. THE CPOE DEVICE PROVIDES NO WARNING THAT A NEW ORDER HAS BEEN ENTERED, AND THE NURSE HAD NOT LOGGED ON TO UPDATE THE NURSE'S TASK SCREEN FOR AT LEAST THREE HOURS. THAT THERE IS NOT A WARNING OR INDICATOR OF NEW ORDERS IS AN INTRINSIC DEFECT IN THE CPOE DEVICE THAT DE FACTO CAUSES DELAYS IN CARE. | 1
4114530 | SEVERAL ORDERS FOR INTRAVENOUS FLUIDS WERE ACTIVE ON THE PATIENT, SIMULTANEOUSLY. ONE OF THE ORDERS WAS FOR D5 +ALCOHOL 10 PERCENT TO BE RUN AT 50 CC/HR. THE SYSTEM OF DEVICES FAILS TO DETECT AND PREVENT LIFE THREATENING ABSURDITY, A SERIOUS DEFECT IN THEIR DESIGN. | 1
4114531 | THERE WAS A FOLEY BLADDER CATHETER IN THIS PATIENT. THE PATIENT HAD A URINARY TRACT INFECTION AND IT WAS DECIDED TO DISCONTINUE THE FOLEY CATHETER. HOWEVER, THE ORDER TO DISCONTINUE THE CATHETER COULD NOT BE ENTERED BECAUSE THERE WAS NOT EVER AN ORDER TO INSERT THE FOLEY CATHETER. THE RIGIDITY OF THE CPOE INSTRUMENTS IS EMBODIED IN THIS CASE, DELAYING TIMELY CARE WHILE SERVING AS AN IMPEDIMENT TO SAFE AND EFFICACIOUS CARE. NOTHING BUT DELAYS AND DISRUPTIONS WITH ONGOING RISK TO THE PATIENTS. | 1
3999613 | THE CPOE ORDERING SYSTEM FOR DOFETILIDE THAT A PT HAS BEEN ON AS AN OUT PATIENT CREATES INNUMERABLE IMPEDIMENTS FOR UNCLEAR REASONS. THE DYSFUNCTION SO ASSOCIATED RESULTS IN DELAY FOR THE PT RECEIVING THE DRUG. | 1
4031302 | ORDERS WERE ENTERED FOR LAB TESTS TO BE RUN ON BLOOD TO BE DRAWN THE NEXT MORNING. THAT WAS NOT DONE, NOR WAS THERE RECONCILIATION TO DETERMINE THAT IT WAS NOT DONE. THE HOSPITAL ADMINISTRATION GETS FUSSED UP AND RETALIATORY ABOUT THE FAILURES OF ITS INFRASTRUCTURE. THERE WAS AN INTERFACE FAILURE AND/OR A FAILURE OF INTEROPERABILITY. THE PATIENT WAS AFFLICTED WITH KIDNEY FAILURE, ANEMIA, NEUROLOGICAL DISEASE, HEART ARRHYTHMIAS AND OTHERS. TO NOT HAVE THE RESULTS IN A TIMELY MANNER PUT THE PATIENT AT UNDUE RISK. WHEN THE RESULTS WERE FINALLY KNOWN, ALTERATION IN THERAPY WAS REQUIRED. | 1
4031301 | CLINICAL LAB BLOOD TESTS FOR THE NEXT MORNING WERE ENTERED INTO THE CPOE SYSTEM TO FURTHER EVALUATE LOW WBC COUNT, RENAL FAILURE, ANEMIA, AND LOW SODIUM. THE BLOOD WAS NOT OBTAINED FOR TESTING DELAYING CARE AND PLACING THE PATIENT AT RISK FOR WORSENING UNDERLYING MEDICAL CONDITIONS. WHO KNEW OF THE NEGLECT? NO ONE, BECAUSE THERE IS NO RECONCILIATION OF ORDERS FOR EXECUTION BY DEVICE OR BY HUMAN. HUMAN REVIEW OF THE CPOE SILOS IS TEDIOUS AND TIME CONSUMING AND FRAUGHT WITH UNCERTAINTY, DUE TO THE INNUMERABLE LINES OF MEANINGFULLY MEANINGLESS GIBBERISH OBFUSCATING THE CRITICAL ELEMENTS. SUCH A NEAR MISS WILL ULTIMATELY RECUR AND THE ASSOCIATED DELAYS WILL CAUSE MEDICAL CATASTROPHE. | 1
4031300 | THERE IS A PROBLEM THAT ORDERS ENTERED ON THE CPOE SYSTEM DO NOT REACH THE RECIPIENT OR DO NOT GET EXECUTED FOR OTHER REASONS. THIS INCLUDES TASK LIST FAILURES, CAMOUFLAGED ORDERS AMIDST MANY OTHERS, AND THE FAILURE OF THE HEALTHCARE PROFESSIONAL RESPONSIBLE FOR EXECUTING THE ORDER TO RECEIVE THE TASK. IN THIS CASE, THERE WAS A DECREASING HEMOGLOBIN LEVEL WITH QUICK PULSE AND MULTIPLE ORGAN FAILURES. THE STOOL TEST TO RULE OUT INTESTINAL BLEED WAS ORDERED AND NOT DONE (DESPITE DIARRHEA DAILY) FOR THREE DAYS. MEDICATIONS AND OTHER INTERVENTIONS WERE POTENTIALLY MISAPPLIED OR NOT AT ALL, PUTTING PATIENT AT RISK. ORDERS THAT LAY IN A SILO OF ORDERS AND NOT EXECUTED ARE COMMONPLACE WITH CPOE SYSTEMS. THE DEVICES ARE DEFECTIVE IN THAT THERE IS NO WARNING THAT TASKS HAVE NOT BEEN COMPLETED. | 1
4031340 | THE PT SUSTAINED AN ARRHYTHMIA ATRIAL FIBRILLATION WITH AN ERRATIC RATE OF HEART CONTRACTIONS. MEDICATIONS WERE ORDERED ON THE CPOE SYSTEM TO CONTROL THE RATE. THERE WAS AN ORDER ENTERED IN TO THE CPOE SYSTEM TO NOTIFY THE PHYSICIAN FOR A RATE THAT STAYED QUICK OVER 100 BEATS PER MINUTE. THE RATE STAYED QUICK FOR HOURS MORE THAN 120 BEATS PER MINUTE, CAUSING HEART AND RESPIRATORY FAILURE. NO ONE WAS INFORMED OF THE SPEEDING HEART RATE, WHICH CONTINUED AND CAUSED SEVERE LIFE THREATENING COMPLICATIONS. THE ORDER OF NOTIFICATION WAS IN A SILO BUT NOT KNOWN TO THE NURSES, AND THE NURSE SO BUSY CLICKING AWAY DOCUMENTING WHATEVER, FAILED TO USE COMMON JUDGEMENT TO NOTIFY OF THE VITAL SIGN ABNORMALITY. SINCE CPOE WAS DEPLOYED, VITAL SIGNS ARE NO LONGER VITAL SINCE THEY GO INTO THE EHR WITHOUT THE NURSES' BEING AWARE AND ARE LARGELY IGNORED, EVEN IF THERE WAS A NOTIFICATION ORDER AS IN THIS CASE. THE EHR IS DEFECTIVE IN THAT ITS DECISION SUPPORT FAILS TO WARN THE NURSE OF THE ORDER TO NOTIFY DESPITE THE ORDER TO DO SO IN THE SETTING OF A TACHYCARDIA. THE DEVICE OBSERVES THE ORDERS OF IMPORTANCE FROM THOSE NEEDING TO CARRY THEM OUT. | 1
4059593 | MEDICATION A THAT CONTROLS BLOOD PRESSURE WAS TO BE INCREASED. THE CPOE ON THIS VENDOR'S DEVICE REQUIRES THAT AN ORDER BE DISCONTINUED AND THEN REORDERED IN SEVERAL CLICKING STEPS. DRUG A WAS DISCONTINUED. MEDICATION A1 IS SAME AS A BUT IS DOUBLE DOSE, AND WAS ORDERED. TWO HOURS LATER, THE PATIENT WAS IN SHOCK AND NEEDED FLUID RESUSCITATION AND PRESSORS. WHEN MEDICATION A1 WAS ORDERED, UNBEKNOWNST TO THE CLICKER, MEDICATION A HAD ALREADY BEEN GIVEN. THE NURSE THEN GAVE MEDICATION A1 TO THE PATIENT, RESULTING IN A TRIPLE DOSE OF WHAT WAS MAINTENANCE. AS FOR THE INCREASE TO MEDICATION A1, NEITHER THE NURSES NOR THE PHARMACY NOR THE CDS DEVICE, RECOGNIZED THAT THE OVERDOSE WAS HAPPENING. | 1
3865731 | DURING A CARDIAC CATH PROCEDURE, TECH WAS ASKED TO RECORD A PULLBACK ON THE PHILLIPS XPER RECORDING SYSTEM. TECH PERFORMED THE NORMAL STEPS TO RECORD THE PULLBACK BUT WHEN SHE WAS ASKED TO LABEL A CHAMBER IN THE HEART ON THE SYSTEM, THE SYSTEM WOULD NOT ALLOW HER TO RECORD THE PULLBACK ON THE XPER SYSTEM IN CATH LAB #2. AFTER MULTIPLE ATTEMPTS TO STOP THE RECORDING, THE SYSTEM WOULD NOT STOP RECORDING AND THE GREEN LIGHT WOULD NOT TURN OFF. THE SYSTEM WAS THEN RESTARTED WITHOUT SUCCESS. THE DECISION WAS MADE TO DO A TOTAL REBOOT. THE PATIENT WAS BEING MONITORED BY THE ANESTHESIA TEAM DURING THE ENTIRE CARDIAC CATH PROCEDURE. AFTER REBOOTING THE SYSTEM, THE PRIOR PULLBACK RECORDING WAS LOST AND HAD TO BE REPEATED. THE TOTAL DELAY IN THE CASE WAS 10 MINUTES. | 1
3957711 | MANUFACTURER REF # (B)(4) ARE RELATED TO THE SAME EVENT. THE CONCOMITANT PRODUCTS: CARTO 3 SYSTEM MODEL # M-4800-01 SERIAL # (B)(4); SMART TOUCH BIDIRECTIONAL MODEL# D-1327-05-S, LOT # 16107500M.><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, THE CUSTOMER GOT HIGH FORCE READING AND THE SIGNAL NOISE OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS AND ALL IC (INTRACARDIAL) RECORDINGS ON BOTH CARTO AND THE RECORDING SYSTEM AT THE SAME TIME. THE ISSUE OCCURRED WHILE PREPARING TO ABLATE. THE CUSTOMER HAD FINISHED CONNECTING THE LEFT VEINS AND HAD THIS ISSUE POP UP WHEN GOING TO THE RIGHT SIDED PULMONARY VEINS. WHEN THEY WENT TO THE RIGHT SIDE AND HAD ONE HIGH FORCE READING AND IT APPEARED AS IF THE CATHETER WAS VISUALLY FLIPPING ACROSS THE SCREEN ON CARTO AND THE PHYSICAL MANIPULATION OF THE CATHETER. REPLACING THE CATHETER CABLE, REBOOTING THE PIU AND CHANGING THE CATHETER DID NOT RESOLVE THE ISSUE. THE PROCEDURE WAS COMPLETED WITH CHANGING OUT THE CATHETER CABLE AGAIN WITHOUT ANY PATIENT CONSEQUENCE. | 1
4233109 | IT WAS REPORTED THAT A PATIENT UNDERWENT AN ATRIAL FLUTTER (AFL) PROCEDURE WITH A C3 INTERFACE CABLE ¿ THERAPEUTIC, AND A BAD NO ECG ALL CHANNELS ¿ BS AND IC ISSUE OCCURRED. IT WAS REPORTED THAT THE STOCKERT 70 SYSTEM WAS CLICKING AND THE TEMPERATURE WAS GOING UP AND DOWN. ALL ECG¿S WERE NOISY. THE NOISE STARTED WHEN THE REDEL CABLE WAS PLUGGED INTO THE PATIENT INTERFACE UNIT (PIU). THE REDEL CABLE AND THE ABLATION ADAPTER WERE REPLACED WITHOUT RESOLUTION. THERE WAS NO CATHETER PLUGGED INTO THE ABLATION CABLE THAT IS PLUGGED INTO THE PIU. WHEN THE C3 INTERFACE CABLE-THERAPEUTIC WAS REMOVED, THE NOISE WENT AWAY AND THERE WAS NO LONGER ANY TEMPERATURE ON THE STOCKERT, AS IT SHOULD BE. THE C3 INTERFACE CABLE-THERAPEUTIC WAS REPROCESSED MANY TIMES. A NEW CABLE WAS PLUGGED IN AND THE ISSUE WAS RESOLVED. THE PROCEDURE WAS CONTINUED WITHOUT PATIENT CONSEQUENCE. UPON REQUEST, ADDITIONAL INFORMATION WAS PROVIDED ON THE EVENT. THE NOISE WAS OBSERVED ON ALL THE BODY SURFACE (BS) AND INTRACARDIAC (IC) CHANNELS ON BOTH THE CARTO SYSTEM AND THE PRUKA RECORDING SYSTEM AND THERE WAS NO SIGNAL AVAILABLE TO MONITOR THE PATIENT¿S HEART RHYTHM. THE ISSUE WAS A BENT PIN IN THE ABLATION SOCKET. IT WAS IN CONTACT WITH THE PIN NEXT TO IT. ONCE IT WAS REPLACED, THE ISSUE WAS RESOLVED. THE TEMPERATURE ISSUE WAS FOUND BEFORE USE. NO ENERGY COULD BE DELIVERED SINCE THERE WAS NO ABLATION CATHETER CONNECTED TO THE STOCKERT GENERATOR. VALUES WERE NOT CHECKED AT THE TIME OF THE ISSUE AS THERE WAS NO ABLATION CATHETER. THE STOCKERT GENERATOR WAS PROGRAMMED WITH THE DEFAULT VALUES FOR THE 8MM CATHETER. SINCE THE NOISE AFFECTED ALL IC AND BS SIGNALS ON BOTH THE CARTO SYSTEM AND THE PRUKA RECORDING SYSTEM AND THERE WAS NO SIGNAL AVAILABLE TO MONITOR THE PATIENT¿S HEART RHYTHM, THIS EVENT IS INDICATIVE OF A REPORTABLE EVENT.><SINCE THE PRODUCT WAS NOT RETURNED FOR ANALYSIS, NO PRODUCT FAILURE ANALYSIS CAN BE CONDUCTED AND NO DETERMINATION OF POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. DEVICE HISTORY RECORD (DHR) REVIEW CANNOT BE CONDUCTED BECAUSE THE NO LOT NUMBER WAS PROVIDED BY THE CUSTOMER. MANUFACTURER'S REF. (B)(4). | 1
3657448 | THE MERGE HEMODYNAMIC RECORDING SYSTEM IS OUR PRIMARY MODE OF MONITORING O2 SAT WHILE PERFORMING MODERATE SEDATION. DURING THE PROCEDURE, IT WAS NOTED THAT THE O2 SAT WAS STAYING AT 98% AND NOT CHANGING. THE WAVEFORM AND PULSE DISPLAY INDICATED THAT THE SENSOR WAS WORKING APPROPRIATELY HOWEVER THE O2 VALUE WAS FROZEN. WHEN THIS WAS NOTICED A SECONDARY O2 SENSOR WAS PLACED AND THE O2 SAT WAS FOUND TO BE IN THE 89-90% RANGE. | 1
3965914 | DURING A CARDIAC CATH PROCEDURE A BLOOD GAS RESULT WAS ENTERED INTO THE PHILIPS XPER SYSTEM IN CATH LAB AND LATER WAS MISSING WHEN THE RESULTS WERE NEEDED PRIOR TO FINISHING THE CASE. THE BLOOD GAS RESULT WAS TYPED INTO THE HEMODYNAMICS PAGE A SECOND TIME AND WAS MISSING FROM THE DATA PAGE APPROXIMATELY 15 MINUTES LATER. DUE TO THE EMERGENT NATURE OF THE CASE, XPER COULD NOT BE CALLED DURING THE PROCEDURE. TWO DAYS LATER ANOTHER BLOOD GAS RESULT WAS PLACED INTO THE SYSTEM AND LATER NOTED THAT IT WAS MISSING WHEN THE TECH TRIED TO REVIEW THE INFORMATION. XPER CUSTOMER SERVICE WAS CALLED ABOUT BOTH OF THESE INCIDENTS AND PLANS TO FOLLOW UP. | 1
3605074 | DURING EP PROCEDURES WE USE CARDIOLAB FOR VISUALIZATION OF THE CATHETERS PLACED WITHIN THE HEART AND TO DOCUMENT EVENTS OF THE PROCEDURE. DURING THIS CASE WHILE PACING FOR THE EP STUDY PORTION OF THE ATRIAL FIB ABLATION, CARDIOLAB DISPLAYED MESSAGE OF SESSION TIMED OUT. WHEN THIS MESSAGE WAS CLICKED OFF WE WERE NOT ABLE TO HAVE ANY RESPONSE FROM CARDIOLAB (SYSTEM FROZEN) INCLUDING NOT BEING ABLE TO PACE THROUGH THE MICROPACE SYSTEM. SHUTDOWN OF SYSTEM WAS PERFORMED WITH ADDITIONAL ERROR MESSAGES WHICH WOULD NOT ALLOW US TO UTILIZE THE SYSTEM. BIOMED WAS NOTIFIED; WHEN ARRIVED A SECOND SHUTDOWN OF THE SYSTEM WAS PERFORMED AND THE SYSTEM WAS ABLE TO BE UTILIZED APPROPRIATELY. APPROXIMATE TIME OF INACCESSIBILITY TO CARDIOLAB SYSTEM WAS FROM 1150 TO 1215. PATIENT WAS ON AN ADDITIONAL PIECE OF EQUIPMENT CALLED CARTO WHICH ALSO ALLOWS THE HEART CATHETERS TO BE VISUALIZED. | 1
3758547 | PHILLIPS WITT HEMODYNAMIC MONITORING SYSTEM WOULD NOT RECORD HEMODYNAMICS. HAD TO REBOOT SYSTEM. CALL MADE TO WITT. ATTENDING REQUESTED SAFETY REPORT BE PLACED. CAUSED DELAY IN PROCEDURE.ONGOING AND UNRESOLVED ISSUES WITH SYSTEM. ADDITIONAL MEDSUN REPORTS TO FOLLOW. REBOOT OF SYSTEM AND TROUBLESHOOTING ARE OCCURRING WHILE PATIENT IS SEDATED AND ON TABLE. MANUFACTURER IS AWARE. SOME OF THE REPORTED ISSUES BY STAFF USING SYSTEM ARE AS FOLLOWS:PREVIOUS PATIENT INFORMATION CAN SHOW UP ON NEXT CASE IF THERE IS NO ACTIVE RECORDING.WHILE DOING CARDIAC OUTPUT, SOMETIME CAN'T SEE THE CURVE, ONLY THE NUMERIC VALUE. SOME OF THE WAVEFORMS WERE DELETED WHILE PERFORMING CARDIAC OUTPUTS.CAL VALUES ARE NOT BE DISPLAYED.HEMOGLOBIN VALUES FALLING OFF.INACCURATE RECORDINGS.CAN'T FIND RECORDED DATA.CANNOT FILTER THROUGH THE SEARCH FEATURE.INJECTION VALUES - IF NOT ALL VALUES ARE ENTERED, RESULTS IN DATA LOSS.VO2 IS INCORRECTLY DISPLAYED.ADDITIONALLY THE FOLLOWING ARE SOME OF THE PROBLEM DESCRIPTIONS LOGGED OVER TIME WITH SOME REFERENCING CALLS PLACED WITH THE MANUFACTURER:WILL NOT SELECT WAVEFORM TRACINGS AND WILL NOT LABEL POINTS ON TRACINGS IN WAVE REVIEW.WILL NOT DELETE LINE ITEMS OR RE-CHART. REBOOT X 2.PULLBACK NO RECORDING.SATS ON HEMO PAGE CHANGED.NOT ABLE TO LABEL ORIGINAL RECORDING/DROPPING INFO/FREEZING UP.BLUE SCREEN!! REBOOTED.WITT NOT SHOWING PRESSURE.SEQUENCE PULLBACK NOT WORKING. | 1
4023629 | RELATED MANUFACTURER REFERENCE 3005188751-2014-00088, 3005188751-2014-00084, 3005188751-2014-00089, 2030404-2014-00077, 2030404-2014-00078, 2030404-2014-00079. FOLLOWING A VENTRICULAR TACHYCARDIA (VT) ABLATION PROCEDURE, THE PATIENT EXPIRED. THIS HIGH RISK SUBSTRATE MODIFICATION VT ABLATION PROCEDURE IN THE LEFT VENTRICLE WAS PERFORMED ON AN ELDERLY MALE WITH SEVERE CARDIOMYOPATHY, A HISTORY OF LARGE LAD TERRITORY MYOCARDIAL INFARCTION, AND AN IMPLANTED CARDIAC PACEMAKER/DEFIBRILLATOR. THE PATIENT PRESENTED TO THE LAB IN A PACED RHYTHM AND ANESTHESIA WAS ADMINISTERED; THE PATIENT SUBSEQUENTLY BECAME HYPOTENSIVE. AN EPINEPHRINE INFUSION WAS INITIATED TO STABILIZE AND MAINTAIN BLOOD PRESSURE. TRANSSEPTAL PUNCTURE WAS PERFORMED WITH TRANSESOPHAGEAL ECHOCARDIOGRAM GUIDANCE USING A BRK TRANSSEPTAL NEEDLE AND AN AGILIS NXT INTRODUCER WAS ADVANCED INTO THE LEFT ATRIUM. FOLLOWING THE TRANSSEPTAL PUNCTURE, A PERICARDIAL EFFUSION WAS NOTED; HOWEVER, THIS WAS NOT SIGNIFICANT AND THE PROCEDURE CONTINUED. A NON-SJM SCREW IN BIPOLAR PACING CATHETER WAS PLACED IN THE RIGHT VENTRICLE AND AN INQUIRY AFOCUS II EP CATHETER WAS ADVANCED INTO THE LEFT VENTRICLE. THE AFOCUS II EP CATHETER WAS USED TO MAP, BOTH TRANSSEPTAL AND VIA A RETROGRADE APPROACH. THE AFOCUS II EP CATHETER WAS DAMAGED AND THEN REPLACED TO CONTINUE MAPPING. DURING COMPLETION OF THE MAP, THE PATIENT DEVELOPED VT, RESULTING IN HEMODYNAMIC INSTABILITY; THEREFORE, THE PATIENT WAS CARDIOVERTED TO SINUS RHYTHM (SR). THE CATHETERS SHIFTED ON THE ENSITE VELOCITY MAPPING SYSTEM DUE TO THE CARDIOVERSION AND SETTINGS CHANGES. RE-MAPPING WAS COMPLETED AND THE AFOCUS II EP CATHETER WAS REPLACED WITH A TACTICATH QUARTZ ABLATION CATHETER. CATHETER DISPLAY ISSUES WERE NOTED AND TROUBLESHOOTING WAS PERFORMED, EVENTUALLY RESOLVING THIS ISSUE. RE-MAPPING WAS PERFORMED AGAIN USING A THERAPY COOL FLEX ABLATION CATHETER. ABLATION WAS SUCCESSFULLY PERFORMED AROUND A VERY LARGE SCAR WHICH STRETCHED FROM THE APEX TO THE BASE OF THE LEFT VENTRICLE. NEAR THE CONCLUSION OF THE PROCEDURE, THE ANESTHESIA TEAM NOTED IT WAS DIFFICULT TO MAINTAIN A STABLE BLOOD PRESSURE AND THE PERICARDIAL EFFUSION NOTED EARLIER HAD INCREASED IN SIZE. PROTAMINE WAS ADMINISTERED AND A PERICARDIAL DRAIN WAS PLACED, WHICH STABILIZED THE PATIENT'S BLOOD PRESSURE. THE PATIENT WAS AWAKENED AND TRANSFERRED TO THE RECOVERY ROOM IN STABLE CONDITION. THE PATIENT THEN BECAME HYPOTENSIVE AND AN EPINEPHRINE BOLUS WAS ADMINISTERED. AN ECHOCARDIOGRAM REVEALED AN EFFUSION WITH NO LEFT VENTRICULAR MOVEMENT. THE PATIENT SUBSEQUENTLY EXPIRED.><TWO DISKS WERE RETURNED CONTAINING THE CASE STUDY AND LOGS. NO RECORDED SEGMENTS WERE MADE OF THE REPORTED EVENT ON EITHER OF THE RETURNED DISK. SYNCHRONIZED CARDIOVERSION CAN CAUSE A TRANSIENT CATHETER SHIFT. AFTER THE IMMEDIATE CIRCUMSTANCE HAS STABILIZED, THERE CAN BE A RESIDUAL CHANGE IN THORACIC IMPEDANCE THAT CAN HAVE AN IMPACT ON LOCATION. THIS EXPECTED BEHAVIOR OF THE VELOCITY SYSTEM IS RELATED TO CHANGES IN THE IMPEDANCE SIGNALS SENSED BY THE AMPLIFIER. | 1
3962906 | BOTH CENTRAL PATIENT MONITORS WENT COMPLETELY BLANK AT 0225. CHARGE NURSE TRIED TO RECTIFY SITUATION WITH REBOOT, WITHOUT SUCCESS. BIOMED IDENTIFIED THAT THE PROBLEM WAS THE UPS BATTERY BACK UP WAS DEAD. | 1
3960935 | USING NELLCOR TECHNOLOGY TO MONITOR SPO2, IT WAS REPORTED THAT THE SPO2 VALUE WAS FREEZING ON THE HEMO MONITOR. IF PULSE RATE IS MEASURED FROM SPO2, PULSE RATE ALSO FREEZES WHICH CAUSES THE HEMO MONITOR TO DISPLAY STALE DATA.><MERGE HEALTHCARE MANUFACTURES MERGE HEMO SOFTWARE TO BE USED IN CONJUNCTION WITH A PATIENT DATA MODULE (PDM) MANUFACTURED BY SCHILLER. (B)(6) REPORTED AN ISSUE TO MERGE WHERE THE SPO2 VALUE WAS FREEZING ON THE MERGE HEMO MONITOR DISPLAY. NORTHEAST GEORGIA USES A PDM WITH NELLCOR SPO2 TECHNOLOGY IN IT. MERGE HEMO ALSO SUPPORTS MASIMO TECHNOLOGY FOR SPO2 MONITORING IN THE PDM. NO COMPLAINTS HAVE BEEN FILED WITH MERGE IN REGARDS TO THE MASIMO TECHNOLOGY. MERGE OFFERED NORTHEAST GEORGIA THE OPPORTUNITY TO HAVE THEIR UNITS REPLACED WITH MASIMO PDMS AND THEY HAVE ACCEPTED THAT. NORTHEAST GEORGIA'S UNITS WERE RETURNED TO MERGE HEALTHCARE FOR EVALUATION. IT WAS DETERMINED THAT THE MERGE HEMO SOFTWARE IS OPERATING AS EXPECTED. THE HEMO MONITOR IS DISPLAYING THE DATA BEING SENT TO IT FROM THE SCHILLER PDM. THE PDM IS NOT UPDATING THE SPO2 VALUE AS IT IS CHANGING AND SENDING TO MERGE HEMO. MERGE HAS NOTIFIED SCHILLER OF THIS ISSUE. IT IS ONLY OCCURRING WITH THE PDMS THAT CONTAIN THE NELLCOR TECHNOLOGY. INTERNALLY, MERGE HAS ISSUED A STOP ORDER, SO NO FURTHER NELLCOR UNITS WILL BE ACCEPTED FROM SCHILLER BY MERGE OR DISTRIBUTED TO THE FIELD BY MERGE. MERGE HAS DRAFTED A CUSTOMER LETTER THAT WILL BE SENT TO ALL SITES THAT HAVE PDMS USING NELLCOR TECHNOLOGY. AS THE ISSUE IS INTERMITTENT AND ALL OTHER VITALS REMAIN MONITORED CORRECTLY, MERGE IS NOT INITIATING A RECALL AT THIS TIME, ONLY ADVISING CUSTOMERS TO BE AWARE. ONCE A ROOT CAUSE IS DETERMINED BY SCHILLER, MERGER MAY TAKE FURTHER ACTION. IN THE ADVISORY LETTER, CUSTOMERS ARE BEING OFFERED THE OPTION TO CALL MERGE SUPPORT AND REPLACE NELLCOR UNITS WITH MASIMO UNITS. | 1
3949057 | THE CARDIAC MONITORING SYSTEM WAS NOT DISPLAYING THE FULL PARAMETER LIST ON REMOTE MONITOR IN THE PROCEDURE ROOM AS WAS DISPLAYED ON CONTROL ROOM MONITOR. CERTAIN LINE ITEM PARAMETERS WERE MISSING ON THE REMOTE MONITOR LIST.NURSE CONTACTED PHILIPS TECH SUPPORT AND FOLLOWED THEIR INSTRUCTIONS. HE TURNED OFF POWER FOR THE CARDIAC MONITORING SYSTEM AND DISCONNECTED THE POWER CORD. HE THEN TURNED ON THE POWER SWITCH MOMENTARILY, TURNED IT BACK OFF, AND THEN RECONNECTED POWER. WHEN POWERED BACK UP THE REMOTE PARAMETERS WERE CORRECTLY DISPLAYED AND MATCHED THE CONTROL ROOM LIST.BASED ON THE COMPLAINT, THE MANUFACTURER IS IN PROGRESS OF PROVIDING SOFTWARE UPGRADE. | 1
3962822 | THE XPER RECORDING SYSTEM WAS USED TO DOCUMENT AN ELECTROPHYSIOLOGY CASE. WHEN TECH TRIED TO LOOK UP THE DOCUMENTATION FOR THE CASE, NOTED THAT THERE WAS NO DOCUMENTATION FOR THE CASE IN THE COMPUTER. TECH THEN CHECKED WITH TWO OTHER TECHS AND THEY BOTH AGREED THE CASE WAS RECORDED IN THE SYSTEM. | 1
3456576 | ACCORDING TO THE CARTO 3 IFU, ¿THE SYNCHRONIZING FUNCTION OF CERTAIN DEFIBRILLATORS MAY BE INCOMPATIBLE WITH THE CARTO® 3 SYSTEM (FOR EXAMPLE, ZOLL DEFIBRILLATORS). WHEN USING EXTERNAL DEFIBRILLATORS WITH AN INCOMPATIBLE SYNCHRONIZING FUNCTION, DISABLE THE CARTO® 3 SYSTEM LOCATION PAD BEFORE ATTEMPTING TO DEFIBRILLATE THE PATIENT.¿ A SUPPLEMENTAL REPORT ON DEVICE EVALUATION WILL BE SUBMITTED ONCE IT IS COMPLETED. (B)(4).><IT WAS REPORTED THAT DURING A VENTRICULAR TACHYCARDIA (VT) MAPPING PROCEDURE USING CARTO3 FOR GUIDANCE, THE PATIENT, WHO HAD BEEN IN SINUS RHYTHM FOR MOST OF THE PROCEDURE, DEVELOPED AN EPISODE OF VENTRICULAR TACHYCARDIA (VT), WHICH THE PHYSICIAN WANTED TERMINATED. AN ATTEMPT WAS MADE TO DELIVER EXTERNAL AN EMERGENCY CARDIOVERSION THERAPY BUT THE DEFIBRILLATOR GAVE AN ERROR MESSAGE ¿EQUIPMENT DISABLED: THERAPY¿. THE EXTERNAL DEFIBRILLATOR PATCHES ATTACHED TO A PHILIPS HEARTSTART XL AND EXTERNAL DEFIBRILLATOR. A NEW CONNECTOR CABLE WAS CONNECTED FROM DEFIBRILLATOR PATCHES TO THE DEFIBRILLATOR AND SECOND DEFIBRILLATOR GENERATOR WAS TRIED (SAME MAKE AND VERSION) BUT THE SAME ERROR MESSAGE OCCURRED AND THE ERROR MESSAGE PERSISTED. WHEN IT WAS CLEAR THERE WAS AN ISSUE WITH THE SECOND DEFIBRILLATOR, THE PHYSICIAN RESUMED ATTEMPTS TO TERMINATE THE VENTRICULAR TACHYCARDIA (VT) WITH OVERDRIVE PACING FROM THE INTERNAL CATHETER ELECTRODE. THE PHYSICIAN SUCCEEDED BY THE TIME THE THIRD DEFIBRILLATOR WAS AVAILABLE AND ATTACHED TO THE PATIENT'S DEFIBRILLATION PADS. DEFIBRILLATION THERAPY WAS NOT REQUIRED FROM THE THIRD DEVICE (OLDER VERSION), BUT IT WAS ATTACHED TO THE PATIENT DEFIBRILLATION PADS AND DID NOT DISPLAY ANY ERROR MESSAGES. THE LOCATION PAD ACTIVATION WAS OFF FROM WHEN THE SECOND DEFIBRILLATOR WAS BROUGHT IN TO THE LAB UNTIL THE PROCEDURE RESUMED. THE PATIENT WAS UNDER GENERAL ANESTHESIA AND THE VENTRICULAR TACHYCARDIA (VT) RATE REMAINED STABLE THROUGHOUT THE EPISODE. DURING TROUBLESHOOTING, THE ORIGINAL FAULTY" DEFIBRILLATOR WAS REATTACHED TO THE PATIENT'S DEFIBRILLATOR PADS AND WITH THE LOCATION PAD STILL ON | 1
3762000 | MANUFACTURER REF # (B)(4). IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, THE SIGNAL NOISE OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS ECGS AND ALL IC (INTRACARDIAC) RECORDINGS ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME. THE ECG PATCHES WERE CHECKED FOR GOOD CONNECTION AND ALL CABLES WERE ROUTED SO NOT TO OVERLAP. BIOSENSE FIELD SERVICE ENGINEERS REPLACED EP OUT CARD, AND AS A RESULT, THE PROBLEM WAS RESOLVED. THERE WERE NO FURTHER ISSUES, AND THE SYSTEM IS FUNCTIONING NORMALLY AND IS READY FOR USE. THE REPLACED EP OUT CARD WAS SENT TO THE CARTO MANUFACTURER (HTC) FOR FURTHER INVESTIGATION. THE CARTO MANUFACTURER TESTED THE CARD AND FOUND NO PROBLEM. THE CARD PASSED ALL RELEVANT TESTS SUCCESSFULLY. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE DHR REVIEW VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES.><(B)(4).><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, THE SIGNAL NOISE OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS ECGS AND ALL IC (INTRACARDIAC) RECORDINGS ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME. THE ECG PATCHES WERE CHECKED FOR GOOD CONNECTION AND ALL CABLES WERE ROUTED SO NOT TO OVERLAP. THE PROCEDURE WAS COMPLETED IN SPITE OF THE NOISE WITHOUT ANY PATIENT CONSEQUENCE. | 1
3898214 | THE HARDWARE INVESTIGATION IS PENDING. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. CONCOMITANT PRODUCTS: UNK. (B)(4).><IT WAS REPORTED THAT WHILE A PATIENT UNDERWENT A CARDIAC ABLATION PROCEDURE WITH A CARTO 3 DEVICE, A PACING DEVICE MALFUNCTION OCCURRED. THE CUSTOMER REPORTED THAT WHEN A TEMPORARY STIMULATOR WAS CONNECTED TO THE PATIENT INTERFACE UNIT (PIU), THE CONNECTION POINT FOR CATHETERS, AND THE OUTPUT WAS INCREASED, THE PACING WAS ROUTED TO THE WRONG CHANNEL. THE PHYSICIAN STATED THAT DUE TO THE PACING ISSUES DESCRIBED, THERE MAY HAVE BEEN ABERRANT HEART RHYTHMS DETECTED, ALTHOUGH THIS ISSUE COULD NOT BE VALIDATED BY ON-SITE FIELD REPRESENTATIVES. IT WAS CONFIRMED THAT THERE WAS NO PATIENT CONSEQUENCE. THE PHYSICIAN CHOSE TO COMPLETE THE CASE WITH MODIFICATIONS, LOWERING THE OUTPUT ON THE STIMULATOR. THE HARDWARE INVESTIGATION HAS BEGUN BUT IT HAS NOT BEEN COMPLETED AT THIS TIME.><B) (B)(4) IT WAS REPORTED THAT WHILE A PATIENT UNDERWENT A CARDIAC ABLATION PROCEDURE WITH A CARTO 3 DEVICE, A PACING DEVICE MALFUNCTION OCCURRED. THE CUSTOMER REPORTED THAT WHEN A TEMPORARY STIMULATOR WAS CONNECTED TO THE PATIENT INTERFACE UNIT (PIU), THE CONNECTION POINT FOR CATHETERS, AND THE OUTPUT WAS INCREASED, THE PACING WAS ROUTED TO THE WRONG CHANNEL. A FIELD SERVICE ENGINEER (FSE) ARRIVED AT THE HOSPITAL AND REPLACED THE PIU, LOCATION PAD (LP) AND POWER UNIT (PU). THE SYSTEM WAS TESTED AND FOUND READY FOR USE WITH THE NEW PARTS AND THE SUSPICIOUS PARTS WERE SENT TO THE DEVICE MANUFACTURER FOR INSPECTION. IT WAS FOUND THAT THERE WAS A FAILED OPTO SWITCH U447 ON THE ECG CARD B057347 WHICH CAUSED THE PACING PROBLEM DESCRIBED IN THE COMPLAINT. THE CUSTOMER COMPLAINT WAS CONFIRMED. AN INTERNAL CORRECTIVE ACTION HAS BEEN OPENED TO INVESTIGATE OPTO SWITCHES AND TVS FAILURES". THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT." | 1
3562851 | INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. MANUFACTURER'S REF. NO: (B)(4).><IT WAS REPORTED THAT IN THE MIDDLE OF THE ATRIAL FLUTTER (AFL) PROCEDURE, ALL THE ELECTROGRAMS ON BOTH THE CARTO 3 SYSTEM AND THE RECORDING SYSTEM BECAME VERY NOISY AND UNREADABLE. THE CALLER WAS INSTRUCTED TO CHECK THE LIMB LEADS OF THE BODY SURFACE. THE CALLER WAS THEN INSTRUCTED TO UNPLUG ALL CATHETERS FROM THE FRONT OF THE PATIENT INTERFACE UNIT (PIU). THE NOISE WAS GONE. THE CALLER WAS THEN INSTRUCTED TO RESEAT ALL THE CATHETERS ONE BY ONE. ALL THE NOISE WAS RESOLVED, BUT THE REF/DECA PORT WAS STICKY" AND THE CABLE DID NOT SEAT EASILY. THERE WAS NO PATIENT INJURY REPORTED IN THIS PROCEDURE. THE SEVERITY OF THE NOISE ON ALL THE CHANNELS ON BOTH THE CARTO 3 SYSTEM AND THE RECORDING SYSTEM AT THE SAME TIME IS INDICATIVE OF A REPORTABLE EVENT." | 1
3583895 | INVESTIGATION STILL IN PROGRESS. A SUPPLEMENTAL REPORT OR DEVICE EVALUATION WILL BE SUBMITTED. (B)(4).><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, THE BS ECG TRUNK CABLE CONNECTOR WAS BROKEN AND THE ECG SIGNALS ON THE CARTO 3 SYSTEM AND ON THE EP RECORDING SYSTEM WERE NOISY. THE LAB STAFF COMPLAINED THAT THE CABLE WAS NOT LONG ENOUGH AND THAT IS WHY THE CABLE BROKE. THE BS-ECG TRUNK CABLE WAS USED FROM THE SAVE THE CASE KIT TO RESOLVE THE ISSUE. THERE WAS NO PATIENT INJURY REPORTED IN THIS PROCEDURE. UPON REQUEST, ADDITIONAL INFORMATION WAS PROVIDED ON THE EVENT. THE ISSUE OCCURRED PROBABLY DURING PRE-OP BUT REALIZED IT WHEN THEY STARTED THE PROCEDURE BECAUSE THE SIGNALS WERE NOISY. THE NOISE OCCURRED ON ALL THE CHANNELS, INCLUDING THE 12 LEADS OF BS ECGS AND ALL IC (INTRACARDIAC) RECORDINGS. THE NOISE OCCURRED ON BOTH THE CARTO 3 SYSTEM AND THE EP RECORDING SYSTEM AT THE SAME TIME. THE NOISE WAS QUITE DISTURBING AND BECAUSE THEIR NEXT PATIENT WAS A VT PATIENT, THEY SAID THEY NEED MUCH CLEARER SIGNALS TO BE CONFIDENT ENOUGH TO EVALUATE THEM. THE PHYSICIAN WAS ABLE TO COMPLETE THE PROCEDURE WITH THE NOISE BECAUSE OF HIS YEARS OF EXPERIENCE AND BEING ABLE TO DIFFERENTIATE THE NOISE FROM NORMAL POTENTIALS. PER THE SEVERITY OF THE NOISE ON ALL SIGNALS ON BOTH THE CARTO 3 SYSTEM AND THE RECORDING SYSTEM AT THE SAME TIME IS INDICATIVE OF A REPORTABLE EVENT. | 1
3700788 | (B)(4).><DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, IT WAS REPORTED THE USER HAD 2 OCCURRENCES OF A MAP SHIFT ON 2 SEPARATE MAPS WITHOUT ANY ERROR CODES BEING DISPLAYED. THE PROCEDURE WAS COMPLETED, SO TROUBLESHOOTING WAS NOT POSSIBLE. ADDITIONAL INFORMATION PROVIDED STATED THE USER NOTICED THE SHIFT BECAUSE ANATOMY DID NOT MATCH. MAP SHIFT OCCURRED DURING MAPPING. IT WAS NOTED WHEN THE LASSO NAV WAS REPOSITIONED INTO THE LSPV, THE POSITION DID NOT CORRESPOND WITH THE PREVIOUSLY COLLECTED ANATOMY. THE MAP SHIFTED ABOUT 1 TO 2 MM.><(B)(4). DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, IT WAS REPORTED THE USER HAD 2 OCCURRENCES OF A MAP SHIFT ON 2 SEPARATE MAPS WITHOUT ANY ERROR CODES BEING DISPLAYED. THE PROCEDURE WAS COMPLETED SO TROUBLESHOOTING WAS NOT POSSIBLE. ADDITIONAL INFORMATION PROVIDED STATED THE USER NOTICED THE SHIFT BECAUSE ANATOMY DID NOT MATCH. MAP SHIFT OCCURRED DURING MAPPING. IT WAS NOTED WHEN THE LASSO NAV WAS REPOSITIONED INTO THE LSPV, THE POSITION DID NOT CORRESPOND WITH THE PREVIOUSLY COLLECTED ANATOMY. THE MAP SHIFTED ABOUT 1 TO 2 MM. THE FILES, WHICH WERE SENT TO DEVICE MANUFACTURER (HTC), DID NOT CONTAIN RECORDINGS NECESSARY FOR INVESTIGATION. THEREFORE, THE INVESTIGATION COULD NOT BEEN PERFORMED. HOWEVER, FIELD SERVICE ENGINEER VISITED THE ACCOUNT AND TESTED THE SYSTEM. THE SYSTEM PASSED FULL PM AND FOUND READY FOR USE. THE BWI REPRESENTATIVE INFORMED THAT THEY PERFORMED MULTIPLE CASES SINCE THIS PROBLEM WAS REPORTED AND THE MAP SHIFT PROBLEM HAS NOT RE-OCCURRED. THE DHR ASSOCIATED WITH CARTO 3 # (B)(4) WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. THE CUSTOMER COMPLAINT WAS NOT CONFIRMED. | 1
3363356 | (B)(4). IT WAS REPORTED THAT THE CARTO 3 SYSTEM HAD NOISE IN IC SIGNALS WHEN RF WAS APPLIED AND ELECTRICAL STIMULATION WAS DONE SIMULTANEOUSLY. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION REGARDING THE EVENT, ADDITIONAL INFORMATION STATED THAT THE SIGNAL NOISE OCCURRED IN ALL CHANNELS, INCLUDING BS ECG AND IC RECORDINGS ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME MAKING THIS EVENT REPORTABLE. THE SYSTEM WAS REMOVED FROM THE ACCOUNT AND SENT TO CARTO MANUFACTURER (HTC) FOR FURTHER INVESTIGATION. HTC TESTED THE SYSTEM AND REPORTED THAT THE COMPLAINT WAS CONFIRMED. OPTO-SWITCHES OF ECG CARD WERE FOUND DEFECTIVE CAUSING NOISE ON CHANNELS 1-2 AND 5-6. THE OPTO-SWITCHES DEFECTIVES WERE REPLACED WITH NEW ONES AND THE ISSUE WAS SOLVED. THE DHR ASSOCIATED WITH CARTO 3 # 11474 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE. AN INTERNAL CORRECTIVE ACTION WAS OPENED TO ADDRESS OPTO-SWITCHES AND TVS FAILURES.><(B)(4).><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB) PROCEDURE, THE CARTO 3 SYSTEM HAD NOISE IN IC (INTRACARDIAL) AND ECG SIGNALS DURING ABLATION AND PACING. AFTER FOLLOW UP WITH THE CUSTOMER TO OBTAIN DETAILED INFORMATION REGARDING THE EVENT, ON (B)(6) 2013 ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE SIGNAL NOISE OCCURRED IN ALL CHANNELS, INCLUDING BS ECG AND IC RECORDINGS ON BOTH CARTO AND EP RECORDING SYSTEMS AT THE SAME TIME MAKING THIS EVENT REPORTABLE. THE PROCEDURE WAS ANATOMICAL MAPPING ON LEFT ATRIUM AND THE PHYSICIAN WAS ABLATING IN THE PULMONARY VEIN OSTIUM. THE PHYSICIAN STATED THAT IN THE FIRSTS 10S OF ABLATION AND PACING (BY A DIAGNOSTIC CATHETER) IT WAS IMPOSSIBLE TO SEE ANY SIGNAL, AFTER STABILIZATION OF THE NOISE, THERE WAS A 0,3MV NOISE IN MAP 1-2. THE PHYSICIAN WAS NOT ABLE TO INTERPRET THE SIGNALS DURING THE FIRSTS 10S, AFTER 10S, 15S, HE DID IT WITH EXTREME DIFFICULTY. AFTER SOME TIME WITH THE SAME NOISE, HE CHOSE TO PACE AND ABLATE ALTERNATELY. WHEN THE PHYSICIAN WAS ABLATING OR PACING ALONE, THERE WAS NOT SIGNIFICANT NOISE. THE PROCEDURE WAS COMPLETED BUT IT TOOK MORE TIME TO FINISH IT. BASED ON THE INFORMATION PROVIDED, THE ALERT DATE WAS CHANGED TO (B)(4) 2013. | 1
3781459 | IT WAS REPORTED THAT DURING IDIOPATHIC VENTRICULAR TACHYCARDIA (IDVT) PROCEDURE THE PHYSICIAN GOT A MAP SHIFT ON THIS CARTO 3 SYSTEM AFTER ABLATING. THERE WAS NO WARNING MESSAGE FOR MAP SHIFT. AFTER MAKING A HIS POINT WITH AN ASSOCIATED HIS EGRAM, THE PHYSICIAN PLACED THE ABLATION CATHETER OUT THE CS AND ABLATED. THE MOVE WAS NOTICED WHEN THE PHYSICIAN COULD NOT PLACE THE CATHETER AT THE INITIAL ABLATION TAGS. THEN HE PLACED THE ABLATION CATHETER AT THE ORIGINAL HIS POINT, BUT THERE WAS NO HIS EGRAM THERE ANYMORE. HE THEN PLACED THE ABLATION CATHETER ABOUT 2CM ABOVE THE HIS AND FOUND WHERE THE HIS WAS AT THAT TIME. THE FLUORO SYSTEM WAS NOT MOVED. THE ISSUE WAS NOT RESOLVED AS THE PROCEDURE WAS COMPLETE WITHOUT ANY PATIENT CONSEQUENCE.><(B)(4).><(B)(4). IT WAS REPORTED THAT DURING IDIOPATHIC VENTRICULAR TACHYCARDIA (IDVT) PROCEDURE THE PHYSICIAN GOT A MAP SHIFT ON THIS CARTO 3 SYSTEM AFTER ABLATING. THERE WAS NO WARNING MESSAGE FOR MAP SHIFT. AFTER MAKING A HIS POINT WITH AN ASSOCIATED HIS EGRAM, THE PHYSICIAN PLACED THE ABLATION CATHETER OUT THE CS AND ABLATED. THE MOVE WAS NOTICED WHEN THE PHYSICIAN COULD NOT PLACE THE CATHETER AT THE INITIAL ABLATION TAGS. THEN HE PLACED THE ABLATION CATHETER AT THE ORIGINAL HIS POINT, BUT THERE WAS NO HIS EGRAM THERE ANYMORE. HE THEN PLACED THE ABLATION CATHETER ABOUT 2CM ABOVE THE HIS AND FOUND WHERE THE HIS WAS AT THAT TIME. THE FLUORO SYSTEM WAS NOT MOVED. THE ISSUE WAS NOT RESOLVED AS THE PROCEDURE WAS COMPLETE WITHOUT ANY PATIENT CONSEQUENCE. ACCORDING TO CLINICAL ACCOUNT SPECIALIST, THIS SHIFTED HAPPENED AFTER CARDIOVERSION. ISSUE WAS RELATED TO PATIENT MOVEMENT. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE DHR REVIEW VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. | 1
3478546 | (B)(4).><IT WAS REPORTED THAT DURING AN ATRIAL FIBRILLATION (AFIB), THE CUSTOMER NOTICED APPROXIMATELY 10 MM MAP SHIFT WITH A DISCREPANCY IN DISTANCE BETWEEN THE LASSO CATHETER AND THE LEFT PULMONARY VEIN. THERE WERE NO ERROR MESSAGES OR ALERTS DISPLAYED ON THE CARTO 3 SYSTEM. MAP SHIFT NOTICED WHEN CHECKING PULMONARY VASCULAR (PV) POST ABLATION. THE ACL FUNCTION WAS IN USE DURING THE CASE. | 1
3504698 | THE INVESTIGATION IS STILL IN PROGRESS. (B)(4).><DURING AN ATRIOVENTRICULAR NODAL REENTRANT TACHYCARDIA PROCEDURE, IT WAS REPORTED THAT THERE WERE NOISE ON THE BODY SURFACE LEADS AND INTRACARDIAL SIGNALS BOTH ON THE CARTO® 3 SYSTEM AND THE RECORDING SYSTEM. THE ISSUE WAS RESOLVED BY REBOOTING THE PATIENT INTERFACE UNITE (PIU) AND TIGHTENING THE LOCATION PAD EXTENSION CABLE. THE PROCEDURE WAS COMPLETED WITH NO PATIENT CONSEQUENCES. ON (B)(4), RECEIVED ADDITIONAL INFORMATION REQUESTED FROM BWI REPRESENTATIVE STATING THAT THE NOISE SIGNAL OCCURRED THRU ALL 12 LEADS AND INTRACARDIAL SIGNALS ON BOTH SYSTEMS AT THE SAME TIME. THE NOISE WAS A WAVE OF 60 CYCLES THRU 12 LEADS AND INTRACARDIAC SIGNALS EVERY 2 TO 3 SECONDS. THE PHYSICIAN STRUGGLED TO INTERPRET BOTH SIGNALS. DUE TO THIS INFORMATION, THIS COMPLAINT BECAME REPORTABLE. | 1
3699275 | DURING AN ATRIAL FIBRILLATION PROCEDURE, IT WAS REPORTED THAT THERE WAS A MAP SHIFT DURING MAPPING WITH NO PATIENT MOVEMENT AND NO ALERT DISPLAYED ON THE CARTO® 3 SYSTEM. THE PROCEDURE WAS COMPLETED CONVENTIONALLY WITHOUT ANY PATIENT CONSEQUENCE. ON (B)(6) RECEIVED ADDITIONAL INFORMATION REQUESTED FROM BWI REPRESENTATIVE STATING THAT THE MAP SHIFT WAS APPROXIMATELY 6 (SIX) MILLIMETERS. AS THERE WERE NO ERROR MESSAGE FOR THE MAP SHIFT ISSUE, THIS MAY BE A POTENTIAL RISK TO THE PATIENT. THEREFORE, THIS COMPLAINT BECAME REPORTABLE.><PLEASE REFER TO SECTION H3 (EVALUATION SUMMARY) FOR INVESTIGATION RESULTS. MANUFACTURER'S REFERENCE # (B)(4) DURING AN ATRIAL FIBRILLATION PROCEDURE, IT WAS REPORTED THAT THERE WAS A MAP SHIFT DURING MAPPING WITH NO PATIENT MOVEMENT AND NO ALERT DISPLAYED ON THE CARTO® 3 SYSTEM. THE PROCEDURE WAS COMPLETED CONVENTIONALLY WITHOUT ANY PATIENT CONSEQUENCE. A DEVICE MANUFACTURER (HTC) REPRESENTATIVE VISITED THE HOSPITAL AND INVESTIGATED THE ISSUE. THE ISSUE WAS REPRODUCED. IT WAS FOUND THAT IT WAS CAUSED BY WORKFLOW (WF): CINE/ IMAGE WERE TAKEN WITH SENSOR-BASED CATHETER DISTANT FROM THE AREA OF INTEREST, AND AS A RESULT THE IMAGE WAS APPLIED TO THE CARTO 3 SPACE IN AN INACCURATE AREA, CAUSING DIFFERENCE BETWEEN THE CARTO MAP AND FLUORO IMAGE TAKEN. HTC REPRESENTATIVE HAS ADVISED ON IMPROVED WF. ONCE PROPOSED WF WAS IMPLEMENTED, NEXT CASES PASSED WITHOUT ISSUES AND THE SYSTEM IS OPERATIONAL. THIS ISSUE WAS ESCALATED FOR PREVENTIVE ACTION. THE DHR ASSOCIATED WITH CARTO 3 # 13453 WAS REVIEWED AND THERE WERE NOT ANY DISCREPANCIES NOTED. THE SYSTEM MET ALL SPECIFICATIONS UPON ITS RELEASE.><A) THE INVESTIGATION IS STILL IN PROGRESS. (B)(4). | 1
3956207 | EVALUATION SUMMARY: FOR INVESTIGATION RESULTS. MANUFACTURER'S REFERENCE # (B)(4). IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO® 3 SYSTEM, AND NOISE WAS REPORTED ON PROXIMAL AND DISTAL ELECTRODE OF THE ABLATION CATHETER IN CARTO SYSTEM AND EP RECORDER. THERE WAS NO PATIENT CONSEQUENCE. THE PROCEDURE WAS PROLONGED FOR 60 MINUTES AND CANCELLED DUE TO THE ISSUE WITHOUT TREATING THE ARRHYTHMIA. THE PROCEDURE WAS BEING PERFORMED IN THE LEFT ATRIUM AND A TRANSSEPTAL PUNCTURE WAS PERFORMED PRIOR TO THE CASE CANCELLATION. THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED. THE FIELD SERVICE ENGINEER (FSE) REPORTED THAT THE ISSUE SOLVED BY REPLACEMENT OF THE BACKPLANE BOARD. SYSTEM IS OPERATIONAL. THE FAULTY BACKPLANE BOARD WAS SENT TO THE DEVICE MANUFACTURER (HTC). HTC RMA REPORTED: THE CUSTOMER COMPLAINT WAS CONFIRMED. THE CARD FAILED AT ISOLATION TEST. SUPPRESSOR DIODES D52, D1005, D63 WERE FOUND FAULTY, THAT CAUSED REPORTED PROBLEM. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED BY THE MANUFACTURER AND NO ANOMALIES WERE NOTED IN MANUFACTURING OR SERVICING OF THIS EQUIPMENT. MANAGEMENT IS NOTIFIED OF FAILURE ANALYSIS THROUGH THE MONTHLY TRENDING REPORTS. NO SIGNIFICANT TRENDS HAVE BEEN IDENTIFIED AT THIS TIME; THEREFORE NO CAPA ACTIVITY IS REQUIRED.><THE HARDWARE INVESTIGATION HAS BEGUN, BUT IT HAS NOT BEEN COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4).><IT WAS REPORTED THAT A PATIENT, UNDERWENT A PROCEDURE WITH A CARTO® 3 SYSTEM, AND NOISE WAS REPORTED ON PROXIMAL AND DISTAL ELECTRODE OF THE ABLATION CATHETER IN CARTO SYSTEM AND EP RECORDER. THERE WAS NO PATIENT CONSEQUENCE. THE PROCEDURE WAS PROLONGED FOR 60 MINUTES AND CANCELLED DUE TO THE ISSUE WITHOUT TREATING THE ARRHYTHMIA. THE PROCEDURE WAS BEING PERFORMED IN THE LEFT ATRIUM AND A TRANSSEPTAL PUNCTURE WAS PERFORMED PRIOR TO THE CASE CANCELLATION. BASED ON THE INFORMATION PROVIDED, STATING THAT THE PATIENT UNDERWENT CARDIAC ABLATION PROCEDURE UNDER GENERAL ANESTHESIA AND TRANSSEPTAL PUNCTURE HAD BEEN PERFORMED PRIOR TO THE CASE TERMINATION. THE PROCEDURE HAD TO BE TERMINATED DUE TO CARTO SYSTEM MALFUNCTION. THEREFORE, CARDIAC ABLATION CATHETERS ARE NOT CONSIDERED AS CONTRIBUTORY TO THE PROCEDURE CANCELLATION. THERE WERE NO CONSEQUENCES TO THE PATIENT. BY PHYSICIAN ASSESSMENT THE RISK TO PATIENT IS CONSIDERED TO BE LOW FOR THIS MALFUNCTION. | 1
4274757 | EVALUATION SUMMARY: VERIFIED THE UNIT SHUTS DOWN AND DISPLAYS THE ERROR MESSAGE ¿RUN TIME ERROR 1450¿ DURING OVER NIGHT BURN-IN TESTING. CONTAMINATION WAS FOUND ON THE MAINBOARD. THE MAINBOARD WAS REPLACED AND COMPLETE BURN-IN TESTING PER EDWARDS LIFESCIENCES¿ PROCEDURE WITH A SIMULATOR WAS PERFORMED WITHOUT ISSUE. THE UNIT ALSO PASSED FINAL TESTING OF THE SYSTEM. THE DEVICE HISTORY RECORD WAS REVIEWED AND THIS DEVICE PASSED ALL MANUFACTURING INSPECTIONS WITH NO NON-CONFORMANCES IDENTIFIED. ADDITIONAL INFORMATION: IT HAS BEEN DETERMINED THAT THE ROOT CAUSE OF THIS EVENT WAS RELATED TO MISHANDLING OF THE DEVICE, AS CONTAMINATION WAS FOUND ON THE MAINBOARD. MOREVOER, PER EDWARDS' PROCEDURE LIFETIME OF THE PRODUCT SPECIFICATION | 1
3656077 | IT WAS REPORTED THAT A PATIENT INCURRED A 3CM LINEAR FULL THICKNESS BURN IN THE RIGHT ATRIUM DURING AN AORTIC ROOT REPLACEMENT AND TRICUSPID VALVE REPAIR. THE BURN WAS ON THE LATERAL ATRIAL WALL ADJACENT TO A THERMAL FILAMENT OF A CCO SWAN GANZ CATHETER. THE SURGEON STATED THAT IT ¿LOOKED LIKE AN ABLATION BURN.¿ THE SURGEON STATED THAT SHE WAS ABLE TO ¿EASILY SUTURE REPAIR THE INJURY.¿ SHE ALSO STATED THAT THE CCO FEATURE OF THE SWAN GANZ CATHETER WAS TURNED OFF WHEN THE PATIENT WENT ON BYPASS AND WORKED WELL AFTER THE CASE. NO PATIENT COMPLICATIONS WERE NOTED. THE PATIENT WAS DISCHARGED ON POD #6.><EVALUATION SUMMARY: EXAMINATION FOUND THAT A CONTINUITY TEST WAS PERFORMED USING A MULTI-METER WHICH RESULTED IN AN OPEN WIRE CONNECTION AT S1-3 TO P1-13 AND S1-1 TO P1-14. THEY ALSO TESTED THE CABLE ON A KNOWN GOOD, CALIBRATED VIGILANCE II MONITOR, RESULTING IN TEST FAILURE WITH ¿FAIL REPLACE CABLE¿ ERROR MESSAGE. THE DEVICE HISTORY RECORD REVIEW WAS NOT PERFORMED BECAUSE THE LOT NUMBER IS UNKNOWN.><THE DEVICE EVALUATION IS ANTICIPATED. HOWEVER THE COMPLIANT COULD NOT BE CONFIRMED WITHOUT THE COMPLETION OF THE PRODUCT EVALUATION. A SUPPLEMENTAL REPORT WILL BE SENT WHEN THE INVESTIGATION IS COMPLETE.><A 70CC2 CABLE WAS RETURNED FOR EVALUATION. SINCE THE CUSTOMER WAS UNCERTAIN WHICH CABLE WAS USED, 3 CABLES MAINTAINED IN THE OR SUITE WERE RETURNED TO EDWARDS FOR EVALUATION. (RELATED COMPLAINT #S (B)(4). THE EVALUATION OF THE STATED COMPLAINT CABLE WAS INDEPENDENT OF THE RETURNED MONITOR TO VERIFY FUNCTIONALITY. VISUAL INSPECTION OF THE CABLE SHOWED NO PHYSICAL DAMAGE. CONTINUITY TESTING USING A CALIBRATED MULTIMETER SHOWED OPEN WIRE CONNECTIONS. DURING FUNCTIONAL TESTING THE MONITOR DISPLAYED ¿FAIL ¿ REPLACE CABLE.¿ THE CABLE WAS FOUND TO BE NON-FUNCTIONAL. ALTHOUGH IT CANNOT BE CONFIRMED, THE CONDITION OF THE RETURNED CABLE IS CONSISTENT WITH IMPROPER HANDLING AND/OR STORAGE BY THE CUSTOMER, AND NOT A PRODUCT MALFUNCTION OR DEFECT. VISUAL INSPECTION OF THE CABLE FROM RELATED COMPLAINT (B)(4) SHOWED NO PHYSICAL DAMAGE. CONTINUITY TESTING USING A CALIBRATED MULTIMETER SHOWED OPEN WIRE CONNECTIONS. DURING FUNCTIONAL TESTING THE MONITOR DISPLAYED ¿FAIL ¿ REPLACE CABLE.¿ THE CABLE WAS FOUND TO BE NON-FUNCTIONAL. THE CABLE WAS PAST ITS USEFUL LIFE. ALTHOUGH IT CANNOT BE CONFIRMED, THE CONDITION OF THE RETURNED CABLE IS CONSISTENT WITH IMPROPER HANDLING AND/OR STORAGE BY THE CUSTOMER, AND NOT A PRODUCT MALFUNCTION OR DEFECT. VISUAL INSPECTION OF THE CABLE FROM RELATED COMPLAINT (B)(4) SHOWED PHYSICAL DAMAGE WITH ONE OF THE CABLE CONNECTORS OVER-MOLD BROKEN FROM THE CABLE. CONTINUITY TESTING USING A CALIBRATED MULTIMETER SHOWED OPEN WIRE CONNECTIONS. DURING FUNCTIONAL TESTING THE MONITOR DISPLAYED ¿FAIL ¿ REPLACE CABLE.¿ THE CABLE WAS FOUND TO BE NON-FUNCTIONAL. THE CABLE WAS PAST ITS USEFUL LIFE. ALTHOUGH IT CANNOT BE CONFIRMED, THE CONDITION OF THE RETURNED CABLE IS CONSISTENT WITH IMPROPER HANDLING AND/OR STORAGE BY THE CUSTOMER, AND NOT A PRODUCT MALFUNCTION OR DEFECT. THE MONITOR, RELATED COMPLAINT (B)(4), WAS EVALUATED. VISUAL EXAMINATION OF THE DEVICE WAS COMPLETED. A MANUAL INSPECTION OF THE DEVICE DID NOT NOTE ANY COSMETIC IRREGULARITIES, DAMAGE OR DIFFERENCES NOTICED FROM A NEW UNIT OF THE SAME REVISION AND PART NUMBER. THE MONITOR PERFORMED NORMALLY IN ALL TEST CONDITIONS. FUNCTIONAL TESTING OF THE SYSTEM WAS CONDUCTED IN VARIOUS FLOW CONDITIONS TO CONFIRM FUNCTIONALITY OF THE THERMAL ELEMENT CONTROLS AND SAFETY SYSTEMS. THE SYSTEM WAS AGGRESSIVELY CHALLENGED BY REDUCING FLOW AROUND THE CATHETER WHILE IN CCO MODE AND THE SYSTEM PERFORMED AS DESIGNED IN ALL TEST CONDITIONS. THE SWAN GANZ CATHETER, RELATED COMPLAINT (B)(4), WAS RECEIVED IN THE EVALUATION LAB FOR VISUAL EXAMINATION, PHOTOS AND X-RAY. VISUAL EXAMINATION DID NOT FIND ANY ABNORMALITIES WITH THE CATHETER BODY, THERMAL FILAMENT, EXTENSION TUBE'S, HUBS OR THERMISTOR AND THERMAL FILAMENT CONNECTORS. THERMAL ELEMENT RESISTANCE VS. TEMPERATURE TEST WAS PERFORMED. THE RESISTANCE OF THE THERMAL ELEMENT OF THE CATHETER AND 5 ADDITIONAL CONTROL UNITS WERE MEASURED IN SPECIFIC FLUID BATH TEMPERATURES. IT WAS CONCLUDED THAT THE RESISTANCE OF THE THERMAL ELEMENT OF THE CATHETER MATCHED THE RESISTANCE THAT WAS INITIALLY PROGRAMMED IN THE CATHETER AS RECORDED IN THE EEPROM VALUES. THERE WERE NO NOTABLE DIFFERENCES BETWEEN THE CATHETER AND THE CONTROL UNITS TESTED. THEREFORE IT WAS CONCLUDED THAT THE RESISTANCE OF THE THERMAL ELEMENT OF THE CATHETER OPERATED WITHIN EXPECTED PARAMETERS. EDWARD¿S PERSONNEL VISITED THE CLINICAL SITE TO GATHER FURTHER INFORMATION REGARDING THE CASE. THE TIMELINES FROM THE ANESTHESIA RECORD AND VIGILANCE II MACHINE WERE REVIEWED. THE ANESTHESIOLOGIST STATED THAT THERE WERE NO PROBLEMS WITH THE INSERTION OF THE CATHETER AND THAT THE PLACEMENT WAS VERIFIED BY ECHO. HE SPOKE ABOUT THE CATHETER BEING PULLED BACK DURING THE REPAIR OF THE PULMONARY ARTERY AND THE REPOSITIONING AFTER COMPLETION OF THE REPAIR. REGARDING THE INJURY, HE STATED THAT THE INJURY WAS ¿VERY SUPERFICIAL¿ AND THERE WAS LITTLE CONCERN ABOUT IT IN THE OR. HE MENTIONED THAT THE INJURY WAS NOT TRANSMURAL AND WAS NOT NOTED UNTIL THE RIGHT ATRIUM WAS OPENED FOR THE TRICUSPID VALVE REPAIR. HE STATED ¿BELIEVE ME, IF THERE WAS ANYTHING OTHER THAN A SUPERFICIAL INJURY, YOU WOULD HAVE GOTTEN A PHONE CALL IMMEDIATELY. HE CONFIRMED THAT THEY CONTINUED TO USE THE CATHETER WITHOUT INCIDENT UNTIL IT WAS REMOVED ON POST OPERATIVE DAY #1. THE PATIENT RECOVERED UNEVENTFULLY. THE PATIENT WAS DISCHARGED ON POST OPERATIVE DAY #6. SEE RELATED MDR SUBMISSIONS: 2015691-2014-00501 (70CC2 CABLE) | 1
4334742 | IT WAS REPORTED THE DISPLAYED VALUES WERE INCORRECT. AFTER THAT, THE CUSTOMER RETRIED TO ZERO ADJUSTMENT BUT IT WAS NOT PERFORMED. THE CABLE WAS REPLACED AND THE SYMPTOM WAS IMPROVED. THE SPECIFIC DETAIL OF THE INCORRECT VALUE IS UNKNOWN. IT IS ALSO UNKNOWN WHETHER ANY ALARM OR ERROR MESSAGE WAS DISPLAYED. NO REPORT OF ANY PATIENT INJURY OR IMPROPER TREATMENT DUE TO THE INACCURATE VALUES.><EVALUATION SUMMARY: PERFORMANCE TESTING WAS COMPLETED WITHOUT ISSUE. NO PHYSICAL DAMAGE WAS FOUND DURING THE VISUAL INSPECTION. ADDITIONAL MANUFACTURER NARRATIVE: THE CABLE CAN REASONABLY BE EXEMPTED AS A POSSIBLE CONTRIBUTOR TO THE CUSTOMER¿S COMPLAINT. IT IS UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE CUSTOMER¿S EXPERIENCE, AS THERE WAS NO EVIDENCE OF ANY FAILURE OF THE DEVICE TO FUNCTION APPROPRIATELY.><THIS DEVICE IS A TYPE OF EDWARDS TRUWAVE CABLE WHICH IS SOLD IN (B)(4) ONLY. THE SPECIFIC BRAND NAME IN (B)(4) AND MODEL NUMBER ARE CURRENTLY UNKNOWN. ADDITIONAL INFORMATION: INACCURATE HEMODYNAMIC READINGS, WITHOUT ANY TYPE OF ALARM, COULD LEAD TO INAPPROPRIATE PATIENT TREATMENT. IN THIS CASE, IT WAS NOT REPORTED WHETHER THERE AN ALARM OCCURRED. THERE IS ALSO NO EVIDENCE OF ANY PATIENT COMPROMISE AS A RESULT OF THE INACCURRATE VALUES. UNFORTUNATELY, THE REPORTED EVENT CANNOT BE CONFIRMED WITHOUT THE SUBJECT DEVICE. THE CABLE IS EXPECTED TO BE RETURNED FOR EVALUATION; HOWEVER, IT HAS NOT BEEN RECEIVED AT THIS TIME. IF THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH OUR FINDINGS. NO FURTHER ACTIONS AT THIS TIME. | 1
4061737 | EVALUATION SUMMARY: THE REPORTED EVENT COULD NOT BE CONFIRMED AFTER 18 HOURS OF BURN-IN SIMULATION TESTING. IN ADDITION, THE SYSTEM-RELATED DEVICES (70CC2 AND CO-SET PROBES) WERE TESTED AND NO FAULT WAS FOUND. PREVENTIVE MAINTENANCE WAS COMPLETED SUCCESSFULLY. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED AND THIS DEVICE PASSED ALL MANUFACTURING INSPECTIONS WITH NO NON-CONFORMANCES IDENTIFIED. DHR REVIEW. IT IS UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE CUSTOMER¿S EXPERIENCE, AS THERE WAS NO ISSUE FOUND WITH THE DEVICE. PER ADDITIONAL FOLLOW UP, THE CUSTOMER INDICATED THE VALUES SEEN WERE NOT COMPLETE ABNORMAL; HOWEVER, HE RETURNED THE DEVICE FOR EVALUATION AS A PRECAUTION. THERE HAVE BEEN NO OTHER OCCURRENCES OF THIS ISSUE SINCE THE MONITOR HAS BEEN RETURNED TO THE CUSTOMER. NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED AT THIS TIME.><AS REPORTED, THE CARDIAC OUTPUT VALUE WAS TOO LOW ON A VIGILANCE 2 MONITOR. THE SPECIFIC VALUES BEING OUTPUTTED WERE NOT PROVIDED. IT IS UNKNOWN IF THERE WAS AN ERROR MESSAGE ON THE MONITOR; THE TECHNICIAN IN CHARGE OF THIS MATERIAL IN THE HOSPITAL DOES NOT RECALL THIS INFORMATION. TROUBLE-SHOOTING WAS PERFORMED; HOWEVER, THE ISSUE WAS NO RESOLVED. NO PATIENT COMPROMISE WAS REPORTED AND NO ADDITIONAL SYSTEM-RELATED DEVICES WERE IDENTIFIED AS SUSPECT.><THE MONITOR HAS NOT BEEN RETURNED FOR EVALUATION AT THIS TIME; HOWEVER, FOLLOW UP ATTEMPTS ARE IN PROGRESS TO OBTAIN THE SUBJECT DEVICE. IT IS UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE CUSTOMER¿S EXPERIENCE. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IN THE EVENT THE DEVICE OR ANY NEW INFORMATION IS RECEIVED. NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED AT THIS TIME. | 1
4192347 | AT 0600, NURSE NOTICED CARDIAC INDEX BEGAN TO DECLINE. AFTER CHECKING MONITORING SYSTEM TO ENSURE ALL CONNECTIONS WERE OKAY AND NO ERRORS ON THE SCREEN WERE PRESENT, THE MONITOR SCREEN SUDDENLY TURNED BLUE WITH AN ERROR CODE AND SHUT OFF. THE MONITOR AND CABLES WERE IMMEDIATELY REPLACED AND THE MONITOR TAGGED. ANOTHER DEVICE WAS OBTAINED AND PATIENT MONITORING CONTINUED. NO PATIENT HARM. | 1
3918335 | ON (B)(6) 2014, AT 07:15 AM, IN THE MEDICAL-SURGICAL STORAGE AREA - NURSE RECEIVED A BURN WHEN PLUGGING IN THE COMPUTER ON WHEELS. NURSE RECEIVED A SMALL BURN ON HER RIGHT MIDDLE FINGER WHILE PLUGGING THE COMPUTER ON WHEELS INTO THE ELECTRICAL SOCKET. NURSE NOTICED A SMALL (APPROX 0.5 CM) OPEN AREA IN THE CORD OF THE COMPUTER ON WHEELS NEXT TO THE PLUG IN POSITION. THE NURSE WAS IMMEDIATELY SENT TO THE EMERGENCY DEPARTMENT FOR EVAL. THE COMPUTER ON WHEELS HAS BEEN TAKING OUT OF SERVICE. BIO-ENGINEERING HAS BEEN NOTIFIED OF THE PROBLEM. THE NURSE WAS SEEN IN THE EMERGENCY DEPARTMENT. SHE HAS A 1CM SECOND DEGREE BURN TO RADIAL ASPECT OF THE FINGER - RIGHT 3RD DIGIT. SHE WAS TREATED WITH ANTIBIOTIC OINTMENT AND BANDAGE PLACED ON FINGER DISCHARGED WITH PRESCRIPTION FOR LIDOCAINE 5% EXTERNAL CREAM EVERY 4 HOURS AS NEEDED FOR PAIN. PT WAS DISCHARGED - STABLE CONDITION. | 1
3835799 | PT WAS ADMITTED TO A HOSPITAL ON (B)(6) 2013 FOR A PARTIAL HYSTERECTOMY. SURGERY WAS PERFORMED ON (B)(6) 2013. AFTER SURGERY PT SHOULD HAVE AWAKENED AFTER APPROXIMATELY 3 HOURS. TEN HOURS FOLLOWING SURGERY PT WAS PLACED IN ICU. DOCTORS EXPLAINED THAT MEDICATION CAUSED HER TO SLEEP LONGER. PT WAS RELEASED FROM ICU THE FOLLOWING DAY. PT NOTICED CLEAR LIQUID LEAKING FROM HER VAGINA AND SHE HAD TROUBLE BREATHING. SHE SAW DIFFERENT DOCTORS 3 TIMES AND THEN FINALLY WENT TO THE EMERGENCY ROOM. CT SCAN SHOWED SHE HAD PNEUMONIA AND WAS PLACED ON ANTIBIOTICS. FIVE DAYS LATER SHE STILL HAD LEAKING AND TROUBLE BREATHING SO SHE WENT TO A DIFFERENT EMERGENCY ROOM. CT SCAN SHOWED DURING THE ORIGINAL SURGERY THE DOCTOR CUT THE URETER. ONCOLOGIST RECOMMENDED THAT SHE CONTACT HER ORIGINAL SURGEON. SHE HAS APPREHENSIONS ABOUT IT BECAUSE HE MESSED UP THE ORIGINAL SURGERY. HER SURGERY IS SCHEDULED FOR NEXT MONTH. SHE HAS LOST HER JOB BECAUSE OF COMPLICATIONS FROM HER SURGERY. SHE HAS NO MEDICAL INSURANCE AND IS LOOKING INTO BRINGING SUIT. IT HAS BEEN ABOUT A YEAR AND SHE IS STILL EXPERIENCING PAIN. | 1
3590667 | VESSEL SEALER BLADE WILL NOT RETRACT WHEN OPENING THE TIP.THIS IS THE 8TH REPORT OF THIS DEVICE AT OUR FACILITY SINCE MARCH OF 2013. | 1
3723671 | DURING THE ROBOTIC SURGERY, THE DOCTOR NOTICED THAT THE CAUTERY TOOL SHE WAS USING WAS NOT FUNCTIONING AS EXPECTED. UPON VIDEO INSPECTION, IT WAS DISCOVERED THAT A SMALL PIECE OF PLASTIC NEAR THE TIP OF THE DEVICE WAS BROKEN OFF SUCH THAT THE TIP OF THE CAUTERY TOOL WAS AT AN ANGLE. THE DOCTOR PROCEEDED TO IRRIGATE AND SUCTION THE SURGICAL SITE IN AN ATTEMPT TO LOCATE THE BROKEN PIECE OF PLASTIC. THE STAFF WAS UNABLE TO LOCATE THE MISSING PIECE OF THE INSTRUMENT. | 1
3670221 | ON (B)(6) 2013, THE CUSTOMER USED THE ROSA DEVICE TO GUIDE THE IMPLANTATION OF 8 DEPTH ELECTRODES FOR A CRANIAL SEEG. DURING THE PROCEDURE, AN INTRACRANIAL HEMORRHAGE OCCURRED; IT WAS DETECTED LATER ON WITH THE POST-OPERATIVE CT-SCAN OF THE PT AND TREATED BY ADDITIONAL SURGERY.><A MEDTECH REPRESENTATIVE PERFORMED NON-CLINICAL TESTS WITH THE ROSA ROBOTIC SYSTEM AT THE HOSPITAL TO FURTHER INVESTIGATE CUSTOMER COMPLAINT. NO INACCURACY WAS FOUND AND SYSTEM WAS FOUND FULLY FUNCTIONAL. CUSTOMER HEAD FRAME COULD NOT HAVE BEEN TESTED YET. INCIDENT STILL UNDER INVESTIGATION. | 1
4081095 | ORTHALIGN REVIEWED THE INJURY WITH THE SURGEON AND DISCUSSED POSSIBLE CAUSES. THE SURGEON NOTED THAT BECAUSE NO COMPLICATIONS WERE OBSERVED DURING THE CASE, IT WOULD PROBABLY NOT BE POSSIBLE TO DETERMINE THE CAUSE. NOTE: THAT REUSE" | 1
4367008 | ON (B)(6) 2014, BAXTER (B)(4) RECEIVED A CALL FROM (B)(6), REQUESTING ASSISTANCE IN RESTORING AN ARCHIVED PATIENT FROM 2012 WITHIN THEIR ABACUS® SOFTWARE, WHICH IS USED WITH THE BAXTER EM2400 EXACTAMIX COMPOUNDER. THE EXACTAMIX COMPOUNDER IS AN AUTOMATED PUMPING SYSTEM THAT COMPOUNDS MULTIPLE STERILE INGREDIENTS INTO A FINISHED SOLUTION, AND IS DESIGNED FOR THE PREPARATION OF TOTAL PARENTERAL NUTRITION (TPN). DURING THE DISCUSSION, THE CUSTOMER NOTED THAT AN ORDER WAS ADMINISTERED PERIPHERALLY TO THE PATIENT AND IT INFILTRATED AND GOT UNDER THE SKIN. THE CUSTOMER STATED THAT A WARM COMPRESS WAS APPLIED TO THE AREA. IT IS UNKNOWN AT THIS TIME IF THE BAXTER PRODUCTS CONTRIBUTED TO THE EVENT; HOWEVER, WE ARE FILING A MDR REPORT BECAUSE, IF REPEATED, AN IV INFILTRATION OF A VESICANT SOLUTION CAN PRODUCE CONDITIONS OF VARYING DEGREES WHICH CAN LEAD TO INJURY REQUIRING MEDICAL INTERVENTION. SHOULD ADDITIONAL INFORMATION REGARDING THIS EVENT BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED.><METHOD: NO TESTING METHODS PERFORMED. RESULTS: NO RESULTS AVAILABLE SINCE NO EVALUATION PERFORMED. CONCLUSIONS: THE CUSTOMER CALL WAS REGARDING RESTORING AN ARCHIVED PATIENT WITHIN THE ABACUS SOFTWARE. THE INFORMATION REGARDING A PATIENT INVOLVED EVENT WAS DISCOVERED THROUGH THE CONVERSATION WITH TECH SUPPORT AND OCCURED UP TO TWO YEARS AGO. SHOULD ADDITIONAL INFORMATION REGARDING THIS EVENT BECOME AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. | 1
4133150 | ON (B)(6) 2014 WE WERE INFORMED OF AN INCIDENT INVOLVING OUR ABACUS® SOFTWARE. ABACUS® IS OUR WINDOWS®-BASED ORDER ENTRY SOFTWARE WHICH IS USED WITH OUR EM2400 EXACTAMIX COMPOUNDER. THE EXACTAMIX COMPOUNDER IS AN AUTOMATED PUMPING SYSTEM THAT COMPOUNDS MULTIPLE STERILE INGREDIENTS INTO A FINISHED SOLUTION, AND IS DESIGNED FOR THE PREPARATION OF TOTAL PARENTERAL NUTRITION (TPN). IN THIS CASE THE CUSTOMER STATED THAT THEIR ABACUS TEMPLATE HAD DEXTROSE GRAMS/KG/DAY AND THE DOCTORS` ORDER WAS GRAMS/BAG AND THEY WOULD CALCULATE TO %/BAG AND ENTER THAT VALUE INTO ABACUS. BAXTER TECHNICAL SUPPORT INFORMED THE CUSTOMER THAT THIS CONVERSION WAS NOT CORRECT. TECHNICAL SUPPORT, ALONG WITH THE BAXTER (B)(4) MANAGER OF MEDICAL AFFAIRS PROVIDED ASSISTANCE TO CORRECT THEIR ABACUS TEMPLATE TO REFLECT THE DENOMINATOR AND NUMERATOR OF HOW THE DOCTOR IS ORDERING THE DEXTROSE. THERE WAS NO ADVERSE EVENT INVOLVED AND NO INTERVENTION WAS REQUIRED; HOWEVER, WE ARE FILING A MDR REPORT BECAUSE, IF REPEATED, THE EVENT COULD REQUIRE INTERVENTION OR CONTRIBUTE TO A SERIOUS PATIENT INJURY. THE EVENT WAS NOT CAUSED BY A DEFECT OR MALFUNCTION OF THE ABACUS SOFTWARE. IF ADDITIONAL INFORMATION REGARDING THIS EVENT BECOMES AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED.><INVESTIGATION SUMMARY: METHOD: SOFTWARE EVALUATION: · THE CUSTOMER STATED THAT THE ABACUS TEMPLATE HAD DEXTROSE GRAMS/KG/DAY AND THE DOCTORS' ORDER WAS GRAMS/BAG SO THEY WOULD CALCULATE TO %/BAG AND ENTER THAT VALUE INTO ABACUS, WHICH RESULTED IN THE INCORRECT DELIVERY OF DEXTROSE TO A PATIENT. · BAXTER TECHNICAL SUPPORT SET UP A FILE EXCHANGE LOGIN SO THE CUSTOMER COULD UPLOAD THEIR DATABASE AND FAX THE DOCTORS' ORDER AND MIX CHECK REPORT FOR REVIEW. RESULTS: DEVICE PERFORMED ACCORDING TO SPECIFICATIONS: · DURING THE REVIEW OF THE DOCUMENTS, THE BAXTER MANAGER OF MEDICAL AFFAIRS DISCOVERED THAT THE DILUTIONAL WARNING LIMIT WAS ONLY ATTACHED TO THE ADULT PATIENT TYPE. · TECHNICAL SUPPORT ASSISTED THE CUSTOMER TO ASSOCIATE THE DILUTIONAL WARNING LIMIT TO THE NEONATAL PATIENT TYPE WHICH RESOLVED THE ISSUE. CONCLUSIONS: DEVICE PERFORMED ACCORDING TO SPECIFICATIONS. USER ERROR CAUSED EVENT: · THERE WAS NO DEFECT OR MALFUNCTION FOUND WITH THE ABACUS SOFTWARE. THE EVENT WAS DUE TO A USER ERROR. SUMMARY: THE INVESTIGATION FOUND THAT THE DILUTIONAL WARNING LIMIT WAS ONLY ATTACHED TO THE ADULT TEMPLATE. TECHNICAL SUPPORT ASSOCIATED THE MAX DILUTION TO THE NEONATAL PATIENT TYPE AND VERIFIED THAT THE WARNING LIMIT WAS FLAGGED PROPERLY. NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION REGARDING THIS EVENT BECOMES AVAILABLE, A FOLLOW-UP REPORT WILL BE SUBMITTED. | 1
3841072 | ><IT WAS REPORTED THAT DURING SERVICING OF THE BRIGHTSPEED CT SYSTEM, A GE EMPLOYEE REMOVED THE GANTRY COVERS. WHEN PUTTING THE FRONT COVER BACK ON, THE EMPLOYEE FELT RESISTANCE. THE EMPLOYEE PUSHED HARDER AND PINCHED THEIR 5TH DIGIT (FINGER) IN BETWEEN THE COVER AND THE CT GANTRY. IN THE EMERGENCY ROOM, A HAND SURGEON REMOVED THE FINGERNAIL, STRAIGHTENED" THE BROKEN DISTAL PHALANGE AND APPLIED A SPLINT AND BANDAGES." | 0
3961367 | A PT WAS GIVEN THEIR SCHEDULED DOSE OF LISINOPRIL BY A NURSE, THEN GIVEN THE SAME DOSE AGAIN BY ANOTHER NURSE 9 MINUTES LATER. THE CLINICAL SOFTWARE IN USE IS CPSI. IT SHOULD PLACE A G ON THE MEDICAL ADMINISTRATION RECORD TO SHOW THAT A DOSE WAS GIVEN, BUT THE SECOND ADMINISTERING NURSE DID NOT SEE IT. IT IS POSSIBLE BOTH NURSES HAD THE PT'S PROFILE UP AT THE SAME TIME, AND IT DOES NOT UPDATE IN REAL TIME. THIS COULD HAVE BEEN AVOIDED BY NOT HAVING TWO NURSES PASSING MEDICATION TO THE SAME PT AT THE SAME TIME. PT COUNSELING PROVIDED: UNK. RELEVANT MATERIALS PROVIDED: NONE. (B)(6). | 0
4218956 | COMPANY REP EVALUATED THE SYSTEM. CORRECTIONS INCLUDED REPLACEMENT OF THE SYSTEM'S POWER SUPPLY, CPU FAN AND MOTHERBOARD. SYSTEM'S ERROR LOGGED WERE CLEARED AND SYSTEM WAS REBOOTED.><CUSTOMER REPORTED THE PANORAMA CENTRAL STATION HAS LOST COMMUNICATION, WHICH MAY HAVE AFFECTED TELEMETRY MONITORING. NO PT INJURY WAS REPORTED. | 0
3813549 | COMPANY REP EVALUATED THE SYSTEM. CORRECTIONS INCLUDED CLEARING THE SYSTEM'S ERROR LOGS AND REBOOTING. COMMUNICATION WAS REESTABLISHED.><CUSTOMER REPORTED THE PANORAMA CENTRAL STATION HAD LOST COMMUNICATION, WHICH MAY HAVE AFFECTED TELEMETRY MONITORING. NO PT INJURY WAS REPORTED. | 0
4319836 | CUSTOMER REPORTED AN ISSUE WITH THE DPM CENTRAL STATION, WHICH MAY HAVE AFFECTED TELEMETRY MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE SYSTEM. CORRECTIONS INCLUDED REPLACEMENT OF THE SYSTEM'S HARD DRIVES. SYSTEM WAS CALIBRATED AND SAFETY TESTED TO FACTORY'S SPECS. | 0
4197365 | TECHNICAL SUPPORT WALKED THE CUSTOMER THROUGH SWAPPING THE HARD DISK DRIVE 2 WITH HDD 1 AND REBOOTED CNS AND APPLICATION AND IT RESUMED PATIENT MONITORING. DEVICE IS BEING SENT IN FOR EVALUATION.><(B)(4).><ADDITIONAL MANUFACTURER NARRATIVE: DEVICE WAS RETURNED AND EVALUATED. CUSTOMER COMPLAINT WAS CONFIRMED. THE UNIT CAME IN AND WOULD NOT BOOT UP COMPLETELY. AFTER SEVERAL TRIES THIS UNIT DID BOOT UP AND INDICATED THAT IT HAD A CORRUPT FILE. THE HARD DRIVES WERE REIMAGED TO FIX THIS ISSUE. THIS UNIT WAS BURNED IN AND TESTED WITH TWO ORGS CONNECTED. THE REPAIRED UNIT WAS SENT BACK TO THE CUSTOMER. | 0
4143675 | DEVICE EXCHANGED AND RETURN OF DEFECTIVE DEVICE REQUESTED FOR FAILURE INVESTIGATION.><(B)(4). | 0
4143659 | (B)(4).><ERROR MESSAGES INDICATIVE OF A HARD DRIVE (HDD) FAILURE. DEFECTIVE DEVICE REQUESTED FOR REPAIR AND FAILURE INVESTIGATION. | 0
3939003 | TECH SUPPORT DISCUSSED OEM RECOMMENDATION FOR CONTROLLED SHUTDOWN RESTART EVERY QUARTER AND REPLACEMENT OF HDD EVERY TWO YEARS. BME FROM STATES THEY ONLY RESTART THE SYSTEMS SEMI-ANNUALLY WHEN PERFORMING CLEAN AND INSPECT. ALSO HE BELIEVES THEY HAVEN'T REPLACED DRIVES IN ABOUT 3 YEARS. HE WILL PERFORM SHUTDOWN/RESTART AND CHECK THEIR RECORDS ON LAST HDD REPLACEMENT.><(B)(4). | 0
4143654 | (B)(4).><THE CUSTOMER REPLACED IT WITH THE SPACE MONITOR THAT THEY HAD AND RETURN OF DEFECTIVE DEVICE REQUESTED FOR FAILURE INVESTIGATION. | 0
4197366 | (B)(4).><ERROR MESSAGES INDICATIVE OF A HARD DRIVE (HDD) FAILURE. WE DID SOME TROUBLESHOOTING WITH THE CUSTOMER AND VERIFIED THAT THE HARD DRIVE IN PORT 1 HAD FAILED. WE GUIDED THE CUSTOMER THROUGH POWERING THE UNIT DOWN AND REMOVING THE BAD DRIVE. INFORMED BME WE WOULD SEND A REPLACEMENT HD TO THEM. CUSTOMER REPLACING THE HARD DISK DRIVE THEMSELVES. REQUESTED RETURN OF FAILED DRIVE FOR FURTHER INVESTIGATION.><MANUFACTURER NARRATIVE: CUSTOMER REPORTED THE CENTRAL MONITORING SYSTEM (CNS) SHUTDOWN ON ITS OWN AND THE UNITS WAS SHOWING AN INTEL MATRIX STORAGE MANAGER ERROR MESSAGE. UNIT WAS ON A UPS AND THE UPS INDICATOR LIGHTS SHOWED NO ISSUES. ERROR MESSAGES INDICATIVE OF HARD DRIVE FAILURE. WE DID SOME TROUBLESHOOTING WITH THE CUSTOMER AND VERIFIED THAT THE HARD DRIVE IN PORT 1 HAD FAILED. WE GUIDED THE CUSTOMER THROUGH POWERING THE UNIT DOWN AND REMOVING THE BAD DRIVE. BIOMEDICAL ENGINEER WAS INFORMED WE WOULD SEND A REPLACEMENT HARD DRIVE TO THEM. CUSTOMER WAS GOING TO REPLACE THE HARD DRIVE THEMSELVES. DUE TO THE AGE OF THIS COMPLAINT NO ADDITIONAL INFORMATION IS AVAILABLE CONCERNING THIS COMPLAINT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | 0
4370541 | THE DEVICE REFERENCED IN THIS REPORT IS BEYOND ITS EXPECTED LIFE. THE RAM WAS REPLACED AND THE REPAIRED UNIT WAS RETURNED TO THE CUSTOMER.><IMP REF # (B)(4). | 0
4143708 | THE CUSTOMER STATED THAT THE CENTRAL MONITORING SYSTEM (CNS) IS BEEPING CONTINUOUSLY AFTER RESTARTING OF THE CNS. MFR REF # 8030225-2014-00037. | 0
3908496 | UPON REBOOT OF THE CENTRAL MONITORING SYSTEM (CNS), THE SYSTEM DOES NOT RESTART. REF MFR REPORT 8030229-2014-00004. | 0
4198591 | THE CENTRAL MONITORING SYSTEM (CNS) FROZE AND UNABLE TO ACCESS PATIENT INFORMATION. MFR REF # 8030229-2014-00052. | 0
4173615 | THE CENTRAL MONITORING SYSTEM (CNS) SCREEN WENT BLANK A FEW TIMES, EACH TIME THE CUSTOMER POWERED THE SCREEN ON THEN OFF. IT CAME BACK ON BUT WOULD GO DARK AGAIN. MFR REF #:8030229-2014-00045. | 0
3838145 | CLINICAL REPORTED THAT THEY HAD A PT THAT CODED AND THE SYSTEM DID NOT ALARM FOR 50 SECONDS AFTER THE EVENT. THEY EXPECTED THE SYSTEM TO ALARM VFIB AT 23:51 PM. REF # MFR 8030229-2014-00001. | 0
4173616 | THE CENTRAL MONITORING SYSTEM (CNS) JUST RECENTLY HAD A HARD DISK DRIVE REPLACEMENT. THE CUSTOMER IS STATING THE UNIT PERIODICALLY FREEZES UP A FEW TIMES A DAY. WHEN THE CNS FREEZES, ALL NUMERICS AND WAVEFORMS STOP. A HARD POWER CYCLE IS REQUIRED FOR THE UNITS TO MONITOR THE BEDS AGAIN. MFR REF#: 8030229-2014-00046. | 0
4347881 | THE CUSTOMER REPORTS CENTRAL MONITORING SYSTEM (CNS) IS DISPLAYING CORRUPT NTFS FILE". MFR REF #8030229-2014-00075." | 0
3952675 | THE CENTRAL MONITORING SYSTEM (CNS) FROZE UP, TECH THEN SHUT DOWN THE PC AND RESTARTED BUT THE CNS WOULD NOT POWER UP. MFR - 8030229-2014-00021. | 0
4173619 | THE CENTRAL MONITORING SYSTEM (CNS) HAS A BLUE SCREEN, AND THERE WAS NO CAUSAL EVENT AS FAR AS THE CUSTOMER KNOWS. UPON REBOOT, THE SCREEN STATES SHOWING AN INACCESSIBLE BOOT DEVICE" MESSAGE. MFR REF#:8030229-2014-00049." | 0
3418272 | CUSTOMER REPORTED THE PANORAMA CENTRAL STATION'S SERVER SHUT DOWN, WHICH MAY HAVE AFFECTED TELEMETRY MONITORING. NO PT INJURY WAS REPORTED.><COMPANY REP EVALUATED THE SYSTEM AND FOUND THAT THE ANTENNA REPEATERS LOST POWER DUE TO AN ISSUE WITH AN UNINTERRUPTED POWER SUPPLY (UPS). CUSTOMER WAS ADVISED TO REPLACE AND MAINTAIN BATTERIES ON ALL UPS. SYSTEM WAS TESTED TO FACTORY'S SPECIFICATIONS AND COMMUNICATION WAS REESTABLISHED. | 0
3606670 | THE PT'S CARDIAC MONITORING SYSTEM FAILED TO RECOGNIZE THE RHYTHM OF VENTRICULAR FIBRILLATION AND FAILED TO ALARM IN RESPONSE TO THE RHYTHM OF VENTRICULAR FIBRILLATION. THE ARRHYTHMIA WAS IMMEDIATELY RECOGNIZED AT THE NURSE'S CENTRAL MONITORING STATION. THE NURSES WERE IMMEDIATELY NOTIFIED. CPR WAS INITIATED WITHOUT DELAY. CPR" RESUSCITATION EFFORTS WERE NOT SUCCESSFUL AND THE PT DIED DURING THE CARDIAC ARREST. FOLLOWING THE CARDIAC ARREST | 0
3535031 | INITIAL REPORTER: THE CUSTOMER'S TELEPHONE NUMBER IS (B)(6). A BECKMAN COULTER FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER'S SITE. THE FSE EVALUATED THE INSTRUMENT AND DETERMINED THAT ROBOT 2 WAS OUT OF ALIGNMENT BY APPROXIMATELY 2 MILLIMETERS, CAUSING THE TUBES TO NOT FIT SOME OF THE POSITIONING RACKS LEADING UP TO THE REPORTED ISSUE. THE FSE DISCOVERED A LOOSE DRIVE BELT UP ON THE GANTRY AND PROCEEDED TO TENSION THE BELT. THE FSE THEN PERFORMED ALL RE-ALIGNMENTS AND VERIFIED THE REPAIR AS PER ESTABLISHED PROCEDURES.><THE CUSTOMER REPORTED THAT THE AUTOMATE¿ 1250 S3I SAMPLE PROCESSING SYSTEM THREW ONE SAMPLE AT A STUDENT WHO WAS STANDING IN FRONT OF THE INSTRUMENT. THE CUSTOMER REPORTED THAT THE SAMPLE SPLASHED ONTO THE STUDENT. THE STUDENT WAS WEARING PERSONAL PROTECTIVE EQUIPMENT CONSISTING OF PROTECTIVE CLOTHING, GOGGLES, WHITE COAT, AND GLOVES AT THE TIME OF THE EVENT BUT BLOOD SPATTERED ON THE STUDENT'S CHIN. THE STUDENT DID NOT SEEK ANY MEDICAL ATTENTION AND THERE WAS NO DIRECT CONTACT OF THE BLOOD TO THE EYES, EARS, NOSE, MOUTH OR MUCOUS MEMBRANE. THERE WAS NO CHANGE OR AFFECT TO PATIENT TREATMENT IN CONNECTION WITH THIS EVENT. | 0
3843572 | THE CUSTOMER REPORTED OBTAINING AN ERRONEOUSLY HIGH POTASSIUM (K) RESULT OF 5.7MMOL/L FROM A HEMOLYZED SAMPLE WHICH WAS RELEASED OUT OF THE LABORATORY DUE TO INCORRECTLY CONFIGURED RULE AT THE REMISOL ADVANCE DATA MGR. BASED UPON THE FALSELY ELEVATED K RESULT, THE CUSTOMER STATED THAT ONE (1) EMERGENCY ROOM PT WAS ADMITTED TO THE HOSPITAL AND ADMINISTERED INSULIN AND SODIUM POLYSTYRENE SULFONATE TO LOWER THE PT'S K LEVEL. FOLLOWING TREATMENT, THE PT WAS REDRAWN AND A K RESULT OF 3.3 MMOL/L WAS OBTAINED. THE PT WAS THEN GIVEN POTASSIUM INTRAVENOUSLY (IV) TO ELEVATE THE PT'S POTASSIUM LEVEL BACK TO THE NORMAL RANGE.><THE CUSTOMER CONFIGURED A RULE AT REMISOL ADVANCE DATA MGR TO ADD A HEMOLYSIS INDEX TEST TO ANY TEST ORDER CONTAINING K. HOWEVER, THE RULE WAS NOT CONFIGURED CORRECTLY AND REMISOL ADVANCE FAILED TO ADD THE HEMOLYSIS INDEX TEST. THE SAMPLE IN QUESTION WAS HEMOLYZED AND A FALSELY ELEVATED K RESULT WAS GENERATED. THE HEMOLYSIS INDEX TEST WOULD HAVE ALERTED THE USER TO THE PRESENCE OF HEMOLYSIS. HOWEVER, SINCE A HEMOLYSIS INDEX TEST WAS NOT RUN TO INDICATE THE PRESENCE OF HEMOLYSIS, THE FALSELY ELEVATED K RESULT WAS RELEASED OUT OF THE LABORATORY. IN CONCLUSION, FAILURE MODE OF THE EVENT IS ATTRIBUTED TO USE ERROR AS THE CUSTOMER INCORRECTLY CONFIGURED THE RULE AT THE REMISOL ADVANCE DATA MGR. APPLICATIONS SPECIALIST ASSISTED THE CUSTOMER WITH RESOLVING THE ISSUE OVER THE TELEPHONE BY CORRECTING THE RULE TO ADD A HEMOLYSIS INDEX TEST. | 0
4250760 | UNDER VERY SPECIFIC AND UNIQUE CIRCUMSTANCES, RESULTS FROM ONE PATIENT COULD BE ADDED TO A DIFFERENT PATIENT, LEADING TO A MISMATCH OF RESULTS FOR A GIVEN PATIENT. THIS IS DUE TO A HIGH LEVEL PROTOCOL (HLP) ISSUE IN JRESULTNET THAT AFFECTS THE MANNER IN WHICH NON TEST - LEVEL FIELD DATA IN THE RESULT (R) RECORD IS PARSED. THIS ONLY AFFECTS CERTAIN DRIVERS. | 0
3842635 | ><ACUDOSE ANESTHESIA CART MACHINE FROZE OR LOCKED WHILE DISPENSING TO PATIENT. TECH HAD STOCKED MEDS INTO CART EARLIER AND IT WORKED AT ITS NORMAL SPEED. LATER THAT MORNING, A CALL WAS RECEIVED FROM A TECH STATING HE HAD DISPENSED TWO ROCURONIUM VIALS TO HIS PATIENT. WHEN THE MACHINE CYCLED TO THE NEXT DRUG (LIDOCAINE), THE SCREEN FROZE. IT WAS IN THAT STATE FOR OVER 10 MINUTES AND AT THAT POINT THE MACHINE WAS REBOOTED. NO FURTHER ISSUES THE REMAINDER OF THE DAY HAVE BEEN REPORTED. THIS FACILITY HAS HAD SIMILAR PROBLEMS WITH THIS DEVICE OVER THE LAST YEAR AND CONTINUES TO SEE PROBLEMS LIKE THIS DESPITE UPGRADE SOFTWARE PATCHES FROM THE MANUFACTURER. | 0
3972800 | PULSE OXIMETRY ON DINAMAP MONITOR GIVING INACCURATE READINGS. PATIENTS DO NOT APPEAR TO HAVE READINGS THAT LOW SO COMPARISON DONE WITH ALTERNATE PULSE OXIMETER AND FOUND THAT THE DINAMAP WAS CONSISTENTLY RUNNING 5-10% LOWER THAN OTHER MONITOR. PROBLEM REMAINED. MFR NOTIFIED AND RESPONSE WAS TRY DIFFERENT FINGER PROBES AND CABLES, WHICH ALSO DID NOT WORK. | 0
4089924 | WELCH ALLYN IS REPORTING THIS EVENT BECAUSE IT WAS ASSOCIATED WITH A SERIOUS INJURY, ALTHOUGH THERE IS NO EVIDENCE OF MALFUNCTION OR CAUSATION AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVAL IS COMPLETE.><WELCH ALLYN RECEIVED AN VIGILANCE REPORT FROM (B)(6) IN REGARDS TO AN ALLEGATION OF A PT THAT HAD SUSTAINED A DEEP VEIN THROMBOSIS (DVT) ON THE RIGHT ARM WHILE USING A AUTOMATED BLOOD PRESSURE DEVICE. THE CUSTOMER DID NOT PROVIDE ANY PT INFO. | 0
3885662 | OVER THE LAST MONTH, WE HAVE HAD 5 EVENTS WHERE THE LUER LOCK CONNECTION ON THE EXTERNAL VENTRICULAR DRAIN (EVD) SYSTEM HAS CRACKED. THIS CONNECTION IS WHAT CONNECTS THE SYSTEM TO THE DRAINAGE BAG. AFTER THE FIRST EVENT, I WAS IN TOUCH WITH THE COMPANY. THEN, WITH THE ADDITIONAL CRACKS, THE REPRESENTATIVE INFORMED ME THAT THE COMPANY RECOMMENDS CLEANING THE CONNECTION WITH ALCOHOL AND OUR HOSPITAL WAS USING CHLORAHEXIDINE. WE CHANGED OUR CLEANING POLICY AND CHANGED OUR CLEANING AGENT TO ALCOHOL. SINCE THIS CHANGE, THERE HAS BEEN ANOTHER LUER LOCK CONNECTION THAT HAS CRACKED.MANUFACTURER RESPONSE: THE REPRESENTATIVE AND COMPANY HAVE BEEN VERY RESPONSIVE AND ARE GOING TO FURTHER INVESTIGATE THE DEFECTIVE PRODUCTS. THE REPRESENTATIVE AND QUALITY TEAM HAVE KEPT IN TOUCH WITH ME. | 0
3876231 | THE PATIENT HAD AN EXTERNAL VENTRICULAR DRAIN (EVD) WITH A BECKER DRAIN SET-UP. DURING A CHANGE IN POSITIONS, THE PROXIMAL TUBING FROM THE PATIENT END SLIPPED OUT OF THE STOPCOCK CEREBRAL SPINAL FLUID (CSF) COLLECTION PORT. AS A RESULT, THE ENTIRE BECKER DRAIN SYSTEM HAD TO BE REPLACED. THE NEUROSURGEON REPORTS THAT HE HAS SEEN THIS BEFORE AND STATED THAT THE GLUE, THAT HOLDS THE CATHETER INTO THE LUER LOCK END, GETS LOOSE AND DOES NOT HOLD WELL. AS A RESULT, THE CATHETER SLIPS OUT.THE MEDTRONIC REPRESENTATIVE PROMPTLY RETURNED MY CALL AND SET UP A TIME TO MEET WITH OUR FACILITY THE FOLLOWING DAY. MEDTRONIC WILL FILE A REPORT ON THIS EVENT AS WELL. | 0
4297378 | ON (B)(4) 2014, BIOMERIEUX WAS INFORMED THAT THE RESULTS OF NEGATIVE BACT/ALERT TESTS WERE ATTACHED TO THE IMPROPER PATIENT. BASED O THE DATA BACKUP THE CUSTOMER PROVIDED, THE HOSPITAL LIS (NOT A BIOMERIEUX PRODUCT) ASSIGNED THE SAME ACCESSION ID TO 2 PT SAMPLES WITHIN A 90 MINUTE PERIOD OF TIME. UNDER NORMAL OPERATING CONDITIONS, ACCESSION IDS ARE UNIQUE TO THE PT BEING TREATED. THIS ALLOWS FOR THE DATA TO BE PROCESSED AND ASSOCIATED WITH THE PROPER PATIENT RECORD WITHIN THE SYSTEMS. THE REPORTED EVENT SUGGESTS THE SYSTEM MAY HAVE MALFUNCTIONED; THEREFORE, AN INVESTIGATION HAS BEEN INITIATED. BASED ON THE DATA BACKUP PROVIDED THE BOTTLES INVOLVED WERE LOADED ON (B)(6) 2014. ALL BOTTLES INVOLVED WERE DETERMINED TO BE NEGATIVE; THEREFORE, NOT IMPACTING PT TREATMENT OR SAFETY IN THIS CASE. FOR POSITIVE SAMPLES, THE INSTRUCTIONS FOR USE INDICATE THAT THE SAMPLE SHOULD BE SUB-CULTURED TO CONFIRMED BACTERIAL GROWTH. DURING THIS TIME, GOOD LABORATORY PRACTICE, INCLUDING THOSE AT THE COMPLAINT'S SITE, WOULD CONFIRM THE PT ASSOCIATED WITH THE POSITIVE BOTTLE. ALTHOUGH THE LIKELIHOOD OF A CONFIRMED POSITIVE SAMPLE BEING ACTED ON BY A HEALTHCARE PROFESSIONAL IS REMOTE, IF GOOD LABORATORY PRACTICE IS NOT OBSERVED IT IS POSSIBLE WERE THIS MALFUNCTION TO RECUR. AN INVESTIGATION IS IN PROCESS AND THE RESULTS/CONCLUSIONS WILL BE DOCUMENTED IN A SUPPLEMENTAL MEDWATCH REPORT. | 0
4251098 | ><WE HAD A PT WHO WAS HAVING NEURO ISSUES, AND WHOSE RESPIRATIONS WERE SLOWING DOWN TO APPROX A RATE OF 5-8 INTERMITTENTLY. HOWEVER THE CENTRAL INFO CENTER (CIC) WAS READING AT ABOUT 30-40 ABOVE THE PT'S ACTUAL RATE. THE MONITOR, BECAUSE IT WAS GIVING A FALSE HIGH READING NEVER ALARMED FOR THIS CONDITION. THE MONITOR AT THE BEDSIDE AND THE CIC BOTH READING RESP RATE: 44 BUT THE PT WAS BREAKING 5/ MIN. RN VERIFIED THIS NUMBER MANUALLY BY USING A PORTABLE CO2 MONITOR, AND CONFIRMED RESP. TO = 5. PT WAS BEING MONITORED ON A B850 S/N# (B)(4), B850 SOFTWARE REV. 2.0.6.2. ALSO IN USE WAS A PT DATA MODULE (PDM), S/N# (B)(4). THE BEDSIDE DEVICE WAS REPORTING TO A CLINICAL INFO CENTER (CIC) S/N# (B)(4), SOFTWARE CA: 5.1.1, OS: 3.0.1, SV:1.1.1. | 0
4030406 | CUSTOMER REPORTED A FAILURE IN AUTO FOCUS WERE NO CELLS WERE SEEN ON THE IQ200.><IQ200 AUTOMATED URINE MICROSCOPY ANALYZER REPORTING FALSE NEGATIVE/FALSE LOW CELLS (RBCS AND WBCS) FOR (B)(6) PATIENT SAMPLES. NONE WERE REPORTED OUT TO LIS OR PATIENT. NO CHANGE IN PATIENT MANAGEMENT, NO LOSS OF PATIENT TO FOLLOW UP OR IMMATURE DISCHARGE WERE REPORTED DUE TO ERRONEOUS RESULT. NO INJURIES WERE REPORTED. PER CUSTOMER, THE CHEM/MICRO CORRELATION ALERT WAS TURNED OFF. THE CUSTOMER TECHNICAL SUPPORT (CTS) INSTRUCTED THE CUSTOMER TO CHANGE THE SAMPLE FILTER; THE IQ200 NOW FOCUSED FINE AND THE QC PASSED. | 0
4030401 | CUSTOMER REPORTED A FAILURE IN AUTO FOCUS WERE NO CELLS WERE SEEN ON THE IQ200.><IQ200 AUTOMATED URINE MICROSCOPY ANALYZER REPORTING FALSE NEGATIVE/FALSE LOW CELLS (RBCS AND WBCS) FOR (B)(6) PATIENT SAMPLES. NONE WERE REPORTED OUT TO LIS OR PATIENT. NO CHANGE IN PATIENT MANAGEMENT, NO LOSS OF PATIENT TO FOLLOW UP OR IMMATURE DISCHARGE WERE REPORTED DUE TO ERRONEOUS RESULT. NO INJURIES WERE REPORTED. PER CUSTOMER, THE CHEM/MICRO CORRELATION ALERT WAS TURNED OFF. THE CUSTOMER TECHNICAL SUPPORT (CTS) INSTRUCTED THE CUSTOMER TO CHANGE THE SAMPLE FILTER; THE IQ200 NOW FOCUSED FINE AND THE QC PASSED. | 0
4030374 | IQ200 AUTOMATED URINE MICROSCOPY ANALYZER REPORTING FALSE NEGATIVE/FALSE LOW CELLS (RBCS AND WBCS) FOR 11 PATIENT SAMPLES. NONE WERE REPORTED OUT TO LIS OR PATIENT. NO CHANGE IN PATIENT MANAGEMENT, NO LOSS OF PATIENT TO FOLLOW UP OR IMMATURE DISCHARGE WERE REPORTED DUE TO ERRONEOUS RESULT. NO INJURIES WERE REPORTED. PER CUSTOMER, THE CHEM/MICRO CORRELATION ALERT WAS TURNED OFF. THE CUSTOMER TECHNICAL SUPPORT (CTS) INSTRUCTED THE CUSTOMER TO CHANGE THE SAMPLE FILTER; THE IQ200 NOW FOCUSED FINE AND THE QC PASSED.><CUSTOMER REPORTED A FAILURE IN AUTO FOCUS WHERE NO CELLS WERE SEEN ON THE IQ200. | 0
4307596 | THE IQ200 URINE MICROSCOPY INSTRUMENT SHORT SAMPLE DETECTOR FAILED TO FLAG EMPTY NEGATIVE CONTROLS. NO PATIENT SAMPLES WERE AFFECTED DURING THIS EVALUATION.><THE FOLLOWING MDRS ARE RELATED TO THIS EVENT: 2023446-2014-00180, 00182, AND 00183. | 0
3807526 | ICHEM VELOCITY URINE CHEMISTRY AND IQ200 URINE MICROSCOPY REPORTING IS COMBINING THE WRONG MICROSCOPY AND CHEMISTRY RESULTS ON THE WORKLIST, POTENTIALLY MIXING UP PT SAMPLES AND REPORTING OUT WRONG PT SAMPLE RESULTS WITH THE WRONG PT, DUE TO AN IWARE FAILURE. NO PT MANAGEMENT CHANGED NOR WAS ANY TREATMENT INITIATED TO PTS BECAUSE THE PROBLEM WAS CAUGHT BEFOREHAND. THE PC WAS REPLACED, IQ PERFORMANCE VERIFICATION CHECKLIST WAS COMPLETED, AUTO-FOCUS AND CONTROLS RAN AND PASSED, AND CUSTOMER RAN SPECIMENS WITHOUT ANY ISSUES. SYSTEM WAS OPERATIONAL.><APPLICATION ENGINEER REPORTED THAT THE IWARE ON A VELOCITY UNIT IS COMBINING THE WRONG MICRO AND CHEM RESULTS ON THE WORKLIST. | 0
4102690 | CUSTOMER REPORTED THE INSTRUMENT WAS NOT ASPIRATING SAMPLES AND RESULT NOT MATCHING WITH MICROSCOPE.><IT WAS REPORTED THAT THE CHEM VELOCITY URINE CHEMISTRY ANALYZER SHOWED FALSE NEGATIVE FOR BLOOD AND LEUKOCYTE ESTERASE ON 3 PT SAMPLES, WHEREAS THE IQ200 URINE MICROSCOPY ANALYZER AND MANUAL MICROSCOPIC METHODS CONFIRMED THAT THOSE SAMPLES HAD THE PRESENCE OF RBCS AND WBCS. NO CHANGE TO PT MANAGEMENT REPORTED. FSE FOUND PROBLEM WITH STRIPS NOT BEING DOSED PROPERLY. FSE CHECKED ALL TUBING CONNECTIONS FROM SAMPLE PROBE TO DFS. DURING PRIMING OF SYSTEM, FSE NOTICED THAT PISTON ON THE SYRINGE STOPPED MOVING, REMOVED SYRINGE AND CHECK FOR DAMAGE SYRINGE PISTON CONES APART, REPLACED SYRINGE ASSEMBLY AND TIGHTEN PROPERLY. FSE PRIMED SYSTEM, RUN QC ALL IN SPECIFICATION. SYSTEM WAS OPERATIONAL. | 0
3799229 | ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED INFO FROM THE COMPLAINT THAT ONE PT HAD TO BE RE-BIOPSIED. THE RE-BIOPSY TOOK PLACE ON (B)(6) 2014. REF # MFR 8020030-2014-00010. | 0
3641218 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING INJURY TO A MEMBER OF THE LABORATORY STAFF WHEN USING A PELORIS TISSUE PROCESSOR. INFORMATION PROVIDED BY THE LABORATORY INDICATES THAT THE STAFF MEMBER WAS SPRAYED IN THE FACE WITH 90% IMS WHILE DRAINING THE REAGENT FROM THE RETORT USING THE REMOTE FILL/DRAIN FUNCTION. THE LABORATORY ALSO REPORTED THE STAFF MEMBER WAS NOT WEARING SAFETY SPECTACLES, BUT DID FOLLOW THE (B)(4) REGULATIONS DURING INTERACTION WITH THE INSTRUMENT. THE STAFF MEMBER ATTENDED THE EMERGENCY ROOM CONSULTED AN OPHTHALMOLOGIST AND EXPERIENCED LOST TIME AT WORK AS A CONSEQUENCE OF THIS EVENT. INVESTIGATION OF THIS COMPLAINT BY LEICA BIOSYSTEMS IS IN PROGRESS. | 0
4280164 | BOTTLE 5 (ETHANOL) WAS REMOVED FROM THE INSTRUMENT AT 08:27AM ON (B)(4) 2014, FOR SUFFICIENT TIME TO COMPLETE MANUAL REPLACEMENT OF THE REAGENT, AND THE PROPERTIES OF THE CORRESPONDING REAGENT STATION WERE REST. RESETTING THE REAGENT STATION SET THE CONCENTRATION TO THE DEFAULT VALUE CONFIGURED IN THE REAGENT TYPES DEFINITIONS (100% IN THIS INSTANCE); AND RESET THE NUMBER OF CASSETTES PROCESSED AND THE NUMBER OF CYCLES AND DAYS TO ZERO. HOWEVER, INFO PROVIDED BY THE LAB INDICATED THAT PRIOR TO COMMENCEMENT OF THE TWO (2) PROCESSING PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED, A USER HAD REPLACED THE REAGENT IN BOTTLE 5 WITH 7% ETHANOL AND SET THE CONCENTRATION TO 100% IN ERROR. BOTTLE 5 (ETHANOL) WAS THEN USED FOR THE FINAL DEHYDRATION STEP OF THE TWO (2) PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. THE MINIMUM FINAL REAGENT CONCENTRATION REQUIRED FOR ETHANOL IS 98%. THE CONSEQUENCES OF USING ETHANOL AT A CONCENTRATION LESS THAN THE MINIMUM REQUIRED FOR THE FINAL DEHYDRATION STEP IN A PROTOCOL IS RE-INTRODUCTION OF WATER INTO THE TISSUE WHICH CANNOT BE DISPLACED IN SUBSEQUENT PROCESSING STEPS; AND CONTAMINATION OF REAGENTS USED IN SUBSEQUENT PROCESSING STEPS ULTIMATELY RESULTING IN SUB-OPTIMAL PROCESSING. THE ROOT CAUSE FOR THE SUB-OPTIMAL TISSUE PROCESSING REPORTED WAS A USER ERROR, WHICH OCCURRED DURING THE REPLACEMENT OF ETHANOL IN BOTTLE 5 AND COMPLETED PRIOR TO COMMENCEMENT OF THE TWO PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED.><LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL TISSUE PROCESSING FROM TWO (2) PROCESSING RUNS. INFO PROVIDED TO A LEICA APPLICATION SPECIALIST/TRAINER BY THE COMPLAINANT INDICATED THAT A USE ERROR WHICH OCCURRED DURING MANUAL REPLACEMENT OF A REGENT PRIOR TO COMMENCEMENT OF THE AFFECTED PROCESSING RUNS, RESULTED IN A REAGENT BOTTLE CONTAINING 70% ETHANOL BEING SET AS 100% ETHANOL IN ERROR. THE COMPLAINANT ADVISED THAT SPECIMENS FROM THE TWO (2) PROCESSING RUN COMPLETED FOLLOWING THIS INCORRECT USER ACTION WERE NOTED TO BE UNDER-PROCESSED EMBEDDING. ON (B)(6) 2014, THE LEICA APPLICATION SPECIALIST/TRAINER CONTACTED THE LAB TO OBTAIN FURTHER SPECIFIC INFO REGARDING THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. THE LAB ADVISED THAT BOTTLE 5 HAD CONTAINED 70% ETHANOL WHEN THE ETHANOL CONCENTRATION CALCULATED BY THE INSTRUMENT SOFTWARE WAS. ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED INFO THAT ALL TISSUE SAMPLES EXHIBITING SUB-OPTIMAL TISSUE PROCESSING WERE DIAGNOSABLE. | 0
4152232 | (B)(6) BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL PROCESSING OF TISSUE IN APPROXIMATELY (B)(6) BLOCKS, WHICH WERE DERIVED FROM (B)(6) CASES. ON (B)(6) 2014, (B)(6) BIOSYSTEMS RECEIVED INFORMATION FROM THE LABORATORY INDICATING THAT AN UNSPECIFIED NUMBER OF PATIENT SAMPLES WERE UNDIAGNOSABLE. CLARIFICATION AS TO WHETHER RE-BIOPSY OF A PATIENT(S) IS REQUIRED AND/OR HAS BEEN PERFORMED HAS NOT BEEN RECEIVED BY (B)(6) BIOSYSTEMS AS OF (B)(6) 2014, DESPITE SEVERAL REQUESTS BEING MADE TO THE LABORATORY. ON (B)(6) 2014, A (B)(6) FIELD SUPPORT SPECIALIST (FSS) ATTENDED THE LABORATORY, IN ORDER TO INVESTIGATE THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. THE COMPLAINANT ADVISED THAT SHE WAS AWARE THAT A MEMBER OF THE LABORATORY STAFF HAD REPLACED A REAGENT INCORRECTLY; ALL THE REAGENTS ON THE INSTRUMENT AT THE TIME OF THE EVENT HAD BEEN REPLACED FOLLOWING IDENTIFICATION OF SUB-OPTIMAL TISSUE PROCESSING AND THE QUALITY OF TISSUE PROCESSING WAS FOUND TO BE SATISFACTORY FOLLOWING THIS ACTION.><EVALUATION OF THE INSTRUMENT LOGS FOUND THAT BOTTLE 11 (XYLENE) WAS REMOVED FROM THE INSTRUMENT FOR 7 SECONDS AT 17:31PM ON (B)(6) 2014 IN RESPONSE TO INFORMATION DISPLAYED IN ASSOCIATED WITH AN ERROR CODE. THE PROPERTIES OF THE CORRESPONDING REAGENT STATION WERE RESET, WHICH SET THE CONCENTRATION TO THE DEFAULT VALUE OF (B)(6) CONFIGURED IN THE REAGENT TYPES DEFINITIONS, AND RESET THE NUMBER OF CASSETTES AND THE NUMBER OF CYCLES AND DAYS TO ZERO. HOWEVER, AS THIS PERIOD IS INSUFFICIENT TO COMPLETE MANUAL REPLACEMENT OF THE REAGENT, THE ACTUAL CONCENTRATION OF THE XYLENE IN BOTTLE 11 REMAINED UNCHANGED AT (B)(6) BECAUSE THE REAGENT HAD NOT BEEN REPLACED. BOTTLE 11 (XYLENE) WAS USED FOR THE FINAL CLEARING STEP OF PROTOCOL FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS LIKELY TO HAVE BEEN DERIVED. THE MINIMUM FINAL REAGENT CONCENTRATION FOR XYLENE IS (B)(6). THE CONSEQUENCES OF USING XYLENE AT A CONCENTRATION LESS THAN THE MINIMUM REQUIRED FOR THE FINAL CLEARING STEP IN A PROTOCOL IS RE-INTRODUCTION OF WATER INTO THE TISSUE WHICH CANNOT BE DISPLACED IN SUBSEQUENT PROCESSING STEPS; AND CONTAMINATION OF THE REAGENTS USED IN THE SUBSEQUENT PROCESSING STEPS ULTIMATELY RESULTING IN SUB-OPTIMAL TISSUE PROCESSING. THE ROOT CAUSE OF THE SUB-OPTIMAL TISSUE PROCESSING REPORTED WAS A USE ERROR, WHICH OCCURRED PRIOR TO COMMENCEMENT OF THE AFFECTED PROTOCOL. SPECIFICALLY, THE USER FAILED TO COMPLETE THE MANUAL REAGENT REPLACEMENT PROCESS IN ACCORDANCE WITH THE MANUFACTURER INSTRUCTIONS DETAILED IN THE (B)(6) PELORIS/PELORIS II USER MANUAL. | 0
3983286 | LEICA BIOSYSTEMS RECEIVED A COMPLAINT THAT PELORIS TISSUE PROCESSOR WAS FAILING TO GENERATE SUFFICIENT PRESSURE TO PUMP LIQUID INTO THE RETORT. DURING ATTENDANCE AT THE LABORATORY TO ASSESS THE OPERATION AND FUNCTION OF THE INSTRUMENT, THE LEICA FIELD SERVICE ENGINEER (FSE) WAS ADVISED BY THE COMPLAINANT THAT THE QUALITY OF TISSUE PROCESSING FOR SAMPLES RUN IN RETORT B WAS SUB-OPTIMAL. INVESTIGATION OF THIS COMPLAINT BY LEICA BIOSYSTEMS IS IN PROGRESS. | 0
3867985 | ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL PROCESSING OF TISSUE IN APPROXIMATELY 110 CASSETTES FROM A RUN IN RETORT A USING PELORIS TISSUE PROCESSOR. THE LABORATORY MEASURED THE CONCENTRATION IN BOTTLE 9, WHICH WAS 84%, IN COMPARISON TO THE CONCENTRATION CALCULATED BY THE INSTRUMENT SOFTWARE OF 100%. ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED INFORMATION FROM THE LABORATORY THAT ALL TISSUE SAMPLES EXHIBITING SUB-OPTIMAL PROCESSING WERE DIAGNOSABLE.><BOTTLE 9 (ETHANOL) MOVED OUT OF CONTACT FROM THE SENSOR FOR APPROX 2 MIN AT 18:49PM ON (B)(6) 2014 DURING THE REAGENT DRAIN PROCESS. WHEN CONTACT WITH THE SENSOR WAS RE-ESTABLISHED, THE <INSERTED BOTTLE CONFIGURATION> DIALOG BOX WOULD HAVE BEEN DISPLAYED. THE INSTRUMENT LOGS SHOW THAT THE USER AFFIRMED IN THE INSTRUMENT SOFTWARE THAT THE ETHANOL CONC WAS TO BE SET TO THE DEFAULT VALUE OF 100% AT 19:04PM ON (B)(6) 2014. RESETTING A REAGENT STATION SETS THE CONCENTRATION TO THE DEFAULT VALUE CONFIGURED IN THE REAGENT TYPES DEFINITION UNLESS AN ALTERNATIVE VALUE IS ENTERED INTO THE INSTRUMENT SOFTWARE BY THE USER; AND RESETS THE NUMBER OF CASSETTES PROCESSED AND THE NUMBER OF CYCLES AND DAYS TO 0. THE ACTUAL ETHANOL CONC OF THE REAGENT IN BOTTLE 9 REMAINED AS 83.7% BECAUSE IT HAD NOT BEEN PHYSICALLY REPLACED. THE INSTRUMENT SOFTWARE USES REAGENT CONC TO SELECT REAGENT STATIONS WHEN EXECUTING A PROTOCOL. BOTTLE 9 WAS USED FOR THE FINAL DEHYDRATION STEP OF THE PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. THE REQUIRED MINIMUM FINAL REAGENT CONC FOR ETHANOL IS 98%. THE CONSEQUENCES OF USING ETHANOL AT LESS THAN THE MINIMUM FINAL CONC REQUIRED ARE REINTRODUCTION OF WATER INTO THE TISSUE WHICH CANNOT BE DISPLACED IN SUBSEQUENT PROCESSING STEPS AND CONTAMINATION OF REAGENTS USED FOR SUBSEQUENT PROCESSING STEPS. THE ROOT CAUSE FOR THE SUB-OPTIMAL TISSUE PROCESSING REPORTED WAS A FAILURE BY THE USER TO COMPLETED THE MANUAL REAGENT REPLACEMENT PROCESS IN ACCORDANCE WITH THE MANUFACTURER INSTRUCTIONS IN THE LEICA PELORIS/PELORISII USER MANUAL. | 0
3897634 | ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED A COMPLAINT REGARDING SUB-OPTIMAL PROCESSING OF TISSUE IN APPROXIMATELY 420 CASSETTES FROM TWO (2) PROCESSING RUNS USING PELORIS II TISSUE PROCESSOR. PRELIMINARY INFO PROVIDED BY THE LABORATORY INDICATES THAT A USER HAD REPLACED THE REAGENT IN BOTTLE 1, WHICH IS DESIGNATED FOR FORMALIN WITH 100% ETHANOL IN ERROR AND REQUESTED ADVICE/ASSISTANCE TO PREVENT THIS EVENT FROM RE-OCCURRING. THE LABORATORY REPORTED THAT ALL THE TISSUE SAMPLES INVOLVED WERE DIAGNOSABLE. TELEPHONE SUPPORT WAS PROVIDED BY A LEICA FIELD SUPPORT SPECIALIST (FSS) ON (B)(6) 2014, IN ORDER TO PROVIDE ASSISTANCE WITH THE INVESTIGATION AND RESOLUTION OF THE CIRCUMSTANCES INVOLVED IN THIS COMPLAINT. THE FSS COMMUNICATED THE IMPORTANCE OF READING LABELS ON THE REAGENT BOTTLES TO MITIGATE RECURRENCE OF THE CIRCUMSTANCES, WHICH RESULTED IN THE SUB-OPTIMAL TISSUE PROCESSING REPORTED IN THIS COMPLAINT.><INFO PROVIDED BY THE LABORATORY THAT A USER HAD PREVIOUSLY REPLACED THE REAGENT IN REAGENT BOTTLE 1 WITH 100% ETHANOL. EVALUATION OF THE INSTRUMENT LOGS PROVIDED AND INFO PROVIDED BY THE LABORATORY INDICATES TWO (2) PROCESSING PROTOCOLS WERE AFFECTED. BOTTLE 1 WAS USED FOR THE FIRST TIME IN THE FIXATION STEP OF THE (B)(4) PROTOCOL STARTED IN RETORT A AT 23:39PM ON (B)(6) 2014 AND BOTTLE 2 WAS USED FOR THE FIRST TIME IN THE FIXATION STEP OF THE (B)(4) PROTOCOL STARTED IN RETORT B AT 17:10PM ON (B)(6) 2014. THIS FINDING SUGGEST A USE ERROR OCCURRED DURING THE REPLACEMENT OF REAGENT IN REAGENT BOTTLES 1 AND 2 WHICH WAS PERFORMED PRIOR TO COMMENCEMENT OF THE PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. INVESTIGATION OF THIS COMPLAINT FOUND THAT THE INSTRUMENT OPERATED WITHIN SPECIFICATION DURING EXECUTION OF THE PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. THE ROOT CAUSE OF THE SUB-OPTIMAL TISSUE PROCESSING REPORTED WAS A USE ERROR, WHICH OCCURRED PRIOR TO COMMENCEMENT OF THE PROTOCOLS FROM WHICH SUB-OPTIMAL TISSUE PROCESSING WAS REPORTED. SPECIFICALLY, THE LABORATORY ADVISED THAT A USER HAD A REPLACED THE REAGENT, WHICH WAS DESIGNATED FROM <FORMALIN>, WITH 100% ETHANOL IN ERROR. THIS INFO WAS PROVIDED TO LEICA BIOSYSTEMS BY THE COMPLAINT. | 0
4073108 | CALIBRATION IF AFFECTED BY SOFTWARE DEFECT.><DAKO HAS FOUND A SOFTWARE DEFECT IN THE CALIBRATION ROUTINE OF DAKOLINK SOFTWARE VERSION 4.0.3, WHICH IS USED IN THE DAKO AUTOSTAINER 48. OUR INVESTIGATION IDENTIFIED THAT THIS MAY HAPPEN ONLY IN CIRCUMSTANCES WHERE THE FOLLOWING EVENTS OCCUR IN THE ORDER LISTED BELOW: 1: DAKOLINK SOFTWARE ON THE AUTOSTAINER 48 HAS BEEN UPGRADED FROM VERSION 3.1 TO 4.0.3 AND 2) EITHER A PREVENTIVE MAINTENANCE OR WASTE PUMP REPLACEMENT OF THE AUTOSTAINER 48 WAS PERFORMED DURING WHICH THE FIELD SERVICE ENGINEER INADVERTENTLY MOVED THE BASE PLATE, AND 3) THE BASEPLATE IS NOT COMPLETELY ALIGNED IN THESE VERY SPECIFIC CIRCUMSTANCES, THE SOFTWARE DEFECT MAY CAUSE THE AUTOSTAINER CONTROL SOFTWARE TO USE OLD CALIBRATION SETTINGS FROM THE PREVIOUS SOFTWARE VERSION 3.1, WHICH MAY LEAD TO: PREMATURE REJECTION OF REAGENT BOTTLES WHICH CAUSES THE INSTRUMENT TO SEND AN ERROR MESSAGE TO THE OPERATOR, AND THEREFORE, HAS NOT PATIENT IMPACT OR INADEQUATE ASPIRATION OF REAGENT, WHICH MAY LEAD TO COMPROMISED STAINING WHICH IF UNDETECTED BY THE OPERATOR AND/OR PATHOLOGIST COULD LEAD TO A FALSE POSITIVE OR FALSE NEGATIVE. IT SHOULD BE NOTED THAT ALL SLIDES GO THROUGH PATHOLOGIST REVIEW PRIOR TO PATIENT RESULT DETERMINATION. HOWEVER, EVEN IF THE ABOVE THREE CONDITIONS ARE MET, THE INSTRUMENT MAY PERFORM AS INTENDED. THE INTERNAL GEOMETRY OF THE DIFFERENT BOTTLES, GEOMETRY OF THE PLATE AND LOCATION OF REAGENT RACK ON THE AUTOSTAINER ARE ALSO FACTORS IN DETERMINING WHETHER A FAILURE WILL OCCUR. NEW INSTRUMENTS ARE NO AFFECTED BY THIS SOFTWARE DEFECT AS THEY ARE MANUFACTURED WITH DAKOLINK SOFTWARE VERSION 4.0.3 AND THERE ARE NO OLD CALIBRATION SETTINGS TO WHICH THE INSTRUMENT CAN REVERT BACK. THE TOTAL AFFECTED INSTRUMENT POPULATION IDENTIFIED ABOVE (SECTION 2) WERE IDENTIFIED BECAUSE THEIR SOFTWARE HAS BEEN UPGRADED TO VERSION 4.0.3 AND EITHER THE PM OR WASTE PUMP REPLACEMENT HAS BEEN PERFORMED. NOT ALL OF THESE INSTRUMENTS WILL MEET THE THIRD REQUIREMENTS FOR THIS POTENTIAL FAILURE BECAUSE FIELD SERVICE VISITS ARE REQUIRED TO DETERMINE WHETHER INSTRUMENT CONTAINS THE BASEPLATE MISALIGNMENT. DAKO ESTIMATES THAT APPROXIMATELY 5% OF INSTRUMENTS IDENTIFIED IN SECTION 2 ABOVE HAVE THE BASEPLATE MISALIGNMENT. THIS PERCENTAGE IS BASED ON THE FACT THAT DAKO HAS INSPECTED APPROXIMATELY ONE HUNDRED (100) IDENTIFIED INSTRUMENTS, AND ONLY FIVE (5) INSTRUMENTS HAVE BEEN IDENTIFIED AS HAVING THE BASEPLATE MISALIGNMENT. | 0
4002745 | IT WAS REPORTED THAT THE IQ200 URINE MICROSCOPY DEVICE DID NOT SHOW CELLS ON CLOUDY SPECIMENS. NO CHANGE IN PATIENT MANAGEMENT WAS REPORTED, HOWEVER, THIS MAY LEAD TO FALSE LOW/FALSE NEGATIVE CELLS ON CLOUDY SAMPLES. FSE REPLACED PROBE AND LAMINA BYPASS VALVE (LBV), RAN REF CALIBRATION AND CONTROLS. SYSTEM WAS OPERATIONAL.><CUSTOMER REPORTED NO CELL" OBSERVED ON SLIGHTLY CLOUDY SAMPLES." | 0
4255795 | FURTHER EVALUATION OF THE ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT AND PHOTOS PROVIDED OF THE SHUTTER COVER, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF THE INSTRUMENT SERVICE HISTORY AND A REVIEW OF PRODUCT LABELING. PHOTOS PROVIDED FOR REVIEW SHOW THE COVER WHERE ONE OF THE PROTECTIVE EDGE MOLDINGS HAD DISLODGED WHEN THE COVER FELL AND HIT THE FLOOR, EXPOSING THE SHARP CORNER OF THE SIDE PIECE; THE FIELD SERVICE ENGINEER (FSE) WAS NOT CERTAIN WHAT PART OF THE COVER CAUSED THE INJURY. THE CLOSE UP PHOTO SHOWS THE TOP PIECE IS BENT AND IS NOT ALIGNED WITH THE SIDE PIECE. A SPOT OF BLOOD IS VISIBLE JUST BELOW THE SHARP CORNER, SUGGESTING THE ENGINEER WAS INJURED AT THAT SPOT. THIS AREA IS NORMALLY COVERED BY RUBBER EDGE MOLDING. IT WAS CONCLUDED THAT THE INJURY SUSTAINED BY THE ENGINEER WAS THE RESULT OF USE ERRORS. REVIEW OF THE INSTRUMENT SERVICE HISTORY SHOWED THAT THE RUBBER EDGE MOLDING TO 6 SHARP CORNERS WAS INSTALLED ON (B)(4) 2013 IN RESPONSE TO A PREVIOUS ISSUE AT THIS SITE. THE PREVIOUS INVESTIGATION IDENTIFIED A DEFICIENCY IN THE RELIABILITY OF THE FRAME IF IT BECAME BENT THROUGH HANDLING, THEREBY CAUSING A SHARP EDGE. THE EDGE IS NOT SHARP AS MANUFACTURED. THE CORNERS OF THE COVER WERE MODIFIED TO BE ROUNDED STARTING IN (B)(4) 2012. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. INSTRUCTION FOR INSTALLATION BY SERVICE PERSONNEL DIRECTS THE EMPLOYEE TO WEAR PROTECTIVE GLOVES AND PROTECTIVE EYE WEAR AND COMPLETE INSTALLATION WITH 2 OR MORE SERVICE PERSONNEL. THE INJURY WAS SUSTAINED WHILE WORKING INDEPENDENTLY. TRACKING AND TRENDING WAS COMPLETED, NO OTHER COMPLAINTS WERE FOUND NOTING INJURIES SUSTAINED FROM THE STORAGE PANEL SINCE THE RUBBER EDGE MOLDING WAS INSTALLED TO EXISTING PANELS AND THE ROUNDED COVERS WERE MANUFACTURED. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY AND NO MALFUNCTION OF THE APS TUBE STORAGE, LN 07L07 WAS IDENTIFIED.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><DURING SERVICE OF THE APS TUBE STORAGE AN ABBOTT FIELD SERVICE ENGINEER WAS INJURED WHEN THE STORAGE SHUTTER COVER FELL DOWN. THE FSE WAS ATTEMPTING TO PLACE THE SHUTTER COVER BACK ON THE INSTRUMENT WHEN THE COVER SLIPPED AND THE INJURY OCCURED. THE INJURY REQUIRED 9 STITCHES ON HIS ARM, A TETANUS SHOT AND OTHER MEDICATION (ANTIBIOTICS AND PAIN KILLERS). IT WAS INDICATED THAT THERE WAS NO LONG TERM ADVERSE IMPACT AFTER RECEIVING TREATMENT. | 0
3813481 | IT WAS DETERMINED THE RACK TRAY WAS MORE THAN TEN YEARS OLD AND A DESIGN CHANGE OCCURRED IN 2005 TO ROUND THE EDGES OF THE TRAY. THIS WAS COMMUNICATED VIA A SERVICE BULLETIN. THIS WAS THE ONLY COMPLAINT RECEIVED SINCE THE DESIGN CHANGE.><THE CUSTOMER WAS UNLOADING RACKS FROM THE OUTPUT BUFFER MODULE OF THEIR MPA. WHILE REMOVING THE 25 RACK TRAY FROM THE MODULE, HE WAS HOLDING ONTO THE END WITH THE TWO LOCATING PRONGS. HE PUT A TEAR" INTO THE END OF HIS FINGER. HE WENT TO THE A/E TO HAVE IT CLEANED AND GLUED AS IT WAS BLEEDING FREELY. HE THEN WENT TO THE HOSPITAL'S OCCUPATIONAL HEALTH DEPARTMENT TO HAVE A BLOOD SAMPLE TAKEN. THE RESULTS CAME BACK ALL CLEAR. THE FINGER WAS HEALING WELL | 0
3483492 | CALLER REPORTS SEEING SMOKE COMING FROM THE MULTIPLATE ANALYZER. NO ADVERSE EVENT REPORTED. REQUESTED RETURN OF SUSPECT SYSTEM AND REPLACEMENT WAS SENT.><THE EVENT OCCURRED IN (B)(6). | 0
3949350 | POSSIBLE FALSE NEGATIVE. LAB REPORTED THREE (3) CASES IN WHICH ABNORMAL CELLS FOUND OUTSIDE OF 22 FIELDS OF VIEW (FOV). NO ABNORMAL CELLS WERE PRESENT WITHIN THE 22 FOV TO PROMPT FULL AUTOSCAN OF THE SLIDE. THE ABNORMAL CELLS WERE FOUND OUTSIDE THE FOV DURING QC. SITE WILL BE MONITORED TO DETERMINE IF EXPECTED OCCURRENCES ARE EXCEEDED. HOLOGIC WILL SCHEDULE A FULL IMAGER PREP REVIEW WITH THE LABORATORY. THIS INCIDENT IS CLASSIFIED AS A RARE EVENT, WHICH IS WHERE THE CELLS IN QUESTION WOULD MATCH OUR EXPERIENCE IN THE CLINICAL TRIAL AS EVIDENCED IN TABLE 8 OF THE THINPREP IMAGING SYSTEM OPERATING SUMMARY AND CLINICAL INFORMATION. | 0
4336984 | PT ON HEMODYNAMIC SUPPORT WITH ABIOMED IMPELLA DEVICE. DEVICE ALARMED LOW FLOW" AND CEASED FUNCTIONING. PT DECOMPENSATED FOLLOWED BY CARDIAC ARREST. PT EXPIRED. PT ON HEMODYNAMIC SUPPORT WITH ABIOMED IMPELLA DEVICE." | 0
3897760 | (B)(4).><THE CUSTOMER STATED THAT A (B)(6) RESULT OF (B)(6) WAS GENERATED. THE SAMPLE WAS REPEATED MULTIPLE TIMES WITH RESULTS IN THE RANGE OF (B)(6). THE PATIENT HAD BEEN RUN ON ANOTHER ANALYZER WITH A (B)(6) RESULT. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT.><A FALSE NEGATIVE ANTI HBC RESULT WAS GENERATED FOR ONE PATIENT ON THE ARCHITECT I2000SR ANALYZER. FIELD SERVICE PERFORMED TROUBLESHOOTING AND DETERMINED THE PROBABLE CAUSE WAS THE WASH BUFFER DISPENSE AT THE WASH ZONES WAS INADEQUATE. SERVICE REPLACED THE VALVE MANIFOLD KITS TO RESOLVE THE ISSUE. NO FURTHER DISCREPANT RESULTS WERE NOTED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A MALFUNCTION / DEFICIENCY. | 0
3655374 | THE CUSTOMER REPORTS A FALSE POSITIVE ARCHITECT STAT TROPONIN-I ASSAY RESULT FOR ONE PATIENT GENERATED ON AN ARCHITECT I1000SR ANALYZER. AN INITIAL RESULT OF 0.243 NG/ML WAS GENERATED ON A PATIENT TRANSPORTED TO THE HOSPITAL IN AN AMBULANCE (DURING TRANSPORT, THE PATIENT WAS ADMINISTERED BRILIQUE AND ASPIRIN). THE CUSTOMER USES A REFERENCE RANGE OF 0.00 TO 0.028 NG/ML. THE PATIENT WAS ADMITTED TO THE ICU WHERE ANOTHER SAMPLE TAKEN APPROXIMATELY FOUR AND ONE-HALF HOURS LATER GENERATED A RESULT OF 0.004 NG/ ML. A CARDIAC CATHETERIZATION WAS PERFORMED BASED ON THE INITIAL ARCHITECT VALUE THAT DID NOT FIND ANY SIGNS OF CARDIAC INJURY. ON (B)(6) 2014 A PLASMA SAMPLE TAKEN FROM THIS PATIENT WAS TESTED AND GENERATED RESULTS OF 0.003 AND 0.005 NG/ML. THE PATIENT WAS DISCHARGED ON (B)(6) 2014 AND IS DOING FINE. THERE IS NO FURTHER IMPACT TO PATIENT MANAGEMENT REPORTED.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><NO RETURNS WERE MADE AVAILABLE FROM THE CUSTOMER SITE FOR THIS EVALUATION. AN ABBOTT FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE AND REPLACED THE SAMPLE PROBE (LIST NUMBER 08C94-42) AND SYRINGE (LIST NUMBER 7-77650-02) DUE TO NORMAL WEAR. SUBSEQUENT INSTRUMENT OPERATIONS AND TEST RESULTS WERE ACCEPTABLE. ACCURACY TESTING WAS PERFORMED WITH AN IN-HOUSE RETAINED LOT OF THE ARCHITECT STAT TROPONIN-I ASSAY, LOT 22947UN13. SIX REPLICATES OF A TROPONIN- I PANEL WERE TESTED ON ONE ARCHITECT ISYSTEMS PLATFORM. ALL TEST VALUES FELL WITHIN THE EXPECTED RANGE OF 0.23 TO 0.43 NG/ML, WHICH DEMONSTRATES THAT THIS ASSAY LOT CAN ACCURATELY DETECT A KNOWN CONCENTRATION OF TROPONIN-I. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT STAT TROPONIN-I ASSAY PACKAGE INSERT AND THE ARCHITECT OPERATIONS MANUAL BOTH CONTAIN INFORMATION TO ADDRESS THE CURRENT CUSTOMER ISSUE. THE CUSTOMER CONTACTED ABBOTT TO REPORT THAT THE CK ASSAY RESULTS FOR THIS PATIENT WERE ELEVATED; HOWEVER, THE DECISION TO PERFORM THE CARDIAC CATHETERIZATION WAS MADE PRIOR TO OBTAINING A NEGATIVE CK-MB ASSAY RESULT. THERE IS NO INFORMATION TO REASONABLY SUGGEST A MALFUNCTION OCCURRED. THE ISSUE WAS ISOLATED TO A SINGLE DISCREET SAMPLE AND THERE IS NO INDICATION OF ANY TROUBLESHOOTING PERFORMED BY THE CUSTOMER WITH THE SUSPECT SAMPLE. NO ADDITIONAL ISSUES WERE IDENTIFIED. | 0
4027249 | (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><REVIEW OF THE COMPLAINT TEXT INDICATES THE CUSTOMER OBSERVED FALSE POSITIVE TOTAL B-HCG RESULTS FOR ONE PATIENT RESULT ON THE ARCHITECT I2000SR, SERIAL NUMBER (B)(4). THE FIELD SERVICE REPRESENTATIVE (FSR) OBSERVED A SERUM AND/OR REAGENT CRYSTAL BUILD UP NEAR THE REACTION VESSEL (RV) AREA. THE FSR DETERMINED THE SINGLE ISOLATED CASE OF FALSE POSITIVE B-HCG RESULTS WAS DUE TO CRYSTALS WHICH CONTAMINATED THE RVS DURING TESTING. CLEANING AND REMOVING THE CRYSTAL BUILD UP WAS DETERMINED TO HAVE RESOLVED THE B-HCG IMPRECISION ISSUE. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND THE ARCHITECT TOTAL B-HCG REAGENT PACKAGE INSERT WERE REVIEWED AND WERE FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A MALFUNCTION / DEFICIENCY.><THE CUSTOMER STATED THAT A FALSE POSITIVE ARCHITECT B-HCG RESULT OF 427.50 MIU/ML WAS GENERATED. THE SAMPLE WAS REPEATED 2 TIMES AND RESULTS OFF < 1.2 MIU/ML WERE GENERATED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. | 0
3856188 | THE CUSTOMER STATED THAT A FALSE POSITIVE ARCHITECT BHCG RESULT OF 417.86 MIU/ML WAS GENERATED FOR A PATIENT SAMPLE ((B)(6)) THAT RETESTED NEGATIVE AT <1.20 MIU/ML IN DUPLICATE. THE SAMPLE WAS TESTED AT A REFERENCE LABORATORY AND THE RESULT WAS NEGATIVE AT <2.00 MIU/ML. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><CUSTOMER RETURNS WERE NOT AVAILABLE FOR INVESTIGATION. AN ABBOTT FIELD SERVICE REPRESENTATIVE INSPECTED THE ARCHITECT I2000SR ANALYZER AND DID NOT IDENTIFY ANY ISSUES AS THE LIKLEY CAUSE OF THE DISCREPANT RESULT. REVIEW OF QUALITY METRICS DID NOT IDENTIFY ANY SIMILAR ISSUES. REVIEW OF HISTORICAL COMPLAINTS ASSOCIATED WITH THE ARCHITECT B-HCG ASSAY AND ARCHITECT B-HCG REAGENT LOT 33907UI00 WAS PERFORMED. NO ADVERSE TRENDS, NON-STATISTICAL TRENDS OR ATYPICAL COMPLAINT ACTIVITY WAS IDENTIFIED. A REVIEW OF LABELING CONCLUDED THAT THE ISSUE IS SUFFICIENTLY ADDRESSED. BASED ON THE DATA AVAILABLE AND THE RESULTS OF THIS INVESTIGATION, NO MALFUNCTION OR DEFICIENCY OF THE ARCHITECT I2000SR INSTRUMENT COULD BE IDENTIFIED.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). | 0
4329693 | THE USER OF A XE-5000 AUTOMATED HEMATOLOGY ANALYZERS (SERIAL NUMBER (B)(4)) REPORTED THAT BASED ON ERRONEOUS LOW HGB RESULT OF 8.4G/DL, PATIENT WITH SAMPLE ID (SID) (B)(6) WAS UNNECESSARILY TRANSFUSED, WITH TWO UNITS OF PACKED RBCS. SID (B)(6) WAS ANALYZED ON (B)(6) AT 07:43 USING THE AUTOMATIC MODE; THEREFORE, THE SAMPLE'S BARCODE WAS READ BY THE ANALYZER'S INTERNAL BARCODE READER. THE OPERATOR RERAN THE SAME SAMPLE AT 07:48 USING THE MANUAL MODE; THE SID WAS ENTERED BY THE OPERATOR. ACCORDING TO THE HOST'S LOG THE OPERATOR INCORRECTLY MODIFIED SID (B)(6); BECAUSE THE OPERATOR MODIFIED THE SID THE WRONG RESULTS TRANSMITTED TO SID (B)(6); THUS, THE WRONG RESULTS REPORTED OUT OF THE LABORATORY. THE ONLY RESULTS SUBMITTED FOR SID (B)(6) WAS RUN ON (B)(6) AND A HGB RESULT OF 12G/DL WAS OBTAINED; IT IS UNKNOWN IF THERE ARE ANY OTHER RESULTS FOR THIS PATIENT. NO KNOWN ADVERSE EFFECT TO THE PATIENT WAS REPORTED. BOTH PATIENTS HAD BEEN ADMITTED TO THE HOSPITAL PRIOR TO THE UNNECESSARY TRANSFUSION; PATIENTS' DIAGNOSIS AND/ OR CONDITIONS ARE UNKNOWN.><THIS EVENT IS BEING REPORTED ON THE BASIS THAT INCORRECT RESULTS LED TO A PATIENT RECEIVING AN UNNECESSARY TRANSFUSION, CARRYING WITH IT THE RISKS INVOLVED WITH RECEIVING BLOOD PRODUCTS: TRANSFUSION REACTION, DEVELOPMENT OF ANTIBODIES, ETC. THE HOST LOG FOR THE EVENTS ON (B)(6) 2014 FROM THE HOST COMPUTER SHOWS SID (B)(6) WAS NOT RECEIVED FOR TESTING UNTIL 07:54; THEREFORE, THERE WERE NO RESULTS FOR SID (B)(6) AT THE TIMES SID (B)(6) WAS ANALYZED AND REPEATED. THE USER CONFIRMED THE EVENT WAS DUE TO AN OPERATOR ERROR. THE USER REPORTED THAT THEY WORKED WITH THEIR QUALITY ASSURANCE DEPARTMENT TO IMPLEMENT THE FOLLOWING CHANGES TO HELP ELIMINATE THIS FROM HAPPENING IN THE FUTURE: THEY WILL RESTRICT THE ABILITY OF THE OPERATOR TO EDIT SAMPLE IDS. EACH OPERATOR WILL BE ASSIGNED THEIR OWN LOG ON. EAC OPERATOR WILL COMPLETE THEIR OWN WORK; THERE WAS A HANDOFF SITUATION WITH HIS SCENARIO. | 0
3802112 | INITIAL REPORTER LOCAL TELEPHONE NUMBER: (B)(6). (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER OBSERVED SMOKE COMING FROM THE CELL-DYN RUBY ANALYZER. THE CUSTOMER ALSO INDICATED THAT THE ELECTRICAL SYSTEM WAS NOISY. THERE WAS NO INJURY TO THE INSTRUMENT OPERATOR OR IMPACT TO PATIENT MANAGEMENT REPORTED.><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, ABBOTT FIELD SERVICE EVALUATION, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. ABBOTT FIELD SERVICE REPLACED THE POWER SUPPLY AND RESOLVED THE ISSUE. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE FOR POWER SPECIFICATIONS AND ELECTRICAL HAZARDS. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY OF THE CELL-DYN RUBY ANALYZER, LIST NUMBER 08H67 WAS IDENTIFIED. | 0
4012448 | THE CUSTOMER STATED THAT THEY SMELLED AN ELECTRICAL BURNING SMELL ON THE CELL-DYN SAPPHIRE AND THEN NOTICED SMOKE COMING FROM THE BACK OF THE INSTRUMENT. THERE WERE NO VISIBLE FLAMES OR SPARKS. THE CUSTOMER COULD HEAR THE ANALYZER FANS RUNNING. THE CUSTOMER TURNED OFF THE ANALYZER AND UNPLUGGED IT FROM THE WALL OUTLET. THERE WAS NO PERSONAL INJURY, IMPACT TO PATIENT MANAGEMENT, OR FACILITY DAMAGE REPORTED.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. | 0
4218452 | (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER REPORTED THAT THE CELLDYN RUBY FANS WERE NOT FUNCTIONING. THE FIELD SERVICE ENGINEER (FSE) FOUND THE PRESENCE OF SOOT ON BOTH OF THE FANS. THE FANS WERE REPLACED AND INSTRUMENT PERFORMANCE WAS VERIFIED. INSPECTION OF THE RETURNED PARTS (FANS) FOUND DISCOLORATION OF THE MAGNETIC COILS OF THE FAN AND THE PRESENCE OF SPIDER WEBS. IN ADDITION, A BURN MARK WAS OBSERVED ON THE PCB, AND THE CENTER COLUMN OF EACH FAN WAS FOUND TO BE OFF-CENTER, INDICATING OVERHEATING; HOWEVER THE CAUSE OF THE OVERHEATING COULD NOT BE DETERMINED. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. TRENDING DATA DOES NOT INDICATE OTHER FANS BEING OFF CENTER. THE CELL-DYN RUBY OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A PRODUCT DEFICIENCY.><THE CUSTOMER INDICATED THAT THE CELL-DYN RUBY ANALYZER FAN WAS NOT WORKING. THE FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE AND OBSERVED THAT THE TOP COVER OF THE CELL-DYN RUBY ANALYZER WAS NOT THE NORMAL TEMPERATURE. THE FSE OPENED THE ANALYZER AND UPON INSPECTION FOUND THAT BOTH FANS WERE COVERED IN SOOT. THE FSE REPLACED THE FANS. THERE WAS NO REPORT OF ANY OTHER DAMAGE OR INJURY. | 0
3977780 | UPON REVIEW, DISCOVERED MANUFACTURER NAME AND MANUFACTURER SITE, WAS INCORRECT IN THE FOLLOWUP SUBMISSION. THIS FOLLOWUP SUBMISSION TO CORRECT SITE OF ABBOTT DIAGNOSTICS DIVISION, (B)(6). THE EVALUATION IS IN PROCESS. A FOLLOW UP SUBMISSION WILL BE PROVIDED WHEN THE EVALUATION IS COMPLETE. EVALUATION IS IN PROCESS.><THE ABBOTT FIELD SERVICE ENGINEER INSTALLED A MOTHERBOARD AND CIRCUIT CARDS ON A CELL-DYN 3200 ANALYZER. AFTER POWERING UP THE INSTRUMENT, THE ABBOTT FIELD SERVICE ENGINEER HEARD A POP AND SAW SMOKE COMING FROM THE CELL-DYN 3200 ANALYZER. NO INJURY WAS REPORTED.><INITIALLY, THE FIELD SERVICE REPRESENTATIVE WAS REQUESTED BY THE CUSTOMER BECAUSE THEIR CELL-DYN 3200 INSTRUMENT WOULD NOT BOOT UP AND WOULD NOT GO TO PRIME. DURING TROUBLESHOOTING THE FIELD SERVICE REPRESENTATIVE FOUND THE ORIGINAL MOTHERBOARD WAS WORN OUT FROM NORMAL USE AND REPLACED IT WITH A NEW MOTHERBOARD. DURING POWER UP, THERE WAS A POP" AND SMOKE WHICH CAME FROM THE MOTHERBOARD. THE BURNT MOTHERBOARD WAS RETURNED FOR EVALUATION. EXAMINATION OF THE RETURNED MOTHERBOARD FOUND EVIDENCE THAT DURING INSTALLATION OF THE NEW MOTHERBOARD | 0
3733590 | THE ACCOUNT PROCESSED PATIENT SPECIMEN ID (B)(6) WITH NO RESULTS (WBC, RBC, HGB, PLT OF 0.00) IN CLOSED MODE USING THE CELL-DYN SAPPHIRE. THE OPERATOR TOOK THE SPECIMEN FROM THE SAMPLE RACK AND PROCESSED IN OPEN MODE WITH FALSELY DEPRESSED WBC OF 0.012 10E9/L WHICH WAS REPORTED OUTSIDE OF THE LABORATORY. THE ACCOUNT REVIEWED THE BLOOD SMEAR AND PERFORMED A MANUAL WBC OF 8.0 10E9/L, WHICH WAS REPORTED OUTSIDE OF THE LAB IN A CORRECTED REPORT. UPON TROUBLESHOOTING, THE ACCOUNT DISCOVERED THE SAMPLE RACK HAD MOVED THROUGH THE CELL-DYN SAPPHIRE. ANALYZER 4 TIMES WITHOUT STOPPING BECAUSE THE RACK FLAG FAILED TO RAISE AND STOP THE RACK AFTER PROCESSING ONCE THROUGH CLOSED MODE. ADDITIONALLY NOTED, THE PRIOR SAMPLE POSITION, SPECIMEN ID 14P016817, WAS A CLOTTED SAMPLE RUN JUST PRIOR TO THE PATIENT SAMPLE IN QUESTION. THE ACCOUNT ACKNOWLEDGES THAT THE OPERATOR SHOULD NOT HAVE RELEASED THE WBC OF 0.012 10E9/L AND IS BEING RETRAINED. THE PATIENT HAS DIFFUSE LARGE B CELL LYMPHOMA. THE PATIENT'S CHEMOTHERAPY WAS DELAYED FOR 1 WEEK DUE TO THE FALSELY DEPRESSED WBC RESULT OF 0.012 10E9/L BEING REPORTED OUTSIDE OF THE LAB. NO OTHER ISSUE FOR THE PATIENT DUE TO DELAY IN TREATMENT WAS REPORTED. DATA LOGS PROVIDED: CLOSED MODE, RUN 1: WBC = 8.96 10E9/L, RBC = 4.13 10E12/L, HGB = 122 G/L, PLT=232 10E9/L; CLOSED MODE, RUN 2: WBC = 9.11 10E9/L, RBC = 4.15 10E12/L, HGB = 121 G/L, PLT=239 10E9/L; CLOSED MODE, RUN 3: WBC = 9.18 10E9/L, RBC = 4.09 10E12/L, HGB = 121 G/L, PLT=238 10E9/L; CLOSED MODE, RUN 4: WBC = 0.00 10E9/L, RBC = 0.00 10E12/L, HGB = 0.00 G/L, PLT=0.00 10E9/L; OPEN MODE, RUN 5: WBC = 0.012 10E9/L, RBC = 4.24 10E12/L, HGB = 121 G/L, PLT=250 10E9/L. MANUAL WBC = 8.0 10E9/L, NO PLATELET CLUMPS.><(B)(4). FIELD SERVICE REPRESENTATIVE (FSR) FOUND THAT SEVERAL RACKS HAD SUB OPTIMAL RACK SWITCHES. IN ADDITION, THE ELEVATED AREA WHICH LIFTS THE MAGNETIC INDICATORS WAS PUSHED BACK, ALLOWING THESE SUB OPTIMAL RACKS TO AVOID THE RAISED AREA AND THUS CYCLE AROUND THE LOAD AND UNLOAD AREA TO BE SAMPLED MANY TIMES. THE FSR ADJUSTED THE RAISED PLATFORM SO THAT THE RACKS COULD NO LONGER CYCLE PAST THE RAISED PLATFORM WITHOUT THE INDICATOR BEING LIFTED. THE DAMAGED RACK INDICATORS WERE REPLACED. THE CAUSE OF THE CONTINUOUS CYCLING WAS ELIMINATED AND SYSTEM IS CURRENTLY WITHIN SPECIFICATION AND SATISFACTORY FOR USE. THE INVESTIGATION INCLUDED REVIEW OF SUBMITTED DATA, PRODUCT HISTORICAL DATA AND PRODUCT LABELING. NO CUSTOMER RETURNS WERE AVAILABLE FOR INVESTIGATION. IT WAS CONCLUDED THAT THE INCIDENT WAS INDICATIVE OF AN OPERATOR ERROR AND TRAINING ISSUE WHEN THE OPERATOR DECIDED TO REPORT THE LAST RUN PERFORMED IN OPEN MODE. A REVIEW OF THE CELL-DYN SAPPHIRE SYSTEM OPERATOR'S MANUAL, SECTION 10, TROUBLESHOOTING AND DIAGNOSTICS, FOUND THE MANUAL PROVIDES THE OPERATOR WITH BASIC TROUBLESHOOTING TECHNIQUES AND TOOLS RELATED TO THIS COMPLAINT INCIDENT. BASED ON THE ANALYSIS OF THE SUBMITTED DATA, THE COMPLAINT INCIDENT WAS AN ISOLATED INCIDENT INVOLVING AN OPERATOR TRAINING ISSUE AND A SERVICE ISSUE. A PRODUCT ISSUE FOR THE CD SAPPHIRE ANALYZER WAS NOT IDENTIFIED. | 0
4329378 | THE CUSTOMER STATED THAT THE CELL DYN SAPPHIRE HAS BEEN UNDERESTIMATING THE CD3 BY APPROXIMATELY 10-20% OVER THE PAST YEAR VERSUS THE BECKMAN COULTER FLOW CYTOMETRY. THE RESULTS PROVIDED BY THE CUSTOMER WERE: (B)(4) CD SAPPHIRE-CD4=36.9%; BECKMAN COULTER CD4=68%; CD SAPPHIRE - CD3=51.5%, CD8=12.6; BECKMAN COULTER CD3=87, CD8=16. THERE WAS NO REPORTED IMPACT TO PATIENT MANAGEMENT. THERE WAS NO ADDITIONAL PATIENT INFORMATION PROVIDED.><(B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING, AND A REVIEW OF SERVICE HISTORY. THE CUSTOMER OBSERVED THE CELL-DYN T-CELL IMMUNO ASSAY (CD 3/4/8), LIST NUMBER 07H22-01, GENERATING LOWER CD 4T % RESULTS WHEN COMPARED TO 2 OTHER NON-ABBOTT SYSTEMS. THE SUBMITTED RESULTS CAME FROM 3 DIFFERENT INSTRUMENTS, REAGENT SYSTEMS, AND PRESUMABLY, OPERATORS. FOR SUBJECT I, TESTED AT (B)(4), THE COMPARISON OF ABSOLUTE RESULTS WAS VERY GOOD. FOR SUBJECT II, THE RESULTS INDICATE THE SAMPLE WAS TESTED AT (B)(4) ONE DAY AFTER THE SAMPLE WAS COLLECTED AND TESTED ON THE CELL-DYN SAPPHIRE. THE DIFFERENCE BETWEEN SAMPLE PROCESSING OF THIS SPECIFIC SAMPLE MAY HAVE AFFECTED THE RESULTS. A REVIEW OF COMPLAINTS DID NOT IDENTIFY ANY SIMILAR ISSUES. A REVIEW OF SERVICE HISTORY FOR THE CELL DYN SAPPHIRE, SERIAL NUMBER (B)(4), ALSO DID NOT FIND ANY SIMILAR ISSUES AS THIS COMPLAINT. A REVIEW OF LABELING FOUND THAT THE PRODUCT LABELING ADEQUATELY ADDRESSED THE CUSTOMER'S ISSUE. BASED ON THE INVESTIGATION A PRODUCT ISSUE WAS NOT IDENTIFIED FOR THE CELL-DYN T-CELL IMMUNO ASSAY (CD 3/4/8), LIST NUMBER 07H22-01. | 0
4147848 | THE CUSTOMER OBSERVED A FALSELY DEPRESSED HEMOGLOBIN RESULT FOR ONE POST-OPERATIVE PATIENT ON THE CELL-DYN SAPPHIRE ANALYZER. THE INITIAL SAMPLE WAS COLLECTED FROM THE LOWER LIMB ARTERIAL BLOOD AND RESULTED WITH A HEMOGLOBIN OF 9.5 G/DL. THE SECOND SAMPLE, COLLECT THE FOLLOWING DAY AND FROM THE LOWER LIMB VENOUS BLOOD, RESULTED WITH A HEMOGLOBIN OF 6.6 G/DL. THE SECOND RESULT DID NOT MATCH THE CLINICAL CONDITION OF THE PATIENT SO THE SAMPLE WAS RECOLLECTED AND RERUN AND RESULTED IN A HEMOGLOBIN OF 9.7 G/DL. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED.><(B)(4). FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS AND A REVIEW OF LABELING. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. THE PRODUCT WAS NOT AVAILABLE FOR RETURN. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
3763448 | THE USER OF CELLPACK 20L REAGENT, PK-30L, CONTACTED THE SYSMEX TECHNICAL SERVICE CENTER (TAC) ON (B)(6), 2014 REGARDING A USER SAFETY ISSUE WHEN HANDLING THE REAGENT CUBE. THE COMPLAINANT INDICATED THAT DURING THE PROCESS OF LIFTING THE REAGENT BY THE PLASTIC STRAP THE OPERATOR TORE THE ROTATOR CUFF IN HIS SHOULDER. THE USER REPORTED THE INCIDENT TO OCCUPATIONAL HEALTH. THE INJURY REQUIRED SURGERY TO CORRECT AND THE OPERATOR WAS ABSENT FROM WORK FOR A SIX WEEK RECOVERY. THE DATE OF OCCURRENCE OF THE INCIDENT IS UNKNOWN. CELLPACK REAGENT IS ASSOCIATED WITH THE XE-5000 ANALYZER. CELLPACK IS A CLASS 1 MEDICAL DEVICE, PRODUCT CODE GIF, (B)(4).><SYSMEX REAGENTS HAVE A WHITE PLASTIC STRAP TIED AROUND THE BOX DESIGNED TO SECURE THE BOX FLAPS AND THE CONTENTS DURING SHIPPING. IT IS NOT MEANT TO PERFORM AS A HANDLE FOR TRANSPORT. THE CURRENT REAGENT BOX CONFIGURATION DOES NOT HAVE CUT-OUTS OR HANDLES. GOOD LABORATORY PRACTICES WOULD INDICATE USING SAFETY MEASURES RECOMMENDED BY THE OCCUPATIONAL SAFETY & HEALTH (OSHA) AND NIOSH (NATIONAL INSTITUTE OF OCCUPATIONAL HEALTH AND SAFETY) TO LIFT HEAVIER OBJECTS SUCH AS THE REAGENT CUBE. FOR EXAMPLE, LIFTING AIDS SUCH AS A DOLLY OR A CART FOR TRANSPORT OR LIFTING WITH ASSISTANCE OF ANOTHER EMPLOYEE WOULD AID THE USER IN TRANSPORTING THE REAGENT CUBE AND AVOIDING POTENTIAL INJURY. THE EMPLOYER IS RESPONSIBLE FOR ASSESSING RISK AND PROVIDING ADEQUATE TRAINING AND ENGINEERING CONTROLS FOR ALL REAGENTS AND OTHER MATERIALS USED IN A LABORATORY. SEVERAL ATTEMPTS WERE MADE BY THE TECHNICAL SERVICE QUALITY SOLUTIONS MANAGER (TSQSM) TO CONTACT THE HOSPITAL'S RISK MANAGEMENT AND HUMAN RESOURCES DEPARTMENT TO OBTAIN ADDITIONAL INFORMATION WITHOUT SUCCESS. IT IS UNKNOWN IF THE OPERATOR SUFFERED PERMANENT INJURY DUE TO THIS INCIDENT. THIS INCIDENT WILL BE REPORTED AS THE OPERATOR ENDURED SERIOUS INJURY TO THE SHOULDER WHICH REQUIRED SURGERY TO CORRECT. | 0
3593320 | THE USER OF AN XN-10 AUTOMATED HEMATOLOGY ANALYZERS (SERIAL NUMBERS (B)(4)) REPORTED THAT ON (B)(6) 2013, SAMPLE ID (B)(6) GENERATED A SPURIOUSLY LOW HEMOGLOBIN (HGB) RESULT OF 61 G/L; THE RESULT WAS REPORTED TO A CLINICIAN. SUBSEQUENTLY, THE PT WAS UNNECESSARILY TRANSFUSED BASED ON THE LOW HGB RESULT. THE SAMPLE WAS RERUN USING THE MANUAL MODE WITH AN OPEN TUBE; A NORMAL HGB RESULT OF 155G/L WAS OBTAINED. THE INITIAL ANALYSIS WAS PERFORMED USING THE AUTOMATIC MODE. NO FLAGS WERE GENERATED AND THE ANALYZER JUDGED THE SAMPLE AS NEGATIVE". AFTER REPORTING THE INITIAL RESULTS | 0
4088003 | THE CUSTOMER REPORTS A HEATER FAILURE ERROR MESSAGE (HTR FAIL) WITH A PRIOR TEMP TOO HIGH" MESSAGE GENERATED ON A TDXFLX ANALYZER. THE CUSTOMER POWERED THE ANALYZER OFF AND THEN BACK ON. THE CUSTOMER SAW VISIBLE SMOKE COMING FROM THE ANALYZER AND IMMEDIATELY POWERED THE ANALYZER BACK OFF. THERE WERE NO INJURIES OR DAMAGE TO THE SURROUNDING LAB ENVIRONMENT REPORTED. A SERVICE CALL WAS REQUESTED.><A TEMPERATURE TOO HIGH MESSAGE WAS GENERATED ON THE TDXFLX INSTRUMENT | 0
3793220 | THERE WAS NO PATIENT INVOLVEMENT IN THIS EVENT. THE DEVICE MALFUNCTIONED BECAUSE THE DEVICE PROMPTS DEVICE SERVICE REQUIRED. | 0
4275478 | MEDIVATORS FIELD SERVICE ENGINEER REPORTED THAT THIS FACILITY WAS POSSIBLY USING ENDOSCOPES ON PATIENTS THAT WERE NOT PROPERLY HIGH LEVEL DISINFECTED.><MEDIVATORS FIELD SERVICE ENGINEER REPORTED THAT THE ORIGINAL ISSUE WITH THE MEDIVATORS DSD-201 AUTOMATED ENDOSCOPE REPROCESSOR MACHINE WAS REPAIRED WITH A VALVE SEAL REPLACEMENT. THE SCOPES/CYCLES RAN ON THE SIDE OF THE MACHINE WHERE THE VALVE NEEDED REPLACEMENT ALL FAILED FOR NO FLUID FLOW" | 0
3874149 | THE FACILITY IS USING OLYMPUS SCOPES AND USING THE REUSABLE OLYMPUS MAJ-855 AUXILLIARY WATER TUBING AT BEDSIDE FOR FLUSHING. THE MEDIVATORS CLINICAL SPECIALIST WAS DOING AN INSERVICE ON THE MEDIVATORS DSD-201 & DSD-EDGE MACHINES AND DISCOVERED THIS FACILITY IS NOT PROPERLY DISINFECTING THE MAJ-855 BETWEEN PATIENT PROCEDURES. THE FACILITY STATED THAT THEY ONLY PUT ALCOHOL THROUGH IT AT THE END OF THE DAY. THE MEDIVATORS CLINICAL SPECIALIST INFORMED THEM THAT THE MAJ-855 NEEDS TO BE REPROCESSED BETWEEN EVERY PATIENT FOR THERE IS POTENTIAL FOR CROSS CONTAMINATION. THE STEPS OF APPROPRIATE REPROCESSING OF THE MAJ-855 IS WRITTEN IN THE OLYMPUS IFU FOR THIS PRODUCT AS WELL AS A MEDIVATORS PRODUCT BULLETIN. TO DATE, THERE ARE NO REPORTS OF PATIENT INJURY OR ILLNESS. THIS COMPLAINT WILL BE MONITORED WITHIN MEDIVATORS COMPLAINT SYSTEM.><THE CASE STATES THAT THE FACILITY IS NOT REPROCESSING THE REUSABLE OLYMPUS MAJ-855 AUXILLIARY WATER TUBING BETWEEN PATIENTS PROCEDURES. MEDIVATORS CLINICAL SPECIALIST DISCOVERED THIS WHILE DOING AN IN-SERVICE ON THE MEDIVATORS DSD -201 AND EDGE AERS. | 0
4012847 | THE CASE STATES THAT THIS FACILITY CALLED MEDIVATORS TO REPORT THAT A PATIENT RECENTLY HAD AN ENDOSCOPY AND IS SUSPECTED TO HAVE THE BOVIE SPONGIFORM ENCEPHALOPATHY (BSE) DISEASE. THE ENDOSCOPE USED WAS REPROCESSED IN THEIR MEDIVATORS DSD EDGE MACHINE. HIGH LEVEL DISINFECTION DOES NOT CLAIM TO INACTIVATE PRIONS AND MEDIVATORS' RECOMMENDATION IS FOR EQUIPMENT TO BE DECOMMISSIONED AND DESTROYED IF BSE EXPOSURE IS SUSPECTED MEDIVATORS RA DEPARTMENT TRIED CONTACTING THIS FACILITY ON JULY 25TH AND AUGUST 1ST. THERE WAS NO ANSWER. MEDIVATORS RA LEFT VOICEMAILS. IT IS SUSPECTED THAT THIS MACHINE IS STILL IN USE AT THIS FACILITY. TO DATE, THERE HAVE BEEN NO CONFIRMED REPORTS OF ADDITIONAL PATIENTS CONTRACTING BSE DISEASE. THIS POTENTIAL HIGH RISK COMPLAINT WILL CONTINUE TO BE MONITORED AND MAINTAINED WITHIN MEDIVATORS COMPLAINT SYSTEM. IF ADDITIONAL INFORMATION IS RECEIVED, THIS COMPLAINT WILL BE REVIEWED.><THE CASE STATES THAT THE FACILITY REPORTED THAT A PATIENT SUSPECTED TO HAVE BOVINE SPONGIFORM ENCEPHALOPATHY (BSE) HAD AN ENDOSCOPY AND THE SCOPE WAS REPROCESSED IN THEIR MEDIVATORS DSD EDGE MACHINE. POTENTIAL CROSS CONTAMINATION IF THIS MACHINE IS STILL IN USE. | 0
3746293 | FACILITY REPROCESSING OLYMPUS MAJ-855 AUXILIARY FLUSHING TUBE IN INCOMPATIBLE CER-1 AUTOMATED ENDOSCOPE REPRESSOR WHICH TUBING AND AUXILIARY WATER CHANNEL OF THE CONNECTED ENDOSCOPE IS POTENTIALLY NOT BEING HIGH LEVEL DISINFECTED. THEREFORE, INADEQUATE OR INAPPROPRIATE REPROCESSING MAY CAUSE PATIENT INJURY VIA CROSS CONTAMINATION.><THE CASE STATES THAT DURING A FACILITY VISIT A MEDIVATORS CLINICAL SPECIALIST OBSERVED THIS FACILITY REPROCESSING THE OLYMPUS MAJ-855 IN A MEDIVATORS CER-1. THE CER-1 AUTOMATED ENDOSCOPE REPROCESSOR IS NOT COMPATIBLE FOR REPROCESSING OF THE OLYMPUS MAJ-855 REUSABLE TUBING AND SUCH REPROCESSING IS IN DIRECT CONTRADICTION TO THE MEDIVATOR'S IFU. MEDIVATORS HAS A PRODUCT BULLETIN INFORMING USERS/CUSTOMERS OF THE MAJ-855 AND IT'S COMPATIBILITY WITH MEDIVATORS MACHINES. THE CLINICAL SPECIALIST ALSO REPORTED THIS FACILITY WAS NOT USING THE DISINFECTANT RESERVOIR THERMOMETER TO VERIFY THE TEMPERATURE OF THE HEATED DISINFECTANT. INSTEAD, THE FACILITY WAS RELYING ON THE DIGITAL SCREEN TEMPERATURE READING. THIS PROCEDURE ALSO DIRECTLY CONTRADICTS MEDIVATORS IFU AND DAILY QUALITY ASSURANCE PROCEDURE FOR THE CER-1. IT WAS ALSO OBSERVED THAT THIS FACILITY WAS USING NON-VALIDATED TUBING AND PARTS TO MODIFY HOOK- UPS TO REPROCESS THEIR ENDOSCOPES. THERE IS POTENTIAL RISK OF ILLNESS AND INJURY TO THE PATIENTS, WHO ARE TO RECEIVE PROCEDURES USING THEIR MAJ-855 AND/OR SCOPES BECAUSE THEY HAVE NOT BEEN DISINFECTED ADEQUATELY. MEDIVATORS COMPANY REPRESENTATIVE HAS BEEN IN COMMUNICATION WITH THE FACILITY IN REGARDS TO TRAINING AND OTHER MACHINES COMPATIBLE TO REPROCESS THE OLYMPUS MAJ-855. ALSO, THE FACILITY WAS EDUCATED REGARDING THE APPROPRIATE TEMPERATURE REFERENCE AND TUBING TO BE USED IN THE CER-1. TO DATE, THERE ARE NO CONFIRMED REPORTS OF PATIENT ILLNESS OR INJURY AS A RESULT OF THESE ISSUES. THIS POTENTIAL COMPLAINT WILL CONTINUE TO BE MONITORED WITHIN MEDIVATOR'S COMPLAINT SYSTEM. | 0
4117431 | HIGH LEVEL DISINFECTANT METRICIDE OPA PLUS WAS USED AS A BACKUP DISINFECTANT FOR THE MEDIVATORS DSD EDGE DUAL BASIN AER (AUTOMATED ENDOSCOPE PREPROCESSOR) WHICH IS ONLY DESIGNED TO USE RAPICIDE PA HIGH LEVEL DISINFECTANT PART A/PART B. OPA PLUS IS AN FDA APPROVED HIGH LEVEL DISINFECTANT (B)(4). THIS CHEMICAL CAN BE REUSED FOR UP TO 14 DAYS. WE DID NOT REUSE THIS CHEMICAL. THE AER PROCESS AT THE FACILITY RUNS 25 MINUTES AT 35 DEGREES CELSIUS. (B)(4). RAPICIDE PA: IS A SINGLE USE, PERACETIC ACID BASED SOLUTION, FDA APPROVED HIGH LEVEL DISINFECTANT THAT REQUIRES NO NEUTRALIZATION AND IS DESIGNED FOR USE IN MEDIVATORS DSD EDGE AER. (B)(4). CORRECTIVE ACTION HAS BEEN TAKEN TO ONLY USE RAPICIDE PA IN THE AER, PER MFR'S INSTRUCTIONS. TO DATE, THERE HAVE BEEN NO REPORTS OF PT INJURY OR ILLNESS AS A RESULT OF OUR PRIOR PROTOCOL. INFECTION RATE RELATED TO GI CASES PRIOR TO METRICIDE OPA PLUS USAGE: 0.05%. PRESENT INFECTION RATE RELATED TO GI CASES: 0%. WE WILL CONTINUE TO MONITOR OUR DATA FOR ANY CHANGES. DATES OR USE: (B)(6) 2013-(B)(6) 2014 (ESTIMATE) (B)(6) 2014-PRESENT. EVENT ABATED AFTER USE STOPPED OR DOSE REDUCED: YES. | 0
3592751 | CUSTOMER REPORTS OBSERVING SMOKE FROM THE UNINTERPRETABLE POWER SUPPLY (UPS) OF THE DEVICE. NO PT WAS PRESENT. NO HARM TO USER.><MANUFACTURE'S INITIAL REPORT DATE: 11/11/2013. FILE NO: (B)(4). CASE NO: (B)(4). ADDITIONAL DATE/FAILURE INVESTIGATION. DEVICE WAS RETURNED TO CARE FUSION FOR FAILURE INVESTIGATION. INVESTIGATION INDICATED THAT THERE WAS EVIDENCE OF FLUID INGRESSION INTO THE UPS WHICH RESULTED IN THE BURNED OUT COMPONENTS AND SMOKE THAT WAS OBSERVED. | 0
3986608 | (B)(6). A BECKMAN COULTER FIELD SERVICE ENGINEER (FSE) EVALUATED THE ANALYZER AND REPLACED THE SHEATH PRESSURE SENSOR MAKING THE UNIT OPERATIONAL. THE BECKMAN COULTER (BEC) INTERNAL IDENTIFIER FOR THIS REPORT IS (B)(4).><CUSTOMER REPORTED A BROKEN SHEATH PRESSURE SENSOR WHEN USING THE NAVIOS FLOW CYTOMETER SYSTEM. A BECKMAN COULTER FIELD SERVICE ENGINEER (FSE) WAS SENT TO THE CUSTOMER'S FACILITY TO EVALUATE THE ANALZYER. THE FSE REVEALED THAT A LEAK WAS ALSO PRESENT. THE LEAK WAS IDENTIFIED AS SHEATH FLUID. THE VOLUME OF THE LEAK WAS ABOUT 100ML AND WAS UNCONTAINED IN THE INSTRUMENT. THE OPERATOR WAS WEARING PERSONAL PROTECTIVE EQUIPMENT OF GLOVES WHEN THE LEAK WAS DISCOVERED. THERE WAS NO BIOHAZARD EXPOSURE TO MUCOUS MEMBRANES OR CUTS. THERE WERE NO ERRONEOUS TEST RESULTS ASSOCIATED WITH THIS EVENT. THERE WAS NO DEATH, INJURY OR AFFECT TO USER OR PATIENT TREATMENT. | 0
4190816 | THE CUSTOMER STATED THAT THE ARCHITECT C4000 ANALYZER GENERATED AN ELEVATED CREATININE RESULT OF 4.11 MG/DL. THE SAMPLE WAS REPEATED ON ANOTHER ANALYZER AND A RESULT OF 0.78 MG/DL WAS GENERATED. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><IT WAS DISCOVERED ON 11/13/2014 THAT THIS MDR (MANUFACTURER'S REPORT # 1415939-2014-00226) WAS FILED AGAINST THE INCORRECT MANUFACTURER SITE: (B)(4) AND SHOULD HAVE BEEN FILED UNDER MANUFACTURER SITE: (B)(4). MDR NUMBER 1628664-2014-00236 HAS BEEN SUBMITTED AND ALL FURTHER INFORMATION WILL BE DOCUMENTED UNDER THAT MDR NUMBER. | 0
3674739 | AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><THE CUSTOMER STATED THAT A FALSELY ELEVATED ARCHITECT MAGNESIUM RESULT OF >7.8 MG/DL WAS GENERATED FOR A PATIENT ON (B)(6) 2014. THE SAME PATIENT SAMPLE RETESTED WITHIN THE NORMAL RANGE AT 1.8 MG/DL ON THE SAME ARCHITECT C8000 ANALYZER AND 1.7/1.8 MG/DL ON TWO DIFFERENT ARCHITECT ANALYZERS ON (B)(6) 2014. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><THE ARCHITECT C8000 MAINTENANCE LOG SHOWS WEEKLY MAINTENANCE PROCEDURES (CLEAN SAMPLE/REAGENT PROBES AND CLEAN CUVETTES WITH DETERGENT) WERE PERFORMED BETWEEN THE INITIAL SUSPECT MAGNESIUM RESULTS ON (B)(6) 2014 AND THE LOWER RETEST RESULTS ON (B)(6) 2014 AND (B)(6) 2014. A REVIEW OF THE SERVICE HISTORY OF ARCHITECT C8000 SERIAL NUMBER (B)(4) FINDS NO ADDITIONAL REPORTS OF ERRATIC OR FALSE HIGH RESULTS AFTER THE WEEKLY MAINTENANCE ITEMS WERE COMPLETED. A REVIEW OF C8000 AND MAGNESIUM ASSAY COMPLAINT AND TRENDING DATA DID NOT IDENTIFY A RELATED PRODUCT ISSUE OR ADVERSE TREND. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND MAGNESIUM ASSAY PACKAGE INSERT PROVIDE ADEQUATE INFORMATION RELATED TO THE CUSTOMER'S REPORTED ISSUE. THE AVAILABLE INFORMATION INDICATES THE FALSE HIGH MAGNESIUM RESULT ISSUE WAS RESOLVED BY PERFORMING WEEKLY MAINTENANCE ITEMS TO CLEAN THE SAMPLE AND REAGENT PROBES AND THE CUVETTES. THE ISSUE WAS RESOLVED PER STANDARD TROUBLESHOOTING PROVIDED IN THE OPERATIONS MANUAL. BASED ON THE INVESTIGATION PERFORMED, NEITHER A MALFUNCTION NOR A DEFICIENCY WERE IDENTIFIED. | 0
4247946 | THE CUSTOMER REVIEWED THE ARCHITECT C16000 SYSTEM MESSAGE HISTORY LOG AND OBSERVED SEVERAL ASPIRATION ERRORS ASSOCIATED WITH THE R1A PROBE (REAGENT PROBE R1A) ON THE DAY OF OCCURRENCE. REPLACEMENT OF THE R1A PROBE RESOLVED THE ISSUE. REVIEW OF THE SUBSEQUENT INSTRUMENT SERVICE HISTORY REVEALED NO ADDITIONAL COMPLAINTS OF ERRATIC OR DISCREPANT RESULTS HAVE BEEN DOCUMENTED. THE R1A PROBE WAS DOCUMENTED AS THE MOST LIKELY CAUSE OF THE RESULT ISSUE WHICH WAS RESOLVED THROUGH NORMAL TROUBLESHOOTING PROCEDURES. THE PART WAS NOT AVAILABLE FOR RETURN. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND TOTAL BILIRUBIN PACKAGE INSERT PROVIDE ADEQUATE INFORMATION WITH REGARD TO SPECIMEN HANDLING, REQUIRED MAINTENANCE, COMPONENT REPLACEMENT, AND TROUBLESHOOTING THE REPORTED ISSUE. COMPLAINT TRACKING/ TRENDING REVIEW OF THE R1A PROBE DID NOT IDENTIFY AN ADVERSE TREND. THERE IS NO INDICATION OF A DEFICIENCY OF THE R1A PROBE OR THE ARCHITECT C16000.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><THE CUSTOMER STATED THAT FALSELY ELEVATED TOTAL BILIRUBIN RESULTS WERE GENERATED ON THE ARCHITECT C16000 ANALYZER. A PHYSICIAN CONTACTED THE CUSTOMER AND STATED THAT HIS PATIENT DOES NOT HAVE SYMPTOMS OF JAUNDICE AND REQUESTED REPEAT TESTING. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED. PATIENT ID: (B)(4): INITIAL RESULT OF 299 UMOL/L RETESTED AT 59 UMOL/L PATIENT ID: (B)(4): INITIAL RESULT OF 305 UMOL/L RETESTED AT 17 UMOL/L | 0
4017143 | THE CUSTOMER OBSERVED VISIBLE SMOKE FROM THE ARCHITECT C8000 ANALYZER. THE CUSTOMER NOTED THAT THE PUMP WILL NOT TURN ON AND HAS NOTED VISIBLE SMOKE AND BURNING SMELL WITH UNIT. NO LIQUID WAS FOUND ANYWHERE INSIDE THE PUMP OR PLUG AREA. THE CUSTOMER CALLED CUSTOMER SERVICE AND TURNED OFF THE INSTRUMENT. THERE WERE NO INJURIES.><FIELD SERVICE IDENTIFIED THE UNINTERRUPTIBLE POWER SUPPLY (UPS) AS THE LIKELY CAUSE OF THE ISSUE; THE VOLTAGE READINGS UPON MEASUREMENT DURING INVESTIGATION WERE NOT CORRECT. THE UPS ARC MODULE UPS, LN 01DD8-07) WAS IDENTIFIED AS THE POWERVAR, MODEL ABCDEF3600, SERIAL NUMBER (B)(4), WHICH IS A LOCALLY SOURCED NON-ABBOTT PRODUCT. THE UPS WAS SERVICED BY THE UPS REPAIR COMPANY. ABBOTT FIELD SERVICE REPLACED THE EXTERNAL WASTE PUMP WHICH WAS DAMAGED IN THE FIRE, AND THE INSTRUMENT IS NOW WORKING AS EXPECTED. A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF SERVICE HISTORY WERE PERFORMED AND NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. BASED ON ALL AVAILABLE INFORMATION AND ABBOTT DIAGNOSTICS' COMPLAINT INVESTIGATION, THE INSTRUMENT PERFORMED AS INTENDED AND NO PRODUCT DEFICIENCY WAS IDENTIFIED.><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). | 0
3690757 | (B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><AN ABBOTT FIELD SERVICE REPRESENTATIVE (FSR) WAS DISPATCHED TO THE ACCOUNT. THE FSR REPLACED A LEAKING REAGENT PROBE TUBING TO RESOLVE THE ISSUE. THERE WERE NO FURTHER DISCREPANT CO2 OR MAGNESIUM RESULTS AFTER THE TUBING REPLACEMENT. CUSTOMER COMPLAINT DATA WAS REVIEWED AND NO ADVERSE TRENDS WERE IDENTIFIED. THE ARCHITECT SYSTEM OPERATIONS MANUAL WAS REVIEWED AND WAS FOUND TO ADEQUATELY ADDRESS THE ISSUE. THE INVESTIGATION DID NOT IDENTIFY A MALFUNCTION / DEFICIENCY.><THE CUSTOMER STATED THAT THE ARCHITECT C4000 ANALYZER GENERATED ELEVATED CO2 AND MAGNESIUM RESULTS. PATIENT 1: AN INITIAL CO2 RESULT OF 36 MMOL/L WAS GENERATED WITH REPEAT RESULTS OF 22 AND 22 MMOL/L. PATIENT (B)(6): AN INITIAL CO2 RESULT OF 30 MMOL/L WAS GENERATED WITH A REPEAT RESULT OF 16 MMOL/L. PATIENT (B)(6): AN INITIAL CO2 RESULT OF 32 MMOL/L WAS GENERATED WITH A REPEAT RESULT OF 24 MMOL/L. PATIENT (B)(6): AN INITIAL MAGNESIUM RESULT OF >7 MG/DL WAS GENERATED WITH A REPEAT RESULT OF 1.78 MG/DL. THERE WAS NO REPORT OF IMPACT TO PATIENT MANAGEMENT. | 0
4096203 | THE ACCOUNT GENERATED A FALSELY ELEVATED MAGNESIUM OF 3.39 MMOL/L ON SPECIMEN ID (B)(6)THAT REPEATED NORMAL RANGE OF 0.80 MMOL/L ON THE ARCHITECT C16000 ANALYZER. THE ACCOUNT QUESTIONED THE ELEVATED MAGNESIUM OF 3.39 MMOL/L BECAUSE THE PREVIOUS RESULT FOR THE PATIENT WAS A LOT LOWER. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><(B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><THE CUSTOMER NOTICED DROPS OF FLUID FALLING INTO THE CUVETTE FROM THE PROBE AT THE CUVETTE WASHER. A FIELD SERVICE REPRESENTATIVE, (FSR) PERFORMED INSPECTIONS AND TROUBLESHOOTING, ALIGNED THE SAMPLE PROBE, CLEANED THE CUVETTE SEGMENT, AND REPLACED THE PERISTALTIC HEAD TUBING. NO CUSTOMER RETURNS WERE AVAILABLE FOR EVALUATION. A PRODUCT LABELING REVIEW FOUND THE ARCHITECT SYSTEM OPERATIONS MANUAL, AND THE MAGNESIUM ASSAY PACKAGE INSERT CONTAIN ADEQUATE INFORMATION FOR THE DESCRIBED ISSUE. A HISTORICAL/QUALITY DATA AND COMPLAINT DATA REVIEW DID NOT IDENTIFY ANY ADVERSE TREND OF A PERISTALTIC HEAD TUBING ISSUE. A COMPLAINT INVESTIGATION SERVICE HISTORY REVIEW FOUND NO SERVICE HISTORY ISSUES WITH ARCHITECT C16000 SERIAL (B)(4). NO ADDITIONAL DISCREPANT RESULT COMPLAINTS DUE TO INSTRUMENT ISSUES, OR COMPLAINTS FOR PERISTALTIC HEAD TUBING ISSUES WERE FOUND TO HAVE BEEN DOCUMENTED FOR ARCHITECT C16000 SERIAL (B)(4). THE MAGNESIUM DISCREPANT RESULT ISSUE WAS RESOLVED THOUGH THE REPLACEMENT OF A USED PERISTALTIC HEAD TUBING. NO SYSTEMIC DEFICIENCY OR MALFUNCTION WAS IDENTIFIED DURING THIS INVESTIGATION. | 0
3682360 | FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT DATA, TROUBLESHOOTING BY ABBOTT FIELD SERVICE, A SEARCH FOR SIMILAR COMPLAINTS, AND A REVIEW OF LABELING. ABBOTT FIELD SERVICE (FS) INVESTIGATED THE ISSUE ON SITE AND FOUND CUVETTE WASHER NOZZLE PAIR 2 WAS BLOCKED CAUSING THE CUVETTES TO OVERFLOW. FS REPLACED THE NOZZLE PAIR TO RESOLVE THE ISSUE. THE CUSTOMER STATED THEY'VE HAD NO FURTHER ISSUES. TRACKING AND TRENDING DID NOT IDENTIFY AN ADVERSE TREND. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. INFORMATION WAS PROVIDED TO THE CUSTOMER INDICATING A LIST OF PROBABLE CAUSES AND CORRECTIVE ACTIONS RELEVANT TO THE ISSUE IS DOCUMENTED IN THE OPERATIONS MANUAL. BASED ON THE AVAILABLE INFORMATION NO PRODUCT DEFICIENCY AND NO MALFUNCTION OF THE ARCHITECT C16000 ANALYZER, LIST NUMBER 03L77, WAS IDENTIFIED.><INITIAL REPORTER ADDRESS: BIOCHEMISTRY DEPT(ABBOTT MSC):(B)(6). (B)(4). AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE AN EVALUATION IS IN-PROCESS.><THE CUSTOMER OBSERVED FALSELY DECREASED CALCIUM RESULTS WHILE USING THE ARCHITECT C16000 ANALYZER. THE CUSTOMER PROVIDED RESULTS FOR 2 PATIENTS, AS FOLLOWS: PATIENT 1: INITIAL 0.51 MMOL/L, RETEST 2.13 MMOL/L. PATIENT 2: INITIAL 0.93 MMOL/L, RETEST 2.33 MMOL/L. THE RESULTS WERE NOT REPORTED FROM THE LABORATORY. THERE WAS NO IMPACT TO PATIENT MANAGEMENT. | 0
3849422 | THE CUSTOMER OBSERVED FALSELY ELEVATED MAGNESIUM RESULTS FOR 3 SAMPLES ON THE ARCHITECT C8000 ANALYZER. THE FOLLOWING DATA WAS PROVIDED (MG/DL): ONE OUTLIER GREATER THAN 7.8, WHEN REPEATED ON AN ALTERNATE INSTRUMENT RESULT WAS NORMAL (NO SPECIFIC RESULT PROVIDED). INITIAL 7.1, REPEAT 3.9. SID (B)(6) INITIAL 2.1, REPEATS 6.1, 2.1, 2.0. THERE WAS NO IMPACT TO PATIENT MANAGEMENT REPORTED.><AN EVALUATION IS IN PROCESS. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4).><FURTHER INVESTIGATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A SEARCH FOR SIMILAR COMPLAINTS, A REVIEW OF LABELING AND INSTRUMENT SERVICE. NO ADVERSE TREND WAS IDENTIFIED FOR THE CUSTOMER'S ISSUE. LABELING WAS REVIEWED AND FOUND TO BE ADEQUATE. FIELD SERVICE REPLACED THE MIXER AND FURTHER OPTIMIZED THE SYSTEM BY REPLACING THE R1 PROBE, CUVETTE DRYER TIP AND MANUALLY WASHED/CLEANED THE CUVETTES. BASED ON THE AVAILABLE INFORMATION, NO PRODUCT DEFICIENCY WAS IDENTIFIED. | 0
3906792 | AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). CONCOMITANT MEDICAL PRODUCTS AND THERAPY DATES: CLINICAL CHEMISTRY CARBON DIOXIDE, LN: 03L80-31, LOT: 44296UQ07.><NO RETURNS WERE AVAILABLE FROM THE CUSTOMER SITE. AN ABBOTT FIELD SERVICE ENGINEER AT THE CUSTOMER SITE PERFORMED THE FOLLOWING: REPLACED THE CUVETTE WASH PUMP BELLOWS AND POPPET VALVES, WATER LINE SEALS, ALL FOUR WASH SOLUTION VALVES, CUVETTE DRYER TIP, AND THE R2 REAGENT PROBE; CLEANED CUVETTE WASH NOZZLE NUMBER 1 AND ALL WASH CUPS; AND ADJUSTED WASH CUP FLOW RATES. A CARRYOVER TEST WAS PERFORMED AFTER TROUBLESHOOTING WAS COMPLETE AND SUBSEQUENT INSTRUMENT OPERATIONS AND ASSAY PERFORMANCE WAS ACCEPTABLE. A REVIEW OF COMPLAINT TRACKING AND TRENDING METRICS WAS PERFORMED AND IDENTIFIED NO ADVERSE TRENDS IN CONJUNCTION WITH THE COMPLAINT ISSUE CURRENTLY UNDER EVALUATION. THE ARCHITECT SYSTEM OPERATIONS MANUAL CONTAINS INFORMATION TO ADDRESS THE CURRENT CUSTOMER ISSUE. BASED ON THE AVAILABLE INFORMATION FROM THE CUSTOMER SITE AND FROM THE RESULTS OF THIS EVALUATION, THERE IS NO EVIDENCE TO REASONABLY SUGGEST A PRODUCT MALFUNCTION OCCURRED. THE ISSUE WAS ADDRESSED THROUGH STANDARD TROUBLESHOOTING PROCEDURES.><THE CUSTOMER REPORTS THE FOLLOWING: AN INITIAL ARCHITECT POTASSIUM ASSAY RESULT OF 9.89 MMOL/L GENERATED ON AN ARCHITECT C8000 ANALYZER ON SAMPLE (B)(6), A MALE PATIENT. THE RESULT DID NOT FIT THE PATIENT'S CLINICAL HISTORY OR OTHER LAB RESULTS AND WAS REPEATED. THE PLASMA SAMPLE RETESTED AT 3.75 MMOL/L. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB. THERE IS NO IMPACT TO PATIENT MANAGEMENT REPORTED. AN INITIAL PLASMA PATIENT SAMPLE GENERATING A CLIN CHEM CO2 ASSAY RESULT OF 39.77 MMOL/L ON AN ARCHITECT C8000 ANALYZER ON SAMPLE (B)(6). THE RESULT DID NOT FIT THE PATIENT'S CLINICAL HISTORY OR OTHER LAB RESULTS AND WAS REPEATED. THE SAMPLE RETESTED AT 26.21 MMOL/L. NO SUSPECT RESULTS WERE REPORTED FROM THE LAB WITH NO PATIENT IMPACT. | 0
4034884 | THE FIELD SERVICE REPRESENTATIVE REPLACED BROKEN CUVETTES TO RESOLVE THE ISSUE. THE ICT PROBE WAS ALSO CLEANED AS IT WAS DIRTY/CONTAMINATED. NO RETURNS WERE AVAILABLE FOR INVESTIGATION. A REVIEW OF DATA FOR THE PERIOD AUGUST 1, 2013 - SEPTEMBER 12, 2014 SHOWED CUVETTES AND THE ICT PROBE ARE USUALLY REPLACED TO ADDRESS FLUIDICS OR RESULTS ISSUES, OR AS PART OF PREVENTIVE MAINTENANCE. NO ADVERSE TRENDS WERE IDENTIFIED FOR THE CUVETTES OR THE ICT PROBE. NO OTHER ISSUES WERE IDENTIFIED FOR INSTRUMENT SERIAL NUMBER (B)(4) THAT MAY HAVE CONTRIBUTED TO THE COMPLAINT ISSUE. A REVIEW OF THE PRODUCT LABELING CONCLUDED THAT THE ISSUE IS SUFFICIENTLY ADDRESSED. BASED ON THE AVAILABLE INFORMATION, A DEFICIENCY WAS NOT IDENTIFIED.><(B)(4). A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS.><THE ACCOUNT GENERATED A FALSELY DEPRESSED ARCHITECT SODIUM RESULT OF 111 ON A PATIENT SAMPLE THAT REPEATED AT 141 ON THE ARCHITECT C16000 ANALYZER. NO UNIT OF MEASURE FOR SODIUM WAS PROVIDED. THE ACCOUNT USES A SODIUM REFERENCE RANGE OF 134 TO 148. NO SPECIFIC PATIENT INFORMATION WAS PROVIDED. NO IMPACT TO PATIENT MANAGEMENT WAS REPORTED. | 0
3904417 | THE CUSTOMER STATED THAT THE ARCHITECT ANALYZER GENERATED A FALSELY ELEVATED MAGNESIUM (MG) RESULT ON ONE PATIENT SAMPLE (SID (B)(6)). RESULTS PROVIDED WERE: INITIAL 3.55MMOL/L / REPEATED 3X = 1.14 / 0.89 / 0.91MMOL/L (NORMAL RANGE FOR MALE IS 0.66-1.07MMOL/L). THERE WAS NO REPORTED IMPACT TO PATIENT MANAGEMENT. THERE WAS NO ADDITIONAL PATIENT INFORMATION PROVIDED.><(B)(4). A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE.><FURTHER EVALUATION OF THE CUSTOMER ISSUE INCLUDED A REVIEW OF THE COMPLAINT TEXT, A REVIEW FOR SIMILAR COMPLAINTS, A REVIEW OF SERVICE RECORDS, AND A REVIEW OF LABELING. THE CUSTOMER OBSERVED QUALITY CONTROL (QC) IMPRECISION FOR LDH AND LIPASE FOLLOWED BY A SAMPLE THAT GENERATED A DISCREPANT HIGH MG RESULT THAT WAS NORMAL WHEN RETESTED ON ARCHITECT, SERIAL NUMBER (B)(4). ABBOTT FIELD SERVICE INVESTIGATED THE ISSUE ON SITE AND RESOLVED THE ISSUE BY TIGHTENING THE R1 SYRINGE PUMP CLAMP (PN 2-89111-02, CLAMP, SYRINGE, SAMP/REAG PUMP). REVIEW OF THE SERVICE HISTORY ON SN (B)(4) DOES NOT SHOW ADDITIONAL REPORTS OF IMPRECISION OR ERRATIC RESULTS. THE SYRINGE PUMP CLAMP PN 2-89111-02 WAS NOT AVAILABLE FOR RETURN AS IT REMAINED IN USE ON (B)(4). A REVIEW OF COMPLAINTS FOR PART NUMBER 2-89111-02 DID NOT IDENTIFY ANY SIMILAR COMPLAINTS OR ANY PRODUCT ISSUES. A REVIEW OF PRODUCT LABELING SHOWS ADEQUATE INFORMATION WITH REGARD TO SPECIMEN HANDLING, REQUIRED MAINTENANCE, COMPONENT REPLACEMENT, AND TROUBLESHOOTING OF THE REPORTED ISSUE. A MALFUNCTION IS NOT INDICATED AS THE DEVICE PERFORMS AS INTENDED WHEN PROPERLY MAINTAINED. BASED ON THE INVESTIGATION PERFORMED A DEFICIENCY IS NOT IDENTIFIED. | 0
3923672 | THE CUSTOMER STATED THAT AN ARCHITECT SODIUM RESULT OF 118 MMOL/L WAS GENERATED FOR A PATIENT WITH HISTORY OF 140 MMOL/L RESULTS. THE SAMPLE WAS RETESTED AFTER REPLACING THE ICT MODULE AND THE RESULT WAS 140 MMOL/L. NO ADVERSE IMPACT TO PATIENT MANAGEMENT WAS REPORTED.><ICT MODULE 131217027 WAS NOT AVAILABLE FOR RETURN. THE ARCHITECT C803079 INSTRUMENT LOGS VERIFY THE ISSUE. THE ICT REFERENCE SOLUTION MILLIVOLT (MV) DATA SHOWS THE SUSPECT SAMPLE HAS A LINGERING EFFECT ON THE SODIUM MEMBRANE. NO OTHER COMPLAINTS AGAINST ICT 131217027 HAVE BEEN RECEIVED. A REVIEW OF COMPLAINT TRACKING AND TRENDING FOR ICT MODULE LN 09D28-03 DID NOT IDENTIFY A PRODUCT ISSUE OR ADVERSE TREND. THE ARCHITECT SYSTEM OPERATIONS MANUAL AND ICT SAMPLE DILUENT PACKAGE INSERT PROVIDE ADEQUATE INFORMATION WITH REGARD TO TROUBLESHOOTING THE REPORTED ISSUE. THE AVAILABLE INFORMATION INDICATES THE SODIUM MEMBRANE OF ICT MODULE 131217027 HAD BEEN EXPOSED TO AGING AND AN UNKNOWN CONTAMINANT WHICH INTERACTS WITH AN INTERFERENT/CONTAMINANT UNIQUE TO SAMPLES FROM THE PATIENT IN QUESTION. THE ISSUE WAS RESOLVED BY INSTALLING A DIFFERENT ICT MODULE OR TESTING THE SAMPLES ON A DIFFERENT ICT MODULE. A MALFUNCTION IS NOT INDICATED AS THE ICT MODULE PERFORMS AS INTENDED WHEN NOT CONTAMINATED. BASED ON THE INVESTIGATION PERFORMED, A DEFICIENCY IS NOT IDENTIFIED.><AN EVALUATION IS IN PROCESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. (B)(4). | 0
4291853 | ADDITIONAL INFORMATION WAS RECEIVED AFTER INITIAL REPORT. THE ACCOUNT USES A LITHIUM NORMAL RANGE OF 0.50 TO 0.70 MMOL/L. ID (B)(6) IS A FEMALE AND INTENDED LITHIUM CONCENTRATION IS 0.60 MMOL/L. THE LITHIUM DOSAGE WAS LOWERED FOR THIS PATIENT DUE TO THE ELEVATED MULTIGENT LITHIUM RESULT. AFTER A WEEK, THE LITHIUM DOSAGE WAS INCREASED TO THE PREVIOUS LEVEL AFTER THE SAMPLE WAS RERUN WITH NO IMPACT TO PATIENT HEALTH.><THE ACCOUNT GENERATED FALSELY ELEVATED MULTIGENT LITHIUM ON 4 PATIENT SAMPLES PROCESSED ON THE ARCHITECT C8000 ANALYZER. THE SAMPLES WERE REPEATED WITH THE LOWER EXPECTED LITHIUM RESULT. THE LITHIUM MEDICATION WAS LOWERED FOR ONE PATIENT DUE TO THE ELEVATED MULTIGENT LITHIUM RESULT, BUT WAS SET ACCORDINGLY AFTER THE SAMPLE WAS RERUN. NO IMPACT TO PATIENT HEALTH WAS REPORTED. NO ADDITIONAL DETAILED PATIENT INFORMATION IS AVAILABLE. DATA PROVIDED: ID (B)(6) (FEMALE): ELEVATED MULTIGENT LITHIUM =1.40 MMOL/L, BUT REPEATED 0.60 MMOL/L (ANOTHER ARCHITECT C8000). ID (B)(6): ELEVATED MULTIGENT LITHIUM =1.13 MMOL/L BUT REPEATED 0.58 AND 0.60 MMOL/L (ANOTHER ARCHITECT C8000). ID (B)(6): ELEVATED MULTIGENT LITHIUM =1.04 MMOL/L BUT REPEATED 0.51 AND 0.53 MMOL/L (ANOTHER ARCHITECT C8000). ID (B)(6): ELEVATED MULTIGENT LITHIUM =1.11 MMOL/L BUT REPEATED 0.70 AND 0.73 MMOL/L (ANOTHER ARCHITECT C8000).><NO CUSTOMER RETURNS WERE AVAILABLE FOR EVALUATION. TROUBLESHOOTING AND ON SITE SERVICE WAS PERFORMED WITH PART REPLACEMENT INCLUDING THE ARCHITECT SAMPLE PROBE. A HISTORICAL / QUALITY DATA AND COMPLAINT REVIEW DID NOT IDENTIFY ANY ADVERSE TREND OF A SAMPLE PROBE DISCREPANT RESULT ISSUE. A SERVICE HISTORY REVIEW FOUND NO ADDITIONAL DISCREPANT RESULT COMPLAINTS AND NO COMPLAINTS FOR SAMPLE PROBE ISSUES HAVE BEEN DOCUMENTED FOR ARCHITECT C8000 SERIAL NUMBER (B)(4), SINCE THE ARCHITECT C8000 WAS SERVICED AND THE SAMPLE PROBE WAS REPLACED. A PRODUCT LABELING REVIEW FOUND THE ARCHITECT SYSTEM OPERATIONS MANUAL, AND THE LITHIUM ASSAY PACKAGE INSERT CONTAIN ADEQUATE INFORMATION FOR THE DESCRIBED ISSUE. THE LITHIUM DISCREPANT RESULT ISSUE WAS RESOLVED THOUGH THE REPLACEMENT OF A SAMPLE PROBE. NO SYSTEMIC PRODUCT ISSUE WAS IDENTIFIED DURING THIS INVESTIGATION.><(B)(4). A FOLLOWUP REPORT WILL BE SUBMITTED WHEN THE EVALUATION IS COMPLETE. EVALUATION IN PROCESS. | 0
3601831 | THIS INSTRUMENT HAS BEEN INVESTIGATED. ON (B)(4) 2014 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND INSPECTED THE INSTRUMENT AND PERFORMED A CAMERA ADJUSTMENT. THE CUSTOMER RAN CONTROLS AND ACCEPTED THE RESULTS. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. (B)(4).><THE CUSTOMER STATES THE ORTHO PROVUE MISSED AN ANTIBODY THAT WAS DETECTED IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. | 0
4216793 | THE OCD REP PROVIDED CTS WITH A SUMMARY OF THE COMPLAINT AND IMAGES OFF OF THE PROVUE AND VISION. UPON INSPECTION OF THE IMAGES PROVIDED, DISTURBANCE OF THE CELL BUTTON WERE NOTED WITH THE VISION RESULTS BUT NONE OFF OF THE ORIGINAL AND REPEAT TESTING OF THE PROVUE. THE DATA COLLECTED AND DOCUMENTATION INDICATES THAT THE MATRIX OF THE FLUID BEING TESTED APPEARS TO BE THE CAUSE OF THE CONCERN YIELDING A WEAK POSITIVE RESULT AGAINST BOTH DONORS IN QUESTION.><THE ORTHO PROVUE RESULTED A FALSE NEGATIVE REACTION FOR CELL#4 AND CELL#7 (HOMOZYGOUS FOR C ANTIGEN) TO THE 0.8% RESOLVE PANEL A WITH A SAMPLE KNOWN TO CONTAIN ANTI-C IN THE PLASMA AND 1+ IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4). | 0
3983619 | CUSTOMER REPORTS PATIENT SAMPLES THAT CLEARLY SHOWED DUAL POPULATION OF CELLS IN THE ANTI-D MICROWELL WERE CALLED 4+ BY THE PROVUE AND WERE NOT FLAGGED AS DP" AS EXPECTED. CUSTOMER REPORTED TWO INCIDENTS ON THIS INSTRUMENT. FIRST INCIDENT | 0
4152032 | OCD REP CONTACTING CTS TO REPORT OF A FALSE NEGATIVE REACTION WITH A SAMPLE KNOWN TO CONTAIN ANTI-S IN THE PLASMA AGAINST CELL#3 AND 4 TO THE 0.8% RESOLVE PANEL C LOT# VRC196 WHEN TESTED ON THE PROVUE.><DATA COLLECTED AND DOCUMENTATION INDICATES THAT THE MATRIX OF THE FLUID BEING TESTED APPEARS TO BE THE CAUSE OF THE CONCERN. THE CONCERN IS ISOLATED TO THE ONE TEST OF THE PATIENT SAMPLE THAT CANNOT BE DUPLICATED. THE INVESTIGATION DETERMINED THAT THE MOST LIKELY ROOT CAUSE OF THIS EVENT WAS SAMPLE RELATED. | 0
4216761 | THE ORTHO PROVUE RESULTED A FALSE NEGATIVE REACTION FOR AN AUTO CONTROL TEST WITH A SAMPLE KNOWN TO CONTAIN ANTI-E IN THE PLASMA AND 1+ IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4).><THE OCD REP PROVIDED CTS WITH A SUMMARY OF THE COMPLAINT AND IMAGES OFF OF THE PROVUE AND VISION. UPON INSPECTION OF THE IMAGES PROVIDED, DISTURBANCE OF THE CELL BUTTON WERE NOTED WITH THE VISION RESULTS BUT NONE OFF OF THE ORIGINAL AND REPEAT TESTING OF THE PROVUE. THE DATA COLLECTED AND DOCUMENTATION INDICATES THAT THE MATRIX OF THE FLUID BEING TESTED APPEARS TO BE THE CAUSE OF THE CONCERN YIELDING A VERY WEAK POSITIVE AUTO CONTROL RESULT . | 0
4216759 | THE OCD REP PROVIDED CTS WITH A SUMMARY OF THE COMPLAINT AND IMAGES OFF OF THE PROVUE AND VISION. UPON INSPECTION OF THE IMAGES PROVIDED, DISTURBANCE OF THE CELL BUTTON WERE NOTED WITH THE VISION RESULTS BUT NONE OFF OF THE PROVUE. DATA COLLECTED AND DOCUMENTATION INDICATES THAT THE MATRIX OF THE FLUID BEING TESTED APPEARS TO BE THE CAUSE OF THE CONCERN. THE CONCERN IS ISOLATED TO THE ONE TEST WITH A WEAK EXPRESSION OF THE ANTI-E ANTIBODY.><THE ORTHO PROVUE REPORTED A FALSE NEGATIVE REACTION FOR THE AHG CROSS MATCH TEST WITH A SAMPLE KNOWN TO CONTAIN ANTI-E IN THE PLASMA AGAINST A DONOR KNOWN TO HAVE THE E ANTIGEN. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4). | 0
3983618 | CUSTOMER REPORTS PATIENT SAMPLES THAT CLEARLY SHOWED DUAL POPULATION OF CELLS IN THE ANTI-D MICROWELL WERE CALLED 4+ BY THE PROVUE AND WERE NOT FLAGGED AS DP" AS EXPECTED. CUSTOMER REPORTED TWO INCIDENTS ON THIS INSTRUMENT. FIRST INCIDENT | 0
3843993 | THE CUSTOMER REPORTS THAT THE PROVUE GAVE A NEGATIVE RESULT FOR A SAMPLE THAT WAS POSITIVE VISUALLY AND IN MANUAL GEL. NO ERRONEOUS RESULTS WERE REPORTED. (B)(4).><ON 5-23-2014 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND VERIFIED THE GRIPPER ADJUSTMENT. THE FE PERFORMED A READER CAMERA ADJUSTMENT AND GENERATED A NEW REFERENCE IMAGE. WADDIAG SUCCESSFULLY PASSED. THE FE VERIFIED CARDS PICTURES, ALL WERE ACCEPTABLE. THE CUSTOMER RAN QC AND ACCEPTED RESULTS. THE CUSTOMER ALSO RAN THE SAME SAMPLE MANUALLY AND USING THE PROVUE. THE CUSTOMER ACCEPTED REPAIR. REPAIRS HAVE RETURNED THIS INSTRUMENT TO EXPECTED OPERATION. | 0
3781041 | THE CUSTOMER REPORTS FALSE NEGATIVE REACTIONS WITH 4 PATIENT SAMPLES WHEN USING THE ORTHO PROVUE IN THE ABO FORWARD TEST. NO ERRONEOUS RESULTS WERE REPORTED. THIS IS INCIDENT 3 OF 4. (B)(4).><THIS INSTRUMENT HAS BEEN INVESTIGATED. ON (B)(4) 2014 AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND CHECKED THE SYRINGE FOR LEAKING, CHECKED THE WASH BLOCK FOR CRACKS, CHECKED ALL BOTTLE CAPS AND FITTINGS, CLEANED HANDLER AND REAGENT CAROUSELS, AND PERFORMED ALL WAD DIAGNOSTICS TESTS. ALL TESTS PASSED. THE INSTRUMENT IS OPERATING AS EXPECTED. NO REPAIRS NEEDED. | 0
4099181 | THE CUSTOMER IS REPORTING FALSE NEGATIVE REACTIONS ON THE PROVUE FOR THE ANTIBODY SCREEN TEST WITH A PATIENT KNOWN TO BE POSITIVE FOR ANTIBODIES. THE PATIENT SAMPLE WAS NEGATIVE ON THE PROVUE YET POSITIVE WHEN USING MANUAL GEL. THE PATIENT HAS A PREVIOUS HISTORY OF BEING POSITIVE FOR ANTIBODIES 4 DAYS EARLIER. THE SAMPLE HAD BEEN TESTED ON THE PROVUE AND WAS POSITIVE WITH THE SELECTOGEN CELLS. ANTI-KELL AND ANTI-LITTLE E WERE IDENTIFIED WITH PANEL A USING MANUAL GEL.><ON (B)(4) 2014, AN OCD FIELD ENGINEER (FE) ARRIVED AT THE CUSTOMER SITE AND VERIFIED THAT CENTRIFUGATION AND INCUBATOR TEMPERATURE WERE CORRECT. FE MADE A CAMERA OPTICAL ADJUSTMENT AND NEW REFERENCE IMAGE. CUSTOMER RAN QC AND SOME 1+ PATIENTS TO COMPARE. ALL QC PASSED AND ALL PROVUE TESTS MATCHED MANUAL WITH THE EXCEPTION OF THE SUBJECT PATIENT. FE SUGGESTED THAT THE PROBLEM IS UNIQUE TO THE SAMPLE IN QUESTION. NO ERRONEOUS RESULTS REPORTED. NO TRANSFUSION OR TREATMENT WAS DELAYED, INITIATED OR ALTERED BECAUSE OF THIS INCIDENT. THE ROOT CAUSE OF THIS EVENT COULD POTENTIALLY BE PATIENT RELATED. | 0
3736082 | THE PT IS UNDERGOING TPE PROCEDURES DUE TO A REJECTED KIDNEY TRANSPLANT. THIS SAME PT HAS UNDERGONE APPROXIMATELY 10 TPE PROCEDURES ON THE AMICUS DEVICE. FOR THIS PARTICULAR PROCEDURE, THE OPERATOR CHANGED THE TARGET FLUID BALANCE FROM 100% TO 90% AFTER 550ML OF WHOLE BLOOD (WB) PROCESSED. THE OPERATOR FURTHER ADJUSTED THE TARGET FLUID BALANCE FROM 90% TO 80% WITHIN MINUTE OF THE FIRST ADJUSTMENT. AN 80% TARGET FLUID BALANCE EQUATES TO LEAVING THE PT HYPOVOLEMIC BY 20% OF THEIR TOTAL BLOOD VOLUME. THE PT HAS PERIPHERAL ACCESS AND WAS SHOWING SIGNS OF INLET LINE OCCLUSIONS. IT IS LIKELY THAT THE OPERATOR INTENDED TO CHANGE THE MAXIMUM WB FLOW RATE, BUT CHANGED THE TARGET FLUID BALANCE INSTEAD. THE MACHINE DID AS INSTRUCTED BY THE OPERATOR AND TARGETED 80% FOR FLUID BALANCE WHICH LED TO THIS INCIDENT. THE NURSE TOLD THE FSE THAT IT WAS VERY LONG DAY AND THIS PROCEDURE AS AROUND 7 PM, SO SHE WAS VERY TIRED. THE PT FAINTED AT THE END OF THE PROCEDURE, HE HAD A BLOOD PRESSURE AT THAT MOMENT OF 65/40. THE TREATMENT THE NURSE GAVE WAS A SALINE SOLUTION BOLUS OF TWO UNITS (TOTAL 1000ML) WITH A RECOVERY BLOOD PRESSURE OF 90/48. AFTER RECOVERY THE PERSON LEFT THE HOSPITAL ASYMPTOMATIC. IT HAS BEEN DIFFICULT FOR FENWAL'S RISK MGR TO SPEAK DIRECTLY TO THE PT; HOWEVER, THE NURSE SPOKE TO HIS WIFE AND SHE STATED HE IS FINE AND WILL GO TO THE NEXT SCHEDULED TOTAL PLASMA EXCHANGE (TPE) TREATMENT ON THE AMICUS MACHINE BECAUSE HE TRUSTS THE AMICUS. | 0
3951126 | ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTEDLY NEGATIVE ANTIBODY SCREEN WHEN TESTING A BLOOD SAMPLE ON A GALILEO ECHO.><IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGES ON THE TESTING INSTRUMENT ON (B)(4) 2014. | 0
3895248 | ON (B)(6) 2014, A CUSTOMER REPORTED UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN USING CAPTURE-R READY-SCREEN (3) ON A GALILEO ECHO INSTRUMENT ON (B)(6) 2014.><IMMUCOR TECHNICAL SUPPORT ASSESSED INSTRUMENT WELL IMAGES VIA A REMOTE ELECTRONIC CONNECTION METHOD ON (B)(4) 2014, BY VIEWING ARCHIVED RESULTS ON AN ARCHIVE DISK. | 0
4034136 | ><THE ARTIFACT WAS IDENTIFIED ON THE GE LIGHTSPEED VCT SCANNER LOCATED AT (B)(6) IN (B)(6). THIS WAS A VERY UNUSUAL SYSTEM DEFECT THAT CAUSED SUBTLE AND INTERMITTENT ARTIFACTS THAT MIMIC SOME TYPE OF PATHOLOGY IN THE BRAIN. | 0
3846013 | IMMUCOR TECHNICAL SUPPORT USED REMOTE ELECTRONIC METHOD TO ASSESS IMAGES OF TEST WELLS ON INSTRUMENT ON (B)(4) 2014. CUSTOMER DID NOT RETURN ANY PATIENT BLOOD SAMPLE OR CUSTOMER PRODUCT TO IMMUCOR FOR TESTING.><ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN USING CAPTURE-R READY-SCREEN (3) ON A GALILEO ECHO INSTRUMENT. | 0
3882209 | IMMUCOR TECHNICAL SUPPORT ASSESSED INSTRUMENT WELL IMAGES VIA A REMOTE ELECTRONIC CONNECTION METHOD.><ON (B)(6) 2014, A CUSTOMER EVENT WAS DOCUMENTED AS AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN USING CAPTURE-R READY-SCREEN (3) ON A GALILEO ECHO INSTRUMENT ON (B)(6) 2014. | 0
4006755 | ON (B)(6) 2014, INTERNAL EVALUATIONS WITH THE ECHO PHENOTYPE ASSAYS SHOWED UN-RESUSPENDED RED CELLS RESULTED IN POSITIVE INTERPRETATIONS.><THE MITIGATION PLAN IS IMPLEMENTATION OF EXPANDED EQUIVOCAL RANGE CUT OFF VALUES FOR THE ASSOCIATED ASSAYS. CUSTOMERS WERE NOTIFIED VIA CUSTOMER COMMUNICATION (B)(4) ON (B)(4) 2014. (B)(4). | 0
4308907 | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT TEST WELL IMAGES ON 11NOV2014. AN IMMUCOR FIELD SERVICE ENGINEER VISITED THE CUSTOMER SITE ON 12NOV2014 TO ASSESS THE INSTRUMENT.><ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN TESTING A BLOOD SAMPLE ON A GALILEO ECHO, WHEN TESTED ON (B)(6) 2014. | 0
4085295 | ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN TESTING A PATIENT SAMPLE ON THE GALILEO ECHO.><CUSTOMER COMMUNICATION (B)(4) ADDRESSES THE ISSUE OF CRRS(3) ASSAYS REPORTING NEGATIVE INTERPRETATION, WHEN THE IMAGES APPEAR VISUALLY POSITIVE. | 0
3567708 | ON (B)(6) 2013, A CUSTOMER REPORTED OBTAINING UNEXPECTED ABO BLOOD MISTYPE WHEN TESTED ON THE GALILEO.><IMMUCOR TECHNICAL SUPPORT USED REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS IMAGES OF WELL RESULTS ON CUSTOMER INSTRUMENT ON (B)(4) 2013. NO PATIENT SAMPLE OR CUSTOMER PRODUCT WAS RETURNED TO IMMUCOR FOR IMMUCOR INVESTIGATION. | 0
4042561 | THE CUSTOMER SITE WAS UNWILLING TO PROVIDE ADDITIONAL INFORMATION OR PARTICIPATE IN TROUBLESHOOTING THE ISSUE. SITE UNWILLING TO TROUBLESHOOT.><ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN TESTING A BLOOD SAMPLE ON A GALILEO ECHO. | 0
3567693 | IMMUCOR TECHNICAL SUPPORT USED REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS IMAGES OF WELL RESULTS ON CUSTOMER INSTRUMENT ON (B)(4) 2013. NO PATIENT SAMPLE OR CUSTOMER PRODUCT WAS RETURNED TO IMMUCOR FOR IMMUCOR INVESTIGATION.><ON (B)(6) 2013, A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE ANTIBODY SCREEN RESULTS WITH CAPTURE-R READY-SCREEN (3), WHEN USED ON THE GALILEO ECHO. | 0
3929094 | THE PRINTER IS NOT MANUFACTURED BY IMMUCOR, BUT IS INSTEAD MANUFACTURED BY HEWLETT PACKARD (HP). IMMUCOR ARRANGED DELIVERY OF A NEW PRINTER (HP LASERJET P2035) FROM THE IMMUCOR WAREHOUSE TO THE CUSTOMER SITE, WHICH WAS THEN INSTALLED SUCCESSFULLY BY THE CUSTOMER, WITH OVER-THE-PHONE ASSISTANCE BY IMMUCOR TECHNICAL SUPPORT ON (B)(4) 2014.><ON (B)(6) 2014, A CUSTOMER EVENT WAS DOCUMENTED THAT THE PRINTER (LINKED TO A GALILEO ECHO INSTRUMENT) EMITTED SMOKE INTO LABORATORY ENVIRONMENT. | 0
3987056 | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE TEST WELL IMAGES ON THE TESTING INSTRUMENT ON (B)(4) 2014. AS IS EXPECTED WITH EXAMPLES OF A2B (WITH ANTI-A1), THE ANTI-A TEST WELL SHOWED VERY SLIGHT AGGLUTINATION, BELOW THE DETECTION LIMIT OF THE GALILEO ECHO (AS DESCRIBED IN THE INSTRUMENT OPERATOR MANUAL). CONSEQUENTLY, AN UNEXPECTED, BUT BELIEVABLE ABO TYPE B WAS GENERATED.><ON (B)(6) 2014, A CUSTOMER REPORTED AN UNEXPECTED ABO MISTYPE WHEN TESTING A BLOOD SAMPLE ON A GALILEO ECHO. | 0
3568498 | A CUSTOMER REPORTED THAT AN RH DISCREPANCY OCCURRED ON THE NEO FOR A PATIENT WITH A HISTORY OF BEING RH NEGATIVE.><THE IMAGE RESULT FILES WERE REVIEWED. INITIAL TESTING SHOWED THAT THE TEST WELL CONTAINING THE ANTI-D REAGENT APPEARED TO DISPLAY A HALF-MOON APPEARANCE OF THE CELL BUTTON WHICH CAN BE ATTRIBUTED TO THE IMPROPER WASHING OF THE PLATE. THE CUSTOMER WAS ADVISED TO REMOVE AND CLEAN THE WASHER MANIFOLD AND TO ADJUST THE WVR. ADDITIONAL TESTING WAS PERFORMED WITH THE SAMPLE AND THE EXPECTED RESULTS WERE OBTAINED. | 0
3479555 | A CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIVITY ON GALILEO ECHO (B)(4).><A REVIEW OF THE IMAGE RESULT FILES SHOWED THAT NEGATIVE REACTIONS FOR TESTING PERFORMED ON (B)(4) APPEARED WEAK POSITIVE. REACTIONS APPEARED AS REPORTED BY THE INSTRUMENT ON ECHO (B)(4). AN IMMUCOR FIELD SERVICE ENGINEER PERFORMED THE UNEXPECTED REACTIONS CHECKLIST. THE 1ML SYRINGE, PBS FILTER, WASHER MANIFOLD, AND PROBE WERE REPLACED; CALIBRATED CAMERA, AND DECONTAMINATED THE PBS CONTAINER AND INSTRUMENT. QC WAS PERFORMED WITH ACCEPTABLE RESULTS. THE INSTRUMENT IS OPERATING ACCORDING TO SPECIFICATIONS. | 0
3446971 | THE EVENT IS REPORTED AS: THE CUSTOMER ALLEGES THAT THE DEVICE WILL NOT HEAT. NO REPORT OF A PATIENT INJURY OR DELAY IN TREATMENT.><THE DEVICE SAMPLE WAS NOT RETURNED FOR EVALUATION AT THE TIME OF THIS REPORT. | 0
3494513 | THE CUSTOMER REPORTED THAT NO AHG WAS ADDED TO SOME OF THE TEST PLATE'S WELLS WHEN TESTING ON TANGO OPTIMO. THE CUSTOMER NOTICED THAT THE AHG WAS NOT ADDED ON SOME OF THE SAMPLES BECAUSE THE BACKGROUND ON THE IMAGES REMAINED WHITE. THE CUSTOMER ALSO MENTIONED THE TANGO OPTIMO RESULTED THESE IMAGES TO BE 2+ POSITIVE WHEN THEY WERE CLEARLY NEGATIVE. THE REPORTED PROBLEM IS A KNOWN ANOMALY OF THE INSTRUMENT. IT'S OCCURENCE IS VERY RARE AND THE DETECTABILITY ON VISUAL EXAMINATION (EVEN WITHOUT FLAGS) IS EXTREMELY HIGH SINCE THE BACKGROUND COLOR REMAINS WHITE INSTEAD OF ADAPTING THE COOMBS COLOR. THUS THE RESIDUAL RISK IS VERY LOW AND TENABLE. THE PROBLEM HAS BEEN ESCALATED TO THE TANGO OPTIMO'S MANUFACTURER FOR ROOT CAUSE ANALYSIS AS PART OF OUR THIRD LEVEL SUPPORT. NO RESULTS ARE AVAILABLE YET.><THIS IS OUR FINAL REPORT ON THIS INCIDENT.><THE CUSTOMER REPORTED THAT NO AHG WAS ADDED TO SOME OF THE TEST PLATE'S WELLS WHEN TESTING ON TANGO OPTIMO. THE CUSTOMER NOTICED THAT THE AHG WAS NOT ADDED ON SOME OF THE SAMPLES BECAUSE THE BACKGROUND ON THE IMAGES REMAINED WHITE. THE CUSTOMER ALSO MENTIONED THE TANGO OPTIMO RESULTED THESE IMAGES TO BE 2+ POSITIVE WHEN THEY WERE CLEARLY NEGATIVE. THE REPORTED PROBLEM IS A KNOWN ANOMALY OF THE INSTRUMENT. IT'S OCCURANCE IS VERY RARE AND THE DETECTABILITY ON VISUAL EXAMINATION (EVEN WITHOUT FLAGS) IS EXTREMELY HIGH SINCE THE BACKGROUND COLOR REMAINS WHITE INSTEAD OF ADAPTING THE COOMBS COLOR. THUS THE RESIDUAL RISK IS VERY LOW AND TENABLE. AFTER EXHAUSTIVE ASSESSMENT OF THE PROBLEM WE ARE STILL NOT IN THE POSITION TO PIN DOWN THE ROOTCAUSE OF THE REPORTED PROBLEM. HOWEVER THE POTENTIAL RISK IS VERY LOW. THE TANGO OPTIMO OFFERS ONLY A PROPOSAL OF AN RESULT INTERPRETATION. EACH TEST RESULT MUST BE VALIDATED BY AN EXPERIENCED OPERATOR. THUS THIS UNUSUAL TEST RESULT COULD NOT PASS UNDETECTED. DIFFERENT FROM OUR INITIAL ASSESSMENT, THIS CASE DOES NOT MEET THE CRITERIA FOR AN MDR.><THIS IS OUR INITIAL REPORT ON THIS INCIDENT. | 0
3750067 | THE CUSTOMER REPORTED DISCREPANT RESULTS WHEN TESTING ON TANGO OPTIMO. FOR SEVERAL PATIENT SAMPLES ANALYZED, THE INSTRUMENT COMBINED THE RESULTS OF THE DAY, BUT ALSO THE RESULTS OF WEEK EARLIER, OF A DIFFERENT PATIENT THAT HAD HAD THE SAME SAMPLE BAR CODE. A REVIEW OF THE INSTRUMENT SPECIFIC DATA CONFIRMED THE CUSTOMER'S COMPLAINT. THE CUSTOMER'S PROCEDURE IS NOT TO USE UNIQUE BARCODE NUMBERS FOR EACH TESTING BUT TO RE-USE THEM AFTER APPROX 30 DAYS. ADDITIONALLY, THE CUSTOMER DOES KEEP VALIDATED RESULTS ON THE DAILY JOURNAL FOR APPROX. ONE MONTH. THE DAILY JOURNAL IS, A IT'S NAME SUGGESTS, INTENDED FOR A SHORT TIME USE. THE REGULAR PROCEDURE WOULD BE TO DELETE VALIDATED TEST RESULTS FROM THE DAILY JOURNAL AFTER THEY ARE EITHER PRINTED OR TRANSFERRED TO THE LAB INFORMATION SYSTEM. IN THIS SPECIFIC CASE, VALIDATED TEST RESULTS ARE KEPT ON THE DAILY JOURNAL UNTIL AN IDENTICAL BARCODE REAPPEARS. RETURNING BARCODES ARE USED ON SAMPLE TUBES CONTAINING MATERIAL OF DIFFERENT PATIENTS AS BEFORE. AS A RESULT, PREVIOUS RESULTS AND CURRENT RESULTS (DIFFERENT PATIENTS) ARE COMBINED AND STORED ON THE DAILY JOURNAL WITH THE NAME OF THE PATIENT WHO WAS TESTED LAST. A THIRD-LEVEL-INVESTIGATION IS STILL ONGOING TO FIND OUT WHETHER ALL THE INSTRUMENT'S SAFETY MEASURES WERE WORKING AND THIS INCIDENT TRUELY IS, AS IT APPEARS, CLEARLY AN USER/HANDLING ERROR.><THIS IS OUR INITIAL REPORT ON THIS INCIDENT.><THIS IS OUR FINAL REPORT ON THIS INCIDENT.><THE CUSTOMER REPORTED DISCREPANT RESULTS WHEN TESTING ON TANGO OPTIMO. FOR SEVERAL PATIENT SAMPLES ANALYZED, THE INSTRUMENT COMBINED THE RESULTS OF THE DAY, BUT ALSO THE RESULTS OF WEEK EARLIER, OF A DIFFERENT PATIENT THAT HAD THE SAME SAMPLE BAR CODE. A REVIEW OF THE INSTRUMENT SPECIFIC DATA CONFIRMED THE CUSTOMER'S COMPLAINT. THE CUSTOMER'S PROCEDURE IS NOT TO USE UNIQUE BARCODE NUMBERS FOR EACH TESTING BUT TO RE-USE THEM AFTER APPROX 30 DAYS. ADDITIONALLY, THE CUSTOMER DOES KEEP VALIDATED RESULTS ON THE DAILY JOURNAL FOR APPROX. ONE MONTH. THE DAILY JOURNAL IS, A IT'S NAME SUGGESTS, INTENDED FOR A SHORT TIME USE. THE REGULAR PROCEDURE WOULD BE TO DELETE VALIDATED TEST RESULTS FROM THE DAILY JOURNAL AFTER THEY ARE EITHER PRINTED OR TRANSFERRED TO THE LAB INFORMATION SYSTEM. IN THIS SPECIFIC CASE, VALIDATED TEST RESULTS ARE KEPT ON THE DAILY JOURNAL UNTIL AN IDENTICAL BARCODE REAPPEARS. RETURNING BARCODES ARE USED ON SAMPLE TUBES CONTAINING MATERIAL OF DIFFERENT PATIENTS AS BEFORE. AS A RESULT, PREVIOUS RESULTS AND CURRENT RESULTS (DIFFERENT PATIENTS) ARE COMBINED AND STORED ON THE DAILY JOURNAL WITH THE NAME OF THE PATIENT WHO WAS TESTED LAST. THE REPORTED PROBLEM COULD BE REPRODUCED BY OUR INTERNATIONAL TECHNICAL SUPPORT BY APPLYING AN INVALID WORKFLOW (DELETING AND RECOVERING THE VERY SAME VALIDATE.PHB). INQUIRIES WITH THE CUSTOMER CONFIRMED THAT THE CUSTOMER APPLIES THE SAME WORKFLOW THAT WAS USED BEFORE AND LEAD TO POSITIVE REPRODUCTION OF THE REPORTED PROBLEM. THE IDENTIFIED ROOT CAUSE FOR THIS REPORTED PROBLEM IS AN ILLEGITIMATE MANIPULATION OUTSIDE THE TANGO OPTIMO APPLICATION SOFTWARE. ALL DATABASES OF THE SYSTEM BELONG TOGETHER AS A UNITY, HENCE IT IS NOT ALLOWED TO DELETE OR REPLACE SINGLE DATABASES. DELETING THE VALIDATE.PHB WHEN THERE ARE STILL RESULTS IN THE DAILY JOURNAL WILL DISCONNECT THE PATIENT DATA FROM THE CORRESPONDING RESULT IN THE DAILY JOURNAL AND ENABLES OVERWRITING IF AN IDENTICAL SAMPLE BARCODE IS USED FOR A DIFFERENT INDIVIDUAL. | 0
4348217 | ><((B)(4)). INOVENT (B)(4) IS UNDER INVESTIGATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITHIN 30 DAYS OF COMPLETION OF INVESTIGATION. | 0
3442410 | CHILDREN'S MEDICAL VENTURES (CHMV) RECEIVED A REPORT FROM A (B)(4) SUPPLIER STATING THAT A SMART MONITOR INFANT APNEA MONITOR HAD FAILED THE DIAGNOSTIC TEST. NO RESULTING EFFECT ON A PATIENT OR POTENTIAL USER HAS BEEN REPORTED. THE DEVICE WAS REPORTEDLY NOT IN USE AT THE TIME OF THE REPORTED EVENT. HOWEVER, IT IS NOT KNOWN IF THE DEVICE ALARMED APPROPRIATELY DURING THE SELF-TEST PROCESS. THE DME HAS BEEN CONTACTED FOR ADDITIONAL INFORMATION AND STATED THAT THEY WERE NOT AWARE OF A COMPLIANT ALLEGATION ASSOCIATED WITH THE DEVICE. ADDITIONALLY, THE (B)(4) STATED THAT THE DEVICES LOCATION WAS NOT CURRENTLY KNOWN. NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. THE ALLEGED FAILURE IS BEING REPORTED DUE TO THE FACT THAT IS UNKNOWN WHETHER OR NOT THE UNIT ALARMED APPROPRIATELY FOR THE REPORTED EVENT.><(B)(4). THE COMPLAINT ISSUE ALLEGED BY THE CUSTOMER WAS NOT ABLE TO BE CONFIRMED BECAUSE THE DEVICE HAS NOT BEEN RETURNED TO THE MFR FOR EVALUATION. THE PURPOSE OF THE SMARTMONITOR 2 FUNCTIONAL SELF-TEST IS TO CHECK THAT ALL THE FEATURES OF THE UNIT ARE FUNCTIONING PROPERLY. A FUNCTIONAL SELF-TEST SHOULD BE PERFORMED AT LEAST ONCE A WEEK OR ACCORDING TO THE INSTRUCTIONS GIVE BY THE HEALTH CARE PROFESSIONAL YOU SHOULD ALSO PERFORM THE TEST: AFTER A LEAD WIRE IS CHANGED. AFTER THE PATIENT CABLE IS CHANGED. THE SMARTMONITOR 2 (SM2) IS INTENDED FOR USE IN CONTINUOUS MONITORING OF HEART BEAT AND RESPIRATION OF INFANT PATIENTS IN A HOME, HOSPITAL, OR PORTABLE ENVIRONMENT. SM2 RECEIVES PHYSIOLOGICAL SIGNALS VIA TRANSDUCERS ATTACHED TO THE PATIENT AND DIRECTLY CONNECTED TO THE MONITOR, OR FROM DEVICES CONNECTED TO THE AUXILIARY INPUTS OF THE MONITOR. PHYSIOLOGICAL SIGNALS (PATIENT EVENTS) AND EQUIPMENT EVENT ARE RECORDED IN NONVOLATILE MEMORY FOR SUBSEQUENT REVIEW AND ANALYSIS BY A HEALTH CARE PROFESSIONAL. DURING MONITORING, WHEN THE PATIENT'S BREATHING EFFORT AND HEART ACTIVITY ARE NOT WITHIN THE SET BOUNDARIES, AND INDICATOR LIGHT COMES ON AND AN ALARM SOUNDS. PATIENT ALARM LIMITS ARE SET BY THE HEALTH CARE PROFESSIONAL BEFORE THE SM2 IS DELIVERED TO THE PATIENT. BASED ON A COMPLETE REVIEW OF THE COMPLAINT ALLEGATION, CHMV HAS DETERMINED THAT NO FURTHER INVESTIGATION ACTIVITY INTO THE REPORTED ISSUE IS POSSIBLE AT THIS TIME. HOWEVER, IF THE DEVICE IS RETURNED FOR INVESTIGATION AND IT IS CONCLUDED THAT A MALFUNCTION OUTSIDE OF THE DESIGN SPECIFICATIONS TOOK PLACE, A FOLLOW-UP, ADDITIONAL INFORMATION REPORT WILL BE FILED. | 0
4135241 | STENT WAS HAVING C-SECTION UNDER GENERAL ANESTHESIA. AT ANESTHESIA WORKSTATION MACHINE, THE ETCO2 TUBING WAS CAUGHT UNDER THE APL VALVE WHICH WAS NOT OBVIOUS VISUALLY, NOR WAS IT OBVIOUS WHEN CLOSING THE APL VALVE. THE CRNA WAS UNABLE TO VENTILATE A PT BECAUSE SHE COULD NOT HOLD PRESSURE (THE BAG DEFLATED EVEN WITH THE APL VALVE COMPLETELY CLOSED). THE CRNA WAS UNABLE TO GIVE POSITIVE PRESSURE VENTILATION WITH VALVE COMPLETELY CLOSED. EVENT OCCURRED WITH INDUCTION WITHOUT ANY HARM TO THE PT. THE MFR WAS MADE AWARE OF THE PROBLEM BY THE HOSPITAL REP IMMEDIATELY AFTER THE INCIDENT. IT WAS REVEALED FROM THE MFR THAT THE DESIGN ISSUE WITH THE ETCO2 GETTING CAUGHT UNDER THE APL VALVE WAS FIRST RECOGNIZED BY THE MFR IN AUGUST 2008 AND THERE IS LITERATURE AVAILABLE ABOUT IT. THE APL VALVE HAS BEEN REDESIGNED ON THE NEWER MACHINES AND THE PROBLEM CORRECTED. THE MFR REP REPORTED THE FIX FOR THE PROBLEM WITH OLDER MODELS LIKE THE HOSPITAL'S (2006 MACHINES) IS TO SECURE THE ETCO2 TUBING AND ROUTING IT OFF THE SIDE OF THE MACHINE AWAY FROM THE APL VALVE WHERE IT DOES NOT INTERFERE AT ALL WITH THE CIRCUIT OR PT CARE. | 0
3934656 | ><THE EVENT OCCURRED ON (B)(6) 2014 WHEN THE PATIENT UNDERWENT A SURGICAL PROCEDURE TO REMOVE A FRAGMENT OF HIS OWN LENS THAT WAS STILL IN PLACE FOLLOWING INITIAL CATARACT SURGERY ON (B)(6) 2014. DURING THIS PROCEDURE, WHILE THE CORNEAL WOUND WAS BEING IRRIGATED WITH STERILE SALINE, THE ANTERIOR CHAMBER CANNULA TIP ON A 3ML LUER LOCK SYRINGE UNEXPECTEDLY CAME OFF AND WENT INTO THE ANTERIOR CHAMBER, THROUGH THE IRIS, AND THROUGH THE SCLERA. THE CANNULA WAS REMOVED AND ACCOUNTED FOR. THERE WAS HEMORRHAGE FROM THE IRIS NOTED IN THE ANTERIOR CHAMBER. THIS REQUIRED THE CONJUNCTIVA TO BE OPENED AND LIGHT CAUTERY FOR HEMOSTASIS WAS PERFORMED. | 0
3927166 | GE CARESCAPE MONITOR - INNOVIAN ANESTH RECORDED FAILED TO RECORD VITAL SIGNS MID-CASE FOR APPROXIMATELY 15 MINUTES - I WAS FOCUSED ON GE TREND VIEW (NUMERIC AND GRAPHIC) AND DID NOT IMMEDIATELY NOTICE THAT RECORDER (OR ITS INTERFACE) WAS NOT WORKING PROPERLY. I HAVE PHOTOS FROM GE MONITOR DOCUMENTING THAT VITAL SIGNS WERE BEING RECORDED AND WERE TRENDING. I HAVE A PHOTO DOCUMENTING NO DATA IN THE ROLLBACK BUFFER." I HAVE A PHOTO DOCUMENTING THE TIME OF THE GE MONITOR WAS THE CORRECT TIME | 0
4100444 | THIS SUPPLEMENTAL REPORT IS SUBMITTED IN ORDER TO CORRECT THE INITIAL REPORT FROM A REPORTABLE SERIOUS INJURY TO A NON-REPORTABLE EVENT. EVALUATION OF THIS PRODUCT EVENT INDICATES THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY. INFORMATION RECEIVED FROM THE FIELD INDICATES THAT THERE WAS NO PATIENT HARM AS A RESULT OF THIS PRODUCT EVENT. THE PATIENT DIED 12 DAYS FOLLOWING THE EVENT AND THE SURGEON INVOLVED WITH THE CASE STATED THAT THE PATIENT DIED DUE TO A RESPIRATORY INFECTION AND THE DEATH WAS NOT RELATED TO THIS PRODUCT EVENT. ADDITIONALLY, EVALUATION OF THE DEVICE EVENT DATA SHOWED THAT THE DEVICE PERFORMED AS EXPECTED, INDICATING NO MALFUNCTION. ADDITIONALLY, A REVIEW OF PRIOR EVENTS INDICATES THAT, FOR THIS PRODUCT OR SIMILAR PRODUCT, THE PRODUCT EVENT IS NOT SUSPECTED TO HAVE BEEN A CAUSE, OR CONTRIBUTORY CAUSE, OF ANY PREVIOUS INCIDENTS THAT HAVE LED TO DEATH OR SERIOUS INJURY. BASED ON THE INFORMATION ABOVE, MEDTRONIC CONCLUDED THAT THIS DEVICE DID NOT CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY AND THE LIKELIHOOD OF THE EVENT CAUSING OR CONTRIBUTING TO A DEATH OR SERIOUS INJURY IF IT WERE TO RECUR IS REMOTE. (B)(4).><MEDTRONIC RECEIVED INFORMATION REPORTING THAT DURING USE OF THIS HMS PLUS IN A PROCEDURE, THE COAGULATION TEST RESULTS DID NOT MEET THE USER'S EXPECTATIONS DUE TO THE APPEARANCE OF CLOTS IN THE PERICARDIUM. THE ACT (ACTIVATED CLOTTING TIME) RESULTS WERE THOUGHT TO BE INACCURATE AND UNRELIABLE. USE OF THE INSTRUMENT WAS CONTINUED WITH NO RESULTING PATIENT IMPACT. IT WAS REPORTED THAT THE PATIENT DIED 12 DAYS LATER FROM RESPIRATORY COMPLICATIONS WHICH WERE NOT LINKED WITH THIS EVENT. THE INSTRUMENT WAS INSPECTED BY MEDTRONIC FRANCE SERVICE AND REMAINS IN USE. ADDITIONAL INFORMATION HAS BEEN REQUESTED. THE DETAILS OF THE EVENT WERE SUMMARIZED AS FOLLOWS: O T=0 NORMAL COAGULATION INJECTION OF 15000 UI HEPARIN O 16MIN ACT =489SEC O 39 MIN ACT=999 SEC O 80 MIN: APPEARANCE OF CLOTS IN THE PERICARDIUM, THE OXYGENATOR (NON-MEDTRONIC) WAS REMOVED DUE TO CLOTS IN THE DEVICE THAT RESTRICTED THE DEVICE. O 85 MIN ACT=335 SEC O INJECTION OF 10000UI HEPARIN O ANTI-COAGULATION WAS REVERSED AT THE END OF THE PROCEDURE BY PROTAMINE WITHOUT ANY PROBLEM. O PATIENT DIED 12 DAYS LATER FROM RESPIRATORY COMPLICATIONS NOT LINKED WITH THE EVENT O PROBLEM: THE RESULTS GIVEN BY THE INSTRUMENT SEEMED TO NOT BE RELIABLE TO THE CUSTOMER. O PATIENT HISTORY: AORTIC DISEASE, ATRIAL FIBRILLATION, TEMPORARY ISCHEMIA, HYPERCHOLESTEROLEMIA, AORTIC VALVE BIOPROSTHESIS PROCEDURE.><.><EVENT INVESTIGATION REMAINS IN PROGRESS AND THE INSTRUMENT REMAINS WITH THE CUSTOMER IN THE FIELD. HOWEVER, PRELIMINARY ANALYSIS PERFORMED IN THE FIELD SHOWED THAT THE DEVICE PERFORMED AS INTENDED. THERE WERE NO PRODUCT ISSUES OR MALFUNCTIONS DETECTED THAT WOULD HAVE IMPACTED THIS EVENT. DEVICE INFORMATION HAS BEEN REQUESTED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE FILED. (B)(4). | 0

## 2015
5048119 | ATTEMPTS HAVE BEEN MADE TO RETRIEVE ADDITIONAL INFORMATION ABOUT THE EVENT AND DEVICE. THE ADDITIONAL INFORMATION WILL REPORTEDLY BE FORWARDED TO DEPUY SYNTHES MITEK HOWEVER IT IS NOT KNOWN IF IT WILL BE RECEIVED WITHIN THE 30 DAY REPORTING REQUIREMENT, THEREFORE DEPUY SYNTHES MITEK WOULD LIKE TO FILE THIS INITIAL MEDWATCH REPORT AT THIS TIME. WHEN AND IF ADDITIONAL INFORMATION IS RECEIVED IT WILL BE REFLECTED IN A FOLLOW-UP MEDWATCH REPORT.><MULTIPLE ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE, HOWEVER THE COMPLAINT DEVICE HAS NOT BEEN RETURNED, THEREFORE IS UNAVAILABLE FOR A PHYSICAL EVALUATION. WE CANNOT DISCERN A ROOT CAUSE FOR THE REPORTED FAILURE MODE, HOWEVER THE DEVICE IS APPROXIMATELY NINE YEARS OLD AND A POSSIBLE ROOT CAUSE COULD BE FAIR WEAR AND TEAR. AT THIS POINT IN TIME, NO CORRECTIVE ACTION IS REQUIRED AND NO FURTHER ACTION IS WARRANTED. HOWEVER, DEPUY SYNTHES MITEK WILL CONTINUE TO TRACK ANY RELATED COMPLAINTS WITHIN THIS DEVICE FAMILY AS A MEANS OF MONITORING THE EXTENT WITH WHICH THIS COMPLAINT IS OBSERVED IN THE FIELD. SHOULD THE COMPLAINT DEVICE EVER BE RECEIVED BACK IN THE FUTURE, THIS FILE WILL BE REOPENED AT THAT TIME AND AN EVALUATION WILL BE PERFORMED AND DOCUMENTED.><IT WAS REPORTED BY THE SALES REP THAT THE PUMP IS BROKEN. IT WAS UNKNOWN HOW THE PROCEDURE WAS COMPLETED. NO PATIENT CONSEQUENCE. DEVICE WILL BE RETURNED AND REPLACED PENDING ADDITIONAL INFORMATION. THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED VIA EMAIL FROM OUR SALES REP ON (B)(6) 2015; THE PROCEDURE WAS EXTENDED OVER THIRTY MINUTES DUE TO THE ISSUES WITH THE FMS PUMP. | 0
5207516 | (B)(4). EVENT OCCURRED IN (B)(6). REPORTING IN THE U.S DUE TO SIMILAR PRODUCT SOLD IN THE U.S. THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. THE MANUFACTURER REVIEWED THE DEVICE HISTORY RECORD, QA TEST DATA, AND RISK ANALYSIS, COMPLAINT HISTORY FOR THE MODEL NUMBER AND SIMILAR MODELS. REVIEW OF THE QA INSPECTION RESULT OF THE UNIT SHOWED THE UNIT HAD NO PROBLEM. THERE WAS NO SIMILAR CASE REPORTED IN THE PAST. THE DEVICE MANUFACTURER PERFORMED AN INVESTIGATION BY EVALUATING RETAINED SAMPLES. THE DEVICE MANUFACTURER EXAMINED THE IGNITION CAPABILITY OF AC ADAPTER FOR FOLLOWING EVENTS: COMPONENT ON SUBSTRATE OR CORD CAUGHT FIRE DUE TO OVERLOAD TO AC ADAPTER; LIGHTING SURGE BROKE INNER COMPONENT AND IT CAUGHT FIRE; IGNITED DUE TO TRACKING PHENOMENON. AS FOR 1, PUT CURRENT LOAD OF 650- 700MA JUST BEFORE VOLTAGE SUPPLY IS SHUT DOWN AT PROTECTION CIRCUIT AND CHECKED TEMPERATURE INCREASE DURING OPERATION OF FOR 30 MINUTES. RECTIFIER DIODE WITH HIGHEST TEMPERATURE ELEVATION WAS SATURATED AT ABOUT 80 DEGREES AND IT WAS NOT LED TO BURNOUT OR IGNITION. (RATING OF THIS AC ADAPTER: 500MA, MAXIMUM LOAD AT ORDINARY USE: 300MA) APPLIED 5A OF ELECTRICAL CURRENT TO AC ADAPTER CORD AND CHECKED TEMPERATURE INCREASE, FOUND THAT IT WAS SATURATED AT ABOUT 60.7 DEGREES, MELT OR IGNITION DID NOT OCCUR. THEREFORE, IT'S JUDGED THAT NO IGNITION HAPPENED DUE TO EVENT 1. AS FOR 2, BROKE AC ADAPTER WITH LIGHTING SURGE, CHECKED SUBSTRATE AND FOUND OUT OPEN FAILURE OF FUSE AND ELECTROLYTIC CAPACITOR WITHOUT BURNOUT. THEREFORE, IT IS JUDGED THAT NO IGNITION HAPPEN DUE TO THE EVENT 2. AS FOR 3, THERE IS NO DESIGN PROBLEM ABOUT TRACKING RESISTANCE SINCE THIS AC ADAPTER IS UL-APPROVED. THEREFORE, IT IS JUDGED THAT IS HAS A LOW POTENTIAL FOR IGNITION DUE TO EVENT 3. THE INVESTIGATION OF THE RETAINED SAMPLES; CONCLUDED THAT RISK ANALYSIS HAS BEEN IMPLEMENTED BECAUSE OF THE FOLLOWING 2 POINTS: NO BURNOUT OR IGNITION OF AC ADAPTER WAS FOUND AT ABOVE LIGHTING SURGE TEST; FROM THE TEST RESULT, WHEN HIGH VOLTAGE APPLIED TO BLOOD PRESSURE MONITOR UNIT (BPM) BURNOUT OF REGULATOR OCCURS AND BPM STOPS OPERATION. IT DOES NOT DAMAGE HUMAN BODY OR PROPERTY, AND IT IS UNLIKELY TO IGNITE. BASED ON THE RISK MANAGEMENT FILES, RISK OF HARM FROM SIMILAR INCIDENTS IS REDUCED TO ACCEPTABLE LEVEL. NO FURTHER INVESTIGATION REQUIRED.><OMRON HEALTHCARE, INC. RECEIVED A LETTER REGARDING A CLAIM FROM (B)(6) ON 10/28/2015. IN THE LETTER, IT WAS REPORTED THAT THE INSURED SUSTAINED SEVERE DAMAGES TO THEIR HOME AS A RESULT OF THE FIRE. THE ENGINEER THAT WAS CALLED IN TO INVESTIGATE THE CAUSE OF FIRE IDENTIFIED THE CAUSE OF THE FIRE AS BEING THE OMRON BLOOD PRESSURE MONITOR WHICH WAS THE ONLY ITEM PLUGGED INTO THE ELECTRICAL OUTLET IN THE MASTER BEDROOM IN THE AREA OF THE ORIGIN OF THE FIRE. THE ENGINEER CONCLUDED: BASED UPON OUR INVESTIGATION TO DATE, WE HAVE FORMED THE FOLLOWING OPINIONS: THE AREA OF FIRE ORIGIN WAS AT THE OMRON BLOOD PRESSURE MONITOR LOCATED IN THE AREA OF FIRE ORIGIN. THE PROBABLE SOURCE OF HEAT FOR IGNITION WAS AN ELECTRICAL FAILURE OF THE POWER CORD FOR THE OMRON BLOOD PRESSURE MONITOR. ADDITIONAL EXAMINATION IS SUGGESTED TO VERIFY. THE PROBABLE FIRST FUEL IGNITED WAS THE POWER CORD INSULATION AND COMBUSTIBLES ON THE WOODEN DESK INCLUDING PAPER/DOCUMENTS. THE FIRE IS CLASSIFIED AS ACCIDENTAL AT THIS TIME BASED ON THE REMAINING PHYSICAL EVIDENCE AND REPORTED INFORMATION. NO OTHER COMPETENT SOURCES OF HEAT FOR IGNITION WERE IDENTIFIED IN THE AREA OF FIRE ORIGIN. ADDITIONAL DESTRUCTIVE EXAMINATION OF THE UNIT IS AN OPTION AT YOUR DISCRETION AFTER THE MANUFACTURER IS PLACED ON NOTICE. AT THIS TIME, OUR CLAIM IS STILL ONGOING. ADDITIONAL COMMUNICATION FROM THE (B)(6) REPRESENTATIVE (DATED 11/02/2015) STATED THEY ARE UNABLE TO PROVIDE COMPLETE ENGINEERING REPORT AT THIS TIME. AND, THEY WERE NOT IN A POSITION TO FORWARD THE EVIDENCE TO OMRON. THEY ALSO STATED, IF OMRON WISHES TO EXAMINE THE EVIDENCE, OMRON MAY HAVE OUR EXPERT CONTACT THEIR ENGINEER TO MAKE ARRANGEMENTS FOR JOINT EXAMINATION.><(B)(4). EVENT OCCURRED IN (B)(6). REPORTING IN THE U.S. DUE TO SIMILAR PRODUCT SOLD IN THE U.S. A REQUEST TO RETRIEVE THE UNIT WAS MADE ON 10/29/2015. THE REQUEST WAS DECLINED BY INSURANCE CLAIMS REPRESENTATIVE. THE BP MONITOR IS RETAINED UNDER CHAIN OF CUSTODY UNTIL FURTHER NOTICE DUE TO INSURANCE CLAIM. THE U.S IMPORTER IS REQUESTING MANUFACTURE OF THE DEVICE TO FURTHER INVESTIGATE THIS INCIDENT. | 0
5098939 | THE CASE STATES THAT A DSD -201 AUTOMATED ENDOSCOPE REPROCESSOR DID NOT HAVE THE SENSORS ENABLED WHEN INSTALLED BY A THIRD PARTY COMPANY. THERE IS POTENTIAL THAT HIGH-LEVEL DISINFECTION OF THE ENDOSCOPES WAS NOT ACHIEVED. POTENTIAL PATIENT CROSS-CONTAMINATION. RA SPOKE WITH (B)(64 FSE CONCERNING THE DISABLED SENSORS IN THE DSD-201 SINCE BEING INSTALLED IN (B)(6) 2014. IT WAS REPORTED THAT THE FACILITY HAD PURCHASED THE MACHINE FROM A THIRD PARTY. THE MACHINE HAD BEEN MOVED FROM A FACILITY IN (B)(6) AND INSTALLED AT THIS FACILITY IN (B)(6). THE THIRD PARTY DID NOT ENABLE THE SENSORS WHEN INSTALLING THE MACHINE. THIS WAS DISCOVERED WHEN THE FACILITY CALLED (B)(4) TECH SERVICE TO REQUEST A SERVICE CALL. WITHOUT TEMPERATURE OR OTHER SENSORS ENABLED, IT IS UNKNOWN WHETHER OR NOT THE ENDOSCOPES BEING REPROCESSED WERE HIGH LEVEL DISINFECTED. IF A 3RD PARTY HAS INVOLVEMENT WITH A (B)(4) PRODUCT, THE WARRANTY AND PRODUCT GUARANTEE ARE VOID. (B)(4) CAN NO LONGER VERIFY PARTS AND SERVICES WHEN IT IS UNKNOWN WHAT WORK OR PART REPLACEMENTS THE THIRD PARTY HAD PERFORMED. DURING THE FSE VISIT, THE SENSORS WERE ENABLED AND THE MACHINE WAS OPERATIONAL. THERE HAVE BEEN NO REPORTS OF PATIENT ILLNESS OR INJURY. THIS INCIDENT WILL BE DOCUMENTED IN (B)(4) COMPLAINT SYSTEM.><THE CASE STATES THAT A DSD -201 AUTOMATED ENDOSCOPE REPROCESSOR DID NOT HAVE THE SENSORS ENABLED WHEN INSTALLED BY A THIRD PARTY VENDOR. THERE IS POTENTIAL THAT HIGH-LEVEL DISINFECTION OF THE ENDOSCOPES WAS NOT ACHIEVED. POTENTIAL PATIENT CROSS-CONTAMINATION. | 0
5117762 | THE CASE STATES THAT THIS FACILITY HAD 2 DSD -201 AUTOMATED ENDOSCOPE REPROCESSORS. THE MACHINES' PROGRAMMED DISINFECTANT CYCLE CONTACT TIME AND TEMPERATURE SETTINGS WERE NOT INLINE WITH THE INSTRUCTIONS FOR METRICIDE OPA HIGH LEVEL DISINFECTANT BEING USED. THERE IS POTENTIAL THAT HIGH-LEVEL DISINFECTION OF THE ENDOSCOPES WAS NOT ACHIEVED, THUS POTENTIAL PATIENT CROSS CONTAMINATION.><THE CASE STATES THAT THIS FACILITY HAD 2 DSD -201 AUTOMATED ENDOSCOPE REPROCESSORS. THE MACHINES' PROGRAMMED DISINFECTANT CYCLE CONTACT TIME AND TEMPERATURE SETTINGS WERE NOT INLINE WITH THE INSTRUCTIONS FOR USE ON METRICIDE OPA HIGH LEVEL DISINFECTANT (HLD) BEING USED. THERE IS POTENTIAL THAT HIGH-LEVEL DISINFECTION OF THE ENDOSCOPES WAS NOT ACHIEVED, THUS POTENTIAL PATIENT CROSS CONTAMINATION. MEDIVATORS FIELD SERVICE TEAM AND SALES REPRESENTATIVE ARE IN CLOSE COMMUNICATION WITH THIS CUSTOMER. IT WAS REPORTED THAT THIS FACILITY HAS CHANGED THE PROGRAMMING ON SEVERAL OCCASIONS TO SHORTEN THE CYCLE TIME. THE FACILITY WAS EDUCATED ON THE IMPORTANCE OF CYCLE TIME/HEATING CHEMISTRY TO ACHIEVE HLD. THERE HAVE BEEN NO REPORTS OF PATIENT ILLNESS OR INJURY. THIS COMPLAINT WILL CONTINUE TO BE MONITORED WITHIN THE MEDIVATORS COMPLAINT SYSTEM. | 0
4922844 | THE FACILITY REPORTED BLOOD DRIPPING OUT OF THE DISTAL END OF THE ENDOSCOPE AFTER HANGING IT UP TO DRY AFTER REPROCESSING IT IN A MEDIVATORS DSD-201 AUTOMATED ENDOSCOPE REPROCESSOR (AER). IT WAS DETERMINED THE FACILITY WAS USING THE INCORRECT HOOK-UP TO REPROCESS THE ENDOSCOPE.><THE FACILITY REPORTED BLOOD DRIPPING OUT OF THE DISTAL END OF THE ENDOSCOPE AFTER HANGING IT UP TO DRY AFTER REPROCESSING IT IN A MEDIVATORS DSD-201 AUTOMATED ENDOSCOPE REPROCESSOR (AER). AFTER INVESTIGATION, IT WAS DETERMINED THE FACILITY WAS USING THE INCORRECT HOOK-UP TO REPROCESS THE ENDOSCOPE. THE ENDOSCOPE INVOLVED WAS NOT USED ON ANY PATIENTS. USING THE INCORRECT HOOK-UP COULD LEAD TO AN ENDOSCOPE NOT BEING PROPERLY DISINFECTED, WHICH COULD LEAD TO POTENTIAL PATIENT CROSS-CONTAMINATION. IT IS UNKNOWN IF THERE HAVE BEEN OTHER INSTANCES WHERE THE INCORRECT HOOK-UP WAS USED TO REPROCESS SCOPES. TO DATE, THERE HAVE BEEN NO REPORTS OF PATIENT ILLNESS OR INJURY. MEDIVATORS CLINICAL STAFF REPORTEDLY HAS APPROACHED THIS FACILITY REGARDING THE IMPORTANCE OF USING THE CORRECT HOOK-UP TO ENSURE ALL CHANNELS ARE HIGH LEVEL DISINFECTED. THE CLINICAL STAFF IS STILL IN CONTACT WITH THE FACILITY. THIS COMPLAINT WILL CONTINUE TO BE MONITORED AND MAINTAINED BY THE MEDIVATORS COMPLAINT HANDLING SYSTEM. | 0
5029200 | A VISUAL INSPECTION OF THE RETURNED DEVICE NOTED NO OBVIOUS SIGNS OF DAMAGE. THE PROBE DID INITIALLY TRACK BUT THE MESSAGE FOR THE BROKEN SENSOR APPEARED AFTER SHAKING IT SLIGHTLY. FURTHER EVALUATION REVEALED THAT THERE IS A BROKEN COIL WITHIN THE SENSOR. A DEFINITIVE ROOT CAUSE CANNOT BE DETERMINED WITH THE INFORMATION PROVIDED. A REVIEW OF THE MANUFACTURING RECORDS DID NOT REVEAL ANY MANUFACTURING OR MATERIAL ABNORMALITIES THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED INCIDENT. AT THIS TIME, WE HAVE NO REASON TO SUSPECT THAT THE PRODUCT FAILED TO MEET ANY PRODUCT SPECIFICATIONS AT THE TIME OF MANUFACTURE. THIS INVESTIGATION COULD NOT VERIFY OR IDENTIFY ANY EVIDENCE OF PRODUCT CONTRIBUTION TO THE REPORTED PROBLEM. NO FURTHER INVESTIGATION IS WARRANTED AT THIS TIME. (B)(4).><.><IT WAS REPORTED THAT A SURGERY WAS EXTENDED BY ONE HOUR DUE TO PROBE SENSOR BEING BROKEN. | 0
4987798 | THE VIG 2 MONITOR WAS MANUFACTURED JUNE 11, 2013 AND REVIEW OF THE DEVICE HISTORY RECORD SUPPORTS THAT THERE WERE NO NON-CONFORMANCES NOTED FOR ANY REASON. EXAMINATION OF THE RETURNED SUSPECT MONITOR WAS UNABLE TO DETECT ANY FAILURE OF THE MONITOR TO PERFORM AS EXPECTED. OVER 36 HOURS OF TESTING PRODUCED EXPECTED RESULTS. NO PHYSICAL DAMAGE WAS OBSERVED DURING THE VISUAL INSPECTION OF THE MONITOR. THE VIG 2 MONITOR MAY BE REASONABLY EXEMPTED AS A CONTRIBUTOR TO THE CUSTOMER COMPLAINT. IT IS UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE REPORTED EVENT, AS NO FAULT COULD BE ASSIGNED TO ANY OF THE SUSPECT DEVICES AND THERE WAS NO EVIDENCE OF ANY FAILURE OF ANY EDWARDS¿ DEVICE TO FUNCTION APPROPRIATELY. THE MONITOR AND CABLE WILL BE RETURNED TO THE CUSTOMER AND NO FURTHER ACTIONS WILL BE PURSUED AT THIS TIME. WE WILL CONTINUE TO MONITOR FOR TRENDS PER PROCEDURE.><IT WAS REPORTED THAT THE CCO DECREASED FROM 4.0L/MIN TO 1.0L/MIN OVER A PERIOD OF 30 MINUTES. THE EXPECTED VALUE WAS 4.0 - 5.0 L/MINUTE. IT IS UNKNOWN IF AN ALARM OR FAULT MESSAGE WAS OBSERVED. AT THE SAME TIME, MONITOR DISPLAYED OTHER PARAMETERS LIKE ECG AND MAP NORMALLY. ICO AND SVO2 WERE NOT MEASURED. WHEN THE DISPLAYED CCO VALUE DECREASED TO 1.0L/MINUTE, MEASUREMENT HAD STOPPED, SO THERE WERE NO PATIENT COMPLICATIONS REPORTED.><ADDITIONAL MANUFACTURER NARRATIVE: THE DEVICE IS EXPECTED TO BE RETURNED FOR EVALUATION; HOWEVER, IT HAS YET TO BE RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY ONCE THE DEVICE IS EVALUATED. OF NOTE, TWO OTHER REPORTS HAVE BEEN SUBMITTED FOR RELATED DEVICES. PLEASE REFERENCE REPORT #2015691-2015-01967 AND #2015691-2015-01993. | 1
4998650 | THE REPORTED DEVICE IS THE SECOND MONITOR OF TWO MONITORS INVOLVED IN THE REPORTED EVENT - AT THE TIME OF THIS SUBMISSION, THE SERIAL NUMBERS FOR EACH MONITOR IS UNKNOWN. ADDITIONAL INFORMATION IS FORTHCOMING. BOTH MONITORS WILL BE REQUESTED FOR RETURN AND THE EVALUATION OF THE SUSPECT DEVICE(S) IS ANTICIPATED. A FOLLOW-UP SUBMISSION WILL COMMUNICATE THE SERIAL NUMBER, RESULT OF THE DEVICE HISTORY REVIEW, THE RETURNED PRODUCT EVALUATION RESULT(S), INVESTIGATION AND CONCLUSION. PLEASE REFER TO MANUFACTURER¿S REPORT NUMBERS 2015691-2015-02020 WITH REFERENCE TO THE SUSPECT FLOTRAC SENSOR, AND 2014591-2015-02022 WITH REFERENCE TO THE OTHER VIGILEO MONITOR REPORTED FOR THIS EVENT.><THE SERIAL NUMBER FOR THE DEVICE IDENTIFIED IN COMPLAINT (B)(4) WHICH CORRELATES TO THIS REPORT, IS (B)(4). THE MONITOR WAS MANUFACTURED MARCH 22, 2010 AND REVIEW OF THE DEVICE HISTORY RECORD SUPPORTS THAT THERE WERE NO NON-CONFORMANCES NOTED FOR ANY REASON. PER EDWARDS¿ PROCEDURE, THE USEFUL LIFE OF THE MONITOR IS 5 YEARS. THE REFERENCED MONITOR HAS EXCEEDED ITS ESTABLISHED USEFUL LIFE. EVALUATION OF THE RETURNED MONITOR WAS UNABLE TO REPLICATE THE VALUES REPORTED BY THE CUSTOMER. OVER 36 HOURS OF TESTING WAS PERFORMED, WITH NO PERFORMANCE ISSUES NOTED FOR ANY REASON. HOWEVER, DURING EXAMINATION OF THE MONITOR, IT WAS NOTICED THAT A RUBBER FOOT WAS MISSING; THE FOOT WAS REPLACED. THE FOREGOING ISSUE WAS NOT ASSOCIATED TO THE CUSTOMER¿S EXPERIENCE REGARDING THE CI VALUE. AN ADDITIONAL MONITOR REFERENCED IN THE COMPLAINT, (B)(4), SERIAL NUMBER (B)(4) WAS ALSO EXEMPTED AS CONTRIBUTOR TO THE CUSTOMER'S REPORT. THE EVALUATION RESULTS FOR THE SECOND MONITOR CONCLUDED THAT THE MONITOR FUNCTIONED AS EXPECTED (SEE MDR 2015691-02022). COMPLAINTS (B)(4) WERE INITIATED FOR THE ASSOCIATED FLOTRAC SENSORS; HOWEVER, THE SENSORS WERE NOT RECEIVED FOR EVALUATION. WITHOUT RETURN OF THE FLOTRAC SENSORS, WE ARE UNABLE TO PERFORM A COMPLETE INVESTIGATION INTO THE ROOT CAUSE OF THE REPORTED EVENT. EDWARDS WILL CONTINUE TO REVIEW AND MONITOR ALL EVENTS. IF ACTION IS REQUIRED, AN APPROPRIATE INVESTIGATION WILL BE PERFORMED.><IT WAS REPORTED THAT INCORRECT CI VALUE WAS SHOWN ON THE VIGILEO MONITOR DURING USE. AROUND 0.7 TO 1.5L WAS SHOWN, ALTHOUGH THE CUSTOMER WAS EXPECTING VALUES OVER 2.0L. THERE WAS NO PROBLEM ZEROING BEFORE USE. THERE WAS ALSO NO PROBLEM NOTED ON THE SHAPE OF THE PRESSURE WAVEFORM AND THE PRESSURE VALUE AND THE WAVEFORM MATCHED. FLOTRAC UNIT AND MONITOR WAS EXCHANGED BUT THE PROBLEM WAS NOT SOLVED. ERROR MESSAGES WERE NOT THERE WERE NO OCCLUSION, LEAKAGE, NOR KINKS OBSERVED. VIGILEO DATA WAS OBTAINED (SEE ATTACHED).THE PATIENT WAS NOT TREATED BASED ON THE INACCURATE VALUES, AND THERE WERE NO PATIENT COMPLICATIONS REPORTED. ADDITIONAL INFORMATION: ADDITIONAL INFORMATION WAS RECEIVED FROM THE EDWARDS SALES REPRESENTATIVE. INITIALLY, THERE WAS NO COMPLAINT FOR THE MONITORS REPORTED FROM THE SALES REP., THE REPORTED COMPLAINT WAS ONLY FOR FLOTRAC. HOWEVER, FOLLOW-UP WITH THE SALES REP. CONFIRMED THAT THE VIGILEO MONITORS WERE ALSO CONSIDERED SUSPECT. TWO ADDITIONAL COMPLAINTS (-2 & -3) WERE INITIATED TO ADDRESS THE SUSPECT VIGILEO MONITORS: INITIAL MONITOR AND 'CHANGED' MONITOR. | 1
5022146 | A FOLLOW-UP SUBMISSION WILL COMMUNICATE THE RESULTS OF THE DEVICE HISTORY REVIEW, RETURNED PRODUCT EVALUATION RESULT(S), INVESTIGATION, AND CONCLUSION.><REVIEW OF THE DEVICE HISTORY RECORD SUPPORTS THAT THERE WERE NO NON-CONFORMANCES NOTED FOR ANY REASON DURING THE MANUFACTURING OF THE MONITOR. THE MONITOR WAS MANUFACTURED 31-JULY-2001; IT IS 14 YEARS-OLD. PER EDWARDS¿ PROCEDURE, THE USEFUL LIFE OF THE MONITOR IS 5 YEARS. THE REFERENCED MONITOR HAS EXCEEDED ITS USEFUL LIFE. THE CUSTOMER¿S COMPLAINT WAS UNABLE TO BE REPLICATED DURING EXAMINATION OF THE RETURNED DEVICE. HOWEVER, ANOTHER ISSUE WAS DETECTED. NVRAM LOCATED ON THE MOTHER BOARD EXCEED THE ¿FIXED TIME LIMIT FOR EXCHANGE.¿ ADDITIONALLY THE DSVO CONNECTOR AND CO CONNECTOR WERE DAMAGED. THE OBSERVED DAMAGE WAS IDENTIFIED AS DAMAGE CONSISTENT WITH WEAR AND TEAR. AS VIGILANCE MONITORS ARE NO LONGER SUPPORTED AND THE REFERENCED MONITOR HAS SIGNIFICANTLY EXCEEDED ITS USEFUL LIFE, THE MONITOR WILL NOT BE REPAIRED NOR RETURNED TO THE CUSTOMER; THE MONITOR WILL BE SCRAPPED. PLEASE REFERENCE MDR 2015691-2015-2280 WITH REFERENCE TO THE CCO CATHETER INTERFACE CABLE WHERE THE CABLE WAS EXEMPTED AS A CONTRIBUTOR TO THIS EVENT. NO FAULT WAS IDENTIFIED FOR THE CABLE ¿ THE DEVICE PERFORMED AS EXPECTED. IT WAS NOTED THAT THE CABLE WAS ALSO PAST ITS USEFUL LIFE EXPECTANCY. IT'S UNKNOWN WHETHER PROCEDURAL PROCESSES OR A SPECIFIC CIRCUMSTANCE PLAYED A ROLE IN THE CUSTOMER¿S EXPERIENCE AS THERE WAS NO INDICATION OR EVIDENCE TO SUGGEST THAT A MANUFACTURING ISSUE WAS RESPONSIBLE FOR THE CUSTOMER'S EXPERIENCE.><IT WAS REPORTED THAT DURING USE OF THE VIGILANCE MONITOR, THE CCO VALUE WAS ABNORMAL - APPROXIMATELY 2L/MIN HIGHER. HOWEVER, ECG AND MAP WERE AS EXPECTED. NO ERROR MESSAGES WERE REPORTED AND THE CUSTOMER WAS NOT PERFORMING SV02 MEASUREMENT WHEN THE ATYPICAL VALUE WAS OBSERVED. THERE WAS NO REPORT OF PATIENT COMPROMISE OR INAPPROPRIATE TREATMENT, AND NO ADDITIONAL SYSTEM-RELATED DEVICES WERE IDENTIFIED AS SUSPECT. | 1
5235687 | A PEDIATRIC HOSPITAL HAS QUESTIONED A PT GLUCOSE SAMPLE ON OPTI-CCA. THE CUSTOMER REPORTS THE FOLLOWING INFO: THE PT WAS ADMITTED TO AN ADULT HOSPITAL WHERE A GLUCOSE LEVEL > 1000 MG/DL WAS RECOVERED (TYPE OF ANALYZER UNK). AN IV LINE (NORMAL SALINE ONLY) WAS STARTED IN THE EU OF THE PT. ANY OTHER TREATMENT IS UNK. THE PEDIATRIC HOSPITAL SENT A AMBULATORY UNIT TO TRANSFER THE PT. PRIOR TO TRANSPORT, THE AMBULATORY TEAM TOOK A FINGER STICK BLOOD SAMPLE FROM HAND OPPOSITE THE IV. THIS WAS MEASURED ON A BATTERY POWERED, PORTABLE OPTI-CCA AND RETURNED A RESULT OF 44 MG/DL. THE RESULT WAS DISCARDED AS ERRONEOUS AND THE PT WAS TRANSPORTED. UPON ARRIVAL AT THE PEDIATRIC HOSPITAL, TWO ADDITIONAL MEASUREMENTS WERE MADE. AN ORTHO CLINICAL VITROS RECOVERED A RESULT >1300 MG/DL AND A J&J SURE STEP PRO RECOVERED A RESULT >500 MG/DL - RESULTS OUTSIDE OF MEASURABLE RANGE. THERE WAS NO KNOWN ADDITIONAL MEASUREMENTS WITH AN OPTI-CCA AND RESULTS, OTHER THAN THE OPTI-CCA, ARE UNCONFIRMED. COPIES OF ANALYZER PRINTOUTS FROM BOTH HOSPITALS HAVE BEEN REQUESTED IN WRITING. ALSO REQUESTED IN WRITING WAS THE RETURN OF THE OPTI-CCA ANALYZER AND ANY UNUSED CASSETTES FROM THE SAME LOT FOR EVALUATION. MANUFACTURER TESTING: IN-HOUSE SAMPLES OF THE SAME LOT OF OPTI CASSETTES WERE VERIFIED TO REPORT HIGH" ON A GLUCOSE SAMPLE OF >400 MG/DL | 0
4991169 | IT WAS REPORTED THE PATIENT HAD PASSED AWAY. THE CAUSE OF DEATH IS UNKNOWN. | 0
5210292 | GYN SURGEON PERFORMING DAVINCI TOTAL LAPAROSCOPIC HYSTERECTOMY. LONG TIP FORCEPS WIRE BROKE DURING SURGERY. NO INJURY TO PATIENT. INSTRUMENT REMOVED AND REPLACED WITH NEW ONE. THE INSTRUMENT WAS SENT TO INTUITIVE FOR EVALUATION.><. | 0
4549162 | DURING ROBOTIC COLON RESECTION PROCEDURE, BLADE IN CENTER OF VESSEL SEALER BECAME EXPOSED. DAVINCI SYSTEM RECOGNIZED MALFUNCTION AND ALARMED IMMEDIATELY. INSTRUMENT REMOVED AND REPLACED WITH NEW INSTRUMENT. NOT REPROCESSED. NO INJURY TO PATIENT. | 0
4834779 | THE DROP DOWN MENU TO SPECIFY ROUTE OF ADMINISTRATION OF MEDICATIONS IS A SOURCE OF INNUMERABLE ERRORS. THE FIRST OPTION IS IM AND IS FREQUENTLY SELECTED BECAUSE OF ITS LOCATION AND CLOSE RESEMBLANCE TO IV. IT IS A FLAW THAT HAS BEEN CARRIED ON BY THE VENDOR FOR YEARS, ENDANGERING HUNDREDS OF PATIENTS, IF NOT THOUSANDS OVER TIME. | 1
5053962 | IV TUBING WITH A DEFECTIVE RYMED VALVE BROKE OFF WITHIN THE IV LINE. PATIENT LEVOPHED DRIP INFUSING LINE WAS IN DATE, BUT RYMED NOT DATED AND HAVE NO WAY OF KNOWING WHEN PUT IN PLACE. LINE WAS FOUND TO BE LEAKING AND WHEN CHECKED THEY FOUND THEM COMING APART. PATIENT WAS VENTED AND NOT PULLING ON LINES. | 0
4834778 | NURSES SPEND MOST OF THEIR TIME ENTERING DATA ON THE EHR AND FAIL TO ASSESS THE DATA THAT THEY ENTER AND FAIL TO KNOW AND COMMUNICATE DATA THAT IS EXTREME, INCLUDING HYPERTENSION, HYPOGLYCEMIA, HYPOTHERMIA, DIARRHEA, AND MANY OTHERS. THE COMMUNICATION IMPEDIMENTS ARE POTENT DUE TO THE HUMAN FACTOR DEFICIENCIES ASSOCIATED WITH CPOE AND EHR DEVICES. | 1
5090500 | A PATIENT ON A VENTILATOR FROM THE ICU WAS TRANSPORTED TO THE MRI SUITE BY A RESPIRATORY THERAPIST AND A STUDENT NURSE. THE VENTILATOR WAS MRI SAFE, HOWEVER, THE CART THAT THE VENT WAS TRANSPORTED ON WAS NOT. THE VENTILATORS ARE USUALLY TRANSPORTED ON THE FOOT OF THE PATIENT'S BED, BUT IN THIS CASE IT WAS NOT. THE NON SAFE CART WAS INADVERTENTLY ALLOWED TO ENTER THE ZONE 3 AREA. THE ZONE 3 AREA IS CLOSELY MONITORED BY THE MRI STAFF SO THAT NO FERROUS EQUIPMENT ENTERS ZONE 3. THE PATIENT REMAINED IN ZONE 2 AND THERE WAS NO INJURY TO THE PATIENT.><. | 0
5059458 | THE (B)(6) USES THE PERKINELMER SCREENING CENTER MODULE OF SPECIMEN GATE (SGSC) SOFTWARE (A CLASS I MEDICAL DEVICE) TO STORE, RETRIEVE, AND PROCESS THE DATA ASSOCIATED WITH SPECIMEN TESTING INCLUDING BUT NOT LIMITED TO PATIENT DEMOGRAPHICS, TESTS ORDERED, TEST RESULTS, TEST RESULT INTERPRETATION, QUALITY CONTROL RESULTS, AND RESULT CODES (FLAGS) THAT MAY BE ASSOCIATED WITH THE SPECIMEN FROM ITS ENTRY INTO THE LABORATORY WORKFLOW UNTIL PATIENT REPORT IS GENERATED AND RELEASED. WHEN SGSC SOFTWARE IS INSTALLED, CONFIGURABLE SETTINGS ARE PROGRAMMED BASED ON THE CUSTOMER PREFERENCES. FOR INSTANCE, EACH LABORATORY ENTERS THEIR SPECIFIC POPULATION BASED REFERENCE RANGES THAT ARE USED TO EVALUATE IF A PATIENT'S TEST RESULTS ARE NORMAL, BELOW OR ABOVE THE REFERENCE RANGE AND REQUIRE FURTHER ACTION SUCH AS REPEAT TESTING. THIS LABORATORY SPECIFIC LOGIC AUTOMATES THE FLOW OF THE SPECIMEN THROUGH THE LABORATORY. THE CUSTOMER DEFINES THE NOMENCLATURE USED FOR THEIR RESULT CODES (FLAGS), AND WHEN THE SOFTWARE ENCOUNTERS THE PREDEFINED RESULT CODE THE SOFTWARE PROCESSES THE SPECIMEN ACCORDING TO THE LOGIC ASSOCIATED WITH THE RESULT CODE. THE NEWBORN SCREENING LABORATORY RECEIVED 2 SPECIMENS FOR THE PATIENT. BY PROTOCOL, TWO SPECIMENS SHOULD BE SENT FOR ALL BABIES BORN IN (B)(6). THE FIRST SPECIMEN WAS RECEIVED ON (B)(6) 2015. THE INITIAL C0 WAS LOW, TRIGGERING A CONFIRMATION REQUEST. THE CONFIRMATION WAS NOT TESTED BY THE LABORATORY, BUT THE CONFIRMATION RECORD WAS ATTACHED TO A SEPARATE ASSAY THAT THE SPECIMEN WAS NOT PUNCHED TO, CAUSING AN INCOMPLETE RESULT FOR THE CONFIRMATION CUD (CARNITINE UPTAKE DEFECT) DISORDER SO THE LABORATORY WAS UNABLE TO PRINT THE REPORT AND CONTACTED PERKINELMER FOR ASSISTANCE. PERKINELMER CANCELLED AND UNATTACHED THE RECORDS THAT WERE INCORRECTLY ASSOCIATED WITH THE SEPARATE ASSAY. A REVIEW ITEM WAS CREATED FOR THIS SPECIMEN ON (B)(6) 2015 AT 10:45 AM PREVENTING THE REPORT FROM BEING PRINTED. THE LABORATORY CLOSED THE REVIEW ITEM ON 01/08/2015 AT 11:28 AM ALLOWING THE REPORT TO BE PRINTED, DISPLAYING NORMAL FOR THE CUD DISORDER EVEN THOUGH TESTING NEEDED TO BE CONFIRMED. THE SECOND SPECIMEN WAS RECEIVED ON (B)(6) 2015, ANALYZED AND REPORTED AS NORMAL FOR CUD DISORDER ON 1/20/2015. THE PATIENT REPORT RELEASED DID NOT CONTAIN ANY ERRONEOUS CONTENT. INTERNAL REFERENCE NUMBER:(B)(4). | 1
5059702 | THE (B)(6) STATE PUBLIC HEALTH NEWBORN SCREENING LABORATORY USES THE PERKINELMER SCREENING CENTER MODULE OF SPECIMEN GATE (SGSC) SOFTWARE (A CLASS I MEDICAL DEVICE) TO STORE, RETRIEVE, AND PROCESS THE DATA ASSOCIATED WITH SPECIMEN TESTING INCLUDING BUT NOT LIMITED TO PATIENT DEMOGRAPHICS, TESTS ORDERED, TEST RESULTS, TEST RESULT INTERPRETATION, QUALITY CONTROL RESULTS, AND RESULT CODES (FLAGS) THAT MAY BE ASSOCIATED WITH THE SPECIMEN FROM ITS ENTRY INTO THE LABORATORY WORKFLOW UNTIL PATIENT REPORT IS GENERATED AND RELEASED. WHEN SGSC SOFTWARE IS INSTALLED, CONFIGURABLE SETTINGS ARE PROGRAMMED BASED ON THE CUSTOMER PREFERENCES. FOR INSTANCE, EACH LABORATORY ENTERS THEIR SPECIFIC POPULATION BASED REFERENCE RANGES THAT ARE USED TO EVALUATE IF A PATIENT'S TEST RESULTS ARE NORMAL, BELOW OR ABOVE THE REFERENCE RANGE AND REQUIRE FURTHER ACTION SUCH AS REPEAT TESTING. THIS LABORATORY SPECIFIC LOGIC AUTOMATES THE FLOW OF THE SPECIMEN THROUGH THE LABORATORY. THE CUSTOMER DEFINES THE NOMENCLATURE USED FOR THEIR RESULT CODES (FLAGS), AND WHEN THE SOFTWARE ENCOUNTERS THE PREDEFINED RESULT CODE THE SOFTWARE PROCESSES THE SPECIMEN ACCORDING TO THE LOGIC ASSOCIATED WITH THE RESULT CODE. THE NEWBORN SCREENING LABORATORY RECEIVED 2 SPECIMENS FOR THE PATIENT. BY PROTOCOL, TWO SPECIMENS SHOULD BE SENT FOR ALL BABIES BORN IN (B)(6). THE FIRST SPECIMEN WAS RECEIVED ON (B)(6) 2014. THE SPECIMEN WAS ACCESSIONED AS SN" WHICH INDICATES A SECOND SPECIMEN. THE LABORATORY DOES NOT RUN IRT AND HGB ON SECOND SPECIMENS. THE LABORATORY RAN ALL OTHER TESTS BASED ON THEIR PROTOCOL. THE SPECIMEN WAS PROMOTED TO REPORTABLE AFTER ALL TESTS WERE COMPLETED ON (B)(6) 2014 AT 10:30AM. THE LABORATORY MUST HAVE REALIZED THAT THERE WAS NO FIRST SPECIMEN FOR THIS BABY AND PUNCHED HEMOGLOBIN AND IRT ON (B)(6) 2014 AT 11:20AM. SINCE THE SPECIMEN HAD ALREADY BEEN PROMOTED TO REPORTABLE STATUS THE MAILER WAS PRINTED ON (B)(6) 2014 AT 4:29PM WHILE HEMOGLOBIN AND IRT WERE STILL IN PROCESS. THE MAILER SHOWED AS TEST PENDING FOR BOTH HGB AND IRT. THE RESULT FOR HGB WAS ORIGINALLY DETERMINED TO BE FA -NORMAL. THE LABORATORY CHANGED THE INTERPRETATION TO FA+OTHER ON (B)(6) 2014 BUT DID NOT RECALCULATE THE ASSAY PREVENTING THE RESULT CODE FROM BEING RECALCULATED USING THE NEW INTERPRETATION. THE ASSAY WAS ACCEPTED ON (B)(6) 2014 AT 10:29AM WITH A NORMAL RESULT CODE. AFTER ACCEPTANCE THE LABORATORY MANUALLY CHANGED THE RESULT CODE FROM NORMAL TO TO BE CONFIRMED ON (B)(6) 2014 AT 10:36AM IN SCREENING CENTER RESULT EDITOR. SCREENING CENTER DOES NOT CREATE REQUESTS WHEN RESULT CODES ARE CHANGED. SINCE THE SPECIMEN HAD ALREADY BEEN REPORTED | 1
5059703 | (B)(6) USES (B)(4) SOFTWARE (A CLASS I MEDICAL DEVICE) TO STORE, RETRIEVE, AND PROCESS THE DATA ASSOCIATED WITH SPECIMEN TESTING INCLUDING BUT NOT LIMITED TO PATIENT DEMOGRAPHICS, TESTS ORDERED, TEST RESULTS, TEST RESULT INTERPRETATION, QUALITY CONTROL RESULTS, AND RESULT CODES (FLAGS) THAT MAY BE ASSOCIATED WITH THE SPECIMEN FROM ITS ENTRY INTO THE LABORATORY WORKFLOW UNTIL PATIENT REPORT IS GENERATED AND RELEASED. WHEN (B)(4) SOFTWARE IS INSTALLED, CONFIGURABLE SETTINGS ARE PROGRAMMED BASED ON THE CUSTOMER PREFERENCES. FOR INSTANCE, EACH LABORATORY ENTERS THEIR SPECIFIC POPULATION BASED REFERENCE RANGES THAT ARE USED TO EVALUATE IF A PATIENT'S TEST RESULTS ARE NORMAL, BELOW OR ABOVE THE REFERENCE RANGE AND REQUIRE FURTHER ACTION SUCH AS REPEAT TESTING. THIS LABORATORY SPECIFIC LOGIC AUTOMATES THE FLOW OF THE SPECIMEN THROUGH THE LABORATORY. THE CUSTOMER DEFINES THE NOMENCLATURE USED FOR THEIR RESULT CODES (FLAGS), AND WHEN THE SOFTWARE ENCOUNTERS THE PREDEFINED RESULT CODE THE SOFTWARE PROCESSES THE SPECIMEN ACCORDING TO THE LOGIC ASSOCIATED WITH THE RESULT CODE. (B)(6) RECEIVED ONE SPECIMEN FOR THE PATIENT. BY PROTOCOL, TWO SPECIMENS SHOULD BE SENT FOR ALL BABIES BORN IN (B)(6) , BUT WE DO NOT SHOW ANOTHER SPECIMEN OF BEING RECEIVED FOR SAID BABY. THE BABY WAS BORN PRIOR TO THE STATE OF (B)(6) IMPLEMENTING IN-STATE NEWBORN SCREENING WITH THEIR OWN LABORATORY ON (B)(6) 2014. THE FIRST AND SECOND SPECIMENS MAY HAVE BEEN TESTED IN (B)(6). THE FIRST SPECIMEN WAS RECEIVED ON (B)(6) 2014. THE INITIAL CO WAS NORMAL, C16 WAS NORMAL, C18 WAS LOW, AND C0/(C16+C18) WAS ELEVATED MAKING THE OVERALL RESULTS FOR CPT1 DISORDER TO BE CONFIRMED, CREATING A REPEAT REQUEST FOR THE DISORDER. THE LAB CANCELED THE REPEAT REQUEST ON (B)(6) 2014. THE REQUEST WAS CANCELED BY THE LAB BECAUSE THE SPECIMEN WAS SENT BACK TO (B)(6) FOR TESTING COMPLETION. CPT1 WAS REPORTED AS NORMAL ON (B)(6) 2014, BUT THE REPORT WAS PRINTED AGAIN ON (B)(6) 20 14. THE PATIENT REPORT RELEASED DID NOT CONTAIN ANY ERRONEOUS CONTENT. (B)(4). | 1
4713745 | SHOCK WAS THE ADMITTING DIAGNOSIS, COMPLICATED BY HYPOKALEMIA. POTASSIUM SUPPLEMENTS WERE ELECTRONICALLY ORDERED AS 20 MEQ P.O. B.I.D., WITH PARAMETER TO BE HELD IF POTASSIUM WAS ABOVE 4.7 MG PERCENT. DESPITE THE SERU, POTASSIUM BEING OVER THAT THRESHOLD, THE POTASSIUM SUPPLEMENT WAS ADMINISTERED BY THE PHARMACIST AND NURSE FOR 3 DOSES, RESULTING IN A LIFE THREATENING POTASSIUM BLOOD LEVEL. THE PART OF THE ORDER LISTING THE PARAMETER WAS SEVERAL LINES FROM THE MAIN ORDER, AND WAS MISSED BY ALL HEALTHCARE PROFESSIONALS PROVIDING CARE TO THAT PATIENT. THE DESIGN FLAW RESULTED IN THE ERROR. | 1
5158608 | THE CUSTOMER REPORTED THAT THE RNS (REMOTE NETWORK SYSTEM OR REMOTE MONITORING SYSTEM) IS DIM AND CANNOT SEE ALL BEDS" SCREEN.><THE DEVICE WAS RETURNED TO NIHON KOHDEN | 1
5305212 | NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE RNS (REMOTE NETWORK SYSTEM OR REMOTE MONITORING SYSTEM) TURNS OFF ON ITS OWN. THE RNS MUST BE MANUALLY TURNED BACK ON AFTER THE POWER LOSS. THE CUSTOMER HAS TRIED NEW COMMUNICATION CABLES. | 1
5337370 | THE CUSTOMER REPORTED THAT THE LCD ON THE RNS (REMOTE NETWORK SYSTEM OR REMOTE MONITORING SYSTEM) HAS FAILED. THE BACKLIGHT ON THE RNS HAS GONE OUT AND DATA CAN NOT BE SEEN ON THE SCREEN. PATIENTS ARE BEING MONITORED ON A CNS (CENTRAL MONITORING SYSTEM) SO PATIENT SAFETY HAS NOT BEEN COMPROMISED.><NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. | 1
5224358 | NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR PART 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.><THE CUSTOMER REPORTED THAT THE RNS (REMOTE NETWORK SYSTEM OR REMOTE MONITORING SYSTEM) WAS EXPERIENCING COMMUNICATION LOSS. | 1
4713787 | THE PATIENT HAS MULTIORGAN FAILURE THOUGH WAS IMPROVING, WEANING FROM THE VENTILATOR, RECOVERING FROM SEPSIS, AND ON 3 DAYS PER WEEK RENAL DIALYSIS. THE PATIENT WAS ORDERED NORMAL SALINE TO RUN AT A RATE OF 75 CC PER HOUR. THIS CPOE ORDER FOR INTRAVENOUS FLUID WAS ENTERED ONE WEEK PRIOR TO THE EVENT, BY A CONSULTANT WHO WAS DOING A PROCEDURE ON THIS PATIENT. THE PATIENT COMPLAINED OF EXACERBATION OF DYSPNEA AND WAS FOUND TO HAVE PULMONARY EDEMA ON CHEST RADIOGRAPHY. ON DETAILED REVIEW OF MAR (MEDICINE LIST), IT WAS DETERMINED THAT THE INTRAVENOUS FLUID FOR THE PROCEDURE WAS NEVER DISCONTINUED AS ITS LISTING IN THE MAR WAS MIDWAY THROUGH MORE THAN A PAGE/SCREEN OF MEDICATIONS AND WAS NOT LISTED IN A CATEGORY HEADED AS INTRAVENOUS MEDICATIONS AND FLUIDS." BECAUSE THE PATIENT WAS ANURIC | 1
4894714 | CABLE EXPOSED IN DA VINCI ROBOT FORCEP ARM. THE EXPOSED CABLE WAS NOTICED BEFORE IT WAS TO BE USED DURING A PROCEDURE AND IT WAS TAKEN OUT OF CIRCULATION AND THE ARM WAS RETURNED TO THE MANUFACTURER.======================MANUFACTURER RESPONSE FOR DAVINCI FORCEP, PK DISSECTING FORCEPS (PER SITE REPORTER).======================HOSPITAL REQUESTS QUARTERLY REPORTS REGARDING DA VINCI ARMS THAT HAVE FAILED. MANUFACTURER SUPPORTIVE AND HELPFUL. | 0
5020766 | THE HOSPITAL IS IN THE PROCESS OF UPGRADING PHYSIOLOGIC MONITORS TO THE PHILIPS PIIC IX VERSION OF THE INTELLIVUE MONITORING SYSTEM. WE HAD BEEN USING A PREVIOUS VERSION OF THE PHILIPS INTELLIVUE SYSTEM. SINCE THE UPGRADE, IT HAS BEEN NOTED THAT THE SYSTEM WILL SOMETIMES BE UNABLE TO CONNECT TELEMETRY MONITORED PATIENTS TO THE CENTRAL STATION. THIS WILL OCCUR WHEN THE PATIENT IS BEING ADMITTED, OR WHEN THE BATTERY IS CHANGED IN THE TELEMETRY TRANSMITTER. MOST FREQUENTLY, THE PROBLEM OCCURS WHEN THE BATTERY IS REPLACED IN THE TELEMETRY TRANSMITTER. THE TRANSMITTER BATTERY NEEDS TO BE REPLACED TWICE PER DAY. THE MP5T BEDSIDE MONITORS WILL DETECT THE SIGNAL FROM THE TELEMETRY TRANSMITTER, BUT THE SIGNAL IS NOT SENT TO THE CENTRAL STATION. THIS LEAVES THE PATIENT UNMONITORED AT THE CENTRAL. THE NURSE REPLACING THE BATTERY MAY NOT BE AWARE OF THE PROBLEM AT THE CENTRAL, SINCE THE BEDSIDE MONITOR IS ABLE TO CONNECT. WE ARE CURRENTLY WORKING WITH PHILIPS TO RESOLVE THE ISSUE. MANUFACTURER RESPONSE FOR CENTRAL STATION, PIIC IX (PER SITE REPORTER): PHILIPS IS CURRENTLY WORKING ON DETERMINING THE CAUSE OF THE ISSUE. MANUFACTURER RESPONSE FOR TELEMETRY TRANSMITTER, INTELLIVUE (PER SITE REPORTER): PHILIPS IS CURRENTLY WORKING WITH US TO DETERMINE THE CAUSE OF THE ISSUE.><. | 1
5210293 | .><DAVINCI MEGA SUTURE NEEDLE DRIVER SNAPPED DURING THE PROCEDURE AND BECAME NON-FUNCTIONAL. | 0
4399172 | ON (B)(6) 2014, LEICA BIOSYSTEMS RECEIVED INFORMATION FROM THE LABORATORY THAT RE-BIOPSY OF THREE PATIENTS WAS REQUIRED. ALL PATIENT IDENTIFIER INFORMATION HAS BEEN REQUESTED, BUT NOT RECEIVED AS OF (B)(6) 2015. LEICA WILL SUBMIT A FOLLOW-UP REPORT IF ADDITIONAL PATIENT IDENTIFIER INFORMATION IS OBTAINED. REFERENCE MFR # (B)(4). | 1
4456020 | IMMUCOR TECHNICAL SUPPORT WAS UNABLE TO OBTAIN REMOTE ELECTRONIC ACCESS TO THE INSTRUMENT AT THE CUSTOMER SITE TO ASSESS THE INSTRUMENT, AND ALSO NO INFORMATION WAS RETURNED FROM THE CUSTOMER SITE IN RESPONSE TO MULTIPLE CONTINUED REQUESTS FROM IMMUCOR TECHNICAL SUPPORT FOR ADDITIONAL INFORMATION. THE ROOT CAUSE IS THEREFORE UNKNOWN, OF WHICH ONE POSSIBLE OPTION IS A WRONGLY LABELED BLOOD SAMPLE. HOWEVER, BASED ON THE DOCUMENTED ABO MISTYPE AT THE CUSTOMER SITE, THIS SITUATION IS THEREFORE REPORTED.MINIMAL INFORMATION FROM CUSTOMER SITE.><ON (B)(6) 2015, A CUSTOMER REPORTED AN UNEXPECTED ABO BLOOD GROUP MISTYPE WITH A GALILEO INSTRUMENT. | 1
4834794 | ON (B)(6) 2015, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY IDENTIFICATION ON A GALILEO ECHO, WHEN TESTED ON (B)(6) 2015.><IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ACCESS THE CUSTOMER ON-SITE INSTRUMENT FILES ON (B)(6) 2015. | 0
4794751 | AN IMMUCOR FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE TO ASSESS THE INSTRUMENT IN QUESTION ON 29APR2015. THE FSE REPLACED THE PROBE AND ALSO THE 1000UL SYRINGE. THE FSE ALSO ADJUSTED THE SOFTWARE WASH X AND WASH Y VALUES. THE FSE ALSO OBTAINED THE APPROPRIATE TEST BATCH INFORMATION AND SUPPLIED THAT INFORMATION BACK TO IMMUCOR TECHNICAL SUPPORT FOR REVIEW. THAT REVIEW DETERMINED THAT BUBBLES WERE OBSERVED IN THE COLOR CHECK WELL IMAGES, MEANING THAT THERE WAS NO SAMPLE LIQUID PIPETTED INTO THE AFFECTED WELLS. THIS WAS AN INSTRUMENT MALFUNCTION. FOLLOWING THE FSE ACTIVITY AT THE SITE, THE INSTRUMENT WAS THEN SUBSEQUENTLY DETERMINED TO BE OPERATING AS EXPECTED.><ON (B)(6) 2015, A CUSTOMER REPORTED SOME UNEXPECTED NEGATIVE ANTIBODY IDENTIFICATION WELLS ON A GALILEO ECHO, WHEN TESTED ON (B)(6) 2015. | 0
5278409 | A REVIEW OF THE IMAGE RESULT FILES SHOWED THAT NEGATIVE RESULTS APPEARED AS REPORTED BY THE INSTRUMENT. CUSTOMER WAS MADE AWARE OF THE LIMITATION IN THE PACKAGE INSERT FOR CAPTURE: NEGATIVE REACTIONS WILL BE OBTAINED IF THE TEST SERUM CONTAINS ANTIBODIES PRESENT IN CONCENTRATIONS TOO LOW TO BE DETECTED BY THE TEST METHODS EMPLOYED. REACTIVITY WAS AT THE LIMIT OF PERFORMANCE ON THE ECHO.><ON (B)(6) 2015, THE CUSTOMER STATED THAT DURING VALIDATION OF ECHO M01788, A SAMPLE KNOW TO CONTAIN ANTI-FYA RESULTED AS NEGATIVE. | 0
4736196 | ON (B)(6) 2015, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY IDENTIFICATION ON A GALILEO ECHO.><AN IMMUCOR FIELD SERVICE ENGINEER (FSE) VISITED THE CUSTOMER SITE TO ASSESS THE INSTRUMENT IN QUESTION ON 06APR2015, 08APR2015, 16APR2015 AND 24APR2015. ON 06APR2015, THE FSE FOUND THE WASHER RESIDUAL VOLUME OUT OF RANGE AND SO IT WAS CORRECTED. ON 08APR2015, 16APR2015 AND 24APR2015, THE FSE INVESTIGATED, UNINSTALLED AND SUBSEQUENTLY RE-INSTALLED THE CAMERA READER. WHILE PERFORMING THE CAMERA CALIBRATION, THE MANUAL GAIN CONTROL WAS FOUND TO BE OUT OF SPECIFICATIONS WHICH WAS THEN CORRECTED IN ORDER TO PASS THE SPECIFICATIONS FOR RGB. | 0
5218392 | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ACCESS THE TEST WELL RESULTS AND IMAGES ON THE CUSTOMER'S INSTRUMENT IN QUESTION.><ON (B)(6) 2015, A CUSTOMER REPORTED AN UNEXPECTED NEGATIVE ANTIBODY SCREEN WHEN TESTING ON A GALILEO ECHO, WHEN TESTED ON (B)(6) 2015. | 0
4387631 | IMMUCOR TECHNICAL SUPPORT USED A REMOTE ELECTRONIC CONNECTION METHOD TO ASSESS THE INSTRUMENT LOG FILES ON (B)(4) 2014. THE IMMUCOR CUSTOMER COMMUNICATIONS THAT DESCRIBE AND DISCUSS THIS SCENARIO (B)(4) WERE PROVIDED TO THE CUSTOMER ON (B)(4) 2014.><ON (B)(6) 2014, A CUSTOMER REPORTED THAT UNEXPECTED DUPLICATE BLOOD SAMPLE IDENTIFICATION NUMBERS WERE TESTED ON A GALILEO ECHO, WHICH YIELDED A SET OF UNEXPECTED ERRONEOUS BLOOD TYPES, WHEN COMPARED TO THE INITIAL CORRECT RESULTS ASSOCIATED WITH THE INITIAL TESTING. | 1
5008456 | A REVIEW OF THE IMAGE RESULT FILES SHOWED THAT CELLS POSITIVE FOR THE (B)(6) RESULTED AS (B)(6) BUT VISUALLY APPEARED (B)(6) ON THE ANTIBODY SCREEN AND ANTIBODY IDENTIFICATION PANELS. CUSTOMERS WERE NOTIFIED OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS.><THE CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS ON THE GALILEO ECHO (M00492) WITH A PATIENT SAMPLE CONTAINING (B)(6). | 0
5223830 | IMMUCOR HAS DETERMINED THAT IMMULINK SOFTWARE VERSION 1.6.1.7 FAILS TO FLAG AND BLOCK RESULTS THAT ARE DISCREPANT WITH A PATIENT'S HISTORICAL RESULTS. THIS IS AN OPTIONAL SAFETY FEATURE INTENDED TO BRING DISCREPANCIES TO LABORATORY PERSONNEL ATTENTION FOR INVESTIGATION PRIOR TO RELEASE OF RESULTS.><IMMUCOR ADVISES USERS OF IMMULINK VERSION 1.6.1.7 TO VERIFY PATIENT RESULTS AGAINST HISTORICAL RESULTS USING DATA FROM PATIENT MEDICAL RECORDS OR USING LABORATORY INFORMATION SYSTEM DATA, RATHER THAN USING IMMULINK FOR THIS PURPOSE. CUSTOMER COMMUNICATION (B)(4) WAS SENT TO CUSTOMERS AND DISTRIBUTORS ON 12 NOV 2015. CORRECTION REPORT 1034569-11/13/2015-003-C IS ATTACHED. | 1
4935259 | THE CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS ON THE GALILEO ECHO (B)(4)WITH A PATIENT SAMPLE CONTAINING ANTI-FYA.><A REVIEW OF THE IMAGE RESULT FILES SHOWED THAT CELLS 1 AND 3 OF THE ANTIBODY SCREEN RESULTED AS NEGATIVE BUT VISUALLY APPEARED TO HAVE SLIGHT RED CELL ADHERENCE. CUSTOMERS WERE NOTIFIED OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS. | 0
5057821 | A REVIEW OF THE IMAGE RESULT FILES FOR INITIAL TESTING SHOWED THAT CELL1 VISUALLY APPEARED TO HAVE SLIGHT RED CELL ADHERENCE WITH A FUZZY BUTTON. CUSTOMERS WERE NOTIFIED OF TECHNICAL COMMUNICATION (B)(4) WHICH ADVISES CUSTOMERS TO VISUALLY INSPECT ALL NEGATIVE ANTIBODY SCREENING AND IDENTIFICATION RESULTS ON THE ECHO PRIOR TO RELEASE OF RESULTS.><CUSTOMER REPORTED OBTAINING UNEXPECTED NEGATIVE REACTIONS ON ECHO M01254 FOR A PATIENT SAMPLE CONTAINING ANTI-K. | 0
4652730 | BOSTON SCIENTIFIC RECEIVED INFORMATION THAT THIS PATIENT CALLED TO REPORT THE LATITUDE REMOTE MONITORING COMMUNICATOR DISPLAYED A YELLOW FLASHING INDICATOR ICON. LATITUDE PATIENT SERVICES EXPLAINED THAT THIS WAS PART OF A COMMUNICATOR FIRMWARE UPGRADE AND WAS NORMAL FUNCTION. SEVERAL DAYS LATER THE PATIENT CALLED AGAIN AS A DIFFERENT ICON ON THE MONITOR WAS FLASHING YELLOW. TROUBLESHOOTING WAS PERFORMED AND IT WAS DETERMINED THAT THERE WAS A DATA ENTRY ERROR MADE BY THE FOLLOWING CLINIC, AND THE PATIENT WOULD NEED TO BE ENROLLED INTO THE LATITUDE SYSTEM AGAIN. AS THE HEALTH CARE PROFESSIONAL (HCP) WAS ATTEMPTING TO RE-ENROLL THIS PATIENT, A MESSAGE DISPLAYED THAT THE IMPLANTED DEVICE AND DATE OF BIRTH COULD NOT BE FOUND. BOSTON SCIENTIFIC NEEDED TO COMPLETE A DATABASE UPDATE TO CORRECT THE INITIAL DATA ENTRY ERROR. ONCE THIS WAS COMPLETED, THE PATIENT WAS THEN SUCCESSFULLY ENROLLED. THERE WERE NO ADVERSE PATIENT EFFECTS AS A RESULT OF THE OBSERVATION.><AS NO FURTHER INFORMATION CONCERNING THIS REPORT IS EXPECTED, OUR INVESTIGATION IS COMPLETE. THIS INVESTIGATION WILL BE UPDATED SHOULD FURTHER INFORMATION BE PROVIDED. | 1
4379934 | ><NEEDLE CAP FOR SYRINGE PUNCTURES TOO READILY, ALLOWING NEEDLESS NEEDLE PUNCTURES TO PROVIDERS. REPEATED EPISODES WITH CAREFUL APPROPRIATE USE. NOTIFIED SUPPLIER. | 0
5284898 | CIRCADIANCE, (B)(4) RECEIVED A CUSTOMER COMPLAINT ON (B)(6) 2015 WHERE IT WAS ALLEGED THAT A SMARTMONITOR 2 UNIT WAS NOT PRODUCING AN AUDIBLE ALARM. THE UNIT WAS SENT BACK TO CIRCADIANCE, (B)(4). UPON INSPECTION OF THE DEVICE, IT WAS DETERMINED THAT THE DEVICE WAS PRODUCING AN AUDIBLE ALARM; HOWEVER, THE ALARM SOUND WAS DISTORTED. THE CUSTOMER STATED THAT THE DEVICE WAS IN USE AT THE TIME THE PROBLEM WAS DISCOVERED AND NO PATIENT OR USER HARM WAS ALLEGED OR HAD OCCURRED. THE ALARM MODULE HAS BEEN RETURNED TO THE MANUFACTURER FOR FURTHER INVESTIGATION. RESULTS OF THE INVESTIGATION WILL BE PROVIDED IN A FOLLOW UP REPORT. THE SMARTMONITOR 2 DEVICE HAS PROVISIONS TO ALLOW THE USER TO DETECT FAILURE OF THE AUDIBLE ALARM. SPECIFICALLY, THE USERS ARE INSTRUCTED THAT THE AUDIBLE ALARM WILL ANNUNCIATE EVERY TIME THAT THE MONITOR IS TURNED ON ALLOWING THE USER TO VERIFY ALARM FUNCTIONALITY PRIOR TO PLACING THE DEVICE INTO SERVICE. ADDITIONALLY, USERS ARE INSTRUCTED TO PERFORM A FULL DEVICE CHECKOUT BY FOLLOWING THE PUBLISHED SMARTMONITOR 2 CHECKOUT PROCEDURE BETWEEN USES ON DIFFERENT PATIENTS OR AT LEAST ANNUALLY IF THE MONITOR HAS NOT BEEN IN USE. THE STEPS OF THE CHECKOUT PROCEDURE FULLY VERIFY ALL OF THE ALARM FUNCTIONS OF THE SMARTMONITOR 2. THESE INSTRUCTIONS AND FEATURES MINIMIZE THE RISK THAT AN AUDIBLE ALARM FAILURE CAN RESULT IN PATIENT INJURY. THE INSTRUCTIONS AND FEATURES NOTIFIED THE CAREGIVER AND SUBSEQUENTLY RETURNED THE DEVICE FOR SERVICING.><CIRCADIANCE, LLC HAS COMPLETED ITS INVESTIGATION OF THE SMARTMONITOR 2 PRODUCT FAILURE DETAILED IN THE INITIAL MEDWATCH SUBMISSION. THE INVESTIGATION CONCLUDED THAT THE IDENTIFIED ISSUE COULD NOT BE SUBSTANTIATED BY THE COMPONENTS MANUFACTURER. TESTING OF THE COMPONENT CONCLUDED THAT IT WAS FOUND TO OPERATE AND PERFORM TO SPECIFICATION. NO OPERATIONAL ISSUES SIGNIFYING A MALFUNCTION WERE RECORDED DURING THE MANUFACTURER'S INVESTIGATION. THE OPERATOR'S MANUAL RECOMMENDS THAT BEFORE YOU USE THE SMART MONITOR 2, TEST TO SEE IF YOU CAN HEAR THE ALARM IN DIFFERENT ROOMS WHILE THERE IS NOISE IN YOUR HOUSE. THE OPERATOR'S MANUAL RECOMMENDS TO ALWAYS KEEP THE AREA IN FRONT OF THE SPEAKER CLEAR AND TO TURN THE MONITOR ON (WITHOUT THE CHILD ATTACHED) TO SOUND THE ALARM, MAKING SURE YOU CAN HEAR THE ALARM IN DIFFERENT AREAS OF YOUR HOME. CIRCADIANCE HAS DETERMINED THAT BASED ON A COMPLETE REVIEW OF THE AVAILABLE INFORMATION, IT IS CONCLUDED THAT USE OF THE DEVICE DOES NOT PRESENT AN INCREASED RISK TO THE END USER OR PATIENT AND THAT NO CORRECTIVE ACTION OR ADDITIONAL INVESTIGATION ACTIVITY IS NECESSARY AT THIS TIME. | 1
4514847 | VESSEL SEALER KNIFE DID NOT FUNCTION PROPERLY. JAWS DID NOT COME TOGETHER. | 0
5273073 | .><WHILE THE SURGEON WAS USING THE MEGA NEEDLE HOLDER, THE CABLE BROKE. NO INJURY TO PT. INSTRUMENT REMOVED FROM STERILE FIELD. ANOTHER USED TO FINISH CASE. ONE LIFE LEFT ON INSTRUMENT. | 0
5038398 | ROBOTIC HYSTERECTOMY BEING PERFORMED. WIRING INSIDE TIP OF NEEDLE DRIVER BROKE DURING THE CASE. MANUFACTURER RESPONSE FOR MEGA SUTURE CUT NEEDLE DRIVER 8 MM, MEGA SUTURE CUT NEEDLE DRIVER 8 MM (PER SITE REPORTER): NONE AT THIS TIME.><. | 0
4798172 | THE ISSUE INVOLVES CERNER RETAIL PHARMACY AND OCCURS WHEN A PRESCRIPTION IS CHANGED PRIOR TO SAVING IT FOR A REFILL REQUEST. WHEN THIS OCCURS, THE SYSTEM DOES NOT PROMPT THE END-USER TO PERFORM THE REQUIRED PHARMACIST VERIFICATION, AND THE USER IS NOT ALERTED THAT THIS VERIFICATION STEP HAS NOT OCCURRED. THIS ISSUE COULD RESULT IN INCORRECT MEDICATIONS BEING DISPENSED. CERNER HAS RECEIVED COMMUNICATION THAT A PATIENT RECEIVED AN INCORRECT MEDICATION. ALTHOUGH THE CERNER SOLUTION DID NOT LEAD TO THE INCORRECT MEDICATION BEING DISPENSED, THE EXPECTED VERIFICATION WHICH COULD HAVE CAUGHT THIS ERROR DID NOT OCCUR. THE OUTCOME OF THE PATIENT WAS NOT COMMUNICATED TO CERNER.><CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON MAY 7, 2015 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE, AN ALTERNATE WORKFLOW, AND NOTIFICATION THAT A SOFTWARE MODIFICATION WAS BEING DEVELOPED TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON AUGUST 20, 2015 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE, AN ALTERNATIVE TO PREVENT THE ISSUE FROM OCCURRING, AND NOTIFICATION THAT A SOFTWARE MODIFICATION IS AVAILABLE TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORPORATION CONSIDERS THE ISSUE TO BE RESOLVED AND NO FURTHER NARRATIVE IS REQUIRED FOR FOLLOW-UP.><THIS REPORT DOCUMENTS INFORMATION RELATED TO AN ISSUE IDENTIFIED IN (B)(6) PHARMACY®. THE ISSUE INVOLVES (B)(6) PHARMACY AND OCCURS WHEN A PRESCRIPTION IS CHANGED PRIOR TO SAVING IT FOR A REFILL REQUEST. WHEN THIS OCCURS, THE SYSTEM DOES NOT PROMPT THE END-USER TO PERFORM THE REQUIRED PHARMACIST VERIFICATION, AND THE USER IS NOT ALERTED THAT THIS VERIFICATION STEP HAS NOT OCCURRED. THIS ISSUE COULD RESULT IN INCORRECT MEDICATIONS BEING DISPENSED. CERNER HAS RECEIVED COMMUNICATION THAT A PATIENT RECEIVED AN INCORRECT MEDICATION. ALTHOUGH THE CERNER SOLUTION DID NOT LEAD TO THE INCORRECT MEDICATION BEING DISPENSED, THE EXPECTED VERIFICATION WHICH COULD HAVE CAUGHT THIS ERROR DID NOT OCCUR. THE OUTCOME OF THE PATIENT WAS NOT COMMUNICATED TO CERNER.><CERNER DISTRIBUTED A PRIORITY REVIEW FLASH NOTIFICATION ON (B)(6) 2015 TO ALL POTENTIALLY IMPACTED CLIENT SITES. THE SOFTWARE NOTIFICATION INCLUDES A DESCRIPTION OF THE ISSUE, AN ALTERNATE WORKFLOW, AND NOTIFICATION THAT A SOFTWARE MODIFICATION IS BEING DEVELOPED TO ADDRESS THE ISSUE FOR ALL SITES THAT COULD BE POTENTIALLY IMPACTED. CERNER CORPORATION WILL PROVIDE A FOLLOW-UP REPORT WHEN THE SOFTWARE MODIFICATION IS AVAILABLE. | 1
4576208 | USER FACILITY REPORTED THAT A WHITE FOREIGN MATERIAL WAS FOUND IN THE ACIST AUTOMATED MANIFOLD KIT, MODEL BT2000, LOT 27414E. THE MATERIAL WAS 3 MM IN SIZE, LOCATED IN THE LOW PRESSURE PERISTALTIC TUBING NEAR THE SALINE SPIKE. THE KIT WAS NOT USED BY THE FACILITY. THE FOREIGN MATERIAL WAS SENT TO A THIRD-PARTY VENDOR FOR FURTHER ANALYSIS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN AN ANALYSIS AND INVESTIGATION HAS BEEN COMPLETED ON THIS DISPOSABLE KIT. | 0
4861580 | A SMALL CAPSULE STUDY DATA RECORDER BATTERY STOPPED AFTER 5 HOURS. THE BATTERY SHOULD LAST 7-8 HOURS. IT CAUSED AN INCOMPLETE STUDY FOR THE PATIENT. | 1
4469211 | (B)(4). THE CUVETTE LOT USED WITH THIS INSTRUMENT IS UNK AT THIS TIME. IT WAS REPORTED THAT THE PT WAS SELF-MEDICATING. THIS MAY HAVE CONTRIBUTED TO THE PROTIME MICROCOAGULATION SYSTEM TO REGISTER LOWER THAN EXPECTED INR RESULTS. ITC HAS REQUESTED ALL DATA REQUIRED FOR FORM 3500A.><USER FACILITY REPORTS LOWER THAN REFERENCE RESULTS WITH THE PROTIME MICROCOAGULATION SYSTEM. PROTIME GENERATED A 2.0 INR RESULT AT THE CLINIC. PT WAS TESTED IN THE EMERGENCY ROOM AND THE RESULT WAS 8.0 INR. THE PT'S UNDERLYING DIAGNOSIS AND THERAPEUTIC RANGE IS UNK, THEREFORE A DEFAULT INR OF 2.0-3.0 IS USED. NO ADVERSE EVENT(S) WERE REPORTED. | 0
4475930 | ANESTHESIA MONITOR ELECTRICAL MALFUNCTION AND MAIN BOARD PRODUCED BLACK SMOKE. FIRE ALARM ACTIVATED PATIENT REMOVED FROM ROOM IMMEDIATELY TRANSPORTED TO PH I. | 1
4538395 | TIP IS SCISSORED.MANUFACTURER RESPONSE FOR INTUITIVE SURGICAL, (BRAND NOT PROVIDED) (PER SITE REPORTER)======================TOOK A DESCRIPTION OF THE PROBLEM AND ISSUED RGA#. | 0
5312991 | THE OPERATOR DISCOVERED THE QUALITY CONTROL (QC) WAS OUT OF RANGE. THE PERFORMANCE OF QC IS USED BY THE OPERATOR TO DETERMINE WHETHER THE ANALYZER IS IN CONTROL AND CAN BE USED FOR PATIENT TESTING. FAILING QC WOULD MAKE THE USER AWARE OF AN ISSUE TO BE RESOLVED PRIOR TO USING THE ANALYZER FOR PATIENT TESTING. THE OPERATOR FOLLOWED GOOD LABORATORY PRACTICE BY REPEATING ALL POTENTIALLY AFFECTED PATIENT SAMPLES AND ISSUING CORRECTED REPORTS AS NEEDED. THE TYPE OF ANTICOAGULANT THERAPY COULD NOT BE CONFIRMED BY THE OPERATOR. THE AMERICAN ASSOCIATION FOR CLINICAL CHEMISTRY (AACC) INDICATES THE NORMAL INR RANGE IS TYPICALLY 0.9-1.1. ON WARFARIN THERAPY THE INR ELEVATES TO BETWEEN 2.0 AND 3.5. MOST HOSPITAL PHARMACIES AND CLINICAL HEMATOLOGY SERVICES WILL HAVE SPECIFIC INR GOALS DETERMINED IN THEIR SPECIFIC TREATMENT PROTOCOLS. THE OPERATOR STATED EACH DOCTOR OR CLINIC HAS THEIR OWN SET OF REFERENCE RANGES AND WAS UNWILLING TO PROVIDE RANGES. PATIENT 1 HAD A DOSE OF ANTICOAGULANT THERAPY HOWEVER THE DOSE SHOULD HAVE BEEN INCREASED BASED UPON THE CORRECTED VALUE FROM REPEAT ANALYSIS. THE OPERATOR'S CONTACT WITH SIEMENS GENERATED A SERVICE CALL. THE SIEMENS FIELD SERVICE ENGINEER (FSE) EXAMINED THE ANALYZER, AND DETERMINED REAGENT PROBE B2 WAS OPERATING TOO HOT. THE ERROR LOG HAD NO TEMPERATURE ERRORS RECORDED. ERROR LOGS OR TEMPERATURE READINGS WERE NOT PROVIDED. IT IS NOT KNOWN WHY THE REAGENT PROBE WAS FOUND TO BE TOO HOT WITHOUT TEMPERATURE ERRORS. THE ISSUE WAS ISOLATED TO THE REAGENT PROBE. REPLACEMENT AND ALIGNMENT OF THE REAGENT PROBE RESOLVED THE ISSUE. TEMPERATURE WAS WITHIN RANGE AFTER PART REPLACEMENT. QC WAS ANALYZED AND WITHIN THE LABORATORY'S ESTABLISHED RANGES. SINCE THE REAGENT PROBE REPLACEMENT, THE ANALYZER HAS BEEN OPERATING WITHIN SPECIFICATIONS. THIS EVENT WILL BE REPORTED ON THE BASIS THAT ERRONEOUS RESULTS MAY HAVE LED TO THE PATIENT NOT RECEIVING APPROPRIATE ANTICOAGULANT THERAPY, CARRYING WITH IT THE RISKS INVOLVED: BLEEDING EPISODES AND THROMBOSIS.><THE OPERATOR STATED ERRONEOUS, ELEVATED INTERNATIONAL NORMALIZED RATIO (INR) RESULTS WERE GENERATED BY THE ANALYZER. THE OPERATOR STATED ON (B)(6) 2015 MORNING QC WAS ANALYZED AND WITHIN RANGE. PATIENT SAMPLES WERE PROCESSED AND RESULTS REPORTED. THE NEXT ANALYSIS OF QC PERFORMED LATER THAT DAY RECOVERED HIGH OUT OF RANGE. THE OPERATOR MADE FRESH REAGENT AND CONTROLS AND REPEATED QC WITH NO CHANGE. THE SAME QC AND REAGENTS WERE USED ON AN ALTERNATE CA-7000 ANALYZER, SERIAL NUMBER NOT PROVIDED, QC RECOVERED WITHIN RANGES. NO ERROR LOGS OR ANALYZER PRINTOUTS WERE PROVIDED. THE LABORATORY REANALYZED 72 PATIENT SAMPLES ON AN ALTERNATE ANALYZER FOR VERIFICATION. THE LABORATORY IDENTIFIED SEVERAL TEST RESULTS WERE REPORTED BETWEEN THE ACCEPTABLE QC RUN AND THE FAILED QC. THE OPERATOR SUBMITTED A RECORD OF AFFECTED PATIENT RESULTS. THE LABORATORY ISSUED CORRECTED REPORTS UPON REPEAT ANALYSIS ON AN ALTERNATE ANALYZER. THERE WAS NO HARM TO ANY OF THE AFFECTED PATIENTS. PATIENT 1 INITIAL INR WAS WITHIN THE THERAPEUTIC RANGE, BUT WAS BELOW THE THERAPEUTIC RANGE WHEN REPEATED ON AN ALTERNATE ANALYZER. THE PATIENT RECEIVED A DOSE OF ANTICOAGULANT TREATMENT. ACCORDING TO THE PHYSICIAN, PATIENT 1 SHOULD HAVE HAD ANTICOAGULANT DOSE INCREASED AND DID NOT. THE OPERATOR WAS UNWILLING TO PROVIDE PT VALUES AND ANTICOAGULANT TREATMENT TYPE WAS ASSUMED TO BE WARFARIN THE OPERATOR WAS NOT ABLE TO CONFIRM. | 0
4588104 | THE BWI FAILURE ANALYSIS LAB RECEIVED THE DEVICE FOR EVALUATION ON 03/20/2015. THE ANALYSIS HAS BEGUN BUT IS NOT COMPLETED AT THIS TIME. WHEN THE INVESTIGATIONAL ANALYSIS HAS BEEN COMPLETED, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED. (B)(4)><MANUFACTURER'S REF. NO: (B)(4) IT WAS REPORTED THAT A PATIENT, (B)(6), MALE, UNDERWENT AN ATRIAL FLUTTER (AFL) PROCEDURE WITH A C3 INTERFACE THERAPEUTIC CABLE AND RECEIVED AN ELECTRICAL SHOCK, BECAME ASYSTOLIC AND THEN WENT INTO ATRIAL FIBRILLATION WHICH REQUIRED HEMODYNAMIC SUPPORT. A VISUAL INSPECTION WAS PERFORMED ON THE EXTENSION CABLE, AND THE CABLE WAS FOUND IN GOOD CONDITION, THERE WAS NO FAILURE THAT CAN RESULT IN A COMPLAINT. PROPER CABLE PINS IN BOTH CONNECTOR PLUG AND RECEPTACLE CONNECTORS, NO BENT PIN OR DAMAGED RECEPTACLE WAS FOUND. NO CRACKS OR SCRATCHES WERE FOUND ON CONNECTOR PLASTIC COVER. ONLY MINOR FINDINGS WERE OBSERVED ON THE CABLE WHICH DO NOT COMPROMISED THE CABLE FUNCTIONALITY. PER THE EVENT, THE CABLE WAS TESTED FOR ELECTRICAL PERFORMANCE, TEMPERATURE RESPONSE AND STOCKERT COMPATIBILITY AND IT WAS FOUND WITHIN SPECIFICATIONS. THE CABLE WAS ALSO EVALUATED FOR CARTO 3. THE TEST CATHETER WAS CONNECTED TO THE REPORTED CABLE WAS RECOGNIZED BY CARTO 3 SYSTEM, NO ERROR MESSAGES WERE DISPLAYED AND THE CATHETER WAS PROPERLY VISUALIZED. THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED AND NO ANOMALIES WERE FOUND RELATED TO THIS COMPLAINT. IN ADDITION, THE DHR REVIEW VERIFIES THAT THE DEVICE WAS MANUFACTURED IN ACCORDANCE WITH DOCUMENTED SPECIFICATION AND PROCEDURES. THE CABLE PASSED ALL SPECIFICATIONS. THE ROOT CAUSE OF THE SERIOUS INJURY REMAINS UNKNOWN.><IT WAS REPORTED THAT A PATIENT, (B)(6) YEAR OLD, MALE, UNDERWENT AN ATRIAL FLUTTER (AFL) PROCEDURE WITH A C3 INTERFACE THERAPEUTIC CABLE AND RECEIVED AN ELECTRICAL SHOCK, BECAME ASYSTOLIC AND THEN WENT INTO ATRIAL FIBRILLATION WHICH REQUIRED HEMODYNAMIC SUPPORT. AN ABLATION CATHETER THAT WAS ALREADY IN THE PATIENT¿S HEART WAS PLUGGED INTO A REPROCESSED MAPPING CABLE WHICH WAS ALREADY CONNECTED TO THE PATIENT INTERFACE UNIT (PIU). UPON CONNECTION, THERE WERE LARGE AMPLITUDE SPIKES AT ABOUT 100 MS AND SIGNAL SATURATION WAS NOTICED THROUGHOUT ALL ELECTROCARDIOGRAMS (ECG), ELECTROGRAMS (EGM) AND MONITORING CHANNELS MAKING IT UNABLE TO SEE ANY ECG SIGNAL. IN ADDITION, THERE WAS A CURRENT LEAKAGE" ALERT ON THE CARTO 3 SYSTEM. THE PATIENT | 0
4774323 | ><THE GREEN PORTION OF THE SUCTION SWAB HAD TO BE EXTRACTED DURING AN EGD PROCEDURE AS IT WAS OBSTRUCTING THE NASOGASTRIC TUBE PLACEMENT. THE FOREIGN BODY WAS BROKEN OFF FROM THE STICK. DATES OF USE: SINCE (B)(6) 2014. | 0
5333026 | CURRENT INFORMATION IS INSUFFICIENT TO PERMIT A CONCLUSION AS TO THE CAUSE OF THE EVENT. THERE ARE WARNINGS IN THE PACKAGE INSERT THAT STATE THAT THIS TYPE OF EVENT CAN OCCUR: A TRAINED MEDICAL PROFESSIONAL DETERMINES FINAL IMPLANT TYPE AND SIZE INTRAOPERATIVELY." "ORTHOSIZE SOFTWARE DOES NOT DETERMINE THE FINAL SIZE AND TYPE OF HARDWARE TO BE IMPLANTED."><THIS FOLLOW-UP REPORT IS BEING FILED TO RELAY ADDITIONAL INFORMATION | 0
4508528 | THE CUSTOMER REPORTED THERE WAS SMOKE EMANATING FROM THE B20. A FIELD ENGINEER CONFIRMED THAT THE POWER SUPPLY UNIT WAS DAMAGED, REPLACED IT WITH A NEW POWER SUPPLY UNIT AND THE B20 WORKED NORMAL. THE DEFECTIVE PART PICTURES SHOWED THE CAPACITOR C5 NEXT TO THE HEAT SINK IS SEEN BULGED WITH NO BURNT COMPONENTS. THE SUPPLIER (MEAN WELL) OF THE POWER SUPPLY PROVIDED FEEDBACK STATING THE DEFECTED CAPACITOR WAS AN ELECTROLYTIC CAPACITOR WHICH IS USED FOR ENERGY STORAGE AND FILTERING. THE WHITE SMOKE WAS FORMED BY EVAPORATION OF THE ELECTROLYTE WHEN THE ELECTROLYTIC CAPACITOR WAS BROKEN. THERE ARE TWO POSSIBLE CAUSES FOR THIS: 1. CAPACITOR QUALITY ISSUE; 2. INPUT VOLTAGE IS TOO HIGH. FOR CAPACITOR QUALITY, THE POWER SUPPLY UNIT HAD PERFORMED FUNCTION TEST UNDER 264VAC POWER INPUT AND BURN-IN TEST UNDER 230VAC POWER INPUT BEFORE SHIPMENT, WHICH CAN SORT THE DEFECT COMPONENT, AND CHECKED THE CAPACITOR FAILURE RECORDS OF THE SERIES MODEL. NO CAPACITOR FAILURE RECORD WAS FOUND, THUS NO EVIDENCE TO SHOW IT IS A CAPACITOR QUALITY ISSUE. FOR INPUT VOLTAGE, THE MAXIMUM INPUT VOLTAGE OF THE POWER SUPPLY UNIT IS 264V AC, AND THE VOLTAGE ACROSS THE CAPACITOR IS THE VOLTAGE AFTER RECTIFIER. THE VOLTAGE ACROSS THE CAPACITOR=1.414*AC INPUT VOLTAGE. THE MAX VOLTAGE ACROSS THE CAPACITOR=1.414*264=373V. THE CAPACITOR MAX VOLTAGE RATING IS 400V, THE VOLTAGE ACROSS THE CAPACITOR UNDER THE RATED INPUT VOLTAGE SPEC OF THE CAPACITOR (400V). FROM THE INDIA SERVICE TEAM FEEDBACK, INDIA'S POWER GRID VOLTAGE IS 230V 50HZ AND THE POWER GRID IS UNSTABLE WHICH MAY CAUSE POWER FLUCTUATION/SURGE. ROOT CAUSE IS UNDETERMINED, THOUGH MOST LIKELY THE INPUT VOLTAGE WAS NOT STABLE WHEN THE CUSTOMER USED THE DEVICE WHICH CAUSED THE VOLTAGE ACROSS THE CAPACITOR TO EXCEED THE SPECIFICATIONS, CAUSING DAMAGE TO THE ELECTROLYTIC CAPACITOR. IN THE USER MANUAL 2053503-002, THERE IS A CAUTION THAT: BEFORE CONNECTING POWER, CHECK VOLTAGE AND FREQUENCY RATINGS OF EQUIPMENT.><NO PATIENT INVOLVEMENT. OCCUPATION OF REPORTER IS UNKNOWN. GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION HAS BEEN COMPLETED.><THE CUSTOMER REPORTS SMOKE WAS EMITTING FROM THE DEVICE. NO PATIENT INVOLVEMENT. | 1
4823622 | ><WHEN WE USE THE PEDS DISPOSABLE MASIMO OXIMETRY SENSORS IN CONJUNCTION WITH THE GE PDM (PATIENT DATA MODULE) AND B850 MONITORS, WE ARE CONSISTENTLY EXPERIENCING 3 FAILURE SITUATIONS:1. FAILURE TO ACQUIRE A SIGNAL FROM THE SENSOR WITHIN A REASONABLE TIME FRAME AT THE BEGINNING OF A CASE.2. COMPLETE LOSS OF SIGNAL DURING A CASE.3. ERRONEOUS DESATURATION SIGNAL DURING A CASE.TO DATE NO SIGNIFICANT PATIENT HARM HAS OCCURRED; HOWEVER WE ARE MOST CONCERNED ABOUT THE POTENTIAL FOR PATIENT HARM THROUGH MISDIAGNOSIS OR MISTREATMENT DUE TO UNRELIABLE EQUIPMENT. MULTIPLE EVENTS HAVE OCCURRED OVER THE LAST 6 MONTHS WITH THIS EQUIPMENT AND THIS TYPE OF SCENARIO. MANUFACTURERS OF DEVICES WERE NOTIFIED AND HAVE COME ON-SITE TO OBSERVE AND WORK WITH STAFF/EQUIPMENT.======================MANUFACTURER RESPONSE FOR MASIMO DISPOSABLE PEDS OXIMETRY SENSOR, MASIMO (PER SITE REPORTER).======================PROVIDED ENGINEERING AND CLINICAL SUPPORT, NEW EQUIPMENT, TRAINING, DATA SUPPORT/ MANUFACTURER HAS CONTINUED TO WORK WITH US TO RESOLVE ISSUES.======================MANUFACTURER RESPONSE FOR GE PDM AND B850 MONITOR, CARESCAPE MONITOR SERIES (PER SITE REPORTER).======================RECEIVED EQUIPMENT, ENGINEERING SUPPORT, DATA COLLECTION, ON-SITE CLINICAL AND ENGINEERING SUPPORT. HAVE BEEN WORKING WITH MANUFACTURER SINCE PROBLEMS IDENTIFIED. | 1
5179796 | .><FORCEP TIPS BECAME UNALIGNED DURING THE PROCEDURE - HAD TO STOP USING. | 0
4437960 | THERE WAS NO PT INVOLVED IN THIS EVENT. THIS DEVICE MALFUNCTIONED BECAUSE THE DEVICE WAS OBSERVED SWITCHING ON AUTOMATICALLY. | 0
5038574 | I AM A (B)(6) YEAR OLD MALE AND TAKE PRESCRIPTION BLOOD PRESSURE MEDICATION DAILY. I AM ALSO A STROKE SURVIVOR FROM A MAJOR STOKE IN 2009 DUE TO BLUNT FORCE TRAUMA; NOT A FLOW ISSUE. TODAY AT 9:15 AM I WAS CHECKING MY BLOOD PRESSURE. I USE AN OMRON BP791IT (HEM-7222-ITZ) MONITOR. IT IS SET TO TAKE MY BLOOD PRESSURE VIA THREE COMPUTERIZED MEASUREMENTS AND AVERAGE. TODAY, IT TOOK THE FIRST READING AND THEN DEFLATED THE CUFF AND RESTARTED AFTER 30 SECONDS. THAT IS NORMAL. THE SECOND READING BEGAN AND WAS OBTAINED AND THE MACHINE STOPPED AS USUAL. THEN AFTER 5 SECONDS WHILE THE CUFF WAS STILL FULLY INFLATED THE MACHINE TURNED ITSELF ON AGAIN AND BEGAN INFLATING THE CUFF WHICH WAS ALREADY FULLY INFLATED. I KIND OF FROZE AS I WAS STARTLED AND IN SHOCK AS THE PRESSURE KEPT INCREASING SUCH THAT IT WAS APPROACHING 275 ON THE MONITOR. AT THAT POINT MY ARM MUSCLE WAS IN PAIN AND IT SNAPPED ME OUT OF MY TRANCE SO I RIPPED THE CUFF OFF FROM MY ARM YET THE MACHINE CONTINUED TO INFLATE THE CUFF FOR ANOTHER (APPROXIMATELY) TEN SECONDS UNTIL IT REPORTED A READING ERROR CODE (E5) WHILE THE CUFF RESTED ON THE TABLE. THE E5 CODE MEANS THAT CLOTHING INTERFERED WITH THE READING AND THE MACHINE STOPPED. I WAS NOT WEARING ANY CLOTHING (SHIRT) TODAY NEAR THE ARM AND I USE THIS MACHINE DAILY AND THE MACHINE STOPPED WHILE NOTHING WAS INSERTED INSIDE THE CUFF EXCEPT AIR. I AM IN PRETTY GOOD SHAPE AS THE BLOOD PRESSURE ISSUE IS RELATED TO A TIGHTENED LUNG AS A RESULT OF ASTHMA. HOWEVER, I DID 75 ARM CURLS YESTERDAY WITH A 10 POUND DEAD WEIGHT SO MY MUSCLE WAS TIGHT TODAY. POSSIBLY THAT CONFUSED THE MACHINE? IF I WAS OLDER THAT MACHINE WOULD HAVE PHYSICALLY HURT SOMEONE. I AM LUCKY I SNAPPED OUT OF MY SHOCK TO WHAT WAS HAPPENING AND TORE THE CUFF OFF MY ARM. I WAS ACTUALLY MESMERIZED THAT THE MACHINE TURNED ITSELF BACK ON AS I WATCHED THE UPPER BLOOD PRESSURE NUMBER ON THE MONITOR APPROACH 275. THIS MACHINE IS FROM HELL AND YOU NEED TO REALLY LOOK INTO IT. WE PURCHASED IT IN (B)(6) 2013. (B)(4). | 0
5323381 | A VITEK GNS-112 CARD LOT P96X WAS TESTED AND INITIALLY WAS CORRECTLY IDENTIFIED BY THE VITEK SYSTEM AS A GNS-112 SUSCEPTIBILITY CARD. THE CARD TYPE IS IDENTIFIED BY A CARD CODE STAMPED ON THE SIDE OF THE CARD. AFTER 5 HOURS INCUBATION, THE CARD WAS INCORRECTLY IDENTIFIED BY THE VITEK READER AS A GNS-114 SUSCEPTIBILITY CARD. AS A RESULT OF THE DIFFERENT CONFIGURATIONS AND CONTENTS FOR THE TWO VITEK CARDS, TEN ANTIBIOTICS WERE ERRONEOUSLY ANALYZED. THIS MAY HAVE CAUSED FALSE SUSCEPTIBLE AND/OR FALSE RESISTANT RESULTS TO BE REPORTED. THE ERROR WAS DETECTED AND THE LAB NOTIFIED THE PHYSICIAN ABOUT THE PROBLEM ON THE SAME DAY PATIENT TREATMENT HAD NOT BEEN CHANGED AS A RESULT OF THE POTENTIALLY INCORRECT LAB RESULTS. IT CONTINUED AS IT WAS PRIOR TO THE LAB TEST. THE LAB WAS DIRECTED TO CLEAN THE OPTICAL ARRAYS OF THE VITEK SYSTEM, TO RETURN THE REMAINING GNS-112 FROM LOT P96X FOR EVALUATION, AND TO NOTIFY US IF ANY ADDITIONAL PROBLEMS OCCURRED.><WE WERE UNABLE TO EVALUATE THE ORIGINAL GNS-112 TEST CARD BECAUSE IT HAD BEEN DISCARDED BY THE CUSTOMER. THE CUSTOMER DID RETURN ALL REMAINING GNS-112 CARDS FROM LOT P96X. THIRTY-EIGHT RETURNED CARDS WERE EVALUATED WITH THREE VISION SYSTEM CAMERAS AND NO PROBLEMS WERE DETECTED WITH THE CARD CODES. VISUALLY, THE CODES APPEARED TO BE GOOD AND WELL CENTERED. THE CARDS WERE ALSO EVALUATED WITH THREE VITEK SYSTEM READERS AND HAD NO FAILURES. THE CUSTOMER HAS REPORTED NO FURTHER PROBLEMS. BASED ON THE DESCRIPTION OF THE EVENT AND THE FACT THAT INITIALLY THE CARD WAS CORRECTLY IDENTIFIED BY THE VITEK READER, WE SUSPECT THAT THE PRESENCE OF DUST ON THE OPTICAL ARRAYS OF THE VITEK READER IS THE PROBABLE EXPLANATION FOR THE PROBLEM. CUSTOMERS ARE PROVIDED WITH A MAINTENANCE CHECKLIST THAT RECOMMENDS THAT THE OPTICAL ARRAYS BE TO PREVENT PROBLEMS OF THIS TYPE. | 0
4713912 | ><OUR CLINICAL INFORMATION SYSTEM (CPOE, EMAR, BCMA) IS XXXX XXXXXX XXXXXX FROM ALLSCRIPTS. WE DO NOT HAVE CONTROL OF THE FORMATTING OF THEIR LABELS. THEIR CANNED FORMAT DOES NOT INCLUDE A LEADING ZERO IN FRONT OF DOSES LESS THAN ONE. WE HAD ONE NEAR MISS WITH A LEVOTHYROXINE INJECTION. WE REPORTED THIS TO ALLSCRIPTS AND THEY HAVE NOT EMBRACED IT AS A PRIORITY. I WILL SEND AN EXAMPLE OF A PENICILLIN G POTASSIUM LABEL IN A SEPARATE EMAIL. PATIENT COUNSELING PROVIDED: UNKNOWN. RELEVANT MATERIALS PROVIDED: NONE. (B)(6). | 1
4960870 | NO PATIENT INCIDENT REPORTED AT THIS MOMENT - THIS ISSUE IS DOCUMENTED AS A PRECAUTION. FIVE RELATED CUSTOMER COMPLAINTS ((B)(4)) WERE REGISTERED AND INVESTIGATED FROM OCT 2014 TO JUNE 2015 FOR RELATED ISSUES. THE INVESTIGATION OF THESE COMPLAINTS LED TO THE DISCOVERY OF ISSUES WITH IMAGE INGESTION. BUGS IDENTIFIED: (B)(4) LDS: LDS: FAILED TO INGEST IMAGE WITH OVERLAYS BELONGING TO 60XX GROUP HAVING ALPHANUMERIC ID (IE, 600A) (B)(4) LDS: ROWS STUCK IN INCOMING_FILES DUE TO FILES MISSING IN THE INCOMING FILES FOLDER. INVESTIGATION: INVESTIGATION OF THE LOGS IN (B)(4) REVEALED THAT SOME OF THE EXPECTED OVERLAY FILES WERE NOT FOUND ON THE FILE SYSTEM; THIS LED TO DISCOVERY OF THE MISSING FILES LEFT BEHIND AT AN EARLIER STAGE OF IMAGE INGESTION. ROOT CAUSE OF THE PROBLEM: DUE TO THE ASYNCHRONOUS NATURE OF IMAGE INGESTION PROCESSING, A FAILURE AT THE EARLY STAGES OF THE INGESTION PROCESS CAN LEAD TO IMAGES NOT BEING VISIBLE TO THE END USER. THERE IS NO INDICATION TO THE USER THAT THIS HAS OCCURRED. REASON FOR POSSIBLE PATIENT SAFETY: THE RADIOLOGIST MAY PROVIDE A DIAGNOSIS WITHOUT BEING AWARE THAT IMAGES ARE MISSING. DESIGN MITIGATION TO REDUCE THE RISK: PROPOSED NEW MITIGATION: THE PACS WILL REPORT PROBLEMS INGESTING IMAGES AND INDICATE SUCH ISSUES TO THE USER ((B)(4)). THIS MITIGATION WILL BE INCLUDED IN A COMING SOFTWARE RELEASE. (B)(4): LDS: EXPOSE BROKEN SERIES TO USER (WARN ABOUT ADDITIONAL IMAGES PENDING INGESTION WHEN VIEWING IMAGES). DETECTABILITY OF PROBLEM: UNTIL THE SOFTWARE UPDATE THAT WILL INCLUDE THE PROPOSED MITIGATION IS INSTALLED ON THE CLIENT'S SYSTEM, THE USER MAY NOT BE AWARE THAT SOME IMAGES MAY BE MISSING. HOWEVER, IF THE CLIENT'S SYSTEM HAS OTHER MECHANISMS TO VERIFY IMAGE COUNT BETWEEN SOURCE AND DESTINATION, THEN THE USER MIGHT BE ABLE TO KNOW ABOUT THE DISCREPANCY.><IN ORDER TO CORRECT THIS ISSUE, INTELERAD WILL: PROVIDE THE AFFECTED CLIENTS WITH A LIST OF STUDIES THAT WERE AFFECTED AND POSSIBLY RECOVER THE MISSING IMAGES. DEPLOY A SOFTWARE UPDATE THAT INCLUDES A FIX AND MITIGATION TO ALL CLIENTS WHERE THE ISSUE COULD POTENTIALLY OCCUR. THIS WILL BE DONE TENTATIVELY. PERIODICALLY AUDIT ALL AFFECTED CLIENT, ANY NEW OCCURRENCES OF THIS ISSUE UNTIL THE MITIGATION (SOFTWARE UPDATE) IS IN PLACE. THE SOFTWARE UPDATE WILL INCLUDE ONLY BUG FIXES - YOUR WORKFLOW WILL NOT BE AFFECTED. THE UPDATE WILL RESULT IN ABOUT 5 MINUTES OF DOWNTIME PER SERVER AS WE PERFORM A FULL SERVICE RESTART. IMPACT ON USERS WILL BE MINIMAL, IF ANY. CORRECTIVE ACTION CALENDAR: FOR ALL CLIENTS WITH PACS VERSION 4-5-1 OR HIGHER, THE SOFTWARE UPDATE WILL BE DEPLOYED WITHIN A TIMEFRAME TO BE AGREED WITH CLIENTS (COMPLETION ESTIMATED BY END OF MARCH 2016). | 1
5273083 | CABLE OF MEGA NEEDLE HOLDER BROKE WHILE SURGEON WAS USING IT. IT WAS REPLACED WITH ANOTHER ONE TO FINISH THE CASE.><. | 0
4644470 | (B)(4). THE INVESTIGATION WAS COMPLETED ON 06/17/2015. PRODUCT ((B)(4)) WAS TESTED AND IS FUNCTIONING ACCORDING TO SPECIFICATION.><ON 03/11/2015, ABBOTT POINT OF CARE (APOC) WAS CONTACTED BY AN ABBOTT IMPLEMENTATION PROJECT MANAGER (IPM) ON BEHALF OF THE CUSTOMER REGARDING RECHARGEABLE BATTERIES THAT BECAME HOT TO THE TOUCH WHILE CHARGING IN THE ANALYZER WHEN PLACED IN THE DOWNLOADER/RECHARGER (DRC). THE CUSTOMER STATES WHEN THE BATTERY IS REMOVED FROM THE HANDHELD AND RECHARGED IN THE DRC BATTERY COMPARTMENT THE BATTERY DOES NOT OVERHEAT. THE ISSUE OCCURRED ON (B)(6) 2015. PER I-STAT SYSTEM MANUAL ART: 714368-00K, REV. DATE: 02-AUG-12: THE DOWNLOADER/RECHARGER CAN RECHARGE A RECHARGEABLE BATTERY IN THE ANALYZER. IF THE ANALYZER CONTAINS A RECHARGEABLE BATTERY, THE BATTERY BEGINS RECHARGING AUTOMATICALLY AS SOON AS THE ANALYZER IS PLACED IN THE DOWNLOADER/RECHARGER. THE DOWNLOADER/RECHARGER ALSO HAS A COMPARTMENT FOR RECHARGING A RECHARGEABLE BATTERY OUTSIDE THE ANALYZER. PLACING AN ANALYZER IN A DOWNLOADER/RECHARGER WILL AUTOMATICALLY INITIATE RECHARGING OF THE RECHARGEABLE BATTERY. THE INDICATOR LIGHT ON TOP OF THE DOWNLOADER/RECHARGER WILL BE GREEN (TRICKLE CHARGE), RED (FAST CHARGE), OR BLINKING RED (FAST CHARGE PENDING) WHEN AN ANALYZER WITH A RECHARGEABLE BATTERY IS PLACED IN THE DOWNLOADER/RECHARGER. AS WELL, PLACING A RECHARGEABLE BATTERY INTO THE RECHARGING COMPARTMENT WILL AUTOMATICALLY INITIATE TRICKLE RECHARGING. THE INDICATOR LIGHT NEAR THE RECHARGING COMPARTMENT WILL BE GREEN WHEN A RECHARGEABLE BATTERY IS PLACED IN THE COMPARTMENT. THE CUSTOMER STATES THEY REMOVE BATTERIES FROM HANDHELDS AND CHARGE THEM IN THE DRC BATTERY COMPARTMENT OVERNIGHT. THE BATTERIES ARE PUT BACK IN THE HANDHELDS NEXT MORNING AND DISTRIBUTED FOR USE. THE ANALYZERS ARE OPERATING AS EXPECTED AND ARE RETURNED AT NIGHT TO RECHARGE THE BATTERIES. THE ISSUE APPEARS TO BE WHEN THE BATTERY IS CHARGED IN THE DRC WHILE IN THE ANALYZER. THERE ARE NO INJURIES ASSOCIATED WITH THIS EVENT. AT THIS TIME THERE IS NO EVIDENCE OF A CONFIRMED MALFUNCTION AND PRODUCT WAS REPLACED AT NO CHARGE. THE CUSTOMER WAS ASKED TO RETURN THE DRC ALONG WITH IT'S ASSOCIATED RECHARGEABLE BATTERY PACK FOR INVESTIGATION.><(B)(4).><. | 0
4684158 | ><MRI COMPATIBLE PORTABLE VENTILATOR MACHINE WAS PULLED INTO THE 3T MRI MAGNET. VENTILATED PT WAS REMOVED FROM THE VENTILATOR AND THE MRI MAGNET. AS THE ASPIRATORY TECH WAS CLEANING AND GETTING THE VENTILATOR READY TO REMOVE IT FROM THE MAGNET ROOM, ZONE IV, THE UNIT WAS PULLED INTO THE BORE. NO ONE WAS INJURED. VERIFIED THAT THE MAGNET SERVO-I VENTILATOR IS NOT MRI COMPATIBLE BUT IS MRI CONDITIONAL UP TO THE 200 GAUSS LINE. | 0
4806216 | THE PT PHONED HOMEDICS TO REPORT HER UNIT (BPA-100E) IS READING TOO HIGH. SHE TOOK HER BPA-100E TO HER DOCTOR'S OFFICE WHERE IT READ HIGHER THAN HER DOCTOR'S BPM DURING A SIDE BY SIDE COMPARISON. THE DATE OF THE VISIT WAS NOT SPECIFIED, AND IT IS UNCLEAR IF THIS DISCREPANCY IS RELATED TO THE SYSTOLIC READING, THE DIASTOLIC READINGS OR BOTH. REF: MFR #3005442893-2015-00002. | 0
4596342 | AFTER THE REPORTED INCIDENT, THE CPU5 BOARD WAS REPLACED. THOUGH THE USED CPU5 BOARD WAS REQUESTED FOR INVESTIGATION, IT WAS SCRAPPED. THE CPU5 BOARD WOULD CONTAIN THE RELEVANT AM LOGS THUS IT IS NOT POSSIBLE TO CONFIRM IF THE S/5 AM ALARMED FOR LOW SPO2 OR NOT. ROOT CAUSE IS UNDETERMINED SINCE THE REPLACED CPU5 BOARD WAS NOT RETURNED.><THE CUSTOMER REPORTED THAT THE CLINICIAN OBSERVED THE S/5 ANESTHESIA MONITOR DISPLAY WAS FROZEN. THE PATIENT WAS ON A VENTILATOR DURING THE TIME OF THE EVENT. THE CUSTOMER ALSO STATED THE PATIENT DID NOT SUFFER ANY NEGATIVE CONSEQUENCES DUE TO THE REPORTED ISSUE.><GE HEALTHCARE'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. THE MANUFACTURE DATE IS UNKNOWN AT THIS TIME. | 1
4366575 | ><THE FAULTY E-NMT-01 MODULE WAS REPLACED IN DECEMBER 2014 WITH A NEW E-NMT-01 MODULE. THE FAULTY E-NMT-01 MODULE WAS RECEIVED WITH THE DESCRIPTION OF THE MODULE SHOWS INCORRECT VALUES". THE FAULTY E-NMT-01 MODULE IS ADDRESSED WITH RECALL NUMBER Z-1640-2014. THE REPORTED ISSUE OF E-NMT-01 MODULE SHOWING INCORRECT VALUES WITH ELECTROSENSOR IS A KNOWN ISSUE." | 0
5214766 | IT WAS REPORTED THAT A MISIDENTIFICATION OF ACINETOBACTER AS SHIGELLA ON NEG COMBO 68 (NC68) PANEL OCCURRED ONE TO TWO MONTHS PRIOR TO REPORTING THE INCIDENT TO BECKMAN COULTER. NO ADDITIONAL DETAILS WERE PROVIDED (I.E. TEST DATE, DATA, BIOTYPE INFORMATION). THE SPECIMEN WAS NOT PROCESSED IN THE WALKAWAY AND THE PANEL WAS MANUALLY ENTERED INTO LABPRO. THE ORGANISM WAS A 24 HOUR ID HOLD AND THE TECHNICIAN EDITED THE SUGAR WELLS (DETAILS NOT PROVIDED), IDENTIFICATION WAS DETERMINED TO BE SHIGELLA WITH THE MANUAL DATA ENTRY. THE SHIGELLA IDENTIFICATION WAS QUESTIONED AS THE MORPHOLOGY OF THE ISOLATE ON THE PLATE WAS NOT CONSISTENT WITH A SHIGELLA SPECIMEN. THE CUSTOMER EDITED THE ID TO ACINETOBACTER BASED ON THE VISUAL APPEARANCE OF THE ISOLATE. THE INCORRECT IDENTIFICATION OF SHIGELLA WAS NOT REPORTED OUT OF THE LAB. THERE WAS NO DEATH, INJURY, OR IMPACT TO PATIENT TREATMENT ATTRIBUTED TO OR CONNECTED WITH THIS EVENT.><A LABPRO SOFTWARE INCORRECTLY ALLOWED THE OPERATOR TO MANUALLY ENTER THE FERMENTER BIOCHEMICAL (SPECIFICALLY GLU, SUC, AND SOR) WHEN MANUALLY READING THE NEG COMBO 68 PANEL AT THE ID HOLD STATUS. THIS RESULTED IN THE ORGANISM INCORRECTLY IDENTIFIED AS A SHIGELLA SPECIES. FIELD ACTION (B)(4) WAS INITIATED TO MITIGATE THE LABPRO SOFTWARE ISSUE. USER ERROR FURTHER CONTRIBUTED TO THIS MISIDENTIFICATION BECAUSE FERMENTER BIOCHEMICALS WERE INPUT AT THE 42-48 HOUR READ. ACCORDING TO THE IFU PANELS HELD FOR THE 42-48 HOUR READ SHOULD ONLY USE THE NON-FERMENTER BIOCHEMICAL REACTIONS IN THEIR BIOTYPE CALCULATION. (B)(4). | 0
4516027 | VESSEL SEALER MALFUNCTION - REGISTERED ON MACHINE WHEN PLUGGED IN BUT WOULD NOT GO COMPLETELY IN THE PORT FOR USE. IT KEPT BEEPING AND THE LIGHT TURNED RED; NOT REPROCESSED. NEW VESSEL SEALER OPENED AND USED. NO HARM TO PATIENT. | 0
4531476 | THE NON-INVASIVE BLOOD PRESSURE (NBP) - DIASTOLIC AND NBP - MEAN HAVE INTERMITTENTLY SWITCHED DATA POINTS (TRANSPOSED) ON THE VITAL SIGNS TAB AND PRINT OUT GRAPH ON THE GE CIC (CENTRAL STATION). NO ISSUES HAVE BEEN IDENTIFIED WITH THE NBP VALUES AT THE BEDSIDE MONITORS THEMSELVES. THIS SITUATION HAS BEEN OBSERVED 3 TIMES AT (B)(6), (B)(6) 2013, (B)(6) 2014, AND (B)(6) 2015. GE HAS PROVIDED NO PERMANENT FIX FOR THIS ISSUE TO DATE. | 1
4489915 | INFORMED BY RELIEF RN THAT THE ROBOTIC BIPOLAR MARYLAND STOPPED FUNCTIONING WHILE IN USE IN A ROBOTIC TOTAL LAPAROSCOPIC HYSTERECTOMY (TLH). THE INSTRUMENT (REF# 420172) WAS SUBSEQUENTLY REMOVED FROM THE FIELD AND REPLACED. AT 1340 THE PROGRASP INSTRUMENT THAT WAS BEING USED IN THE SAME PROCEDURE MALFUNCTIONED AND WOULD ONLY PARTIALLY OPEN. THE LATTER INSTRUMENT (REF# 420093) WAS ALSO REMOVED FROM THE FIELD AND REPLACED. BOTH INSTRUMENTS WERE SENT TO CPD FOR DECONTAMINATION AND SEQUESTERED AT OR FRONT DESK TO BE FORWARDED TO RISK. THERE WAS NO INJURY TO THE PATIENT NOTED. | 0
4703976 | DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED.><IT WAS REPORTED BY THE CUSTOMER THAT THE ASYSTOLE ALARM OCCURS ONLY ACOUSTICALLY ON THE CENTRAL STATION AND ON THE DECT PHONE. OPTICALLY NOTHING IS SHOWN ON THE CENTRAL STATION. THERE CAN BE NO ASSIGNMENT OF THE ALARM SOURCE (PATIENT ASSIGNMENT). IF YOU KNOW WHICH PATIENT INITIATED THE ALARM, YOU CAN SEE THE ALARM IN THE EXTENDED PATIENT DISPLAY. ACCORDING TO INFO OF THE CUSTOMER IT OCCURRED ON (B)(6) 2014 ABOUT 16:00 -17:00 O'CLOCK ON BED 11. THERE WAS NO PATIENT INJURY REPORTED. (B)(4). | 1
4703992 | IT WAS REPORTED THAT AN M300 SPONTANEOUSLY SELF-DISCHARGED WHILE BEING USED ON A PATIENT. IT WAS CHARGED SUFFICIENTLY. THE MONITOR THEN STARTED UP AGAIN WITHOUT INTERVENTION FROM THE STAFF MEMBERS WHO WERE WITH THE PATIENT. THIS MONITOR HAD RECENTLY HAD A NEW BATTERY FITTED. THERE WERE ANECDOTAL REPORTS OF THIS HAPPENING PREVIOUSLY, BUT THE M300 WAS CHECKED OUT AND FOUND NOT TO BE FAULTY. THERE WAS NO PATIENT INJURY REPORTED. (B)(4).><DRAEGER IS STILL INVESTIGATING THE REPORTED INCIDENT. A F/U REPORT WILL BE SUBMITTED AS SOON AS THE INVESTIGATION HAS BEEN COMPLETED. | 1
5262297 | PLEASE SEE INITIAL REPORT.><THE INVESTIGATION WAS STARTED BUT HAS NOT YET BEEN CONCLUDED. THE INVESTIGATION RESULT WILL BE REPORTED IN THE FOLLOW UP-REPORT.><THE AFFECTED M540 AND LOG FILES FROM THE DAY OF THE EVENT WERE PROVIDED TO DRÄGER FOR THE INVESTIGATION. THERE WAS NO INDICATION OF A HARDWARE FAILURE IN THE LOG FILES. LOG ENTRIES SHOW THAT THE SPEAKER VOLUME WAS HELD AT 50% FOR THE MAJORITY OF THE DAY EXCEPT AT 11:29, WHEN THEY WERE INCREASED TO 100%. ABOUT 40 SECONDS PRIOR TO INCREASING THE VOLUME TO 100%, THE USER INITIATED AN AUDIO PAUSE EVENT. AUDIO PAUSE SUPPRESSES ALL AUDIBLE ALARMS FOR TWO MINUTES OR UNTIL THE AUDIO PAUSE BUTTON IS PRESSED AGAIN. ANY CHANGE OF SPEAKER VOLUME DOES NOT AFFECT THE AUDIO PAUSE TIME. THERE WERE SEVERAL AUDIO PAUSE EVENTS THROUGHOUT THE REPORTED DATE OF THE EVENT. IT CANNOT BE DETERMINED IF THIS IS THE CAUSE OF THE LACK OF AUDIBLE ALARM AS THE TIME OF THE EVENT WAS NOT REPORTED. A FULL SYSTEM TEST WAS PERFORMED TO CHECK ALL FUNCTIONALITIES OF THE AFFECTED M540 INCLUDING AUDIO TESTS. IT PASSED ALL TESTS AND THERE WERE NO ISSUE WITH THE SPEAKER. REPORTEDLY THE M540 WORKED CORRECTLY AFTER BEING RESTARTED BY THE USER. NEITHER THE LOG FILES NOR THE AFFECTED M540 INDICATED A DEVICE MALFUNCTION FOR THE REPORTED DATE OF EVENT. THE REPORTED PROBLEM COULD NOT BE REPRODUCED. DUE TO THIS A PRECISE ROOT CAUSE COULD NOT BE DETERMINED. THE M540 INCLUDES A SPEAKER AUDIO-TEST DURING STARTUP FOR USER CONFIRMATION OF PROPER OPERATION OF THE DEVICE AND THE IFU FURTHER INSTRUCTS THE USER ON HOW TO VERIFY AUDIO TONES. VISUAL ALARMS WILL STILL CORRECTLY BE DISPLAYED ON THE M540 TO INFORM THE USER. IN A NETWORKED ENVIRONMENT THE CENTRAL STATION WILL ADDITIONALLY ALARM VISUALLY AND AUDIBLY. RISK ANALYSIS WAS REVIEWED AND IT WAS DETERMINED THAT THERE ARE NO NEW OR REVISED RISKS.><IT WAS REPORTED THAT WHEN AN ALARM OCCURS THERE WILL BE NO ALARM SOUND AT THE M540. EVEN IF THE ALARM IS SET AT 100% VOLUME, THERE IS NO ACOUSTIC WARNING. THE CORRECT MESSAGE AND THE ALARM ARRIVE AT THE CENTRAL STATION AND IT ALARMS VISUALLY AND ACOUSTICALLY. THE ERROR COULD BE CORRECTED BY RESTARTING THE M540. | 1
4465959 | PATIENT BEING MONITORED AT THE CENTRAL MONITORING UNIT ON A SPACELAB REMOTE TELEMETRY HAD A 4 SECOND RUN OF VENTRICULAR TACHYCARDIA THAT ACTIVATED A SYSTEM ALARM. THE MONITOR TECH WAS DISTRACTED, DEALING WITH ANOTHER TELEMETRY ISSUE, AND DID NOT SEE THE ALARM NOTIFICATION AT THE TIME OF OCCURRENCE. THE MONITOR TECH LATER REVIEWED THE ALARM LOG AND IDENTIFIED THE ISSUE AND REPORTED THE RHYTHM CHANGE TO THE PATIENT'S NURSE. SEVERAL HOURS LATER, THE PATIENT WENT INTO THE SAME ARRHYTHMIA AND SUBSEQUENTLY DETERIORATED TO VENTRICULAR FIBRILLATION/ASYSTOLE. SPACELAB DOES NOT MAINTAIN A VISUAL SECTOR ALARM THAT REMAINS RESIDENT WITH EPISODIC RHYTHM EVENTS EVEN WHEN NOT ACKNOWLEDGED BY THE USER. DUE TO THIS SOFTWARE ALERT FLAW MONITOR TECHS MUST ROUTINELY CHECK THE SPACELAB ALARM LOG TO IDENTIFY ANY MISSED ALARM ACTIVITY. THIS FURTHER PREVENTS AND DISTRACTS MONITORING PERSONNEL FROM CONTINUOUSLY MONITORING PATIENT PHYSIOLOGIC DATA. THE ALARM STATUS THAT DISPLAYS, REMAINS AUDIBLE AND VISUAL FOR A PERIOD OF APPROX 30 SECONDS THEN DISAPPEARS FROM THE MONITOR SCREEN. THIS PATIENT SAFETY RISK DELAYS OR PREVENTS REAL TIME INTERVENTIONS FOR PATIENT CARE THAT COULD ULTIMATELY LEAD TO LIFE THREATENING OUTCOMES. | 1
4945121 | ACUDOSE MACHINE CONTINUES TO MALFUNCTION. NARCOTIC POCKETS ARE STILL AVAILABLE (GREEN LIGHT) TO BE OPENED WITHOUT ASKING FOR THE MEDICATION. PHARMACY PERSONNEL ARE AWARE AND DRAWER IS TO BE REPLACED. NO ACCESS TO DRAWER ONE FOR THE STAFF TO ACCESS MEDICATION, THUS THE NURSES ARE REQUIRED TO WALK TO A DIFFERENT UNIT FOR MEDICATION INCREASING THE TIME INTERVAL BETWEEN THE TIME PATIENT CALLS FOR MEDICATION AND ADMINISTRATION IS DONE. | 1
5027258 | .><DURING H&N (HEAD AND NECK) ROBOTIC PROCEDURE FOR A GLOSSECTOMY, MONOPOLAR CAUTERY WAS USED INTRAORALLY. ON THE VIDEO, IT WAS NOTED THAT THE SPATULA TIP WAS COMING OFF THE 5MM INTUITIVE S TORS MONOPOLAR CAUTERY INSTRUMENT. REMOVED THE INSTRUMENT FROM THE ORAL CAVITY AND UPON INSPECTION, THE SPATULA TIP HAD MELTED OFF WHERE IT CONNECTED TO THE INSTRUMENT. THE TIP OF THE INSTRUMENT HAD ALSO BURNED OFF. NO HARM TO PATIENT OR STAFF. EQUIPMENT REPLACED AND PROCEDURE COMPLETED. | 0
4580113 | VESSEL SEALER BLADE WAS EXPOSED. IT HAD TO BE REMOVED AND ANOTHER ONE OPENED. NO HARM TO PATIENT. | 0
4877596 | WHILE USING A MEGA NEEDLE DRIVER DURING A ROBOTIC PROCEDURE, THE DRIVER BROKE AND A WIRE CAME DISCONNECTED FROM THE WRIST OF THE DRIVER. THE NEEDLE DRIVER WAS REMOVED FROM THE PATIENT SAFELY AND WITHOUT ANY HARM TO THE PATIENT. AT THE CONCLUSION OF THE CASE IT WAS NOTED THAT THE ARM DID HAVE REMAINING LIVES, BUT IT IS UNKNOWN HOW MANY REMAIN.======================MANUFACTURER RESPONSE FOR NEEDLE DRIVER, MEGA NEEDLE DRIVER (PER SITE REPORTER).======================WE GET FEEDBACK REGARDING THE DA VINCI ROBOTIC ARMS ON A QUARTERLY BASIS. WE WILL RECEIVE CREDIT FOR THOSE THAT HAD REMAINING LIVES AND WERE NOT MISUSED OR MISHANDLED. | 0
4814490 | THE SURGEON WAS USING A MEGA SUTURECUT NEEDLE DRIVER AND NOTICED A WIRE WAS EXPOSED ON THE TIP OF INSTRUMENT FROM THE JOINT. REMOVED INSTRUMENT AND GOT A NEW ONE. NO HARM TO THE PATIENT. NOT REPROCESSED. | 0
5238605 | .><THE PROPRIETARY CABLE THAT CONNECTS THE FUJI DIGITAL PLATE (MODEL DR-ID 601SE) TO THE FUJI WORKSTATIONS BREAKS EASILY. THIS IS THE FOURTH TIME THE CABLE BROKE DURING NORMAL ROUTINE USE. THERE IS A 7-PIN CONNECTOR THAT HAS A SMALL, BLACK PLASTIC CASE. THE BLACK PLASTIC AROUND THE PINS BREAKS EASILY. MANUFACTURER RESPONSE FOR FUJI DIGITAL RADIOGRAPHY PLATE, DR-ID 601SE (PER SITE REPORTER): REPLACEMENT OF THE CABLE DURING TECHNICAL SERVICE. STATED OTHER FACILITIES ARE ABLE TO USE PRODUCT WITHOUT BREAKAGE. | 1
5004013 | THE ROBOTIC GRASPER INSTRUMENT CABLE BROKE DURING USE.><. | 0
4607819 | ON (B)(4) 2015 PER THE USER FACILITY'S BIOMEDICAL ENGINEER (BIOMED), THE UNIT WORKS AND HE BELIEVES HE HAD AIR IN THE LINE. ON (B)(4) 2015 PER THE BIOMED, HE CORRECTED PROBLEM BY PRIMING THE LINES. IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS AND/OR CONCLUSION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.><IT WAS REPORTED THAT THE HEATER COOLER UNIT STOPPED PUMPING ON THE RIGHT CHANNEL. NO OTHER DETAILS REGARDING THE NATURE OF THIS EVENT WERE PROVIDED. | 0
4774504 | (B)(4). THE REPORTED COMPLAINT WAS CONFIRMED. THE FIELD SERVICE REPRESENTATIVE (FSR) REPLACED CHANNEL B MIXING VALVE AND VERIFIED SAFETY AND PERFORMANCE ACCORDING TO THE NOTICE OF FIELD CORRECTION (NFC). THE UNIT OPERATED TO MANUFACTURER SPECIFICATIONS AND WAS RETURNED TO CLINICAL USE. IF ADDITIONAL INFORMATION BECOMES AVAILABLE ON THIS COMPLAINT THAT WOULD ALTER THE FACTS AND/OR CONCLUSION, A SUPPLEMENTAL REPORT WILL BE FILED ACCORDINGLY.><IT WAS REPORTED THAT DURING PRE-CARDIOPULMONARY BYPASS PROCEDURE, THAT CHANNEL B WAS NOT WORKING ON THE HEATER COOLER UNIT. AS A RESULT, AN ALTERNATE DEVICE WAS EMPLOYED. THE SURGICAL PROCEDURE WAS COMPLETED SUCCESSFULLY. THERE WAS NO DELAY, NO BLOOD LOSS, NOR ADVERSE CONSEQUENCES TO THE PATIENT. | 0
4426875 | IT WAS REPORTED THAT DURING SET-UP OF THE DEVICE FOR A CARDIOPULMONARY BYPASS PROCEDURE, THERE WAS AN E0D" ERROR CODE HEATER COLLER UNIT. AS A RESULT | 0
4463990 | (B)(4). THE USER FACILITY REPORTED THAT A PATIENT WAS BURNED WHILE CONNECTED TO THE DEVICE. THE USER FACILITY CLAIMED THAT THE PATIENT SUSTAINED BURN MARKS AT THE SPO2 APPLICATION AREA AND THAT A REUSABLE SPO2 CLIP WAS BEING USED. THE SPO2 CLIP WAS APPLIED OVER THE PATIENT'S PINKY TOE AND WAS IN PLACE FOR ROUGHLY 1 HOUR AND 15 MINUTES DURING MRI SCANNING. THE USER FACILITY REPORTED THAT, WHEN THE SPO2 CLIP WAS REMOVED FROM THE PATIENT'S FOOT, A BLISTER APPEARED. THE BLISTER WAS LOCATED ON THE BOTTOM OF THE PATIENT'S FOOT, UNDERNEATH THE PINKY TOE AND WAS DESCRIBED BY THE USER FACILITY AS 1/2 OF A DIME IN SIZE. THE USER FACILITY INDICATED THAT THE PATIENT RECEIVED FIRST AID CARE, WHICH CONSISTED OF THE APPLICATION OF A TRIPLE ANTIBIOTIC OINTMENT, AND THAT NO FOLLOW-UP TREATMENT WAS EXPECTED. THE DEVICE MANUFACTURER DOES NOT CONSIDER THE APPLICATION OF A TRIPLE ANTIBIOTIC OINTMENT TO BE EMERGENT CARE AND, AS SUCH, IS NOT CONSIDERING THIS REPORT TO BE A SERIOUS INJURY. THE USER FACILITY ALSO REPORTED THAT, PRIOR TO MRI IMAGING, THE SPO2 CLIP WAS ON THE PATIENT'S OTHER FOOT FOR A SHORT PERIOD OF TIME. UPON REMOVAL OF THE SPO2 CLIP FROM THE PATIENT'S OTHER FOOT, A RED AREA WAS NOTED ON THE BOTTOM OF THE PATIENT'S FOOT AS WELL (NO BLISTER). SINCE THE RED AREA WAS NOTED ON THE PATIENT'S OTHER FOOT AFTER REMOVAL OF THE SPO2 CLIP AND NO MRI SCANNING HAD TAKEN PLACE, IT IS NOT BELIEVED AT THIS TIME THAT THE RESPONSE THE PATIENT DISPLAYED WAS THE RESULT OF A BURN. THE USER FACILITY HAS ASKED THE DEVICE MANUFACTURER TO EVALUATE THE WIRELESS SPO2 MODULE AND SPO2 CLIP. SINCE THIS EVALUATION IS CURRENTLY ON-GOING, WE ARE REPORTING THIS INCIDENT IN ORDER TO MEET INITIAL REPORTING DEADLINES.><THE USER FACILITY REPORTED THAT A PATIENT WAS BURNED WHILE CONNECTED TO THE DEVICE. | 0
5239268 | (B)(4). A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.><THE CUSTOMER INITIALLY CALLED IN AND WANTED TO KNOW IF DATA HISTORY CAN BE VIEWED. SUBSEQUENT TO THE INITIAL CALL, ADDITIONAL INFORMATION WAS RECEIVED FROM THE CUSTOMER OUTLINING AN INCIDENT WITH A PATIENT THAT WAS FOUND PULSELESS AND NONRESPONSIVE. INITIALLY, THERE WAS CONCERN THAT THE DEVICE'S ALARM DID NOT SOUND. A CODE WAS CALLED AND THE PATIENT WAS RESUSCITATED AND TRANSFERRED TO THE ICU. IN THE ICU, THE PATIENT SUBSEQUENTLY CODED AGAIN AND ULTIMATELY EXPIRED. THE CAUSE OF THE PATIENT'S DEATH IS NOT KNOWN. THE DEVICE WAS BEING USED FOR PATIENT MONITORING DURING THE INCIDENT.><TESTING OF THE DEVICE WAS PERFORMED BY HILL-ROM. WHEN THE DEVICE WAS RECEIVED FROM HOSPITAL QUARANTINE, HILL-ROM FOUND THAT ECG AND RESPIRATION MONITORING ONLY WERE CONFIGURED ON THE UNIT. WHEN TESTED, HILL-ROM FOUND THAT THE ALARM SOUNDED FOR ECG WHEN THE READINGS WENT ABOVE AND BELOW THE SET LIMITS. WHEN RECEIVED FROM HOSPITAL QUARANTINE, HILL-ROM FOUND THAT THE RESPIRATION ALARM LIMIT FOR THE LOW ALARM WAS TURNED OFF, THEREFORE, NO LOW LEVEL ALARM ACTIVATED WHEN TESTED. ACCORDING TO (B)(6), MSN, RN (MANAGER, RISK & REGULATORY COMPLIANCE AT THE HOSPITAL), ¿THE OUTCOME OF OUR INTERNAL INVESTIGATION WAS THAT OUR CLINICAL ENGINEERING DEPT. AS WELL AS HILL-ROM BIO MED TECHNICIAN COULD FIND NO DEFECTS OR PROBLEMS WITH THE MONITOR IN QUESTION. NO ACTIONS WERE TAKEN IN REGARD TO THE MONITOR AS NO ISSUES WERE FOUND WITH IT.¿ DURING THE INITIAL PHONE CALL, THE CUSTOMER WAS PROVIDED WITH THE FOLLOWING INFORMATION FROM THE DEVICE¿S IFU: GRAPHICAL TRENDS: GRAPHICAL TREND ALLOWS VIEW OF COMPILATION OF DATA COLLECTED FOR GIVEN PARAMETER OVER DEFINED TIME PERIOD. TIME PERIOD CHOICES ARE 1, 8, AND 24 HOURS FOR EACH MONITORED PARAMETER. TECHNICAL SUPPORT ALSO EMAILED THE CUSTOMER THE IFU. THE DEVICE REMAINS AT THE CUSTOMER SITE. A SEARCH IN ONEEMS FOUND NO FURTHER RELATED CALLS. THE CUSTOMER REPORTED THAT A PATIENT THAT WAS FOUND PULSELESS AND NONRESPONSIVE. INITIALLY, THERE WAS CONCERN THAT THE DEVICE'S ALARM DID NOT SOUND. A CODE WAS CALLED AND THE PATIENT WAS RESUSCITATED AND TRANSFERRED TO THE ICU. IN THE ICU, THE PATIENT SUBSEQUENTLY CODED AGAIN AND ULTIMATELY EXPIRED. THE CAUSE OF THE PATIENT'S DEATH IS NOT KNOWN. HILL-ROM¿S EVALUATION OF THE DEVICE DID NOT CONFIRM A PRODUCT MALFUNCTION. THE HOSPITAL¿S RISK MANAGER REPORTED THAT THE OUTCOME OF THEIR INTERNAL INVESTIGATION WAS THAT NO DEFECTS OR PROBLEMS WITH THE MONITOR IN QUESTION COULD BE FOUND. | 1
4412400 | THERE WERE NO OPERATOR OR PT INJURY ASSOCIATED WITH THIS INCIDENT, BUT AN OBSERVATION DESCRIBED BY A CUSTOMER/DENTIST DURING ROUTINE USE OF X-RAY IMAGES. IMAGES CANNOT BE LOST BECAUSE IMAGES ARE BACK-UP BY THE TWAIN INTERFACE. IMAGES CAN BE EASILY RETRIEVED AND ASSIGNED TO THE PROPER PT BY DENTIST BASED ON THEIR CLINICAL EXPERIENCE AND SPECIFIC KNOWLEDGE OF THE PT'S DENTAL ANATOMY.><OBSERVED THE POTENTIAL OF THE UTILITY SOFTWARE DRIVER TO PRESENT X-RAY AND PLACE X-RAY IMAGES IN A PC FOLDER THAT COULD THEN BE CAPTURED BY THE DENTIST AND MATCHED TO A NEW OR WRONG PT. DENTIST WOULD DETECT AND MATCH THE PROPER IMAGE TO A PT BASED THEIR CLINICAL EXAMINATION AND COMPLETE CLINICAL ASSESSMENT BEFORE USING PANORAMIC X-RAY IMAGES. | 1
5201312 | IT IS COMMON PROBLEM THAT TESTS ORDERED WITH CPOE INSTRUMENT AS PART OF THE EHR ARE NOT DONE IN A TIMELY MANNER, IF AT ALL. THIS IS DUE TO A FLAW IN THE CPOE SYSTEMS IN THAT IT FAILS TO ADEQUATELY NOTIFY THE RECIPIENT ANCILLARY SERVICE THAT A TEST HAS BEEN ORDERED. IN THIS CASE, THE PATIENT WITH ATRIAL FIBRILLATION WITH MYOCARDIAL INFARCTION WAS ORDERED. IN THIS CASE, THE PATIENT WITH ATRIAL FIBRILLATION WITH MYOCARDIAL INFARCTION WAS ORDERED AN EKG TO ASSESS ACUTE MYOCARDIAL ISCHEMIA AND IT WAS NOT DONE FOR 14 HOURS AFTER IT WAS TO HAVE BEEN DONE, DELAYING DIAGNOSIS AND SPECIFIC TREATMENTS IN A CRITICALLY ILL PATIENT. | 1
4846860 | EP SYSTEM WOULD COMPLETELY SHUT DOWN WITHOUT WARNING. SYSTEM REBOOTED AND TURNED BACK ON 15 MINUTES LATER, THUS CAUSING CARE OF PATIENT BEING DELAYED WHILE WAITING FOR EQUIPMENT TO RESPOND. | 1
4656193 | DURING DAVINCI TLH, DOCTOR WAS USING THE VESSEL SEALER. MIDWAY THROUGH, THE VESSEL SEALER MACHINE ALARMED AND AN ERROR MESSAGE APPEARED ON THE VISION TOWER SCREEN READING BLADE EXPOSURE ERROR". INSTRUMENT WAS REMOVED FROM FIELD AND NEW VESSEL SEALER WAS OPENED. NO HARM TO PATIENT." | 0
4713835 | REPORTER WANTED TO PROVIDE US WITH INFORMATION REGARDING A PROBLEM THAT'S SPIRALING OUT OF CONTROL. HE IS VERY CONCERNED AND FRUSTRATED BY THE LACK OF STANDARDIZATION AND LACK OF PROPER LANGUAGE / COMMUNICATION BETWEEN COMPUTER SYSTEMS. HE SAID IN HIS STORE, AN INDEPENDENT PHARMACY (THAT USES MCKESSON), E-PRESCRIBED ORDERS ARE NOT MATCHING DUE TO A RECENT CHANGE FROM A LARGE HEALTHCARE ORGANIZATION. HE NOTED THAT THERE IS CURRENTLY AN ADDITION OF A SPACE BETWEEN A DOSE EG 50MG IS NOT 50 MG WHICH IS CAUSING THE ORDERS NOT TO MATCH UP. PHARMACY STAFF IS NOW FORCED TO PICK FROM A LIST WHICH MIGHT HAVE UP TO ITEM DIFFERENT OPTIONS. AS YOU CAN IMAGINE, THEY NOW HAVE MANY ERRORS THAT HAVE RESULTED DUE TO PICKING" WRONG. HE SAID DUE TO MEDICARE EXPECTING STORES TO REPORT EVERY ERROR ON A MEDICARE FUNDED PATIENT | 1
4920879 | DURING THE PROCEDURE THE BLADE BROKE ON THE VESSEL SEALER. IT WAS REPLACED AND THE PROCEDURE CONTINUED. | 0
4713843 | THIS IS ANOTHER EXAMPLE OF BAD DATA FORM FDB AND MCKESSON PARAGON. IT'S THE SUGGESTED DOSES OF MAGNESIUM CHLORIDE ORAL TABLETS. THIS OCCURS WHEN TRYING TO UPDATE A PATIENT'S HOME MED LIST. I HAVE FORWARDED THIS INFORMATION TO FDS AND MCKESSON PARAGON AS WELL. (B)(4). | 1
4778585 | THE SURGEON PERFORMED A MAKOPLASTY TOTAL HIP ARTHROPLASTY USING THE ROBOTIC ARM INTERACTIVE ORTHOPEDIC SYSTEM (RIO). DURING THE SURGERY, BONE REGISTRATION FAILED DUE TO INCORRECT SEGMENTATION. THE CASE WAS SUCCESSFULLY COMPLETED USING MANUAL INSTRUMENTS AND THERE WAS NO HARM TO THE PATIENT.><REPORTED EVENT: AN EVENT REGARDING A FAILED THA REGISTRATION INVOLVING RIO® THA SOFTWARE APPLICATION CATALOG 205634 WAS REPORTED. METHOD & RESULTS: DEVICE HISTORY REVIEW: NOT PERFORMED AS THE REPORTED DEVICE IS SOFTWARE. RIO SERIAL NUMBER NOT REPORTED. COMPLAINT HISTORY: BASED ON THE DEVICE IDENTIFICATION (PN 205634) THE COMPLAINT DATABASES WERE REVIEWED FROM 2011 TO PRESENT FOR SIMILAR REPORTED EVENTS REGARDING FAILED THA REGISTRATION. THERE WAS ONLY 2 OTHER REPORTED EVENT (ACTIONS 2039 AND 753). CONCLUSION: DEVICE INSPECTION COULD NOT BE PERFORMED, NO SESSION DATA WAS PROVIDED. THREE COMMUNICATION ATTEMPTS WERE MADE. SESSION FILES WERE NOT PROVIDED FOR EVALUATION.><THE SURGEON PERFORMED A MAKOPLASTY TOTAL HIP ARTHROPLASTY USING THE ROBOTIC ARM INTERACTIVE ORTHOPEDIC SYSTEM (RIO). DURING THE SURGERY, BONE REGISTRATION FAILED DUE TO INCORRECT SEGMENTATION. THE CASE WAS SUCCESSFULLY COMPLETED USING MANUAL INSTRUMENTS AND THERE WAS NO HARM TO THE PATIENT.><AS PART OF NORMAL COMPLAINT FOLLOW-UP, AN EVALUATION OF THE EVENT IS BEING COMPLETED AT MAKO SURGICAL. A SUPPLEMENTAL REPORT WILL BE FILED WHEN ADDITIONAL INFORMATION IS OBTAINED. | 0
5013890 | .><CABLE AT TIP OF DA VINCI PROGRASP FORCEPS BROKE WHILE BEING USED BY SURGEON DURING VENTRAL HERNIA REPAIR. NO HARM TO PATIENT AND ANOTHER GRASPER WAS USED TO COMPLETE THE PROCEDURE. | 0
4834844 | ON (B)(6) 2015, A 12 BEAT V TACH EPISODE DID NOT ALARM THE ICS ALARM SYSTEM AS WE WOULD EXPECT. UPON FACILITY INVESTIGATION, THIS FAILURE POINT HAS BEEN IDENTIFIED TEN ADDITIONAL TIMES OVER A TWO MONTH TIME PERIOD. EACH OF THESE EVENTS HAS BEEN REPORTED TO SPACELABS AND TO DATE THEY HAVE NOT PROVIDED OUR FACILITY WITH TIMELY RESPONSES, AN ACTIONABLE IMMEDIATE SOLUTION OR ASSURANCE THAT THESE MISSED EVENTS WILL NOT HAVE A NEGATIVE IMPACT ON OUR PTS. | 1
5273096 | .><BOWEL GRASPER TIP BROKE DURING PROCEDURE. WHEN BOWEL GRASPER WAS REMOVED FROM THE PATIENT AND EXAMINED, SMALL METAL FRAGMENT FOUND LOOSE IN GRASPER. PATIENT HAD X-RAY WHICH SHOWED NO METAL FRAGMENTS. | 0
4508603 | WE HAD A PATIENT ADMITTED INTO THE HOSPITAL THAT TAKES ARANESP (DARBEPOETIN) AT HOME. THE DOCTOR ENTERED THE ORDER TO BE FILLED BY PHARMACY VIA (B)(4). THE ORDER CAME ACROSS AS FOLLOWS: DARBEPOETIN 0.4ML IV QWED (MEDREC) DARBEPOETIN 60MCG/ 0.3ML (ARANESP) 0.4ML IV SYRINGE QWED......" THE ISSUE IS ARANESP HAS DIFFERENT CONCENTRATIONS: 40MCG =0.4ML SYRINGE. 60MCG= 0.3ML SYRINGE THEREFORE | 1
4684071 | PT UNDERWENT CRANIECTOMY AND VENTRICULOSTOMY PLACEMENT ON (B)(6) 2015. ON (B)(6) 2015, PT NOTED TO BE CONFUSED. CT OF THE HEAD DEMONSTRATED INTRAVENTRICULAR AIR. ABX INITIATED. CSF RESULTS CAME BACK NEGATIVE AND PT'S MENTAL STATUS IMPROVED. | 0
4495483 | ON (B)(6) 2014, ROSA DEVICE HAS BEEN USED TO GUIDE THE IMPLANTATION OF 13 DEPTH ELECTRODES FOR A CRANIAL SEEG. ON (B)(6) 2014, USER IN NEUROLOGY DEPARTMENT SENT US A LETTER ABOUT AN INCIDENT OCCURRED ON (B)(6) 2014 FOLLOWING THE SURGERY. THE INCIDENT WAS A SUBDURAL BLEEDING ON A (B)(6) MALE PT.><A MEDTECH DEVICE ENGINEER PERFORMED QC/PERFORMANCE VERIFICATION TESTING WITH THE ROSA ROBOTIC SYSTEM AT THE HOSPITAL TO FURTHER INVESTIGATE CUSTOMER COMPLAINT. NO INACCURACY WAS FOUND AND SYSTEM WAS FOUND FULLY FUNCTIONAL. CUSTOMER HEAD FRAME WAS ALSO TESTED AND FOUND. FUNCTIONAL. NO FURTHER ACTION WILL BE TAKEN. INVESTIGATION IS CLOSED. | 0
5251988 | APPROPRIATE ELECTRICAL SAFETY MEASURES (ISOLATION TRANSFORMER) ARE PRESENT IN THE TRUE DEFINITION SCANNER SYSTEM. THE 3M TRUE DEFINITION SCANNER INSTRUCTIONS FOR USE DOES CONTAIN THE WARNING, 'TO REDUCE THE RISKS ASSOCIATED WITH HAZARDOUS VOLTAGE AND FIRE - USE ONLY A PROPERLY GROUNDED POWER OUTLET; DO NOT USE EXTENSION CORDS OR MULTIPLE PORTABLE POWER SOCKET OUTLETS.'><DURING A SCANNING PROCEDURE USING A 3M TRUE DEFINITION SCANNER, A PATIENT EXPERIENCED A SHOCK AND PAIN INSIDE THE MOUTH. IT WAS CONFIRMED THAT THE POWER OUTLET, WHERE THE SCANNER WAS CONNECTED AT THE TIME THE INCIDENT OCCURRED, WAS NOT PROPERLY GROUNDED. AFTER THE SCANNER WAS PLUGGED INTO THE POWER OUTLET WHERE THE GROUND WAS PROVEN TO BE WIRED AND FUNCTIONING PROPERLY, NO MORE SHOCKS OCCURRED. THE DENTIST WAS ADVISED OF THE INVESTIGATION FINDINGS AND AGREED THAT THE INCIDENT WAS CAUSED BY THE FAULTY POWER OUTLET AT THE OFFICE; AN ELECTRICIAN WAS CALLED AND FIXED ALL POWER OUTLETS IN THE OFFICE. | 0
4602930 | DURING ROBOTIC ASSISTED LAPAROSCOPIC DISTAL PANCREATECTOMY, ROBOTIC HARMONIC ENDOWRIST WAS BEING USED FOR CAUTERY. PA-C NOTICED A PIECE OF THE INSULATED INNER JAW OF DISPOSABLE HARMONIC CAME LOOSE, AND SUCTIONED WITH LAPAROSCOPIC SUCTION. INSTRUMENT WAS REMOVED AND PASSED OFF THE FIELD WITH ORIGINAL WRAPPER. FIELD VISUALIZED FOR REMAINING PIECES, WITH NONE FOUND. PA-C AFFIRMED SUCTIONING THE LOOSE PIECE WITH SUCTION TIP. A NEW INSERT WAS PASSED UP AND USED FOR REMAINDER OF CASE. | 0
5001434 | AT APPROX 11:57AM, ALARMS STARTED SOUNDING. THE PT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATING ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL MONITOR. AFTER ABUT 15 MINUTES, THE PROBLEM WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PT MONITOR WAS REMOVED FROM THE NETWORK AND THE MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
5001445 | AT APPROX 11:57AM, ALARMS STARTING SOUNDING. THE PT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATING ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL MONITOR. AFTER ABOUT 15 MINUTES THE PROBLEM WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PT MONITOR WAS REMOVED FROM THE NETWORK AND THE PT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
5011248 | THERE WAS TREATMENT FOR END STAGE HEART FAILURE THAT INCLUDED POTASSIUM SPARING DIURETICS, POTASSIUM SUPPLEMENTS, AND OTHER MEDICATIONS LISTED ON HEART FAILURE THERAPY GUIDELINES. INITIALLY, THERE WAS SYNCOPE WITH VENTRICULAR TACHYCARDIA WHILE THE POTASSIUM WAS LESS THAN NORMAL. IT WAS NOT UNTIL 24 HOURS AFTER A POTASSIUM TEST RESULT WAS COMPLETED THAT WAS 6 MEQ/L ( UP FROM 5 MEQ/L ONE DAY EARLIER) THAT ANYONE NOTICED THIS LIFE THREATENING METABOLIC PROBLEM. THE TREATMENT AND MEDICATION ADJUSTMENTS RESULTED IN A PROLONGED HOSPITALIZATION. THE FAILURE OF THE DEVICE TO WARN OF NEW RESULTS RENDERS THE EHR MDDS NOT UNLIKE A HEART MONITOR, THAT REQUIRES CHRONIC VIGILANCE TO AVOID MISSING A CRITICAL ARRHYTHMIA. THE EHR MDDS GENERALLY HAS ZERO VIGILANCE. | 1
4846867 | WHILE USING, THE DOCTOR NOTICED THE WIRES STARTING TO FRAY ON DEVICE. THE SURGEON STOPPED USING AND REMOVED IMMEDIATELY. THIS IS A MULTIPLE USE ITEM. IT HAD FOUR USES LEFT ON IT. | 0
5128810 | FUJIFILM MEDICAL SYSTEMS U.S.A., INC. (FMSU) RECENTLY COMPLETED A RETROSPECTIVE REVIEW OF COMPLAINT FILES FROM 2013 THROUGH 2015 AND HAS DETERMINED THAT THIS COMPLAINT MEETS THE THRESHOLD FOR REPORTING. FMSU IS THEREFORE REPORTING IT AT THIS TIME. AN EXAM AT (B)(6) HOSPITAL IN (B)(6), WAS ENTERED MANUALLY ON THE MODALITY AND THE ACCESSION NUMBER WAS ENTERED INCORRECTLY BY THE TECHNOLOGIST. AS A RESULT, THE EXAM WAS SENT FROM THE MODALITY TO THE PATIENT EXAM FOLDER IN SYNAPSE MATCHING THE ACCESSION NUMBER THAT WAS ENTERED. THIS PATIENT FOLDER HOWEVER WAS NOT FOR THE CORRECT PATIENT. SYNAPSE HAS FUNCTIONALITY DESIGNED TO MITIGATE THIS TYPE OF EVENT. WHEN STUDIES ARE RECEIVED FROM A MODALITY, SYNAPSE PULLS THE PATIENT INFORMATION FROM THE DICOM HEADER. SYNAPSE THEN TRIES TO MATCH THE PATIENT INFORMATION, E.G., THE PATIENT ID, PATIENT NAME, ETC. WITH THE ACCESSION NUMBER. WHEN SYNAPSE FINDS A POSSIBLE MISMATCH IT CREATES AN ANOMALY ALERT. ANYTIME THE EXAM IS OPENED AN ANOMALY ALERT WILL OPEN ON THE USER'S MONITOR SCREEN INDICATING THAT AN ANOMALY EXISTS FOR THIS EXAM. THE ANOMALY ALERT IDENTIFIED IN THE EVENT ABOVE WOULD INDICATE MISMATCH PATIENT NAME." A SCREENSHOT OF AN ANOMALY ALERT IS PROVIDED FOR YOUR REFERENCE (THIS IS NOT FROM THE PARTICULAR EVENT). THE USER HAS THE ABILITY TO SELECT A CHECKBOX NOT TO SHOW ANOMALIES WHILE IN THEIR CURRENT SESSION WHICH WAS DONE IN THIS INSTANCE. AFTER BEING NOTIFIED OF THE ANOMALY AND TURNING OFF THE WARNING THE RADIOLOGIST STILL WENT AHEAD AND DICTATED THE RESULTS OF PATIENT A'S EXAM TO THE PATIENT B'S EXAM FOLDER; AS A RESULT THERE WAS AN ALLEGED MISDIAGNOSIS OF PATIENT B. FMSU IS NOT AWARE OF WHETHER TREATMENT WAS PROVIDED ON THE BASIS OF THE MISDIAGNOSIS OR OF THE OUTCOME. EXAMS WITH ANOMALIES ARE ALSO COPIED TO AN ANOMALY FOLDER. THE PACS ADMINISTRATOR SHOULD BE REVIEWING THE ANOMALY FOLDER PERIODICALLY TO RESOLVE ANOMALIES. OR | 1
5001402 | AT APPROXIMATELY 11:57AM ALARMS STARTED SOUNDING. THE PATIENT MONITOR IN THE OPERATING ROOM WENT DEAD. TRIED SEVERAL TIMES TO REBOOT MONITOR. PATIENT MONITOR TRIED TO REBOOT THEN WOULD SHUT DOWN AGAIN. A BACKUP MONITOR WAS TAKEN INTO THE OPERATING ROOM. TROUBLESHOOTING ENSUED ON THE ORIGINAL MONITOR. AFTER ABOUT 15 MINUTES THE PROBLEM WAS DETERMINED TO BE A NETWORK RELATED ISSUE. THE PATIENT MONITOR WAS REMOVED FROM THE NETWORK AND THE PATIENT MONITOR BOOTED UP AND RESUMED NORMAL OPERATION. | 1
4438989 | RECEIVED LETTER ABOUT CONDITIONS THAT INSTRUMENT MAY NOT BE USED DURING THESE CONDITIONS. FOLLOWING INSTRUCTIONS OF WARNING LETTER TO REPORT ONLINE. | 0
4949031 | ><PATIENT HAD A MEDIASTINOSCOPY, THORACOTOMY, AND LOBECTOMY THE DAY PRIOR TO THIS OCCURRENCE; ON THE NIGHT OF SURGERY, THE PATIENT WAS HYPOTENSIVE WITH BLOODY DRAINAGE > 1LITER OUT OF CHEST TUBE IN CRITICAL CARE UNIT. PATIENT WAS TAKEN EMERGENTLY TO SURGERY FOR RE-EXPLORATION OF BLEEDING AND 2 UNITS PACKED RED CELLS WERE ORDERED. WHILE IN SURGERY, THE BLEEDING SOURCE WAS IDENTIFIED, HEMOCLIPS APPLIED WITH GOOD HEMOSTASIS. AT THIS POINT IN THE OPERATION, ANESTHESIA REPORTED OXYGENATION ISSUES (DECREASED OXYGEN SATURATION) WHICH RAPIDLY DETERIORATED. ANESTHESIOLOGIST NOTED A SIGNIFICANT LEAK DEVELOPED WITH POSITIVE PRESSURE VENTILATION. TROUBLESHOOTING OF THE CIRCUIT AND VENTILATOR REVEALED NO APPARENT CAUSE. PATIENT REQUIRED RE-INTUBATION, SUCCESSFUL ON SECOND ATTEMPT. PATIENT HAD ASYSTOLE AND CPR WAS PERFORMED SUCCESSFULLY WITH RETURN OF SPONTANEOUS CIRCULATION. PATIENT HAD A PERIOD OF 1-2 MINUTES OF A NON-PERFUSING RHYTHM DURING THE CARDIAC ARREST. PROCEDURE WAS COMPLETED AND PATIENT WAS TAKEN TO ICU IN CRITICAL BUT STABLE CONDITION INTUBATED. THE FOLLOWING DAY THE PATIENT WAS EXTUBATED AND STABLE. PATIENT HAD AN UNEVENTFUL POSTOP COURSE AND WAS DISCHARGED ON POST-OP DAY 6 IN GOOD CONDITION. ADDITIONAL NOTES: ANESTHESIA MACHINE, ALL PATIENT MONITORING EQUIPMENT, AND MONITORING CABLES WERE PLACED IN SEQUESTER AT THE HOSPITAL AFTER THIS OCCURRENCE. THE PATIENT CIRCUIT WAS NOT RECOVERED. CLINICAL ENGINEER (BIOTECH) WAS CALLED TO SURGERY FOR THIS VENTILATOR FAILURE DURING THIS PROCEDURE. ANESTHESIOLOGIST STATED THE MACHINE HAD AN AUTO VENTILATION FAILURE DURING THE PROCEDURE. THE MACHINE WAS PLACED INTO MANUAL VENTILATION MODE AND ALSO FAILED. THE MACHINE DISPLAYED A LEAK ALARM INDICATING A LEAK IN THE SYSTEM. EVENTUALLY THE LEAK ALARM CLEARED AND THE CASE WAS COMPLETED. GE HEALTHCARE WAS NOTIFIED OF THIS EVENT; WE REQUESTED THAT GE HEALTHCARE DO AN ON-SITE WITNESSED DOWNLOAD TO OBTAIN THE LOG RECORDS AND THE PERMANENT HISTORY OF THE ANESTHESIA MACHINE. THE ROOT CAUSE OF THE EQUIPMENT FAILURE HAS NOT BEEN DETERMINED YET. ====================== MANUFACTURER RESPONSE FOR ANESTHESIA MACHINE, AVANCECS2PRO (PER SITE REPORTER)======================GE HEALTHCARE REP CAME ON-SITE AT OUR REQUEST TO DOWNLOAD THE ALARM LOG, ERROR LOG, EVENT LOG, AND KEY PRESS LOG. WE WILL REQUEST THAT GE HEALTHCARE COME ON-SITE TO DO WITNESSED NON-DESTRUCTIVE TESTING OF THE ANESTHESIA MACHINE. WE ARE WORKING WITH THE MANUFACTURER TO DETERMINE THE ROOT CAUSE OF THIS EVENT. | 0
5201306 | REPEATEDLY, ABNORMAL SERUM POTASSIUM LEVELS COME IN TO THE EHR BUT NO ONE KNOWS THEY ARE THERE OR THAT THEY ARE ABNORMAL. THERE IS A DEFICIENCY IN THIS MEDICAL DEVICE IN THAT IT FAILS TO PROVIDE NOTIFICATION THAT NEW RESULTS HAVE APPEARED OR THAT THEY ARE ABNORMAL. CONSEQUENTLY, THE NURSES ADMINISTER MEDICATIONS THAT EXACERBATE THE ABNORMAL LAB VALUE. FOR INSTANCE, POTASSIUM AND/OR POTASSIUM SPARING MEDICATIONS ARE ADMINISTERED WHEN THE POTASSIUM IS ALREADY TOO HIGH, PUTTING THE PATIENT AT EXTREME RISK, AND REQUIRING CORRECTIVE TREATMENTS. | 1
4506998 | DURING ROBOTIC SURGERY USING ELECTROSURGICAL ROBOTIC SCISSOR, A DISPOSABLE TIP WAS ON THE SCISSOR TO PREVENT ANY SPREAD OF THE POWER, BUT THE SCISSOR WAS SMOKING UNDER THE TIP. WE REPLACED THE SCISSOR AND THE TIP WITH NO FURTHER INCIDENT. MOST LIKELY A DEFECTIVE DISPOSABLE SCISSOR TIP WAS THE CAUSE OF THE ELECTROSURGICAL POWER SPREAD. | 0
5059465 | THE (B)(6) STATE PUBLIC HEALTH NEWBORN SCREENING LABORATORY USES THE PERKINELMER SCREENING CENTER MODULE OF SPECIMEN GATE (SGSC) SOFTWARE (A CLASS I MEDICAL DEVICE) TO STORE, RETRIEVE, AND PROCESS THE DATA ASSOCIATED WITH SPECIMEN TESTING INCLUDING BUT NOT LIMITED TO PATIENT DEMOGRAPHICS, TESTS ORDERED, TEST RESULTS, TEST RESULT INTERPRETATION, QUALITY CONTROL RESULTS, AND RESULT CODES (FLAGS) THAT MAY BE ASSOCIATED WITH THE SPECIMEN FROM ITS ENTRY INTO THE LABORATORY WORKFLOW UNTIL PATIENT REPORT IS GENERATED AND RELEASED. WHEN SGSC SOFTWARE IS INSTALLED, CONFIGURABLE SETTINGS ARE PROGRAMMED BASED ON THE CUSTOMER PREFERENCES. FOR INSTANCE, EACH LABORATORY ENTERS THEIR SPECIFIC POPULATION BASED REFERENCE RANGES THAT ARE USED TO EVALUATE IF A PATIENT'S TEST RESULTS ARE NORMAL, BELOW OR ABOVE THE REFERENCE RANGE AND REQUIRE FURTHER ACTION SUCH AS REPEAT TESTING. THIS LABORATORY SPECIFIC LOGIC AUTOMATES THE FLOW OF THE SPECIMEN THROUGH THE LABORATORY. THE CUSTOMER DEFINES THE NOMENCLATURE USED FOR THEIR RESULT CODES (FLAGS), AND WHEN THE SOFTWARE ENCOUNTERS THE PREDEFINED RESULT CODE THE SOFTWARE PROCESSES THE SPECIMEN ACCORDING TO THE LOGIC ASSOCIATED WITH THE RESULT CODE. THE NEWBORN SCREENING LABORATORY RECEIVED 3 SPECIMENS FOR THE PATIENT. BY PROTOCOL, TWO SPECIMENS SHOULD BE SENT FOR ALL BABIES BORN IN (B)(6). THE FIRST SPECIMEN WAS RECEIVED ON 1/02/2015. THE INITIAL MET WAS PANIC HIGH, TRIGGERING A CONFIRMATION REQUEST. THE CONFIRMATION WAS NOT TESTED BY THE LABORATORY, BUT THE CONFIRMATION RECORD WAS ATTACHED TO A SEPARATE ASSAY THAT THE SPECIMEN WAS NOT PUNCHED TO, CAUSING INCOMPLETE RESULTS FOR THE CONFIRMATION HCY (HOMOCYSTINURIA) SO THE LABORATORY WAS UNABLE TO PRINT THE REPORT AND CONTACTED PERKINELMER FOR ASSISTANCE. PERKINELMER CANCELLED AND UNATTACHED THE RECORDS THAT WERE INCORRECTLY ASSOCIATED WITH THE SEPARATE ASSAY. A REVIEW ITEM WAS CREATED FOR THIS SPECIMEN ON (B)(6) 2015 AT 10:45 AM PREVENTING THE REPORT FROM BEING PRINTED. THE LABORATORY CLOSED THE REVIEW ITEM ON 01/08/2015 AT 11:25 AM ALLOWING THE REPORT TO BE PRINTED, DISPLAYING NORMAL FOR THE HCY DISORDER EVEN THOUGH TESTING NEEDED TO BE CONFIRMED. THE SECOND SPECIMEN WAS RECEIVED ON 01/05/2015, ANALYZED AND REPORTED AS NORMAL FOR HCY DISORDER ON 1/07/2015. THE THIRD SPECIMEN WAS RECEIVED ON 01/30/2015, ANALYZED AND REPORTED AS NORMAL FOR HCY DISORDER ON 2/03/2015. THE PATIENT REPORT RELEASED DID NOT CONTAIN ANY ERRONEOUS CONTENT. (B)(4). | 1
5065924 | SIEMENS HEALTHCARE DIAGNOSTICS HAS IDENTIFIED FIVE FEATURE ISSUES WITH THE SYNGO LAB DATA MANAGER. THE ISSUES PERTAIN TO RESULT UNIT CONVERSIONS, QUALITY CONTROL PROCESSING, VIRUS PROTECTION, SYSTEM PERFORMANCE DEGRADATION, AND ORDERS RECEIVED FROM LIS BEING REJECTED. AN URGENT MEDICAL DEVICE CORRECTION (UMDC) ENTITLED SYNGO LAB DATA MANAGER SYSTEM SOFTWARE ISSUES" WAS SENT TO CUSTOMERS IN THE UNITED STATES IN AUGUST 2015. AN URGENT FIELD SAFETY NOTICE (UFSN) ENTITLED "SYNGO LAB DATA MANAGER SYSTEM SOFTWARE ISSUES" WAS SENT TO CUSTOMERS OUTSIDE OF THE UNITED STATES IN AUGUST 2015. THE UMDC AND UFSN DESCRIBE THE ISSUES AND PROVIDE ACTIONS THE CUSTOMER CAN TAKE TO PREVENT AND MITIGATE THE ISSUES.><A SIEMENS CUSTOMER CARE CENTER (CCC) SPECIALIST FOR A CUSTOMER SITE DISCOVERED THAT THE SYNGO LAB DATA MANAGER WOULD BECOME SLOW AND NOT ACCEPT A QUERY OR DOWNLOAD FROM THE LABORATORY INFORMATICS SYSTEM. THERE ARE NO KNOWN REPORTS OF ADVERSE HEALTH CONSEQUENCES DUE TO THIS ISSUE." | 1